## Jul'25
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22503907&pc=OZP | 2032227-2025-212200 | 2025-06-19 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-15 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Heart Failure/Congestive Heart Failure | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER'S NEXT OF KIN INFORMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025, IN THE HOSPITAL. THE CAUSE OF DEATH WAS HEART FAILURE DUE TO DIABETES COMPLICATIONS. THE LAST KNOWN PRODUCTS FOR THIS CUSTOMER WERE AS FOLLOWS: MMT-1884, MMT-7841ZN, UNK_RESERVOIR, UNOMEDICAL, AND UNK_SENSOR. TROUBLESHOOTING WAS PARTIALLY PERFORMED BUT LATER DECLINED. IT IS UNKNOWN WHETHER THE PUMP WAS WORN AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR UNK_SENSOR. NO PRODUCT RETURN IS REQUIRED FOR UNK_RESERVOIR. NO PRODUCT RETURN IS REQUIRED FOR MMT-7841ZN. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22505089&pc=DSP | 2249723-2025-0003034 | 2025-07-01 04:00 | Death | DATASCOPE CORP. | 2025-07-15 04:00 | DSP | CS100 INTRA-AORTIC BALLOON PUMP, ENGLISH, 220V | Incorrect, Inadequate or Imprecise Result or Readings | Insufficient Information | K031636 |  | 1 | Event Description: IT WAS REPORTED ON (B)(6) 2025, THAT CS100 INTRA-AORTIC BALLOON PUMP (IABP) SHOWED ELECTRICAL TEST FAILD CODE #118 DURING USE ON PATIENT. THE PATIENT HAS DIED. CUSTOMER ALSO REPORTED THAT A SYSTEM FAILURE ERROR OCCURRED ON (B)(6) 2025. Manufacturer Narrative: DUE TO CHARACTER LIMITATION IN E1, 1. FULL EVENT SITE NAME IS (B)(6). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22505564&pc=DSQ | 2916596-2025-04605 | 2025-06-16 04:00 | Death | THORATEC CORPORATION | 2025-07-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. ADDITIONAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE PUMP WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV D, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV A, ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A: PATIENT INFORMATION HAS NOT YET BEEN PROVIDED BY THE SITE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22505641&pc=OZD | 1220648-2025-30327 | 2025-06-27 04:00 | Death | ABIOMED, INC. | 2025-07-15 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WITH ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK (CGS) WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT. THE COMPLAINT REPORTED THAT A PATIENT PRESENTED FOR PRIMARY PERCUTANEOUS CORONARY INTERVENTION YESTERDAY AND WAS EXPLANTED IN THE CARDIAC CATHETERIZATION LAB. THE PATIENT WENT TO THE FLOOR AND DECOMPENSATED OVERNIGHT. THE IMPELLA CP WAS IMPLANTED DUE TO LOW CARDIAC OUTPUT/CARDIAC INDEX AND CGS. DURING SUPPORT IT WAS NOTED THAT THE PATIENT HAD MULTIPLE ROUNDS OF VENTRICULAR TACHYCARDIA (VT) REQUIRING SHOCKS. ULTIMATELY, THE FAMILY WITHDREW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22505930&pc=DSQ | 2916596-2025-04509 | 2025-04-02 04:00 | Death | THORATEC CORPORATION | 2025-07-15 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH PATIENT OUTCOME THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: DUE TO PATIENT PRIVACY LAWS NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22505957&pc=PNY | 2024168-2025-07245 | 2013-12-01 05:00 | Death | ABBOTT VASCULAR | 2025-07-15 04:00 | PNY | ABSORB BIORESORBABLE VASCULAR SCAFFOLD SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P150023 |  | 1 | Event Description: THE POTENTIAL BENEFITS OF BIOABSORBABLE STENTS CAN BE BETTER ASSESSED OVER THE LONG TERM. THE IMPLANTATION OF BIORESORBABLE SCAFFOLD (BRS) IN SITUATIONS WITH OFF-LABEL INDICATIONS PROVIDES REAL WORLD INSIGHTS INTO HOW CLINICAL EVENTS DIFFER IN CONTRAST TO STANDARD PROVED INDICATIONS. THE ARTICLE PROVIDES LONG-TERM FOLLOW-UP DATA ABOUT THE USE OF BIORESORBABLE SCAFFOLD (BRS) FOR OFF LABEL COMPARED WITH APPROVED INDICATIONS. BETWEEN NOVEMBER 2013 AND JANUARY 2016, 3264 PATIENTS WERE ENROLLED IN THE GERMAN AUSTRIAN ABSORB REGLSTRY (GABI-R). ADVERSE PATIENT EFFECTS INCLUDED THE FOLLOWING: DEATH, MAJOR ADVERSE CARDIAC EVENT, SCAFFOLD THROMBOSIS, TARGET LESION FAILURE, TARGET VESSEL FAILURE AND RE-HOSPITALIZATION. THE ARTICLE CONCLUDED THAT THE OFF-LABEL USE OF BRS IS ASSOCIATED WITH A HIGHER RATE OF STENT THROMBOSIS IN THE SHORT TERM AND IN THE LONG TERM WITH HIGHER MACE EVENTS CONSIDERING MORE COMPLEX LESIONS AND A HIGHER MORBIDITY. IN THE LONG TERM, THERE ARE NO DIFFERENCES REGARDING STENT THROMBOSIS. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED " ABSORB® BIORESORBABLE SCAFFOLD IN ¿ESTABLISHED¿ VERSUS ¿OFF LABEL¿ CORONARY LESIONS: 5 YEAR DATA FROM THE GABI R® REGISTRY.¿ NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE ABSORB BIORESORBABLE VASCULAR SCAFFOLD SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. LITERATURE ATTACHMENT: ARTICLE TITLE: "ABSORB® BIORESORBABLE SCAFFOLD IN ¿ESTABLISHED¿ VERSUS ¿OFF LABEL¿ CORONARY LESIONS: 5 YEAR DATA FROM THE GABI R® REGISTRY" B2, B3, 3, D6A: ESTIMATED DATE. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT ADVERSE EFFECTS REPORTED IN THE ARTICLE IS CAPTURED UNDER SEPARATE MEDWATCH REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22506387&pc=CAW | 2518422-2025-107438 | 2025-05-12 04:00 | Death | RESPIRONICS, INC. | 2025-07-15 04:00 | CAW | EVERFLO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K061261 |  | 1 | Event Description: THE RELATIVE OF A PATIENT INQUIRED ABOUT RETURNING A PATIENT'S DEVICE REPORTING THE PATIENT PASSED AWAY. THE PATIENT'S CAUSE OF DEATH WAS NOT PROVIDED. ADDITIONAL INFORMATION IS BEING REQUESTED FROM THE REPORTER SURROUNDING THE DETAILS OF THE PATIENT'S DEATH, AND IF THERE WAS ANY KNOWN DEVICE ISSUE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22506534&pc=OZD | 1220648-2025-30339 | 2025-06-27 04:00 | Death | ABIOMED, INC. | 2025-07-15 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT, THE PATIENT EXPERIENCED OOZING FROM THE ACCESS SITE. BLOOD PRODUCTS WERE ADMINISTERED AND THE SITE WAS REDRESSED. FOLLOWING CLINICAL DETERIORATION, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22506688&pc=CBF | MW5172645 | 2025-07-09 04:00 | Death | OXYHEALTH, LLC. | 2025-07-15 04:00 | CBF | FORTIUS 420 HYPERBARIC CHAMBER | Fire | Insufficient Information |  |  | 1 | Event Description: THE DECEDENT WAS THE FOUNDER OF THE (B)(6) WHERE THE INCIDENT OCCURRED AND HAD RECENTLY COMPLETED A SAFETY AND CERTIFICATION COURSE PROVIDED BY THE INTERNATIONAL HYPERBARIC ASSOCIATION, INC. (IHA), A NON-PROFIT ORGANIZATION AFFILIATED WITH OXYHEALTH, AND THE INTERNATIONAL BOARD OF UNDERSEA MEDICINE (IBUM). INITIAL REPORT: A MALE CLINIC FOUNDER DIED IN A FLASH FIRE INSIDE AN OXYHEALTH FORTIUS 420 HYPERBARIC CHAMBER DURING SELF-TREATMENT. THE FIRE'S CAUSE IS STILL UNDER INVESTIGATION. THE INDIVIDUAL HAD RECENTLY COMPLETED SAFETY CERTIFICATION TRAINING THROUGH THE INTERNATIONAL HYPERBARIC ASSOCIATION (IHA) AND THE INTERNATIONAL BOARD OF UNDERSEA MEDICINE (IBUM). UPDATE: I AM ATTACHING PHOTOS OF THE HYPERBARIC CHAMBERS THAT WERE IN USE AT THE CLINIC PRIOR TO THE INCIDENT. THESE IMAGES CLEARLY SHOW THE CHAMBERS BEARING BOTH THE COMPANY NAME, OXYHEALTH, AND THE MODEL DESIGNATION, FORTIUS 420. NOTABLY, THE FDA LOGO APPEARS IN TWO LOCATIONS: ON THE EXTERIOR CONTROL PANEL AND ON THE INTERIOR WALL OF THE CHAMBER NEAR THE INTERNAL CONTROLS. THIS USE OF FDA MARKS DOES NOT APPEAR TO BE AUTHORIZED AND MAY CONSTITUTE A VIOLATION OF FEDERAL LAW.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22507391&pc=MND | 3005099803-2025-03197 | 2025-06-19 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-15 04:00 | MND | SPEEDBAND SUPERVIEW SUPER 7 | Use of Device Problem; Failure to Fire | Cardiac Arrest | 510K EXEMPT |  | 1 | Event Description: NOTE: THIS REPORT PERTAINS TO ONE OF THREE SPEEDBAND SUPERVIEW SUPER 7 DEVICES THAT WERE USED IN THE SAME PATIENT AND PROCEDURE. IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT THREE SPEEDBAND SUPERVIEW SUPER 7 WERE USED DURING AN OESOPHAGOGASTRODUODENOSCOPY (OGD) PROCEDURE FOR OESOPHAGEAL VARICES PERFORMED ON (B)(6) 2025. DURING THE PROCEDURE, IT WAS REPORTED THAT THE BANDING KITS WOULD NOT DEPLOY. ADDITIONALLY, THE SAME PROBLEM HAPPENED ON THE OTHER TWO SPEEDBAND SUPERVIEW SUPER 7 DEVICES. ON (B)(6) 2025, ADDITIONAL INFORMATION WAS RECEIVED THAT THE TRIPWIRE WAS NOT SECURED ON THE SLOT. THE PATIENT EXPERIENCED A CARDIAC ARREST AND PASSED AWAY. HOWEVER, IT WAS NOT CONFIRMED IF THE DEVICES CONTRIBUTED TO THE PATIENT COMPLICATIONS. NOTE: IT WAS REPORTED THAT THE TRIP WIRE WAS UNABLE TO BE SECURED IN THE SLOT ON THE HANDLE UNIT; HOWEVER, PER THE INSTRUCTIONS FOR USE (IFU), SECURE TRIP WIRE IN SLOT ON HANDLE UNIT. NO FURTHER INFORMATION HAS BEEN OBTAINED DESPITE GOOD-FAITH EFFORTS.Event Description: NOTE: THIS REPORT PERTAINS TO ONE OF THREE SPEEDBAND SUPERVIEW SUPER 7 DEVICES THAT WERE USED IN THE SAME PATIENT AND PROCEDURE. PROCEDURE SUMMARY: IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT THREE SPEEDBAND SUPERVIEW SUPER 7 WERE USED DURING AN OESOPHAGOGASTRODUODENOSCOPY (OGD) PROCEDURE FOR OESOPHAGEAL VARICES PERFORMED ON (B)(6) 2025. EVENT SUMMARY: DURING THE PROCEDURE, IT WAS REPORTED THAT THE BANDING KITS WOULD NOT DEPLOY. ADDITIONALLY, THE SAME PROBLEM HAPPENED ON THE OTHER TWO SPEEDBAND SUPERVIEW SUPER 7 DEVICES. ON JUNE 27, 2025, ADDITIONAL INFORMATION WAS RECEIVED THAT THE TRIPWIRE WAS NOT SECURED ON THE SLOT. THE PATIENT EXPERIENCED A CARDIAC ARREST AND PASSED AWAY. HOWEVER, IT WAS NOT CONFIRMED IF THE DEVICES CONTRIBUTED TO THE PATIENT COMPLICATIONS. NOTE: IT WAS REPORTED THAT THE TRIP WIRE WAS UNABLE TO BE SECURED IN THE SLOT ON THE HANDLE UNIT; HOWEVER, PER THE INSTRUCTIONS FOR USE (IFU), SECURE TRIP WIRE IN SLOT ON HANDLE UNIT. ADDITIONAL INFORMATION RECEIVED ON JULY 22, 2025. THE DEVICE WAS USED TO TREAT ACTIVELY BLEEDING OESOPHAGEAL VARICES. ALTHOUGH THE SECOND TWO BANDING DEVICES WERE SUCCESSFUL, THE PATIENT WAS NOT STABLE FOLLOWING THE PROCEDURE AND EXPERIENCED CARDIAC ARREST AND DIED. THE PATIENT DIED ON THE SAME DAY AS THE PROCEDURE. WHILE THE EXACT CAUSE OF DEATH IS NOT CONFIRMED DUE TO THE ABSENCE OF A DEATH CERTIFICATE, IT IS BELIEVED TO BE DUE TO SEVERE LIVER PROBLEMS, COAGULOPATHY, AND VARICES. THE PATIENT ALSO HAD A SIGNIFICANT CO-MORBIDITY, END-STAGE LIVER CIRRHOSIS, WHICH LIKELY CONTRIBUTED TO HER DEATH. Manufacturer Narrative: BLOCK B2: THE EXACT DATE OF DEATH WAS NOT REPORTED. IT WAS ESTIMATED TO BE JUNE 19, 2025, BASED ON EVENT DATE REPORTED. BLOCK H6: IMDRF DEVICE CODE A050501 CAPTURES THE REPORTABLE EVENT OF BANDS UNABLE TO DEPLOY. IMDRF PATIENT CODE E0602 CAPTURES THE REPORTABLE PATIENT COMPLICATION OF CARDIAC ARREST. IMDRF IMPACT CODE F02 CAPTURES THE REPORTED DEATH. Manufacturer Narrative: BLOCK B2: THE EXACT DATE OF DEATH WAS NOT REPORTED. IT WAS ESTIMATED TO BE JUNE 19, 2025, BASED ON EVENT DATE REPORTED. BLOCK H6: IMDRF DEVICE CODE A050501 CAPTURES THE REPORTABLE EVENT OF BANDS UNABLE TO DEPLOY. IMDRF PATIENT CODE E0602 CAPTURES THE REPORTABLE PATIENT COMPLICATION OF CARDIAC ARREST. IMDRF IMPACT CODE F02 CAPTURES THE REPORTED DEATH. BLOCK H2 (ADDITIONAL INFORMATION): BLOCK B5 HAS BEEN UPDATED BASED ON THE ADDITIONAL INFORMATION RECEIVED ON JULY 22, 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22507423&pc=LWS | MW5172659 | 2024-10-24 04:00 | Death | MEDTRONIC, INC. | 2025-07-15 04:00 | LWS | EVERY XT DR | Failure to Deliver Shock/Stimulation; Pacing Problem | Myocardial Infarction |  |  | 1 | Event Description: MY MOTHER WAS HAVING A HEART ATTACK ON (B)(6) 2024. HER PACEMAKER/DEFIBRILLATOR DID NOT PACE OR SHOCK HER HEART AT NO POINT BEFORE HER HEART STARTED HAVING A PROBLEM NOR AT THE TIME. SHE PASSED AWAY! I HAD MEDTRONIC RUN A TEST ON THE DEVICE THAT LAST WEEK UP TO HER DEATH, AND THEY STATED THE TEST SHOWED NO HEART PROBLEMS NOR THAT DAY SHE HAD HER HEART ATTACK. THERE WAS NO SHOCKS OR PACING. I HAD THEM TURN THE DEVICE OFF THAT IS WHY IT WAS IN MY POSSESSION; I COULDN'T HAVE MY MOTHER CREMATED WITH IT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22507908&pc=PYX | 1220648-2025-30353 | 2025-06-27 04:00 | Death | ABIOMED, INC. | 2025-07-15 04:00 | PYX | IMPELLA RP FLEX | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia Stroke | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT DEVELOPED AN ACCESS SITE BLEED FOLLOWING REMOVAL OF THE INTRODUCER AND ADVANCEMENT OF THE REPOSITIONING SHEATH. TWO PURSE STRING SUTURES WERE ADDED TO ACHIEVE HEMOSTASIS AT THE SITE. ROUGHLY FIVE DAYS INTO SUPPORT, THE PATIENT SUFFERED AN ISCHEMIC STROKE. SUPPORT CONTINUED FOR TWO MORE DAYS, AT WHICH POINT THE DECISION WAS MADE TO WITHDRAW CARE AND THAT PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND STROKE/CVA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ACCESS SITE BLEEDING: THE ROOT CAUSE OF ACCESS SITE BLEEDING IS DETERMINED TO BE USE ISSUE BECAUSE THE HEMOSTASIS WAS ACHIEVED BY PLACING 2 ADDITIONAL PURSE STRING SUTURES. STROKE/CVA: THE ROOT CAUSE OF THE STROKE/CVA WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA RP FLEX WITH SMARTASSIST INSTRUCTIONS FOR USE & CLINICAL REFERENCE MANUAL SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP FLEX SMARTASSIST SYSTEM CATHETER USE OF THE REPOSITIONING SHEATH AND THE 23 FR PEEL-AWAY INTRODUCER ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE INTERNAL JUGULAR VEIN. CLOSE AND DRESS THE WOUND .¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22508024&pc=MND | 3005099803-2025-03196 | 2025-06-19 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-15 04:00 | MND | SPEEDBAND SUPERVIEW SUPER 7 | Use of Device Problem; Failure to Fire | Cardiac Arrest | 510K EXEMPT |  | 1 | Event Description: NOTE: THIS REPORT PERTAINS TO ONE OF THREE SPEEDBAND SUPERVIEW SUPER 7 DEVICES THAT WERE USED IN THE SAME PATIENT AND PROCEDURE. IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT THREE SPEEDBAND SUPERVIEW SUPER 7 WERE USED DURING AN OESOPHAGOGASTRODUODENOSCOPY (OGD) PROCEDURE FOR OESOPHAGEAL VARICES PERFORMED ON (B)(6) 2025. DURING THE PROCEDURE, IT WAS REPORTED THAT THE BANDING KITS WOULD NOT DEPLOY. ADDITIONALLY, THE SAME PROBLEM HAPPENED ON THE OTHER TWO SPEEDBAND SUPERVIEW SUPER 7 DEVICES. ON JUNE 27, 2025, ADDITIONAL INFORMATION WAS RECEIVED THAT THE TRIPWIRE WAS NOT SECURED ON THE SLOT. THE PATIENT EXPERIENCED A CARDIAC ARREST AND PASSED AWAY. HOWEVER, IT WAS NOT CONFIRMED IF THE DEVICES CONTRIBUTED TO THE PATIENT COMPLICATIONS. NOTE: IT WAS REPORTED THAT THE TRIP WIRE WAS UNABLE TO BE SECURED IN THE SLOT ON THE HANDLE UNIT; HOWEVER, PER THE INSTRUCTIONS FOR USE (IFU), SECURE TRIP WIRE IN SLOT ON HANDLE UNIT. NO FURTHER INFORMATION HAS BEEN OBTAINED DESPITE GOOD-FAITH EFFORTS.Event Description: NOTE: THIS REPORT PERTAINS TO ONE OF THREE SPEEDBAND SUPERVIEW SUPER 7 DEVICES THAT WERE USED IN THE SAME PATIENT AND PROCEDURE. PROCEDURE SUMMARY: IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT THREE SPEEDBAND SUPERVIEW SUPER 7 WERE USED DURING AN OESOPHAGOGASTRODUODENOSCOPY (OGD) PROCEDURE FOR OESOPHAGEAL VARICES PERFORMED ON (B)(6) 2025. EVENT SUMMARY: DURING THE PROCEDURE, IT WAS REPORTED THAT THE BANDING KITS WOULD NOT DEPLOY. ADDITIONALLY, THE SAME PROBLEM HAPPENED ON THE OTHER TWO SPEEDBAND SUPERVIEW SUPER 7 DEVICES. ON JUNE 27, 2025, ADDITIONAL INFORMATION WAS RECEIVED THAT THE TRIPWIRE WAS NOT SECURED ON THE SLOT. THE PATIENT EXPERIENCED A CARDIAC ARREST AND PASSED AWAY. HOWEVER, IT WAS NOT CONFIRMED IF THE DEVICES CONTRIBUTED TO THE PATIENT COMPLICATIONS. NOTE: IT WAS REPORTED THAT THE TRIP WIRE WAS UNABLE TO BE SECURED IN THE SLOT ON THE HANDLE UNIT; HOWEVER, PER THE INSTRUCTIONS FOR USE (IFU), SECURE TRIP WIRE IN SLOT ON HANDLE UNIT. ADDITIONAL INFORMATION RECEIVED ON JULY 22,2025. WAS THE DEVICE USED TO TREAT ACTIVELY BLEEDING OESOPHAGEAL VARICES? YES. IT WAS STATED THAT THE SECOND TWO BANDING DEVICES WERE SUCCESSFUL, WAS THE PATIENT STABLE AFTER THE PROCEDURE? NO, ARRESTED AND DIED DURING PROCEDURE. HOW MANY DAYS POST PROCEDURE DID THE PATIENT DIE? AS ABOVE. WHAT WAS THE CAUSE OF DEATH? I HAVE NOT SEEN DEATH CERTIFICATE BUT IN MY OPINION SEVERE LIVER FAILURE / COAGULOPATHY / VARICES. DID THE PATIENT HAVE ANY CO-MORBIDITIES THAT CONTRIBUTED TO HER DEATH? YES -END STAGE LIVER CIRRHOSIS. Manufacturer Narrative: BLOCK B2: THE EXACT DATE OF DEATH WAS NOT REPORTED. IT WAS ESTIMATED TO BE (B)(6) 2025, BASED ON EVENT DATE REPORTED. BLOCK H6: IMDRF DEVICE CODE A050501 CAPTURES THE REPORTABLE EVENT OF BANDS UNABLE TO DEPLOY. IMDRF PATIENT CODE E0602 CAPTURES THE REPORTABLE PATIENT COMPLICATION OF CARDIAC ARREST. IMDRF IMPACT CODE F02 CAPTURES THE REPORTED DEATH. Manufacturer Narrative: BLOCK B2: THE EXACT DATE OF DEATH WAS NOT REPORTED. IT WAS ESTIMATED TO BE JUNE 19, 2025, BASED ON EVENT DATE REPORTED. BLOCK H6: IMDRF DEVICE CODE A050501 CAPTURES THE REPORTABLE EVENT OF BANDS UNABLE TO DEPLOY. IMDRF PATIENT CODE E0602 CAPTURES THE REPORTABLE PATIENT COMPLICATION OF CARDIAC ARREST. IMDRF IMPACT CODE F02 CAPTURES THE REPORTED DEATH. BLOCK H2 (ADDITIONAL INFORMATION): BLOCK B5 HAS BEEN UPDATED BASED ON THE ADDITIONAL INFORMATION RECEIVED ON JULY 22, 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22508028&pc=MND | 3005099803-2025-03199 | 2025-06-19 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-15 04:00 | MND | SPEEDBAND SUPERVIEW SUPER 7 | Use of Device Problem; Failure to Fire | Cardiac Arrest | 510K EXEMPT |  | 1 | Event Description: NOTE: THIS REPORT PERTAINS TO ONE OF THREE SPEEDBAND SUPERVIEW SUPER 7 DEVICES THAT WERE USED IN THE SAME PATIENT AND PROCEDURE. IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT THREE SPEEDBAND SUPERVIEW SUPER 7 WERE USED DURING AN OESOPHAGOGASTRODUODENOSCOPY (OGD) PROCEDURE FOR OESOPHAGEAL VARICES PERFORMED ON (B)(6) 2025. DURING THE PROCEDURE, IT WAS REPORTED THAT THE BANDING KITS WOULD NOT DEPLOY. ADDITIONALLY, THE SAME PROBLEM HAPPENED ON THE OTHER TWO SPEEDBAND SUPERVIEW SUPER 7 DEVICES. ON JUNE 27, 2025, ADDITIONAL INFORMATION WAS RECEIVED THAT THE TRIPWIRE WAS NOT SECURED ON THE SLOT. THE PATIENT EXPERIENCED A CARDIAC ARREST AND PASSED AWAY. HOWEVER, IT WAS NOT CONFIRMED IF THE DEVICES CONTRIBUTED TO THE PATIENT COMPLICATIONS. NOTE: IT WAS REPORTED THAT THE TRIP WIRE WAS UNABLE TO BE SECURED IN THE SLOT ON THE HANDLE UNIT; HOWEVER, PER THE INSTRUCTIONS FOR USE (IFU), SECURE TRIP WIRE IN SLOT ON HANDLE UNIT. NO FURTHER INFORMATION HAS BEEN OBTAINED DESPITE GOOD-FAITH EFFORTS.Event Description: NOTE: THIS REPORT PERTAINS TO ONE OF THREE SPEEDBAND SUPERVIEW SUPER 7 DEVICES THAT WERE USED IN THE SAME PATIENT AND PROCEDURE. PROCEDURE SUMMARY: IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT THREE SPEEDBAND SUPERVIEW SUPER 7 WERE USED DURING AN OESOPHAGOGASTRODUODENOSCOPY (OGD) PROCEDURE FOR OESOPHAGEAL VARICES PERFORMED ON (B)(6) 2025. EVENT SUMMARY: DURING THE PROCEDURE, IT WAS REPORTED THAT THE BANDING KITS WOULD NOT DEPLOY. ADDITIONALLY, THE SAME PROBLEM HAPPENED ON THE OTHER TWO SPEEDBAND SUPERVIEW SUPER 7 DEVICES. ON JUNE 27, 2025, ADDITIONAL INFORMATION WAS RECEIVED THAT THE TRIPWIRE WAS NOT SECURED ON THE SLOT. THE PATIENT EXPERIENCED A CARDIAC ARREST AND PASSED AWAY. HOWEVER, IT WAS NOT CONFIRMED IF THE DEVICES CONTRIBUTED TO THE PATIENT COMPLICATIONS. NOTE: IT WAS REPORTED THAT THE TRIP WIRE WAS UNABLE TO BE SECURED IN THE SLOT ON THE HANDLE UNIT; HOWEVER, PER THE INSTRUCTIONS FOR USE (IFU), SECURE TRIP WIRE IN SLOT ON HANDLE UNIT. ADDITIONAL INFORMATION RECEIVED ON JULY 22,2025. WAS THE DEVICE USED TO TREAT ACTIVELY BLEEDING OESOPHAGEAL VARICES? YES. IT WAS STATED THAT THE SECOND TWO BANDING DEVICES WERE SUCCESSFUL, WAS THE PATIENT STABLE AFTER THE PROCEDURE? NO, ARRESTED AND DIED DURING PROCEDURE. HOW MANY DAYS POST PROCEDURE DID THE PATIENT DIE? AS ABOVE. WHAT WAS THE CAUSE OF DEATH? I HAVE NOT SEEN DEATH CERTIFICATE BUT IN MY OPINION SEVERE LIVER FAILURE / COAGULOPATHY / VARICES. DID THE PATIENT HAVE ANY CO-MORBIDITIES THAT CONTRIBUTED TO HER DEATH? YES -END STAGE LIVER CIRRHOSIS. Manufacturer Narrative: BLOCK B2: THE EXACT DATE OF DEATH WAS NOT REPORTED. IT WAS ESTIMATED TO BE JUNE 19, 2025, BASED ON EVENT DATE REPORTED. BLOCK H6: IMDRF DEVICE CODE A050501 CAPTURES THE REPORTABLE EVENT OF BANDS UNABLE TO DEPLOY. IMDRF PATIENT CODE E0602 CAPTURES THE REPORTABLE PATIENT COMPLICATION OF CARDIAC ARREST. IMDRF IMPACT CODE F02 CAPTURES THE REPORTED DEATH. Manufacturer Narrative: BLOCK B2: THE EXACT DATE OF DEATH WAS NOT REPORTED. IT WAS ESTIMATED TO BE JUNE 19, 2025, BASED ON EVENT DATE REPORTED. BLOCK H6: IMDRF DEVICE CODE A050501 CAPTURES THE REPORTABLE EVENT OF BANDS UNABLE TO DEPLOY. IMDRF PATIENT CODE E0602 CAPTURES THE REPORTABLE PATIENT COMPLICATION OF CARDIAC ARREST. IMDRF IMPACT CODE F02 CAPTURES THE REPORTED DEATH. BLOCK H2 (ADDITIONAL INFORMATION): BLOCK B5 HAS BEEN UPDATED BASED ON THE ADDITIONAL INFORMATION RECEIVED ON JULY 22, 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22508289&pc=OZD | 1220648-2025-30356 | 2025-06-28 04:00 | Death | ABIOMED, INC. | 2025-07-15 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND LOST PULSES IN THE LEFT LOWER EXTREMITY. THE IMPELLA CP WAS REMOVED AND REPLACED. NO PATIENT HARM WAS REPORTED.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. PER MEDICAL SAFETY OFFICER REVIEW USE OF THE DEVICE CANNOT CONCLUSIVELY BE ASSOCIATED WITH THE OUTCOME. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP PUMP 1 AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. B2 DEATH AND DATE OF DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30356. B5 MSO STATEMENT THE FIRST IMPELLA CP STARTED A CASCADE OF EVENTS, FOR FURTHER EVENTS RELATED TO MANUFACTURER DEVICE REPORT 1220648-2025-30356, SEE MANUFACTURER DEVICE REPORT 1220648-2025-30359. H1 TYPE OF REPORT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. H6 HEALTH EFFECT - IMPACT CODE 1802 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30356. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22508369&pc=NRY | 3011370111-2025-00065 | 2024-12-23 05:00 | Death | NEURAVI LTD. | 2025-07-15 04:00 | NRY | EMBOTRAP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Perforation of Vessels | K193063 |  | 1 | Event Description: ON 25-JUN-2025, THE HEALTHCARE PROFESSIONAL REPORTED THAT SHE WAS MADE AWARE OF AN EVENT THAT OCCURRED ON 23-DEC-2024. THE PROCEDURE WAS A MECHANICAL THROMBECTOMY. THE 86-YEAR-OLD PATIENT PRESENTED WITH ¿LOTS OF COMORBIDITIES¿ PRESENTED WITH AN OCCLUSION IN THE DISTAL M2 SEGMENT OF THE MIDDLE CEREBRAL ARTERY (MCA). IT WAS REPORTED THAT ¿IT WAS CONSIDERED WHETHER PATIENT WAS A SUITABLE CANDIDATE FOR THROMBECTOMY. [THE] PATIENT HAD LOTS OF VESSEL ATROPHIES. VESSELS [DID NOT] LOOK LIKE THEY WOULD BE SUITABLE FOR USING A STENT RETRIEVER, SO THE FIRST PASS WAS PERFORMED WITH ASPIRATION ONLY. THE FIRST PASS WAS PERFORMED USING A RED 62 REPERFUSION CATHETER (PENUMBRA) WAS NOT SUCCESSFUL, NO [CLOT] WAS REMOVED. THE SECOND PASS WAS PERFORMED VIA CO-ASPIRATION, USING A 5MM X 22MM EMBOTRAP III REVASCULARIZATION DEVICE (ET307522 / 24J087AV) DELIVERED VIA THE PHENOM¿ 021 MICROCATHETER (MEDTRONIC), THE RED 62 REPERFUSION CATHETER.¿ IT WAS REPORTED THAT ONE PASS OF THE EMBOTRAP III DEVICE AND THE CLOT WAS REMOVED, AND THE VESSEL WAS OPENED, ¿BUT THE VESSEL EVULSED AND A BLEED WAS SEEN JUST PAST THE BEND OF THE M1/M2 SEGMENTS.¿ THE PHYSICIAN PERFORMED COIL EMBOLIZATION TO STOP THE BLEED, BUT THE PATIENT EXPIRED FOLLOWING THE PROCEDURE. THE NEUROINTERVENTIONALIST COMMENTED THAT ¿THE OUTCOME WAS NOT CAUSED BY EMBOTRAP AND THAT HE WOULD HAVE HAD THE SAME OUTCOME WITH ANY STENT RETRIEVER THAT WAS USED. THE OUTCOME WAS PATIENT-RELATED AND NOT DEVICE-RELATED. THE PATIENT WAS PROBABLY NOT SUITABLE FOR PROCEDURE, ELDERLY, VERY UNWELL WITH LOTS OF COMORBIDITIES AND FRIABLE VESSELS.¿ Manufacturer Narrative: MANUFACTURER¿S REF. NO: (B)(4). INFORMATION REGARDING PATIENT IDENTIFIER, DATE OF BIRTH, SEX, GENDER, WEIGHT, RACE, AND ETHNICITY WERE NOT PROVIDED. SECTION B.2: DATE OF DEATH: THE DATE OF DEATH IS NOT KNOWN / REPORTED. SECTION E.1: THE INITIAL REPORTER PHONE IS NOT AVAILABLE / REPORTED. BASED ON COMPLAINT INFORMATION, THE DEVICE IS NOT AVAILABLE TO BE RETURNED FOR ANALYSIS. A REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH THIS LOT (24J087AV) TOP AND SUB ASSEMBLY PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESSES RELATED TO THE REPORTED COMPLAINT. THERE WERE NO INTERNAL ACTIONS RELATED TO DEVICE MANUFACTURE OR INSPECTION. PRODUCT ANALYSIS CANNOT BE CONDUCTED AS THE PRODUCT WAS NOT RETURNED FOR ANALYSIS. NO DETERMINATION OF CAUSES AND POSSIBLE CONTRIBUTING FACTORS COULD BE MADE. AS SUCH, THE INVESTIGATION WILL BE CLOSED. THERE WAS NO DEVICE PERFORMANCE ISSUE OR DEVICE MALFUNCTION REPORTED DURING THE PROCEDURE. BASED ON THE MANUFACTURING DOCUMENTATION REVIEW, THERE IS NO INDICATION THAT THE EVENT IS RELATED TO THE DEVICE MANUFACTURING PROCESS. AS PART OF THE POST MARKET SURVEILLANCE PROGRAM, INFORMATION FROM THIS COMPLAINT IS TRENDED FOR STATISTICAL SIGNALS AND CORRECTIVE / PREVENTIVE ACTION MAY BE TRIGGERED LATER. SINCE THERE WAS NO EVIDENCE TO SUGGEST THE EVENT WAS RELATED TO A MANUFACTURING OR DESIGN ISSUE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. VESSEL INJURY AND CEREBRAL HEMORRHAGE ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE EMBOTRAP III DEVICE AND ARE MENTIONED IN THE INSTRUCTIONS FOR USE (IFU) AS SUCH. THERE WAS NO ALLEGED QUALITY ISSUE RELATED TO THE USED DEVICES, AS THE DEVICES PERFORMED AS INTENDED. THERE ARE PATIENT AND PROCEDURAL FACTORS, INCLUDING VESSEL CHARACTERISTICS (I.E., ATROPHY, FRIABLE VESSELS), ADVANCED AGE WITH MULTIPLE COMORBIDITIES, DEVICE INTERACTION, AND OPERATOR TECHNIQUES, THAT MAY HAVE CONTRIBUTED TO THE REPORTED EVENTS RATHER THAN THE DESIGN OR MANUFACTURE OF THE DEVICE. ADDITIONALLY, THE PHYSICIAN STATED THE EVENTS WERE ATTRIBUTED TO PATIENT-RELATED FACTORS AND THE PATIENT WAS PROBABLY NOT SUITABLE FOR PROCEDURE. HOWEVER, SINCE THE EVENTS OCCURRED DURING THE USE OF THE EMBOTRAP DEVICE, THE CORRELATION BETWEEN EVENTS TO THE DEVICE USED AS A CONTRIBUTING FACTOR CANNOT BE RULED OUT. ADDITIONALLY, THE EVENTS REQUIRED A SURGICAL INTERVENTION, AND THE PATIENT ULTIMATELY EXPIRED SHORTLY AFTER THE PROCEDURE. BASED ON THIS INFORMATION, THIS EVENT MEETS US FDA REPORTING CRITERIA UNDER 21 CFR 803 WITH A CLASSIFICATION OF ¿DEATH.¿ THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. THE MANUFACTURER WILL SUBMIT A SUPPLEMENTAL REPORT IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT. ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. Manufacturer Narrative: MANUFACTURER¿S REF. NO: (B)(4). THE PURPOSE OF THIS MDR SUBMISSION IS TO INCLUDE THE ADDITIONAL EVENT INFORMATION RECEIVED ON 24-JUL-2025. [ADDITIONAL INFORMATION]: ON 24-JUL-2025, ADDITIONAL INFORMATION WAS RECEIVED. THE INFORMATION CONFIRMED THE DATE OF DEATH WAS (B)(6) 2024. THE SAME DAY AS THE PROCEDURE. THE PATIENT WAS MOVED TO THE ICU FOLLOWING THE PROCEDURE BUT LATER DIED THE SAME DAY. THE OFFICIAL CAUSE OF DEATH WAS BRAIN HEMORRHAGE. THE INFORMATION INCLUDED ADDITIONAL COMMENT FROM THE PHYSICIAN: ¿VESSEL EVULSED WHEN USING STENT RETRIEVER (EMBOTRAP) PLUS ASPIRATION ON THE SECOND PASS RESULTING IN A HEMORRHAGE. EMBOTRAP WAS THE STENT RETRIEVER BEING USED AND THEREFORE CONTRIBUTED TO THE VESSEL EVULSION AND SUBSEQUENT HEMORRHAGE, BUT IT IS THE PHYSICIAN¿S OPINION THAT ANY STENT RETRIEVER THAT WAS USED WOULD HAVE PRODUCED THE SAME OUTCOME.¿ PER THE ADDITIONAL INFORMATION, THE PATIENT WAS 86-YEARS-OLD, ELDERLY IN FRAIL CONDITION WITH THE FOLLOWING COMORBIDITIES: PERIPHERAL VASCULAR DISEASE, ATRIAL FIBRILLATION, DIABETES, AND CALCIFIED ARTERIES. UPDATED SECTIONS: B.4, B.7, G.3, G.6, H.2, AND H.11. THE MANUFACTURER WILL SUBMIT A SUPPLEMENTAL REPORT IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT. ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22508643&pc=OZD | 1220648-2025-30357 | 2025-06-28 04:00 | Death | ABIOMED, INC. | 2025-07-15 04:00 | OZD | IMPELLA CP | Malposition of Device | Thrombocytopenia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND WAS AT RISK FOR HEPARIN INDUCED THROMBOCYTOPENIA SO PLATELETS WERE TRANSFUSED AND SYSTEMIC ANTICOAGULATION WAS ON HOLD. LATER, THE PATIENT WAS ROLLED FOR A BED CHANGE. DURING THE MOVEMENT, THE PUMP MOVED INTO THE AORTA. THE PUMP WAS REPOSITIONED BUT THE PATIENT CODED AND EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES AND THROMBOCYTOPENIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. POSITIONING ISSUE: THE ROOT CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY DUE TO PATIENT MOVEMENT AS EVIDENCED BY CLINICAL DETAIL OF PUMP SHIFTING WHEN PATIENT WAS ROLLED OVER FOR BED CHANGE AND 'IMPELLA POSITION IN AORTA' ALARM TRIGGERING. THROMBOCYTOPENIA: THE CAUSE OF THE THROMBOCYTOPENIA WAS NOT DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22508672&pc=DSQ | 2916596-2025-04392 | 2025-06-20 04:00 | Death | THORATEC CORPORATION | 2025-07-15 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN UNKNOWN CAUSE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22508687&pc=DSQ | 2916596-2025-04602 | 2025-07-01 04:00 | Death | THORATEC CORPORATION | 2025-07-15 04:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Dyspnea; Muscle Weakness; Pneumonia; Dizziness; Lethargy | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD WEAKNESS, LETHARGY, SHORTNESS OF BREATH, AND LIGHTHEADEDNESS IN THE SETTING OF PULMONARY EMBOLISM AND PNEUMONIA WITH SUBSEQUENT TRANSITION TO COMFORT CARE AT THE PATIENT'S REQUEST. THE PATIENT PASSED AWAY ON COMFORT CARE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LVAS IFU, REV. A IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING PERIPHERAL THROMBOEMBOLIC EVENTS AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT,¿ LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. UNDER ¿ANTICOAGULATION," THIS SECTION ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22509478&pc=DSQ | 2916596-2025-04530 | 2025-07-01 04:00 | Death | THORATEC CORPORATION | 2025-07-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS NOT WAKING UP. A COMPUTED TOMOGRAPHY (CT) SCAN "LOOKED EMBOLIC". THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO AN EMBOLIC STROKE. THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WAS NOT RETURNED FOR EVALUATION. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22509864&pc=NVN | 2182208-2025-03428 | 2025-06-01 04:00 | Death | MEDTRONIC, INC. | 2025-07-15 04:00 | NVN | SELECTSECURE MRI SURESCAN | Appropriate Term/Code Not Available | Heart Failure/Congestive Heart Failure | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING PARTIAL PERFORATIONS AFTER LEFT BUNDLE BRANCH AREA PACING. THE STUDY INCLUDED HUMAN PARTICIPANTS AND AN EXPERIMENTAL SWINE STUDY. THE AUTHORS DESCRIBED ELEVEN PATIENT DEATHS; THE CAUSES OF DEATH WERE DUE TO NONCARDIOVASCULAR CAUSES IN NINE AND TWO WERE DUE TO HEART FAILURE. THERE IS NO ALLEGATION OF LEAD-DEATH RELATEDNESS INDICATED IN THE ARTICLE; HOWEVER, THE CAUSE OF DEATH AND LEAD-DEATH RELATEDNESS HAS BEEN REQUESTED, BUT NOT YET RECEIVED. THERE WERE PATIENTS WHO EXPERIENCED COMPLETE AND PARTIAL PERFORATIONS WHICH REQUIRED LEAD REPOSITIONING. ONE PATIENT DEVELOPED A CEREBRAL EMBOLISM CAUSED BY ATRIAL FIBRILLATION (AF) DURING THE FOLLOW-UP PERIOD. THERE WERE LEADS WHICH EXHIBITED LOSS OF CAPTURE, INCREASED CAPTURE THRESHOLDS, AND DECREASED IMPEDANCE. IN THE SWINE MODELS, FOUR LEADS WERE IMPLANTED WITH TWO PARTIAL PERFORATIONS AND ONE COMPLETE PERFORATION. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CLINICAL OUTCOMES AND ELECTROPHYSIOLOGICAL CHARACTERISTICS OF PARTIAL PERFORATION AFTER LEFT BUNDLE BRANCH AREA PACING. JACC CLINICAL ELECTROPHYSIOLOGY. 2025; 11:1279¿1289. DOI: 10.1016/J.JACEP.2025.01.016 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22509890&pc=DSQ | 2916596-2025-04424 | 2025-06-01 04:00 | Death | THORATEC CORPORATION | 2025-07-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: SECTION A: PATIENT INFORMATION NOT YET PROVIDED. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. THIS TYPE OF REPORTED EVENT WILL CONTINUE TO BE MONITORED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22510208&pc=MLV | 2135147-2025-03866 | 2025-06-20 04:00 | Death | ABBOTT MEDICAL | 2025-07-15 04:00 | MLV | AMPLATZER POST-INFARCT VSD OCCLUDER | Adverse Event Without Identified Device or Use Problem | Pericardial Effusion | P040040 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 24MM AMPLATZER POST-INFARCT VENTRICULAR SEPTAL DEFECT (VSD) OCCLUDER. IT WAS NOTED DURING THE PROCEDURE THAT THE PATIENT DEVELOPED A CARDIAC EFFUSION. THE DECISION WAS MADE TO REMOVED THE OCCLUDER FROM THE PATIENT'S BODY WITHOUT DISCONNECTING FROM THE DELIVERY CABLE. THE PATIENT PASSED AWAY ON AN UNKNOWN DATE DUE TO AN UNKNOWN REASON. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE IS NOT RETURNING FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22510234&pc=DSQ | 2916596-2025-04573 | 2025-03-01 05:00 | Death | THORATEC CORPORATION | 2025-07-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Abdominal Pain | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT IN (B)(6) 2025 WITH A VACUUM-ASSISTED WOUND CLOSURE AND COMPLAINED OF WORSENING ABDOMINAL PAIN. THE PATIENT PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: SECTION B3: EVENT DATE WAS ESTIMATED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22511033&pc=DSQ | 2916596-2025-04498 | 2025-06-05 04:00 | Death | THORATEC CORPORATION | 2025-07-15 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER IMPLANTATION OF HEARTMATE 3, THE PATIENT DEVELOPED HYPOTENSION AND REQUIRED HIGH DOSE OF CATECHOLAMINES WITH EMPRESSINE. ECHOCARDIOGRAM CONFIRMED GOOD POSITION PF THE LEFT VENTRICULAR ASSIST DEVICE (LVAD), THE RIGHT VENTRICLE WAS SEVERELY IMPAIRED, WHICH INDICATED FOR TEMPORARY RIGHT VENTRICULAR ASSIST DEVICE (RVAD). DUE TO THE PATIENT'S CRITICAL CONDITION, IMPLANTATION OF THE RVAD WAS UNSUCCESSFUL. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22511757&pc=MKJ | 1220908-2025-02533 | 2025-06-19 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-07-15 04:00 | MKJ | AED PLUS | Incorrect Interpretation of Signal | Insufficient Information | P160015 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE GAVE A "NO SHOCK ADVISED" PROMPT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: JUSTIFICATION FOR NO UDI: THIS DEVICE WAS MANUFACTURED BEFORE THE UDI REQUIREMENTS. EVALUATION RESULTS: ZOLL MEDICAL CORPORATION EVALUATED THE DEVICE AND THE DEVICE PERFORMED TO SPECIFICATION. THE DEVICE WAS PUT THROUGH EXTENSIVE TESTING INCLUDING FUNCTIONAL TESTING WITHOUT DUPLICATING A MALFUNCTION. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. REVIEW OF THE CLINICAL LOGS OBSERVED TESTING VIA A SIMULATOR INSTEAD OF THE PATIENT EVENT BASED OFF THE RECORDED IMPEDANCE READINGS. THE TESTING HAD OVERWRITTEN THE CLINICAL LOG OF THE EVENT AS THE AED PLUS IS ONLY DESIGNED TO HOLD ONE CLINICAL LOG. WITHOUT THE CLINICAL FILES FROM THE EVENT, WE WERE UNABLE TO CONFIRM WHY THE DEVICE ADVISED NO SHOCK DURING THE EVENT. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22511814&pc=DSQ | 2916596-2025-04697 | 2025-06-20 04:00 | Death | THORATEC CORPORATION | 2025-07-15 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI SYSTEM ORGAN FAILURE. THIS WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE DEVICE OPERATED AS EXPECTED. DETAILS LEADING UP TO THE MULTI SYSTEM ORGAN FAILURE WERE NOT AVAILABLE. Manufacturer Narrative: H6: 3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION H6: HEALTH EFFECT - CLINICAL CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER: (B)(6) AND THE REPORTED MULTISYSTEM ORGAN FAILURE (MSOF) AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22511822&pc=DSQ | 2916596-2025-04425 | 2025-06-12 04:00 | Death | THORATEC CORPORATION | 2025-07-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO INTRACEREBRAL HEMORRHAGE. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22511836&pc=DSQ | 2916596-2025-04672 | 2025-07-01 04:00 | Death | THORATEC CORPORATION | 2025-07-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE (MSOF). THE PATIENT HAD RECEIVED A MASSIVE TRANSFUSION DUE TO BLEEDING WITH MSOF. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE SOURCE OF THE BLEEDING WAS INTRAPERICARDIAL BLEEDING WHICH WAS FELT TO BE FROM THE AORTA BUT COULD NOT BE IDENTIFIED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THE CURRENT REVISIONS OF THE IFU CAN ALSO BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22511847&pc=DSQ | 2916596-2025-04558 | 2025-07-02 04:00 | Death | THORATEC CORPORATION | 2025-07-15 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Muscle Weakness; Ischemia Stroke | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO ARTERIAL ISCHEMIC STROKE, MIDDLE CEREBRAL ARTERY (MCA) LEFT AND ACUTE. ADDITIONAL INFORMATION REPORTED THE PATIENT SYMPTOMS WERE DECREASED RESPONSIVENESS, LEFT-SIDED GAZE, FLACCID RIGHT SIDE. THERE WERE NO CHANGS TO THE PATIENT OR ANTICOAGULATION THAT MAY HAVE CONTRIBUTED TO THE STROKE. THE PATIENT UNDERWENT THROMBECTOMY. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: SECTION G2 CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP WAS NOT RETURNED. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22512071&pc=NPT | 2015691-2025-05751 | 2025-06-19 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-15 04:00 | NPT | SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture; Pericardial Effusion | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS AFFILIATES IN ITALY, DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE USING A 29MM SAPIEN 3 VALVE VIA THE TRANSFEMORAL APPROACH, AN ANNULAR RUPTURE OCCURRED, RESULTING IN A PERICARDIAL EFFUSION. DESPITE MULTIPLE ATTEMPTS TO MANAGE THE EFFUSION THROUGH PERICARDIOCENTESIS, THE PATIENT UNFORTUNATELY PASSED AWAY. AS PER MEDICAL OPINION, THE ROOT CAUSE OF THE EVENT WAS CALCIFICATION OF THE ANNULUS. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THAT CALCIUM IN THE ANNULUS MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22512130&pc=NPT | 2025587-2025-05038 | 2025-03-29 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-15 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Cardiac Arrest; Chest Pain; Cardiac Enzyme Elevation; Renal Failure; Ventricular Fibrillation; Syncope/Fainting | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 YEAR AND 6 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT WAS HOSPITALIZED AND PRESENTED WITH CHEST PAIN ALONG WITH ELEVATED TROPONIN I (TNI) LEVELS, WHICH CAUSED A NON-ST ELEVATION MYOCARDIAL INFARCTION (NSTEMI) TO BE SUSPECTED. AS A RESULT, A CORONARY ANGIOGRAPHY (CAG) WAS PERFORMED; HOWEVER, NO NEW MAJOR ARTERY BLOCKAGES WERE NOTED. SHORTLY THEREAFTER, THE PATIENT¿S CONSCIOUSNESS DECREASED, AND VENTRICULAR FIBRILLATION (VF) OCCURRED. SUBSEQUENTLY, CPR WAS THEN INITIATED, AND THE HEARTBEAT WAS RESTORED AFTER ONE SHOCK AT 150 JOULES (DC150J). FOLLOWING THIS, INTUBATION, ARTIFICIAL RESPIRATOR MANAGEMENT, AND ANTIARRHYTHMIC DRUG ADMINISTRATION (AMD) WERE INITIATED. IN ADDITION, AN ECHOCARDIOGRAM (ECG) AND COMPUTED TOMOGRAPHY (CT) SCAN WERE PERFORMED, WHICH REVEALED NO SIGNIFICANT CHANGES, AND THE CAUSE OF VF REMAINED UNKNOWN. CONTINUOUS HEMODIAFILTRATION (CHDF) WAS THEN STARTED; HOWEVER, THE PATIENT'S CONDITION REMAINED POOR, AND MULTIPLE CARDIAC ARRESTS OCCURRED. DESPITE RESUSCITATION EFFORTS, THE PATIENT DEVELOPED FAILURE IN MULTIPLE ORGANS AND DIED ONE DAY AFTER HOSPITALIZATION. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS VF, ALONG WITH CHRONIC RENAL FAILURE. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22512399&pc=NVZ | 2182208-2025-03431 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-07-15 04:00 | NVZ | MEDTRONIC IPG | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING EPICARDIAL LEAD PERFORMANCE IN PEDIATRIC PATIENTS WITH CONGENITAL HEART DISEASE (CHD). THE AUTHORS DESCRIBED PATIENT DEATHS WITHIN THE ONE-YEAR STUDY PERIOD; THE CAUSES OF DEATH WERE RELATED TO COMPLICATIONS OF COMPLEX CHD (AND NOT A FAILURE OF PACING), ONE WAS DUE TO BRUGADA-RELATED ICD STORM, AND OTHER UNKNOWN REASONS. THERE WERE LEADS WHICH EXHIBITED HIGH THRESHOLDS, POLARITY CHANGES, AND WERE PROGRAMMED OFF OR REMOVED AND REPLACED. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS UNKNOWN/6 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: EPICARDIAL LEAD PERFORMANCE TRENDS IN PEDIATRIC AND CONGENITAL HEART DISEASE. WORLD JOURNAL FOR PEDIATRIC AND CONGENITAL HEART SURGERY. 2025. VOL. 16(3) 384-387. DOI: 10.1177/21501351241293722 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22512404&pc=DTB | 2182208-2025-03432 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-07-15 04:00 | DTB | CAPSURE EPICARDIAL LEAD | Appropriate Term/Code Not Available | Insufficient Information | P950024 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING EPICARDIAL LEAD PERFORMANCE IN PEDIATRIC PATIENTS WITH CONGENITAL HEART DISEASE (CHD). THE AUTHORS DESCRIBED PATIENT DEATHS WITHIN THE ONE-YEAR STUDY PERIOD; THE CAUSES OF DEATH WERE RELATED TO COMPLICATIONS OF COMPLEX CHD (AND NOT A FAILURE OF PACING), ONE WAS DUE TO BRUGADA-RELATED ICD STORM, AND OTHER UNKNOWN REASONS. THERE WERE LEADS WHICH EXHIBITED HIGH THRESHOLDS, POLARITY CHANGES, AND WERE PROGRAMMED OFF OR REMOVED AND REPLACED. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS UNKNOWN/6 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: EPICARDIAL LEAD PERFORMANCE TRENDS IN PEDIATRIC AND CONGENITAL HEART DISEASE. WORLD JOURNAL FOR PEDIATRIC AND CONGENITAL HEART SURGERY. 2025. VOL. 16(3) 384-387. DOI: 10.1177/21501351241293722 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22513046&pc=DSQ | 2916596-2025-04608 | 2025-07-02 04:00 | Death | THORATEC CORPORATION | 2025-07-15 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A KNOWN AORTIC DISSECTION. THE PATIENT CODED AT HOME. THEY WERE TRANSPORTED TO THE HOSPITAL BY EMERGENCY MEDICAL SERVICES (EMS) AND WAS DECLARED DEAD UPON ARRIVAL. THE VENTRICULAR ASSIST DEVICE (VAD) FUNCTIONED AS INTENDED AND THERE WERE NO ALARMS NOTED. THE INTENSIVE CARE UNIT (ICU) NURSE WAS INSTRUCTED ON HOW TO TURN OFF THE PUMP. NO OTHER VAD INTERVENTIONS WERE NEEDED. NO AUTOPSY WAS REQUESTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE PUMP WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION D4: EXPIRATION DATE AND MANUFACTURE DATE. (H4) WILL BE PROVIDED UPON INVESTIGATION. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22513089&pc=LWS | 9614453-2025-02737 | 2025-06-15 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-07-15 04:00 | LWS | COBALT¿ XT DR MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED, AND A CAUSE OF DEATH WAS NOT PROVIDED. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS IMPLANTED 25 DAYS PRIOR TO THE PATIENT¿S DEATH. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID 6935M62 LOT# SERIAL # (B)(6); IMPLANTED: (B)(6) 2025; EXPLANTED: PRODUCT ID 5076-52 LOT # SERIAL# (B)(6); IMPLANTED: (B)(6) 2025; EXPLANTED: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22513118&pc=LWS | 2649622-2025-19144 | 2025-06-15 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-15 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED, AND A CAUSE OF DEATH WAS NOT PROVIDED. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS IMPLANTED 25 DAYS PRIOR TO THE PATIENT¿S DEATH. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: DDPA2D4, SERIAL#: (B)(6), IMPLANTED: (B)(6) 2025, PRODUCT ID: 5076-52, SERIAL#: (B)(6), IMPLANTED: (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22513120&pc=DTB | 2649622-2025-19145 | 2025-06-15 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-15 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED, AND A CAUSE OF DEATH WAS NOT PROVIDED. THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS IMPLANTED 25 DAYS PRIOR TO THE PATIENT¿S DEATH. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID DDPA2D4 LOT# SERIAL# (B)(6) IMPLANTED: (B)(6) 2025: PRODUCT ID 6935M62 LOT# SERIAL# (B)(6) IMPLANTED: (B)(6) 2025: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22513904&pc=FKX | 0008030665-2025-01593 | 2025-07-02 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-07-15 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K222318 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY SUDDENLY. THE CAUSE OF DEATH WAS UNKNOWN. IT WAS NOT REPORTED THAT THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF DEATH. ADDITIONAL INFORMATION REQUESTED BUT HAS NOT BEEN OBTAINED.Event Description: IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY SUDDENLY. THE CAUSE OF DEATH WAS UNKNOWN. IT WAS NOT REPORTED THAT THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF DEATH. ADDITIONAL INFORMATION REQUESTED BUT HAS NOT BEEN OBTAINED. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN D9 AND H3. CORRECTION PROVIDED IN G4. PLANT INVESTIGATION: A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR NO SHOWED SIGNS OF PHYSICAL DAMAGE. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. SYSTEM AIR LEAK TEST PASSED. VALVE ACTUATION TEST PASSED. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THERE WERE NO VISUAL DISCREPANCIES ENCOUNTERED DURING THE INTERNAL INSPECTION. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: THERE IS NO INDICATION OF A SERIOUS INJURY, PATIENT DEATH, OR OTHER ADVERSE EVENT RELATED TO A FRESENIUS PRODUCT OR OTHER ISSUE WARRANTING FURTHER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22515514&pc=NJK | 3007797756-2025-00025 | 2025-04-09 04:00 | Death | PULMONX CORPORATION | 2025-07-15 04:00 | NJK | ZEPHYR ENDOBRONCHIAL VALVE (EBV) | Adverse Event Without Identified Device or Use Problem | Pneumonia; Respiratory Failure | P180002 |  | 1 | Event Description: THE PATIENT RECEIVED 4 VALVES ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT WAS ADMITTED TO THE HOSPITAL WITH COPD EXACERBATIONS, PNEUMONIA AND RAPID DECLINE. THE PATIENT WAS RELEASED FROM THE HOSPITAL ON (B)(6) 2025. PATIENT EXPIRED (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22515674&pc=QBJ | 3004753838-2025-185414 | 2023-08-09 04:00 | Death | DEXCOM, INC. | 2025-07-15 04:00 | QBJ | DEXCOM G6 CONTINUOUS GLUCOSE MONITORING SYSTEM | Insufficient Information | Cardiac Arrest; Low Blood Pressure/ Hypotension; Renal Failure; Diabetic Ketoacidosis; Loss of consciousness | DEN170088 |  | 1 | Event Description: DECEDENT¿S ATTORNEY FILED A LAWSUIT AGAINST DEXCOM THAT ALLEGES THE PATIENT WITH TYPE 1 DIABETES WAS A G6 USER WITH AN INTEGRATED OMNIPOD 5. SHE HAD USED OP5 SINCE (B)(6) 2022. THE PATIENT ALLEGEDLY EXPERIENCED HIGH GLUCOSE READINGS ON (B)(6) 2023. SHE DIRECTED HER OMNIPOD TO ADMINISTER SEVERAL DOSES OF INSULIN. BELIEVING THAT THE INSULIN HAD BEEN DELIVERED, SHE WENT TO BED. THE PATIENT'S SPOUSE ALLEGEDLY SLEPT IN THE SAME ROOM WITH HER THAT EVENING TO MONITOR HER FOR HYPOGLYCEMIC SHOCK FOLLOWING THE TREATMENT FOR HYPERGLYCEMIA. THE SPOUSE REPORTEDLY WOKE SOMETIME BETWEEN 0600-0700 WHILE THE PATIENT WAS STILL SLEEPING. THE SPOUSE SUBSEQUENTLY FELL BACK ASLEEP ON THE COUCH AND WAS WAKENED BY A PHONE CALL AT 0900. AT THAT TIME, THE SPOUSE CHECKED ON THE PATIENT AND FOUND HER COLLAPSED ON THE BEDROOM FLOOR AND CALLED EMS. SHE WAS DIAGNOSED AND TREATED FOR SEVERE DIABETIC KETOACIDOSIS, ACUTE KIDNEY INJURY, AND HYPOTENSION. DURING HER HOSPITALIZATION, SHE EXPERIENCED BRADYCARDIC PULSELESS ELECTRICAL ACTIVITY -PEA ARREST, NECESSITATING ADVANCED CARDIAC LIFE SUPPORT - ACLS AND INTUBATION. HER CONDITION WAS FURTHER COMPLICATED BY ACUTE HYPOXIC RESPIRATORY FAILURE, POSSIBLY DUE TO ASPIRATION, SUSPECTED TYPE II MYOCARDIAL INFARCTION IN THE CONTEXT OF RESPIRATORY FAILURE, AND DIABETIC KETOACIDOSIS (DKA). DESPITE AGGRESSIVE TREATMENT, INCLUDING INSULIN AND HEPARIN DRIPS, AND EMPIRIC ANTIBIOTIC THERAPY, THE PATIENT'S CONDITION ALLEGEDLY DETERIORATED, LEADING TO HER DEATH ON (B)(6) 2023. AT SOME POINT, THE DECEDENT¿S PHYSICIAN AND INTENSIVE CARE UNIT (ICU) NURSE ALLEGEDLY DISCOVERED THE OMNIPOD HAD A BENT CANNULA, WHICH PREVENTED INSULIN FROM BEING DELIVERED. DATA HAS BEEN REQUESTED BUT IS PENDING EVALUATION. A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION. Manufacturer Narrative: (B)(4). DIABETES MELLITUS IS A KNOWN CAUSE OF HYPERGLYCEMIA. H6 HEALTH EFFECT - CLINICAL CODE - E0742 RESPIRATORY FAILURE. H6 HEALTH EFFECT - CLINICAL CODE - E0704 ASPIRATION/INHALATION. H6 HEALTH EFFECT - CLINICAL CODE - E0612 ISCHEMIC HEART DISEASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22515962&pc=NPT | 2015691-2025-05765 | 2024-11-16 05:00 | Death | EDWARDS LIFESCIENCES | 2025-07-15 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P140031 |  | 1 | Event Description: AS REPORTED FROM GERMANY AS A PART OF THE RETAVI STUDY, A PATIENT WITH NYHA CLASS IV HEART FAILURE AND DYSPNEA, WHO HAD PREVIOUSLY RECEIVED A 23 MM SAPIEN 3 VALVE IN THE AORTIC POSITION, UNDERWENT A VALVE-IN-VALVE TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) VIA TRANSFEMORAL APPROACH 8 YEARS AND 11 MONTHS AFTER THE INITIAL IMPLANTATION. DUE TO RAPIDLY ADVANCING CARDIOGENIC SHOCK, THE HEART TEAM MADE THE DECISION TO PERFORM AN EMERGENT SALVAGE REDO-TAVR WITHOUT PRIOR CT OR CORONARY ANGIOGRAPHY. A 23MM SAPIEN 3 ULTRA VALVE WAS IMPLANTED WITHIN THE EXISTING TRANSCATHETER HEART VALVE. ON POST OPERATIVE DAY 1, THE PATIENT EXPERIENCED AN ISCHEMIC STROKE WITH APHASIA. THE PATIENT WAS TRANSFERRED TO A REHABILITATION FACILITY. DESPITE PROLONGED CARE AND PALLIATIVE SUPPORT, THE STROKE PROGRESSED TO FATAL, AND THE PATIENT PASSED AWAY 45 DAYS POST-IMPLANTATION. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRAPROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION THAT A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PATIENT FACTORS (FEMALE, AGE OF 89 YEARS, HISTORY OF DIABETES) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22515969&pc=DSQ | 2916596-2025-04541 | 2025-06-12 04:00 | Death | THORATEC CORPORATION | 2025-07-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN UNKNOWN CAUSE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION REGARDING THIS EVENT WERE SUBMITTED TO THE ACCOUNT; HOWEVER, NO RESPONSE HAS BEEN RECEIVED AT THIS TIME. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22491344&pc=OZP | 2032227-2025-211216 | 2025-06-13 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-14 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Pain; Loss of consciousness | P160017 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY AT THE HOSPITAL ON (B)(6) 2025. THE CUSTOMER WAS ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS HEMORRHAGIC STROKE, NEUROPATHY, SEVERE BACK PAIN AND TYPE 1 DIABETES. THE CUSTOMER¿S BLOOD GLUCOSE WAS 154 MG/DL. THE EVENT INVOLVED PRODUCT(S) MMT-342G, MMT-442AK, MMT-1884. TROUBLESHOOTING WAS PERFORMED. THE CUSTOMER WAS WEARING THE INSULIN PUMP AT THE TIME OF DEATH AND PUMP WAS WORN AT TIME OF PASSING. THE CUSTOMER WAS USING SENSORS. THE INSULIN PUMP WAS LAST WORN ON UNKNOWN, PRIOR TO PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-342G, MMT-442AK. MMT-1884 WAS REQUESTED AND CUSTOMER RESPONSE WAS THE DEVICE WILL BE RETURNED. Manufacturer Narrative: THE PUMP PASSED THE FUNCTIONAL TESTS, INCLUDING THE SELF TEST, SLEEP CURRENT MEASUREMENT, ACTIVE CURRENT MEASUREMENT, REWIND TEST, PRIME/SEATING TEST, BASIC OCCLUSION TEST, OCCLUSION TEST, FORCE SENSOR TEST, DISPLACEMENT TEST AND DAT WITHIN SPECIFICATION AT .08785 INCHES. SUCCESSFULLY DOWNLOADED HISTORY FILES AND TRACES USING THUMP. SUCCESSFULLY UPLOADED PUMP TO CARELINK. ON THE SS EVENT DATE OF 13-JUN-2025, THERE IS NO UNEXPECTED ALARMS/SUSPENDS RECORDED IN THE FORMATTED HISTORY FILE. ON THE SAP NOTE EVENT DATE OF 15-JUN-2025, THERE IS AUTOSUSPEND (12) RECORDED IN THE FORMATTED HISTORY FILE ON (B)(6) 2025 05:01:00.000 AND (B)(6) 2025 05:11:00.000. UPON CHECKING THE PUMP DIAGNOSTIC TRACE FILE, AUTO SUSPEND ALARM WAS EXPECTED SINCE NO KEYS PRESSED IN THE TIME DURATION SET (12 HOURS) FOR AUTO SUSPEND. UNABLE TO VERIFY DAILY TOTAL OF BASAL/BOLUS AND ALL INSULIN DELIVERED ON THE SS EVENT DATE OF 13-JUN-2025 AND SAP NOTE EVENT DATE OF 15-JUN-2025 LISTED ON SMARTSOLVE DUE TO INSUFFICIENT DATA IN THE TRACE/HISTORY FILES. IN FURTHER REVIEW OF THE FORMATTED HISTORY FILE 1 WEEK PRIOR TO THE SS EVENT DATE OF 13-JUN-2025 AND SAP NOTE EVENT DATE OF 15-JUN-2025, THESE PUMP ERROR(S)/ALARM(S) WERE NOTED: LOSTSENSOR1ALERT (780) WAS FOUND ON: (B)(6) 2025 03:24:00.000. (B)(6) 2025 04:01:00.000, (B)(6) 2025 04:11:00.000. (B)(6) 2025 22:00:00.000. (B)(6) 2025 20:22:00.000, (B)(6) 2025 20:32:00.000. (B)(6) 2025 10:25:00.000, (B)(6) 2025 10:35:00.000. (B)(6) 2025 16:58:00.000, (B)(6) 2025 17:08:00.000. LOSTSENSOR2ALERT (781) WAS FOUND ON: (B)(6) 2025 10:51:00.000. (B)(6) 2025 10:51:00.000. (B)(6) 2025 18:55:00.000. THE PUMP WAS PROGRAMMED WITH A TEST GUARDIAN LINK (3) TRANSMITTER AND A GLUCOSE SENSOR SIMULATOR. THE PUMP CONNECTED SUCCESSFULLY TO THE TRANSMITTER AND DISPLAYED ¿TRANSMITTER CONNECTION SUCCESSFUL¿. THE PUMP COMMUNICATED PROPERLY WITH GLUCOSE SENSOR SIMULATOR AND DISPLAYED THE CALIBRATE YOUR SENSOR ALARM PROPERLY AFTER COMPLETION OF THE WARM UP. THE PUMP CALIBRATED AND DISPLAYED THE PROGRAMMED VALUE OF 240 MG/DL PROPERLY ON THE DISPLAY GRAPH. NO LOST SENSOR ALERT OR UNEXPECTED SENSOR ERRORS OR ANOMALIES WERE NOTED DURING TESTING. INSERT BATTERY ALARM WAS FOUND ON: (B)(6) 2025 19:27:46.000 (B)(6) 2025 12:14:40.000 (B)(6) 2025 07:52:35.000 (B)(6) 2025 08:03:00.000, (B)(6) 2025 08:27:18.000 (B)(6) 2025 08:37:00.000 LOW BATTERY ALERT WAS FOUND ON: (B)(6) 2025 09:44:00.000 REPLACE BATTERY ALERT WAS FOUND ON: (B)(6) 2025 09:44:00.000 (B)(6) 2025 19:25:00.000 PUMP ERROR 23 ALARM WAS FOUND ON: (B)(6) 2025 08:03:05.000 (B)(6) 2025 08:38:04.000 POWER LOSS ALARM WAS FOUND ON: (B)(6) 2025 08:03:21.000, (B)(6) 2025 08:03:29.000 (B)(6) 2025 08:38:14.000, (B)(6) 2025 08:38:22.000 POWER MANAGEMENT GRAPH WAS SUCCESSFULLY GENERATED. THE POWER MANAGEMENT TOOL CONFIRMED THE UNLOADED VOLTAGE (UL VLITH) AND LOADED VOLTAGE (LOADED VLITH) WERE WITHIN SPEC RANGE. INSERT BATTERY ALARM WAS EXPECTED SINCE THE BATTERY WAS REMOVED FROM THE PUMP. EARLIEST POWER DATA AVAILABLE PER THE POWER MANAGEMENT TOOL/DETAIL TRACE FILE IS ON (B)(6) 2025 AT 15:21:59.000. THERE WAS NO POWER DATA AVAILABLE FOR THE EVENT DATE OF 08-JUN-2025. UNABLE TO CHECK POWER DATA FOR LOW BATTERY ALERT AND REPLACE BATTERY ALERT. PUMP ERROR 23 ALARM AND POWER LOSS ALARM WERE EXPECTED SINCE THE BATTERY WAS REMOVED FOR MORE THAN 10 MINUTES. NO UNEXPECTED LOW BATTERY ALERT, REPLACE BATTERY ALERT, PUMP ERROR 23 ALARM AND POWER LOSS ALARM NOTED DURING TESTING. THE PUMP WAS CUT OPEN TO PERFORM VISUAL INSPECTION AND FOUND NO EVIDENCE OF PHYSICAL OR MOISTURE DAMAGE ON THE PCBA 1, PCBA 2, FORCE SENSOR, MOTOR AND VIBRATOR ASSEMBLY NOTED. FORCE SENSOR ZERO OFFSET WITHIN SPECIFICATION (23.32 MV). THE PUMP WAS RECEIVED WITHOUT A BATTERY. THE PUMP WAS RECEIVED WITH A BATTERY CAP. NO CRACKED/DAMAGED BATTERY CAP NOTED DURING VISUAL INSPECTION. THE MOLDING KNIT LINE DAMAGE (NEAR THE BELT CLIP PLATE) WAS ACCEPTABLE PER CASE ASSEMBLY COSMETIC SPECIFICATION, (B)(4). THE SC1 CAP LOCKS PROPERLY INTO THE RESERVOIR COMPARTMENT. THE FOLLOWING WERE NOTED DURING VISUAL INSPECTION: A SCRATCHED CASE. THE PUMP PASSED ALL THE REQUIRED TESTING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. THE DEVICE HAS BEEN RETURNED, BUT NOT YET EVALUATED. FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22491447&pc=OZP | 2032227-2025-211252 | 2024-12-15 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-14 04:00 | OZP | 670G INSULIN PUMP MMT-1780KL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia | P150001 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY AT THE HOSPITAL ON (B)(6) 2025. THE CUSTOMER WAS ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS HYPERGLYCEMIA, WAS HOSPITALIZED DATED ON (B)(6) 2024. THE CUSTOMER¿S BLOOD GLUCOSE WAS 405 MG/DL. THE EVENT INVOLVED PRODUCT(S) UNK_RESERVOIR, UNOMEDICAL, MMT-1780KL. TROUBLESHOOTING WAS PERFORMED. THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF DEATH. THE CUSTOMER WAS USING SENSORS. THE INSULIN PUMP WAS LAST WORN ON (B)(6) 2024, PRIOR TO PASSING. NO PRODUCT RETURN IS REQUIRED FOR UNK_RESERVOIR, UNOMEDICAL, MMT-1780KL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22492818&pc=OZD | 1220648-2025-30281 | 2025-06-17 04:00 | Death | ABIOMED, INC. | 2025-07-14 04:00 | OZD | IMPELLA CP | Malposition of Device | Ischemia; Tachycardia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED VENTRICULAR TACHYCARDIA. THE PATIENT WAS DEFIBRILLATED. THE PUMP WAS FOUND TO BE MISPOSITIONED SO IT WAS EXPLANTED. UPON EXPLANT, THE PATIENT LOST PULSES IN HER LEG FOR WHICH SHE UNDERWENT AN EMBOLECTOMY. THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED THE NEXT DAY. Manufacturer Narrative: H6 COMPONENT CODE 3038 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30281. Manufacturer Narrative: MEDICAL SAFETY OFFICER REVIEW: THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22493016&pc=OZD | 1220648-2025-30284 | 2025-06-25 04:00 | Death | ABIOMED, INC. | 2025-07-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND LOST DOPPLER DISTAL PULSES AND HIS TOES WERE MOTTLED. THE PATIENT'S FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22493127&pc=MIH | 2017233-2025-06410 | 2025-05-07 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-07-14 04:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P040043 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, PATIENT UNDERWENT AN UNKNOWN ENDOVASCULAR PROCEDURE UTILIZING A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM. REPORTEDLY, IT IS UNKNOWN, IF DURING OR AFTER PROCEDURE BUT PATIENT PASSED AWAY THE SAME DAY. FIELD SALES ASSOCIATE WAS NOT PRESENT FOR THIS PROCEDURE. ON (B)(6) 2025, PER MEDICAL DEVICE TRACKING SYSTEM THAT A DEVICE CARD WAS PROVIDED BY THE HOSPITAL WITH DEVICE INFORMATION AND NOTE OF PATIENT BEING DECEASED. Manufacturer Narrative: H6: D15: THERE IS INSUFFICIENT INFORMATION AVAILABLE FOR GORE TO REASONABLY DRAW CONCLUSIONS RELATED TO ASPECTS OF THE EVENT, THEREFORE CONCLUSION CODE ¿D15: CAUSE NOT ESTABLISHED¿ IS BEING USED. INSUFFICIENT INFORMATION MAY INCLUDE LIMITED OR MISSING RELEVANT MEDICAL RECORDS, INVOLVEMENT OF MULTIPLE IMPLANTED DEVICES (INCLUDING NON-GORE DEVICES) IN THE FIELD OF TREATMENT, PATIENT NON-COMPLIANCE, AND/OR A GENERAL LACK OF AVAILABLE DETAIL OR SPECIFICITY RELATED TO AN ADVERSE EVENT AND/OR DEVICE." H6: CODE C19 - A REVIEW OF THE MANUFACTURING RECORDS INDICATED THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE B20 -THE DEVICE REMAINS IMPLANTED AND WAS THEREFORE NOT AVAILABLE FOR ENGINEERING EVALUATION BY GORE. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22493373&pc=DSQ | 2916596-2025-04479 | 2025-06-22 04:00 | Death | THORATEC CORPORATION | 2025-07-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. NO PRODUCTS WERE EXCHANGED. THE DEVICE WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED MULTISYSTEM ORGAN FAILURE (MSOF) AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22493411&pc=OZD | 1220648-2025-30288 | 2025-06-25 04:00 | Death | ABIOMED, INC. | 2025-07-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT ON IMPELLA CP SUPPORT AND DEVELOPED HEMATOMA AND BLEEDING FROM THE RIGHT GROIN. THE IMPELLA ACCESS SITE WAS ON THE LEFT SIDE. THE RIGHT SIDE HAD PREVIOUS ILIAC STENTING. HEPARIN WAS WITHHELD AND BLOOD PRODUCTS WERE TRANSFUSED. THE PATIENT ALSO SHOWED SIGNS OF ISCHEMIA TO BILATERAL LOWER EXTREMITIES AND RIGHT HAND HAD RISING LACTATE. THE PATIENT WAS MADE COMFORT CARE AND LATER EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL, LIMB ISCHEMIA, AND ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ACCESS SITE BLEEDING: THE PATIENT IS HAVING BLEEDING FROM MULTIPLE SITES. THE ROOT CAUSE OF ACCESS SITE BLEEDING IS DETERMINED TO BE PATIENT CONDITION BECAUSE THE PATIENT IS BLEEDING FROM MULTIPLE SITES. LIMB ISCHEMIA: THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. VASCULAR DAMAGE - PERIPHERAL VESSEL: THE ROOT CAUSE OF THE VASCULAR DAMAGE WAS UNABLE TO BE DETERMINED AS INSUFFICIENT CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED FOR ANALYSIS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22493469&pc=MOZ | MW5172541 | 2025-06-26 04:00 | Death | FERRING PHARMACEUTICALS INC. | 2025-07-14 04:00 | MOZ | EUFLEXXA 20 MG/2 ML (6=3 SYR) | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: CAREGIVER REPORTED THAT THE PT PASSED AWAY ON (B)(6) 2025. FREQ: INJECT 1 SYRINGE INTRA-ARTICULARLY INTO RIGHT KNEE EVERY WEEK FOR 3 WEEKS. PRESCRIBER INFO: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22493536&pc=DQA | 2936999-2025-00374 | 2022-07-01 04:00 | Death | COVIDIEN LLC | 2025-07-14 04:00 | DQA | NELLCOR | Device Alarm System | Cardiac Arrest | K142865 |  | 1 | Event Description: IT WAS REPORTED THAT DURING USE, THE MONITOR FAILED TO ALARM, RESULTING IN THE PATIENT'S CARDIAC ARREST AND SUBSEQUENT DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22493542&pc=MOZ | MW5172542 |  | Death | ANIKA THERAPEUTICS, INC. | 2025-07-14 04:00 | MOZ | ORTHOVISC | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: THE PATIENT'S DATE OF DEATH (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22493594&pc=QZI | MW5172543 | 2025-07-03 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-14 04:00 | QZI | FARAWAVE PULSED FIELD ABLATION (PFA) SYSTEM / FARAPULSE GEN | Erratic or Intermittent Display | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Pericardial Effusion |  |  | 1 | Event Description: IT WAS REPORTED THAT THE CARTO 3 SYSTEM WAS DISPLAYING ERROR 1006, BACK PATCH SENSOR ERROR, INTERMITTENTLY AT FIRST, THEN CONSTANT. THE CALLER WAS ADVISED TO REPLACE THE BACK PATCH SENSOR CABLES AND TO REBOOT THE CARTO 3 SYSTEM. THE PIU WAS TURNED OFF WHEN THE ANESTHESIOLOGIST ADVISED THAT THE PATIENT WAS HYPOTENSIVE. THE PHYSICIAN OPTED TO FORGO TROUBLESHOOTING AT THIS TIME. CAS WILL CALL BACK. CAS CALLED BACK STATING THAT AFTER THE FIRST CALL, A STAT ECHO WAS ORDERED. THE PATIENT WAS IN TAMPONADE AND CPR WAS STARTED (12:45 CENTRAL TIME). A PERICARDIOCENTESIS WAS PERFORMED. UNKNOWN AMOUNT OF FLUID REMOVED. BLOOD AND PLASMA PRODUCTS WERE GIVEN ALONG WITH ACLS CODE PROTOCOL. CARDIOTHORACIC SURGERY WAS CALLED. THE PATIENT EXPIRED (14:10 CST) BEFORE SURGERY WAS AVAILABLE. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA OF A FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22493884&pc=DRA | 2135147-2025-03827 | 2025-06-19 04:00 | Death | ABBOTT MEDICAL | 2025-07-14 04:00 | DRA | MITRACLIP | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Perforation | K190167 |  | 1 | Event Description: IT WAS REPORTED THIS WAS A MITRACLIP PROCEDURE TO TREAT MITRAL REGURGITATION (MR) WITH A GRADE OF 4. AFTER A SUCCESSFUL CLIP DEPLOYMENT AND REMOVAL OF THE STEERABLE GUIDE CATHETER (SGC), THE PATIENT'S BLOOD PRESSURE DROPPED WITHIN MINUTES, LEADING TO CPR. A VEIN RUPTURE WAS SUSPECTED, AND EMERGENCY SURGICAL CONVERSION WAS PERFORMED, HOWEVER THE PATIENT DID NOT SURVIVE. THE PATIENT PASSED AWAY ON THE SAME DATE.Event Description: IT WAS REPORTED THIS WAS A MITRACLIP PROCEDURE TO TREAT MITRAL REGURGITATION (MR) WITH A GRADE OF 4. AFTER A SUCCESSFUL CLIP DEPLOYMENT AND REMOVAL OF THE STEERABLE GUIDE CATHETER (SGC), THE PATIENT'S BLOOD PRESSURE DROPPED WITHIN MINUTES, LEADING TO CPR. A VEIN RUPTURE WAS SUSPECTED, AND EMERGENCY SURGICAL CONVERSION WAS PERFORMED, HOWEVER THE PATIENT DID NOT SURVIVE. THE PATIENT PASSED AWAY ON THE SAME DATE. Manufacturer Narrative: THE DEVICE WILL NOT BE RETURNED FOR EVALUATION; THE DEVICE WAS REPORTEDLY DISCARDED. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR THE REPORTED PERFORATION, HYPOTENSION, AND DEATH. PERFORATION, HYPOTENSION, AND DEATH, LISTED IN THE INSTRUCTIONS FOR USE, ARE KNOWN POSSIBLE COMPLICATION ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED UNEXPECTED MEDICAL INTERVENTION AND SURGICAL INTERVENTION WERE RESULTS OF CASE SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22493900&pc=OZD | 1220648-2025-30292 | 2025-06-24 04:00 | Death | ABIOMED, INC. | 2025-07-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD AN ELDERLY PATIENT TRANSFERRED FROM COMMUNITY TO TERTIARY CENTER WITH ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK. THE IMPELLA CP PUMP WAS PLACED BUT THERE WAS VENTRICULAR TACHYCARDIA ON DAY AFTER IMPLANT TREATED BY DEFIBRILLATION. PUMP REMAINED ON FOR 3 DAYS AND THEN EXPIRED ON SUPPORT. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22494256&pc=BZD | 2518422-2025-107261 | 2025-06-21 04:00 | Death | RESPIRONICS, INC. | 2025-07-14 04:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 | 320500985 | 1 | Event Description: MANUFACTURER RECEIVED A CALL FROM THE PATIENT'S DAUGHTER IN LAW, REQUESTING REPLACEMENT DEVICE SENT TO HER AS HER MOTHER-IN-LAW HAS PASSED AWAY. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22494753&pc=DSQ | 2916596-2025-04259 | 2022-10-16 04:00 | Death | THORATEC CORPORATION | 2025-07-14 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis; Multiple Organ Failure; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIORGAN FAILURE (MOF) AND CHRONIC DRIVELINE INFECTION. THE PATIENT SUFFERED FROM A PROGRESSIVE MOF WHICH WAS RELATED TO THE OUTCOME. CHRONIC DRIVELINE INFECTION WAS ALSO NOTED BUT IT WAS NOT RELATED TO THE DEATH BUT WITH SEPSIS. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) OPERATED AS EXPECTED AND THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22494952&pc=NPT | 2135147-2025-03832 | 2025-06-16 04:00 | Death | ABBOTT MEDICAL | 2025-07-14 04:00 | NPT | NAVITOR VISION | Adverse Event Without Identified Device or Use Problem | Nausea; Ventricular Fibrillation; Heart Block | P190023 |  | 1 | Event Description: IT WAS REPORTED ON 16 JUNE 2025 A 35MM NAVITOR VISION VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM. THE LENGTH OF THE MEMBRANOUS SEPTUM WAS 1.4MM. A TEMPORARY PACEMAKER WAS INSERTED TO CONTROL PACING. THE VALVE WAS IMPLANTED. THE FINAL IMPLANT DEPTH WAS 5MM FROM THE NON-CORONARY CUSP (NCC). AT THE END OF THE PROCEDURE THE PATIENT HAD A LEFT BUNDLE BRANCH BLOCK (LBBB). THE TEMPORARY PACEMAKER WAS LEFT INDWELLING. THE FOLLOWING DAY ECHOCARDIOGRAM SHOWED A MILD PERIVALVULAR LEAK. AFTER THE TEMPORARY PACEMAKER WAS REMOVED THE PATIENT FELT NAUSEOUS, THEREFORE THE TEMPORARY PACEMAKER WAS LEFT IN, AND THE PATIENT WAS KEPT OVERNIGHT FOR MONITORING. ON (B)(6) 2025 THE PATIENT WENT INTO VENTRICULAR FIBRILLATION AND PASSED AWAY. THE CAUSE OF DEATH WAS VENTRICULAR FIBRILLATION. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT OF ARRHYTHMIA AND PATIENT DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED AND WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON ALL AVAILABLE INFORMATION, A CAUSE FOR THE REPORTED DEATH COULD NOT BE CONCLUSIVELY DETERMINED. IT IS POSSIBLE THAT THE EVENT IS RELATED TO THE PATIENT CONDITION PRIOR TO IMPLANT. HOWEVER, THIS CANNOT BE CONFIRMED. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS UNDER CASE-SPECIFIC CIRCUMSTANCES, AS TEMPORARY PACEMAKER WAS IMPLANTED FOR HEART BLOCK. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. FROM MEDICAL REVIEW: "BASED ON THE INFORMATION AVAILABLE, THERE IS NO INDICATION OF DEVICE MALFUNCTION. SHOULD ADDITIONAL CLINICAL DATA OR IMAGING BECOME AVAILABLE, THIS MEDICAL REVIEW WILL BE UPDATED ACCORDINGLY."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22494967&pc=OZD | 1220648-2025-30302 | 2025-06-24 04:00 | Death | ABIOMED, INC. | 2025-07-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED A IMPELLA CP FOR THE SUPPORT OF THE PATIENT ADMITTED IN WITH ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK. THE PATIENT HAD RIGHT FEMORAL ARTERIAL TRACT BLEED REQUIRING TRANSFUSIONS OVERNIGHT AND IN THE MORNING. FEM STOP WAS PLACED AND THE DOCTOR PLACED A STITCH BEDSIDE AND BLEEDING STOPPED. LATER IT WAS REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER. ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22494972&pc=MVK | 3008642652-2025-06780 | 2025-06-28 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-14 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Labelling, Instructions for Use or Training Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. PER CLINICAL REVIEW OF THE CONTINUOUS ECG RECORDING, THE DEVICE WAS STARTED UP AT 10:35:07 ON (B)(6) 2025. AN ARRHYTHMIA WAS DETECTED AT 05:41:08 ON (B)(6) 2025 WITH RESPONSE BUTTON USE. ECG SHOWS SINUS RHYTHM @ 60 BPM DEGRADING TO VF WITH VARYING AMPLITUDES. THE DETECTION STOPPED AT 05:41:34. THE DEVICE SHUT DOWN AT 05:41:44 ON (B)(6) 2025. THE DEVICE PROPERLY DETECTED VF. RESPONSE BUTTON USE PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. Manufacturer Narrative: THE MONITOR AND ELECTRODE BELT HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION. THE REVIEW OF THE DATA INDICATED THAT THE DEVICE POWERED ON NORMALLY AND WAS ABLE TO ACQUIRE THE PATIENT¿S ECG SIGNAL ON THE LAST DAY OF USE CAPTURED IN THE DATA DOWNLOAD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22495359&pc=DSQ | 2916596-2025-04666 | 2025-06-09 04:00 | Death | THORATEC CORPORATION | 2025-07-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Renal Failure; Respiratory Failure; Mitral Valve Insufficiency/ Regurgitation; Tricuspid Valve Insufficiency/ Regurgitation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT, STATUS POST HEARTMATE 3 PLACEMENT ON (B)(6) 2024, EXPERIENCED SEVERE RIGHT VENTRICULAR DYSFUNCTION AND SEVERE TRICUSPID REGURGITATION (TR), STATUS POST ST. JUDE CARDIAC RESYNCHRONIZATION THERAPY WITH DEFIBRILLATOR (CRT-D). THE PATIENT ALSO EXPERIENCED RECURRENT HEMOPTYSIS AND RIGHT LOWER LOBE COLLAPSE, PRIOR TRACHEOSTOMY STATUS POST DECANNULATION (B)(6) 2025, WITH SEVERE MITRAL REGURGITATION (MR) STATUS POST MECHANICAL MITRAL VALVE REGURGITATION (MVR) (ON WARFARIN), PAROXYSMAL ATRIAL FIBRILLATION/ATRIAL FLUTTER (PAF/AFL) STATUS ABLATION IN 2016, WITH HISTORY OF RENAL FAILURE OFF HEMODIALYSIS. THE PATIENT PRESENTED ON (B)(6) 2025 WITH WORSENING DIFFICULTY BREATHING FOR ONE-WEEK, ACUTE HYPOXIC RESPIRATORY FAILURE AND POSSIBLE FOOD IMPACTION. GASTROINTESTINAL (GI) CONSULTED, ESOPHAGRAM AND ACUTE ABDOMINAL SERIES WERE UNREMARKABLE, WITH PLANS FOR OUTPATIENT MOTILITY STUDIES. OVERNIGHT INTO (B)(6) 2025, THE PATIENT WAS NOTED TO BE MORE SOMNOLENT WITH VENOUS BLOOD GAS (VBG) OF 7.20/70/40 AND CHEST X-RAY WITH RIGHT LUNG BASE CONSOLIDATION CONCERNING FOR ASPIRATION PNEUMONIA. VANCOMYCIN/ZOSYN WAS STARTED. COMPUTED TOMOGRAPHY (CT) HEAD SHOWED INTERVAL DEVELOPMENT OF SUBARACHNOID HEMORRHAGE ON LEFT SIDE. NEUROSURGERY WAS CONSULTED. TOMOGRAPHY ANGIOGRAPHY (CTA) OF THE HEAD AND NECK SHOWED NO VASCULAR STENOSIS OR ANEURYSM. ON (B)(6) 2025, REPEAT CT HEAD SHOWED INTERVAL INCREASED BURDEN OF SUBARACHNOID HEMORRHAGE (SAH) IN LEFT HEMISPHERE AND NEW SAH IN RIGHT ANTERIOR SYLVIAN FISSURE; THE PATIENT WAS GIVEN VITAMIN K TO CORRECT RISING INTERNATIONAL NORMALIZED RATIO (INR). ON (B)(6) 2025, NEUROLOGY WAS CONSULTED WITH ELECTROENCEPHALOGRAM MONITORING DUE TO CONCERN FOR CHANGE IN MENTAL STATUS. EEG NON-CONTRIBUTORY WAS STOPPED. ON (B)(6) 2025, PATIENT APPEARED TO WAKE UP MORE, WAS MAKING EYE CONTACT AND INTERMITTENTLY FOLLOWING COMMANDS, ALTHOUGH MENTAL STATUS WAS WAXING AND WANING. (B)(6) 2025, THEY WERE OBSERVED TO HAVE DILATED, FIXED LEFT PUPIL AT 06:00 PROMPTING REPEAT CT HEAD THAT SHOWED WORSENING INTRACRANIAL HEMORRHAGE WITH MIDLINE SHIFT AND UNCAL HERNIATION. NEUROSURGERY WAS REVIEWED AND FELT THAT THIS WAS A NON-SURVIVABLE NEUROLOGIC INJURY WITHOUT CHANCE OF MEANINGFUL RECOVERY. GOALS OF CARE DISCUSSION HAD WITH FAMILY, WITH PLANS TO ULTIMATELY TRANSITION TO COMFORT CARE ONCE FAMILY HAD A CHANCE TO VISIT. INTO (B)(6) 2025, BLOOD PRESSURE WAS GRADUALLY DECLINING. LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS TURNED OFF AT APPROXIMATELY 3:30. THE PATIENT PASSED AWAY AT 3:36 PM ON (B)(6) 2025.Event Description: REGARDING TO STROKE, THE PATIENT EXPERIENCED ALTERED MENTAL STATUS. THE MITRAL VALVE REGURGITATION/TRICUSPID REGURGITATION EXISTED PRE-LVAD IMPLANT. IT WAS UNKNOWN WHETHER THERE WERE CHANGES TO PATIENT CONDITION OR ANTICOAGULATION STATUS THAT MAY HAVE CONTRIBUTED TO THE SUBARACHNOID HEMORRHAGE. THE DEATH WAS NOT DIRECTLY CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22495450&pc=DSQ | 2916596-2025-04341 | 2025-06-23 04:00 | Death | THORATEC CORPORATION | 2025-07-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ON MULTIPLE VASOPRESSORS AND HAD ELEVATED CENTRAL VENOUS PRESSURE (CVP). THE PATIENT REQUIRED MECHANICAL CIRCULATORY SUPPORT (MCS) VIA EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). RESUSCITATION WAS ATTEMPTED. THERE WAS SUSPECTED RIGHT VENTRICULAR (RV) FAILURE. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PUMP WAS NOT EXPLANTED. AN AUTOPSY WOULD BE PERFORMED. THE PUMP WAS DEACTIVATED AND WOULD NOT BE RETURNED. THE OUTCOME WAS CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE PARAMETERS WERE WITHIN NORMAL LIMITS FOR THE PATIENT. LOG FILES WERE SENT FOR REVIEW AFTER THE PATIENT PASSED AWAY. THE EVENT LOG CAPTURED INTERMITTENT LOW FLOW EVENTS ON (B)(6) 2025 AT 14:58 THROUGH 23:38 WITH FLOW NOTED AS LOW AS 1.4 LITERS PER MINUTE (LPM). THESE LOWER FLOWS APPEARED TO BE PATIENT RELATED. AFTER THIS TIME THE FIXED SPEED WAS DROPPED TO 4000 REVOLUTIONS PER MINUTE (RPM) THAT CAUSED A LOW SPEED EVENT AND ADDITIONAL LOW FLOW EVENTS. THE DRIVELINE WAS DISCONNECTED ON (B)(6) 2025 AT 01:27. IT WAS NOTED THAT THE PATIENT DID NOT HAVE LEFT VENTRICULAR ASSIST DEVICE (LVAD) ISSUES FROM IMPLANT UNTIL SUNDAY AT 6PM. THE PATIENT'S FLOW WAS CONSISTENTLY AROUND 4-4.5 LPM. THEY STARTED TO GO DOWN THE HILL SUBSEQUENTLY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: ANALYSIS OF THE SUBMITTED LOG FILES CONFIRMED THE REPORTED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILES CONTAINED EVENTS FROM (B)(6) 2025. LOW FLOW HAZARD ALARMS WERE CAPTURED ON (B)(6) 2025, AND THE ESTIMATED FLOW WAS CAPTURED DECREASING TO 0 LITERS PER MINUTE (LPM). NO OTHER EVENTS OR ALARMS WERE CAPTURED, AND THE DEVICE APPEARED TO OPERATE AS INTENDED AT THE FIXED SPEED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THIS SECTION ALSO ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 6 OF THE IFU, "PATIENT CARE AND MANAGEMENT" (UNDER "RIGHT HEART FAILURE"), DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22495640&pc=NPT | 2015691-2025-05683 | 2025-06-19 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-14 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIVESCIENCE'S AFFILIATE IN FRANCE, REGARDING A 26MM SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IMPLANT IN AORTIC POSITION BY TRANSFEMORAL APPROACH, DURING THE PROCEDURE, AFTER VALVE IMPLANT, AN ANNULAR RUPTURE OCCURRED DUE TO SEVERE CALCIFICATION. A SECOND VALVE WAS IMPLANTED IN ORDER TO SEAL THE BREACH FOLLOWED BY A PERICARDIAL DRAINAGE. AT THE END, PATIENT PASSED AWAY DUE TO CARDIAC TAMPONADE FOLLOWED BY CARDIAC ARREST. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PROCEDURAL FACTORS (INTERACTION WITH CALCIUM) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22496131&pc=DSQ | 2916596-2025-04708 | 2025-02-05 05:00 | Death | THORATEC CORPORATION | 2025-07-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Liver Damage/Dysfunction; Tachycardia; Ventricular Fibrillation; Heart Failure/Congestive Heart Failure; Peripheral Edema | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO HEART FAILURE.Event Description: ON (B)(6) 2025, THE PATIENT EXPERIENCED RIGHT HEART FAILURE, WITH PERIPHERAL EDEMA. THIS EVENT WAS NOT DEVICE RELATED. CARDIAC CATHETERIZATION WAS PERFORMED. THEIR CENTRAL VENOUS PRESSURE (CVP) WAS MMHG. PULMONARY ARTERY (PA) SYSTOLIC PRESSURE WAS11 MMHG. PULMONARY CAPILLARY WEDGE PRESSURE (PCW) WAS 12 MMHG. THEY WERE ADMITTED FROM ON (B)(6) 2025.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT WAS READMITTED ON (B)(6) 2025. THEIR CENTRAL VENOUS PRESSURE (CVP) WAS GREATER THAN 18MMHG WITH ASCITES AND PERIPHERAL EDEMA. THE PATIENT ALSO PRESENTED WITH SIGNS OF LIVER DYSFUNCTION. MEDICATION WAS ADJUSTED AS TREATMENT. TOTAL BILIRUBIN WAS 6.1 MG/DL. GLUTAMIC OXALOACETIC TRANSAMINASE/ASPARTATE AMINOTRANSFERASE. (GOT/AST) WAS 36 U/L AND GLUTAMIC OXALOACETIC TRANSAMINASE/ALANINE AMINOTRANSFERASE (GOT/ALT) WAS 19 U/L. THE PATIENT REMAINED ADMITTED FOR RIGHT HEART FAILURE AND LIVER DYSFUNCTION UNTIL (B)(6) 2025 WHEN THEY PASSED AWAY. THE CAUSE OF DEATH WAS CONSIDERED TO BE DETERIORATION OF GENERAL CONDITION DUE TO RIGHT HEART FAILURE AND LIVER FAILURE AFTER RIGHT VENTRICULAR ASSIST DEVICE RVAD WITHDRAWAL. OTHER CAUSES INCLUDED VENTRICULAR TACHYCARDIA OR VENTRICULAR FIBRILLATION. Manufacturer Narrative: SECTION A: DATE OF BIRTH CORRECTED. SECTION E1: CUSTOMER SITE WAS (B)(6). MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. A IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, HEPATIC DYSFUNCTION, CARDIAC ARRHYTHMIA, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 OF THE IFU (UNDER ¿RIGHT HEART FAILURE¿) STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THIS SECTION (UNDER ¿CAUTION!¿) ALSO EXPLAINS THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION D4: MODEL AND CATALOG NUMBERS CORRECTED NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22496413&pc=NIK | MW5172586 | 2025-04-06 04:00 | Death | ABBOTT MEDICAL | 2025-07-14 04:00 | NIK | ABBOTT CRT-D DEVICE | Failure to Deliver Shock/Stimulation; Device Sensing Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: PATIENT WAS TRANSCUTANEOUSLY PACED DUE TO POOR DEVICE FUNCTION. DEVICE WAS INTERROGATED BY ABBOTT REP AND DEVICE WAS NOT RECOGNIZING VFIB/VTACH AND SUBSEQUENTLY NOT DEFIBRILLATING THE PATIENT. CARDIOLOGY WAS AT BEDSIDE AS WELL AND TRANSVENOUS PACING WAS DISCUSSED. RIJ (RIGHT INTERNAL JUGULAR) CENTRAL LINE INTRODUCER (FOR PLANNED TRANSVENOUS PACING), RIGHT FEMORAL CVC (CENTRAL VENOUS CATHETER), AND LEFT FEMORAL ARTERIAL LINE WERE PLACED AND PATIENT CODED AGAIN SHORTLY THEREAFTER. ACLS WAS INITIATED AGAIN AND ROSC (RETURN OF SPONTANEOUS CIRCULATION) WAS OBTAINED AFTER DEFIBRILLATION. THE PATIENT'S FAMILY WAS BROUGHT TO BEDSIDE, AND THEY DECIDED TO OPT FOR DNR (DO NOT RESUSCITATE). PATIENT GRADUALLY BECAME HYPOTENSIVE (DESPITE MAXIMUM DOSES OF 3 PRESSORS AND IVF (INTRAVENOUS FLUID) BOLUSES) AND HE SUBSEQUENTLY PASSED AT 1646. THIS PATIENT IS A (B)(6) MALE, HISTORY OF CORONARY ARTERY DISEASE, PRIOR CORONARY BYPASS SURGERY. PRIOR STENT PLACEMENT, LAST PERFORMED 2023. HISTORY OF VALVE IN VALVE TAVR (TRANSCATHETER AORTIC VALVE REPLACEMENT). HISTORY OF PAROXYSMAL ATRIAL FIBRILLATION, PRIOR ELECTRICAL CARDIOVERSION LAST PERFORMED (B)(6) 2024. PRESENTLY IN SINUS RHYTHM. HISTORY OF BIV ICD (BIVENTRICULAR IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR) IMPLANT, ABBOTT CRT-D (CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR) DEVICE LAST INTERROGATED (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22496474&pc=DSQ | 2916596-2025-04533 | 2025-06-27 04:00 | Death | THORATEC CORPORATION | 2025-07-14 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Septic Shock; Multiple Organ Failure; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: THE PATIENT PASSED AWAY DUE TO MULTIPLE ORGAN FAILURE RELATED TO RIGHT VENTRICULAR FAILURE AND SEPTIC SHOCK. THE DEATH WAS NOT CONSIDERED DEVICE RELATED. THERE WERE NO PROBLEMS WITH THE DEVICE OR THE SURGICAL PROCEDURE. HOWEVER, THE COMPLICATION OF RIGHT VENTRICULAR FAILURE IMPACTED THE DEATH. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22496499&pc=DSQ | 2916596-2025-04664 | 2025-06-20 04:00 | Death | THORATEC CORPORATION | 2025-07-14 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY, DUE TO UNKNOWN CAUSE. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22496564&pc=FKX | 0008030665-2025-01583 | 2025-06-28 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-07-14 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K222318 |  | 1 | Event Description: IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT WAS RECEIVING DIALYSIS AND EXPIRED WHILE IN TREATMENT. THE NURSE WAS ADVISED TO DISCONTINUE THE USE OF THE CYCLER. A REPLACEMENT CYCLER WAS ISSUED. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DISCHARGE SUMMARY, HOSPITAL RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, TIMELINE OF EVENTS, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL.Event Description: IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT WAS RECEIVING DIALYSIS AND EXPIRED WHILE IN TREATMENT. THE NURSE WAS ADVISED TO DISCONTINUE THE USE OF THE CYCLER. A REPLACEMENT CYCLER WAS ISSUED. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DISCHARGE SUMMARY, HOSPITAL RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, TIMELINE OF EVENTS, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER WITH NO PHYSICAL DAMAGE NOTED. ALTHOUGH THERE WAS EVIDENCE OF DRIED FLUID PRESENT WITHIN THE CASSETTE COMPARTMENT, THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED THE CASSETTE MEMBRANE. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THE CYCLER UNDERWENT AND PASSED A SYSTEM AIR LEAK TEST, VALVE ACTUATION TEST, AND TEACH PUMPS TEST. THERE WERE NO VISUAL DISCREPANCIES ENCOUNTERED DURING THE INTERNAL INSPECTION. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED. Manufacturer Narrative: CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENT OCCURRED. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. IT SHOULD BE NOTED THAT LIMITED CLINICAL FOLLOW-UP PRECLUDED A MORE COMPREHENSIVE ASSESSMENT/INVESTIGATION. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CONTRIBUTORY ROLE IN THE PATIENT¿S EXPIRATION. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION OCCURRED. HOWEVER, GIVEN THE UNKNOWN CAUSE OF DEATH, PAST MEDICAL HISTORY, SUMMARY/TIMELINE OF EVENTS, DISCHARGE SUMMARY, AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT EXCLUDE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22499134&pc=OZD | 1220648-2025-30319 | 2025-06-26 04:00 | Death | ABIOMED, INC. | 2025-07-14 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Sepsis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT BECAME SEPTIC DURING IMPELLA 5.5 SUPPORT. THE ORIGIN OF THE SEPSIS COULD NOT BE DETERMINED. IT WAS ADDITIONALLY REPORTED THAT THE PATIENT EXPERIENCED A GASTROINTESTINAL BLEED THAT WAS BELIEVED TO BE UNRELATED TO THE PUMP. SUPPORT WAS CONTINUED UNINTERRUPTED. AFTER THIRTEEN DAYS OF SUPPORT, THE PATIENT'S UNDERLYING CONDITION DID NOT IMPROVE AND THEREFORE CARE WAS WITHDRAWN. THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS . ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22499457&pc=DSQ | 2916596-2025-04406 | 2025-06-08 04:00 | Death | THORATEC CORPORATION | 2025-07-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ischemia; Obstruction/Occlusion | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE VENTRICULAR ASSIST DEVICE (VAD) WAS DEACTIVATED ON COMFORT CARE DUE TO MESENTERIC ISCHEMIA. THE CAUSE OF MESENTERIC ISCHEMIA WAS OCCLUSION OF CELIAC AND THE SUPERIOR MESENTERIC ARTERY (SMA). THE PATIENT EXPERIENCED SYMPTOMS OF ALTERED MENTAL STATUS (AMS). THE EVENT WAS TREATED MEDICALLY, BUT THE PATIENT WAS NOT A SURGICAL CANDIDATE. THE DEVICE OPERATED AS EXPECTED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING THROMBOEMBOLISM, NEUROLOGICAL EVENTS (NOT STROKE RELATED), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿, LISTS THROMBOEMBOLISM AND NEUROLOGICAL DYSFUNCTION AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22499519&pc=QFG | 3019004087-2025-00397 | 2025-02-23 05:00 | Death | BETA BIONICS, INC | 2025-07-14 04:00 | QFG | ILET BIONIC PANCREAS | Adverse Event Without Identified Device or Use Problem | Hypoglycemia | K231485 |  | 1 | Event Description: ON 14-APR-2025, BETA BIONICS WAS CONTACTED BY (B)(6) REGARDING THE DEATH OF A 56-YEAR-OLD MALE WHO HAD BEEN USING THE ILET. THE MEDICAL EXAMINER REQUESTED DEVICE EVALUATION, AS THE USER WAS WEARING THE DEVICE AT THE TIME OF PASSING. THE ILET WAS SUBSEQUENTLY RETURNED TO BETA BIONICS FOR INVESTIGATION. Manufacturer Narrative: NO DEVICE MALFUNCTIONS WERE IDENTIFIED; THE ILET FUNCTIONED WITHIN SPECIFICATION. INSULIN DELIVERY AND SUSPENSION OCCURRED AS EXPECTED IN RESPONSE TO CGM VALUES AND MEAL ANNOUNCEMENTS. PREDICTIVE AND THRESHOLD-BASED HYPOGLYCEMIA ALARMS WERE APPROPRIATELY TRIGGERED BUT NOT ACKNOWLEDGED. DESPITE AUTOMATIC SUSPENSION OF INSULIN DELIVERY IN RESPONSE TO FALLING GLUCOSE LEVELS, SUSTAINED HYPOGLYCEMIA (CGM <54MG/DL) WAS OBSERVED BEGINNING AT APPROXIMATELY 00:19 ON (B)(6) 2025. A CGM SENSOR ISSUE WAS SUBSEQUENTLY REPORTED, FOLLOWED BY SENSOR FAILURE AND LOSS OF GLUCOSE DATA. BASED ON ALL AVAILABLE INFORMATION, THE ILET SYSTEM OPERATED AS INTENDED. THERE IS NO EVIDENCE TO SUGGEST THAT THE ILET CAUSED OR CONTRIBUTED TO THE USER'S DEATH. THE DOCUMENTED CAUSE OF DEATH WAS HYPERTENSIVE AND ATHEROSCLEROTIC CORONARY HEART DISEASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22500148&pc=NPS | 2135147-2025-03837 | 2025-06-19 04:00 | Death | ABBOTT MEDICAL | 2025-07-14 04:00 | NPS | TRICLIP G4 SYSTEM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Heart Failure/Congestive Heart Failure | P230007 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A TRICLIP PROCEDURE WAS PERFORMED TO TREAT TRICUSPID REGURGITATION (TR) WITH A GRADE OF 5+. TWO TRICLIP XTW CLIPS WERE DEPLOYED ON THE TRICUSPID VALVE, REDUCING TR TO A GRADE OF 1+. THE CLIP WAS NOTED TO BE STABLE ON THE LEAFLETS. IT WAS NOTED THAT THE PATIENT LOOKED GOOD WITH GOOD TR REDUCTION. HOWEVER, A FEW HOURS AFTER THE PROCEDURE, WHILE IN RECOVERY, THE PATIENT¿S SYSTOLIC BLOOD PRESSURE DROPPED FROM 130 TO 40. THE PATIENT WAS THEN TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). ON (B)(6) 2025, THE PATIENT DIED. IN THE PHYSICIAN'S OPINION, THE TRICLIP DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT'S DEATH. ONE WEEK PRIOR TO THE TRICLIP PROCEDURE, THE PATIENT HAD A PLEURAL EFFUSION, ASCITES, AND HAD BEEN HOSPITALIZED. THE PATIENT UNDERGOING ANESTHESIA LIKELY CONTRIBUTED TO THE PATIENT DEATH. IT WAS STATED THAT THE CAUSE OF DEATH PER ATTENDING INTENSIVIST: MYXEDEMA COMA, ADRENAL INSUFFICIENCY, ACUTE BLOOD LOSS ANEMIA, DISSEMINATED INTRAVASCULAR COAGULATION (DIC), ACUTE RESPIRATORY DISTRESS SYNDROME (ARDS), ACUTE PULMONARY EDEMA, NON TRAUMATIC ACUTE KIDNEY INJURY (AKI) FROM ACUTE TUBULAR NECROSIS (ATN), LACTIC ACIDOSIS, HYPOGLYCMIA ON A BACKGROUND OF CHRONIC DIASTOLIC HEART FAILURE, CHRONIC KIDNEY DISEASE (CKD) IV, CIRRHOSIS, AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD). Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR THE REPORTED HEART FAILURE AND HYPOTENSION. HEART FAILURE/CONGESTIVE HEART FAILURE, DEATH, AND HYPOTENSION ARE LISTED IN THE INSTRUCTIONS FOR USE IS A KNOWN POSSIBLE COMPLICATION ASSOCIATED WITH TRICLIP PROCEDURES. THE REPORTED DEATH APPEARS TO BE RELATED TO PATIENT/PROCEDURE CONDITIONS. THE REPORTED HOSPITALIZATION WAS A RESULT OF CASE SPECIFIC CIRCUMSTANCE. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22500205&pc=DXY | 2182208-2025-03393 | 2024-02-07 05:00 | Death | MEDTRONIC, INC. | 2025-07-14 04:00 | DXY | MEDTRONIC LEAD | Appropriate Term/Code Not Available | Unspecified Infection; Renal Failure; Liver Failure |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CARDIAC RESYNCHRONIZATION THERAPY RESPONSE IN HEART FAILURE PATIENTS. THE AUTHORS DESCRIBED PATIENT DEATHS; CAUSES OF DEATH INCLUDED INFECTIONS, POST-OPERATIVE COMPLICATIONS, CHRONIC KIDNEY FAILURE, LIVER FAILURE, AND OTHER UNKNOWN CAUSES DESCRIBED AS CARDIOVASCULAR OR ALL-CAUSE MORTALITY. PATIENTS WERE IMPLANTED WITH TRANSVENOUS LEADS UNLESS IT WAS NOT FEASIBLE, THEN THEY WERE IMPLANTED WITH AN EPICARDIAL LEAD. EPICARDIAL LEAD PLACEMENT OCCURRED DUE TO TRANSVENOUS LEAD DISLODGMENT, DEVICE ENDOCARDITIS WITH LEAD EXTRACTION, HIGH THRESHOLDS, SUBCLAVIAN VEIN OCCLUSION/STENOSIS, OR CORONARY SINUS DISSECTION. THERE WERE PATIENTS WHO EXPERIENCED PNEUMOTHORAX POST-OPERATIVELY AND OTHER COMPLICATIONS THAT OCCURRED IN THE FOLLOW UP PERIOD WHICH INCLUDED TWIDDLER'S SYNDROME, HEART FAILURE HOSPITALIZATIONS, DEVICE ENDOCARDITIS, AND LEAD FRACTURE. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/68 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: ASSESSING CARDIAC RESYNCHRONIZATION THERAPY RESPONSE IN HEART FAILURE PATIENTS: A COMPARATIVE ANALYSIS OF EFFICACY AND OUTCOMES BETWEEN TRANSVENOUS AND EPICARDIAL LEADS. MONALDI ARCHIVES FOR CHEST DISEASE. 2025; 95:2845. DOI: 10.4081/MONALDI.2024.2845. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22500207&pc=NIK | 2182208-2025-03392 | 2024-02-07 05:00 | Death | MEDTRONIC, INC. | 2025-07-14 04:00 | NIK | MEDTRONIC CRT-D | Appropriate Term/Code Not Available | Unspecified Infection; Renal Failure; Liver Failure |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CARDIAC RESYNCHRONIZATION THERAPY RESPONSE IN HEART FAILURE PATIENTS. THE AUTHORS DESCRIBED PATIENT DEATHS; CAUSES OF DEATH INCLUDED INFECTIONS, POST-OPERATIVE COMPLICATIONS, CHRONIC KIDNEY FAILURE, LIVER FAILURE, AND OTHER UNKNOWN CAUSES DESCRIBED AS CARDIOVASCULAR OR ALL-CAUSE MORTALITY. PATIENTS WERE IMPLANTED WITH TRANSVENOUS LEADS UNLESS IT WAS NOT FEASIBLE, THEN THEY WERE IMPLANTED WITH AN EPICARDIAL LEAD. EPICARDIAL LEAD PLACEMENT OCCURRED DUE TO TRANSVENOUS LEAD DISLODGMENT, DEVICE ENDOCARDITIS WITH LEAD EXTRACTION, HIGH THRESHOLDS, SUBCLAVIAN VEIN OCCLUSION/STENOSIS, OR CORONARY SINUS DISSECTION. THERE WERE PATIENTS WHO EXPERIENCED PNEUMOTHORAX POST-OPERATIVELY AND OTHER COMPLICATIONS THAT OCCURRED IN THE FOLLOW UP PERIOD WHICH INCLUDED TWIDDLER'S SYNDROME, HEART FAILURE HOSPITALIZATIONS, DEVICE ENDOCARDITIS, AND LEAD FRACTURE. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/68 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: ASSESSING CARDIAC RESYNCHRONIZATION THERAPY RESPONSE IN HEART FAILURE PATIENTS: A COMPARATIVE ANALYSIS OF EFFICACY AND OUTCOMES BETWEEN TRANSVENOUS AND EPICARDIAL LEADS. MONALDI ARCHIVES FOR CHEST DISEASE. 2025; 95:2845. DOI: 10.4081/MONALDI.2024.2845. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22500282&pc=OZD | 1220648-2025-30320 | 2025-06-23 04:00 | Death | ABIOMED, INC. | 2025-07-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Ventricular Fibrillation; Asystole | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ON IMPELLA CP SUPPORT EXPERIENCED VENTRICULAR TACHYCARDIA WHICH PROGRESSED TO VENTRICULAR FIBRILLATION AND ASYSTOLE. THE PATIENT HAD SPONTANEOUS RETURN OF SPONTANEOUS CIRCULATION. THE PATIENT WAS THEN PUT ON HIGHER VASOACTIVE SUPPORT. HOWEVER, AFTER 6 DAYS OF SUPPORT THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED ARRHYTHMIA HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE ARRYTHMIA COULD NOT BE DETERMINED. Manufacturer Narrative: A4 AND A5 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST SYSTEM LISTS VENTRICULAR ARRYTHMIA AS AN ADVERSE EVENT THAT OCCURRED WHILE IN USE WITH THE IMPELLA DEVICE. IT STATES THAT "ONLY A FRACTION OF THESE RATES WERE ATTRIBUTED TO THE IMPELLA DEVICE AND THE EVENTS RESOLVED WITHOUT RESIDUAL EFFECT IN MOST OF THE CASES, UNLESS THE EVENT OF DEATH OCCURRED. OVERALL, THE RESULTS DID NOT SHOW ANY EVIDENCE OF INCREASED MORBIDITY ASSOCIATED WITH THE IMPELLA SUPPORT IN CARDIOMYOPATHY, MYOCARDITIS, AND PERIPARTUM CARDIOMYOPATHY (PPCM) PATIENTS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22500637&pc=NHL | 3006630150-2025-05477 | 2025-03-28 04:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-07-14 04:00 | NHL | VERCISE GENUS? | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P150031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ONE WEEK FOLLOWING THEIR DEEP BRAIN STIMULATION (DBS) IMPLANT PROCEDURE. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22500714&pc=NKM | 2135147-2025-03846 | 2008-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-14 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "PROGNOSTIC VALUE OF NT-PROBNP IN PATIENTS WITH PRIMARY MITRAL REGURGITATION UNDERGOING TRANSCATHETER EDGE-TO-EDGE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE ASSOCIATION BETWEEN NT-PROBNP AND OUTCOMES AND EXPLORED ITS ADDITIVE VALUE TO THE MITRAL REGURGITATION INTERNATIONAL DATABASE (MIDA) SCORE. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT NT-PROBNP, BUT NOT THE MIDA SCORE, WAS INDEPENDENTLY ASSOCIATED WITH DEATH OR HEART FAILURE HOSPITALIZATIONS WITHIN 3 YEARS IN M-TEER-TREATED PMR PATIENTS. INCORPORATING NT-PROBNP LEVELS INTO CLINICAL ASSESSMENT MAY IMPROVE RISK STRATIFICATION AND POTENTIALLY SUPPORTS EARLIER INTERVENTION AT LOWER NT-PROBNP LEVELS TO OPTIMIZE OUTCOMES. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS CHRISTOS ILIADIS AT UNIVERSITY HOSPITAL COLOGNE INSTITUTION WITH CORRESPONDING EMAIL CHRISTOS.ILIADIS@UK-KOELN.DE]. THE TIME FRAME OF THE STUDY WAS JANUARY 2008 TO DECEMBER 2022. A TOTAL OF 1382 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 1382 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, CORONARY ARTERY DISEASE, ATRIAL FIBRILLATION, CHRONIC LUNG DISEASE, HISTORY OF STROKE, MYOCARDIAL INFARCTION, AND CARDIAC SURGERY. (B)(6), UNK MITRACLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, STROKE, CARDIOGENIC SHOCK, BLEEDING, ACUTE KIDNEY INJURY, EMBOLISM, EXPULSION, SURGICAL INTERVENTION, SINGLE LEAFLET DEVICE ATTACHMENT. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22500731&pc=DSQ | 2916596-2025-04552 | 2025-04-29 04:00 | Death | THORATEC CORPORATION | 2025-07-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. ADDITIONAL INFORMATION WAS NO PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22500792&pc=NPT | 2025587-2025-04988 | 2025-07-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-14 04:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Positioning; Perivalvular Leak; Patient Device Interaction Problem | Non specific EKG/ECG Changes; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE, THE PATIENT HAD TORTUOUS FEMORAL ARTERIES AND PROTRUDING CALCIUM IN THE ABDOMINAL AORTA AND AORTIC ARCH. A PRE-DILATION WAS PERFORMED WITH A 26 MILLIMETER (MM) NON-MEDTRONIC BALLOON. THE DELIVERY CATHETER SYSTEM (DCS) WAS UNABLE TO TRAVERSE THE AORTIC ARCH AND MET RESISTANCE PROXIMAL OF THE LEFT SUBCLAVIAN ARTERY DESPITE BEING POSITIONED IN DIFFERENT ANGLES. THE GUIDEWIRE WAS SWITCHED. THE DCS WAS STUCK AGAIN AT THE POSITION DESPITE BEING POSITIONED AGAIN IN DIFFERENT ANGLES. FLUOROSCOPY DID NOT REVEAL AN INJURY TO THE AORTIC ARCH. THE NON-MEDTRONIC SHEATH WAS EXCHANGED FOR A 22 FR 65CM LENGTH SHEATH. THE DCS WAS UNABLE TO CROSS THE AORTIC ARCH AGAIN DESPITE THE DIFFERENT POSITIONING ANGLES. THE DCS WAS REMOVED AND FLUSHED. THE DCS WAS INSERTED THROUGH THE NON-MEDTRONIC SHEATH AND WAS UNABLE TO CROSS THE AORTIC ARCH AGAIN DESPITE THE DIFFERENT POSITIONING ANGLES. A NEW SYSTEM WAS PREPPED AND THE VALVE WAS ABLE TO CROSS USING A SNARE. THE PATIENT WAS HYPOTENSIVE WITH SEVERE AORTIC INSUFFICIENCY AND WAS RESPONSIVE TO MEDICAL MANAGEMENT. THE VALVE WAS DEPLOYED INITIALLY TOO SHALLOW AND WAS RECAPTURED. A SECOND AND FINAL DEPLOYMENT ATTEMPT WAS AT 4 MM ON THE NON-CORONARY CUSP (NCC) AND 6 MM ON THE LEFT CORONARY CUSP (LCC). TRACE-MILDPARAVALVULAR LEAK (PVL) WAS REPORTED. AN AORTOGRAM NOTED NO INJURY TO THE AORTIC ARCH OR ABDOMINAL AORTA. THE SHEATH TO THE LEFT FEM ORAL ARTERY PARTIALLY DISLODGED CAUSING A HEMATOMA WHICH WAS ALLEVIATED WITH MANUAL COMPRESSION. TRANSIENT LEFT BUNDLE BRANCH BLOCK (LBBB) WAS NOTED. NO VASCULAR ISSUES TO THE RIGHT FEMORAL ARTERY WAS NOTED. A POST IMPLANT COMPUTED TOMOGRAPHY (CT) SHOWED NO DISSECTION, TEAR, OR INJURY TO THE AORTA. THE PATIENT BECAME UNSTABLE FOLLOWING THE PROCEDURE AND DIED THE FOLLOWING DAY DUE TO BLEEDING OF AN UNKNOWN ORIGIN. IT WAS REPORTED A COMPUTED TOMOGRAPHY (CT) SHOWED PROTRUDING CALCIUM ON THE ABDOMINAL AORTA. A POST OPERATIVE CT SHOWED THE RETROPERITONEAL BLEED WAS LIKELY FROM THE ABDOMINAL AORTA.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE, THE PATIENT HAD TORTUOUS FEMORAL ARTERIES AND PROTRUDING CALCIUM IN THE ABDOMINAL AORTA AND AORTIC ARCH. A PRE-DILATION WAS PERFORMED WITH A 26MM NON-MEDTRONIC BALLOON. THE DELIVERY CATHETER SYSTEM (DCS) WAS UNABLE TO TRAVERSE THE AORTIC ARCH AND MET RESISTANCE PROXIMAL OF THE LEFT SUBCLAVIAN ARTERY DESPITE BEING POSITIONED IN DIFFERENT ANGLES. THE GUIDEWIRE WAS SWITCHED. THE DCS WAS STUCK AGAIN AT THE POSITION DESPITE BEING POSITIONED AGAIN IN DIFFERENT ANGLES. FLUOROSCOPY DID NOT REVEAL AN INJURY TO THE AORTIC ARCH. THE NON-MEDTRONIC SHEATH WAS EXCHANGED FOR A 22FR 65CM LENGTH SHEATH. THE DCS WAS UNABLE TO CROSS THE AORTIC ARCH AGAIN DESPITE THE DIFFERENT POSITIONING ANGLES. THE DCS WAS REMOVED AND FLUSHED. THE DCS WAS INSERTED THROUGH THE NON-MEDTRONIC SHEATH AND WAS UNABLE TO CROSS THE AORTIC ARCH AGAIN DESPITE THE DIFFERENT POSITIONING ANGLES. A NEW SYSTEM WAS PREPPED AND THE VALVE WAS ABLE TO CROSS USING A SNARE. THE PATIENT WAS HYPOTENSIVE WITH SEVERE AORTIC INSUFFICIENCY AND WAS RESPONSIVE TO MEDICAL MANAGEMENT. THE VALVE WAS DEPLOYED INITIALLY TOO SHALLOW AND WAS RECAPTURED. A SECOND AND FINAL DEPLOYMENT ATTEMPT WAS AT 4 MM ON THE NON-CORONARY CUSP AND 6MM ON THE LEFT CORONARY CUSP. TRACE-MILD PARAVALVULAR LEAK WAS REPORTED. AN AORTOGRAM NOTED NO INJURY TO THE AORTIC ARCH OR ABDOMINAL AORTA. THE SHEATH TO THE LEFT FEMORAL ARTERY PARTIALLY DISLODGED CAUSING A HEMATOMA WHICH WAS ALLEVIATED WITH MANUAL COMPRESSION. TRANSIENT LEFT BUNDLE BRANCH BLOCK WAS NOTED. NO VASCULAR ISSUES TO THE RIGHT FEMORAL ARTERY WAS NOTED. A POST IMPLANT COMPUTED TOMOGRAPHY (CT) SHOWED NO DISSECTION, TEAR, OR INJURY TO THE AORTA. THE PATIENT BECAME UNSTABLE FOLLOWING THE PROCEDURE AND DIED THE FOLLOWING DAY DUE TO BLEEDING OF AN UNKNOWN ORIGIN. IT WAS REPORTED A CT SHOWED PROTRUDING CALCIUM ON THE ABDOMINAL AORTA. A POST OPERATIVE CT SHOWED THE RETROPERITONEAL BLEED WAS LIKELY FROM THE ABDOMINAL AORTA. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE PRE-IMPLANT BALLOON DILATION CAUSED THE SEVERE CENTRAL AORTIC INSUFFICIENCY. PER THE PHYSICIAN, A POTENTIAL CAUSE OF THE BLEED WAS DUE TO MULTIPLE SHEATH AND WIRE CHANGES; HOWEVER, THE BLEED WAS NOT SEEN ON THE FINAL AORTOGRAM. THE PHYSICIAN INDICATED THE BLEED MAY HAVE BEEN CAUSED BY THE SHEATH AND CLARIFIED NEITHER OF THE DCS CONTRIBUTED TO THE BLEED AS NEITHER DCS INTERACTED WITH THE PATIENT'S ANATOMY WHERE THE BLEED WAS LOCATED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS A POSSIBLE RETROPERITONEAL BLEED IN THE SETTING OF PROTRUDING CALCIUM AND THE PATIENT'S FAMILY'S WITHDRAWAL OF CARE. PER THE PHYSICIAN, THE MEDTRONIC VALVE AND DCS DID NOT CAUSE THE PATIENT'S DEATH. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: E VFXPLUS-34, SERIAL/LOT #: (B)(6), UBD: 23-DEC-2026, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22501002&pc=DTB | 2182208-2025-03397 | 2025-06-01 04:00 | Death | MEDTRONIC, INC. | 2025-07-14 04:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P930039 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING QRS MORPHOLOGY IN PACEMAKER IMPLANTS. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE. THERE WERE PATIENTS WHO EXPERIENCED ONE PNEUMOTHORAX, AND ONE POCKET INFECTION. THERE WERE THREE LEADS WHICH EXHIBITED LEAD DISPLACEMENTS. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/77 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: THE QRS INTERVAL AFTER PACEMAKER IMPLANT: AN INDEPENDENT MORTALITY RISK FACTOR. CRITICAL PATHWAYS IN CARDIOLOGY. 2025; 24: E0385. DOI: 10.1097/HPC.0000000000000385. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22501017&pc=DSQ | 2916596-2025-04577 | 2025-06-09 04:00 | Death | THORATEC CORPORATION | 2025-07-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Pumping Stopped | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND DECEASED BY FAMILY ON (B)(6) 2025. THE PATIENT'S FAMILY STATED THEY SPOKE WITH THE PATIENT ON (B)(6) 2025. THE SYSTEM CONTROLLER WAS INTERROGATED, AND THE LAST DATA WAS ON (B)(6) 2025 WHICH SHOWED PUMP STOP.Event Description: IT WAS LATER REPORTED THAT THE CAUSE OF DEATH AND PUMP STOP WAS UNKNOWN. IT WAS NOTED THAT IT WAS PRESUMED TO BE THAT THE PATIENT FELL ASLEEP ON THEIR BATTERIES, THEY WERE HARD OF HEARING AND THEIR SON STATED THAT THE PATIENT DID NOT HAVE THEIR HEARING AIDS IN WHEN THEY WERE DISCOVERED AND WAS ON BATTERIES. IT WAS ALSO STATED THAT THE DEVICE OPERATED AS EXPECTED ALTHOUGH IT WAS UNKNOW IF DEATH WAS CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22501067&pc=NPT | 2025587-2025-04987 | 2025-07-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-14 04:00 | NPT | EVOLUT FX DCS | Failure to Advance; Patient Device Interaction Problem | Hemorrhage/Bleeding | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE, THE PATIENT HAD TORTUOUS FEMORAL ARTERIES AND PROTRUDING CALCIUM IN THE ABDOMINAL AORTA AND AORTIC ARCH. A PRE-DILATION WAS PERFORMED WITH A 26 MILLIMETER (MM) NON-MEDTRONIC BALLOON. THE DELIVERY CATHETER SYSTEM (DCS) WAS UNABLE TO TRAVERSE THE AORTIC ARCH AND MET RESISTANCE PROXIMAL OF THE LEFT SUBCLAVIAN ARTERY DESPITE BEING POSITIONED IN DIFFERENT ANGLES. THE GUIDEWIRE WAS SWITCHED. THE DCS WAS STUCK AGAIN AT THE POSITION DESPITE BEING POSITIONED AGAIN IN DIFFERENT ANGLES. FLUOROSCOPY DID NOT REVEAL AN INJURY TO THE AORTIC ARCH. THE NON-MEDTRONIC SHEATH WAS EXCHANGED FOR A 22 FR 65 CM LENGTH SHEATH. THE DCS WAS UNABLE TO CROSS THE AORTIC ARCH AGAIN DESPITE THE DIFFERENT POSITIONING ANGLES. THE DCS WAS REMOVED AND FLUSHED. THE DCS WAS INSERTED THROUGH THE NON-MEDTRONIC SHEATH AND WAS UNABLE TO CROSS THE AORTIC ARCH AGAIN DESPITE THE DIFFERENT POSITIONING ANGLES. A NEW SYSTEM WAS PREPPED AND THE VALVE WAS ABLE TO CROSS USING A SNARE. THE PATIENT WAS HYPOTENSIVE WITH SEVERE AORTIC INSUFFICIENCY AND WAS RESPONSIVE TO MEDICAL MANAGEMENT. THE VALVE WAS DEPLOYED INITIALLY TOO SHALLOW AND WAS RECAPTURED. A SECOND AND FINAL DEPLOYMENT ATTEMPT WAS AT 4 MM ON THE NON-CORONARY CUSP (NCC) AND 6 MM ON THE LEFT CORONARY CUSP (LCC). TRACE-MILD PARAVALVULAR LEAK (PVL) WAS REPORTED. AN AORTOGRAM NOTED NO INJURY TO THE AORTIC ARCH OR ABDOMINAL AORTA. THE SHEATH TO THE LEFT FE MORAL ARTERY PARTIALLY DISLODGED CAUSING A HEMATOMA WHICH WAS ALLEVIATED WITH MANUAL COMPRESSION. TRANSIENT LEFT BUNDLE BRANCH BLOCK (LBBB) WAS NOTED. NO VASCULAR ISSUES TO THE RIGHT FEMORAL ARTERY WAS NOTED. A POST IMPLANT COMPUTED TOMOGRAPHY (CT) SHOWED NO DISSECTION, TEAR, OR INJURY TO THE AORTA. THE PATIENT BECAME UNSTABLE FOLLOWING THE PROCEDURE AND DIED THE FOLLOWING DAY DUE TO BLEEDING OF AN UNKNOWN ORIGIN. IT WAS REPORTED A COMPUTED TOMOGRAPHY (CT) SHOWED PROTRUDING CALCIUM ON THE ABDOMINAL AORTA. A POST OPERATIVE CT SHOWED THE RETROPERITONEAL BLEED WAS LIKELY FROM THE ABDOMINAL AORTA.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE, THE PATIENT HAD TORTUOUS FEMORAL ARTERIES AND PROTRUDING CALCIUM IN THE ABDOMINAL AORTA AND AORTIC ARCH. A PRE-DILATION WAS PERFORMED WITH A 26MM NON-MEDTRONIC BALLOON. THE DELIVERY CATHETER SYSTEM (DCS) WAS UNABLE TO TRAVERSE THE AORTIC ARCH AND MET RESISTANCE PROXIMAL OF THE LEFT SUBCLAVIAN ARTERY DESPITE BEING POSITIONED IN DIFFERENT ANGLES. THE GUIDEWIRE WAS SWITCHED. THE DCS WAS STUCK AGAIN AT THE POSITION DESPITE BEING POSITIONED AGAIN IN DIFFERENT ANGLES. FLUOROSCOPY DID NOT REVEAL AN INJURY TO THE AORTIC ARCH. THE NON-MEDTRONIC SHEATH WAS EXCHANGED FOR A 22FR 65CM LENGTH SHEATH. THE DCS WAS UNABLE TO CROSS THE AORTIC ARCH AGAIN DESPITE THE DIFFERENT POSITIONING ANGLES. THE DCS WAS REMOVED AND FLUSHED. THE DCS WAS INSERTED THROUGH THE NON-MEDTRONIC SHEATH AND WAS UNABLE TO CROSS THE AORTIC ARCH AGAIN DESPITE THE DIFFERENT POSITIONING ANGLES. A NEW SYSTEM WAS PREPPED AND THE VALVE WAS ABLE TO CROSS USING A SNARE. THE PATIENT WAS HYPOTENSIVE WITH SEVERE AORTIC INSUFFICIENCY AND WAS RESPONSIVE TO MEDICAL MANAGEMENT. THE VALVE WAS DEPLOYED INITIALLY TOO SHALLOW AND WAS RECAPTURED. A SECOND AND FINAL DEPLOYMENT ATTEMPT WAS AT 4 MM ON THE NON-CORONARY CUSP AND 6MM ON THE LEFT CORONARY CUSP. TRACE-MILD PARAVALVULAR LEAK WAS REPORTED. AN AORTOGRAM NOTED NO INJURY TO THE AORTIC ARCH OR ABDOMINAL AORTA. THE SHEATH TO THE LEFT FEMORAL ARTERY PARTIALLY DISLODGED CAUSING A HEMATOMA WHICH WAS ALLEVIATED WITH MANUAL COMPRESSION. TRANSIENT LEFT BUNDLE BRANCH BLOCK WAS NOTED. NO VASCULAR ISSUES TO THE RIGHT FEMORAL ARTERY WAS NOTED. A POST IMPLANT COMPUTED TOMOGRAPHY (CT) SHOWED NO DISSECTION, TEAR, OR INJURY TO THE AORTA. THE PATIENT BECAME UNSTABLE FOLLOWING THE PROCEDURE AND DIED THE FOLLOWING DAY DUE TO BLEEDING OF AN UNKNOWN ORIGIN. IT WAS REPORTED A CT SHOWED PROTRUDING CALCIUM ON THE ABDOMINAL AORTA. A POST OPERATIVE CT SHOWED THE RETROPERITONEAL BLEED WAS LIKELY FROM THE ABDOMINAL AORTA. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE PRE-IMPLANT BALLOON DILATION CAUSED THE SEVERE CENTRAL AORTIC INSUFFICIENCY. PER THE PHYSICIAN, A POTENTIAL CAUSE OF THE BLEED WAS DUE TO MULTIPLE SHEATH AND WIRE CHANGES; HOWEVER, THE BLEED WAS NOT SEEN ON THE FINAL AORTOGRAM. THE PHYSICIAN INDICATED THE BLEED MAY HAVE BEEN CAUSED BY THE SHEATH AND CLARIFIED NEITHER OF THE DCS CONTRIBUTED TO THE BLEED AS NEITHER DCS INTERACTED WITH THE PATIENT'S ANATOMY WHERE THE BLEED WAS LOCATED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS A POSSIBLE RETROPERITONEAL BLEED IN THE SETTING OF PROTRUDING CALCIUM AND THE PATIENT'S FAMILY'S WITHDRAWAL OF CARE. PER THE PHYSICIAN, THE MEDTRONIC VALVE AND DCS DID NOT CAUSE THE PATIENT'S DEATH. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22501150&pc=DTB | 2182208-2025-03396 | 2025-06-01 04:00 | Death | MEDTRONIC, INC. | 2025-07-14 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P930039 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING QRS MORPHOLOGY IN PACEMAKER IMPLANTS. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE. THERE WERE PATIENTS WHO EXPERIENCED ONE PNEUMOTHORAX, AND ONE POCKET INFECTION. THERE WERE THREE LEADS WHICH EXHIBITED LEAD DISPLACEMENTS. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/77 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: THE QRS INTERVAL AFTER PACEMAKER IMPLANT: AN INDEPENDENT MORTALITY RISK FACTOR. CRITICAL PATHWAYS IN CARDIOLOGY. 2025; 24: E0385. DOI: 10.1097/HPC.0000000000000385. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22501494&pc=OAE | 3008452825-2025-00341 | 2025-06-05 04:00 | Death | ABBOTT MEDICAL | 2025-07-14 04:00 | OAE | TACTIFLEX¿ ABLATION CATHETER, SENSOR ENABLED¿ | Adverse Event Without Identified Device or Use Problem | Perforation of Esophagus | P220013 |  | 1 | Event Description: ONE MONTH FOLLOWING AN ATRIAL FIBRILLATION PROCEDURE, THE PATIENT EXPIRED DUE TO AN ESOPHAGEAL FISTULA. CAUSES AND CONTRIBUTING FACTORS ARE CURRENTLY UNKNOWN AND THERE HAS BEEN NO ALLEGED DEVICE INVOLVEMENT. Manufacturer Narrative: MODEL/LOT NUMBER INFORMATION WAS NOT ABLE TO BE OBTAINED FOR THIS DEVICE. THEREFORE, FULL UDI INFORMATION(D4) AND 510K(G3) ARE NOT AVAILABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: D4, G3, H2, H4.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22502342&pc=NPT | 2025587-2025-04993 | 2025-06-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-14 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Thrombosis/Thrombus | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED DUE TO THE PATIENT'S FIBROUS AND TIGHT ANATOMY. FOLLOWING THE DEPLOYMENT OF THE TRANSCATHETER VALVE, THROMBUS DEVELOPED, RESULTING IN A LEFT MAIN CORONARY OCCLUSION. SUBSEQUENTLY, AFTER 30 MINUTES OF UNKNOWN TREATMENT, THE PROCEDURE WAS CONVERTED TO OPEN HEART SURGERY, DURING WHICH THE OCCLUSION WAS OPENED AND PERFUSION WAS RESTORED; HOWEVER, THE PATIENT ULTIMATELY DIED. IT WAS NOTED THAT A 2 HOUR PROCEDURAL DELAY OCCURRED AS A RESULT OF THE EVENT. PER THE PHYSICIAN, THROMBUS FORMATION CAUSED THE DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE VALVE WAS IMPLANTED IN THE PATIENT'S NATIVE ANNULUS. THE THROMBUS WAS LOCATED IN THE PATIENT'S LEFT MAIN CORONARY ARTERY AND PERCUTANEOUS CORONARY INTERVENTION (PCI) WAS SUBSEQUENTLY PERFORMED INSTEAD OF OPEN HEART SURGERY. HEPARIN WAS ADMINISTERED DURING THE PROCEDURE AND THE PATIENT'S ACTIVATED CLOTTING TIME (ACT) WAS MONITORED EVERY 5-10 MINUTES WITH THE ACT NEVER BELOW 250. IT WAS NOTED THAT THE PATIENT HAD ATRIAL FIBRILLATION AND THE PATIENT'S BLOOD THINNER MEDICATION WAS DISCONTINUED 1 WEEK PRIOR TO IMPLANT. PER THE PHYSICIAN, THE DELIVERY CATHETER SYSTEM (DCS) COULD BE RULED OUT AS A CONTRIBUTING FACTOR TO THE DEATH. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {D-EVOLUTFX-2329}; PRODUCT LOT/SERIAL NUMBER {UNNOWN}; PRODUCT TYPE: {0195 HEART VALVES}; IMPLANT DATE {NA}; EXPLANT DATE {NA}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22502969&pc=NVN | 2017865-2025-93064 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-14 04:00 | NVN | TENDRIL SDX LEAD | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PRESENTED WITH AN INFECTION NOTED ON THE SYSTEM. IT WAS THEN NOTED THAT THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. NO FURTHER INTERVENTION WAS REPORTED. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN Manufacturer Narrative: THE REPORTED EVENT INDICATED INFECTION. AS RECEIVED, A COMPLETE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. OTHER DAMAGE FOUND IS CONSISTENT WITH PROCEDURAL DAMAGE. DURING ANALYSIS, A FAILURE EVENT WAS OBSERVED WHICH WAS UNRELATED TO THE REPORTED EVENT. FINAL \VISUAL EXAMINATION FOUND FOUR EXTERNAL INSULATION ABRASIONS BREACHING THE OUTER INSULATION EXPOSING THE OUTER COIL, ONE AT PROXIMAL REGION CONSISTENT WITH FRICTION TO THE DEVICE CAN, AND THREE AT THE DISTAL REGION CONSISTENT WITH FRICTION TO ANOTHER DEVICE OR FEATURE OF PATIENT¿S ANATOMY. CORRECTION: D9: FIELD SHOULD BE MARKED (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22502970&pc=NIK | 2017865-2025-93062 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-14 04:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PRESENTED WITH AN INFECTION NOTED ON THE SYSTEM. IT WAS THEN NOTED THAT THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. NO FURTHER INTERVENTION WAS REPORTED. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. THE PRODUCT WAS RETURNED AND VISUAL INSPECTION WAS NORMAL. VISUAL EXAMINATION FOUND NO MALFUNCTIONS. THE CAUSE OF INFECTION COULD NOT BE TRACED TO THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22502972&pc=NVN | 2017865-2025-93061 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-14 04:00 | NVN | TENDRIL SDX LEAD | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PRESENTED WITH AN INFECTION NOTED ON THE SYSTEM. IT WAS THEN NOTED THAT THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. NO FURTHER INTERVENTION WAS REPORTED. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: THE REPORTED EVENT INDICATED INFECTION. AS RECEIVED, A PARTIAL LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. OTHER DAMAGE FOUND IS CONSISTENT WITH PROCEDURAL DAMAGE. DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. DURING ANALYSIS, A FAILURE EVENT WAS OBSERVED WHICH WAS UNRELATED TO THE REPORTED EVENT. FINAL VISUAL EXAMINATION FOUND THREE EXTERNAL INSULATION ABRASION AT THE DISTAL REGION EXPOSING THE OUTER COIL CONSISTENT WITH FRICTION TO ANOTHER DEVICE OR FEATURE OF PATIENT¿S ANATOMY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22502973&pc=NVY | 2017865-2025-93060 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-14 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PRESENTED WITH AN INFECTION NOTED ON THE SYSTEM. IT WAS THEN NOTED THAT THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. NO FURTHER INTERVENTION WAS REPORTED. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. THE PRODUCT WAS RETURNED AND VISUAL INSPECTION WAS NORMAL. VISUAL EXAMINATION FOUND NO MALFUNCTIONS. THE CAUSE OF INFECTION COULD NOT BE TRACED TO THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22503401&pc=NIK | 2017865-2025-93055 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-14 04:00 | NIK | QUADRA ASSURA CRT-D QUAD RF HV | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PRESENTED WITH AN INFECTION NOTED ON THE SYSTEM. IT WAS THEN NOTED THAT THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. NO FURTHER INTERVENTION WAS REPORTED. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22503599&pc=DSQ | 2916596-2025-04480 | 2025-06-27 04:00 | Death | THORATEC CORPORATION | 2025-07-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIORGAN FAILURE. THE PATIENT WAS TRANSITIONED TO COMFORT CARE. THE PUMP FUNCTIONED AS INTENDED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WAS NOT RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22503616&pc=NAY | 2955842-2025-29761 | 2025-06-18 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-07-14 04:00 | NAY | DAVINCI XI | Insufficient Information | Asystole; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT EXPIRED DURING A DA VINCI-ASSISTED PARAESOPHAGEAL HIATAL HERNIA SURGICAL PROCEDURE. THE SURGEON REPORTED THAT THE PATIENT HAD A HISTORY OF LUNG CANCER WITH RADIATION AND CHEMOTHERAPY TREATMENTS RESULTING IN DIFFICULT, SCARRED, AND WEAK TISSUE ANATOMY. WHILE DISSECTING TISSUE NEAR THE VAGUS NERVE, CONTROLLED BLEEDING OCCURRED AND PRESSURE WAS APPLIED. THE PATIENT THEN EXPERIENCED BRADYCARDIA FROM A VASOVAGAL RESPONSE THAT DETERIORATED TO ASYSTOLE. THE BLEEDING SITE WAS PACKED, THE ROBOT WAS EMERGENTLY UNDOCKED, AND THE PROCEDURE WAS CONVERTED TO OPEN. CHEST COMPRESSIONS WERE INITIATED. DESPITE RESUSCITATION EFFORTS THE PATIENT REMAINED IN ASYSTOLE; NO CARDIAC RHYTHM WAS RESTORED AND THE PATIENT EXPIRED. THE SURGEON STATED THAT NO MALFUNCTIONS OF ANY INTUITIVE PRODUCTS OCCURRED DURING THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM SHOW THE SYSTEM FUNCTIONED NORMALLY AND THERE WERE NO INDICATIONS RELATING TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, TWO CADIERE FORCEPS, PERMANENT CAUTERY HOOK, VESSEL SEALER EXTEND, SUCTION IRRIGATOR, 8MM 0 DEGREE ENDOSCOPE. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) INVOLVED WITH THE REPORTED EVENT FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS COMPLAINT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A BRADYCARDIC EVENT LIKELY FROM A VAGAL RESPONSE DURING A HIATAL HERNIA REPAIR. THIS LED TO ASYSTOLE AND THE PATIENT EXPIRED DESPITE RESUSCITATIVE EFFORTS. NO ALLEGATION IS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22503625&pc=NAY | 2955842-2025-29256 | 2025-06-16 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-07-14 04:00 | NAY | DAVINCI 5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Multiple Organ Failure | K232610 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI-ASSISTED PANCREATICODUODENECTOMY SURGICAL PROCEDURE FOR A BENIGN TUMOR, THE PATIENT EXPIRED POSTOPERATIVELY. THE ROBOTIC PROCEDURE WAS COMPLETED AS PLANNED WITHOUT ANY INTRAOPERATIVE COMPLICATIONS. BLOOD LOSS WAS MINIMAL AND THE SURGICAL FIELD WAS "DRY" AT THE END OF THE PROCEDURE. THE SURGEON REPORTED THAT THE PATIENT REQUIRED A SUBSEQUENT OPEN LAPAROTOMY A FEW HOURS POSTOPERATIVELY DUE TO BLOOD PRESSURE COMPLICATIONS INDICATING BLEEDING. AN OPEN LAPAROTOMY WAS PERFORMED AND THE BLEEDING WAS IDENTIFIED FROM A SMALL BLOOD VESSEL. THE BLEEDING WAS STOPPED WITH UNSPECIFIED INTERVENTIONS, AND THE ANASTOMOSIS WAS REDONE. THE PATIENT ULTIMATELY EXPIRED THREE DAYS POSTOPERATIVELY IN THE INTENSIVE CARE UNIT DUE TO MULTISYSTEM ORGAN FAILURE. THE SURGEON REPORTED THAT THE PATIENT¿S POSTOPERATIVE COURSE WAS RELATED TO THE MANAGEMENT OF MEDICAL CARE AND DOES NOT ATTRIBUTE THE PATIENT'S POSTOPERATIVE COURSE TO BE RELATED TO THE INITIAL DA VINCI PROCEDURE. NO SPECIFIC DETAILS WERE PROVIDED. Manufacturer Narrative: REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS WERE LOGGED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON DETERMINED THAT THE SYSTEM FUNCTIONED NORMALLY, NO INTRAOPERATIVE EVENTS OR ISSUES WERE FOUND. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: ENDOSCOPE 30 DEGREE, TIP-UP FENESTRATED GRASPER, FORCE BIPOLAR, TWO CLIP APPLIERS, MONOPOLAR CURVED SCISSORS, PERMANENT CAUTERY HOOK, LARGE SUTURECUT NEEDLE DRIVER, FORCE FEEDBACK MARYLAND BIPOLAR FORCEPS, VESSEL SEALER EXTEND, SUREFORM 60 STAPLER, SUREFORM 30 CURVED-TIP STAPLER, SUREFORM 45 STAPLER, SUCTION IRRIGATOR. SITE HISTORY REVIEW FOUND NO COMPLAINTS FOR THE INSTRUMENTS USED IN THE PROCEDURE, AND THE MULTIPLE-USE DEVICES WERE USED IN SUBSEQUENT PROCEDURES AFTER THE EVENT DATE WITH NO REPORTED COMPLAINTS. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE INSTRUMENTS CONFIRMED THAT THERE WERE NO RELATED NON-CONFORMANCES DOCUMENTED. A REVIEW OF THE VESSEL SEALER EXTEND LOGS FOUND THE INSTRUMENT WAS INSTALLED ONCE AND PASSED HOMING. THE CUT COMPLETE AND SEAL EVENTS WERE COMPLETED WITH NO RELEVANT ERRORS NOTED. STAPLER LOG REVIEWS OF THE STAPLER INSTRUMENTS FOUND THE SUREFORM 60 STAPLER WAS INSTALLED ONCE AND FIRED ONE BLUE RELOAD. THE SUREFORM 30 CURVED-TIP STAPLER INSTRUMENT WAS INSTALLED FIVE TIMES AND FIRED THREE WHITE RELOADS. THE SUREFORM 45 STAPLER INSTRUMENT WAS INSTALLED FIVE TIMES AND FIRED THREE BLUE RELOADS AND ONE WHITE RELOAD. THERE WERE NO STAPLER RELATED ERRORS IN THE LOGS. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT CONCLUDED THAT THE PATIENT IN THIS REPORT DEVELOPED A DELAYED BLEED AFTER A ROBOTIC PANCREATICODUODENECTOMY. WHAT SPECIFIC VESSEL BLED AND HOW OR WHY IT MAY HAVE BLED WAS NOT PROVIDED. THE AMOUNT OF BLEEDING AND THE ISSUES THAT RESULTED IN THE BLEEDING WERE NOT PROVIDED. THE SPECIFIC ANASTOMOSIS THAT WAS REDONE WAS NOT PROVIDED. PERFORMING THE ANASTOMOSIS REVISION MAY INDICATE THAT IT WAS THE SITE OF BLEEDING; HOWEVER, THE SPECIFIC SITE OF BLEEDING WAS NOT DEFINED. THE PATIENT DEVELOPED MULTISYSTEM ORGAN FAILURE AND EXPIRED 3 DAYS LATER; HOW THE ORGAN FAILURE DEVELOPED WAS NOT PROVIDED. NO ALLEGATION WAS MADE ABOUT ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22503709&pc=DSQ | 2916596-2025-04526 | 2025-07-01 04:00 | Death | THORATEC CORPORATION | 2025-07-14 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cyanosis; Syncope/Fainting | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND UNCONSCIOUS. THE PATIENT WAS HEMODYNAMICALLY STABLE AND THEN FOUND CYANOTIC AND UNCONSCIOUS SOME MINUTES AFTER. IT WAS REPORTED THAT THE PATIENT HAD A VENTRICULAR PACED HEART RHYTHM. NO TACHYCARDIA ARRHYTHMIAS WERE REPORTED. THE SYSTEM CONTROLLER WAS EXCHANGED TO THE BACKUP SYSTEM CONTROLLER WITHOUT RESTORING FLOW. LOG FILES WERE OBTAINED FROM THE SYSTEM CONTROLLER. THERE WERE LOW FLOW ALARMS WITH AN INCREASE IN PULSATILITY INDEX (PI). FLOW IN RANGE OF 3.9-0 LITERS/MIN IN JUST A FEW MINUTES. THE PATIENT EXPIRED SUDDENLY, THE CAUSE WAS UNKNOWN AND LOW FLOW ALARMS WERE OBSERVED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: EVALUATION OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), DID NOT REVEAL ANY DEVICE-RELATED ISSUES THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. A SPECIFIC CAUSE FOR THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. FURTHERMORE, REVIEW OF THE SUBMITTED LOG FILES AND LOG FILES RETRIEVED FROM THE RETURNED DEVICE CONFIRMED THE REPORTED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE FULLY INTACT. THE MODULAR CABLE WAS RETURNED CONNECTED TO THE PUMP CABLE. APPROXIMATELY 8.5¿ OF THE OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP COVER OUTLET PORT WITH THE OUTFLOW GRAFT BEND RELIEF ENGAGED TO THE GRAFT HARDWARE. THE APICAL CUFF WAS RETURNED SECURED WITH THE CUFF LOCK FULLY ENGAGED. UPON DISASSEMBLY OF THE PUMP BODY, EXAMINATION OF THE BLOOD-CONTACTING SURFACES REVEALED NO EVIDENCE OF DEVELOPED DEPOSITIONS OR THROMBUS FORMATIONS THAT WOULD HAVE CONTRIBUTED TO A FUNCTIONAL ISSUE. VISUAL INSPECTION OF THE PUMP ROTOR AND ROTOR WELL DID NOT REVEAL ANY OBVIOUS SURFACE SCRATCHES OR DEFECTS. THE SUBMITTED CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 28JUN2025 THROUGH 01JUL2025. THE FILE CAPTURED A SUSTAINED LOW FLOW HAZARD ALARM ON 01JUL2025. A SPECIFIC CAUSE FOR THE LOW FLOW EVENTS COULD NOT BE CONCLUSIVELY DETERMINED. TOWARD THE END OF THE FILE, THE DRIVELINE WAS DISCONNECTED RESULTING IN A PUMP STOP. THE CONTROLLER WAS THEN SHUT DOWN. THESE EVENTS APPEAR CONSISTENT WITH THE REPORTED EVENT OF THE PATIENT¿S EXPIRATION. NO OTHER NOTABLE EVENTS OR ALARMS REGARDING PUMP FUNCTION WERE CAPTURED, AND THE SYSTEM APPEARED TO BE OPERATING AS INTENDED PRIOR TO THE DISCONNECTION OF THE DRIVELINE. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) EVENT LOG FILE RETRIEVED FROM THE RETURNED DEVICE CAPTURED LOW FLOW EVENTS ON 01JUL2025, CONSISTENT WITH THE EVENTS OBSERVED IN THE SUBMITTED CONTROLLER EVENT LOG FILE. NO OTHER NOTABLE EVENTS WERE OBSERVED IN THE LVAD LOG FILES, AND THE PUMP APPEARED TO BE OPERATING AS INTENDED. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. B ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THIS SECTION ALSO ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22503713&pc=NAY | 2955842-2025-29254 | 2025-04-11 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-07-14 04:00 | NAY | DAVINCI 5 | Adverse Event Without Identified Device or Use Problem | Aspiration/Inhalation; Pulmonary Edema | K232610 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI-ASSISTED PARAESOPHAGEAL HIATAL HERNIA SURGICAL PROCEDURE, THE PATIENT EXPIRED DUE TO COMPLICATIONS OF AN ASPIRATION EVENT. THE SURGEON STATED THAT THE HIATAL HERNIA WAS LARGE, BUT THE ROBOTIC PROCEDURE WAS COMPLETED SUCCESSFULLY WITHOUT ANY INTRAOPERATIVE COMPLICATIONS, AND WITHOUT ERRORS OR ISSUES RELATED TO THE DA VINCI SYSTEM. THE PATIENT WAS RECOVERING; HOWEVER, REQUIRED PROLONGED INTUBATION DUE TO PULMONARY EDEMA. THREE WEEKS POST-PROCEDURE, THE PATIENT WAS STILL INTUBATED, AND DURING AN ENDOTRACHEAL TUBE EXCHANGE, THE PATIENT ASPIRATED AND ULTIMATELY EXPIRED DUE TO COMPLICATIONS FROM THE ASPIRATION EVENT. THE SURGEON BELIEVES THE PROLONGED HOSPITALIZATION AND INTUBATION WERE RELATED TO THE PATIENT'S LARGE HIATAL HERNIA AND PULMONARY EDEMA. THE PATIENT¿S MEDICAL HISTORY AND SPECIFIC ORIGIN OF THE PULMONARY EDEMA WERE NOT PROVIDED. Manufacturer Narrative: REVIEW OF THE SYSTEM LOGS FOUND THE SYSTEM FUNCTIONED NORMALLY AND THERE WERE NO ERRORS FOUND RELATED TO THIS EVENT. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM DETERMINED IT FUNCTIONED NORMALLY, NO INTRAOPERATIVE EVENTS OR ISSUES WERE FOUND. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: ENDOSCOPE 30 DEGREE, FORCE FEEDBACK CADIERE FORCEPS, FORCE FEEDBACK FENESTRATED BIPOLAR FORCEPS, FORCE FEEDBACK MEGA SUTURECUT NEEDLE DRIVER, VESSEL SEALER EXTEND, SUCTION IRRIGATOR. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THESE PRODUCTS. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE INSTRUMENTS CONFIRMED THAT THERE WERE NO RELATED NON-CONFORMANCES DOCUMENTED. A REVIEW OF THE VESSEL SEALER EXTEND LOGS FOUND THE INSTRUMENT WAS INSTALLED ONCE AND PASSED HOMING. THE CUT COMPLETE AND SEAL EVENTS WERE COMPLETED WITH NO RELEVANT ERRORS NOTED. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AS A RESULT OF AN ASPIRATION EVENT WHILE IN THE ICU SEVERAL WEEKS AFTER A ROBOTIC HIATAL HERNIA REPAIR. THIS TYPE OF POST OPERATIVE COMPLICATION CAN OCCUR WITH INTUBATED PATIENTS AND HAS NO RELATIONSHIP TO THE ORIGINAL PROCEDURE. THERE IS NO ALLEGATION AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION B2 DATE OF DEATH IS ESTIMATED. THE SURGEON STATED THAT THE PATIENT'S DATE OF DEATH WAS UNKNOWN, APPROXIMATELY THREE WEEKS POST-PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22476542&pc=OZP | 2032227-2025-209468 | 2025-01-23 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-11 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Coma | P160017 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2025, AT HOME, AND THE CAUSE OF DEATH WAS DIABETES COMPLICATIONS. THE CUSTOMER EXPERIENCED HYPERGLYCEMIA. THE BLOOD GLUCOSE VALUE WAS 1100 MG/DL AT THE TIME OF ADMISSION TO THE HOSPITAL. THE CUSTOMER'S HYPERGLYCEMIA WAS TREATED WITH HOSPITALIZATION, AND THE CUSTOMER WAS ADMITTED TO THE HOSPITAL DUE TO NO LONGER SWALLOWING ANY FOOD AND WENT INTO A COMA. THE EVENT INVOLVED PRODUCTS UNK-SENSOR, MMT-1884, UNOMEDICAL, AND MMT-332A. TROUBLESHOOTING WAS PARTIALLY PERFORMED. THE CUSTOMER'S PASSING WAS DUE TO DIABETES COMPLICATIONS. THE CUSTOMER WORE THE PUMP AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR UNK-SENSOR, MMT-1884, UNOMEDICAL, AND MMT-332A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22476601&pc=NAY | 2955842-2025-29442 | 2025-06-18 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-07-11 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Insufficient Information | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN UNSPECIFIED DA VINCI-ASSISTED PROCEDURE, THE PATIENT LATER PRESENTED FOR UNSPECIFIED REASONS, UNDERWENT AN EXPLORATORY LAPAROTOMY, AND SUBSEQUENTLY EXPIRED. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT NO RESPONSE HAS BEEN RECEIVED. Manufacturer Narrative: INSUFFICIENT INFORMATION IS AVAILABLE FOR PROCEDURE VERIFICATION; THEREFORE, NO DA VINCI SYSTEM OR INSTRUMENT LOGS ARE AVAILABLE FOR REVIEW. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING AN UNKNOWN ROBOTIC PROCEDURE. NO ADDITIONAL INFORMATION IS PROVIDED ABOUT THE PROCEDURE OR ANY INTRAOPERATIVE ISSUES. THE EVENTS THAT LED TO THEIR REOPERATION, THE FINDINGS OF THE REOPERATION, AND THE SUBSEQUENT ISSUES THAT LED TO THEIR DEATH AND CAUSE OF DEATH ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22476750&pc=NAY | 2955842-2025-29255 | 2025-06-19 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-07-11 04:00 | NAY | DAVINCI XI | Insufficient Information | Hemorrhage/Bleeding | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LIVER RESECTION SURGICAL PROCEDURE, THE PATIENT EXPIRED DUE TO AN INJURY TO THE VENA CAVA, WHICH RESULTED IN A SIGNIFICANT AMOUNT OF BLOOD LOSS. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION HAVE BEEN MADE; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. Manufacturer Narrative: REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS WERE LOGGED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS OCCURRED. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, MEGA SUTURECUT NEEDLE DRIVER, MONOPOLAR CURVED SCISSORS, FENESTRATED BIPOLAR FORCEPS, PROGRASP FORCEPS, MARYLAND BIPOLAR FORCEPS, VESSEL SEALER EXTEND, SUREFORM 60 STAPLER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THE PRODUCTS. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE INSTRUMENTS CONFIRMED THAT THERE WERE NO RELATED NON-CONFORMANCES DOCUMENTED. STAPLER LOG REVIEW SHOWS THE SUREFORM 60 STAPLER INSTRUMENT WAS INSTALLED ON THE SYSTEM TWICE AND SUCCESSFULLY FIRED 2 BLUE RELOADS. THERE WERE NO STAPLER RELATED ERRORS IN THE LOGS. A REVIEW OF THE VESSEL SEALER EXTEND INSTRUMENT AND THE E-100 GENERATOR LOGS FOUND THE INSTRUMENT WAS INSTALLED ONCE AND PASSED HOMING. THE CUT COMPLETE AND SEAL EVENTS WERE COMPLETED WITH NO RELEVANT ERRORS NOTED. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT IN THE REPORT DIED AS THE RESULT OF BLEEDING FROM AN INJURY TO THE VENA CAVA DURING A LIVER RESECTION. HOW THIS EVENT OCCURRED OR DETAILS OF WHAT MAY HAVE LED TO THE INJURY ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22476830&pc=DSQ | 2916596-2025-04450 | 2025-06-20 04:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO BRAIN DEATH. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22477249&pc=QZI | 2124215-2025-46404 | 2025-01-23 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-11 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Pneumonia; Hypervolemia; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER, THE PATIENT PASSED AWAY. IT WAS REPORTED THAT APPROXIMATELY 2 AND A HALF MONTHS AFTER THE PROCEDURE, THE PATIENT PASSED AWAY. PER THE PATIENT'S LAST HOSPITAL VISIT, THEY WERE EXPERIENCING AN ACUTE KIDNEY INJURY, PNEUMONIA, AND FLUID OVERLOAD. IT WAS NOT REPORTED IF ANY OF THESE COMPLICATIONS WERE CONSIDERED A RESULT OF THE PROCEDURE, SO THE TREATMENT CANNOT BE RULED OUT AS A CAUSE. THE CATHETER IS PRESUMED TO HAVE BEEN DISPOSED BY THE FACILITY AND WILL NOT BE RETURNED AS THE DEVICE PERFORMED AS EXPECTED DURING THE PROCEDURE AND THERE WERE NO REPORTS OF ANY PERFORMANCE CONCERNS, AND THEY ARE INSTRUCTED TO DISPOSE OF THE DEVICE AFTER THE PROCEDURE. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. GOOD FAITH EFFORTS TO OBTAIN THE INFORMATION ARE IN PROGRESS. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE FILED IF THE INFORMATION IS RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22477384&pc=NVN | 2649622-2025-18695 | 2025-06-23 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-11 04:00 | NVN | CAPSUREFIX LEAD | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED AN INFECTION. IT WAS FURTHER REPORTED THAT VEGETATION WAS OBSERVED. THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS EXPLANTED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS NOTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: ADDR01 IPG 5076-45 LEAD; THERAPY DATE: (B)(6) 2017. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22477397&pc=NVN | 2649622-2025-18696 | 2025-06-23 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-11 04:00 | NVN | CAPSUREFIX LEAD | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED AN INFECTION. IT WAS FURTHER REPORTED THAT VEGETATION WAS OBSERVED. THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS EXPLANTED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS NOTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: ADDR01 IPG 5076-45 LEAD; THERAPY DATE: (B)(6) 2017. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22477502&pc=BZD | 2518422-2025-047989 | 2024-04-23 04:00 | Death | RESPIRONICS, INC. | 2025-07-11 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS EVALUATED, THE CUSTOMER RESPONDED, BUT UNABLE TO GATHER ADDITIONAL INFORMATION RELATED DEVICE RETURN. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22477655&pc=NVZ | 3004209178-2025-11790 | 2025-06-23 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-11 04:00 | NVZ | ADAPTA | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED AN INFECTION. IT WAS FURTHER REPORTED THAT VEGETATION WAS OBSERVED. THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS EXPLANTED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS NOTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: 4068-52 LEAD, 4068-45 LEAD; THERAPY DATE: (B)(6) 1998. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22477660&pc=DTB | 2649622-2025-18697 | 2025-06-23 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-11 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED AN INFECTION. IT WAS FURTHER REPORTED THAT VEGETATION WAS OBSERVED. THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM WAS EXPLANTED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS NOTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: 4068-52 LEAD, 4068-45 LEAD; THERAPY DATE: (B)(6) 1998. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22478090&pc=OZP | 2032227-2025-210016 | 2025-06-18 04:00 | Death | MEDTRONIC MINIMED | 2025-07-11 04:00 | OZP | UNKNOWN NGP | Adverse Event Without Identified Device or Use Problem | Hypoglycemia; Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER HAD PASSED AWAY AND THE CAUSE OF DEATH WAS HYPOGLYCEMIA. THE CUSTOMER REPORTED THE HYPOGLYCEMIA WAS TREATED WITH SEVERAL AMPULES OF GLUCAGON AND EMERGENCY MEDICAL SERVICES. THE EVENT INVOLVED PRODUCTS UNK_NGP, UNOMEDICAL, UNK_TRANSMITTER, MMT-332A, AND UNK_SENSOR. TROUBLESHOOTING WAS NOT PERFORMED, AND IT WAS UNKNOWN WHETHER THE CUSTOMER HAS BEEN USING THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED LOW BLOOD GLUCOSE EVENT OR NOT. IT WAS UNKNOWN WHETHER THE CUSTOMER WAS USED THE AUTO MODE FEATURE AT THE TIME OF THE EVENT OR NOT. IT WAS UNKNOWN WHETHER THE PUMP WAS WORN AT THE TIME OF PASSING OR NOT. NO PRODUCT RETURN IS REQUIRED FOR UNK_NGP, UNOMEDICAL, UNK_TRANSMITTER, MMT-332A, AND UNK_SENSOR. Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22478244&pc=DSQ | 2916596-2025-04416 | 2025-05-24 04:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22478299&pc=DSQ | 2916596-2025-04508 | 2025-06-30 04:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED, AND THERE IS NO INFORMATION REGARDING THE OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV C, IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22478377&pc=DSQ | 2916596-2025-04556 | 2025-04-10 04:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS THAT THE PATIENT NEVER RECOVERED AFTER IMPLANTATION. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED, NO RECOVERING AFTER IMPLANTATION. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV C, IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22478390&pc=JXG | 9612501-2025-01809 | 2025-05-22 04:00 | Death | MEDTRONIC DOMINICANA | 2025-07-11 04:00 | JXG | UNKNOWN CATHETER | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: CHIA-NING CHANG, CHIUNG-CHYI SHEN, MENG-YIN YANG, WEN-YU CHENG, AND CHIH-MING LAI. SAFETY AND EFFICACY OF STEREOTACTIC ASPIRATION WITH FIBRINOLYSIS FOR SUPRATENTORIAL SPONTANEOUS INTRACEREBRAL HEMORRHAGES: A SINGLE-CENTER EXPERIENCE. JOURNAL OF CLINICAL MEDICINE 14 (2025). DOI: 10.3390/JCM14113636 ABSTRACT: BACKGROUND/OBJECTIVES: IN RECENT YEARS, STEREOTACTIC ASPIRATION FOLLOWED BY FIBRINOLYSIS HAS BEEN ACCEPTED AS BEING A LESS INVASIVE AND MORE EFFECTIVE TREATMENT FOR SPONTANEOUS INTRACEREBRAL HEMORRHAGE (ICH). THE AIM OF THIS STUDY WAS TO EVALUATE THE SAFETY AND CLINICAL OUTCOMES OF FRAMELESS STEREOTACTIC ASPIRATION AND FIBRINOLYSIS USING UROKINASE IN A SINGLE MEDICAL CENTER. METHODS: THIS STUDY INCLUDED 123 PATIENTS WITH SPONTANEOUS SUPRATENTORIAL ICH WHO WERE TREATED WITH STEREOTACTIC ASPIRATION AND SUBSEQUENT FIBRINOLYSIS USING UROKINASE. THEIR CLINICAL STATUS, RADIOLOGICAL IMAGES, AND FUNCTIONAL OUTCOMES WERE ASSESSED. RESULTS: UNFAVORABLE OUTCOMES AT DISCHARGE WERE ASSOCIATED WITH EACH PATIENT¿S PREOPERATIVE GLASCOW COMA SCORE, AS WELL AS THEIR INITIAL AND RESIDUAL VOLUMES OF HEMATOMA. LOW MORTALITY AND MINIMAL COMPLICATIONS OF REBLEEDING WERE ALSO RECORDED. CONCLUSIONS: THE RESULTS REVEALED THAT STEREOTACTIC ASPIRATION AND SUBSEQUENT FIBRINOLYSIS WITH UROKINASE APPEARED TO BE A SAFE AND FEASIBLE TREATMENT MODALITY FOR TREATING ICH. FURTHER STUDIES ARE STILL NEEDED IN ORDER TO BETTER ASSESS THE OPTIMAL THERAPEUTIC WINDOW, THROMBOLYTIC DOSAGE, LONG-TERM EV ALUATION, AND CONTROLLED COMPARISONS OF MORTALITY, AS WELL AS DISABILITY OUTCOMES IN TREATED AND UNTREATED PATIENTS. REPORTED EVENTS - 3 PATIENTS DIED DURING ADMISSION. SEE ATTACHED LITERATURE ARTICLE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22478561&pc=OZD | 1220648-2025-30238 | 2025-06-23 04:00 | Death | ABIOMED, INC. | 2025-07-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP HAD ACCESS SITE BLEEDING. BLOOD PRODUCTS WERE TRANSFUSED BUT THE PATIENT'S HEMODYNAMICS DECLINED AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22478739&pc=OZD | 1220648-2025-30240 | 2025-06-23 04:00 | Death | ABIOMED, INC. | 2025-07-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP DEVELOPED A CLOT IN THE PERICARDIAL SAC SPACE. THE PATIENT'S FAMILY DECLINED SURGERY AND CARE WAS WITHDRAWN. THE PATIENT SUBSEQUENTLY PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: A2, A4, AND A5 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22478801&pc=OZP | 2032227-2025-210175 | 2025-06-13 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-11 04:00 | OZP | MINIMED 770G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Renal Failure | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY AT HOME ON (B)(6) 2025. THE CAUSE OF DEATH WAS RENAL FAILURE AND DIABETES. THE CUSTOMER¿S BLOOD GLUCOSE VALUE WAS 120 MG/DL AT THE TIME OF THE INCIDENT. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: UNOMEDICAL, MMT-1880, AND MMT-332A. TROUBLESHOOTING WAS PARTIALLY PERFORMED FOR DECEASED REPORTING. THE CUSTOMER WAS WEARING AN INSULIN PUMP AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL, MMT-1880, OR MMT-332A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22479043&pc=OZD | 1220648-2025-30242 | 2025-01-16 05:00 | Death | ABIOMED, INC. | 2025-07-11 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT ON IMPELLA 5.5 SUPPORT PASSED AWAY FROM A STROKE FOLLOWING IMPELLA EXPLANT. IT WAS BELIEVED THAT A THROMBUS DISLODGED DURING EXPLANT, RESULTING IN THE NOTED STROKE AND RESULTING DEATH. NO FURTHER DETAILS WERE AVAILABLE FROM THE SITE. Manufacturer Narrative: FURTHER INFORMATION NOTED THE PHYSICIAN DOES NOT REMEMBER EXACTLY WHEN THE EVENT OCCURRED (BEST ESTIMATE PROVIDED) AND CERTAIN DETAILS WERE NOT KNOWN. PATIENT- AND PRODUCT-SPECIFIC INFORMATION AND IMPLANT/EXPLANT DATES ARE UNAVAILABLE, NOT KNOWN; THEREFORE, RESPECTIVE FIELDS REFLECT "UNKNOWN" OR LEFT BLANK ACCORDINGLY. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22479324&pc=OZD | 1220648-2025-30245 | 2025-06-24 04:00 | Death | ABIOMED, INC. | 2025-07-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem; Positioning Problem | Arrhythmia; Atrial Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED, A PATIENT UNDERGOING HIGH-RISK PERCUTANEOUS CORONARY INTERVENTION IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) EXPERIENCED ARRHYTHMIC EVENTS REQUIRING INTERVENTION; HOWEVER, THE PATIENT WOULD SUBSEQUENTLY EXPIRED. THE PATIENT WAS ADMITTED WITH COMPLAINT OF SHORTNESS OF BREATH FOR ONE WEEK AND WENT INTO PULSELESS ELECTRICAL ACTIVITY. LEFT HEART CATHETERIZATION REVEALED TRICUSPID VALVE DISEASE. THE PATIENT WAS A SURGICAL TURNDOWN AND WAS PLANNED FOR HRPCI. THE IMPELLA CP DEVICE WAS IMPLANTED AHEAD OF THE PCI THAT INCLUDED DRUG-ELUTING STENTS TO THE LEFT MAIN, LEFT ANTERIOR DESCENDING, DIAGONAL AND CIRCUMFLEX ARTERIES. DURING THE PROCEDURE, THE PATIENT EXPERIENCED LOW BLOOD PRESSURE. LEVOPHED AND MILRINONE WERE STARTED WITH THE DECISION TO LEAVE THE IMPELLA CP IMPLANTED. THE PATIENT WAS SENT TO THE UNIT ON IMPELLA MCS. OVERNIGHT, THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA AND WAS CARDIOVERTED TO NORMAL SINUS RHYTHM. THE IMPELLA CP DEVICE PERFORMANCE LEVEL WAS DECREASED TO P-3 WITH PLANS TO EXPLANT. THE PATIENT'S NATIVE CARDIAC OUTPUT WAS GRADUALLY IMPROVING, AND THE PATIENT WAS WEANING TOWARDS EXPLANT OF IMPELLA CP DEVICE; HOWEVER, THE PATIENT WAS NOTED TO BE IN ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE. A BOLUS OF AMIODARONE WAS GIVEN, AND DRIP STARTED. ADDITIONALLY, THE PATIENT WAS NOTED TO HAVE A MODERATE SIZED NOSEBLEED AND WAS GIVEN AFRIN. AN OTOLARYNGOLOGIST WAS CONSULTED. EARLIER THE NEXT MORNING, THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA, CARDIOVERTED THREE TIMES AND STARTED ON A LIDOCAINE DRIP. THE PATIENT WAS SEDATED WITH PROPOFOL AND INTUBATED, WHICH APPEARED TO BREAK THE VENTRICULAR TACHYCARDIA. HOWEVER, SHORTLY LATER IN THIS MORNING, THE PATIENT WENT BACK INTO VENTRICULAR TACHYCARDIA AND PULSELESS ELECTRICAL ACTIVITY. CARDIOPULMONARY RESUSCITATION WAS PERFORMED FOR 30 MINUTES. THE PATIENT RECEIVED MULTIPLE SHOCKS, MULTIPLE DOSES OF EPINEPHRINE AND STARTED ON EPINEPHRINE DRIP AND GIVEN MULTIPLE DOSES OF AMIODARONE AND LIDOCAINE. THE PATIENT'S HEMOGLOBIN HAD DROPPED TO 8.5 FROM 10.2 STATUS POST NOSEBLEED. WITH DEVELOPMENT OF INCESSANT VENTRICULAR ARRHYTHMIAS DESPITE EXTENSIVE EFFORTS AT RESUSCITATION, IT COULD NOT BE CONTROLLED AND ULTIMATELY THE PATIENT EXPIRED IN THE MIDST OF PROLONGED ADVANCED CARDIOPULMONARY LIFE SUPPORT ATTEMPT.Event Description: ADDITIONAL INFORMATION RECEIVED FROM LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY THAT FOLLOWING IMPELLA ASSISTED PERCUTANEOUS CORONARY INTERVENTION (PCI) PROCEDURE THE TRANSTHORACIC ECHOCARDIOGRAM (TTE) SHOWED IMPELLA WAS TOO DEEP. IMPELLA WAS REPOSITIONED. HEMODYNAMICS IMMEDIATELY IMPROVED AND IMPELLA WAS TURNED TO P3. Manufacturer Narrative: THE INVESTIGATION HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED. PER THE CLINICAL INVESTIGATION, THE ROOT CAUSE OF POSITIONING ISSUE AND VT/VF/AT/AF WAS UNABLE TO BE DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED FOR REVIEW. Manufacturer Narrative: B1 UPDATED TO REFLECT BOTH ADVERSE EVENT AND PRODUCT PROBLEM BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. B5 UPDATED WITH ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. H6 ADDITIONAL FAILURE MODE ADDED BASED ON THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. THE IMPELLA PUMP WAS NOT RETURNED FOR INVESTIGATION. SHOULD NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. G2 REPORT SOURCE; STUDY WAS MISSING FROM MANUFACTURER DEVICE REPORT 1220648-2025-30245. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITES STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22479554&pc=MVK | 3008642652-2025-06742 | 2025-06-26 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-11 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Patient-Device Incompatibility | Skin Inflammation/ Irritation | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT HAD PASSED AWAY ON (B)(6) 2025. THE PATIENT'S NURSE REPORTED THAT THE PATIENT WAS NOT WEARING THE LIFEVEST DUE TO THE SKIN BREAKDOWN ON THE EQUIPMENT ON THE LEFT UPPER SIDE QUADRANT BY THE ARMPIT. IT WAS ALSO REPORTED THAT THE PATIENT HAD STAGE 4 SKIN CANCER WHICH MAY HAVE BEEN EXACERBATED BY THE LIFEVEST. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22479756&pc=OZD | 1220648-2025-30251 | 2025-06-29 04:00 | Death | ABIOMED, INC. | 2025-07-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT EXPERIENCED AN UNCONTROLLED BLEED DURING IMPELLA CP SUPPORT. THE PATIENT SUBSEQUENTLY PASSED AWAY ON SUPPORT DUE TO MULTIPLE ORGAN FAILURE. AS A RESULT OF HEMORRHAGIC SHOCK. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22479848&pc=DSQ | 2916596-2025-04420 | 2020-01-01 05:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿THE USE OF THE MULTISENSOR HEARTLOGIC ALGORITHM FOR HEART FAILURE REMOTE MONITORING IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH ARRHYTHMIA, GASTROINTESTINAL BLEEDING, HYPOVOLEMIA, PAIN, ORTHOSTASIS, DEATH, HEART FAILURE, AND HOSPITALIZATION. THIS MULTICENTER RETROSPECTIVE STUDY EVALUATED 14 PATIENTS EQUIPPED WITH A HEARTLOGIC CAPABLE DEVICE WHO WERE IMPLANTED BETWEEN 2018 AND 2020. THE STUDY¿S OBJECTIVE WAS TO CHARACTERIZE THE USE OF HEARTLOGIC MONITORING AMONG PATIENTS IMPLANTED WITH A DURABLE LEFT VENTRICULAR ASSIST DEVICE. THE STUDY MONITORED THE REASONS FOR HOSPITALIZATIONS AND DIVIDED THEM INTO 2 GROUPS, HEART FAILURE (HF) RELATED HOSPITALIZATIONS AND NON-HF RELATED HOSPITALIZATIONS. THE SITE DEFINED HF RELATED HOSPITALIZATIONS AS PATIENTS REPORTING AT LEAST ONE SYMPTOM OF HF, DEMONSTRATING AT LEAST ONE OBJECTIVE SIGN OF DECOMPENSATED HF (IMAGING, BIOMARKER, PHYSICAL EXAMINATION, OR HEMODYNAMIC EVIDENCE OF CONGESTION), AND HAD RECEIVED AN ESCALATION IN DIURETIC THERAPY WITHIN 48 HOURS OF ADMISSION. THERE WERE 29 HOSPITALIZATIONS OVER THE STUDY PERIOD, OF WHICH 10 WERE HF RELATED HOSPITALIZATIONS, AND 19 WERE NON-HF RELATED HOSPITALIZATIONS. 7 PATIENTS EXPERIENCED ARRHYTHMIA, 5 PATIENTS EXPERIENCED AN INFECTION, 2 PATIENTS EXPERIENCED GASTROINTESTINAL BLEEDING, 2 PATIENTS EXPERIENCING HYPOVOLEMIA, 2 PATIENTS EXPERIENCED PAIN, AND 1 PATIENT EXPERIENCED ORTHOSTASIS. 3 PATIENTS PASSED AWAY. IN CONCLUSION, THE STUDY FOUND THAT HEARTLOGIC MAY HOLD PROMISE FOR PREDICTION OF NON-HF ASSOCIATED HOSPITALIZATIONS, WITH LESS CERTAINTY REGARDING HF-ASSOCIATED HOSPITALIZATIONS, IN LVAD PATIENTS. FURTHER INVESTIGATION WAS NOTED TO BE WARRANTED IN LARGER, PROSPECTIVE COHORTS. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED, AND A DIRECT CORRELATION WITH THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO HEARTMATE 3 LVAD¿S WERE RETURNED FOR EVALUATION. DEVICE HISTORY RECORDS COULD NOT BE REVIEWED DUE TO THE SERIAL NUMBER OF THE LVAD BEING UNKNOWN. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B2: SPECIFIC DEATH DATE IS DOCUMENTED AS UNKNOWN. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01JAN2020 AS DATA WAS COLLECTED BETWEEN 2018 AND 2020. ARTICLE INFORMATION: LERMAN, J. B., CYR, D. D., CHISWELL, K., TOBIN, R. S., FUDIM, M., POKORNEY, S. D., MENTZ, R. J., & SAMSKY, M. D. (2023). THE USE OF THE MULTISENSOR HEARTLOGIC ALGORITHM FOR HEART FAILURE REMOTE MONITORING IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES. ASAIO JOURNAL (AMERICAN SOCIETY FOR ARTIFICIAL INTERNAL ORGANS : 1992), 69(7), E351¿E353. HTTPS://DOI.ORG/10.1097/MAT.0000000000001946. DUKE UNIVERSITY SCHOOL OF MEDICINE, DURHAM, NORTH CAROLINA. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22479887&pc=OZD | 1220648-2025-30253 | 2025-06-23 04:00 | Death | ABIOMED, INC. | 2025-07-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP AND EXPERIENCED ACCESS SITE BLEEDING. HEPARIN WAS REMOVED FROM THE PURGE AND PRESSURE WAS HELD. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF ACCESS SITE BLEEDING IS DETERMINED TO BE PATIENT MOVEMENT BECAUSE THE BLEEDING WAS NOTED RIGHT AFTER THE PATIENT TRANSFER AND HEMOSTASIS WAS ACHIEVED BY PRESSURE DRESSING. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER. ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22479897&pc=OZD | 1220648-2025-30254 | 2025-06-21 04:00 | Death | ABIOMED, INC. | 2025-07-11 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ON IMPELLA 5.5 SUPPORT DEVELOPED BLEEDING AND HEPARIN WAS STOPPED. IT WAS FURTHER REPORTED THAT THE PLAN WAS TO EXPLANT THE IMPELLA BUT WHEN P LEVEL WAS REDUCED TO P2, THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA AND RECEIVED AMIO. THE PATIENT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: WARNINGS & CAUTIONS: CAUTIONS. ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22479908&pc=MSD | 1282497-2025-00498 | 2025-03-07 05:00 | Death | COVIDIEN LLC | 2025-07-11 04:00 | MSD | UNKNOWN DIALYSIS | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: PLI 10: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A PROSPECTIVE COHORT QUASI-EXPERIMENTAL STUDY ENROLLED A TOTAL OF 67 PATIENTS WITH ESRD (END-STAGE RENAL DISEASE) REQUIRING DIALYSIS WHO WERE RANDOMLY ASSIGNED TO EITHER PD (PERITONEAL DIALYSIS) OR HD (HEMODIALYSIS) TREATMENT REGIMENS. FROM THE POOL OF POTENTIAL PARTICIPANTS, 34 PATIENTS UNDERGOING HEMODIALYSIS AND 33 RECEIVING PERITONEAL DIALYSIS WERE IDENTIFIED. HOWEVER, SEVEN INDIVIDUALS (FOUR FROM THE HEMODIALYSIS GROUP AND THREE FROM THE PERITONEAL DIALYSIS GROUP) WERE EXCLUDED FROM THE STUDY, WHICH IS WHY THE FINAL COHORT CONSISTED OF 60 ESRD PATIENTS. ALL 60 PATIENTS HAD STAGE 5 CKD (CHRONIC KIDNEY DISEASE), AS EVIDENCED BY A CREATININE CLEARANCE OF LESS THAN 15 ML/MIN (MILLILITERS PER MINUTE), AND REQUIRED IMMEDIATE, UNSCHEDULED DIALYSIS THERAPY. COMORBIDITIES OBSERVED AMONG ALL PATIENTS ENROLLED IN THE STUDY INCLUDED DIABETES, HYPERTENSION, CARDIOVASCULAR DISEASE, CHRONIC HEART FAILURE (NEW YORK HEART ASSOCIATION [NYHA] STAGE III¿IV), CEREBROVASCULAR DISEASE, AUTOIMMUNE DISEASE, MALIGNANCY, HYPERVOLEMIA, AND UREMIA. NOTABLY, AMONG THE PD GROUP, 5 PATIENTS HAD PRE-EXISTING CHRONIC HEART FAILURE. THE PRIMARY OBJECTIVE WAS TO COMPARE MORTALITY RATES BETWEEN THE TWO GROUPS, WHILE SECONDARY OBJECTIVES INCLUDED COMPARING COMPLICATION RATES AND OVERALL HEALTHCARE COSTS. TWO-CUFF CURLED TENCKHOFF PD CATHETERS WERE INSERTED PERCUTANEOUSLY INTO 30 PATIENTS BY THE DIALYSIS UNIT NEPHROLOGISTS. TWO PD PATIENTS (6.7%) DIED. OVERALL MORTALITY RATES AFTER 3 MONTHS WERE 6.7% FOR PD. PLI 20: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A PROSPECTIVE COHORT QUASI-EXPERIMENTAL STUDY ENROLLED A TOTAL OF 67 PATIENTS WITH ESRD (END-STAGE RENAL DISEASE) REQUIRING DIALYSIS WHO WERE RANDOMLY ASSIGNED TO EITHER PD (PERITONEAL DIALYSIS) OR HD (HEMODIALYSIS) TREATMENT REGIMENS. FROM THE POOL OF POTENTIAL PARTICIPANTS, 34 PATIENTS UNDERGOING HEMODIALYSIS AND 33 RECEIVING PERITONEAL DIALYSIS WERE IDENTIFIED. HOWEVER, SEVEN INDIVIDUALS (FOUR FROM THE HEMODIALYSIS GROUP AND THREE FROM THE PERITONEAL DIALYSIS GROUP) WERE EXCLUDED FROM THE STUDY, WHICH IS WHY THE FINAL COHORT CONSISTED OF 60 ESRD PATIENTS. ALL 60 PATIENTS HAD STAGE 5 CKD (CHRONIC KIDNEY DISEASE), AS EVIDENCED BY A CREATININE CLEARANCE OF LESS THAN 15 ML/MIN (MILLILITERS PER MINUTE), AND REQUIRED IMMEDIATE, UNSCHEDULED DIALYSIS THERAPY. COMORBIDITIES OBSERVED AMONG ALL PATIENTS ENROLLED IN THE STUDY INCLUDED DIABETES, HYPERTENSION, CARDIOVASCULAR DISEASE, CHRONIC HEART FAILURE (NEW YORK HEART ASSOCIATION [NYHA] STAGE III¿IV), CEREBROVASCULAR DISEASE, AUTOIMMUNE DISEASE, MALIGNANCY, HYPERVOLEMIA, AND UREMIA. NOTABLY, AMONG THE PD GROUP, 5 PATIENTS HAD PRE-EXISTING CHRONIC HEART FAILURE. THE PRIMARY OBJECTIVE WAS TO COMPARE MORTALITY RATES BETWEEN THE TWO GROUPS, WHILE SECONDARY OBJECTIVES INCLUDED COMPARING COMPLICATION RATES AND OVERALL HEALTHCARE COSTS. TWO-CUFF CURLED TENCKHOFF PD CATHETERS WERE INSERTED PERCUTANEOUSLY INTO 30 PATIENTS BY THE DIALYSIS UNIT NEPHROLOGISTS. IN THE PD GROUP, INFECTIOUS COMPLICATIONS INCLUDED EXIT-SITE INFECTIONS IN 2 PATIENTS (6.7%) AND PERITONITIS IN 2 PATIENTS (6.7%), WHICH WAS SUCCESSFULLY TREATED WITH ANTIBIOTICS WITHOUT THE NEED FOR CATHETER REMOVAL. BACTEREMIA WAS REPORTED IN 1 PD PATIENT (3.3%). HOSPITALIZATION DAYS POST-DIALYSIS INITIATION AVERAGED 6.2 ± 3.3 FOR PD PATIENTS, WITH AN AVERAGE OF 4 RE-HOSPITALIZATIONS DURING THE OBSERVATION PERIOD. PLI 30: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A PROSPECTIVE COHORT QUASI-EXPERIMENTAL STUDY ENROLLED A TOTAL OF 67 PATIENTS WITH ESRD (END-STAGE RENAL DISEASE) REQUIRING DIALYSIS WHO WERE RANDOMLY ASSIGNED TO EITHER PD (PERITONEAL DIALYSIS) OR HD (HEMODIALYSIS) TREATMENT REGIMENS. FROM THE POOL OF POTENTIAL PARTICIPANTS, 34 PATIENTS UNDERGOING HEMODIALYSIS AND 33 RECEIVING PERITONEAL DIALYSIS WERE IDENTIFIED. HOWEVER, SEVEN INDIVIDUALS (FOUR FROM THE HEMODIALYSIS GROUP AND THREE FROM THE PERITONEAL DIALYSIS GROUP) WERE EXCLUDED FROM THE STUDY, WHICH IS WHY THE FINAL COHORT CONSISTED OF 60 ESRD PATIENTS. ALL 60 PATIENTS HAD STAGE 5 CKD (CHRONIC KIDNEY DISEASE), AS EVIDENCED BY A CREATININE CLEARANCE OF LESS THAN 15 ML/MIN (MILLILITERS PER MINUTE), AND REQUIRED IMMEDIATE, UNSCHEDULED DIALYSIS THERAPY. COMORBIDITIES OBSERVED AMONG ALL PATIENTS ENROLLED IN THE STUDY INCLUDED DIABETES, HYPERTENSION, CARDIOVASCULAR DISEASE, CHRONIC HEART FAILURE (NEW YORK HEART ASSOCIATION [NYHA] STAGE III¿IV), CEREBROVASCULAR DISEASE, AUTOIMMUNE DISEASE, MALIGNANCY, HYPERVOLEMIA, AND UREMIA. NOTABLY, AMONG THE PD GROUP, 5 PATIENTS HAD PRE-EXISTING CHRONIC HEART FAILURE. THE PRIMARY OBJECTIVE WAS TO COMPARE MORTALITY RATES BETWEEN THE TWO GROUPS, WHILE SECONDARY OBJECTIVES INCLUDED COMPARING COMPLICATION RATES AND OVERALL HEALTHCARE COSTS. TWO-CUFF CURLED TENCKHOFF PD CATHETERS WERE INSERTED PERCUTANEOUSLY INTO 30 PATIENTS BY THE DIALYSIS UNIT NEPHROLOGISTS. MECHANICAL COMPLICATIONS IN PD PATIENTS INCLUDED CATHETER TIP MIGRATION, DIALYSATE LEAKS, AND CATHETEROBSTRUCTION. IN THE PD GROUP, 4 PATIENTS (13.3%) EXPERIENCED CATHETER LEAKAGE, AND 4 PATIENTS (13.3%) EXPERIENCED CATHETER OBSTRUCTION AS THE PRIMARY MECHANICAL COMPLICATIONS AFTER 2 WEEKS OF THERAPY, BUT THESE WERE SUCCESSFULLY MANAGED WITHOUT TREATMENT INTERRUPTION. THERE WAS NO REPORTED PATIENT OUTCOME. PLI 40: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A PROSPECTIVE COHORT QUASI-EXPERIMENTAL STUDY ENROLLED A TOTAL OF 67 PATIENTS WITH ESRD (END-STAGE RENAL DISEASE) REQUIRING DIALYSIS WHO WERE RANDOMLY ASSIGNED TO EITHER PD (PERITONEAL DIALYSIS) OR HD (HEMODIALYSIS) TREATMENT REGIMENS. FROM THE POOL OF POTENTIAL PARTICIPANTS, 34 PATIENTS UNDERGOING HEMODIALYSIS AND 33 RECEIVING PERITONEAL DIALYSIS WERE IDENTIFIED. HOWEVER, SEVEN INDIVIDUALS (FOUR FROM THE HEMODIALYSIS GROUP AND THREE FROM THE PERITONEAL DIALYSIS GROUP) WERE EXCLUDED FROM THE STUDY, WHICH IS WHY THE FINAL COHORT CONSISTED OF 60 ESRD PATIENTS. ALL 60 PATIENTS HAD STAGE 5 CKD (CHRONIC KIDNEY DISEASE), AS EVIDENCED BY A CREATININE CLEARANCE OF LESS THAN 15 ML/MIN (MILLILITERS PER MINUTE), AND REQUIRED IMMEDIATE, UNSCHEDULED DIALYSIS THERAPY. COMORBIDITIES OBSERVED AMONG ALL PATIENTS ENROLLED IN THE STUDY INCLUDED DIABETES, HYPERTENSION, CARDIOVASCULAR DISEASE, CHRONIC HEART FAILURE (NEW YORK HEART ASSOCIATION [NYHA] STAGE III¿IV), CEREBROVASCULAR DISEASE, AUTOIMMUNE DISEASE, MALIGNANCY, HYPERVOLEMIA, AND UREMIA. NOTABLY, AMONG THE PD GROUP, 5 PATIENTS HAD PRE-EXISTING CHRONIC HEART FAILURE. THE PRIMARY OBJECTIVE WAS TO COMPARE MORTALITY RATES BETWEEN THE TWO GROUPS, WHILE SECONDARY OBJECTIVES INCLUDED COMPARING COMPLICATION RATES AND OVERALL HEALTHCARE COSTS. ALL CENTRAL VENOUS CATHETERS (CVCS) USED FOR HEMODIALYSIS WERE 11.5 FRENCH DOUBLE-LUMEN MAHURKAR CATHETERS. THREE HD PATIENTS (10.0%) DIED. PLI 50: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A PROSPECTIVE COHORT QUASI-EXPERIMENTAL STUDY ENROLLED A TOTAL OF 67 PATIENTS WITH ESRD (END-STAGE RENAL DISEASE) REQUIRING DIALYSIS WHO WERE RANDOMLY ASSIGNED TO EITHER PD (PERITONEAL DIALYSIS) OR HD (HEMODIALYSIS) TREATMENT REGIMENS. FROM THE POOL OF POTENTIAL PARTICIPANTS, 34 PATIENTS UNDERGOING HEMODIALYSIS AND 33 RECEIVING PERITONEAL DIALYSIS WERE IDENTIFIED. HOWEVER, SEVEN INDIVIDUALS (FOUR FROM THE HEMODIALYSIS GROUP AND THREE FROM THE PERITONEAL DIALYSIS GROUP) WERE EXCLUDED FROM THE STUDY, WHICH IS WHY THE FINAL COHORT CONSISTED OF 60 ESRD PATIENTS. ALL 60 PATIENTS HAD STAGE 5 CKD (CHRONIC KIDNEY DISEASE), AS EVIDENCED BY A CREATININE CLEARANCE OF LESS THAN 15 ML/MIN (MILLILITERS PER MINUTE), AND REQUIRED IMMEDIATE, UNSCHEDULED DIALYSIS THERAPY. COMORBIDITIES OBSERVED AMONG ALL PATIENTS ENROLLED IN THE STUDY INCLUDED DIABETES, HYPERTENSION, CARDIOVASCULAR DISEASE, CHRONIC HEART FAILURE (NEW YORK HEART ASSOCIATION [NYHA] STAGE III¿IV), CEREBROVASCULAR DISEASE, AUTOIMMUNE DISEASE, MALIGNANCY, HYPERVOLEMIA, AND UREMIA. NOTABLY, AMONG THE PD GROUP, 5 PATIENTS HAD PRE-EXISTING CHRONIC HEART FAILURE. THE PRIMARY OBJECTIVE WAS TO COMPARE MORTALITY RATES BETWEEN THE TWO GROUPS, WHILE SECONDARY OBJECTIVES INCLUDED COMPARING COMPLICATION RATES AND OVERALL HEALTHCARE COSTS. ALL CENTRAL VENOUS CATHETERS (CVCS) USED FOR HEMODIALYSIS WERE 11.5 FRENCH DOUBLE-LUMEN MAHURKAR CATHETERS. FOR HEMODIALYSIS (HD) PATIENTS, CATHETER-RELATED INFECTIONS, INCLUDING CVC-RELATED BACTEREMIA OR SEPSIS, WERE DIAGNOSED WHEN BLOOD CULTURES PERFORMED IN RESPONSE TO A SIGNIFICANT INCREASE IN C-REACTIVE PROTEIN (CRP) OR PATIENT DETERIORATION YIELDED POSITIVE RESULTS, OR WHEN BLOOD CULTURES DRAWN DIRECTLY FROM THE CATHETER WERE POSITIVE ALONG WITH CLINICAL AND LABORATORY EVIDENCE OF INFECTION AND SEPSIS. INFECTIOUS COMPLICATIONS INCLUDED EXIT-SITE INFECTIONS IN 6 PATIENTS (20%) AND BACTEREMIA IN 7 PATIENTS (23.3%) (P = 0.008). CATHETER INFECTION OCCURRED IN 6 PATIENTS (20.0%) (P <(><<)> 0.01), NECESSITATING CATHETER AND SITE REPLACEMENT IN 4 PATIENTS (13.3%). PATIENTS WERE HOSPITALIZED FOR AN AVERAGE OF 13.6 ± 4.1 DAYS POST-DIALYSIS INITIATION, WITH AN AVERAGE OF 5 RE-HOSPITALIZATIONS DURING THE OBSERVATION PERIOD. PLI 60: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A PROSPECTIVE COHORT QUASI-EXPERIMENTAL STUDY ENROLLED A TOTAL OF 67 PATIENTS WITH ESRD (END-STAGE RENAL DISEASE) REQUIRING DIALYSIS WHO WERE RANDOMLY ASSIGNED TO EITHER PD (PERITONEAL DIALYSIS) OR HD (HEMODIALYSIS) TREATMENT REGIMENS. FROM THE POOL OF POTENTIAL PARTICIPANTS, 34 PATIENTS UNDERGOING HEMODIALYSIS AND 33 RECEIVING PERITONEAL DIALYSIS WERE IDENTIFIED. HOWEVER, SEVEN INDIVIDUALS (FOUR FROM THE HEMODIALYSIS GROUP AND THREE FROM THE PERITONEAL DIALYSIS GROUP) WERE EXCLUDED FROM THE STUDY, WHICH IS WHY THE FINAL COHORT CONSISTED OF 60 ESRD PATIENTS. ALL 60 PATIENTS HAD STAGE 5 CKD (CHRONIC KIDNEY DISEASE), AS EVIDENCED BY A CREATININE CLEARANCE OF LESS THAN 15 ML/MIN (MILLILITERS PER MINUTE), AND REQUIRED IMMEDIATE, UNSCHEDULED DIALYSIS THERAPY. COMORBIDITIES OBSERVED AMONG ALL PATIENTS ENROLLED IN THE STUDY INCLUDED DIABETES, HYPERTENSION, CARDIOVASCULAR DISEASE, CHRONIC HEART FAILURE (NEW YORK HEART ASSOCIATION [NYHA] STAGE III¿IV), CEREBROVASCULAR DISEASE, AUTOIMMUNE DISEASE, MALIGNANCY, HYPERVOLEMIA, AND UREMIA. NOTABLY, AMONG THE PD GROUP, 5 PATIENTS HAD PRE-EXISTING CHRONIC HEART FAILURE. THE PRIMARY OBJECTIVE WAS TO COMPARE MORTALITY RATES BETWEEN THE TWO GROUPS, WHILE SECONDARY OBJECTIVES INCLUDED COMPARING COMPLICATION RATES AND OVERALL HEALTHCARE COSTS. ALL CENTRAL VENOUS CATHETERS (CVCS) USED FOR HEMODIALYSIS WERE 11.5 FRENCH DOUBLE-LUMEN MAHURKAR CATHETERS. IN THE HD GROUP, CATHETER OBSTRUCTION (PARTIAL OR TOTAL) OCCURRED AS THE INITIAL COMPLICATION IN 3 PATIENTS (10%) AND WAS RESOLVED THROUGH CATHETER MANIPULATION. CATHETER LEAKAGE WAS REPORTED IN 5 PATIENTS (16.7%). THERE WAS NO REPORTED PATIENT OUTCOME. URGENT UNPLANNED PERITONEAL DIALYSIS VERSUS HEMODIALYSIS IN END-STAGE RENAL DISEASE: COSTS AND OUTCOMES: ABDULLAH ALHWIESH, 2025, THERAPEUTIC APHERESIS AND DIALYSIS. Manufacturer Narrative: ABDULLAH ALHWIESH, IBRAHIEM SAEED ABDUL-RAHMAN, MOHAMMED AHMED NASR EL-DIN, MOAZ ABDULGALIL, MOHAMMED ABDELMEGEED SAKR, KHADIJA M. A LSHEHABI, NADIA AL-AUDAH URGENT UNPLANNED PERITONEAL DIALYSIS VERSUS HEMODIALYSIS IN END-STAGE RENAL DISEASE: COSTS AND OUTCOMES THER APHER DIAL. 2025;29(3):437¿46. HTTPS://DOI.ORG/10. 1111/1744-9987.70007 D10 - CONCOMITANT PRODUCT: PERITONEAL CATHETER LOT UNKNOWN PERITONEAL CATHETER LOT UNKNOWN PERITONEAL CATHETER LOT UNKNOWN DIALYSIS UNKNOWN LOT UNKNOWN DIALYSIS UNKNOWN LOT UNKNOWN MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22480043&pc=OZP | 2032227-2025-210357 | 2025-06-18 04:00 | Death | MEDTRONIC MINIMED | 2025-07-11 04:00 | OZP | UNKNOWN NGP | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER HAD PASSED AWAY AND THE CAUSE OF DEATH WAS HYPOGLYCEMIA AND HYPERGLYCEMIA. THE CUSTOMER REPORTED THE HYPERGLYCEMIA TREATED WITH INSULIN PUMP (SUBCUTANEOUS INSULIN INFUSION) AND HYPOGLYCEMIA. THE EVENT INVOLVED PRODUCTS UNK_NGP, UNOMEDICAL, MMT-332A, AND UNK_SENSOR. TROUBLESHOOTING WAS NOT PERFORMED, AND IT WAS UNKNOWN WHETHER THE CUSTOMER HAS BEEN USING THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED LOW BLOOD GLUCOSE EVENT OR NOT. IT WAS UNKNOWN WHETHER THE CUSTOMER WAS USED THE AUTO MODE FEATURE AT THE TIME OF THE EVENT OR NOT. IT WAS UNKNOWN WHETHER THE PUMP WAS WORN AT THE TIME OF PASSING OR NOT. NO PRODUCT RETURN IS REQUIRED FOR UNK_NGP, UNOMEDICAL, MMT-332A, AND UNK_SENSOR. Manufacturer Narrative: CURRENTLY, IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE, THEREFORE, CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22480127&pc=NPT | 2025587-2025-04936 | 2025-04-14 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-11 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Disseminated Intravascular Coagulation (DIC); Liver Damage/Dysfunction; Renal Failure; Sepsis; Urinary Tract Infection; Multiple Organ Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT 1 MONTH FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT'S LEVEL OF CONSCIOUSNESS SUDDENLY WORSENED. SUBSEQUENTLY, THE PATIENT WAS TRANSPORTED TO THE EMERGENCY DEPARTMENT BY AMBULANCE AND WAS ADMITTED TO THE HOSPITAL AS HEART FAILURE AND SEPSIS WERE SUSPECTED. DURING HOSPITALIZATION FOR SEPSIS DUE TO URINARY TRACT INFECTION, THE PATIENT WAS ADMINISTERED ANTIBIOTICS. THE PATIENT DEVELOPED MULTIPLE ORGAN FAILURE INCLUDING IMPAIRED CONSCIOUSNESS, RENAL DYSFUNCTION, HEPATIC DYSFUNCTION, AND DISSEMINATED INTRAVASCULAR COAGULATION (DIC); HOWEVER, THE PATIENT DIED DUE TO SEPSIS. PER THE PHYSICIAN, THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND DEVICE OR IMPLANT PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22480143&pc=DSQ | 2916596-2025-04356 | 2025-06-20 04:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE WAS NOT PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22480233&pc=DSQ | 2916596-2025-04404 | 2025-03-28 04:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE WAS UNKNOWN. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22480287&pc=LJT | 3006260740-2025-04744 | 2025-06-20 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-07-11 04:00 | LJT | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Insufficient Information | K122899 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITIGATION PROCESS THAT SOMETIME LATER PORT PLACEMENT PROCEDURE, THE PATIENT DEVELOPED WITH UNSPECIFIED INFECTION AND RESULTED IN DEATH. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. Manufacturer Narrative: H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, HENCE DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22480539&pc=DSQ | 2916596-2025-04457 | 2024-02-05 05:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH PATIENT OUTCOME THAT THE PATIENT PASSED AWAY FROM MULTI ORGAN FAILURE (MOF). THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED, THE DEVICE OPERATING AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE PATIENT'S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED, AND THE DEVICE OPERATED AS EXPECTED. IT WAS COMMUNICATED THAT THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22481181&pc=MVK | 3008642652-2025-06748 | 2025-06-13 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-11 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 10:46:59 ON (B)(6) 2025. AT 22:38:38, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS AF @ 70 BPM DEGRADING TO ASYSTOLE WITH UNCONDUCTED P WAVES. AT 22:39:29, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH UNCONDUCTED P WAVES. POST SHOCK RHYTHM WAS ASYSTOLE WITH UNCONDUCTED P WAVES. THE ELECTRODE BELT WAS DISCONNECTED AT 22:55:54 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE ELECTRODE BELT HAS BEEN COMPLETED. THE ELECTRODE BELT ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION. THE REVIEW OF THE DATA INDICATED THAT THE DEVICE POWERED ON NORMALLY AND WAS ABLE TO ACQUIRE THE PATIENT¿S ECG SIGNAL ON THE LAST DAY OF USE CAPTURED IN THE DATA DOWNLOAD. NO DEFICIENCIES ALLEGED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22481182&pc=DQO | 2024168-2025-07134 | 2025-06-26 04:00 | Death | LIGHTLAB IMAGING, INC. | 2025-07-11 04:00 | DQO | DRAGONFLY OPSTAR¿ | Difficult to Advance | Cardiac Arrest; Vascular Dissection | K192019 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE RIGHT CORONARY ARTERY WITH HEAVY CALCIFICATION, HEAVY TORTUOSITY AND 80% STENOSIS. THE DRAGONFLY OPSTAR IMAGING CATHETER WAS ADVANCED TO THE LESION WITH RESISTANCE NOTED BUT WAS ABLE TO CROSS THE LESION AND PERFORM PULLBACK. A DISSECTION WAS SUSPECTED TO BE CAUSED BY THE DRAGONFLY AND ORBITAL ATHERECTOMY WAS PERFORMED. THERE WAS A CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE DUE TO THE DISSECTION AND UNSPECIFIED TREATMENT AND THE PATIENT WAS UNSTABLE AND SYMPTOMATIC THROUGHOUT THE PROCEDURE. THE PATIENT EXPERIENCED CARDIAC ARREST AND EXPIRED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. D4 - A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22481209&pc=OZD | 1220648-2025-30261 | 2025-06-25 04:00 | Death | ABIOMED, INC. | 2025-07-11 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Pumping Stopped; Application Program Problem | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THE IMPELLA CP WAS PLACED AND THERE WAS A POSSIBLE PUMP STOP WITH A CONTROLLER FAILURE ALARM. THE AIC WAS REPLACED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN RECEIVED BUT HAS NOT YET BEEN EVALUATED. WHEN THE DEVICE INVESTIGATION HAS BEEN COMPLETED OR ANY NEW INFORMATION HAS BEEN RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA 5.5 WITH SMARTASSIST & IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: TROUBLESHOOTING THE PURGE SYSTEM: ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED: ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22481303&pc=NVN | 2182207-2025-01787 | 2025-01-01 05:00 | Death | MEDTRONIC NEUROMODULATION | 2025-07-11 04:00 | NVN | SELECTSECURE MRI SURESCAN | Appropriate Term/Code Not Available | Renal Failure; Sepsis; Heart Failure/Congestive Heart Failure | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CARDIAC RESYNCHRONIZATION THERAPY (CRT) AND LEFT BUNDLE BRANCH PACING (LBBP). THE AUTHORS DESCRIBED PATIENT DEATHS; THE CAUSES OF DEATH WERE RENAL DYSFUNCTION, HEART FAILURE, AND SEPSIS. THERE WERE PATIENTS WHO EXPERIENCED WORSENING HEART FAILURE WITH HOSPITALIZATIONS, ONE PATIENT WITH POLYMORPHIC VENTRICULAR TACHYCARDIA (VT) DUE TO HYPOKALEMIA IN THE POSTOPERATIVE PERIOD, PNEUMOTHORAX AFTER IMPLANTATION IN ONE PATIENT, AND ONE WITH INAPPROPRIATE THERAPY DUE TO T-WAVE OVERSENSING (TWOS) IN WHICH REPROGRAMMING WAS PERFORMED. THERE WERE TWO LEADS WHICH EXHIBITED LOSS OF LBBP CAPTURE DURING FOLLOW-UP WHICH WERE REPOSITIONED AND ONE LEAD WITH AN INCREASE IN THRESHOLD WHICH WAS REPROGRAMMED. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS FEMALE/60 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CARDIAC RESYNCHRONIZATION THERAPY RISK STRATIFIED BY MAGNETIC RESONANCE IMAGING AND OPTIMIZED BY LEFT BUNDLE BRANCH PACING: RESULTS FROM LONG-TERM FOLLOW-UP OF (B)(6) STUDY. HEART RHYTHM. 2025; 1¿10. DOI: 10.1016/J.HRTHM.2025.04.040. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22481339&pc=LWS | 2182207-2025-01788 | 2025-01-01 05:00 | Death | MEDTRONIC NEUROMODULATION | 2025-07-11 04:00 | LWS | SPRINT QUATTRO SECURE MRI SURESCAN | Appropriate Term/Code Not Available | Heart Failure/Congestive Heart Failure | P920015 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CARDIAC RESYNCHRONIZATION THERAPY (CRT) AND LEFT BUNDLE BRANCH PACING (LBBP). THE AUTHORS DESCRIBED PATIENT DEATHS; THE CAUSES OF DEATH WERE RENAL DYSFUNCTION, HEART FAILURE, AND SEPSIS. THERE WERE PATIENTS WHO EXPERIENCED WORSENING HEART FAILURE WITH HOSPITALIZATIONS, ONE PATIENT WITH POLYMORPHIC VENTRICULAR TACHYCARDIA (VT) DUE TO HYPOKALEMIA IN THE POSTOPERATIVE PERIOD, PNEUMOTHORAX AFTER IMPLANTATION IN ONE PATIENT, AND ONE WITH INAPPROPRIATE THERAPY DUE TO T-WAVE OVERSENSING (TWOS) IN WHICH REPROGRAMMING WAS PERFORMED. THERE WERE TWO LEADS WHICH EXHIBITED LOSS OF LBBP CAPTURE DURING FOLLOW-UP WHICH WERE REPOSITIONED AND ONE LEAD WITH AN INCREASE IN THRESHOLD WHICH WAS REPROGRAMMED. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS FEMALE/60 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CARDIAC RESYNCHRONIZATION THERAPY RISK STRATIFIED BY MAGNETIC RESONANCE IMAGING AND OPTIMIZED BY LEFT BUNDLE BRANCH PACING: RESULTS FROM LONG-TERM FOLLOW-UP OF MADURAI LBBP STUDY. HEART RHYTHM. 2025; 1¿10. DOI: 10.1016/J.HRTHM.2025.04.040. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22481350&pc=OZD | 1220648-2025-30262 | 2025-06-14 04:00 | Death | ABIOMED, INC. | 2025-07-11 04:00 | OZD | IMPELLA 5.5 | Malposition of Device | Atrial Fibrillation; Hemolysis; Tachycardia; Ventricular Fibrillation; Ectopic Heartbeat | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT (MCS). AFTER APPROXIMATELY EIGHT HOURS, THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5 FOR CONTINUATION IN THERAPY. HOWEVER, THE FIRST IMPELLA 5.5 IMPLANTED MALFUNCTIONED AND THEREFORE WAS EXPLANTED AND REPLACED. THE REPLACEMENT IMPELLA 5.5 WAS POSITIONED UNDER TRANSESOPHAGEAL ECHOCARDIOGRAPHIC GUIDANCE AND RAMPED ACHIEVING A FLOW OF 4.3 LITERS PER MINUTE. THE PATIENT WAS SENT TO THE UNIT INTUBATED AND ON IMPELLA 5.5 MCS. THE FOLLOWING MORNING, THE PATIENT REMAINED INTUBATED. HOWEVER, DID NOT REQUIRE SEDATION AT THIS TIME. AN ECHOCARDIOGRAM WAS COMPLETED, AND EJECTION FRACTION WAS NOTED AS 15-20% WITH RIGHT VENTRICULAR SYSTOLIC FUNCTION REDUCED. THE PLAN WAS TO TAKE THE PATIENT TO THE CATHETERIZATION LAB FOR LEFT AND RIGHT HEART CATHETERIZATION. OVERNIGHT THE PATIENT HAD REQUIRED ADDITIONAL BLOOD PRODUCTS, VOLUME RESUSCITATION, AND MULTIPLE AMPS OF BICARBONATE. THE PATIENT WAS IN ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE AND SUSTAINED VENTRICULAR TACHYCARDIA. THE CARDIOLOGIST GAVE MULTIPLE AMIODARONE BOLUSES WITH AMIODARONE DRIP, AND ATTEMPTED BETA BLOCKERS INTRAVENOUS. HOWEVER, RATE CONTROL WAS UNABLE TO BE ACHIEVED. THE PHYSICIAN ALSO ATTEMPTED TO CARDIOVERT THE PATIENT TWO TIMES. THE PATIENT WAS EVENTUALLY CARDIOVERTED BACK INTO SINUS RHYTHM AND AN IMPELLA RP FLEX WAS IMPLANTED (PATIENT NOW ON BIPELLA SUPPORT). IT WAS NOTED THAT SINCE PLACING THE IMPELLA RP FLEX AND CARDIOVERTING, THE PATIENT HAS BEEN ABLE TO WEAN PRESSOR REQUIREMENTS. THE PATIENT WAS CONTINUED ON BIPELLA SUPPORT PER THE NOTED PLAN. AT THIS TIME, THE PATIENT WAS SEDATED ALONG WITH INTUBATION. THE PATIENT WAS NOTED TO MOVE ALL EXTREMITIES WITH PURPOSEFUL MOVEMENT NOTED TO BILATERAL UPPER EXTREMITIES. THROUGHOUT THE DAY, THE PATIENT WAS NOTED TO HAVE ECTOPY AND LATER THAT EVENING HAD TWO RUNS OF VENTRICULAR TACHYCARDIA AND WAS ABLE TO BE SHOCKED TO NORMAL SINUS RHYTHM EACH TIME. THE IMPELLA 5.5 MEASURED 4.4 CENTIMETERS FROM THE AORTIC VALVE ANNULUS. THE LEFT VENTRICLE CHAMBER APPEARED NARROW, AND THEREFORE, THE PATIENT WAS REPOSITIONED IN THE BED TO ENSURE NO CONTACT WITH THE WALL. THE INLET TIP APPEARED CLOSE TO THE LATERAL WALL BUT TOUCHING AT THE TIME OF THE ECHOCARDIOGRAM. WHEN THE PATIENT WAS FLAT ON HIS BACK, THE IMPELLA 5.5 INLET WAS IN MID-VENTRICULAR SPACE. THE ECHOCARDIOGRAM TECHNICIAN HAD COMPLETED THE EXAM WHEN THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA AND THEN VENTRICULAR FIBRILLATION. THE PATIENT WAS SHOCKED, AND THE PATIENT REMAINED IN VENTRICULAR FIBRILLATION DESPITE THE SECOND SHOCK. MEDICATION AND FLUIDS WERE ADMINISTERED, AND THE PATIENT WAS SHOCKED BACK TO SINUS RHYTHM. THE IMPELLA 5.5 WAS SLIGHTLY ROTATED AND LOCKED AT 34.5 CENTIMETERS MARKING, MEASURING 4.2 CENTIMETERS FROM THE AORTIC VALVE ANNULUS. THE PATIENT WAS IN SINUS RHYTHM WITH NO FURTHER RUNS OF VENTRICULAR TACHYCARDIA. HOWEVER, ECTOPY WAS STILL NOTED AT TIMES. THE FOLLOWING DAY, IT WAS NOTED THE PATIENT¿S HEART RATE WAS IRREGULAR. THERE WAS ECTOPY AT TIMES. THERE WERE NO MORE VENTRICULAR TACHYCARDIA AND VENTRICULAR FIBRILLATION ISSUES AFTER THE EVENT OVERNIGHT. URINE WAS GREEN IN THE FOLEY TUBING/BAG AND URINE OUTPUT HAS BEEN LESS THAN 30 MILLILITERS PER HOUR TO NO URINE FOR THE LAST FIVE HOURS. BUMEX DRIP WAS STARTED OVERNIGHT. CREATININE CONTINUED INCREASING AND A DISCUSSION WAS MADE FOR POSSIBLY STARTING CONTINUOUS RENAL REPLACEMENT THERAPY (WHICH WAS STARTED). LABS WERE DRAWN AND WERE HEMOLYZED. THE IMPELLA 5.5 AND IMPELLA RP FLEX WERE WEANED BY ONE PERFORMANCE LEVEL. LABS WERE DRAWN AGAIN. THE NEXT DAY, THE PATIENT REMAINED ON MECHANICAL VENTILATION AND SEDATED WITH DIALYSIS RUNNING. DRESSINGS FOR BOTH IMPELLA 5.5 AND IMPELLA RP FLEX WERE NOTED CLEAN, DRY, AND INTACT. A CHEST X-RAY SHOWED THE IMPELLA RP FLEX WAS IN THE PROPER POSITION ABOVE THE PULMONIC VALVE. LABS CAME BACK HEMOLYZED. OVERNIGHT, NO ECTOPY WAS AND ONE UNIT OF BLOOD AND 500 OF ALBUMIN WERE GIVEN. ECHOCARDIOGRAM THIS DAY IN THE MORNING SHOWED LEFT VENTRICULAR EJECTION FRACTION OF AROUND 10%. THE PLAN OF CARE WAS TO CONTINUE THE PATIENT ON BIPELLA (IMPELLA 5.5 AND IMPELLA RP FLEX) SUPPORT. HOWEVER SUBSEQUENTLY, THE PATIENT¿S FAMILY WITHDREW CARE. THE IMPELLA 5.5 AND IMPELLA RP FLEX DEVICES WERE BOTH TURNED OFF ALONG WITH INTRAVENOUS DRIPS AND MECHANICAL VENTILATION. THE PATIENT THEREAFTER EXPIRED. Manufacturer Narrative: MEDICAL SAFETY OFFICER REVIEWED AND DETERMINED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA 5.5 (SERIAL NUMBER (B)(6)) AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THE IS ONE OF TWO REPORTS ASSOCIATED WITH THE PATIENT'S IMPLANT EXPERIENCE. REFER TO MANUFACTURING REPORT NUMBER 1220648-2025-29934 FOR THE REPORT ON THE INITIALLY IMPLANTED IMPELLA 5.5 DEVICE SERIAL NUMBER (B)(6). THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION FOR THE HEMOLYSIS, ECTOPY, AND POSITIONING ISSUES HAS BEEN COMPLETED. CLINICAL REPORTS RUNS OF VT (VENTRICULAR TACHYCARDIA), ECTOPY, ATRIAL FIBRILLATION (AF), AND VENTRICULAR FIBRILLATION (VF) THROUGHOUT THE CASE. PUMP APPEARED TO BE TOUCHING LATERAL WALL OF LV DURING SOME ECHOS AND WHILE THE PATIENT WAS IN DIFFERENT POSITIONS IN BED. AMIO BOLUSES, BETA BLOCKERS, AND MAGNESIUM WERE GIVEN, CARDIOVERSION WAS ATTEMPTED, AND PATIENT WAS SHOCKED MULTIPLE TIMES THROUGHOUT THE CASE. DATA LOGS ARE NOT APPLICABLE AND PRODUCT WAS NOT RETURNED. THE ROOT CAUSE OF THE VT, ECTOPY, AF, AND VF WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. CLINICAL DETAILS NOTE THAT THE PHYSICIAN CHECKED PUMP POSITIONING VIA ECHO DUE TO RUNS OF VT. THE 5.5 MEASURED 4.4 CM FROM AV ANNULUS. THE LV CHAMBER APPEARED NARROW AND THEREFORE, PATIENT WAS REPOSITIONED IN BED TO ENSURE NO CONTACT WITH WALL. INLET TIP APPEARED CLOSE TO LATERAL WALL BUT WAS TOUCHING AT THE TIME OF ECHO. WHEN PATIENT IS FLAT ON BACK, THE INLET IS MID-VENTRICULAR. THE 5.5 WAS SLIGHTLY ROTATED AND MEASURED 4.2 CM FROM AV ANNULUS. ANOTHER ECHO THE NEXT DAY SHOWED INLET MEASURING 4.0 CM FROM AV ANNULUS. DATA LOG REVIEW SHOWED NO MAJOR POSITIONING ISSUES EXCEPT POSITION UNKNOWN ALARMS DURING SUCTION. THE ROOT CAUSE OF THE POSITIONING ISSUES WAS MOST LIKELY PATIENT CONDITION RELATED AS THE CLINICAL NOTED THAT THE PATIENT'S LV WAS NARROW AND THE PATIENT'S POSITION IN BED AFFECTED THE PUMP'S PROXIMITY TO THE LV WALL. CLINICAL DETAILS STATE THAT LABS WERE HEMOLYZED AND IT IS UNKNOWN IF THE HEMOLYSIS RESOLVED. EVENTUALLY CARE WAS WITHDRAWN. DATA LOG REVIEW SHOWED SUCTION WITH POSITION UNKNOWN ALARMS. THERE WAS ALSO A SLIGHT DOWNWARD TREND IN PS. THE ROOT CAUSE OF THE HEMOLYSIS WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND IT IS UNKNOWN IF/WHEN THE HEMOLYSIS RESOLVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22481758&pc=DSQ | 2916596-2025-04459 | 2024-03-08 05:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIPLE ORGAN FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE PATIENT¿S MULTISYSTEM ORGAN FAILURE; HOWEVER, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: H6- 3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22482043&pc=DSQ | 2916596-2025-04660 | 2025-06-16 04:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | THORATEC® HEARTMATE 3¿¿ LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Renal Failure; Electrolyte Imbalance | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS SEEN IN THE CLINIC ON (B)(6) 2025 AND HAD WORSENING FUNCTIONAL STATUS OVER THE PAST FEW WEEKS. THERE WERE CONCERNING LABS WHICH NOTED POTASSIUM LEVEL AT 2.3 MILLIEQUIVALENTS PER LITER. THE PATIENT WAS GIVEN POTASSIUM CHLORIDE (K-DUR) 40 MILLIEQUIVALENTS (MEQ) IN CLINIC. THE PATIENT WAS ADMITTED FOR INFECTIOUS WORKUP VS FAILURE TO THRIVE. WORKUP FOR REVERSABLE CAUSES WAS UNSUCCESSFUL. THE PATIENT'S FAMILY HAD ELECTED TO GO HOME WITH HOSPICE. THEY WERE AWARE THEY COULD REQUEST THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) TO BE TURNED OFF. AT THE TIME, THEY DID NOT WANT THE IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR (ICD) OR LVAD TURNED OFF. THE FAMILY SAID THEY WOULD CALL THE ICD REPRESENTATIVE WHEN THEY GOT HOME. THE PATIENT PASSED AWAY ON (B)(6) 2025 ON HOSPICE. THE CAUSE OF DEATH WAS FAILURE TO THRIVE AND RENAL FAILURE. THE OUTCOME WAS CONSIDERED NOT PUMP RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE PATIENT¿S INFECTION COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. FURTHER, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, AND PUMP POCKET), RENAL DYSFUNCTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. FURTHERMORE, SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22482321&pc=DSQ | 2916596-2025-04576 | 2025-06-25 04:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AND HAD A HISTORY OF RECENT MITRAL VALVE REPLACEMENT (MVR), STENTS, AND THROMBUS ON THE MITRAL VALVE (MV) PRE-IMPLANT. A CENTRIMAG RIGHT VENTRICULAR ASSIST DEVICE (RVAD) VIA A THIRD-PARTY CATHETER (PROTEK DUO) WAS PLACED DUE TO RIGHT VENTRICULAR (RV) DYSFUNCTION. ON (B)(6) 2025, THE PATIENT PASSED AWAY DUE TO A LARGE MID-CEREBRAL ARTERY (MCA) STROKE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22483000&pc=DSQ | 2916596-2025-04629 | 2025-05-01 04:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Abscess; Bacterial Infection; Purulent Discharge; Renal Failure; Unspecified Tissue Injury | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD BEEN DISCHARGED ON DIALYSIS IN (B)(6). THE PATIENT HAD DEVELOPED END STAGE RENAL DISEASE WITH HEMODIALYSIS POST OPERATION. IN (B)(6), THE PATIENT WAS ADMITTED DUE TO A CHEST ABSCESS AND UNDERWENT AN INCISION AND DRAINAGE (I&D) OF THEIR LEFT CHEST. THE I&D REVEALED A LEFT CHEST WALL WOUND WITH PURULENT DRAINAGE AND FRIABLE NECROTIC GRANULATION TISSUE THAT TRACKED DEEP SUPERIOR INTO LIKELY PRIOR LEFT CHEST TUBE SITE INTO THE PLEURAL CAVITY. THE INFERIOR STERNAL WOUND WAS JUST IMMEDIATELY BELOW THE XIPHOID WITH PURULENT DRAINAGE AND FRIABLE NECROTIC GRANULATION TISSUE. THE WOUND TRACKED DEEP TO MEDIASTINUM. ALTHOUGH THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) PUMP COMPONENTS WERE NOT VISUALIZED, IT WAS LIKELY THE PATIENT HAD AN LVAD DEVICE INFECTION DUE TO THE LOCATION/PATHS OF THE WOUNDS. TWO WOUND VACUUMS WERE APPLIED. BONE WAS NOT INVOLVED. THE PATIENT THEN NEEDED A PERIPHERALLY INSERTED CENTRAL CATHETER (PICC) LINE AND LONG TERM INTRAVENOUS (IV) ANTIBIOTIC (CEFIPIME) DUE TO PSEUDOMONAS. THE PATIENT WAS ABLE TO COME OFF DIALYSIS WITH IMPROVEMENT OF KIDNEY FUNCTION AND BEGAN MAKING ADEQUATE URINE. PERMCATH WAS REMOVED AND THE PATIENT WAS DISCHARGED WITH PICC, IV ANTIBIOTICS, AND WOUND VACUUMS. THE PATIENT'S MOBILITY HAD SEVERELY DECLINED SINCE (B)(6) AND THE FURTHEST THEY COULD WALK WAS A FEW STEPS TO THE SCALE. IN (B)(6), THE PATIENT WAS TRANSITIONED TO PALLIATIVE CARE DUE TO LACK OF QUALITY OF LIFE RELATED TO ILLNESS. DESPITE THE PATIENT MAKING URINE, THEIR KIDNEY FUNCTION BEGAN TO DECLINE, AND IT WAS DETERMINED THAT DIALYSIS WAS FUTILE AND WAS NOT RESTARTED. THE PATIENT'S KIDNEY FUNCTION CONTINUED TO WORSEN. PALLIATIVE CARE WAS CONSULTED, AND THE PATIENT WAS MOVED TO DO NOT RESUSCITATE/DO NOT INTUBATE (DNR/DNI) AND HOSPICE CARE. IT WAS NOTED THAT THE PATIENT'S CARE WAS ALSO COMPLICATED BY THE PATIENT'S SPOUSE'S RESISTANCE TO THE PATIENT COMING HOME ON DIALYSIS. THE PATIENT PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: SECTION A4: PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22483345&pc=OZD | 1220648-2025-30271 | 2025-06-23 04:00 | Death | ABIOMED, INC. | 2025-07-11 04:00 | OZD | IMPELLA 5.5 | Malposition of Device | Hemorrhage/Bleeding; Cardiac Perforation | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPELLA 5.5 WAS IMPLANTED IN A PATIENT AND WAS BENT IN THE WRONG DIRECTION. THE PHYSICIAN ATTEMPTED TO REPOSITION THE PUMP AND A LEFT VENTRICLE PERFORATION WAS NOTED. THE VENTRICLE WAS REPAIRED. LATER, THE PATIENT EXHIBITED BLEEDING FROM THE CHEST TUBES AND A WASHOUT WAS PERFORMED. FIVE UNITS OF PACKED RED BLOOD CELLS AND THREE UNITS OF FRESH FROZEN PLASMA WERE PLANNED TO BE GIVEN. TWO DAYS LATER, THE PLAN WAS FOR THE PATIENT TO GO TO THE OPERATING ROOM FOR A WASHOUT. THE FOLLOWING DAY, THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM TO FIX THE LEFT VENTRICLE PERFORATION THAT WAS NOT PROPERLY FIXED. ULTIMATELY, CARE WAS WITHDRAWN THAT DAY AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS. ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION FOR THE VENTRICLE PERFORATION AND THE POSITIONING ISSUE HAS BEEN COMPLETED. CLINICAL DETAILS REPORT THAT THE PUMP WAS NOTED TO BE ANGLED IN THE WRONG DIRECTION, SO THEY ATTEMPTED TO ROTATE/REPOSITION. DURING THIS EVENT, THE VENTRICLE WAS PERFORATED. THE PUMP WAS PUT IN SURGICAL MODE AND THE PERFORATION WAS REPAIRED. IT WAS ALSO REPORTED THAT THE PUMP WAS INITIALLY POSITIONED SHALLOW INTENTIONALLY BECAUSE THE PATIENT HAD A SMALL LV. DATA LOGS WERE REVIEWED AND JUST AFTER THE PUMP WAS STARTED, THE POSITION IN AORTA ALARM TRIGGERED FOR ~20 SECONDS. TWO MINUTES LATER, THE PUMP WAS PUT IN SURGICAL MODE. THIS POSITION IN AORTA ALARM LIKELY OCCURRED WHILE THE PHYSICIAN WAS ATTEMPTING TO REPOSITION/MANIPULATE THE PUMP, BASED ON HOW DATA LOGS ALIGN WITH THE CLINICAL TIMELINE REPORTED. PRODUCT WAS NOT RETURNED FOR INVESTIGATION. BASED ON THE CLINICAL DETAIL THAT THE PUMP WAS INTENTIONALLY POSITIONED SHALLOW WHICH LIKELY THEN MADE IT MORE DIFFICULT TO REPOSITION THE PUMP, THE MOST LIKELY ROOT CAUSE OF THE VENTRICLE PERFORATION IS PATIENT CONDITION, FOR THE POSITIONING ISSUES, CLINICAL DETAILS REPORT THAT THE PUMP WAS INTENTIONALLY POSITIONED SHALLOW DUE TO THE PATIENT'S ANATOMY (SMALL LV). SHORTLY AFTER INITIAL POSITIONING, IT WAS NOTED THAT THE PUMP WAS ANGLED IN THE IMPROPER DIRECTION, SO THE PHYSICIAN ATTEMPTED TO REPOSITION. DURING THIS REPOSITIONING ATTEMPT, THE VENTRICLE WAS PERFORATED, AND AFTER THE PERFORATION WAS REPAIRED, THE PHYSICIAN ELECTED TO INTENTIONALLY POSITION THE PUMP SHALLOW AGAIN DUE TO THE PATIENT'S ANATOMY. DATA LOGS WERE REVIEWED AND JUST AFTER THE PUMP WAS STARTED, THE POSITION IN AORTA ALARM TRIGGERED FOR ~20 SECONDS. TWO MINUTES LATER, THE PUMP WAS PUT IN SURGICAL MODE. THIS POSITION IN AORTA ALARM LIKELY OCCURRED WHILE THE PHYSICIAN WAS ATTEMPTING TO REPOSITION/MANIPULATE THE PUMP, BASED ON HOW DATA LOGS ALIGN WITH THE CLINICAL TIMELINE REPORTED. PRODUCT WAS NOT RETURNED FOR INVESTIGATION. BASED ON THE CLINICAL DETAILS PROVIDED, THERE WAS DIFFICULTY INITIALLY POSITIONING THE PUMP AND SUBSEQUENTLY REPOSITIONING THE PUMP DUE TO THE PATIENT'S SMALL LV. THEREFORE, THE ROOT CAUSE OF THE POSITIONING ISSUES WAS MOST LIKELY PATIENT CONDITION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22483649&pc=DSQ | 2916596-2025-04365 | 2024-01-15 05:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhagic Stroke | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY SECONDARY TO AN ACUTE SUBDURAL HEMATOMA. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22483992&pc=NKM | 2135147-2025-03804 | 2021-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-11 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "DIASTOLOGY AND MITRACLIP® OUTCOMES: MULTICENTER REAL-WORLD EVIDENCE STUDY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE ASSOCIATION BETWEEN THE PRE-PROCEDURAL ECHOCARDIOGRAPHIC PARAMETERS OF DIASTOLIC FUNCTION (DF) AND ONE-YEAR OUTCOMES AFTER MC PLACEMENT. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT PATIENTS THAT FAILED TO REACH THE PEE HAD A LOWER PRE-PROCEDURAL MV DECELERATION TIME OF 176.88 MS (164.14¿189.62); HOWEVER, NO ASSOCIATION WAS OBSERVED BETWEEN MV DECELERATION TIME AND A FAILURE TO REACH THE PEE IN THE MULTIVARIATE REGRESSION ANALYSIS. THE PRE-PROCEDURAL E/A RATIO HAD A SIGNIFICANT ASSOCIATION WITH SYMPTOM WORSENING AFTER MULTIVARIATE REGRESSION ANALYSIS: OR: 1.61 (1.13¿2.29). THE AUC FOR THE E/A RATIO IN THE SYMPTOM-WORSENING GROUP WAS 0.603, MAKING IT A MORE MODERATE PREDICTOR THAN RANDOM GUESSING FOR THE FAILURE TO REACH THE PEE. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS VIVEK JOSEPH VARUGHESE AT PRISMA HEALTH INSTITUTION WITH CORRESPONDING EMAIL VIVEKJVARUGHESE@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS JANUARY 2021 TO MARCH 2024. A TOTAL OF 224 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 224 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HISTORY OF ATRIAL FIBRILLATION AND MYOCARDIAL INFARCTION. (B)(6), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HISTORY OF ATRIAL FIBRILLATION AND MYOCARDIAL INFARCTION. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22484286&pc=DSQ | 2916596-2025-04366 | 2025-06-24 04:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | THORATEC® HEARTMATE 3¿¿ LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22484389&pc=OZD | 1220648-2025-30276 | 2025-06-25 04:00 | Death | ABIOMED, INC. | 2025-07-11 04:00 | OZD | IMPELLA 5.5 | Malposition of Device | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA 5.5 SUPPORT SHOWED A THALAMIC INFARCT ON HEAD CT. THE PATIENT UNDERWENT AN EMBOLECTOMY. THE PUMP WAS POSITIONED IN THE PAPILLARY MUSCLE AND THE PATIENT UNDERWENT A WASHOUT AND REPOSITIONING OF THE PUMP. THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF STROKE/CVA, POSITIONING ISSUES, AND VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: A5 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: & IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22484505&pc=OZD | 1220648-2025-30277 | 2025-06-24 04:00 | Death | ABIOMED, INC. | 2025-07-11 04:00 | OZD | IMPELLA 5.5 | Failure to Advance | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT HAD AN ATTEMPTED IMPELLA 5.5 IMPLANT AND THE PUMP WAS UNABLE TO ADVANCE INTO THE LEFT VENTRICLE (LV). THE DOCTOR REMOVED THE IMPELLA 5.5 AND WAS GOING TO REATTEMPT ACCESS USING THE IMPELLA WIRE. DURING THIS TIME THE PATIENT'S EXTRACORPOREAL MEMBRANEOXYGENATION (ECMO) FLOWS WERE NOT SUSTAINING GREATER THAN ONE LITER. ECHOCARDIOGRAM SHOWED NO VOLUME IN THE HEART. THE PHYSICIAN DISCUSSED OPTIONS WITH PERFUSION AND ANESTHESIA REGARDING ADDITIONAL VOLUME AND VOLUME REQUIREMENTS. MEAN ARTERIAL PRESSURE HAD DROPPED SIGNIFICANTLY AND ECMO FLOWS REMAINED LOW, DESPITE VOLUME AND PRESSERS. BLOOD PRESSURE CONTINUED TO DROP. THE PATIENT EXPIRED AT THIS TIME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22484540&pc=NPT | 2135147-2025-03816 | 2013-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-11 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P190023 |  | 1 | Event Description: THE ARTICLE, "SELF-EXPANDABLE TRANSCATHETER AORTIC VALVES IN PATIENTS WITH SMALL AORTIC ANNULUS: THE SWEDEHEART REGISTRY", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, MULTICENTER STUDY TO COMPARE AN INTRA-ANNULAR SELF-EXPANDING VALVES (IA SEVS) TRANSCATHETER HEART VALVES (THVS) WITH CONTEMPORARY SUPRA-ANNULAR SELF-EXPANDING VALVES (SA SEVS) REGARDING MORTALITY OUTCOMES, DEVICE SUCCESS, TECHNICAL SUCCESS, AND HEMODYNAMIC PERFORMANCE. DEVICES INCLUDED IN THE STUDY WERE EVOLUT, ACURATE, PORTICO, AND NAVITOR. THE ARTICLE CONCLUDED THAT SA AND IA SEVS SHOW COMPARABLE CLINICAL OUTCOMES AND EXCELLENT HEMODYNAMIC PERFORMANCE IN PATIENTS WITH SMALL AORTIC ANNULI. FURTHER STUDIES, INCLUDING RANDOMIZED TRIALS, ARE NEEDED TO PROVIDE CLEARER GUIDANCE ON VALVE SELECTION. [THE PRIMARY AND CORRESPONDING AUTHOR WAS ANTROS LOUCA, UNIVERSITY OF GOTHENBURG, DEPARTMENT OF MOLECULAR AND CLINICAL MEDICINE 41 345, GOTHENBURG, SWEDEN, WITH CORRESPONDING EMAIL: ANTROS.LOUCA@VGREGION.SE] THE TIME FRAME OF THE STUDY WAS FROM 01 JANUARY 2013 TO 31 DECEMBER 2022. A TOTAL OF 1068 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 102 (9.6%) RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 81.2 YEARS AND THE MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES MELLITUS, CHRONIC KIDNEY DISEASE, AORTIC/MITRAL REGURGITATION, BICUSPID AORTIC VALVE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF SELF-EXPANDABLE TRANSCATHETER AORTIC VALVES IN PATIENTS WITH SMALL AORTIC ANNULUS: THE SWEDEHEART REGISTRY WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES MELLITUS, CHRONIC KIDNEY DISEASE, AORTIC/MITRAL REGURGITATION, BICUSPID AORTIC VALVE. SOME OF THE COMPLICATIONS REPORTED WERE AORTIC VALVE STENOSIS, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: SELF-EXPANDABLE TRANSCATHETER AORTIC VALVES IN PATIENTS WITH SMALL AORTIC ANNULUS: THE SWEDEHEART REGISTRY. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: SELF-EXPANDABLE TRANSCATHETER AORTIC VALVES IN PATIENTS WITH SMALL AORTIC ANNULUS: THE SWEDEHEART REGISTRY
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22484560&pc=DSQ | 2916596-2025-04411 | 2024-07-07 04:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, TO DATE, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22484570&pc=NIK | 9614453-2025-02680 | 2025-06-10 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-07-11 04:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Fever; Low Blood Pressure/ Hypotension; Respiratory Failure | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT FOR ALERTED MENTAL STATE. THE PATIENT WAS SUBSEQUENTLY SEDATED AND INTUBATED DUE TO ACUTE HYPOXIC RESPIRATORY FAILURE AND ANTIARRHYTHMIC MEDICATIONS WERE ADMINISTERED. THE PATIENT WAS NOTED WITH FEVER, TACHYCARDIC, HYPOTENSIVE WITH HIGH BLOOD CELL COUNT AND CHEST COMPUTED TOMOGRAPHY (CT SCAN) SHOWED BIBASILAR ATELECTASIS. THE PATIENT WAS STARTED ON ANTIBIOTICS AND VASOPRESSORS. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) PERFORMED SHOWED RIGHT VENTRICULAR (RV) LEAD THROMBUS. ADDITIONALLY, IT WAS NOTED THAT THE PATIENT REQUIRED INCREASING AMOUNTS OF VASOPRESSORS DUE TO LOW BLOOD PRESSURE. AFTER A HIGH DECLINE IN PATIENTS¿ CLINICAL STATUS, THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WAS EXPLANTED DUE TO SEPSIS CONCERNS. CULTURES WERE PERFORMED AND WERE POSITIVE FOR METHICILLIN SENSITIVE STAPHYLOCOCCUS AUREUS (MSSA) BACTEREMIA. THE PATIENT¿S VASOPRESSOR REQUIREMENTS DECREASED AFTER REMOVAL OF THE CRT-D SYSTEM. THE PATIENT¿S CLINICAL COURSE CONTINUED TO DECLINE, AND THE PATIENT¿S FAMILY CONFIRMED DO NOT RESUSCITATE (DNR) STATUS AND OPTED FOR COMFORT CARE. THE PATIENT WAS EXTUBATED AND PASSED AWAY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. THE DEVICE WAS RETURNED AND SCREENING ANALYSIS WAS PERFORMED, BUT NO ISSUE WAS IDENTIFIED REQUIRING FULL ANALYSIS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22484581&pc=DTB | 2649622-2025-18797 | 2025-06-10 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-11 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Fever; Low Blood Pressure/ Hypotension; Respiratory Failure | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT FOR ALERTED MENTAL STATE. THE PATIENT WAS SUBSEQUENTLY SEDATED AND INTUBATED DUE TO ACUTE HYPOXIC RESPIRATORY FAILURE AND ANTIARRHYTHMIC MEDICATIONS WERE ADMINISTERED. THE PATIENT WAS NOTED WITH FEVER, TACHYCARDIC, HYPOTENSIVE WITH HIGH BLOOD CELL COUNT AND CHEST COMPUTED TOMOGRAPHY (CT SCAN) SHOWED BIBASILAR ATELECTASIS. THE PATIENT WAS STARTED ON ANTIBIOTICS AND VASOPRESSORS. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) PERFORMED SHOWED RIGHT VENTRICULAR (RV) LEAD THROMBUS. ADDITIONALLY, IT WAS NOTED THAT THE PATIENT REQUIRED INCREASING AMOUNTS OF VASOPRESSORS DUE TO LOW BLOOD PRESSURE. AFTER A HIGH DECLINE IN PATIENTS¿ CLINICAL STATUS, THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WAS EXPLANTED DUE TO SEPSIS CONCERNS. CULTURES WERE PERFORMED AND WERE POSITIVE FOR METHICILLIN SENSITIVE STAPHYLOCOCCUS AUREUS (MSSA) BACTEREMIA. THE PATIENT¿S VASOPRESSOR REQUIREMENTS DECREASED AFTER REMOVAL OF THE CRT-D SYSTEM. THE PATIENT¿S CLINICAL COURSE CONTINUED TO DECLINE, AND THE PATIENT¿S FAMILY CONFIRMED DO NOT RESUSCITATE (DNR) STATUS AND OPTED FOR COMFORT CARE. THE PATIENT WAS EXTUBATED AND PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22484584&pc=DTB | 2649622-2025-18798 | 2025-06-14 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-11 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Under-Sensing | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT REVIEW OF A REMOTE TRANSMISSION SHOWED APPARENT RIGHT ATRIAL (RA) LEAD UNDERSENSING TRIGGERING DEVICE-CLASSIFIED FALSE TERMINATION OF ATRIAL TACHYCARDIA/ATRIAL FIBRILLATION (AT/AF) EVENTS AT TIMES. THE REMOTE TRANSMISSION ALSO NOTED PRIOR PHRENIC NERVE STIMULATION FROM THE LEFT VENTRICULAR (LV) LEAD. INFORMATION WAS RECEIVED THAT THE PATIENT WAS DECEASED FOUR DAYS LATER. A CAUSE OF DEATH WAS UNABLE TO BE OBTAINED AT THE TIME OF SUBMISSION. Manufacturer Narrative: CONTINUATION OF D10: 5076-58 LEAD AND 5076-52 LEAD IMPLANTED: (B)(6) 2016. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22484586&pc=DSQ | 2916596-2025-04519 | 2025-04-28 04:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO THE VENTRICULAR ASSIST DEVICE BEING DEACTIVATED WHILE ON COMFORT CARE. THE OUTCOME WAS CONSIDERED DEVICE OR THERAPY RELATED AND WAS DUE TO RIGHT VENTRICULAR FAILURE. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22484596&pc=DSQ | 2916596-2025-04472 | 2025-03-26 04:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22484597&pc=LWS | 2649622-2025-18799 | 2025-06-10 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-11 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Fever; Low Blood Pressure/ Hypotension; Respiratory Failure; Thrombosis/Thrombus | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT FOR ALERT MENTAL STATE. THE PATIENT WAS SUBSEQUENTLY SE DATED AND INTUBATED DUE TO ACUTE HYPOXIC RESPIRATORY FAILURE AND ANTIARRHYTHMIC MEDICATIONS WERE ADMINISTERED. THE PATIENT WAS NOTED WITH FEVER, TACHYCARDIC, HYPOTENSIVE WITH HIGH BLOOD CELL COUNT; AND CHEST COMPUTED TOMOGRAPHY (CT SCAN) SHOWED BIBASILAR ATELECTASIS. THE PATIENT WAS STARTED ON ANTIBIOTICS AND VASOPRESSORS. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) PERFORMED SHOWED RIGHT VENTRICULAR (RV) LEAD THROMBUS. ADDITIONALLY, IT WAS NOTED THAT THE PATIENT REQUIRED INCREASING AMOUNTS OF VASOPRESSORS DUE TO LOW BLOOD PRESSURE. AFTER A HIGH DECLINE IN PATIENT'S CLINICAL STATUS, THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WAS EXPLANTED DUE TO SEPSIS CONCERNS. CULTURES WERE PERFORMED AND WERE POSITIVE FOR METHICILLIN SENSITIVE STAPHYLOCOCCUS AUREUS (MSSA) BACTEREMIA. THE PATIENT¿S VASOPRESSOR REQUIREMENTS DECREASED AFTER REMOVAL OF THE CRT-D SYSTEM. THE PATIENT¿S CLINICAL COURSE CONTINUED TO DECLINE, AND THE PATIENT¿S FAMILY CONFIRMED DO NOT RESUSCITATE (DNR) STATUS AND OPTED FOR COMFORT CARE. THE PATIENT WAS EXTUBATED AND PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22484606&pc=OJX | 2649622-2025-18800 | 2025-06-10 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-11 04:00 | OJX | ATTAIN PERFORMA STRAIGHT MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Fever; Low Blood Pressure/ Hypotension; Respiratory Failure | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT FOR ALERTED MENTAL STATE. THE PATIENT WAS SUBSEQUENTLY SE DATED AND INTUBATED DUE TO ACUTE HYPOXIC RESPIRATORY FAILURE AND ANTIARRHYTHMIC MEDICATIONS WERE ADMINISTERED. THE PATIENT WAS NOTED WITH FEVER, TACHYCARDIC, HYPOTENSIVE WITH HIGH BLOOD CELL COUNT AND CHEST COMPUTED TOMOGRAPHY (CT SCAN) SHOWED BIBASILAR ATELECTASIS. THE PATIENT WAS STARTED ON ANTIBIOTICS AND VASOPRESSORS. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) PERFORMED SHOWED RIGHT VENTRICULAR (RV) LEAD THROMBUS. ADDITIONALLY, IT WAS NOTED THAT THE PATIENT REQUIRED INCREASING AMOUNTS OF VASOPRESSORS DUE TO LOW BLOOD PRESSURE. AFTER A HIGH DECLINE IN PATIENTS¿ CLINICAL STATUS, THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WAS EXPLANTED DUE TO SEPSIS CONCERNS. CULTURES WERE PERFORMED AND WERE POSITIVE FOR METHICILLIN SENSITIVE STAPHYLOCOCCUS AUREUS (MSSA) BACTEREMIA. THE PATIENT¿S VASOPRESSOR REQUIREMENTS DECREASED AFTER REMOVAL OF THE CRT-D SYSTEM. THE PATIENT¿S CLINICAL COURSE CONTINUED TO DECLINE, AND THE PATIENT¿S FAMILY CONFIRMED DO NOT RESUSCITATE (DNR) STATUS AND OPTED FOR COMFORT CARE. THE PATIENT WAS EXTUBATED AND PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22484930&pc=OZD | 1220648-2025-30280 | 2025-07-01 04:00 | Death | ABIOMED, INC. | 2025-07-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS NOTED THAT THE PATIENT HAD INTERMITTENT EPISODES OF CARDIAC ARREST PRIOR TO IMPELLA PLACEMENT. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT IN CRITICAL CONDITION. HE REMAINED IN PERSISTENT VENTRICULAR TACHYCARDIA (VT) AND VENTRICULAR FIBRILLATION (VF). ALTHOUGH HE WAS SHOCKED AND BRIEFLY CONVERTED, HE REPEATEDLY REVERTED TO VT/VF. FOLLOWING CLINICAL DETERIORATION, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22485102&pc=DSQ | 2916596-2025-04460 | 2025-06-20 04:00 | Death | THORATEC CORPORATION | 2025-07-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED, AND NO ADDITIONAL INFORMATION WAS PROVIDED AS TO WHY. AN AUTOPSY WAS NOT PERFORMED. THE PUMP WAS NOT EXPLANTED.Event Description: ADDITIONAL INFORMATION RECEIVED STATED THAT THERE WAS SEPSIS THAT LED UP TO DEATH. THERE WAS ALSO A SUSPECTED DRUG OVERDOSE. THE DEVICE HAD OPERATED AS EXPECTED WITH NO ALARMS. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22485439&pc=QZI | 3012520654-2025-00092 | 2025-07-02 04:00 | Death | MEDTRONIC, INC. | 2025-07-11 04:00 | QZI | AFFERA | Adverse Event Without Identified Device or Use Problem | Laceration(s); Myocardial Infarction; Sepsis; Tachycardia; Ventricular Fibrillation; Cardiac Tamponade; Cardiogenic Shock; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P240013 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THIS STUDY AIMED TO ASSESS FEASIBILITY, SAFETY PROFILE AS WELL AS PATIENT OUTCOMES AFTER VENTRICUL AR ARRHYTHMIA (VA) ABLATION WITH A LATTICE-TIP CATHETER IN A MULTICENTER EUROPEAN REGISTRY. THE AUTHORS DESCRIBED FOUR PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE MULTI-ORGAN FAILURE (2), SEPSIS(2) AND VT STORM (1). ONE PATIENT DIED IN INTENSIVE CARE UNIT (I CU). THERE WERE PATIENTS WHO EXPERIENCED INTRAOPERATIVE COMPLICATIONS WHICH INCLUDED TWO MAJOR BLEEDING (ONE REQUIRING BLOOD TRANSFUSION AND ONE REQUIRING ADDITIONAL VASCULAR SURGERY), ONE TAMPONADE DUE TO EPICARDIAL ACCESS WITH CORONARY LACERATION AND MYOCARDIAL INFARCTION, ONE CARDIOGENIC SHOCK (REQUIRED CARDIOPULMONARY RESUSCITATION (CPR) AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). THIS PATIENT EVENTUALLY RECOVERED AND WAS DISCHARGED.), TWO THROMBO-EMBOLIC EVENTS - STROKE (FULL NEUROLOGIC RECOVERY), ONE VENTRICULAR FIBRILLATION (VF), ATRIAL FIBRILLATION (AF) WAS INDUCED, THREE FEMORAL ACCESS ISSUES, ONE ATELECTASIA, ONE PERICARDITIS, EIGHT TRANSIENT ST ELEVATION/DEPRESSION POST APPLICATION, CORONARY SPASMS, THREE BIOLOGICAL HEMOLYSIS, ONE MECHANICAL LEFT BUNDLE BRANCH BLOCK (LBBB) AND ONE LV CHORDAE RUPTURE BY CATHETER MANIPULATION BUT WITHOUT CONSEQUENCE -ENTRAPMENT (NO CHANGE IN MITRAL VALVE INSUFFICIENCY WAS NOTED.). TREATMENTS INCLUDED MEDICATIONS FOR THE ST ELEVATIONS/SPASMS AND PERCUTANEOUS CORONARY INTERVENTION (PCI) FOR THE MYOCARDIAL INFARCTION. POSTOPERATIVE COMPLICATIONS INCLUDED TWO HEART FAILURES AFTER DAY 1, ONE MYOCARDIAL INFARCTION, ONE SEPSIS AND ONE MAJOR GASTRO-INTESTINAL BLEED. RECURRENCE OF VENTRICULAR TACHYCARDIA (VT) AND VENTRICULAR FIBRILLATION (VF) WERE NOTED. THE CATHETER EXHIBITED INFERENCE AND A RUPTURED SHAFT. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/59 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: SAFETY AND EFFICACY OF A LATTICE-TIP CATHETER FOR VENTRICULAR ARRHYTHMIA 1 ABLATION: 2 THE AFFERA VENTRICULAR ARRHYTHMIA ABLATION (AVAAR) REGISTRY. EP EUROPACE ¿ EHJ ARRHYTHMIAS AND ELECTROPHYSIOLOGY; 2025; EUAF139 DOI.ORG/10.1093/EUROPACE/EUAF13. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22485587&pc=LWS | 2124215-2025-46542 | 2025-06-13 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-11 04:00 | LWS | VIGILANT EL ICD DR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS NEWLY IMPLANTED IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM HAD PASSED AWAY ON (B)(6) 2025. EARLIER IN THE DAY, THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS IMPLANTED SUCCESSFULLY, COMPLETE WITH A NORMAL PRE-DISCHARGE CHECK WHILE IN THE HOSPITAL. LATER IN THE DAY, THE PATIENT SUFFERED A CARDIAC ARREST AND PASSED AWAY. THERE WERE NO ALERTS OR EPISODES NOTED IN THE ICD HISTORY DURING A POST-MORTEM INTERROGATION. ALL EVIDENCE INDICATES THE ICD SYSTEM REMAINS IN SITU.Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS NEWLY IMPLANTED IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM HAD PASSED AWAY ON (B)(6) 2025. EARLIER IN THE DAY, THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS IMPLANTED SUCCESSFULLY, COMPLETE WITH A NORMAL PRE-DISCHARGE CHECK WHILE IN THE HOSPITAL. LATER IN THE DAY, THE PATIENT SUFFERED A CARDIAC ARREST AND PASSED AWAY. THERE WERE NO ALERTS OR EPISODES NOTED IN THE ICD HISTORY DURING A POST-MORTEM INTERROGATION. ALL EVIDENCE INDICATES THE ICD SYSTEM REMAINS IN SITU. ADDITIONAL INFORMATION PROVIDED FROM THE FIELD INDICATED THE PHYSICIAN ASSOCIATED WITH THIS EVENT DID NOT BELIEVE EITHER THE ICD OR THE PROCEDURE TO ORIGINALLY IMPLANT IT HAD CAUSED OR CONTRIBUTED TO THE PATIENT'S PASSING THE FOLLOWING DAY. NO ADVERSE PATIENT EFFECTS RELATING TO THE ICD WERE REPORTED. Manufacturer Narrative: AS NO FURTHER INFORMATION REGARDING THIS EVENT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS REPORT WILL BE UPDATED SHOULD ADDITIONAL INFORMATION BE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22487101&pc=CBK | 2518422-2025-107339 | 2025-07-08 04:00 | Death | RESPIRONICS, INC. | 2025-07-11 04:00 | CBK | TRILOGY EVO, O2, JAPAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING WHEN CHANGING SETTINGS THAT THE VALUE HAD BECOME THE MAXIMUM VALUE ON A TRILOGY EVO O2 DEVICE. THE SALES REPRESENTATIVE VISITED THE CUSTOMER SITE AND THAT "AFTER TURNING THE POWER OFF, IT WAS AFTER A WHILE." WHEN YOU TOUCH AN ITEM THAT YOU WANT TO CHANGE BUT IT BECAME A MAXIMUM VALUE WHEN TOUCHING AN ITEM THAT WANTED TO BE CHANGED." THE SALES REPRESENTATIVE GOES ON TO STATE, "IT DID NOT RESPOND WHEN TOUCHING IT, BUT IT WAS IN A SITUATION LIKE "+" PR "+- WAS BEING PRESSED ALL THE TIME". THE SALES REPRESENTATIVE VISITED THE ME, "THERE WAS AN INQUIRY REGARDING THE CARE ORCHESTRATOR DATA. SETTINGS CHANGED ON 8TH, BUT IT WAS NOT RECORDED IN THE ELECTRONIC MEDICAL RECORDS." THE PATIENT HAD PASSED AWAY. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: UPDATED H10/H11, PATIENT IDENTIFIER, DATE OF BIRTH, AND RELEVANT TEST/LAB DATA FIELDS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING WHEN CHANGING SETTINGS THAT THE VALUE HAD BECOME THE MAXIMUM VALUE ON A TRILOGY EVO O2 DEVICE. THE SALES REPRESENTATIVE VISITED THE CUSTOMER SITE AND THAT "AFTER TURNING THE POWER OFF, IT WAS AFTER A WHILE." WHEN YOU TOUCH AN ITEM THAT YOU WANT TO CHANGE BUT IT BECAME A MAXIMUM VALUE WHEN TOUCHING AN ITEM THAT WANTED TO BE CHANGED." THE SALES REPRESENTATIVE GOES ON TO STATE, "IT DID NOT RESPOND WHEN TOUCHING IT, BUT IT WAS IN A SITUATION LIKE "+" PR "+- WAS BEING PRESSED ALL THE TIME". THE SALES REPRESENTATIVE VISITED THE ME, "THERE WAS AN INQUIRY REGARDING THE CARE ORCHESTRATOR DATA. SETTINGS CHANGED ON 8TH, BUT IT WAS NOT RECORDED IN THE ELECTRONIC MEDICAL RECORDS." THE PATIENT HAD PASSED AWAY. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER¿S COMPLAINT WAS NOT CONFIRMED DURING AN OPERATIONAL CHECK OF THE DEVICE. THE MANUFACTURER'S SERVICE TECHNICIAN WAS ABLE TO CONFIRM "THAT DURING CHANGES TO EACH SETTING, ADJUSTMENTS COULD BE MADE BOTH "+"AND "-" ACCORDING TO THE TOUCH OPERATIONS ON THE DISPLAY, WITHOUT ANY FORCED CHANGES TO THE SETTINGS. THE MANUFACTURER'S SERVICE TECHNICIAN REVIEWED THE DEVICE'S ERROR LOG AND OBSERVED NO ERRORS INDICATING ABNORMALITIES FOR THIS DEVICE. THE MANUFACTURER'S SERVICE TECHNICIAN DISASSEMBLED THE DEVICE AND PERFORMED AN INTERNAL INSPECTION OF THE DEVICE. THE MANUFACTURER'S SERVICE TECHNICIAN CONFIRMED THERE WERE NO ABNORMALITIES FOR THIS DEVICE. THE MANUFACTURER'S SERVICE TECHNICIAN PERFORMED A FUNCTIONAL TEST ON THE DEVICE. THE DEVICE PASSED THE FUNCTIONAL TEST. THE MANUFACTURER'S SERVICE TECHNICIAN "REQUEST FOR SOFTWARE IMPROVEMENT HAS BEEN MADE. AS FOR THE CHANGE IN SETTING VALUES AT THE SPECIFIC DATE AND TIME, IT IS BELIEVED THAT THE CHANGE WAS MADE MANUALLY, AS THE CHANGE IS NOT APPLIED UNLESS THE CONFIRM BUTTON IS PRESSED EVEN IF THE ISSUE OCCURS. ADDITIONAL FINDINGS NOT RELATED TO THE MALFUNCTION/ISSUE THE SILVER STICKER AROUND THE SILENCE BUTTON WAS GONE. THE MANUFACTURER'S SERVICE TECHNICIAN REPLACED THE VANITY COVER TO ADDRESS THIS ISSUE. THE MANUFACTURER'S SERVICE TECHNICIAN PROACTIVELY REPLACED THE AIR INLET FOAM FILTER AND PARTICLE FILTER FOR THIS DEVICE. AFTERWARDS, THE MANUFACTURER'S SERVICE TECHNICIAN PERFORMED THE FINAL AND RUN-IN TESTS, ALONG WITH THE BATTERY AND VALVE CHECKS. THE MANUFACTURER'S SERVICE TECHNICIAN CONFIRMED THE DEVICE WAS OPERATIONAL, AND IT WAS CLEANED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22487428&pc=QBJ | 2954323-2025-26785 | 2025-05-08 04:00 | Death | ABBOTT DIABETES CARE INC | 2025-07-11 04:00 | QBJ | FREESTYLE LIBRE 3 PLUS | Low Readings | Diabetic Ketoacidosis; Loss of consciousness | K223435 |  | 1 | Event Description: ON 02-JULY-2025, ADC RECEIVED AN EMAIL FROM A CAREGIVER, WHO REPORTED A LOW READINGS ISSUE WITH THE ADC DEVICE. IT WAS REPORTED THAT THE CUSTOMER HAD BEEN RECEIVING UNSPECIFIED ¿NORMAL¿ READINGS ON THE ADC DEVICE LEADING UP TO THE EVENT, AND THEREFORE THE CUSTOMER DID NOT ADMINISTER INSULIN FOR SEVERAL DAYS AND SUBSEQUENTLY LOST CONSCIOUSNESS. THE CUSTOMER WAS ADMITTED TO THE HOSPITAL WITH A LABORATORY BLOOD GLUCOSE RESULT OF 1400 MG/DL AND DIAGNOSED WITH DIABETIC KETOACIDOSIS. THE CAREGIVER STATED THAT DURING THE HOSPITALIZATION, THE CUSTOMER PASSED AWAY ON (B)(6) 2025. NO INFORMATION WAS PROVIDED AS TO WHAT TREATMENT WAS PROVIDED DURING THE HOSPITALIZATION NOR THE CAUSE OF DEATH. NO AUTOPSY REPORT OR DEATH CERTIFICATE HAS BEEN PROVIDED AT THIS TIME. GIVEN THE INFORMATION PRESENTED AT THIS TIME, IT IS INCONCLUSIVE THAT THE ADC DEVICE HAS LED TO OR CONTRIBUTED TO THE REPORTED DEATH. ADC CUSTOMER SERVICE IS CURRENTLY IN THE PROCESS OF MAKING ADDITIONAL ATTEMPTS TO GATHER FURTHER INFORMATION REGARDING THE REPORTED DEATH. A FOLLOW-UP REPORT WILL BE PROVIDED IF ADC OBTAINS ANY ADDITIONAL INFORMATION REGARDING THIS EVENT.Event Description: ON 02-JULY-2025, ADC RECEIVED AN EMAIL FROM A CAREGIVER, WHO REPORTED A LOW READINGS ISSUE WITH THE ADC DEVICE. IT WAS REPORTED THAT THE CUSTOMER HAD BEEN RECEIVING UNSPECIFIED ¿NORMAL¿ READINGS ON THE ADC DEVICE LEADING UP TO THE EVENT, AND THEREFORE THE CUSTOMER DID NOT ADMINISTER INSULIN FOR SEVERAL DAYS AND SUBSEQUENTLY LOST CONSCIOUSNESS. THE CUSTOMER WAS ADMITTED TO THE HOSPITAL WITH A LABORATORY BLOOD GLUCOSE RESULT OF 1400 MG/DL AND DIAGNOSED WITH DIABETIC KETOACIDOSIS. THE CAREGIVER STATED THAT DURING THE HOSPITALIZATION, THE CUSTOMER PASSED AWAY ON (B)(6) 2025. NO INFORMATION WAS PROVIDED AS TO WHAT TREATMENT WAS PROVIDED DURING THE HOSPITALIZATION NOR THE CAUSE OF DEATH. NO AUTOPSY REPORT OR DEATH CERTIFICATE HAS BEEN PROVIDED AT THIS TIME. GIVEN THE INFORMATION PRESENTED AT THIS TIME, IT IS INCONCLUSIVE THAT THE ADC DEVICE HAS LED TO OR CONTRIBUTED TO THE REPORTED DEATH. ADC CUSTOMER SERVICE IS CURRENTLY IN THE PROCESS OF MAKING ADDITIONAL ATTEMPTS TO GATHER FURTHER INFORMATION REGARDING THE REPORTED DEATH. A FOLLOW-UP REPORT WILL BE PROVIDED IF ADC OBTAINS ANY ADDITIONAL INFORMATION REGARDING THIS EVENT. ON 15-JULY-2025, ADC RECEIVED ADDITIONAL INFORMATION REGARDING THIS EVENT. THE CAREGIVER REPORTED A GLUCOSE READING OF 128 MG/DL OBTAINED FROM THE ADC DEVICE WHILE THE CUSTOMER WAS IN THE HOSPITAL¿S INTENSIVE CARE UNIT. IT IS UNKNOWN WHETHER THE CAREGIVER NOTED A TREND ARROW ON THE DEVICE. ADDITIONALLY, IT IS UNKNOWN IF ANY COMPARISON READINGS WERE TAKEN WITHIN 10 MINUTES AT THE TIME OF THE MEDICAL EVENT. THE CUSTOMER REPORTEDLY DID NOT RECEIVE THIRD-PARTY TREATMENT FROM ANYONE OTHER THAN THEMSELVES. A DEATH CERTIFICATE HAS BEEN ISSUED; HOWEVER THE DEATH CERTIFICATE DOES NOT SPECIFY AN OFFICIAL CAUSE OF DEATH, AND THE CAREGIVER IS NOT AWARE OF WHAT THE OFFICIAL CAUSE MAY BE. Manufacturer Narrative: THIS REPORT IS RELATED TO FDA MFR REPORT # 2954323-2025-18852. ADC RECEIVED ADDITIONAL INFORMATION ON 02-JULY-2025 IN WHICH IT WAS REPORTED THAT THE ADVERSE EVENT WAS RELATED TO A LOW READINGS ISSUE WITH THE ADC DEVICE. IT WAS PREVIOUSLY REPORTED TO ADC THAT THE ADVERSE EVENT WAS RELATED TO A SENSOR ERROR MESSAGE; THIS EVENT WAS SUBMITTED TO THE FDA UNDER MFR REPORT # 2954323-2025-18852. PER THE ADDITIONAL INFORMATION RECEIVED, NO FURTHER REPORTS WILL BE SUBMITTED UNDER FDA MFR REPORT # 2954323-2025-18852. THE DEVICE IN QUESTION WAS REMOVED AT THE HOSPITAL AND ITS CURRENT STATUS IS UNKNOWN. AN INVESTIGATION HAS BEEN PERFORMED FOR THE REPORTED COMPLAINT, AND IT HAS BEEN DETERMINED THAT THERE WAS NO INDICATION THAT THE PRODUCT DID NOT MEET SPECIFICATION. THE DHRS (DEVICE HISTORY RECORD) FOR THE PRODUCT WAS REVIEWED AND THE DHRS SHOWED THE PRODUCT MET SPECIFICATIONS PRIOR TO RELEASE TO DISTRIBUTION. THE MOST PROBABLE ROOT CAUSES ASSOCIATED WITH THIS FAILURE MODE ARE DISCONNECTED, FAULTY OR DAMAGED COMPONENTS, SOFTWARE/DATA CORRUPTION, OR MISUSE. HOWEVER, MITIGATIONS ARE IN PLACE TO REDUCE AND PREVENT SUCH ISSUES. ALL VITAL MANUFACTURING STEPS ARE VALIDATED, MONITORED, AND VERIFIED DURING MANUFACTURING TO ENSURE THE SYSTEM IS IN CONFORMANCE WITH THE VERIFIED DESIGN. ALL VITAL FUNCTIONS ARE MONITORED BY THE SYSTEM AND, WHEN NECESSARY, FUNCTION IS SUSPENDED TO SAFEGUARD AGAINST INACCURATE RESULTS. LABELING IS PROVIDED TO INSTRUCT THE USER ON THE INTENDED USE OF ALL VITAL PARTS OF THE SYSTEM TO MINIMIZE MISUSE. ALL COMPLAINTS AND COMPLAINT TRENDS ARE INVESTIGATED TO DETERMINE IF THERE IS A PRODUCT DEFECT/ DEFICIENCY. IF A PRODUCT DEFECT/ DEFICIENCY IS IDENTIFIED, A RISK EVALUATION IS COMPLETED AND COMPARED TO THE RISK MANAGEMENT REPORT, TO ENSURE THE RISK PROFILE HAS NOT CHANGED. ADDITIONALLY, AS A PART OF ABBOTT¿S POST-MARKET SURVEILLANCE PROCESS, ALL RISK EVALUATIONS WITH ASSOCIATED COMPLAINT DATA ARE REVIEWED ANNUALLY TO DETERMINE IF THE RISK PROFILES HAVE CHANGED AS COMPARED TO THE PRODUCT RISK MANAGEMENT REPORTS. THESE MONITORING PROCESSES ENSURE THAT ALL PRODUCT RISK PROFILES REMAIN ACCEPTABLE AND HAVE A POSITIVE BENEFIT/ RISK RATIO. AN EXTENDED MANUFACTURING REVIEW WILL BE PERFORMED AND A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. IF PRODUCT IS RETURNED, A PHYSICAL INVESTIGATION WILL BE PERFORMED PER ADC'S ESTABLISHED PROCESSES AND PROCEDURES AND A REPORT WILL BE SUBMITTED. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED. Manufacturer Narrative: THE FOLLOWING SECTION HAS BEEN UPDATED: B5 - DESCRIBE EVENT OR PROBLEM THIS REPORT IS RELATED TO FDA MFR REPORT # 2954323-2025-18852. ADC RECEIVED ADDITIONAL INFORMATION ON 02-JULY-2025 IN WHICH IT WAS REPORTED THAT THE ADVERSE EVENT WAS RELATED TO A LOW READINGS ISSUE WITH THE ADC DEVICE. IT WAS PREVIOUSLY REPORTED TO ADC THAT THE ADVERSE EVENT WAS RELATED TO A SENSOR ERROR MESSAGE; THIS EVENT WAS SUBMITTED TO THE FDA UNDER MFR REPORT # 2954323-2025-18852. PER THE ADDITIONAL INFORMATION RECEIVED, NO FURTHER REPORTS WILL BE SUBMITTED UNDER FDA MFR REPORT # 2954323-2025-18852. THE DEVICE IN QUESTION WAS REMOVED AT THE HOSPITAL AND ITS CURRENT STATUS IS UNKNOWN. AN EXTENDED MANUFACTURING REVIEW OF THE REPORTED PRODUCT HAS BEEN PERFORMED. LOT 1001015875 PERFORMED WITHIN SPECIFICATION FOR ALL MEASUREMENTS DURING THE SENSOR MANUFACTURING AND RELEASE TESTING PROCESS. THERE WERE NO NON-CONFORMANCES IN THE SAME TIME PERIOD THAT WAS RELATED TO MANUFACTURE OF LOT 1001015875. ALL MEASUREMENTS REVIEWED WERE WITHIN SPECIFICATIONS. THE DHRS (DEVICE HISTORY RECORD) FOR THE PRODUCT WAS REVIEWED AND THE DHRS SHOWED THE PRODUCT MET SPECIFICATIONS PRIOR TO RELEASE TO DISTRIBUTION. THE MOST PROBABLE ROOT CAUSES ASSOCIATED WITH THIS FAILURE MODE ARE DISCONNECTED, FAULTY OR DAMAGED COMPONENTS, SOFTWARE/DATA CORRUPTION, OR MISUSE. HOWEVER, MITIGATIONS ARE IN PLACE TO REDUCE AND PREVENT SUCH ISSUES. ALL VITAL MANUFACTURING STEPS ARE VALIDATED, MONITORED, AND VERIFIED DURING MANUFACTURING TO ENSURE THE SYSTEM IS IN CONFORMANCE WITH THE VERIFIED DESIGN. ALL VITAL FUNCTIONS ARE MONITORED BY THE SYSTEM AND, WHEN NECESSARY, FUNCTION IS SUSPENDED TO SAFEGUARD AGAINST INACCURATE RESULTS. LABELING IS PROVIDED TO INSTRUCT THE USER ON THE INTENDED USE OF ALL VITAL PARTS OF THE SYSTEM TO MINIMIZE MISUSE. ALL COMPLAINTS AND COMPLAINT TRENDS ARE INVESTIGATED TO DETERMINE IF THERE IS A PRODUCT DEFECT/ DEFICIENCY. IF A PRODUCT DEFECT/ DEFICIENCY IS IDENTIFIED, A RISK EVALUATION IS COMPLETED AND COMPARED TO THE RISK MANAGEMENT REPORT, TO ENSURE THE RISK PROFILE HAS NOT CHANGED. ADDITIONALLY, AS A PART OF ABBOTT¿S POST-MARKET SURVEILLANCE PROCESS, ALL RISK EVALUATIONS WITH ASSOCIATED COMPLAINT DATA ARE REVIEWED ANNUALLY TO DETERMINE IF THE RISK PROFILES HAVE CHANGED AS COMPARED TO THE PRODUCT RISK MANAGEMENT REPORTS. THESE MONITORING PROCESSES ENSURE THAT ALL PRODUCT RISK PROFILES REMAIN ACCEPTABLE AND HAVE A POSITIVE BENEFIT/ RISK RATIO. IF PRODUCT IS RETURNED, A PHYSICAL INVESTIGATION WILL BE PERFORMED PER ADC'S ESTABLISHED PROCESSES AND PROCEDURES AND A REPORT WILL BE SUBMITTED UPON COMPLETION OF INVESTIGATION. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22487469&pc=PNJ | 9612164-2025-03414 | 2023-12-17 05:00 | Death | MEDTRONIC IRELAND | 2025-07-11 04:00 | PNJ | MICRA AV | Appropriate Term/Code Not Available | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE PATIENT PASSED AWAY DUE TO THE FAILURE OF LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). THEY FURTHER STATED THAT THE PATIENT WENT TO THE HOSPITAL AND THE DEVICE WAS REPROGRAMMED. ONE DAY LATER, PATIENT PASSED AWAY.Event Description: IT WAS FURTHER REPORTED THAT THERE WAS NO PERFORMANCE ISSUE WITH THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). CAUSE OF DEATH WAS HYPERCAPNIC RESPIRATORY FAILURE AND TORSADE DE POINT, PATIENT WAS DNR. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22487554&pc=NVN | 2017865-2025-92982 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-11 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22487559&pc=NVN | 2017865-2025-92981 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-11 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22487613&pc=LWP | 2017865-2025-92980 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-11 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22466759&pc=NVZ | 3008973940-2025-04979 | 2025-06-01 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-07-10 04:00 | NVZ | SENSIA IPG | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. IT WAS ALSO REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED LOW IMPEDANCE.Event Description: IT WAS NOTED THAT THE DEATH WAS NOTED TO BE UNRELATED TO THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22466773&pc=DXY | 2182208-2025-03331 | 2025-06-01 04:00 | Death | MEDTRONIC, INC. | 2025-07-10 04:00 | DXY | CAPSURE VDD-2 LEAD | Low impedance | Insufficient Information | P890003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. IT WAS ALSO REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED LOW IMPEDANCE.Event Description: IT WAS NOTED THAT THE DEATH WAS NOTED TO BE UNRELATED TO THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22466786&pc=MKJ | 1218058-2025-00114 | 2025-06-26 04:00 | Death | BIO-DETEK INCORPORATED | 2025-07-10 04:00 | MKJ | PRO PADZ BISPHASIC MFE (1 PA) | Sparking; Noise, Audible | Cardiac Arrest | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A 62-YEAR-OLD MALE PATIENT IN CARDIAC ARREST, AFTER DELIVERING THE SELECTED ENERGY, THE ANTERIOR ELECTRODE PAD SPARKED AND POPPED. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER SET OF ELECTRODE PADS TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: JUSTIFICATION FOR NO UDI: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22466806&pc=MKJ | 1220908-2025-02544 | 2025-05-22 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-07-10 04:00 | MKJ | X SERIES | Unable to Obtain Readings | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A 3-MONTH-OLD PATIENT (GENDER UNKNOWN), THE DEVICE WAS UNABLE TO OBTAIN AN ECG SIGNAL VIA ELECTRODE PADS. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED; HOWEVER, THEY DO NOT BELIEVE THIS WAS DEVICE ASSOCIATED. Manufacturer Narrative: JUSTIFICATION FOR NO UDI: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. Manufacturer Narrative: THE DEVICE AND MULTIFUNCTION CABLE (MFC) WERE RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. THE CUSTOMER'S REPORT WAS OBSERVED DURING REVIEW OF THE DEVICE DATA LOGS. HOWEVER, THE DEVICE WAS PUT THROUGH EXTENSIVE TESTING USING DEFIB SIMULATOR TESTING, DEFIB CYCLE TESTING, AND BENCH HANDLING WITHOUT DUPLICATING THE REPORT. AN INTERNAL INSPECTION FOUND NO DISCREPANCIES. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. THE ELECTRODE PADS USED WERE NOT RETURNED AS PART OF THIS INVESTIGATION. REVIEW OF THE DEVICE LOG SHOWED AN ECG THAT APPEARED INTERMITTENT. "PADS OFF" MESSAGES WERE OBSERVED, WHICH INDICATE THE DEVICE RECOGNIZED THAT PADS WERE ATTACHED BUT IT WAS UNABLE TO OBTAIN A VALID PATIENT IMPEDANCE READING. THIS IS TYPICALLY CAUSED BY POOR COUPLING BETWEEN THE PADS AND THE PATIENT BUT MAY ALSO BE CAUSED BY AN ISSUE WITH THE PADS OR POOR CONNECTION BETWEEN THE DEVICE AND ACCESSORIES. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22466837&pc=GDW | 1219930-2025-02898 | 2025-06-11 04:00 | Death | COVIDIEN SURGICAL | 2025-07-10 04:00 | GDW | SIGNIA | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Laceration(s) | K191070 |  | 1 | Event Description: IT WAS REPORTED THAT WHEN CLAMPING A SEGMENT OF THE LEFT PULMONARY ARTERY (A1/3), BLEEDING OCCURRED AFTER REMOVAL OF THE MAGAZINE AT THE PROXIMAL CLAMP LINE, WHICH RAPIDLY INCREASED AND NECESSITATED AN EMERGENCY THORACOTOMY. DUE TO MASSIVE BLEEDING WITH CARDIAC ARREST, THE PATIENT DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22466840&pc=OEV | 9610691-2025-00018 | 2025-06-12 04:00 | Death | AMBU A/S | 2025-07-10 04:00 | OEV | SPUR II | Incomplete or Inadequate Connection | Hypoxia |  |  | 1 | Event Description: A CONNECTION ISSUE WAS OBSERVED BETWEEN THE O2 CONNECTOR AND THE TUBING DURING USE. THE DEVICE INVOLVED WAS AN AIRPRODUCT OXYGEN CYLINDER, WITH A FLOW RATE OF 15 L/MIN. THE SPUR II TUBING NO LONGER CONNECTS PROPERLY, WHICH IS A NEW ISSUE. PREVIOUSLY, THE SPUR II STAYED IN PLACE WITHOUT ANY PROBLEM. IT WAS TESTED ON 7 DIFFERENT CYLINDERS, BUT THE CONNECTION REMAINED UNSTABLE. DURING THE LAST USE, A THIRD PERSON HAD TO HOLD BOTH THE CYLINDER AND THE RESUSCITATOR TO ALLOW PROPER FUNCTIONING, ALTHOUGH THIS IS NOT THE NORMAL INTENDED USE OF THE DEVICE. THE PATIENT UNFORTUNATELY PASSED AWAY. THEY HAD AN UNDERLYING MEDICAL CONDITION, SO WE CANNOT DIRECTLY ATTRIBUTE THE OUTCOME TO THE SPUR II ISSUE. HOWEVER, DUE TO THE CONNECTION PROBLEM, A THIRD PERSON WAS REQUIRED TO HOLD BOTH THE CYLINDER AND THE RESUSCITATOR TO ENSURE IT FUNCTIONED CORRECTLY. THIS SITUATION REDUCED THE NUMBER OF RESCUERS AVAILABLE TO ASSIST WITH OTHER CRITICAL ACTIONS DURING THE INTERVENTION. Manufacturer Narrative: THERE WAS NO SAMPLE RETURNED FOR INVESTIGATION, THEREFORE, THE REPORTED FAILURE ABOUT SPUR II HAS CONNECTION ISSUE WITH AIRPRODUCT OXYGEN CYLINDER COULD NOT BE VERIFIED. SPUR II OXYGEN TUBE CONNECTOR IS DESIGNED TO BE A FEMALE CONNECTOR THAT IS COMPATIBLE WITH A CONNECTOR WHICH IS IN COMPLIANCE WITH EN13544-2. THE CONNECTOR IS MADE BY INJECTION WHICH HAS BEEN VERIFIED PER PROCESS VALIDATION. AFTER INJECTION, QC PERFORMS SAMPLE DIMENSION MEASUREMENT ON INNER DIAMETER. IF IT IS OUT OF SPECIFICATION, IT SHOULD BE EASILY DETECTED. AFTER REVIEWING PRODUCTION RECORDS, NO ABNORMALITY WAS FOUND. THE IFU STATES THAT THE SPUR II CONNECTORS FOLLOW EN 13544-2 STANDARD AND ALSO "WHEN APPLYING EXTERNAL DEVICES TO THE RESUSCITATOR, MAKE SURE TO TEST FOR FUNCTIONALITY AND CONSULT THE INSTRUCTIONS FOR USE ACCOMPANYING THE EXTERNAL DEVICE". SPUR II OXYGEN TUBE CONNECTOR AND AIRPRODUCT OXYGEN CYLINDER CONNECTOR FOLLOW DIFFERENT STANDARDS, THEREFORE THEY ARE INCOMPATIBLE. WE WILL KEEP MONITORING THE MARKET FOR SIMILAR INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22466860&pc=FTL | 1213643-2025-00605 | 2016-08-09 04:00 | Death | DAVOL INC., SUB. C.R. BARD, INC. -1213643 | 2025-07-10 04:00 | FTL | MESH - VENTRALEX | Patient-Device Incompatibility | Abscess; Bacterial Infection; Fistula | K132441 |  | 1 | Event Description: PER INFORMATION PROVIDED: ON (B)(6) 2008 ¿ PATIENT WAS DIAGNOSED WITH RECURRENT VENTRAL HERNIA ABOVE UMBILICUS THEREBY UNDERWENT OPEN REPAIR WITH VENTRALEX MESH (DEVICE #1, #2). PER OP NOTES, ¿A VENTRALEX PATCH (DEVICE #1) WAS PLACED INTO PREPERITONEAL POCKET AND SUTURED. THEN, PRUDENT TROCAR SITE WAS REPAIRED USING VENTRALEX PATCH (DEVICE #2) PLACED BETWEEN THE FASCIA AND THE PERITONEUM AND SECURED USING SUTURES CATCHING A PORTION OF UNDERLYING MESH (DEVICE #1).¿ ON (B)(6) 2015 ¿ PATIENT WAS DIAGNOSED WITH CELLULITIS AND ABSCESS THEREBY UNDERWENT INCISION & DRAINAGE OF ABSCESS. PER NOTES, ¿A SMALL AMOUNT OF SEROUS FLUID WAS OBTAINED, AND ANTIBIOTICS WERE PROVIDED.¿ ON (B)(6) 2016 ¿ PATIENT WAS DIAGNOSED WITH INFECTED UMBILICAL MESH WITH CUTANEOUS FISTULA THEREBY UNDERWENT LAPAROSCOPIC REPAIR WITH THE REMOVAL OF MESHES (DEVICE #1, #2). PER OP NOTES, ¿A OLD CUTANEOUS DEFECTS ERODED THROUGH THE ABDOMINAL WALL AND MESH (DEVICE#1, #2) WERE NOTED. THE MESH (DEVICE #1, #2) IMBEDDED INTO FASCIA WAS REMOVED COMPLETELY. PERITONEUM HAD CHRONIC GRANULATION TISSUE WAS EXCISED.¿ ATTORNEY ALLEGES THAT THE PATIENT PASSED AWAY ON (B)(6) 2017. Manufacturer Narrative: BASED ON THE MEDICAL RECORD REVIEW, POST IMPLANT OF VENTRALEX MESHES (DEVICE #1, #2), PATIENT HAD BACTERIAL INFECTION, CUTANEOUS FISTULA, CELLULITIS, ABSCESS, THEREBY UNDERWENT LAPAROSCOPIC REPAIR WITH THE REMOVAL OF MESHES (DEVICE #1, #2). PER PLAINTIFF PROFILE FORM, PATIENT HAD ABSCESS, SEROMA, CHRONIC INFECTION, CELLULITIS, ABDOMINAL DRAINING AND ALLEGES PATIENT DEATH. NO INFORMATION REGARDING PATIENT¿S DEATH IN THE MEDICAL RECORDS PROVIDED TO DATE AND NO AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED." ROOT CAUSE IS INCONCLUSIVE; NO CONCLUSION CAN BE MADE AS TO THE EXTENT TO WHICH THE IMPLANT MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S POSTOPERATIVE COURSE. NOTE, THE MANUFACTURING DATE (15-FEB-2008) IS CONSIDERED TO BE A BEST ESTIMATE. THIS EMDR REPRESENTS THE BARD MESH - VENTRALEX (DEVICE #2). AN ADDITIONAL SUPPLEMENTAL EMDR WAS SUBMITTED TO REPRESENT THE BARD MESH - VENTRALEX (DEVICE #1). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22467166&pc=DSQ | 2916596-2025-04378 | 2023-10-29 04:00 | Death | THORATEC CORPORATION | 2025-07-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED OUTCOME; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED BY THE ACCOUNT. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22467479&pc=LWS | 9614453-2025-02639 | 2025-06-16 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-07-10 04:00 | LWS | COBALT¿ XT DR MRI SURESCAN¿ | Appropriate Term/Code Not Available | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S SON THAT THERE WAS AN "INCIDENT" WITH THE PATIENT'S IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD). THE PATIENT PASSED AWAY APPROXIMATELY A WEEK LATER. CAUSE OF DEATH INFORMATION WAS REQUESTED BUT WAS NOT RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22467526&pc=LHQ | MW5172494 | 2024-03-18 04:00 | Death | UNK | 2025-07-10 04:00 | LHQ | THERMAL IMAGING | Improper or Incorrect Procedure or Method | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: THERMAL IMAGING / FUNCTIONAL INFRARED DEVICE MISUSE - MY WIFE'S PROVIDERS REPRESENTED THERMAL IMAGING AS AN ALTERNATIVE THAT THEY COULD IDENTIFY BREAST CANCER WITH THEIR THERMAL IMAGING USE AND DID NOT ADVISE THAT THERMAL IMAGING IS ONLY AN ADJUNCTIVE METHOD. WEBSITE ((B)(6) IS THE OWNER - (B)(6)) STATES "WELCOME TO (B)(6) LEADING HEALTHCARE IMAGING SERVICE. WE SPECIALIZE IN AN INNOVATIVE, SAFE ALTERNATIVE TO CONVENTIONAL BODY IMAGING AND HEALTH SCREENINGS. OUR PERSONAL AND PROFESSIONAL SERVICES PROVIDE DIVERSE BODY IMAGING CHOICES FOR MEN, WOMEN, AND CHILDREN; A FULL BODY SCAN, A PARTIAL BODY SCAN OR BREAST IMAGING. THERMOGRAPHY IS A RELIABLE METHOD TO DETECT THE SIGNS OF HEALTH CONDITIONS BEFORE THEY BECOME A PROBLEM ALLOWING YOU TO MAKE CHANGES THAT WILL HELP YOUR BODY HEAL." REPORT FROM (B)(6) 2024, VISIT SHOWED SHE WAS AT RISK OF DEVELOPING AGGRESSIVE TISSUE, BUT DID NOT INDICATE SHE HAD ANY AGGRESSIVE TISSUE. NO REFERRAL WAS MADE. I HAVE SIGNIFICANT DOCUMENTATION. ON THE REPORTS, THE "READER" IS (B)(6). HE CONSULTED WITH MY WIFE ON THE IMAGING RUN BY (B)(6). I FOUND OUT AFTER THE FACT THAT HE WAS PREVIOUSLY WARNED IN CONJUNCTION WITH (B)(6) IN 2019 BY THE FDA ABOUT THEIR MISUSE OF THERMAL IMAGING. (B)(6) WERE THE SUBJECT OF ABC'S GOOD MORNING AMERICA, NIGHTLINE, AND ONLINE ARTICLES FOR THEIR MISUSE OF THERMAL IMAGING. (B)(6) DOCTOR OF CHIROPRACTIC LICENSE WAS PULLED BY STATE OF (B)(6) IN 2023 BECAUSE OF HIS VIOLATIONS. HE STILL KEPT THE (B)(6) INITIALS ON THE REPORTS PROVIDED TO MY WIFE IN 2024. HE ALSO REPRESENTED HIMSELF AS A FUNCTIONAL MEDICINE DOCTOR AND (B)(6), HE, AND HIS OTHER COLLEAGUES ALL REFER TO HIM AS "DOCTOR" IN THEIR COMMUNICATIONS TO MY WIFE. WE BECAME ESPECIALLY AWARE OF THEIR INCOMPETENCE AFTER A FOLLOW-UP SCAN FOR (B)(6) 2024, SHOWED MASSIVE CHANGES TO THE BREAST AND (B)(6) RECOMMENDED FOLLOWING UP IN ANOTHER YEAR. PDFS OF REPORTS CAN BE PROVIDED. HORRIFIC MISDIRECTION AND MINIMIZATION OF SYMPTOMS PROVIDED. BREAST CANCER METASTASIZED BECAUSE OF THIS AND KILLED HER ON (B)(6) 2024.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22468268&pc=MVK | 3008642652-2025-06698 | 2025-03-29 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-10 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 22:40:14 ON (B)(6) 2025. AT 02:03:49 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 02:04:26, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 02:04:57, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 02:51:43. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION. THE REVIEW OF THE DATA INDICATED THAT THE DEVICE POWERED ON NORMALLY AND WAS ABLE TO ACQUIRE THE PATIENT¿S ECG SIGNAL ON THE LAST DAY OF USE CAPTURED IN THE DATA DOWNLOAD. NO DEFICIENCIES ALLEGED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22468393&pc=DTB | 2649622-2025-18603 | 2025-06-10 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-10 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Impedance Problem | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS FURTHER NOTED THAT THE PATIENT PASSED AWAY.Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED OUT OF RANGE IMPEDANCE PREVENTING THE PATIENT FROM HAVING A MAGNETIC RESONANCE IMAGING (MRI) BY LOCKING OUT THE IMPLANTABLE PULSE GENERATOR (IPG). THE RV LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 383069 LEAD WAS IMPLANTED ON (B)(6) 2017. 5076-52 LEAD WAS IMPLANTED ON (B)(6) 2016. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22469008&pc=DSQ | 2916596-2025-04279 | 2025-06-17 04:00 | Death | THORATEC CORPORATION | 2025-07-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT CHOSE TO DEACTIVATE THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) AND WAS TRANSITIONED TO COMFORT CARE. THE PATIENT PASSED AWAY DUE TO SEPSIS. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS REPORTED THAT THE ONSET DATE OF THE SEPSIS WAS (B)(6) 2025. THE INFECTION SOURCED FROM A URINARY TRACT INFECTION (UTI) AND OSTEOMYELITIS. THE PATIENT EXPERIENCED HYPOTENSION, ALTERED MENTAL STATUS, AND POSITIVE URINE CULTURES FOR YEAST. THE PATIENT WAS TREATED WITH VANCOMYCIN AND ZOSYN INITIALLY, WHICH WAS LATER CHANGED TO MEROPENEM AND MICAFUNGIN. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(4), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(4), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING NEUROLOGICAL EVENTS (NOT STROKE RELATED) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿, LISTS NEUROLOGICAL DYSFUNCTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22469859&pc=MCX | 3004742232-2025-00199 | 2025-06-21 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-07-10 04:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Perforation of Vessels | P130005 |  | 1 | Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A DIFFUSELY CALCIFIED RIGHT CORONARY ARTERY (RCA). EIGHT LOW SPEED TREATMENTS WERE PERFORMED FOLLOWED BY THREE HIGH SPEED TREATMENTS. TOWARDS THE END OF THE PROCEDURE, THE OAD CROWN CAUSED A PERFORATION IN THE RCA. THE PHYSICIAN ATTEMPTED TO SEAL THE PERFORATION USING A PROLONGED BALLOON DILATION INITIALLY AND LATER BY DEPLOYING A COVERED STENT BUT UNFORTUNATELY THIS WAS UNSUCCESSFUL. THE PATIENT'S CONDITION DETERIORATED AND THE PATIENT WENT INTO CARDIAC ARREST. THE PATIENT EXPIRED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22469928&pc=MSX | 3009156722-2025-00012 | 2025-04-24 04:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-07-10 04:00 | MSX | BENEVISION DISTRIBUTED MONITORING SYSTEM (DMS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K220058 |  | 1 | Event Description: THE CUSTOMER IS NOT REPORTING A PRODUCT MALFUNCTION OF A MINDRAY DEVICE. THE CUSTOMER INDICATED THAT THEY WERE TRYING TO RETRIEVE DISCHARGED PATIENT INFORMATION BUT COULD NOT LOCATE IT AND IS REQUESTING HELP IN RETRIEVING PATIENT'S DISCHARGED DATA AS THE PATIENT EXPIRED ON (B)(6) 2025. Manufacturer Narrative: MINDRAY FIELD SERVICE REPRESENTATIVE VISITED THE FACILITY AND COLLECTED BENEVISION DMS LOGS, BUT DISCHARGED PATIENT'S DATA WAS NOT AVAILABLE. MINDRAY PERFORMED A REVIEW OF THE AVAILABLE SYSTEM LOGS, AND NO DEVICE PERFORMANCE ISSUES WERE IDENTIFIED. THE SYSTEM PERFORMED PER SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22469935&pc=MSX | 2221819-2025-00012 | 2025-04-24 04:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-07-10 04:00 | MSX | BENEVISION DISTRIBUTED MONITORING SYSTEM (DMS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE CUSTOMER IS NOT REPORTING A PRODUCT MALFUNCTION OF A MINDRAY DEVICE. THE CUSTOMER INDICATED THAT THEY WERE TRYING TO RETRIEVE DISCHARGED PATIENT INFORMATION BUT COULD NOT LOCATE IT AND IS REQUESTING HELP IN RETRIEVING PATIENT'S DISCHARGED DATA AS THE PATIENT EXPIRED ON (B)(6) 2025. Manufacturer Narrative: MINDRAY FIELD SERVICE REPRESENTATIVE VISITED THE FACILITY AND COLLECTED BENEVISION DMS LOGS, BUT DISCHARGED PATIENT'S DATA WAS NOT AVAILABLE. MINDRAY PERFORMED A REVIEW OF THE AVAILABLE SYSTEM LOGS, AND NO DEVICE PERFORMANCE ISSUES WERE IDENTIFIED. THE SYSTEM PERFORMED PER SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22470270&pc=QZI | 2124215-2025-46111 | 2025-06-11 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-10 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Loss of consciousness; Ischemia Stroke |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A STROKE DURING A PULSED FIELD ABLATION PROCEDURE. A CT SCAN CONFIRMED THE STROKE. THE DEVICE IS NOT EXPECTED TO RETURN AS IT WAS DISPOSED, THEREFORE THE LOT NUMBER IS NOT AVAILABLE.Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A STROKE DURING A PULSED FIELD ABLATION PROCEDURE. A CT SCAN CONFIRMED THE STROKE. THE DEVICE IS NOT EXPECTED TO RETURN AS IT WAS DISPOSED, THEREFORE THE LOT NUMBER IS NOT AVAILABLE. IT WAS FURTHER REPORTED, A 72-YEAR-OLD FEMALE WITH A HISTORY OF PAROXYSMAL ATRIAL FIBRILLATION (AF) AND NO STRUCTURAL HEART ABNORMALITIES OR REDUCED EJECTION FRACTION UNDERWENT A PLANNED PROCEDURE. THE INTERVENTION WAS PERFORMED UNDER UNINTERRUPTED ORAL ANTICOAGULATION, AND A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS CONDUCTED PRE-PROCEDURE TO RULE OUT THE PRESENCE OF THROMBUS. THE PROCEDURE HAD PREVIOUSLY BEEN RESCHEDULED FROM THE END OF MAY TO JUNE DUE TO A CONSOLE NOT FUNCTIONING AS EXPECTED AT THE ACCOUNT. THE PATIENT WAS NEVER PREPPED FOR THE PROCEDURE IN MAY AND DID NOT RECEIVE ANY ANESTHESIA. ON THE DAY OF THE PROCEDURE, ACTIVATED CLOTTING TIME (ACT) WAS MAINTAINED ABOVE 300 SECONDS WITH HEPARIN ADMINISTERED PRIOR TO THE TRANSSEPTAL PUNCTURE. THE VERSACROSS SYSTEM WAS USED FOR THE TRANSSEPTAL ACCESS. LEFT ATRIAL DWELL TIME WAS KEPT UNDER 15 MINUTES. PULMONARY VEIN ISOLATION (PVI) WAS PERFORMED EXCLUSIVELY, REQUIRING A TOTAL OF 48 APPLICATIONS. FOLLOWING PVI, A CAVOTRICUSPID ISTHMUS (CTI) LINE WAS CREATED USING A RADIOFREQUENCY (RF) CATHETER. THE ONLY ADDITIONAL CATHETER PLACED WAS IN THE CORONARY SINUS. POST-PROCEDURE, THE PATIENT EXPERIENCED DIFFICULTY AWAKENING FROM ANESTHESIA AND WAS CHALLENGING TO EXTUBATE. A CT SCAN FOLLOWED BY AN MRI REVEALED MULTIPLE CEREBRAL INFARCTS CONSISTENT WITH AN EMBOLIC SHOWER. THE PATIENT DID NOT REGAIN CONSCIOUSNESS FOLLOWING GENERAL ANESTHESIA. NO TREATMENT WAS ADMINISTERED BEYOND INTRACRANIAL PRESSURE (ICP) RELIEF, AS THE STROKE INVOLVED MULTIPLE AREAS WITH ASSOCIATED MASS EFFECTS. THE PATIENT HAS SINCE BEEN TRANSITIONED TO COMFORT CARE, WITH NO EXPECTATION OF RECOVERY. THE PHYSICIAN WAS UNABLE TO DETERMINE THE CAUSE OF THE STROKE. THE TEE WAS REVIEWED AND THOROUGHLY EXAMINED, SHOWING NO EVIDENCE OF A LEFT ATRIAL APPENDAGE (LAA) THROMBUS. IT IS HIGHLY UNLIKELY THAT THE CEREBROVASCULAR ACCIDENT (CVA) WAS DUE TO A MISSED THROMBUS, AS THE INFARCTS WERE DISTRIBUTED THROUGHOUT THE BRAIN. THE FARAWAVE LOT NUMBER HAS BEEN PROVIDED. IT WAS FURTHER REPORTED THAT THE PATIENT HAD PASSED AWAY AND THE DATE OF DEATH IS NOT AVAILABLE. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. Manufacturer Narrative: ADDITIONAL INFORMATION IN SECTION B2 OUTCOMES ATTRIB TO ADV EVENT, SECTION B5 DESCRIBE EVENT OR PROBLEM, SECTION D4 LOT NUMBER, SECTION H6 PATIENT CODES AND SECTION H6 IMPACT CODES. CORRECTION TO SECTION B3 DATE OF EVENT. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22470480&pc=QBJ | 2954323-2025-26495 | 2025-06-19 04:00 | Death | ABBOTT DIABETES CARE INC | 2025-07-10 04:00 | QBJ | FREESTYLE LIBRE 3 PLUS | Low Readings | Dehydration; Hyperglycemia; Dizziness | K223435 |  | 1 | Event Description: ABBOTT DIABETES CARE (ADC) WAS NOTIFIED ON JUNE 30, 2025 BY A HEALTH CARE PROVIDER (HCP) THAT A CUSTOMER, WHO WAS PREGNANT, WAS RECEIVING FLAT AND LOW SENSOR VALUES FOR SEVERAL DAYS FROM AN ADC DEVICE. THE HCP FURTHER REPORTED THAT THE CUSTOMER EXPERIENCED SEVERE KETOACIDOSIS RESULTING IN FETAL DEATH. ON (B)(6) 2025, THE CUSTOMER WAS CONTACTED BY ADC AND REPORTED THAT ON (B)(6) 2025 AROUND 8:00 PM THEY EXPERIENCED SYMPTOMS OF DISCOMFORT, DRY MOUTH, DIZZINESS AND THE FIRE DEPARTMENT WERE CALLED WHERE A BLOOD GLUCOSE READING OF ¿BETWEEN 400 TO 600 MG/DL¿ WAS OBTAINED DURING TRANSPORT TO THE HOSPITAL. REPORTEDLY, THE CUSTOMER WAS TREATED BY THE FIREMEN AND AT THE HOSPITAL IN "REANIMATION SERVICE" (CRITICAL CARE) FOR A DIAGNOSIS OF HYPERGLYCEMIA; HOWEVER, TREATMENT DETAILS OR ADDITIONAL GLUCOSE READINGS WERE NOT PROVIDED. IT IS UNKNOWN WHETHER THE CUSTOMER NOTED A TREND ARROW ON THE DEVICE AT THE TIME OF THE EVENT. THE CUSTOMER INDICATED THE DEATH OF HER BABY OCCURRED ON (B)(6) 2025 WHILE SHE WAS IN "REANIMATION SERVICE" OF THE HOSPITAL. NO AUTOPSY REPORT OR DEATH CERTIFICATE HAS BEEN PROVIDED AT THIS TIME. GIVEN THE INFORMATION PRESENTED AT THIS TIME, IT IS INCONCLUSIVE THAT THE ADC DEVICE HAS LED TO OR CONTRIBUTED TO THE REPORTED FETAL DEATH. A FOLLOW-UP REPORT WILL BE SENT IF ADC OBTAINS ANY ADDITIONAL INFORMATION REGARDING THIS EVENT.Event Description: AN ANSM USER REPORT WAS RECEIVED THAT OBTAINED THE FOLLOWING INFORMATION: A HEALTH CARE PROVIDER (HCP) THAT A CUSTOMER, WHO WAS PREGNANT, WAS RECEIVING FLAT AND LOW SENSOR VALUES FOR SEVERAL DAYS FROM AN ADC DEVICE. THE HCP FURTHER REPORTED THAT THE CUSTOMER EXPERIENCED SEVERE KETOACIDOSIS RESULTING IN FETAL DEATH. ON (B)(6) 2025, THE CUSTOMER WAS CONTACTED BY ADC AND REPORTED THAT ON (B)(6) 2025 AROUND 8:00 PM THEY EXPERIENCED SYMPTOMS OF DISCOMFORT, DRY MOUTH, DIZZINESS AND THE FIRE DEPARTMENT WERE CALLED WHERE A BLOOD GLUCOSE READING OF ¿BETWEEN 400 TO 600 MG/DL¿ WAS OBTAINED DURING TRANSPORT TO THE HOSPITAL. REPORTEDLY, THE CUSTOMER WAS TREATED BY THE FIREMEN AND AT THE HOSPITAL IN "REANIMATION SERVICE" (CRITICAL CARE) FOR A DIAGNOSIS OF HYPERGLYCEMIA; HOWEVER, TREATMENT DETAILS OR ADDITIONAL GLUCOSE READINGS WERE NOT PROVIDED. IT IS UNKNOWN WHETHER THE CUSTOMER NOTED A TREND ARROW ON THE DEVICE AT THE TIME OF THE EVENT. THE CUSTOMER INDICATED THE DEATH OF HER BABY OCCURRED ON (B)(6) 2025 WHILE SHE WAS IN "REANIMATION SERVICE" OF THE HOSPITAL. NO AUTOPSY REPORT OR DEATH CERTIFICATE HAS BEEN PROVIDED AT THIS TIME. GIVEN THE INFORMATION PRESENTED AT THIS TIME, IT IS INCONCLUSIVE THAT THE ADC DEVICE HAS LED TO OR CONTRIBUTED TO THE REPORTED FETAL DEATH. Manufacturer Narrative: AN EXTENDED MANUFACTURING REVIEW OF THE REPORTED PRODUCT HAS BEEN PERFORMED. LOT 1001001995 PERFORMED WITHIN SPECIFICATION FOR ALL MEASUREMENTS DURING THE SENSOR MANUFACTURING AND RELEASE TESTING PROCESS. THERE WERE NO NON-CONFORMANCES IN THE SAME TIME PERIOD THAT WAS RELATED TO MANUFACTURE OF LOT 1001001995. ALL MEASUREMENTS REVIEWED WERE WITHIN SPECIFICATIONS. THE MOST PROBABLE ROOT CAUSES ASSOCIATED WITH THIS FAILURE MODE ARE DISCONNECTED, FAULTY OR DAMAGED COMPONENTS, SOFTWARE/DATA CORRUPTION, OR MISUSE. HOWEVER, MITIGATIONS ARE IN PLACE TO REDUCE AND PREVENT SUCH ISSUES. ALL VITAL MANUFACTURING STEPS ARE VALIDATED, MONITORED, AND VERIFIED DURING MANUFACTURING TO ENSURE THE SYSTEM IS IN CONFORMANCE WITH THE VERIFIED DESIGN. ALL VITAL FUNCTIONS ARE MONITORED BY THE SYSTEM AND, WHEN NECESSARY, FUNCTION IS SUSPENDED TO SAFEGUARD AGAINST INACCURATE RESULTS. LABELING IS PROVIDED TO INSTRUCT THE USER ON THE INTENDED USE OF ALL VITAL PARTS OF THE SYSTEM TO MINIMIZE MISUSE. ALL COMPLAINTS AND COMPLAINT TRENDS ARE INVESTIGATED TO DETERMINE IF THERE IS A PRODUCT DEFECT/ DEFICIENCY. IF A PRODUCT DEFECT/ DEFICIENCY IS IDENTIFIED, A RISK EVALUATION IS COMPLETED AND COMPARED TO THE RISK MANAGEMENT REPORT, TO ENSURE THE RISK PROFILE HAS NOT CHANGED. ADDITIONALLY, AS A PART OF ABBOTT¿S POST-MARKET SURVEILLANCE PROCESS, ALL RISK EVALUATIONS WITH ASSOCIATED COMPLAINT DATA ARE REVIEWED ANNUALLY TO DETERMINE IF THE RISK PROFILES HAVE CHANGED AS COMPARED TO THE PRODUCT RISK MANAGEMENT REPORTS. THESE MONITORING PROCESSES ENSURE THAT ALL PRODUCT RISK PROFILES REMAIN ACCEPTABLE AND HAVE A POSITIVE BENEFIT/ RISK RATIO. THE REPORTED PRODUCT IS NOT EXPECTED TO BE RETURNED AS THE CUSTOMER INDICATED THE DEVICE WAS DISCARDED. IN THE EVENT THAT PRODUCT IS RECEIVED, A PHYSICAL INVESTIGATION WILL BE PERFORMED PER ADC'S ESTABLISHED PROCESSES AND PROCEDURES AND A REPORT WILL BE SUBMITTED UPON COMPLETION OF INVESTIGATION. ABBOTT DIABETES CARE RECEIVED AN ANSM USER REPORT (R2518288) IN REGARD TO THE SAME INCIDENT. PLEASE REFER TO ADC CASE ID (B)(4) FOR ALL SUBSEQUENT REPORTS RELATED TO THIS INCIDENT. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED. Manufacturer Narrative: THE REPORTED PRODUCT IS NOT EXPECTED TO BE RETURNED AS THE CUSTOMER INDICATED THE DEVICE WAS DISCARDED. IN THE ABSENCE OF A RETURNED PRODUCT, AN INVESTIGATION HAS BEEN PERFORMED. AN INVESTIGATION HAS BEEN PERFORMED FOR THE REPORTED COMPLAINT AND THERE WAS NO INDICATION THAT THE PRODUCT DID NOT MEET SPECIFICATION. THE DEVICE HISTORY RECORDS (DHRS) FOR THE PRODUCT WERE REVIEWED AND THE DHRS SHOWED THE PRODUCT PASSED ALL TESTS PRIOR TO RELEASE. THE MOST PROBABLE ROOT CAUSES ASSOCIATED WITH THIS FAILURE MODE ARE DISCONNECTED, FAULTY OR DAMAGED COMPONENTS, SOFTWARE/DATA CORRUPTION, OR MISUSE. HOWEVER, MITIGATION'S ARE IN PLACE TO REDUCE AND PREVENT SUCH ISSUES. ALL VITAL MANUFACTURING STEPS ARE VALIDATED, MONITORED, AND VERIFIED DURING MANUFACTURING TO ENSURE THE SYSTEM IS IN CONFORMANCE WITH THE VERIFIED DESIGN. ALL VITAL FUNCTIONS ARE MONITORED BY THE SYSTEM AND, WHEN NECESSARY, FUNCTION IS SUSPENDED TO SAFEGUARD AGAINST INACCURATE RESULTS. LABELING IS PROVIDED TO INSTRUCT THE USER ON THE INTENDED USE OF ALL VITAL PARTS OF THE SYSTEM TO MINIMIZE MISUSE. ALL COMPLAINTS AND COMPLAINT TRENDS ARE INVESTIGATED TO DETERMINE IF THERE IS A PRODUCT DEFECT/ DEFICIENCY. IF A PRODUCT DEFECT/ DEFICIENCY IS IDENTIFIED, A RISK EVALUATION IS COMPLETED AND COMPARED TO THE RISK MANAGEMENT REPORT, TO ENSURE THE RISK PROFILE HAS NOT CHANGED. ADDITIONALLY, AS A PART OF ABBOTT¿S POST-MARKET SURVEILLANCE PROCESS, ALL RISK EVALUATIONS WITH ASSOCIATED COMPLAINT DATA ARE REVIEWED ANNUALLY TO DETERMINE IF THE RISK PROFILES HAVE CHANGED AS COMPARED TO THE PRODUCT RISK MANAGEMENT REPORTS. THESE MONITORING PROCESSES ENSURE THAT ALL PRODUCT RISK PROFILES REMAIN ACCEPTABLE AND HAVE A POSITIVE BENEFIT/ RISK RATIO. THE REPORTED PRODUCT IS NOT EXPECTED TO BE RETURNED AS THE CUSTOMER INDICATED THE DEVICE WAS DISCARDED. IN THE EVENT THAT PRODUCT IS RECEIVED, A PHYSICAL INVESTIGATION WILL BE PERFORMED PER ADC'S ESTABLISHED PROCESSES AND PROCEDURES AND A REPORT WILL BE SUBMITTED UPON COMPLETION OF INVESTIGATION. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22471347&pc=OZD | 1220648-2025-30219 | 2025-06-22 04:00 | Death | ABIOMED, INC. | 2025-07-10 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED PATIENT HAD A CP PUMP PLACED TO SUPPORT THE PATIENT ADMITTED IN WITH ACUTE MYOCARDIAL INFARCTION COMPLICATED BY CARDIOGENIC SHOCK. THE PUMP WAS PLACED BUT, THERE WERE ISSUES WITH SUCTION AND LIMB ISCHEMIA. SUCTION WAS REMEDIED BY REPOSITIONING THE PUMP AND GIVING VOLUME. THE ISCHEMIC LEG WAS TREATED BY RE-PERFUSION FROM THE EXTRACORPOREAL MEMBRANE OXYGENATION. AFTER 66 HOURS OF SUPPORT THE TEAM WITHDREW CARE AND PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22471432&pc=NPT | 2025587-2025-04903 | 2025-04-16 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-10 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Pneumonia; Septic Shock | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT WAS HOSPITALIZED DUE TO WORSENING CONGESTIVE HEART FAILURE (CHF), AND HAD PRE-EXISTING MITRAL STENOSIS (MS) AND TRICUSPID REGURGITATION (TR). TREATMENT FOR THE MS AND TR WERE PLANNED TO OCCUR AFTER THE TRANSCATHETER VALVE IMPLANT PROCEDURE. DURING THE IMPLANT PROCEDURE OF THE TRANSCATHETER AORTIC VALVE, HYPOTENSION WAS OBSERVED AND INTRA-AORTIC BALLOON PUMP (IABP)/EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WERE PERFORMED. ADDITIONALLY, ORAL AND NASAL HEMORRHAGING WERE OBSERVED ON THE SAME DAY AS THE IMPLANT OF THE VALVE. IT WAS REPORTED THAT 2 DAYS FOLLOWING THE IMPLANT OF THE TRANSCATHETER VALVE, THE MITRAL VALVE REPLACEMENT AND TRICUSPID VALVE REPAIR WERE PERFORMED SUCCESSFULLY; HOWEVER, THE PATIENT COULD NOT BE WEANED OFF OF THE CIRCULATORY SUPPORT DEVICES. A TRACHEOSTOMY WAS PERFORMED 12 DAYS FOLLOWING THE IMPLANT OF THE VALVE; HOWEVER, THE PATIENT'S OXYGENATION LEVELS DECREASED AND BACTERIAL PNEUMONIA WAS SUSPECTED FOLLOWING FURTHER EXAMINATION. THE PATIENT WAS TREATED WITH ANTIBIOTICS; HOWEVER, THE INFLAMMATION PERSISTED AND SPUTUM CULTURES DID NOT DETECT ANY BACTERIA. THEREFORE, THE PATIENT WAS TREATED WITH STEROIDS AS POST-BACTERIAL PNEUMONIA ORGANIZING PNEUMONIA WAS SUSPECTED. DURING THIS PROCESS, BACTEREMIA RECURRED AND THE PATIENT'S CONDITION DID NOT IMPROVE. APPROXIMATELY 1 MONTH AND 3 WEEKS FOLLOWING THE IMPLANT OF THE TRANSCATHETER AORTIC VALVE, THE PATIENT DIED DUE TO SEPTIC SHOCK CAUSED BY PNEUMONIA. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. CONTINUATION OF D10: PRODUCT ID {D-EVOLUTFX-2329}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEARTVALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22471936&pc=NIK | 2182207-2025-01772 | 2025-01-01 05:00 | Death | MEDTRONIC NEUROMODULATION | 2025-07-10 04:00 | NIK | MEDTRONIC CRT | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CONDUCTION SYSTEM PACING (CSP) VERSUS BIVENTRICULAR PACING (BIVP). THE AUTHORS DESCRIBED PATIENT DEATHS IN BOTH CSP AND BIVP GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE. THERE WERE PATIENTS IN BOTH GROUPS WHO EXPERIENCED HEART FAILURE HOSPITALIZATIONS. CSP COMPLICATIONS INCLUDED ONE INFECTION, ONE PATIENT WITH VENTRICULAR FIBRILLATION (VF) DURING THE IMPLANT PROCEDURE WHICH REQUIRED ADVANCED CARDIOPULMONARY RESUSCITATION, LEAD DISLODGMENTS, AND ONE LEAD DISCONNECTION WHICH ALL REQUIRED INTERVENTION. BIVP COMPLICATIONS INCLUDED ONE PATIENT EACH WITH PERFORATION, HEMATOMA, INFECTION, AND PHRENIC NERVE STIMULATION, AND LEAD DISLODGMENTS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/69 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CLINICAL RESPONSE TO RESYNCHRONIZATION THERAPY: CONDUCTION SYSTEM PACING VS BIVENTRICULAR PACING. CONSYST-CRT TRIAL. JOURNAL OF THE AMERICAN COLLEGE OF CARDIOLOGY: CLINICAL ELECTROPHYSIOLOGY. 2025. 1-12. DOI: 10.1016/J.JACEP.2025.03.024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22472053&pc=NPT | 2025587-2025-04905 | 2025-05-31 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-10 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Encephalopathy | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH AND 3 WEEKS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT COLLAPSED AND WAS TRANSPORTED IN A STATE OF CARDIOPULMONARY ARREST (CPA). CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED WITH THE ADMINISTRATION OF NASAL OXYGEN, AND SPONTANEOUS BREATHING WAS RESTORED. APPROXIMATELY 2 MONTHS FOLLOWING THE IMPLANT OF THE VALVE, AT A DIFFERENT HOSPITAL, THE PATIENT DEVELOPED HYPOXIC ENCEPHALOPATHY AND DIED DUE TO SENILITY. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22472175&pc=OEV | 1220828-2025-00018 | 2025-06-12 04:00 | Death | AMBU A/S | 2025-07-10 04:00 | OEV | SPUR II | Incomplete or Inadequate Connection | Hypoxia |  |  | 1 | Event Description: A CONNECTION ISSUE WAS OBSERVED BETWEEN THE O2 CONNECTOR AND THE TUBING DURING USE. THE DEVICE INVOLVED WAS AN AIRPRODUCT OXYGEN CYLINDER, WITH A FLOW RATE OF 15 L/MIN. THE SPUR II TUBING NO LONGER CONNECTS PROPERLY, WHICH IS A NEW ISSUE. PREVIOUSLY, THE SPUR II STAYED IN PLACE WITHOUT ANY PROBLEM. IT WAS TESTED ON 7 DIFFERENT CYLINDERS, BUT THE CONNECTION REMAINED UNSTABLE. DURING THE LAST USE, A THIRD PERSON HAD TO HOLD BOTH THE CYLINDER AND THE RESUSCITATOR TO ALLOW PROPER FUNCTIONING, ALTHOUGH THIS IS NOT THE NORMAL INTENDED USE OF THE DEVICE. THE PATIENT UNFORTUNATELY PASSED AWAY. THEY HAD AN UNDERLYING MEDICAL CONDITION, SO WE CANNOT DIRECTLY ATTRIBUTE THE OUTCOME TO THE SPUR II ISSUE. HOWEVER, DUE TO THE CONNECTION PROBLEM, A THIRD PERSON WAS REQUIRED TO HOLD BOTH THE CYLINDER AND THE RESUSCITATOR TO ENSURE IT FUNCTIONED CORRECTLY. THIS SITUATION REDUCED THE NUMBER OF RESCUERS AVAILABLE TO ASSIST WITH OTHER CRITICAL ACTIONS DURING THE INTERVENTION. Manufacturer Narrative: THERE WAS NO SAMPLE RETURNED FOR INVESTIGATION, THEREFORE, THE REPORTED FAILURE ABOUT SPUR II HAS CONNECTION ISSUE WITH AIRPRODUCT OXYGEN CYLINDER COULD NOT BE VERIFIED. SPUR II OXYGEN TUBE CONNECTOR IS DESIGNED TO BE A FEMALE CONNECTOR THAT IS COMPATIBLE WITH A CONNECTOR WHICH IS IN COMPLIANCE WITH EN13544-2. THE CONNECTOR IS MADE BY INJECTION WHICH HAS BEEN VERIFIED PER PROCESS VALIDATION. AFTER INJECTION, QC PERFORMS SAMPLE DIMENSION MEASUREMENT ON INNER DIAMETER. IF IT IS OUT OF SPECIFICATION, IT SHOULD BE EASILY DETECTED. AFTER REVIEWING PRODUCTION RECORDS, NO ABNORMALITY WAS FOUND. THE IFU STATES THAT THE SPUR II CONNECTORS FOLLOW EN 13544-2 STANDARD AND ALSO "WHEN APPLYING EXTERNAL DEVICES TO THE RESUSCITATOR, MAKE SURE TO TEST FOR FUNCTIONALITY AND CONSULT THE INSTRUCTIONS FOR USE ACCOMPANYING THE EXTERNAL DEVICE". SPUR II OXYGEN TUBE CONNECTOR AND AIRPRODUCT OXYGEN CYLINDER CONNECTOR FOLLOW DIFFERENT STANDARDS, THEREFORE THEY ARE INCOMPATIBLE. WE WILL KEEP MONITORING THE MARKET FOR SIMILAR INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22472226&pc=DXY | 2182207-2025-01773 | 2025-01-01 05:00 | Death | MEDTRONIC NEUROMODULATION | 2025-07-10 04:00 | DXY | MEDTRONIC LEAD | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CONDUCTION SYSTEM PACING (CSP) VERSUS BIVENTRICULAR PACING (BIVP). THE AUTHORS DESCRIBED PATIENT DEATHS IN BOTH CSP AND BIVP GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE. THERE WERE PATIENTS IN BOTH GROUPS WHO EXPERIENCED HEART FAILURE HOSPITALIZATIONS. CSP COMPLICATIONS INCLUDED ONE INFECTION, ONE PATIENT WITH VENTRICULAR FIBRILLATION (VF) DURING THE IMPLANT PROCEDURE WHICH REQUIRED ADVANCED CARDIOPULMONARY RESUSCITATION, LEAD DISLODGMENTS, AND ONE LEAD DISCONNECTION WHICH ALL REQUIRED INTERVENTION. BIVP COMPLICATIONS INCLUDED ONE PATIENT EACH WITH PERFORATION, HEMATOMA, INFECTION, AND PHRENIC NERVE STIMULATION, AND LEAD DISLODGMENTS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/69 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CLINICAL RESPONSE TO RESYNCHRONIZATION THERAPY: CONDUCTION SYSTEM PACING VS BIVENTRICULAR PACING. CONSIST-CRT TRIAL. JOURNAL OF THE AMERICAN COLLEGE OF CARDIOLOGY: CLINICAL ELECTROPHYSIOLOGY. 2025. 1-12. DOI: 10.1016/J.JACEP.2025.03.024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22472320&pc=NVN | 2182207-2025-01774 | 2025-01-01 05:00 | Death | MEDTRONIC NEUROMODULATION | 2025-07-10 04:00 | NVN | SELECTSECURE MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CONDUCTION SYSTEM PACING (CSP) VERSUS BIVENTRICULAR PACING (BIVP). THE AUTHORS DESCRIBED PATIENT DEATHS IN BOTH CSP AND BIVP GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE. THERE WERE PATIENTS IN BOTH GROUPS WHO EXPERIENCED HEART FAILURE HOSPITALIZATIONS. CSP COMPLICATIONS INCLUDED ONE INFECTION, ONE PATIENT WITH VENTRICULAR FIBRILLATION (VF) DURING THE IMPLANT PROCEDURE WHICH REQUIRED ADVANCED CARDIOPULMONARY RESUSCITATION, LEAD DISLODGMENTS, AND ONE LEAD DISCONNECTION WHICH ALL REQUIRED INTERVENTION. BIVP COMPLICATIONS INCLUDED ONE PATIENT EACH WITH PERFORATION, HEMATOMA, INFECTION, AND PHRENIC NERVE STIMULATION, AND LEAD DISLODGMENTS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/69 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CLINICAL RESPONSE TO RESYNCHRONIZATION THERAPY: CONDUCTION SYSTEM PACING VS BIVENTRICULAR PACING. CONSYST-CRT TRIAL. JOURNAL OF THE AMERICAN COLLEGE OF CARDIOLOGY: CLINICAL ELECTROPHYSIOLOGY. 2025. 1-12. DOI: 10.1016/J.JACEP.2025.03.024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22472478&pc=NPW | 2015691-2025-05608 | 2025-06-19 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-10 04:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Malposition of Device | Pericardial Effusion; Cardiovascular Insufficiency | P230013 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION OF A TRANSCATHETER TRICUSPID VALVE REPLACEMENT (TTVR) PROCEDURE INVOLVING THE EVOQUE SYSTEM. AN INTRAPROCEDURAL DEATH OCCURRED. THE PATIENT WENT INTO PEA IN THE PROCEDURE. ACUTE RV DYSFUNCTION SUSPECTED MID DEPLOYMENT. CPR WAS PERFORMED PRIOR TO VALVE DEPLOYMENT. THERE WAS SEPTAL PLASTERING AND >15MM LEAFLET TETHERING OF ANTERIOR LEAFLET AND AN INABILITY TO NAVIGATE THE WIRE TO RV FREE SPACE. ADDITIONAL INFORMATION REPORTED THAT INITIALLY THERE WAS GOOD WIRE POSITIONING, BUT IT WAS LOST DURING DILATION, CAUSING THE WIRE TO POP TOWARDS THE ANTERIOR RVOT AREA. IT WAS DECIDED TO REPOSITION WITH THE DELIVERY SYSTEM, WHICH WAS SUCCESSFULLY DONE. SMOKE (SWIRLING BLOOD FLOW) WAS NOTICED IN THE RVOT, ACCOMPANIED BY A BLOOD PRESSURE DROP. UPON NOTICING THE PRESSURE DROP, THE PHYSICIAN CHECKED FOR EFFUSION AND NOTED SEVERELY REDUCED LV AND RV FUNCTION. THERE WAS A PRE-PROCEDURE EFFUSION, AND IT SEEMED TO INCREASE. SWIRLING BLOOD AND POTENTIAL CLOT FORMATION IN THE LV WAS ALSO NOTED. IT WAS DECIDED TO PROCEED WITH VALVE DEPLOYMENT DESPITE THE HEMODYNAMIC CRISIS. AFTER CPR AND MEDICATION, THE VALVE WAS SUCCESSFULLY DEPLOYED WITH FULL LEAFLET CAPTURE. THE DELIVERY SYSTEM WAS SAFELY REMOVED, AND CPR WAS CONTINUED. Manufacturer Narrative: THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED/ADDED: B4, G3, G6, H2, AND H11. ADDITIONAL INFORMATION: PER INTERNAL IMAGING REVIEW, THE GUIDEWIRE APPEARS TO HAVE PERFORATED THE RIGHT VENTRICULAR WALL. FLUOROSCOPY SHOWS THE NOSECONE EXTENDING BEYOND THE RV LEAD ANCHOR. A LARGE PERICARDIAL EFFUSION WITH TAMPONADE DEVELOPED DURING ANCHOR DEPLOYMENT. NO DEVICE MALFUNCTION IDENTIFIED. ROOT CAUSE ATTRIBUTED TO PROCEDURAL FACTORS-SPECIFICALLY, INADEQUATE GUIDEWIRE MANAGEMENT AND DEPTH CONTROL. A POTENTIAL CONTRIBUTING FACTOR INCLUDES PATIENT-SPECIFIC ANATOMY, NOTABLY RV TRABECULATIONS. THE MANUFACTURER'S INVESTIGATION IS ONGOING AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22472630&pc=DSQ | 2916596-2025-04617 | 2025-06-09 04:00 | Death | THORATEC CORPORATION | 2025-07-10 04:00 | DSQ | HEARTMATE 3 LVAS | Adverse Event Without Identified Device or Use Problem | Dyspnea; Fever; Confusion/ Disorientation; Unspecified Respiratory Problem | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE 87-YEAR-OLD FEMALE PRESENTED FROM AN OUTSIDE HOSPITAL EMERGENCY ROOM AFTER FAMILY HAD CALLED WITH INCREASED CONCERN OF CONFUSION AND DIFFICULTY BREATHING. THE PATIENT WAS TRANSFERRED TO THEIR IMPLANTING HOSPITAL WITH WORSENING SHORTNESS OF BREATH, COUGH WITH CLEAR SPUTUM AND CONFUSION. LEUKOCYTOSIS WITH LEFT SHIFT WAS REPORTED WITH FEVER AND DELIRIUM. PER THE PHYSICIAN, IT SOUNDED MORE LIKE AN INFECTIOUS PROCESS OVERALL. UPON ARRIVAL, LABS WERE SENT AND A CHEST X-RAY (CXR) WAS COMPLETED FOR FURTHER EVALUATION, WHICH WAS REMARKABLE. BLOOD CULTURES AND URINE STUDIES WERE COMPLETED. AFTER DISCUSSIONS WITH THE OVERNIGHT TEAM, THE PATIENT AND FAMILY CHOSE TO CHANGE THE CODE STATUS TO DO NOT ATTEMPT RESUSCITATION/CONTINUE OTHER TREATMENTS. THE PATIENT WAS GIVEN INTRAVENOUS LASIX (FUROSEMIDE) WITH GOOD URINE OUTPUT BUT MENTAL STATUS CONTINUED TO DECLINE TO WHERE THEY WERE ONLY ALERT TO A PERSON. THE DECISION WAS MADE TO TRANSITION TO DO NOT RESUSCITATE AND TURN OFF THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) TO BE CONSISTENT WITH THE PATIENT'S PREVIOUSLY VERBALIZED WISHES TO FAMILY, NOT WANTING "TO LIVE THIS WAY". IMPLANTABLE CARDIAC DEFIBRILLATOR TACHYTHERAPIES WERE TURNED OFF. THE PATIENT WAS STARTED ON A MORPHINE DRIP FOR COMFORT AND THE LVAD WAS TURNED OFF AT 1512. TIME OF DEATH WAS 1516 ON (B)(6) 2025 WITH FAMILY AT BEDSIDE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING OTHER NEUROLOGICAL EVENTS (NOT STROKE-RELATED) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS NEUROLOGICAL DYSFUNCTION AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22472926&pc=NPT | 2025587-2025-04909 | 2025-06-26 04:00 | Death | MEDTRONIC IRELAND | 2025-07-10 04:00 | NPT | EVOLUT FX DCS | Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Rupture; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS EVFXPLUS-29, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED TWICE USING A 20 MM NON-MEDTRONIC BALLOON. THE VALVE AND DELIVERY CATHETER SYSTEM WERE INSERTED INTO THE PATIENT AND ADVANCED TO THE ANNULUS. UPON DEPLOYMENT TO 80%, THE PATIENT'S BLOOD PRESSURE DROPPED. AS A RESULT, LIGHT CARDIOPULMONARY RESUSCITATION (CPR) COMPRESSIONS WERE COMMENCED DURING THE FINAL DEPLOYMENT. SUBSEQUENTLY, A PERICARDIAL EFFUSION WAS IDENTIFIED AND PERICARDIOCENTESIS WAS PERFORMED AS AN INTERVENTION. CARDIOPULMONARY RESUSCITATION CONTINUED WITH "CODE BLUE" PROTOCOL INITIATED. THE PATIENT DIED DESPITE INTERVENTION. PER THE PHYSICIAN, THE PROCEDURE CONTRIBUTED TO THE DEATH. IT IS UNKNOWN WHETHER THE DEVICE CONTRIBUTED TO THE CAUSE OF DEATH.Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS EVFXPLUS-29, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED TWICE USING A 20 MM NON-MEDTRONIC BALLOON. THE VALVE AND DELIVERY CATHETER SYSTEM WERE INSERTED INTO THE PATIENT AND ADVANCED TO THE ANNULUS. UPON DEPLOYMENT TO 80%, THE PATIENT'S BLOOD PRESSURE DROPPED. AS A RESULT, LIGHT CARDIOPULMONARY RESUSCITATION (CPR) COMPRESSIONS WERE COMMENCED DURING THE FINAL DEPLOYMENT. SUBSEQUENTLY, A PERICARDIAL EFFUSION WAS IDENTIFIED AND PERICARDIOCENTESIS WAS PERFORMED AS AN INTERVENTION. CARDIOPULMONARY RESUSCITATION CONTINUED WITH "CODE BLUE" PROTOCOL INITIATED. THE PATIENT DIED DESPITE INTERVENTION. PER THE PHYSICIAN, THE PROCEDURE CONTRIBUTED TO THE DEATH. IT IS UNKNOWN WHETHER THE DEVICE CONTRIBUTED TO THE CAUSE OF DEATH. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE CAUSE OF DEATH WAS AN ANNULAR RUPTURE. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT FX PLUS VALVE; PRODUCT ID EVFXPLUS-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. ADDED SECOND PARAGRAPH. D3. D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22473142&pc=MFE | 2029214-2025-01559 | 2024-10-01 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-07-10 04:00 | MFE | ONYX | Adverse Event Without Identified Device or Use Problem | Rupture |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ILIAC BRANCH DEVICES IN THE REPAIR OF RUPTURED AORTO-ILIAC ANEURYSMS: A MULTICENTER STUDY THE FOLLOWING MEDTRONIC DEVICES WERE USED: ONYX LIQUID EMBOLIC AGENT DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. AMONG ALL PATIENTS ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDES: TECHNICAL SUCCESS WAS ACHIEVED IN ALL CASES, BUT 10 (42%) PATIENTS DIED WITHIN 30 DAYS. THE OVERALL SURVIVAL AT 12 AND 36 MONTHS WAS 48±11% AND 38±10%, RESPECTIVELY. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: KARELIS, A., SONESSON, B., GALLITTO, E., TSILIMPARIS, N., FORSELL, C., LEONE, N., SILINGARDI, R., MESNARD, T., SOBOCINSKI, J., ISERNIA, G., RESCH, T., GARGIULO, M., DIAS, N.V.. ILIAC BRANCH DEVICES IN THE REPAIR OF RUPTURED AORTO-ILIAC ANEURYSMS: A MULTICENTER STUDY. JOURNAL OF ENDOVASCULAR THERAPY VOL 3 2024. DOI: 10.1177/15266028221149922 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. PATIENT AGE AND SEX IS REFLECTIVE OF THE MEAN PATIENT DATA IN THE ARTICLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22473672&pc=QFG | 3013756811-2025-164272 | 2025-06-08 04:00 | Death | TANDEM DIABETES CARE | 2025-07-10 04:00 | QFG | TANDEM MOBI INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Diabetic Ketoacidosis | K233044 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025 DUE TO A HEART ATTACK AND DIABETIC KETOACIDOSIS. THE CUSTOMER WAS WEARING THE PUMP AT THE TIME OF DEATH. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22474411&pc=DSQ | 2916596-2025-04310 | 2025-04-30 04:00 | Death | THORATEC CORPORATION | 2025-07-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Septic Shock; Coma; Hydrocephalus; Hemorrhagic Stroke; Embolism/Embolus; Drug Resistant Bacterial Infection; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PUMP WAS NOT EXPLANTED.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT HAD A MAJOR BACTERIAL INFECTION "INSIDE THE PUMP" WITH POSITIVE BLOOD CULTURES. THE PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2025 AND TREATED WITH DRUG THERAPY. DURING THE ADMISSION ON (B)(6) 2025 THE PATIENT HAD AN EMBOLISM IN THEIR RIGHT CEREBRAL HEMISPHERE WHICH WAS DIAGNOSED WITH A COMPUTED TOMOGRAPHY (CT) SCAN. THE PATIENT HAD A STROKE AND WENT INTO A COMA. THIS NEUROLOGICAL DYSFUNCTION WAS DUE TO CEREBROVASCULAR DISORDER ASSOCIATED WITH SEPTIC SHOCK FROM THE PUMP INFECTION. THE SEPTIC SHOCK LED TO A CEREBRAL INFARCTION WHICH LED TO A CEREBRAL HEMORRHAGE EVENTUALLY LEADING TO HYDROCEPHALUS, RESULTING IN THE INABILITY TO MAINTAIN CIRCULATORY DYNAMICS. THE PATIENT WAS TREATED WITH DRUG THERAPY, INTUBATION AND MECHANICAL VENTILATION. THE PATIENT SUBSEQUENTLY PASSED AWAY. THE DEVICE WAS FUNCTIONING PROPERLY AT THE TIME OF DEATH.Event Description: IT WAS ADDITIONALLY REPORTED THAT AT AROUND 11:00 ON (B)(6) 2025 WHEN THE PATIENT WENT SHOPPING, THEY COMPLAINED OF CHEST PAIN AND LOOKED DISTRESSED. LATER, AT AROUND 17:45, THE PATIENT WAS FOUND LYING ON THEIR STOMACH AND REQUESTED AN AMBULANCE. WHEN THE AMBULANCE ARRIVED, THERE WAS ANISOCORIA. THE PATIENT WAS TRANSPORTED TO A NEARBY OUTSIDE HOSPITAL. THE HEAD COMPUTED TOMOGRAPHY (CT) SHOWED NO CEREBRAL HEMORRHAGE OR CEREBRAL INFARCTION. THE PATIENT WAS TRANSFERRED TO THE CLINIC UNDER INTUBATION. AT THE TIME OF TRANSFER, AN INCREASED INFLAMMATORY REACTION WAS OBSERVED, AND A RELAPSE OF SEPSIS WAS SUSPECTED. THE BLOOD CULTURE PERFORMED ON ADMISSION DETECTED CORYNEBACTERIUM. DISTURBANCE OF CONSCIOUSNESS WAS PROLONGED. ON (B)(6) 2025, A HEAD CT SHOWED SUSPECTED INFARCTION DUE TO INFECTIVE ENDOCARDITIS (IE) COMPLICATED BY INFECTION. AN ELECTROENCEPHALOGRAM (EEG) SHOWED SLOW ALPHA IN THE CEREBRUM, AND THE PATIENT WAS DIAGNOSED AS NOT BRAIN DEAD. THE FOLLOWING DAY, A HEAD CT SHOWED CEREBRAL HEMORRHAGE AND HYDROCEPHALUS. THE FAMILY WAS NOTIFIED, AND IT WAS DECIDED NOT TO AGGRESSIVELY TREAT THE PATIENT. ON (B)(6) 2025, IN THE MORNING, BLOOD PRESSURE DECREASED, AND THE FAMILY REQUESTED TO STOP THE PUMP. THE PUMP WAS STOPPED. BLOOD PRESSURE DECREASED AND BRADYCARDIA DEVELOPED, AND THEN CARDIAC ARREST OCCURRED. DEATH WAS CONFIRMED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATION FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION", LISTS POTENTIAL ADVERSE EVENTS INCLUDING INFECTION (LOCAL, DRIVELINE OR PUMP POCKET INFECTION), SEPSIS, STROKE, BLEEDING, OTHER NEUROLOGICAL EVENTS (NOT STROKE-RELATED), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER "ANTICOAGULATION"), OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INR RANGE FOR PATIENTS USING THE HEARTMATE 3 LVAS, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. THIS SECTION ALSO LISTS INFECTION AND NEUROLOGICAL DYSFUNCTION AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING INFECTION PREVENTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION E: CORRECTION. SECTION G2: CORRECTION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22475067&pc=NPT | 2015691-2025-05613 | 2025-06-17 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-07-10 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem; Device Stenosis |  | P140031 |  | 1 | Event Description: AS REPORTED FROM PATIENT SUPPORT, ON THE SAME DAY FOLLOWING A TRANSFEMORAL AORTIC VALVE REPLACEMENT INVOLVING A 20MM SAPIEN 3 ULTRA VALVE IMPLANTED WITHIN A MECHANICAL VALVE, THE PATIENT PASSED AWAY. THE PATIENT PREVIOUSLY HAD AN AORTIC VALVE REPLACEMENT WITH THE PERIMOUNT MAGNA EASE 9 YEARS AND 6 MONTHS PRIOR AND WENT FOR A SCHEDULED VALVE-IN-VALVE TAVR WITH THE IMPLANTATION OF THE SAPIEN 3 ULTRA AT THE SAME HOSPITAL. AFTER THE PATIENT CAME OUT OF THE PROCEDURE AND WAS IN THE RECOVERY UNIT, THE PATIENT WENT ACUTELY UNRESPONSIVE AND WAS RUSHED TO THE OPERATING ROOM. PER THE CARE ADVOCATE, 'THEY WORKED ON HER FOR A WHILE BUT WERE UNABLE TO BRING HER BACK'. EVENTUALLY, THE PATIENT PASSED AWAY IN THE HOSPITAL THAT SAME DAY. AFTERWARDS, THE HEART TEAM TOLD THE CARE ADVOCATE THAT THEY BELIEVE 'A PIECE OF PLAQUE WAS DISLODGED FROM THE PRIOR VALVE AND CAUSED HER TO HAVE A HEART ATTACK'. ACCORDING TO PATIENT'S CERTIFICATE, CAUSE OF DEATH STATES 'LEFT MAIN CORONARY ARTERY OCCLUSION DUE TO SEVERE AORTIC STENOSIS'.Event Description: AS REPORTED FROM PATIENT SUPPORT, THE PATIENT'S CARE ADVOCATE RECEIVED THE PATIENT'S IMPLANT CARDS BUT NOTIFIED PATIENT SUPPORT OF THE PATIENT'S PASSING THAT OCCURRED ON THE SAME DAY FOLLOWING THE TRANSFEMORAL AORTIC VALVE REPLACEMENT INVOLVING A 20MM SAPIEN 3 ULTRA VALVE IMPLANTED WITHIN A MECHANICAL VALVE. SHE EXPLAINED THAT THE PATIENT PREVIOUSLY HAD AN AORTIC VALVE REPLACEMENT WITH THE PERIMOUNT MAGNA EASE 9 YEARS AND 6 MONTHS PRIOR. RECENTLY, THE PATIENT WENT FOR A SCHEDULED VALVE-IN-VALVE TAVR WITH THE IMPLANTATION OF THE SAPIEN 3 ULTRA AT THE SAME HOSPITAL. ACCORDING TO CARE ADVOCATE, THE INTERVENTIONALIST INITIALLY NOTIFIED THAT THE PATIENT CAME OUT OF THE PROCEDURE AND WAS IN THE RECOVERY UNIT. HOWEVER, THE PATIENT WENT ACUTELY UNRESPONSIVE AND WAS RUSHED TO THE OPERATING ROOM. PER CARE ADVOCATE, 'THEY WORKED ON HER FOR A WHILE BUT WERE UNABLE TO BRING HER BACK'. EVENTUALLY, THE PATIENT PASSED AWAY IN THE HOSPITAL THAT SAME DAY. AFTERWARDS, THE HEART TEAM TOLD THE CARE ADVOCATE THAT THEY BELIEVE 'A PIECE OF PLAQUE WAS DISLODGED FROM THE PRIOR VALVE AND CAUSED HER TO HAVE A HEART ATTACK'. ACCORDING TO PATIENT'S CERTIFICATE, CAUSE OF DEATH STATES 'LEFT MAIN CORONARY ARTERY OCCLUSION DUE TO SEVERE AORTIC STENOSIS'. THE CARE ADVOCATE DID NOT ASK TO SPEAK WITH OUR PHYSICIANS AND DID NOT MAKE ANY ALLEGATIONS TO OUR DEVICES. PER THE MEDICAL RECORD, AFTER THE PATIENT UNDERWENT A VALVE-IN-VALVE PROCEDURE THE VALVE WAS SUCCESSFULLY DEPLOYED WITH NO PARAVALVULAR LEAK FOLLOWING ONE POST-DILATION. NO DEVICE MALFUNCTIONS OR PROCEDURAL COMPLICATIONS WERE REPORTED. THE PATIENT WAS TRANSFERRED TO THE ICU IN STABLE CONDITION. WHILE IN THE ICU, THE PATIENT EXPERIENCED A PULSELESS ELECTRICAL ACTIVITY (PEA) ARREST AND WAS URGENTLY TAKEN BACK TO THE CATHETERIZATION LAB. A TOTAL OCCLUSION OF THE LEFT ANTERIOR DESCENDING (LAD) ARTERY WAS OBSERVED, WHICH HAD NOT BEEN PRESENT IMMEDIATELY POST-TAVR. THE PHYSICIAN DETERMINED THAT THE OCCLUSION WAS NOT AMENABLE TO INTERVENTION, AS THEY WERE UNABLE TO PASS A WIRE. BY THAT TIME, THE PATIENT WAS NO LONGER A CANDIDATE FOR STERNOTOMY OR EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) DUE TO LACK OF PERFUSION. THE FAMILY WAS UPDATED, AND THE PATIENT PASSED AWAY SHORTLY THEREAFTER. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED, DUE TO MEDICAL RECORDS RECEIVED. B4, G3, AND H2 HAVE BEEN UPDATED. B1, B2, B5, B7, H2, H6: DEVICE PROBLEM, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS AND H11 HAVE BEEN CORRECTED. PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ALTHOUGH THE EXACT CAUSE IS NOT KNOWN BASED ON THE AVAILABLE INFORMATION, INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT FACTORS (DIABETES, HYPERTENSION, HYPERLIPIDEMIA) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. THIS IS ONE OF TWO MANUFACTURING REPORTS BEING SUBMITTED FOR THIS CASE. PLEASE REFERENCE RELATED MANUFACTURER REPORT. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED, DUE TO ACCIDENTAL OMISSION OF TYPE OF INVESTIGATION FOUND DURING ADMINISTRATIVE REVIEW. H6: INVESTIGATION FINDINGS AND H11 HAVE BEEN CORRECTED. B4, G3, G6, H2 HAVE BEEN UPDATED. PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THIS EVENT MAY HAVE BEEN CAUSED BY A PIECE OF PLAQUE THAT BECAME DISLODGED FROM THE PRIOR STENOTIC SURGICAL VALVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED, FOLLOWING ADMINISTRATIVE REVIEW. B4, G3, G6 AND H2 HAVE BEEN UPDATED. H11 HAS BEEN CORRECTED. PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THIS EVENT MAY HAVE BEEN CAUSED BY A PIECE OF PLAQUE THAT BECAME DISLODGED FROM THE PRIOR STENOTIC SURGICAL VALVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. THIS IS ONE OF TWO MANUFACTURING REPORTS BEING SUBMITTED FOR THIS CASE. PLEASE REFERENCE RELATED MANUFACTURER REPORT. Manufacturer Narrative: THE INVESTIGATION FOR THIS REPORT IS ONGOING. THIS IS ONE OF TWO MANUFACTURING REPORTS BEING SUBMITTED FOR THIS CASE. PLEASE REFERENCE RELATED MANUFACTURER REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22475080&pc=PNJ | 9612164-2025-03400 | 2025-06-20 04:00 | Death | MEDTRONIC IRELAND | 2025-07-10 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade; Cardiac Perforation | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT INTRA OPERATIVELY DURING THE PLACEMENT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE PATIENT EXPERIENCED CARDIAC TAMPONADE AND CARDIAC PERFORATION. BLOOD WAS DRAINED. PATIENT LEFT THE OPERATING ROOM UNDER STABLE CONDITION. LATER, PATIENT CODED IN INTENSIVE CARE UNIT (ICU) AFTER THE PROCEDURE, EXPERIENCED PULSELESS ELECTRICAL ACTIVITY (PEA) ARREST. EVENTUALLY, THE PATIENT PASSED AWAY. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6)). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22475082&pc=FKX | 0008030665-2025-01570 | 2025-06-28 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-07-10 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K181108 |  | 1 | Event Description: A NURSE REPORTED A PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT). DURING THE INTAKE PROCESS, IT WAS REPORTED THE PATIENT WAS UNDERGOING CCPD THERAPY WHEN PATIENT EXPIRED. FOLLOW-UP WITH THE PATIENTS¿ HOME PROGRAM MANAGER (HPM) CONFIRMED THE PATIENT EXPIRED AT HOME WHILE UNDERGOING CCPD THERAPY ON (B)(6) 2025. THE PATIENT REPORTEDLY EXPIRED WHILE SLEEPING AND WAS DISCOVERED BY THE PATIENT¿S SPOUSE. THE HPM REPORTED THE PATIENT¿S PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST AND WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION.Event Description: A NURSE REPORTED A PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT). DURING THE INTAKE PROCESS, IT WAS REPORTED THE PATIENT WAS UNDERGOING CCPD THERAPY WHEN PATIENT EXPIRED. FOLLOW-UP WITH THE PATIENTS¿ HOME PROGRAM MANAGER (HPM) CONFIRMED THE PATIENT EXPIRED AT HOME WHILE UNDERGOING CCPD THERAPY ON (B)(6) 2025. THE PATIENT REPORTEDLY EXPIRED WHILE SLEEPING AND WAS DISCOVERED BY THE PATIENT¿S SPOUSE. THE HPM REPORTED THE PATIENT¿S PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST AND WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. Manufacturer Narrative: CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING CCPD THERAPY WHEN HE EXPIRED. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY CANNOT FIRMLY BE ESTABLISHED. HOWEVER, THE HPM REPORTED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (INCLUDING SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. Manufacturer Narrative: ADDITIONAL INFORMATION: D.9.,H.3. PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL ANALYSIS. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGN OF PHYSICAL DAMAGE.THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT.THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA.HERE WERE NOT BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE.SYSTEM AIR LEAK TEST PASSED.VALVE ACTUATION TEST PASSED.A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THERE WERE VISUAL INDICATIONS OF DRIED FLUID UNDER THE PUMP ASSEMBLY ON THE BOTTOM COVER. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED.THERE WERE NOT DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S). UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22476424&pc=OZP | 2032227-2025-209445 | 2025-06-17 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-10 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Chest Pain; Myocardial Infarction; Paresis | P160017 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY AT THE HOSPITAL ON (B)(6) 2025. THE CUSTOMER WAS ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS DIABETES AND GASTROPARESIS, BUT THE CHEST PAIN AND HEART ATTACK LED TO HOSPITALIZATION. THE CUSTOMER¿S BLOOD GLUCOSE WAS 200 MG/DL. THE EVENT INVOLVED PRODUCT(S) MMT-1884, MMT-242A, MMT-7841ZN, MMT-7040MA, AND MMT-332A. TROUBLESHOOTING WAS PERFORMED. THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF DEATH. THE CUSTOMER WAS USING SENSORS. THE INSULIN PUMP WAS LAST WORN ON (B)(6) 2025, PRIOR TO PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-242A, MMT-7841ZN, MMT-7040MA, AND MMT-332A. MMT-1884 WAS REQUESTED, AND THE CUSTOMER RESPONSE WAS THE DEVICE WILL BE RETURNED. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22455354&pc=LPB | 2029046-2025-02227 | 2025-06-12 04:00 | Death | BIOSENSE WEBSTER INC | 2025-07-09 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade | P030031 |  | 1 | Event Description: THE REPORT INDICATED THAT DURING CARDIAC ABLATION FOR AFIB ¿ PERSISTENT WITH A THERMOCOOL SMARTTOUCH SF CATHETER, THE PATIENT EXPERIENCED A DROP IN BLOOD PRESSURE, TAMPONADE TREATED WITH PERICARDIAL DRAINAGE. THE PATIENT HAD CARDIAC ARREST, REQUIRED CARDIAC MASSAGE, RESUSCITATION WITHOUT SUCCESS, AND THE PATIENT DIED. DURING CARDIAC ABLATION, THE PATIENT HAD TAMPONADE DETECTED DUE TO A DROP IN PATIENT BLOOD PRESSURE. TRANSTHORACIC ULTRASOUND CONFIRMED TAMPONADE AND PERICARDIAL DRAINAGE. THE PATIENT HAD CARDIAC ARREST WHICH OCCURRED ON THE TABLE. CARDIAC MASSAGE AND RESUSCITATION WERE PERFORMED WITHOUT SUCCESS. THE PATIENT ULTIMATELY PASSED AWAY. POSSIBLE CAUSE NOTED WAS ABLATION IN THE CORONARY SINUS. ¿BIOSENSE MATERIAL IS NOT INCRIMINATED.¿ ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THAT THE FORCE VISUALIZATION FEATURES USED WERE GRAPH, DASHBOARD, VECTOR, AND VISITAG. THE VISITAG MODULE PARAMETERS FOR STABILITY USED WERE 3 MM - 3SEC. THE COLOR OPTION USED PROSPECTIVELY WAS TAG INDEX, TAG SIZE 3. THE NUMBER OF PERFORMED RADIOFREQUENCY (RF) ABLATIONS WERE 97 IN THE LEFT ATRIUM, CORONARY SINUS 15, AND RIGHT ATRIUM 12. THERE WAS NO EVIDENCE OF STEAM POP, APART FROM BLOOD PRESSURE DROP. THE FLOW SETTINGS WERE HIGH FLOW BETWEEN 8 AND 15 ML, LOW FLOW 2ML, PRE RF TIME 2SECS, POST RF TIME 1SEC, AND LOW FLUID WARNING 100ML. THE FILTER FOR VISITAG USED WERE FORCE: 3G. FORCE OVER TIME: 25%, AND TAG INDEX: LOW 350 - MAX 450. CORRECT CATHETER SETTINGS WERE SELECTED ON THE GENERATOR. THE PUMP SWITCHED FROM ¿LOW¿ TO ¿HIGH¿ FLOW DURING ABLATION. NO ERROR MESSAGES WERE OBSERVED ON BIOSENSE WEBSTER EQUIPMENT DURING THE PROCEDURE. THE PATIENT DID NOT REQUIRE CARDIAC SURGERY. ONE TRANSSEPTAL PUNCTURE SITE WAS PERFORMED DURING THE PROCEDURE. Manufacturer Narrative: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER LOT 31626438L AND NO INTERNAL ACTION RELATED TO THE COMPLAINT WAS FOUND DURING THE REVIEW. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22456075&pc=HRS | 0009613350-2025-00558 | 2024-08-27 04:00 | Death | ZIMMER GMBH | 2025-07-09 04:00 | HRS | UNKNOWN NCB PLATE | Adverse Event Without Identified Device or Use Problem |  | NA |  | 1 | Event Description: A JOURNAL ARTICLE WAS RETRIEVED FROM ORTHOPAEDICS & TRAUMATOLOGY: SURGERY & RESEARCH (2025) THAT REPORTED A RETROSPECTIVE STUDY FROM FRANCE. THE PURPOSE OF THE STUDY WAS TO PRESENT THE RESULTS OF NCB-PP® LOCKING PLATES IN THE MANAGEMENT OF PERIPROSTHETIC FEMORAL FRACTURES. THE STUDY REVIEWED 89 PATIENTS WITH PERIPROSTHETIC FEMUR FRACTURES BETWEEN JANUARY 2014 AND SEPTEMBER 2022 (74 THA, 11 TKA, AND 4 INTERPROSTHETIC). PATIENTS WERE THEN DIVIDED INTO 2 GROUPS ACCORDING TO THE TIME TO POSTOPERATIVE FULL WEIGHT BEARING: "IMMEDIATE" (30) AND "DELAYED" (59) (A MINIMUM OF 6 WEEKS POST-OPERATIVELY). ALL FRACTURES UNDERWENT ISOLATED NCB-PP® PLATE OSTEOSYNTHESIS USING THE ORIF (OPEN REDUCTION INTERNAL FIXATION) TECHNIQUE VIA A SUB-VASTUS LATERAL APPROACH WITH AN NCB-PP PLATE AND AN UNSPECIFIED NUMBER OF SCREWS. ADDITIONALLY, 38 PATIENTS ALSO RECEIVED CERCLAGE-READY CABLES (ZIMMER BIOMET). THE STUDY POPULATION HAD A MEAN AGE OF 81 YEARS AT TIME OF SURGERY (RANGE 28-99); (62 FEMALES, 27 MALES). MEAN FOLLOW-UP TIME WAS 14.6 MONTHS (RANGE 12-65 MONTHS). THE STUDY REPORTED THAT 1 PATIENT ACQUIRED AN INFECTION AND UNDERWENT SURGICAL LAVAGE WITH APPROPRIATE ANTIBIOTIC THERAPY. THE PATIENT EXPIRED AT 6 MONTHS POSTOP. ATTEMPTS HAVE BEEN MADE, AND NO FURTHER INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: (B)(4). B2: THE PATIENT EXPIRED AT 6 MONTHS POSTOP. THE DATE IS CURRENTLY UNKNOWN. ATTEMPTS HAVE BEEN MADE TO GATHER THIS INFORMATION, AND NO FURTHER INFORMATION HAS BEEN PROVIDED. B3: EVENT DATE IS THE PUBLICATION DATE OF THE ARTICLE. D4: THE PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER IS NOT APPLICABLE AS THE DEVICE'S PRODUCT NUMBER IS UNKNOWN. D10: UNSPECIFIED NUMBER OF UNKNOWN SCREWS. G2: FRANCE. JOURNAL ARTICLE CITATION: BATISTE SANTONI ET AL., ORTHOPAEDICS & TRAUMATOLOGY: SURGERY & RESEARCH, HTTPS://DOI.ORG/10.1016/J.OTSR.2024.104009. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MDR WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22456137&pc=HWC | 0009613350-2025-00559 | 2024-08-27 04:00 | Death | ZIMMER GMBH | 2025-07-09 04:00 | HWC | UNKNOWN SCREWS | Adverse Event Without Identified Device or Use Problem |  | NA |  | 1 | Event Description: A JOURNAL ARTICLE WAS RETRIEVED FROM ORTHOPAEDICS & TRAUMATOLOGY: SURGERY & RESEARCH (2025) THAT REPORTED A RETROSPECTIVE STUDY FROM FRANCE. THE PURPOSE OF THE STUDY WAS TO PRESENT THE RESULTS OF NCB-PP® LOCKING PLATES IN THE MANAGEMENT OF PERIPROSTHETIC FEMORAL FRACTURES. THE STUDY REVIEWED 89 PATIENTS WITH PERIPROSTHETIC FEMUR FRACTURES BETWEEN JANUARY 2014 AND SEPTEMBER 2022 (74 THA, 11 TKA, AND 4 INTERPROSTHETIC). PATIENTS WERE THEN DIVIDED INTO 2 GROUPS ACCORDING TO THE TIME TO POSTOPERATIVE FULL WEIGHT BEARING: "IMMEDIATE" (30) AND "DELAYED" (59) (A MINIMUM OF 6 WEEKS POST-OPERATIVELY). ALL FRACTURES UNDERWENT ISOLATED NCB-PP® PLATE OSTEOSYNTHESIS USING THE ORIF (OPEN REDUCTION INTERNAL FIXATION) TECHNIQUE VIA A SUB-VASTUS LATERAL APPROACH WITH AN NCB-PP PLATE AND AN UNSPECIFIED NUMBER OF SCREWS. ADDITIONALLY, 38 PATIENTS ALSO RECEIVED CERCLAGE-READY CABLES (ZIMMER BIOMET). THE STUDY POPULATION HAD A MEAN AGE OF 81 YEARS AT TIME OF SURGERY (RANGE 28-99); (62 FEMALES, 27 MALES). MEAN FOLLOW-UP TIME WAS 14.6 MONTHS (RANGE 12-65 MONTHS). THE STUDY REPORTED THAT 1 PATIENT ACQUIRED AN INFECTION AND UNDERWENT SURGICAL LAVAGE WITH APPROPRIATE ANTIBIOTIC THERAPY. THE PATIENT EXPIRED AT 6 MONTHS POSTOP. ATTEMPTS HAVE BEEN MADE, AND NO FURTHER INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: (B)(4). B2: THE PATIENT EXPIRED AT 6 MONTHS POSTOP. THE DATE IS CURRENTLY UNKNOWN. ATTEMPTS HAVE BEEN MADE TO GATHER THIS INFORMATION, AND NO FURTHER INFORMATION HAS BEEN PROVIDED. B3: EVENT DATE IS THE PUBLICATION DATE OF THE ARTICLE. D4: THE PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER IS NOT APPLICABLE AS THE DEVICE'S PRODUCT NUMBER IS UNKNOWN. D10: UNKNOWN NBC PLATE G2: FRANCE JOURNAL ARTICLE CITATION: BATISTE SANTONI ET AL., ORTHOPAEDICS & TRAUMATOLOGY: SURGERY & RESEARCH, HTTPS://DOI.ORG/10.1016/J.OTSR.2024.104009. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MDR WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22456224&pc=HRS | 0009613350-2025-00560 | 2024-10-02 04:00 | Death | ZIMMER GMBH | 2025-07-09 04:00 | HRS | UNKNOWN NCB PLATE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | NA |  | 7 | Event Description: A JOURNAL ARTICLE WAS RETRIEVED FROM THE JOURNAL OF BONE AND JOINT SURGERY (2025) THAT REPORTED A STUDY CONDUCTED IN CANADA. THE PURPOSE OF THE STUDY WAS TO TEST THE HYPOTHESIS IN A COMPARATIVE EFFECTIVENESS CLINICAL TRIAL THAT FAR CORTICAL LOCKING (FCL) CONSTRUCTS LEAD TO IMPROVED FRACTURE-HEALING AS COMPARED WITH THAT ACHIEVED WITH TRADITIONAL LOCKING PLATES. THE TRIAL WAS AN INTERNATIONAL, MULTICENTER, RANDOMIZED CONTROLLED CLINICAL TRIAL PREFORMED ACROSS 16 PARTICIPATING CENTERS IN CANADA AND THE UNITED STATES. THE STUDY REVIEWED 167 PATIENTS (84 FCL SCREWS/83 STANDARD SCREWS) WITHIN A BROAD, HETEROGENOUS POPULATION THAT EXPERIENCED AO/OTA TYPE 33A OR 33C DISTAL FEMORAL FRACTURES. PATIENTS WERE RANDOMLY ASSIGNED TO TREATMENT WITH FCL SCREW FIXATION (NCB DISTAL FEMUR PLATE WITH NCB MOTIONLOC SCREWS; ZIMMER BIOMET) OR STANDARD SCREW FIXATION (NCB DISTAL FEMUR PLATE WITH STANDARD NCB SCREWS; ZIMMER BIOMET). THE FCL GROUP REQUIRED THE USE OF 4 BICORTICAL MOTIONLOC SCREWS FOR FIXATION IN THE SHAFT, WITH NO SPECIFIED WORKING LENGTH. THE STANDARD GROUP ALLOWED LOCKING SCREWS, NON-LOCKING SCREWS, AND HYBRID CONSTRUCTS (I.E., A MIXTURE OF LOCKING AND NON-LOCKING SCREWS) AND ANY WORKING LENGTH. THE INDICATION FOR SURGERY WAS A MECHANISM OF INJURY THAT INCLUDES THE FOLLOWING FOR THE FCL SCREW GROUP/STANDARD SCREW GROUP: DRIVER/PASSENGER IN MOTOR VEHICLE ACCIDENT (9/14), PEDESTRIAN OF MOTOR VEHICLE ACCIDENT (5/5), MOTORCYCLE ACCIDENT (8/4), FALL (58/54), DIRECT TRAUMA (1/2), AND OTHER (3/1). THE STUDY POPULATION HAD A MEAN AGE, AT THE TIME OF SURGERY, OF 63.0 ± 15.4 YEARS FOR THE FCL SCREW GROUP AND 63.9 ± 17.3 YEARS FOR THE STANDARD SCREW GROUP; MALE/FEMALE RATIO OF 27/57 WITHIN THE FCL SCREW GROUP AND 26/57 WITHIN THE STANDARD SCREW GROUP. FOLLOW-UP WAS CONDUCTED AT THREE MONTHS, SIX MONTHS, TWELVE MONTHS AFTER FIXATION. A JOURNAL ARTICLE REPORTED SEVEN PATIENTS WITHIN THE FCL SCREW GROUP EXPIRED DUE TO AN UNKNOWN CAUSE WITHIN AN UNKNOWN TIMEFRAME. ATTEMPTS HAVE BEEN MADE, AND NO FURTHER INFORMATION HAS BEEN PROVIDED.Event Description: ATTEMPTS HAVE BEEN MADE, AND NO FURTHER INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: (B)(4). B3: THE EVENT DATE IS THE PUBLICATION DATE OF THE ARTICLE. D4: THE PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER IS NOT APPLICABLE AS THE DEVICE'S PRODUCT NUMBER IS UNKNOWN. D10: UNKNOWN BICORTICAL MOTIONLOC SCREW X4 E1: THE TRIAL WAS AN INTERNATIONAL, MULTICENTER, RANDOMIZED CONTROLLED CLINICAL TRIAL PERFORMED ACROSS 16 PARTICIPATING CENTERS IN CANADA AND THE UNITED STATES. THE METHODS CENTER, LOCATED AT THE UNIVERSITY OF BRITISH COLUMBIA IN VANCOUVER, CANADA, COORDINATED THE TRIAL. G2: CANADA JOURNAL ARTICLE CITATION: FAR CORTICAL LOCKING VERSUS STANDARD CONSTRUCTS FOR LOCKED PLATE FIXATION IN THE TREATMENT OF ACUTE, DISPLACED FRACTURES OF THE DISTAL FEMUR. LEFAIVRE, K A.; SLOBOGEAN, G; O¿HARA, N.; O¿BRIEN, P J.; THE CANADIAN ORTHOPAEDIC TRAUMA SOCIETY (COTS) INVESTIGATORS., ET AL. (2024) HTTP://DX.DOI.ORG/10.2106/JBJS.23.01390. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MDR WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND CORRECTED INFORMATION. D4: THE PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER IS NOT APPLICABLE AS THE DEVICE'S PRODUCT NUMBER IS UNKNOWN. ATTEMPTS HAVE BEEN MADE TO GATHER ALL PRODUCT IDENTIFICATION INFORMATION, AND NO FURTHER INFORMATION HAS BEEN PROVIDED. NO PRODUCT WAS RETURNED, AND NO PICTURES WERE PROVIDED; THEREFORE, A PRODUCT EVALUATION COULD NOT BE PERFORMED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS COULD NOT BE PERFORMED DUE TO MISSING LOT NUMBER. A REVIEW OF THE COMPLAINT HISTORY COULD NOT BE PERFORMED DUE TO MISSING REFERENCE AND LOT NUMBERS. BASED ON THE JOURNAL ARTICLE, THE REPORTED EVENT CAN BE CONFIRMED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED. SINCE THE COMPONENTS ARE UNKNOWN, WHETHER OR TO WHAT EXTENT THIS MAY HAVE CONTRIBUTED TO THE REPORTED EVENT REMAINS UNKNOWN. IF ANY FURTHER INFORMATION IS FOUND THAT WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22456243&pc=HWC | 0009613350-2025-00561 | 2024-10-02 04:00 | Death | ZIMMER GMBH | 2025-07-09 04:00 | HWC | UNK BICORTICAL MOTIONLOC SCREWS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | NA |  | 7 | Event Description: A JOURNAL ARTICLE WAS RETRIEVED FROM THE JOURNAL OF BONE AND JOINT SURGERY (2025) THAT REPORTED A STUDY CONDUCTED IN CANADA. THE PURPOSE OF THE STUDY WAS TO TEST THE HYPOTHESIS IN A COMPARATIVE EFFECTIVENESS CLINICAL TRIAL THAT FAR CORTICAL LOCKING (FCL) CONSTRUCTS LEAD TO IMPROVED FRACTURE-HEALING AS COMPARED WITH THAT ACHIEVED WITH TRADITIONAL LOCKING PLATES. THE TRIAL WAS AN INTERNATIONAL, MULTICENTER, RANDOMIZED CONTROLLED CLINICAL TRIAL PREFORMED ACROSS 16 PARTICIPATING CENTERS IN CANADA AND THE UNITED STATES. THE STUDY REVIEWED 167 PATIENTS (84 FCL SCREWS/83 STANDARD SCREWS) WITHIN A BROAD, HETEROGENOUS POPULATION THAT EXPERIENCED AO/OTA TYPE 33A OR 33C DISTAL FEMORAL FRACTURES. PATIENTS WERE RANDOMLY ASSIGNED TO TREATMENT WITH FCL SCREW FIXATION (NCB DISTAL FEMUR PLATE WITH NCB MOTIONLOC SCREWS; ZIMMER BIOMET) OR STANDARD SCREW FIXATION (NCB DISTAL FEMUR PLATE WITH STANDARD NCB SCREWS; ZIMMER BIOMET). THE FCL GROUP REQUIRED THE USE OF 4 BICORTICAL MOTIONLOC SCREWS FOR FIXATION IN THE SHAFT, WITH NO SPECIFIED WORKING LENGTH. THE STANDARD GROUP ALLOWED LOCKING SCREWS, NON-LOCKING SCREWS, AND HYBRID CONSTRUCTS (I.E., A MIXTURE OF LOCKING AND NON-LOCKING SCREWS) AND ANY WORKING LENGTH. THE INDICATION FOR SURGERY WAS A MECHANISM OF INJURY THAT INCLUDES THE FOLLOWING FOR THE FCL SCREW GROUP/STANDARD SCREW GROUP: DRIVER/PASSENGER IN MOTOR VEHICLE ACCIDENT (9/14), PEDESTRIAN OF MOTOR VEHICLE ACCIDENT (5/5), MOTORCYCLE ACCIDENT (8/4), FALL (58/54), DIRECT TRAUMA (1/2), AND OTHER (3/1). THE STUDY POPULATION HAD A MEAN AGE, AT THE TIME OF SURGERY, OF 63.0 ± 15.4 YEARS FOR THE FCL SCREW GROUP AND 63.9 ± 17.3 YEARS FOR THE STANDARD SCREW GROUP; MALE/FEMALE RATIO OF 27/57 WITHIN THE FCL SCREW GROUP AND 26/57 WITHIN THE STANDARD SCREW GROUP. FOLLOW-UP WAS CONDUCTED AT THREE MONTHS, SIX MONTHS, TWELVE MONTHS AFTER FIXATION. A JOURNAL ARTICLE REPORTED SEVEN PATIENTS WITHIN THE FCL SCREW GROUP EXPIRED DUE TO AN UNKNOWN CAUSE WITHIN AN UNKNOWN TIMEFRAME. ATTEMPTS HAVE BEEN MADE, AND NO FURTHER INFORMATION HAS BEEN PROVIDED.Event Description: ATTEMPTS HAVE BEEN MADE, AND NO FURTHER INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: (B)(4). THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND CORRECTED INFORMATION. D4: THE PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER IS NOT APPLICABLE AS THE DEVICE'S PRODUCT NUMBER IS UNKNOWN. ATTEMPTS HAVE BEEN MADE TO GATHER ALL PRODUCT IDENTIFICATION INFORMATION, AND NO FURTHER INFORMATION HAS BEEN PROVIDED. NO PRODUCT WAS RETURNED, AND NO PICTURES WERE PROVIDED; THEREFORE, A PRODUCT EVALUATION COULD NOT BE PERFORMED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS COULD NOT BE PERFORMED DUE TO MISSING LOT NUMBER. A REVIEW OF THE COMPLAINT HISTORY COULD NOT BE PERFORMED DUE TO MISSING REFERENCE AND LOT NUMBERS. BASED ON THE JOURNAL ARTICLE, THE REPORTED EVENT CAN BE CONFIRMED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED. SINCE THE COMPONENTS ARE UNKNOWN, WHETHER OR TO WHAT EXTENT THIS MAY HAVE CONTRIBUTED TO THE REPORTED EVENT REMAINS UNKNOWN. IF ANY FURTHER INFORMATION IS FOUND THAT WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS. Manufacturer Narrative: (B)(4). B3: THE EVENT DATE IS THE PUBLICATION DATE OF THE ARTICLE. D1: UNKNOWN BICORTICAL MOTIONLOC SCREWS X4. D4: THE PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER IS NOT APPLICABLE AS THE DEVICE'S PRODUCT NUMBER IS UNKNOWN. D10: UNKNOWN NCB PLATE. E1: THE TRIAL WAS AN INTERNATIONAL, MULTICENTER, RANDOMIZED CONTROLLED CLINICAL TRIAL PERFORMED ACROSS 16 PARTICIPATING CENTERS IN CANADA AND THE UNITED STATES. THE METHODS CENTER, LOCATED AT THE (B)(6) CANADA, COORDINATED THE TRIAL. G2: CANADA. JOURNAL ARTICLE CITATION: FAR CORTICAL LOCKING VERSUS STANDARD CONSTRUCTS FOR LOCKED PLATE FIXATION IN THE TREATMENT OF ACUTE, DISPLACED FRACTURES OF THE DISTAL FEMUR. LEFAIVRE, K A.; SLOBOGEAN, G; O¿HARA, N N.; O¿BRIEN, P J.; THE CANADIAN ORTHOPAEDIC TRAUMA SOCIETY (COTS) INVESTIGATORS., ET AL. (2024) HTTP://DX.DOI.ORG/10.2106/JBJS.23.01390. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MDR WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22456264&pc=HRS | 0009613350-2025-00562 | 2024-10-02 04:00 | Death | ZIMMER GMBH | 2025-07-09 04:00 | HRS | UNKNOWN NCB PLATE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | NA |  | 4 | Event Description: A JOURNAL ARTICLE WAS RETRIEVED FROM THE JOURNAL OF BONE AND JOINT SURGERY (2025) THAT REPORTED A STUDY CONDUCTED IN CANADA. THE PURPOSE OF THE STUDY WAS TO TEST THE HYPOTHESIS IN A COMPARATIVE EFFECTIVENESS CLINICAL TRIAL THAT FAR CORTICAL LOCKING (FCL) CONSTRUCTS LEAD TO IMPROVED FRACTURE-HEALING AS COMPARED WITH THAT ACHIEVED WITH TRADITIONAL LOCKING PLATES. THE TRIAL WAS AN INTERNATIONAL, MULTICENTER, RANDOMIZED CONTROLLED CLINICAL TRIAL PREFORMED ACROSS 16 PARTICIPATING CENTERS IN CANADA AND THE UNITED STATES. THE STUDY REVIEWED 167 PATIENTS (84 FCL SCREWS/83 STANDARD SCREWS) WITHIN A BROAD, HETEROGENOUS POPULATION THAT EXPERIENCED AO/OTA TYPE 33A OR 33C DISTAL FEMORAL FRACTURES. PATIENTS WERE RANDOMLY ASSIGNED TO TREATMENT WITH FCL SCREW FIXATION (NCB DISTAL FEMUR PLATE WITH NCB MOTIONLOC SCREWS; ZIMMER BIOMET) OR STANDARD SCREW FIXATION (NCB DISTAL FEMUR PLATE WITH STANDARD NCB SCREWS; ZIMMER BIOMET). THE FCL GROUP REQUIRED THE USE OF 4 BICORTICAL MOTIONLOC SCREWS FOR FIXATION IN THE SHAFT, WITH NO SPECIFIED WORKING LENGTH. THE STANDARD GROUP ALLOWED LOCKING SCREWS, NON-LOCKING SCREWS, AND HYBRID CONSTRUCTS (I.E., A MIXTURE OF LOCKING AND NON-LOCKING SCREWS) AND ANY WORKING LENGTH. THE INDICATION FOR SURGERY WAS A MECHANISM OF INJURY THAT INCLUDES THE FOLLOWING FOR THE FCL SCREW GROUP/STANDARD SCREW GROUP: DRIVER/PASSENGER IN MOTOR VEHICLE ACCIDENT (9/14), PEDESTRIAN OF MOTOR VEHICLE ACCIDENT (5/5), MOTORCYCLE ACCIDENT (8/4), FALL (58/54), DIRECT TRAUMA (1/2), AND OTHER (3/1). THE STUDY POPULATION HAD A MEAN AGE, AT THE TIME OF SURGERY, OF 63.0 ± 15.4 YEARS FOR THE FCL SCREW GROUP AND 63.9 ± 17.3 YEARS FOR THE STANDARD SCREW GROUP; MALE/FEMALE RATIO OF 27/57 WITHIN THE FCL SCREW GROUP AND 26/57 WITHIN THE STANDARD SCREW GROUP . FOLLOW-UP WAS CONDUCTED AT THREE MONTHS, SIX MONTHS, TWELVE MONTHS AFTER FIXATION. A JOURNAL ARTICLE REPORTED FOUR PATIENTS WITHIN THE STANDARD SCREW GROUP EXPIRED DUE TO AN UNKNOWN CAUSE WITHIN AN UNKNOWN TIMEFRAME. ATTEMPTS HAVE BEEN MADE, AND NO FURTHER INFORMATION HAS BEEN PROVIDED.Event Description: ATTEMPTS HAVE BEEN MADE, AND NO FURTHER INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: (B)(4). THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND CORRECTED INFORMATION. THE FOLLOWING SECTIONS WERE UPDATED: B4, B5, G3, G6, H2, H6, H11. D4: THE PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER IS NOT APPLICABLE AS THE DEVICE'S PRODUCT NUMBER IS UNKNOWN. ATTEMPTS HAVE BEEN MADE TO GATHER ALL PRODUCT IDENTIFICATION INFORMATION, AND NO FURTHER INFORMATION HAS BEEN PROVIDED. NO PRODUCT WAS RETURNED, AND NO PICTURES WERE PROVIDED; THEREFORE, A PRODUCT EVALUATION COULD NOT BE PERFORMED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS COULD NOT BE PERFORMED DUE TO MISSING LOT NUMBER. A REVIEW OF THE COMPLAINT HISTORY COULD NOT BE PERFORMED DUE TO MISSING REFERENCE AND LOT NUMBERS. BASED ON THE JOURNAL ARTICLE, THE REPORTED EVENT CAN BE CONFIRMED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED. IF ANY FURTHER INFORMATION IS FOUND THAT WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS. Manufacturer Narrative: (B)(4). B3: EVENT DATE IS THE PUBLICATION DATE OF THE ARTICLE. D4: THE PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER IS NOT APPLICABLE AS THE DEVICE'S PRODUCT NUMBER IS UNKNOWN. D10: UNKNOWN AND UNSPECIFIED NUMBER OF STANDARD NCB SCREWS. E1: THE TRIAL WAS AN INTERNATIONAL, MULTICENTER, RANDOMIZED CONTROLLED CLINICAL TRIAL PERFORMED ACROSS 16 PARTICIPATING CENTERS IN CANADA AND THE UNITED STATES. THE METHODS CENTER, LOCATED AT THE (B)(6) IN VANCOUVER, CANADA, COORDINATED THE TRIAL. G2: CANADA. JOURNAL ARTICLE CITATION: FAR CORTICAL LOCKING VERSUS STANDARD CONSTRUCTS FOR LOCKED PLATE FIXATION IN THE TREATMENT OF ACUTE, DISPLACED FRACTURES OF THE DISTAL FEMUR. LEFAIVRE, K A.; SLOBOGEAN, G; O¿HARA, N.; O¿BRIEN, P J.; THE CANADIAN ORTHOPAEDIC TRAUMA SOCIETY (COTS) INVESTIGATORS., ET AL. (2024) HTTP://DX.DOI.ORG/10.2106/JBJS.23.01390. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MDR WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22456275&pc=HWC | 0009613350-2025-00563 | 2024-10-02 04:00 | Death | ZIMMER GMBH | 2025-07-09 04:00 | HWC | UNKNOWN STD NCB SCREWS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | NA |  | 4 | Event Description: A JOURNAL ARTICLE WAS RETRIEVED FROM THE JOURNAL OF BONE AND JOINT SURGERY (2025) THAT REPORTED A STUDY CONDUCTED IN CANADA. THE PURPOSE OF THE STUDY WAS TO TEST THE HYPOTHESIS IN A COMPARATIVE EFFECTIVENESS CLINICAL TRIAL THAT FAR CORTICAL LOCKING (FCL) CONSTRUCTS LEAD TO IMPROVED FRACTURE-HEALING AS COMPARED WITH THAT ACHIEVED WITH TRADITIONAL LOCKING PLATES. THE TRIAL WAS AN INTERNATIONAL, MULTICENTER, RANDOMIZED CONTROLLED CLINICAL TRIAL PREFORMED ACROSS 16 PARTICIPATING CENTERS IN CANADA AND THE UNITED STATES. THE STUDY REVIEWED 167 PATIENTS (84 FCL SCREWS/83 STANDARD SCREWS) WITHIN A BROAD, HETEROGENOUS POPULATION THAT EXPERIENCED AO/OTA TYPE 33A OR 33C DISTAL FEMORAL FRACTURES. PATIENTS WERE RANDOMLY ASSIGNED TO TREATMENT WITH FCL SCREW FIXATION (NCB DISTAL FEMUR PLATE WITH NCB MOTIONLOC SCREWS; ZIMMER BIOMET) OR STANDARD SCREW FIXATION (NCB DISTAL FEMUR PLATE WITH STANDARD NCB SCREWS; ZIMMER BIOMET). THE FCL GROUP REQUIRED THE USE OF 4 BICORTICAL MOTIONLOC SCREWS FOR FIXATION IN THE SHAFT, WITH NO SPECIFIED WORKING LENGTH. THE STANDARD GROUP ALLOWED LOCKING SCREWS, NON-LOCKING SCREWS, AND HYBRID CONSTRUCTS (I.E., A MIXTURE OF LOCKING AND NON-LOCKING SCREWS) AND ANY WORKING LENGTH. THE INDICATION FOR SURGERY WAS A MECHANISM OF INJURY THAT INCLUDES THE FOLLOWING FOR THE FCL SCREW GROUP/STANDARD SCREW GROUP: DRIVER/PASSENGER IN MOTOR VEHICLE ACCIDENT (9/14), PEDESTRIAN OF MOTOR VEHICLE ACCIDENT (5/5), MOTORCYCLE ACCIDENT (8/4), FALL (58/54), DIRECT TRAUMA (1/2), AND OTHER (3/1). THE STUDY POPULATION HAD A MEAN AGE, AT THE TIME OF SURGERY, OF 63.0 ± 15.4 YEARS FOR THE FCL SCREW GROUP AND 63.9 ± 17.3 YEARS FOR THE STANDARD SCREW GROUP; MALE/FEMALE RATIO OF 27/57 WITHIN THE FCL SCREW GROUP AND 26/57 WITHIN THE STANDARD SCREW GROUP. FOLLOW-UP WAS CONDUCTED AT THREE MONTHS, SIX MONTHS, TWELVE MONTHS AFTER FIXATION. A JOURNAL ARTICLE REPORTED FOUR PATIENTS WITHIN THE STANDARD SCREW GROUP EXPIRED DUE TO AN UNKNOWN CAUSE WITHIN AN UNKNOWN TIMEFRAME. ATTEMPTS HAVE BEEN MADE, AND NO FURTHER INFORMATION HAS BEEN PROVIDED.Event Description: ATTEMPTS HAVE BEEN MADE, AND NO FURTHER INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: (B)(4). B3: THE EVENT DATE IS THE PUBLICATION DATE OF THE ARTICLE. D1: UNKNOWN AND UNSPECIFIED NUMBER OF STANDARD NCB SCREWS D4: THE PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER IS NOT APPLICABLE AS THE DEVICE'S PRODUCT NUMBER IS UNKNOWN. D10: UNKNOWN NCB PLATE E1: THE TRIAL WAS AN INTERNATIONAL, MULTICENTER, RANDOMIZED CONTROLLED CLINICAL TRIAL PERFORMED ACROSS 16 PARTICIPATING CENTERS IN CANADA AND THE UNITED STATES. THE METHODS CENTER, LOCATED AT THE UNIVERSITY OF BRITISH COLUMBIA IN VANCOUVER, CANADA, COORDINATED THE TRIAL. G2: CANADA JOURNAL ARTICLE CITATION: FAR CORTICAL LOCKING VERSUS STANDARD CONSTRUCTS FOR LOCKED PLATE FIXATION IN THE TREATMENT OF ACUTE, DISPLACED FRACTURES OF THE DISTAL FEMUR. LEFAIVRE, K A.; SLOBOGEAN, G; O¿HARA, N N.; O¿BRIEN, P J.; THE CANADIAN ORTHOPAEDIC TRAUMA SOCIETY (COTS) INVESTIGATORS. , ET AL. (2024) HTTP://DX.DOI.ORG/10.2106/JBJS.23.01390 THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MDR WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND CORRECTED INFORMATION. D4: THE PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER IS NOT APPLICABLE AS THE DEVICE'S PRODUCT NUMBER IS UNKNOWN. ATTEMPTS HAVE BEEN MADE TO GATHER ALL PRODUCT IDENTIFICATION INFORMATION, AND NO FURTHER INFORMATION HAS BEEN PROVIDED. NO PRODUCT WAS RETURNED, AND NO PICTURES WERE PROVIDED; THEREFORE, A PRODUCT EVALUATION COULD NOT BE PERFORMED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS COULD NOT BE PERFORMED DUE TO MISSING LOT NUMBER. A REVIEW OF THE COMPLAINT HISTORY COULD NOT BE PERFORMED DUE TO MISSING REFERENCE AND LOT NUMBERS. BASED ON THE JOURNAL ARTICLE, THE REPORTED EVENT CAN BE CONFIRMED. A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED. IF ANY FURTHER INFORMATION IS FOUND THAT WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22456690&pc=DSQ | 2916596-2025-04278 | 2025-05-08 04:00 | Death | THORATEC CORPORATION | 2025-07-09 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Fall; Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT SUFFERED A SUBARACHNOID BLEEDING (SAB) AFTER FALLING DOWN THE STAIRS. A COMPUTED TOMOGRAPHY (CT) SCAN WAS DONE. THIS SCAN CONFIRMED THE BLEEDING. THE BLEEDING WAS TREATED CONSERVATIVELY. THE FALL AS NOT DUE TO LVAD THERAPY. THE PATIENT PASSED AWAY DUE TO THE SAB. THE DEATH WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS BLEEDING AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO RANGE, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22456710&pc=LWS | 3003120897-2025-00001 | 2025-06-10 04:00 | Death | MEDTRONIC SOFAMOR DANEK USA, INC. | 2025-07-09 04:00 | LWS | PROTECTA VR | Insufficient Information | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS RETURNED TO THE MANUFACTURER AFTER THE PATIENT¿S DEATH WITH PAPERWORK INDICATING THE DEVICE HAD NOT BEEN FUNCTIONING WITHIN NORMAL TOLERANCES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT HAS NOT BEEN RECEIVED. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. ANALYSIS OF THE DEVICE REVEALED NORMAL BATTERY DEPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22457134&pc=LPB | 2029046-2025-02230 | 2025-05-15 04:00 | Death | BIOSENSE WEBSTER INC | 2025-07-09 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Fistula | P030031 |  | 1 | Event Description: IT WAS REPORTED 3 TO 4 WEEKS AFTER PULMONARY VEIN ISOLATION (PVI) ABLATION PROCEDURE WITH A THERMOCOOL SMARTTOUCH SF CATHETER THE PATIENT EXPERIENCED STROKE AND EXPIRED. THE PATIENT EXPERIENCED DEATH APPROXIMATELY 3-4 WEEKS POST PVI ABLATION AFTER RETURNING TO THE HOSPITAL WITH A STROKE. A COMPUTED TOMOGRAPHY (CT) SCAN WAS COMPLETED TO RULE OUT A FISTULA, BUT NOTHING WAS FOUND. PATIENT PASSED AWAY OVER THE WEEKEND, APPROXIMATELY JUNE 14. LIMITED INFORMATION IS AVAILABLE, BUT THE PATIENT WAS FOUND TO HAVE REMAINED IN ATRIAL FIBRILLATION THE FOLLOWING DAY AND SO WAS TREATED FOR REDO PVI ABLATION WITH RADIOFREQUENCY WITHIN 48 HOURS OF THE FIRST PFA ABLATION. THE FAMILY REFUSED AUTOPSY. Manufacturer Narrative: ON 9-JUL-2025, IT WAS NOTICED THE FOLLOWING INFORMATION WAS INADVERTENTLY OMITTED FROM THE 3500A INITIAL MEDWATCH REPORT: ADDITIONAL INFORMATION WAS RECEIVED INDICATING THE PHYSICIAN¿S OPINION FOR CAUSE OF DEATH WAS SUSPECTED A FISTULA BUT OVERALL, IT WAS NOT CONFIRMED. THE INITIAL CASE DATE WAS (B)(6), FOR PULMONARY VEIN ISOLATION (PVI) WITH VARIPULSE USING 30 ML IRRIGATION DURING ABLATION. UNDERWENT PVI WITH VARIPULSE AND POST ABLATION, AT THE HOTEL THAT NIGHT, WENT INTO PERSISTENT AF. RETURNED TO THE HOSPITAL NEXT DAY, REPEATED UNSUCCESSFUL CARDIOVERSIONS. THREE DAYS AFTER THE INITIAL ABLATION PROCEDURE, THE PATIENT WAS BROUGHT BACK FOR REDO ABLATION WITH RADIOFREQUENCY (RF). RIGHT POSTERIOR CARINA, LEFT INFERIOR PULMONARY VEIN (LIPV) FOUND TO BE RECONNECTED. THE PHYSICIANS RE-ISOLATED THE VEINS AND POSTERIOR BOX WITH RADIOFREQUENCY ABLATION. THE PATIENT WAS CARDIOVERTED AT END OF PROCEDURE, AND THE PATIENT LEFT IN SINUS RHYTHM. TWO WEEKS LATER THE PHYSICIAN FOLLOWS UP, AND THE PATIENT WAS DOING WELL. THREE WEEKS LATER THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT WITH STROKE, AND ONE SYMPTOM WAS APPARENTLY ESOPHAGEAL BLEEDING. THE PATIENT DIED IN THE HOSPITAL, AND THE FAMILY DENIED AUTOPSY. THE NEUROLOGICAL IMPAIRMENT EVENT OCCURRED WAS CEREBROVASCULAR ACCIDENT (CVA) /STROKE AND TRANSIENT ISCHEMIC ATTACK (TIA). A CT SCAN IMAGING WAS PERFORMED TO CONFIRM THE STROKE. NO EVIDENCE OF CHAR DURING THE PROCEDURE. NO EVIDENCE OF BLOOD THROMBUS/CLOT DURING THE PROCEDURE. THE CORRECT CATHETER SETTINGS WERE SELECTED ON THE GENERATOR. NO ISSUES WITH FLOW RATE CHANGE AT THE START OF ABLATION. THE EVENT DESCRIPTION INDICATED THAT THE INITIAL ABLATION PROCEDURE WAS PERFORMED ON (B)(6), BUT NOT THE ADVERSE EVENT. THE DEVICES USED WERE THERMOCOOL SMARTTOUCH SF CATHETER AND NGEN GENERATOR. AS SUCH, THE EVENT HAS BEEN DETERMINED TO BE MDR REPORTABLE AGAINST THE THERMOCOOL SMARTTOUCH SF CATHETER AND THE NGEN RF GENERATOR. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REF. # (B)(4). Manufacturer Narrative: DEVICE INVESTIGATION DETAILS: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. HOWEVER, IF THE PRODUCT OR PRODUCT ID NUMBERS ARE RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. NOTE: FOR FIELD D4. UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REF # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22457496&pc=OZD | 1220648-2025-30162 | 2025-06-14 04:00 | Death | ABIOMED, INC. | 2025-07-09 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE PATIENT WAS INITIALLY PRESENTED TO AN OUTSIDE EMERGENCY DEPARTMENT WITH COMPLAINTS OF CHEST PAIN. A COMPUTED TOMOGRAPHY (CT) SCAN REVEALED A TYPE A AORTIC DISSECTION WITH A DILATED AORTIC ROOT MEASURING 6.5 CENTIMETERS. DUE TO LEFT VENTRICULAR DISTENSION WHILE ON VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO), THE DECISION WAS MADE TO PLACE A PERCUTANEOUS IMPELLA CP FOR LEFT VENTRICULAR UNLOADING. DURING SUPPORT, THE PATIENT DEVELOPED BLEEDING AROUND THE REPOSITIONING SHEATH. A FEM-STOP WAS APPLIED FOR HEMOSTASIS, AND ONE UNIT OF PACKED RED BLOOD CELLS WAS TRANSFUSED. THE IMPELLA CP ALSO EXHIBITED INTERMITTENT SUCTION EVENTS WHICH WERE MANAGED WITH BOLUSES OF 25% ALBUMIN AND BY ADJUSTING THE P-LEVEL SETTINGS TO RELIEVE SUCTION. DUE TO PERSISTENT HEMODYNAMIC INSTABILITY, THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5. THE PATIENT WAS ON IMPELLA 5.5 SUPPORT FOR 12 DAYS AND CARE WAS ULTIMATELY WITHDRAWN, AND THE PATIENT EXPIRED. Manufacturer Narrative: MEDICAL SAFETY OFFICER REVIEW: THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP DEVICE USED AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW AND ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ACCESS SITE BLEEDING: THE CAUSE OF REPORTED BLEEDING ISSUE MOST LIKELY PATIENT CONDITION RELATED SINCE BLEEDING OCCURRED IN MULTIPLE SITES. LOW PUMP FLOW: THE CAUSE OF LOW PUMP FLOW ISSUE MOST LIKELY WAS PATIENT CONDITION RELATED BECAUSE IMPELLA WAS USED AS A VENT FOR ECMO IN ADDITION TO THE BOLUS OF 25% ALBUMIN, AND ADJUSTING P-LEVEL TO BREAK SUCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22457633&pc=OZD | 1220648-2025-30163 | 2025-06-20 04:00 | Death | ABIOMED, INC. | 2025-07-09 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Defective Device | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THEY ENCOUNTERED ISSUES WITH THE AUTOMATED IMPELLA CONTROLLER (AIC). AS THE TEAM WAS PLACING AN IMPELLA CP AND AS THE NURSE WENT TO GO INSERT THE PURGE CASSETTE INTO THE AIC, IT WOULD NOT CLICK INTO PLACE ALL THE WAY. THE TEAM HAD TO SWITCH AIC'S AND IT THEN INSERTED JUST FINE WITH SUCCESSFUL PUMP PLACEMENT. THE PATIENT WAS TRANSFERRED AND THE TEAM DISCUSSED TREATMENT OPTIONS. ULTIMATELY THERE WAS NOTHING ELSE THAT COULD BE DONE FOR THE PATIENT AND THE FAMILY DECIDED TO WITHDRAW CARE. THE PATIENT EXPIRED AND THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF AUTOMATED IMPELLA CONTROLLER (AIC) - PURGE DISC/FLAG ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS RETURNED FOR EVALUATION. DATA REVIEW: CONSOLE LOGS WERE PARTIALLY CONSISTENT WITH WHAT WAS REPORTED IN THE COMPLAINT. ACCORDING TO THE IMC LOGS: THE USER INSERTED THE PURGE CASSETTE DURING STEP 2 OF THE PURGE SETUP WIZARD. AT THIS TIME THE PURGE DISC WAS NOT INSERTED YET AND THE PURGE DISC HELP SCREEN WAS DISPLAYED. THE USER THEN REMOVED THE PURGE CASSETTE AND INSERTED THE PURGE DISC SHORTLY AFTER. NUMEROUS CONSECUTIVE ERRORS WERE OBSERVED AFTER THE PURGE DISC WAS INSERTED WHICH IS INDICATIVE OF THE USER LIKELY HAVING ISSUES RE-INSERTING THE CASSETTE INTO THE PURGE ASSEMBLY DEVICE ANALYSIS: CONSOLE WAS TESTED AFTER ARRIVING IN FIELD SERVICE. THE REPORTED PURGE CASSETTE INSERTION ISSUE WAS UNABLE TO BE REPRODUCED. CONCLUSION: THE ROOT CAUSE OF THE REPORTED PURGE CASSETTE INSERTION ISSUE COULD NOT BE DETERMINED DUE TO THE INABILITY TO REPRODUCE. THE CONSOLE OPERATED AS INTENDED WHILE TESTING ON AN ENGINEERING LAB BENCH WITH A KNOWN GOOD PURGE CASSETTE KIT. H5 WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30163. H8 WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30163.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22457756&pc=OZD | 1220648-2025-30164 | 2025-06-20 04:00 | Death | ABIOMED, INC. | 2025-07-09 04:00 | OZD | IMPELLA CP | Device Sensing Problem | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PLACEMENT SIGNAL ON AN IMPELLA CP PUMP BECAME UNRELIABLE DURING PATIENT SUPPORT. THE PLACEMENT MONITORING WAS DISABLED IN RESPONSE TO THE NOTED ISSUE AND SUPPORT WAS CONTINUED UNINTERRUPTED. TWO DAYS LATER, THE PATIENT WAS MADE PALLIATIVE CARE BY THEIR FAMILY AND PASSED AWAY ON SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF OPTICAL SIGNAL ISSUE HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED FOR EVALUATION. DATA LOGS WERE REVIEWED AND LT LOG OF CASE SHOWED SUDDEN INCREASE IN PLACEMENT SIGNAL OUT OF RANGE, WITH OPTICAL 5S PARAMETERS CONSISTENT WITH OSCILLATING CONTRAST. CLINICAL DETAILS NOTED THAT A SUDDEN PLACEMENT SIGNAL NOT RELIABLE (PSNR) ALARM WAS TRIGGERED AND NO ABNORMALITIES WITH PATIENT WERE NOTED AS CLINICAL TEAM WAS IN THE ROOM AND MONITORING THE PATIENT WHEN IT OCCURRED. UPON REVIEW OF RETURNED DATA LOGS, IT IS APPARENT THAT THE CONSOLE IS THE POTENTIAL CAUSE OF THE ISSUE. THE ROOT CAUSE OF THE OPTICAL SIGNAL ISSUE COULD NOT BE DEFINITIVELY DETERMINED AS NO PRODUCT WAS RETURNED FOR EVALUATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22457913&pc=OZD | 1220648-2025-30165 | 2025-06-20 04:00 | Death | ABIOMED, INC. | 2025-07-09 04:00 | OZD | IMPELLA CP | Failure to Advance; Patient-Device Incompatibility | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT IMPLANTATION OF AN IMPELLA CP WAS UNSUCCESSFUL DUE TO SEVERE AORTIC TORTUOSITY IN THE PATIENT'S ANATOMY. THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION SINCE THE PLACEMENT OF PUMP WAS ABORTED DUE TO SEVERE AORTIC TORTUOSITY. D6B EXPLANT DATE WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT. Manufacturer Narrative: A2, A3, A4, AND A5 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI STATES THE FOLLOWING. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: CAUTIONS. ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22458000&pc=LWS | 2649622-2025-18479 | 2025-05-04 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-09 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Device Sensing Problem; Impedance Problem; High Capture Threshold; Unstable Capture Threshold | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT AN INTERNAL REVIEW OF DATA TRANSMITTED FROM THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) INDICATED THAT THE RIGHT VENTRICULAR (RV) LEAD DISPLAYED VARIATION IN PACING IMPEDANCE, DEFIBRILLATION IMPEDANCE, THRESHOLDS, AND SENSED INTRINSIC ACTIVITY. THE LEAD ALSO HAD HIGH THRESHOLDS. IT WAS NOTED THAT THE PATIENT SUBSEQUENTLY EXPIRED A FEW DAYS LATER, THE CAUSE OF DEATH WAS REQUESTED BUT WAS NOT PROVIDED. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED THE IMPEDANCE TREND OF THE RIGHT VENTRICULAR DEFIBRILLATION COIL WAS VARIABLE. ANALYSIS OF THE DEVICE MEMORY INDICATED THE IMPEDANCE TREND OF THE RIGHT VENTRICULAR PACING LEAD WAS VARIABLE. ANALYSIS OF THE DEVICE MEMORY INDICATED PACING CAPTURE THRESHOLD IN THE RIGHT VENTRICLE WAS ELEVATED. ANALYSIS OF THE DEVICE MEMORY INDICATED PACING CAPTURE THRESHOLD IN THE RIGHT VENTRICLE WAS UNSTABLE. ANALYSIS OF THE DEVICE MEMORY INDICATED DIMINISHED RIGHT VENTRICULAR SENSING. ANALYSIS OF THE DEVICE MEMORY INDICATED THAT THE VENTRICULAR TACHYARRHYTHMIA THERAPY ENERGY WAS INSUFFICIENT. ANALYSIS OF THE DEVICE MEMORY OBSERVED NOISE ON THE ELECTROGRAM WAVEFORMS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22458281&pc=NVZ | 3008973940-2025-04959 | 2025-04-05 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-07-09 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED APPROXIMATELY TWO MONTHS POST IMPLANT THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED LOW IMPEDANCE. IT WAS FURTHER REPORTED THAT THE PATIENT WAS DECEASED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT'S DEATH WAS NOT LINKED TO THE PERFORMANCE ISSUE OF THE IPG SYSTEM AND THE PATIENT WAS MADE COMFORT MEASURES. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22458554&pc=DTB | 2649622-2025-18492 | 2025-04-05 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-09 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED APPROXIMATELY TWO MONTHS POST IMPLANT THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED LOW IMPEDANCE. IT WAS FURTHER REPORTED THAT THE PATIENT WAS DECEASED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT'S DEATH WAS NOT LINKED TO THE PERFORMANCE ISSUE OF THE IPG SYSTEM AND THE PATIENT WAS MADE COMFORT MEASURES. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22459329&pc=OZD | 1220648-2025-30173 | 2025-06-06 04:00 | Death | ABIOMED, INC. | 2025-07-09 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Stroke/CVA; Hematoma; Hemolysis; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED AN IMPELLA 5.5 AFTER CARE ESCALATION FROM AN IMPELLA CP. THE PATIENT HAD BEEN ADMITTED IN WITH CARDIOMYOPATHY AND CARDIOGENIC SHOCK. ON THE DAY AFTER IMPLANT, THE PURGE PRESSURE ROSE PROMPTING THE TEAM TO GIVE TISSUE PLASMINOGEN INTO THE PURGE. ON DAY 10, A LARGE HEMATOMA IN THE RIGHT AXILLARY SOFT TISSUES WAS NOTED AND THERE WAS A POSSIBLE PSEUDOANEURYSM. THE TEAM TREATED BY INFUSING BLOOD PRODUCTS. ADDITIONALLY, THERE WAS A SUBARACHNOID BLEED. THIS PROMPTED THE TEAM TO STOP HEPARIN AND SWITCH TO SODIUM BICARBONATE. THE PUMP REMAINED ON DESPITE THIS UNTIL THE FAMILY MADE THE CHOICE TO WITHDRAW CARE. MEDICAL SAFETY OFFICER REVIEW: MULTIPLE PUMPS WHICH INCLUDED BIPELLA SUPPORT WITH IMPELLA RP AND IMPELLA 5.5 SUPPORTED THE PATIENT. IMPELLA CP AND IMPELLA RP RELATED EVENTS ARE SERIOUS INJURY TYPE OF REPORTABLE EVENTS. THE DEMISE OUTCOME IS ASSOCIATED WITH THE IMPELLA 5.5 AS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDED THE IMPELLA 5.5 DEVICE USED AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME; LAST DEVICE USED AND THE PATIENT HAD HEMOLYSIS, EXTENSIVE REPAIR AND BLOOD TRANSFUSION WHICH IS MORE LIKELY TO BE RELATED TO THE DEMISE OUTCOME. Manufacturer Narrative: MEDICAL SAFETY OFFICER REVIEW: MULTIPLE PUMPS WHICH INCLUDED BIPELLA SUPPORT WITH IMPELLA RP AND IMPELLA 5.5 SUPPORTED THE PATIENT. IMPELLA CP AND IMPELLA RP RELATED EVENTS ARE SERIOUS INJURY TYPE OF REPORTABLE EVENTS. THE DEMISE OUTCOME IS ASSOCIATED WITH THE IMPELLA 5.5 AS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDED THE IMPELLA 5.5 DEVICE USED AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME; LAST DEVICE USED AND THE PATIENT HAD HEMOLYSIS, EXTENSIVE REPAIR AND BLOOD TRANSFUSION WHICH IS MORE LIKELY TO BE RELATED TO THE DEMISE OUTCOME. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER. ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION. ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ IMPELLA 5.5 WITH SMARTASSIST. SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER. ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING, STROKE/CVA, HIGH PURGE PRESSURE, HEMOLYSIS, AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ACCESS SITE BLEEDING: NO PRODUCT WAS RETURNED AND DATA LOGS ARE NOT APPLICABLE. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING WAS NOT DETERMINED SINCE INSUFFICIENT CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED. STROKE/CVA: THE ROOT CAUSE OF THE STROKE/CVA WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. HIGH PURGE PRESSURE (HPP): THE DATA LOG REVIEW DID NOT SHOW ANY HIGH PURGE PRESSURE ALARMS UNTIL (B)(6) 2025. ON (B)(6) 2025, THERE WAS A BRIEF SPIKE IN PURGE PRESSURE THAT RESOLVED WITHIN 4 MINUTES. A FEW HOURS AFTER THIS SPIKE, THERE WAS A SHIFT UPWARD IN PURGE PRESSURE FROM 400 MMHG TO 600 MMHG. THERE WERE SOME LONG BAG/CASSETTE CHANGES THROUGHOUT THE CASE, BUT NONE DIRECTLY PRIOR TO THE HPP EVENT. PURGE PRESSURE WAS STABLE FROM (B)(6) 2025 UNTIL THE HPP EVENT ON (B)(6) 2025. NO PROBLEM WAS FOUND WITH THE 5.5 DURING THE TIME OF THE COMPLAINT BASED ON DATA LOG REVIEW. HEMOLYSIS: DATA LOG REVIEW SHOWED SUCTION ALARMS THROUGHOUT THE CASE, BUT MOST PRESENT BETWEEN (B)(6) 2025. THERE ARE NO MAJOR POSITIONING ISSUES BASED ON THE DATA LOGS. THERE IS A MOTOR CURRENT SPIKE ON (B)(6) 2025. THE PLACEMENT SIGNAL VARIES THROUGHOUT THE CASE WITH NO CLEAR TREND. NO PRODUCT WAS RETURNED. THE ROOT CAUSE OF THE HEMOLYSIS WAS NOT DETERMINED SINCE NO PRODUCT WAS RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. VF/VT/AF/AT: THE ROOT CAUSE OF THE VF/VT/AF/AT WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. H6 HEALTH EFFECT - CLINICAL CODE 1770 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30173.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22459677&pc=LJT | 3012307300-2025-08127 | 2023-03-15 04:00 | Death | ICU MEDICAL INC | 2025-07-09 04:00 | LJT | PORT-A-CATH IMPLANTABLE ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K072657 |  | 1 | Event Description: ICU MEDICAL LEARNED OF A CIVIL LAWSUIT FILED IMPLICATING A PORT-A-CATH II, 21-4455-24, LOT 3829680. THE LAWSUIT ALLEGES THE FOLLOWING. THE PATIENT, (B)(6), UNDERWENT PLACEMENT OF A PORT IN THE LEFT CHEST, ON (B)(6) 2021 FOR VEIN ACCESS. THE PATIENT PRESENTED TO (B)(6) WITH SHORTNESS OF BREATH. DOCTORS DETERMINED SHE HAD AN INFECTION WITH ACCUMULATION OF FLUID IN THE LUNGS. THE PATIENT'S PORT WAS EXPLANTED IMMEDIATELY AFTER THIS DETERMINATION ON OR ABOUT (B)(6) 2023. THE PATIENT PASSED AWAY ON (B)(6) 2023. Manufacturer Narrative: NO DEVICE WAS RETURNED FOR INVESTIGATION. DUE TO THIS, NO DEVICE ISSUE WAS DETERMINED. IF THE DEVICE IS RETURNED FOR INVESTIGATION, THE DEVICE WILL BE EXAMINED WITH THE RESULTS SENT IN A SUPPLEMENTAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22459859&pc=OZD | 1220648-2025-30175 | 2025-06-20 04:00 | Death | ABIOMED, INC. | 2025-07-09 04:00 | OZD | IMPELLA CP | Malposition of Device | Tachycardia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND EXPERIENCED TWO EPISODES OF VENTRICULAR TACHYCARDIA AND WAS SUCCESSFULLY CARDIOVERTED. AFTERWARDS, THE PUMP HAD TO BE SLIGHTLY REPOSITIONED DUE TO ASPIRATION ALARMS. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22459917&pc=DSQ | 2916596-2025-04236 | 2025-05-29 04:00 | Death | THORATEC CORPORATION | 2025-07-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO A HEMORRHAGIC STROKE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. PUMP WAS NOT EXPLANTED. ADDITIONAL INFORMATION IS UNKNOWN.Event Description: ADDITIONAL INFORMATION WAS PROVIDED ON 29JUL2025 THAT THE PATIENT PRESENTED TO THE EMERGENCY ROOM (ER) VIA AMBULANCE ON (B)(6) 2025 DUE TO ALTERED MENTAL STATUS. COMPUTED TOMOGRAPHY (CT) SHOWED A 5.7 X 3.1 X 3 CM INTRAPARENCHYMAL HEMORRHAGE CENTERED IN THE RIGHT OCCIPITAL LOBE WHICH CONTINUED TO WORSEN OVER THE SUBSEQUENT DAYS, RESULTING IN AN 18 MM MIDLINE SHIFT AND SUBFALCINE HERNIATION. THE FAMILY MADE THE DECISION TO TRANSITION THE PATIENT TO COMFORT CARE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE HAD OPERATED AS EXPECTED. THE PATIENT HAD A HISTORY OF INTRACRANIAL HEMORRHAGE WHILE ANTICOAGULATED WITH WARFARIN. WARFARIN WAS STOPPED. THE PATIENT REMAINED OFF ANTICOAGULATION FOR SEVERAL YEARS UNTIL EARLIER IN 2025 WHEN THE PATIENT DEVELOPED ATRIAL FIBRILLATION, AND APIXABAN WAS STARTED IN LIEU OF WARFARIN. PER THE SITE, THIS CHANGE IN ANTICOAGULATION STATUS MAY HAVE CONTRIBUTED TO THE HEMORRHAGIC STROKE. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: SECTION B: UPDATED. H6 - ADVERSE EVENT PROBLEM CODES UPDATED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22460228&pc=DSQ | 2916596-2025-04250 | 2025-05-30 04:00 | Death | THORATEC CORPORATION | 2025-07-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Stroke/CVA; Cardiogenic Shock; Pseudoaneurysm; Hemorrhagic Stroke; Embolism/Embolus | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS BACTEREMIC AND PASSED AWAY.Event Description: ADDITIONAL INFORMATION RECEIVED STATED THAT THE PATIENT DEVELOPED REFRACTORY BACTEREMIA WITH METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) RESISTANT TO CONVENTIONAL THERAPY AND DESPITE THE USE OF MULTIPLE ANTIBACTERIAL AGENTS CONCURRENTLY, BACTEREMIA WAS UNABLE TO BE CLEARED. SUSPECT DEVICE WAS SEEDED. CULTURES IDENTIFIED INCLUDED MRSA. THE PATIENT ALSO SUFFERED MULTIPLE CEREBROVASCULAR ACCIDENTS (CVAS) WITH SUSPECT SEPTIC EMBOLI. AFTER DISCUSSION WITH FAMILY WITH REGARDS FOR CONCERN FOR INFECTED PUMP AND REFRACTORY BACTEREMIA, THE PATIENT AND FAMILY ELECTED TO TRANSITION TO COMFORT MEASURES AND TERMINATE SUPPORT. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE WAS LIKELY INFECTED BUT NOT A FACTOR IN THE DEATH OF THE PATIENT. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22460920&pc=DSQ | 2916596-2025-04261 | 2025-06-15 04:00 | Death | THORATEC CORPORATION | 2025-07-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. ADDITIONAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS UNKNOWN WHETHER THE PATIENT'S OUTCOME WAS CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22460963&pc=DRA | 2124215-2025-45845 | 2025-06-12 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-09 04:00 | DRA | FARADRIVE STEERABLE SHEATH | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Perforation; Pericardial Effusion; Unspecified Mental, Emotional or Behavioural Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED PULSES ELECTRICAL ACTIVITY (PEA) DURING A CARDIAC ARREST. POST PULSED FIELD ABLATION PROCEDURE, AND UPON ARRIVAL TO THE POST CARE UNIT, THE PATIENT HAD A CARDIAC ARREST. A CODE WAS CALLED, AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. IT WAS SUSPECTED THAT THE CARDIAC ARREST MIGHT HAVE BEEN CAUSED BY THE REVERSAL OF PROTEINS. ADDITIONALLY, THE PATIENT WAS THEN TRANSFERRED INTO THE OPERATING ROOM AND AN EFFUSION WAS OBSERVED. THE FARADRIVE SHEATH HAD CAUSED A HOLE IN A LINEAR FASHION IN THE RIGHT ATRIUM FROM THE LEVEL OF THE RIGHT INFERIOR PULMONARY VEIN (RIPV) TO THE INFERIOR VENA CAVA (IVC). THE PATIENT RIGHT ATRIUM TISSUE WAS THIN, BUT IT WAS CLEARLY CAUSED WHEN REMOVING THE SHEATH POST PROCEDURE. THE DEVICE IS NOT EXPECTED TO RETURN AS IT WAS DISPOSED. IT WAS FURTHER REPORTED THAT THE SHEATH WAS DISPOSED THEREFORE THE LOT NUMBER IS NOT AVAILABLE. THE TREATMENT ADMINISTERED TO RESOLVE THE EFFUSION INITIALLY INVOLVED PERFORMING TAP IN THE POST-CARE UNIT. HOWEVER, DUE TO THE PATIENTS HABITUS, THIS WAS UNSUCCESSFUL, AND THE PROCEDURE WAS MOVED TO THE OPERATIVE ROOM FOR A PERICARDIAL WINDOW. THE EFFUSION WAS IDENTIFIED ONLY AFTER OPENING THE CHEST IN THE OPERATIVE ROOM. THE PATIENT WAS STABILIZED AND IS CURRENTLY ALIVE. THEY HAVE NOT BEEN DISCHARGED AND REMAIN AN INPATIENT AT THE HOSPITAL. THE PFA PROCEDURE WAS PERFORMED TO TREAT PAROXYSMAL ATRIAL FIBRILLATION. NO PHYSICAL DAMAGE WAS OBSERVED WITH ANY BOSTON SCIENTIFIC DEVICES PRIOR TO THE PROCEDURE, AND NO EFFUSION WAS NOTED IMMEDIATELY POST-PROCEDURE.Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED PULSES ELECTRICAL ACTIVITY (PEA) DURING A CARDIAC ARREST. POST PULSED FIELD ABLATION PROCEDURE, AND UPON ARRIVAL TO THE POST CARE UNIT, THE PATIENT HAD A CARDIAC ARREST. A CODE WAS CALLED, AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. IT WAS SUSPECTED THAT THE CARDIAC ARREST MIGHT HAVE BEEN CAUSED BY THE REVERSAL OF PROTEINS. ADDITIONALLY, THE PATIENT WAS THEN TRANSFERRED INTO THE OPERATING ROOM AND AN EFFUSION WAS OBSERVED. THE FARADRIVE SHEATH HAD CAUSED A HOLE IN A LINEAR FASHION IN THE RIGHT ATRIUM FROM THE LEVEL OF THE RIGHT INFERIOR PULMONARY VEIN (RIPV) TO THE INFERIOR VENA CAVA (IVC). THE PATIENT RIGHT ATRIUM TISSUE WAS THIN, BUT IT WAS CLEARLY CAUSED WHEN REMOVING THE SHEATH POST PROCEDURE. THE DEVICE IS NOT EXPECTED TO RETURN AS IT WAS DISPOSED. IT WAS FURTHER REPORTED THAT THE SHEATH WAS DISPOSED THEREFORE THE LOT NUMBER IS NOT AVAILABLE. THE TREATMENT ADMINISTERED TO RESOLVE THE EFFUSION INITIALLY INVOLVED PERFORMING TAP IN THE POST-CARE UNIT. HOWEVER, DUE TO THE PATIENTS HABITUS, THIS WAS UNSUCCESSFUL, AND THE PROCEDURE WAS MOVED TO THE OPERATIVE ROOM FOR A PERICARDIAL WINDOW. THE EFFUSION WAS IDENTIFIED ONLY AFTER OPENING THE CHEST IN THE OPERATIVE ROOM, IT WAS DETERMINED THE SHEATH CAUSED UPON REMOVAL. ALSO NOTED PATIENTS RA TISSUE WAS ON THE THINNER SIDE. THE PATIENT WAS STABILIZED AND IS CURRENTLY ALIVE. THEY HAVE NOT BEEN DISCHARGED AND REMAIN AN INPATIENT AT THE HOSPITAL. THE PFA PROCEDURE WAS PERFORMED TO TREAT PAROXYSMAL ATRIAL FIBRILLATION. NO PHYSICAL DAMAGE WAS OBSERVED WITH ANY BOSTON SCIENTIFIC DEVICES PRIOR TO THE PROCEDURE, AND NO EFFUSION WAS NOTED IMMEDIATELY POST-PROCEDURE. IT WAS FURTHER REPORTED THAT THE PATIENT HAD PASSED AWAY. THE SHEATH DAMAGE WAS REPAIRED, AND THE PATIENT WAS RECOVERING AND HEMODYNAMICALLY STABLE. SUBSEQUENTLY, SHE DEVELOPED SOME NEUROLOGICAL CONCERNS, PROMPTING A CTA, WHICH REVEALED NO ABNORMALITIES. DESPITE THIS, SHE REMAINED INTUBATED. THE PHYSICIAN NOTED THAT SHE WAS EXPECTED TO RECOVER, ALTHOUGH IT WOULD HAVE TAKEN SOME TIME. HOWEVER, THE FAMILY CHOSE TO WITHDRAW CARE, AND THE PATIENT PASSED AWAY. IT WAS FURTHER REPORTED THAT THE DATE OF DEATH WAS NOT AVAILABLE. THE CAUSE OF DEATH WAS THE WITHDRAWAL OF CARE BY THE FAMILY. THE PATIENTS MEDICAL HISTORY AND CONTRIBUTING CONDITIONS INCLUDED PERSISTENT ATRIAL FIBRILLATION, A HISTORY OF SMOKING, AND AN OVERALL UNHEALTHY LIFESTYLE. NO AUTOPSY WAS PERFORMED. IT WAS FURTHER REPORTED THAT THERE WERE NO IMAGING AVAILABLE AS THE PROCEDURE WAS SATISFACTORY. THERE WAS NO ABLATION PERFORMED ON THE RIGHT ATRIUM AND THE TRANSEPTAL LOCATION WAS MID ANTERIOR. PER CUSTOMER, THE SURGICAL REPORT IS NOT AVAILABLE FOR REVIEW. Manufacturer Narrative: ADDITIONAL INFORMATION FOR SECTION A1, A2, A3 PATIENT INFORMATION, SECTION B2 OUTCOMES ATTRIB TO ADV EVENT, SECTION B2 DATE OF DEATH, SECTION B5 DESCRIBE EVENT OR PROBLEM AND SECTION H6 IMPACT CODE. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AS THE DEVICE WAS DISPOSED, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AS THE DEVICE WAS DISPOSED, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22461129&pc=DSQ | 2916596-2025-04506 | 2025-06-20 04:00 | Death | THORATEC CORPORATION | 2025-07-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS COMPASSIONATE WITHDRAWAL. Manufacturer Narrative: SECTION A4: THE PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22461224&pc=DSQ | 2916596-2025-04384 | 2025-06-20 04:00 | Death | THORATEC CORPORATION | 2025-07-09 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO MULTI-ORGAN FAILURE (MOF). THE PATIENT WAS ADMITTED DUE TO HEART FAILURE AND FOUND TO HAVE HIGH WHITE BLOOD CELL COUNT (WBC). THE PATIENT WAS CACHEXIC. THEY WERE TREATED WITH ANTIBIOTICS. ECHOCARDIOGRAM (ECHO) SUGGESTED LEFT VENTRICULAR EJECTION FRACTION (LVEF) 7%, CLOTS IN RIGHT ATRIUM, LEFT VENTRICLE, AND THROMBUS IN PULMONARY ARTERY (PA) WAS FOUND. THE PATIENT WAS SUPPORTED WITH INTRAVENOUS (IV) INOTROPES AND TOTAL PARENTERAL NUTRITION (TPN). LATER, THE PATIENT REQUIRED INTRA-AORTIC BALLOON PUMP (IABP) AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). PATIENT'S WBC COUNT WAS FOUND NORMAL. LIVER BILIRUBIN LEVELS WERE INCREASING. THE PATIENT WAS ALSO LISTED FOR TRANSPLANTATION, BUT THEY WERE UNABLE TO GET A HEART. THE PATIENT WAS IMPLANTED WITH HEARTMATE 3 (HM3) AS RESCUE. POST OPERATIVELY, THE PATIENT WENT INTO MULTI ORGAN FAILURE. THEY WERE SUPPORTED WITH RIGHT VENTRICULAR ASSIST DEVICE (RVAD), BUT IT DID NOT IMPROVE. THE DEVICE WORKED AS EXPECTED AND DID NOT CAUSE OR CONTRIBUTE TO MOF. RIGHT HEART FAILURE EXISTED PRIOR TO HM3 IMPLANT, WHICH WAS NOT CAUSED OR CONTRIBUTED BY ANY DEVICE RELATED ISSUE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), THROMBOSIS, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THIS SECTION, UNDER ¿ANTICOAGULATION¿, PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A. PATIENT INFORMATION. PT IDENTIFIER AND DATE OF BIRTH NOT REPORTED BY CUSTOMER. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22461446&pc=BTO | 22461446 | 2025-06-26 04:00 | Death | COVIDIEN | 2025-07-09 04:00 | BTO | SHILEY ET TUBE CONNECTOR | Break | Cardiac Arrest; Diminished Pulse Pressure |  |  | 1 | Event Description: PT. ARRIVED AT HOSPITAL VIA AIR AMBULANCE FROM (B)(6) HOSPITAL IN (B)(6). UPON CHANGING THE VENTILATOR TUBING FROM TRANSPORT MACHINE TO OUR VENTILATOR, THE HUB CONNECTING THE ETT SNAPPED AT THE LEVEL OF THE ETT. WE EMERGENTLY SET UP FOR ETT EXCHANGE. AIRWAY EXCHANGE COMPLETED WITH A BOUGIE, RIGHT AS PT BECAME PULSELESS. CPR IMMEDIATELY STARTED AND PT BAGGED VIA NEW ETT. EVENT OCCURRED AT APPROX 1140. PATIENT DIED AT 1537. NOTE: PRODUCT CAME FROM (B)(6) HOSPITAL, (B)(6). THEY PROVIDED ME WITH THE FOLLOWING INFORMATION ABOUT THE PRODUCTS: 100636 - CIRCUIT VENT TUBING. ET TUBE CONNECTORS - SHILEY MANUFACTURER: COVIDIEN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22461876&pc=MCX | 3004742232-2025-00195 | 2025-06-18 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-07-09 04:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest; Obstruction/Occlusion; Embolism/Embolus | P130005 |  | 1 | Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A LEFT ANTERIOR DESCENDING ARTERY (LAD). FOUR LOW SPEED TREATMENTS OF 25 SECONDS EACH, WITH 25 SECONDS OF RESTING TIME, WERE PERFORMED IN THE PROXIMAL LAD. AFTER THE FOURTH TREATMENT, IT WAS OBSERVED THE PATIENT HAD TRANSIENT SLOW FLOW. 100 MCG OF ADENOSINE INTRA-CORONARY WAS ADMINISTERED SUCCESSFULLY AS FLOW WAS RESTORED. THE PHYSICIAN THEN PROCEEDED TO TREAT THE MID LAD. THREE LOW SPEED TREATMENTS OF 20 SECONDS EACH, WITH LONGER RESTING TIME, WERE PERFORMED. THERE WERE SLIGHT ECG CHANGES NOTED. THE ECG CHANGES SETTLED POST ORBITAL ATHERECTOMY. THE PHYSICIAN THEN DILATED THE MID LAD WITH A NON-ABBOTT BALLOON - 3 X 12 MM WITH LOW INFLATION. AT THIS TIME, THE PATIENT HAD CHANGES TO ECG AND BLOOD PRESSURE DECREASED DURING INFLATION AND POST INFLATION. METARAMINOL WAS REQUIRED. THE PATIENT STABILIZED FOR A SHORT PERIOD BUT CHANGES TO ECG AND DECREASED BLOOD PRESSURE CONTINUED REQUIRING CARDIOPULMONARY RESUSCITATION AND INTRAVENOUS (IV) METARAMINOL AND IV ADRENALINE ADMINISTERED. THE PATIENT WAS STABILIZED AND INTUBATED. THE PHYSICIAN WAS ABLE TO SUCCESSFULLY DELIVER AND TREAT THE LEFT MAIN ARTERY TO PROXIMAL LAD, AND PROXIMAL LAD TO MID LAD WITH A NON-ABBOTT DRUG-ELUTING STENTS. THERE WAS NO OVERLAPPING BETWEEN THE IMPLANTED STENTS. POST IMPLANT OF THE STENTS, THE PATIENT'S BLOOD PRESSURE DECREASED AND THE PATIENT WENT INTO VENTRICULAR FIBRILLATION. THE PATIENT EXPIRED. IN THE PHYSICIAN'S OPINION, THE CAUSE OF THE SLOW FLOW WAS DUE TO MICROEMBOLIZATION OF CALCIUM AND POSSIBLY THROMBUS ON CALCIUM NODULE. IT WAS INDICATED THE LOW SPEED ORBITAL ATHERECTOMY TREATMENT LED TO THE SLOW FLOW. IN THE PHYSICIAN'S OPINION, THE PRIMARY AND SECONDARY CAUSE OF PATIENT DEATH WAS CARDIOGENIC SHOCK DUE TO ISCHEMIA FROM HIGH RISK PERCUTANEOUS CORONARY INTERVENTION.Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A LEFT ANTERIOR DESCENDING ARTERY (LAD). FOUR LOW SPEED TREATMENTS OF 25 SECONDS EACH, WITH 25 SECONDS OF RESTING TIME, WERE PERFORMED IN THE PROXIMAL LAD. AFTER THE FOURTH TREATMENT, IT WAS OBSERVED THE PATIENT HAD SLOW FLOW. 100 MCG OF ADENOSINE INTRA-CORONARY WAS ADMINISTERED SUCCESSFULLY AS FLOW WAS RESTORED. THE PHYSICIAN THEN PROCEEDED TO TREAT THE MID LAD. THREE LOW SPEED TREATMENTS OF 20 SECONDS EACH, WITH LONGER RESTING TIME, WERE PERFORMED. THERE WERE SLIGHT ECG CHANGES NOTED. THE ECG CHANGES SETTLED POST ORBITAL ATHERECTOMY. THE PHYSICIAN THEN DILATED THE MID LAD WITH A NON-ABBOTT BALLOON - 3 X 12 MM WITH LOW INFLATION. AT THIS TIME, THE PATIENT HAD CHANGES TO ECG AND BLOOD PRESSURE DECREASED DURING INFLATION AND POST INFLATION. AROMINE WAS REQUIRED. THE PATIENT STABILIZED FOR A SHORT PERIOD BUT CHANGES TO ECG AND DECREASED BLOOD PRESSURE CONTINUED REQUIRING CARDIOPULMONARY RESUSCITATION AND INTRAVENOUS (IV) AROMINE AND IV ADRENALINE ADMINISTERED. THE PATIENT WAS STABILIZED AND INTUBATED. THE PHYSICIAN WAS ABLE TO SUCCESSFULLY DELIVER AND TREAT THE LEFT MAIN ARTERY TO PROXIMAL LAD, AND PROXIMAL LAD TO MID LAD WITH A NON-ABBOTT DRUG-ELUTING STENTS. THERE WAS NO OVERLAPPING BETWEEN THE IMPLANTED STENTS. POST IMPLANT OF THE STENTS, THE PATIENT'S BLOOD PRESSURE DECREASED AND THE PATIENT WENT INTO VENTRICULAR FIBRILLATION. THE PATIENT EXPIRED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THE SLOW FLOW WAS TRANSIENT AND RESPONSIVE TO INTRACORONARY ADENOSINE. IN THE PHYSICIAN'S OPINION, THE CAUSE OF THE SLOW FLOW WAS DUE TO MICROEMBOLIZATION OF CALCIUM AND POSSIBLY THROMBUS ON CALCIUM NODULE. IT WAS INDICATED THE LOW SPEED ORBITAL ATHERECTOMY TREATMENT LED TO THE SLOW FLOW. IN THE PHYSICIAN'S OPINION, THE PRIMARY AND SECONDARY CAUSE OF PATIENT DEATH WAS CARDIOGENIC SHOCK DUE TO ISCHEMIA FROM HIGH RISK PCI. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED PATIENT ARRHYTHMIA, CARDIAC ARREST, OBSTRUCTION, EMBOLISM, MEDICATION, UNEXPECTED MEDICAL INTERVENTION, AND DEATH APPEAR TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE IT IS POSSIBLE THAT THE REPORTED PATIENT ARRHYTHMIA, CARDIAC ARREST, OBSTRUCTION, EMBOLISM, MEDICATION, UNEXPECTED MEDICAL INTERVENTION, AND DEATH ARE A RESULT OF THE PATIENT¿S DISEASE SEVERITY COMBINED WITH USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. H6: CODES 4118, 3233, AND 11 NO LONGER APPLY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22461937&pc=NVY | 2017865-2025-91650 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-09 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND UNCONSCIOUS. THE PATIENT WAS TAKEN TO THE EMERGENCY ROOM AND EXPIRED SHORTLY AFTER DUE TO CARDIAC ARREST. THE PHYSICIAN DID NOT BELIEVE THAT A MALFUNCTION OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV) OR RIGHT ATRIAL (RA) LEADS, CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22461970&pc=DSQ | 2916596-2025-04271 | 2025-06-18 04:00 | Death | THORATEC CORPORATION | 2025-07-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Ischemia | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS A BRIEF POWER OUTAGE AT THE HOSPITAL ON 19JUN2025 AROUND 18:00. THE PATIENT WAS CONNECTED TO THE POWER MODULE (PM) WHICH WAS PLUGGED INTO THE RED EMERGENCY OUTLET. THE PATIENT AROUND 18:00 WENT INTO PULSELESS ELECTRICAL ACTIVITY (PEA). THE SITE WAS UNSURE IF THE PATIENT WAS CONNECTED TO THE PM OR BATTERIES AT THAT TIME AND WERE UNSURE IF THE POWER OUTAGE LED TO ANYTHING HAPPENING WITH THE LEFT VENTRICULAR ASSIST DEVICE (LVAD). IT WAS NOTED THAT THERE WERE 3 OTHER LVADS IN THE HOSPITAL THAT DID NOT HAVE ANY ISSUES. LOG FILE REVIEW WAS REQUESTED, AND THE ONLY EVENTS CAPTURED ON 18JUN2025 WERE PULSATILITY INDEX (PI) EVENTS AND POWER CABLE DISCONNECTS. ON (B)(6) 2025 AT THE 18:00 HOUR TIME FRAME THE PATIENT WAS CONNECTED TO THE PM. AT THIS TIME OR AT ANY TIME IN THE LOG FILE THERE WAS NO LOSS OF POWER TO THE PM. THE LOG FILE CAPTURED LOW FLOW EVENTS ON 19JUN2025. THE REASON FOR THE PATIENT'S PULSELESS ELECTRICAL ACTIVITY WAS NOT KNOWN, AND NO DEVICE MALFUNCTION WAS IDENTIFIED. THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO ISCHEMIC BOWEL. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. REVIEW OF THE SUBMITTED LOG FILES CONFIRMED INTERMITTENT LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THESE FINDINGS COULD NOT BE CONCLUSIVELY DETERMINED. REVIEW OF THE SUBMITTED LOG FILES CONFIRMED INTERMITTENT LOW FLOW ALARMS ON (B)(6) 2025. ELEVATED PULSATILITY INDEX (PI) VALUES WERE ALSO NOTED DURING THE LOW FLOW EVENTS. THE SYSTEM APPEARED TO BE OPERATING AS INTENDED AT THE SET SPEED, AND THERE WERE NO OTHER NOTABLE ALARMS ACTIVE IN THE LOG FILES. IT WAS REPORTED THAT HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED THROUGH THE MANUFACTURING EXECUTION SYSTEM (MES) AND SHOWED NO RELEVANT DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS VENOUS THROMBOEMBOLISM, ARTERIAL NON-CENTRAL NERVOUS SYSTEM (CNS) THROMBOEMBOLISM, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ OUTLINES INDICATIONS OF THROMBUS AS WELL AS HOW TO RESPOND TO SUCH EVENTS. THIS SECTION ALSO PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE. SECTION 1 OF THE IFU PROVIDES AN EXPLANATION OF ALL PUMP PARAMETERS, INCLUDING FLOW, POWER, AND PULSATILITY INDEX (PI). THIS SECTION EXPLAINS THAT PUMP FLOW IS A CALCULATED VALUE THAT IS ESTIMATED BASED ON PUMP POWER. SECTION 1 EXPLAINS THAT IN GENERAL, THE MAGNITUDE OF THE PI VALUE IS RELATED TO THE AMOUNT OF ASSISTANCE PROVIDED BY THE PUMP. HIGHER VALUES INDICATE MORE VENTRICULAR FILLING AND HIGHER PULSATILITY (IE, THE PUMP IS PROVIDING LESS SUPPORT TO THE LEFT VENTRICLE). LOWER VALUES INDICATE LESS VENTRICULAR FILLING AND LOWER PULSATILITY (IE, THE PUMP IS PROVIDING GREATER SUPPORT AND FURTHER UNLOADING THE VENTRICLE). SECTION 4 ¿HEARTMATE TOUCH COMMUNICATION SYSTEM¿ STATES THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.5 LPM (LITERS PER MINUTE) AND EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7 ¿ALARMS AND TROUBLESHOOTING¿ DESCRIBES ALARM CONDITIONS, INCLUDING THE LOW FLOW HAZARD, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. THE HEARTMATE 3 LVAS PATIENT HANDBOOK, IS CURRENTLY AVAILABLE. SECTION 5 "ALARMS AND TROUBLESHOOTING" OUTLINES SYSTEM CONTROLLER ALARMS AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. THIS DOCUMENT INSTRUCTS THE USER THAT IN THE EVENT OF A LOW FLOW HAZARD ALARM, CALL YOUR HOSPITAL CONTACT IMMEDIATELY FOR DIAGNOSIS AND INSTRUCTIONS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: CORRECTION: DATE OF DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22462282&pc=DSQ | 2916596-2025-04200 | 2023-01-01 05:00 | Death | THORATEC CORPORATION | 2025-07-09 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "GENDER DIFFERENCE AND OUTCOMES IN HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE SUPPORT: RESULTS FROM THE J-MACS REGISTRY" THAT HEARTMATE 3 LVAS MAY BE ASSOCIATED WITH BLEEDING, INFECTION, ARRHYTHMIAS, RIGHT HEART FAILURE, NEUROLOGICAL DISORDER, RENAL DYSFUNCTION, AORTIC REGURGITATION, DEVICE MALFUNCTION AND DEATH. PATIENTS WITH ADVANCED HEART FAILURE WHO UNDERWENT HEARTMATE 3 LVAD IMPLANTATION BETWEEN 2019 AND 2023 AND WERE PROSPECTIVELY REGISTERED IN THE JAPANESE REGISTRY FOR MECHANICALLY ASSISTED CIRCULATORY SUPPORT WERE INCLUDED. CUMULATIVE MORTALITY, CHANGES IN NEW YORK HEART ASSOCIATION (NYHA) FUNCTIONAL CLASS, AND ADVERSE EVENTS LEADING TO REHOSPITALIZATION WERE RETROSPECTIVELY COMPARED BETWEEN GENDERS. A TOTAL OF 465 PATIENTS WHO RECEIVED HEARTMATE 3 IMPLANTATION (MEDIAN AGE: 50 YEARS; WOMEN: 26.5%) WERE ANALYZED. WOMEN HAD SIGNIFICANTLY SMALLER BSA (1.45 M2 VERSUS 1.70 M2 MEDIAN, P < 0.001) AND LOWER SYSTOLIC BLOOD PRESSURE (85 MMHG VERSUS 89 MMHG MEDIAN, P = 0.019). THE MEDIAN INDEX HOSPITAL STAY WAS LONGER FOR WOMEN THAN FOR MEN (85 DAYS VERSUS 75 DAYS, P = 0.030). THERE WAS NO SIGNIFICANT DIFFERENCE IN CUMULATIVE MORTALITY BETWEEN GENDERS DURING A MEDIAN FOLLOW-UP OF 964 DAYS (P = 0.499). WOMEN EXHIBITED A HIGHER RISK OF REHOSPITALIZATION DUE TO MAJOR BLEEDING EVENTS, WITH AN ADJUSTED HAZARD RATIO OF 3.343 (95% CONFIDENCE INTERVAL: 1.225¿10.018, P = 0.031). THE BREAKDOWN OF ADVERSE EVENTS THAT REQUIRED HOSPITALIZATION WERE MAJOR INFECTION, ARRHYTHMIAS, MAJOR BLEEDING, RIGHT HEART FAILURE, NEUROLOGICAL DISORDER, RENAL DYSFUNCTION, AORTIC REGURGITATION, AND DEVICE MALFUNCTION. OUT OF PATIENTS INCLUDED IN THE STUDY; 91 EXPERIENCED MAJOR INFECTION (35.7%), 34 EXPERIENCED ARRHYTHMIA (13.3%), 27 EXPERIENCED MAJOR BLEEDING (10.6%), 25 HAD RIGHT HEART FAILURE (9.8%), AND 13 HAD NEUROLOGICAL DISORDER (5.1%). RENAL DISFUNCTION OCCURRED IN 7 PATIENTS (2.7%), AORTIC REGURGITATION OCCURRED IN 4 (1.6%) AND DEVICE MALFUNCTION OCCURRED IN 3 PATIENTS (1.2%). THE CUMULATIVE INCIDENCE OF MAJOR BLEEDING WAS SIGNIFICANTLY HIGHER IN WOMEN THAN MEN (16.7% VERSUS 6.3%, P = 0.004; FIG. 5A), WHEREAS THERE WERE NO SIGNIFICANT DIFFERENCES IN THOSE OF OTHER EVENTS BETWEEN WOMEN AND MEN (P > 0.05 FOR ALL.)ALTHOUGH WOMEN DEMONSTRATED AN INCREASED RISK OF MAJOR BLEEDING REQUIRING HOSPITALIZATION, THIS DID NOT AFFECT OVERALL MORTALITY FOLLOWING HEARTMATE 3 IMPLANTATION. Manufacturer Narrative: B5: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS 01 JAN 2023 AS PATIENTS WERE IMPLANTED BETWEEN 2019 AND 2023. SECOND DEPARTMENT OF INTERNAL MEDICINE, UNIVERSITY OF TOYAMA, TOYAMA, JAPAN DEPARTMENT OF CARDIOVASCULAR SURGERY, CHIBA UNIVERSITY GRADUATE SCHOOL OF MEDICINE, CHIBA, JAPAN DIVISION OF CARDIOVASCULAR SURGERY, TOHOKU UNIVERSITY, SENDAI, JAPAN DEPARTMENT OF CARDIAC SURGERY, UNIVERSITY OF TOKYO, TOKYO, JAPAN. M. NAKAMURA, T. IMAMURA, M. NAKAGAITO, ET AL., GENDER DIFFERENCE AND OUTCOMES IN HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE SUPPORT: RESULTS FROM THE J-MACS REGISTRY, (2025), HTTPS://DOI.ORG/10.1016/J.JJCC.2025.06.010. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) DEVICES AND THE REPORTED OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. A IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE REPORTED INFORMATION WAS OBTAINED THROUGH A LITERATURE REVIEW. THE HEARTMATE 3 DEVICE SERIAL NUMBERS AND OTHER SPECIFIC CASE/PATIENT INFORMATION ARE NOT AVAILABLE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22462307&pc=PNJ | 2017865-2025-91752 | 2025-06-11 04:00 | Death | ABBOTT | 2025-07-09 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER THE LEADLESS (LP) PACEMAKER IMPLANT WAS COMPLETED, THE PATIENT EXPERIENCED CARDIAC ARREST RESULTING IN PULSELESS HEART ACTIVITY. A THORACOTOMY AND CARDIAC MASSAGE WERE PERFORMED. DURING THE CARDIAC MASSAGE, THE LP DISLODGED AND MIGRATED TO THE PULMONARY ARTERY. THE LP WAS UNABLE TO BE RETRIEVED WITH A RETRIEVAL CATHETER DUE TO A STRETCHED HELIX. THE PACEMAKER WAS SURGICALLY REMOVED. THE PATIENT WAS ALERT WITH NORMAL VITAL SIGNS FOLLOWING THE PROCEDURE, HOWEVER, PASSED AWAY A FEW DAYS LATER. THE PHYSICIAN ALLEGED THE PATIENT PASSED AWAY DUE TO A PULMONARY EMBOLISM AND THE PROCEDURE TIME MAY HAVE CONTRIBUTED TO THE PATIENT'S PASSING. THE PHYSICIAN ALLEGED THAT THE ABBOTT PRODUCTS USED WERE NOT RELATED TO THE CARDIAC ARREST OR THE PATIENT DEATH. ADDITIONALLY, IT WAS STATED THAT THE PATIENT EXPERIENCED A PLEURAL EFFUSION. IT IS UNKNOWN IF THE PULMONARY EMBOLISM OR PLEURAL EFFUSION WAS RELATED TO THE ABBOTT PRODUCTS OR THE PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22462310&pc=PNJ | 2017865-2025-91750 | 2025-06-11 04:00 | Death | ABBOTT | 2025-07-09 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Stretched; Connection Problem; Device Dislodged or Dislocated | Arrhythmia; Embolism/Embolus; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER THE LEADLESS (LP) PACEMAKER IMPLANT WAS COMPLETED, THE PATIENT EXPERIENCED CARDIAC ARREST RESULTING IN PULSELESS HEART ACTIVITY. A THORACOTOMY AND CARDIAC MASSAGE WERE PERFORMED. DURING THE CARDIAC MASSAGE, THE LP DISLODGED AND MIGRATED TO THE PULMONARY ARTERY. THE LP WAS UNABLE TO BE RETRIEVED WITH A RETRIEVAL CATHETER DUE TO A STRETCHED HELIX. THE PACEMAKER WAS SURGICALLY REMOVED. THE PATIENT WAS ALERT WITH NORMAL VITAL SIGNS FOLLOWING THE PROCEDURE, HOWEVER, PASSED AWAY A FEW DAYS LATER. THE PHYSICIAN ALLEGED THE PATIENT PASSED AWAY DUE TO A PULMONARY EMBOLISM AND THE PROCEDURE TIME MAY HAVE CONTRIBUTED TO THE PATIENT'S PASSING. THE PHYSICIAN ALLEGED THAT THE ABBOTT PRODUCTS USED WERE NOT RELATED TO THE CARDIAC ARREST OR THE PATIENT DEATH. ADDITIONALLY, IT WAS STATED THAT THE PATIENT EXPERIENCED A PLEURAL EFFUSION. IT IS UNKNOWN IF THE PULMONARY EMBOLISM OR PLEURAL EFFUSION WAS RELATED TO THE ABBOTT PRODUCTS OR THE PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22462311&pc=MMX | 2017865-2025-91753 | 2025-06-11 04:00 | Death | ABBOTT | 2025-07-09 04:00 | MMX | AVEIR¿ RETRIEVAL SYSTEM CATHETER | Connection Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K213494 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER THE LEADLESS (LP) PACEMAKER IMPLANT WAS COMPLETED, THE PATIENT EXPERIENCED CARDIAC ARREST RESULTING IN PULSELESS HEART ACTIVITY. A THORACOTOMY AND CARDIAC MASSAGE WERE PERFORMED. DURING THE CARDIAC MASSAGE, THE LP DISLODGED AND MIGRATED TO THE PULMONARY ARTERY. THE LP WAS UNABLE TO BE RETRIEVED WITH A RETRIEVAL CATHETER DUE TO A STRETCHED HELIX. THE PACEMAKER WAS SURGICALLY REMOVED. THE PATIENT WAS ALERT WITH NORMAL VITAL SIGNS FOLLOWING THE PROCEDURE, HOWEVER, PASSED AWAY A FEW DAYS LATER. THE PHYSICIAN ALLEGED THE PATIENT PASSED AWAY DUE TO A PULMONARY EMBOLISM AND THE PROCEDURE TIME MAY HAVE CONTRIBUTED TO THE PATIENT'S PASSING. THE PHYSICIAN ALLEGED THAT THE ABBOTT PRODUCTS USED WERE NOT RELATED TO THE CARDIAC ARREST OR THE PATIENT DEATH. ADDITIONALLY, IT WAS STATED THAT THE PATIENT EXPERIENCED A PLEURAL EFFUSION. IT IS UNKNOWN IF THE PULMONARY EMBOLISM OR PLEURAL EFFUSION WAS RELATED TO THE ABBOTT PRODUCTS OR THE PROCEDURE. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. D4 - PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER CANNOT BE PROVIDED AS A PRODUCT IDENTIFIER COULD NOT BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22462316&pc=DSQ | 2916596-2025-04401 | 2023-07-01 04:00 | Death | THORATEC CORPORATION | 2025-07-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿CLINICAL OUTCOMES OF LEFT VENTRICULAR ASSIST DEVICE BLEEDING COMPLICATION¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH PUMP EXCHANGE, BLEEDING, DEATH, AND TRANSPLANT. THIS RETROSPECTIVE STUDY EVALUATED 78 PATIENTS WHO UNDERWENT CONTINUOUS FLOW LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION AND HEMODYNAMIC RAMP TEST AFTER LVAD IMPLANTATION BETWEEN JUL2017 AND JUL2023. PATIENTS WHOSE HEMODYNAMIC RAMP TESTS WERE NOT PERFORMED BEFORE HOSPITAL DISCHARGE AND THOSE IMPLANTED WITH A RIGHT VENTRICULAR ASSIST DEVICE/BIVENTRICULAR ASSIST DEVICES WERE EXCLUDED. ADDITIONALLY, 8 PATIENTS WERE EXCLUDED FROM THE STUDY DUE TO PERIOPERATIVE DEATH AFTER LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION AND 2 PATIENTS WERE EXCLUDED DUE TO PUMP EXCHANGE DURING THE PERIOPERATIVE PERIOD. 2 PATIENTS WERE ALSO EXCLUDED FROM THE TRIAL DUE TO HAVING CATECHOLAMINE SUPPORT. OF NOTE, 19 PATIENTS WITH A HEARTMATE II AND 20 PATIENTS WITH A DEVICE OTHER THAN THE HEARTMATE II OR 3 WERE ALSO INCLUDED IN THE ANALYSIS RESULTS, PATIENT BACKGROUND INFORMATION, AND OUTCOMES. THE STUDY INVESTIGATED THE PREOPERATIVE CHARACTERISTICS AND HEMODYNAMICS AT HEMODYNAMIC RAMP TESTS IN TERMS OF THE OCCURRENCE OF BLEEDING COMPLICATIONS. THE STUDY HYPOTHESIZED THAT HEMODYNAMICS AFTER LVAD IMPLANTATION MAY INFLUENCE THE OCCURRENCE OF BLEEDING COMPLICATIONS AFTER LVAD IMPLANTATION. THE AVERAGE PATIENT AGE AT TIME OF LVAD IMPLANTATION WAS 53 (41-59, P=0.12). 21 PATIENTS WERE FEMALE, 57 WERE MALE. THE MEAN FOLLOW UP DURATION WAS 1015+/- 576 DAYS. 13 PATIENTS HAD BLEEDING COMPLICATIONS AFTER LVAD IMPLANT. BLEEDING COMPLICATIONS WERE IDENTIFIED AS EVENTS THAT RESULTED IN HOSPITALIZATION FOR INVASIVE TREATMENTS, SUCH AS ENDOSCOPY, SURGERY, AND INTERVENTIONAL RADIOLOGY, AND/OR NECESSITATED THE TRANSFUSION OF AT LEAST 2 UNITS OF PACKED RED BLOOD CELLS WITHIN 24 HOURS OF THE EVENT. THE RATES OF FREEDOM FROM BLEEDING COMPLICATIONS AT 1, 3, AND 5 YEARS WERE 94%, 85%, AND 74%. 6 PATIENTS EXPERIENCED GASTROINTESTINAL (GI) BLEEDING, 3 PATIENTS EXPERIENCED NOSE BLEEDING, AND 2 PATIENTS EXPERIENCED AN INTRAPERITONEAL HEMORRHAGE. 1 PATIENT EXPERIENCED HEMATURIA, AND 1 PATIENT EXPERIENCED A FEMORAL SUBCUTANEOUS HEMATOMA. OF THE 6 PATIENTS THAT HAD GI BLEEDS, 3 UNDERWENT ENDOSCOPIC HEMOSTASIS AND 1 UNDERWENT INTERVENTIONAL RADIOLOGY FOR HEMOSTASIS. 1 OF THE 3 PATIENTS WITH A NOSEBLEED UNDERWENT HEMOSTASIS IN THE KIESELBACH AREA. OF THE 2 PATIENTS THAT EXPERIENCED INTRAPERITONEAL HEMORRHAGE, 1 UNDERWENT OPEN ABDOMINAL SURGERY AND THE OTHER UNDERWENT LAPAROSCOPIC SURGERY FOR HEMOSTASIS. ECHOCARDIOGRAMS OF THE PATIENTS BOTH WITH AND WITHOUT BLEEDING WERE TAKEN. THE CENTER FOUND THAT HIGH PREOPERATIVE BRACHIAL-ANKLE PULSE WAVE VELOCITY (BAPWV) AND HIGH SYSTEMIC VASCULAR RESISTANCE (SVR) IN THE HEMODYNAMIC RAMP TEST WERE SIGNIFICANTLY ASSOCIATED WITH BLEEDING COMPLICATIONS AFTER LVAD IMPLANTATION. ARTERIOSCLEROSIS WAS ALSO FOUND TO BE A RISK FACTOR FOR BLEEDING COMPLICATIONS AFTER LVAD IMPLANTATION. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN, DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01JUL2023 AS DATA WAS COLLECTED BETWEEN JUL2017 AND JUL2023. SECTION E1: REPORTER PHONE NUMBER (B)(6). AUTHOR INFORMATION: IMAOKA, S., YOSHIOKA, D., SAITO, S., KAWAMURA, T., KAWAMURA, A., TODA, K., & MIYAGAWA, S. (2025). CLINICAL OUTCOMES OF LEFT VENTRICULAR ASSIST DEVICE BLEEDING COMPLICATION. ASAIO JOURNAL (AMERICAN SOCIETY FOR ARTIFICIAL INTERNAL ORGANS : 1992), 71(3), 229¿234. HTTPS://DOI.ORG/10.1097/MAT.0000000000002317 OSAKA UNIVERSITY GRADUATE SCHOOL OF MEDICINE, SUITA SHI, OSAKA FU, JAPAN. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. A IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22462624&pc=KGC | 9611594-2025-00144 |  | Death | AVANOS MEDICAL INC. | 2025-07-09 04:00 | KGC | INTRODUCER KIT FOR GASTROSTOMY FEEDING TUBE | Malposition of Device | Insufficient Information | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED, THE FACILITY REPORTED TWO CASED OF DEATH RELATED TO THE INTRODUCER KIT PROCEDURE. PER ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2025, THERE WAS ONLY ONE CASE OF DEATH RELATED TO THE INTRODUCER KIT PROCEDURE. CUSTOMER ¿ESTIMATED THAT THE SET WAS GOOD BUT THE TECHNIQUE WAS FAILING. WHY, HE DID NOT KNOW. THE OUTCOME WAS THAT VENTRICLE WAS PUNCTURED AND PATIENT DIED TO THESE CONSEQUENCES.¿ Manufacturer Narrative: THE ACTUAL COMPLAINT PRODUCT WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD IS NOT POSSIBLE AS NO LOT NUMBER WAS PROVIDED; THEREFORE, THE UDI-PI IS UNAVAILABLE. ALL INFORMATION REASONABLY KNOWN AS OF 09 JUL 2025 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AVANOS MEDICAL INC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AVANOS MEDICAL INC. AVANOS MEDICAL INC. HAS NO INDEPENDENT KNOWLEDGE OF THE REPORTED EVENT BUT IS RELAYING THE INFORMATION PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AVANOS MEDICAL COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN AVANOS MEDICAL INC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22462626&pc=KGC | 9611594-2025-00145 | 2025-03-27 04:00 | Death | AVANOS MEDICAL INC. | 2025-07-09 04:00 | KGC | INTRODUCER KIT FOR GASTROSTOMY FEEDING TUBE | Break; Malposition of Device | Aspiration Pneumonitis | K080253 |  | 1 | Event Description: IT WAS REPORTED ON (B)(6) 2025, A PATIENT DIED SHORTLY AFTER GASTRONOMY. THE FACILITY BELIEVE IT IS UNLIKELY THAT AN ISSUE WITH THE KIT CONTRIBUTED TO THE SUTURES FAILING. PER ADDITIONAL INFORMATION RECEIVED ON REPORTABLE ALERT DATE (B)(6) 2025, "PATIENT UNDERWENT ELECTIVE INSERTION OF 14F MICKEY BUTTON USING THE INTRODUCER KIT ON (B)(6). COMMENCED ON ROUTINE FEEDS POST-OPERATIVELY AND DISCHARGED HOME 48 HOURS LATER. REPRESENTED TO LOCAL HOSPITAL AFTER 4 DAYS WITH SUSPECTED PNEUMONIA. UNDERWENT CHANGE IN WATER IN THE BALLOON ON (B)(6) (WAS RECOMMENDED TO NOT PERFORM BEFORE (B)(6)) - UNSURE IF ANY CHECKS TOOK PLACE LOCALLY TO CONFIRM POSITION OF BALLOON - RECOMMENCED ON FEEDS AND DETERIORATED CLINICALLY - CT SCAN SUSPICIOUS OF BALLOON INFLATION OUTSIDE THE STOMACH - LAPAROTOMY ON (B)(6) - CONFIRMED BALLOON OUTSIDE STOMACH - T-FASTENERS CAME AWAY EASILY FROM ABDOMINAL WALL WHEN DRESSING WAS REMOVED AT TIME OF LAPAROTOMY. AT LAPAROTOMY, DEFECT IN STOMACH CLOSED AND FREKA GASTROSTOMY SITED. PATIENT REMAINED IN PICU/HDU ENVIRONMENT FOR 6-7 WEEKS POST-PROCEDURE. DEVELOPED AN ASPIRATION PNEUMONITIS FOLLOWING A FOLLOW-THROUGH STUDY TO ASSESS FOR AN OBSTRUCTION - SUBSEQUENTLY DEMISED 24 HOURS LATER ON (B)(6) 2025. INCIDENT REVIEW MEETING COMPLETED AT (B)(6) WITH LOCAL HOSPITAL." ENDOSCOPY WAS USED AT THE TIME OF INITIAL SURGERY TO VERIFY TUBE PLACEMENT. THE BALLOON INFLATION VOLUME WAS 5ML. Manufacturer Narrative: THE ACTUAL COMPLAINT PRODUCT WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD IS NOT POSSIBLE AS NO LOT NUMBER WAS PROVIDED; THEREFORE, THE UDI-PI IS UNAVAILABLE. ALL INFORMATION REASONABLY KNOWN AS OF 09 JUL 2025 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AVANOS MEDICAL INC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AVANOS MEDICAL INC. AVANOS MEDICAL INC. HAS NO INDEPENDENT KNOWLEDGE OF THE REPORTED EVENT BUT IS RELAYING THE INFORMATION PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AVANOS MEDICAL COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN AVANOS MEDICAL INC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22464234&pc=NPT | 2015691-2025-05587 | 2025-06-19 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-09 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY OUR AFFILIATES IN CANADA, IT WAS A CASE WITH A 29MM SAPIEN 3 ULTRA RESILIA VALVE, IN AORTIC POSITION BY TRANSFEMORAL APPROACH. DURING THE PROCEDURE, THE VALVE WAS DEPLOYED IN THE INTENDED POSITION. HOWEVER, A FEW MINUTES AFTER THE PROCEDURE, THE PATIENT EXPERIENCED SIGNIFICANT BLEEDING AND THE BLOOD PRESSURE WAS DECREASING. DESPITE THE ADMINISTRATION OF A BLOOD TRANSFUSION, THE PATIENT RAPIDLY DECOMPENSATED DUE TO AN ANNULAR RUPTURE. THE PATIENT PASSED AWAY. THE PERCEIVED ROOT CAUSE OF THE ANNULAR RUPTURE WAS EXCESSIVE CALCIUM IN THE ANNULUS. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE EXCESSIVE CALCIUM IN THE ANNULUS CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22464242&pc=NVY | 2017865-2025-91764 | 2025-06-24 04:00 | Death | ABBOTT | 2025-07-09 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Failure to Capture; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO HEART FAILURE. LOSS OF CAPTURE WAS OBSERVED ON THE DEVICE AND THE RIGHT VENTRICULAR (RV) LEAD. THERE WAS NO ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22464249&pc=LWS | 2017865-2025-91763 | 2025-06-24 04:00 | Death | ABBOTT | 2025-07-09 04:00 | LWS | ELLIPSE VR ICD, US | Failure to Capture; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO HEART FAILURE. LOSS OF CAPTURE WAS OBSERVED ON THE DEVICE AND THE RIGHT VENTRICULAR (RV) LEAD. THERE WAS NO ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22464539&pc=MIH | 3007284313-2025-04092 | 2025-04-01 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-07-09 04:00 | MIH | NOT OTHERWISE SPECIFIED (NOS) | Insufficient Information | Unspecified Infection |  |  | 1 | Event Description: PATEL, R., FATIMA, J., KANAMORI, L., ODERICH, G., SAQIB, N., TREFFALLS, R., LIMA, G., MENDES, B. (2025) A MULTI-INSTITUTIONAL COLLABORATION EVALUATING THE USE OF ILIAC BRANCH ENDOPROSTHESIS TO REPAIR INTERNAL ILIAC ARTERY ANEURYSMS. JOURNAL OF VASCULAR SURGERY, 81(4), E43. THE FOLLOWING WAS LISTED AMONG THE ABSTRACTS OF THE 2025 SOCIETY FOR CLINICAL VASCULAR SURGERY ANNUAL SYMPOSIUM: A MULTI-INSTITUTE RETROSPECTIVE REVIEW WAS PERFORMED FOR PATIENTS THAT UNDERWENT PLACEMENT OF GORE®¿EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS [IBE] FOR INTERNAL ILIAC ARTERY ANEURYSMS BETWEEN JANUARY 2013 AND JUNE 2024. THERE WERE 110 IBES PLACED IN 95 PATIENTS [MEAN AGE 73.8 YEARS OLD, 92% MALE] DURING THE STUDY PERIOD. TECHNICAL SUCCESS WAS ACHIEVED IN 98% OF IBES DEPLOYED (108/110). THERE WERE NO IN-HOSPITAL MORTALITIES AND THERE WAS ONE AORTIC-RELATED DEATH RELATED TO ENDOGRAFT INFECTION NECESSITATING EXPLANTATION 11 MONTHS AFTER IMPLANTATION. A TOTAL OF THREE PATIENTS DEVELOPED BUTTOCK CLAUDICATION AFTER IBE DURING INITIAL HOSPITALIZATION THAT ALL REQUIRED EXTENSIVE SIDE BRANCH EMBOLIZATION. IN TERMS OF INTERNAL ILIAC BRANCH SELECTION, 50% WERE THE GORE IBE BRANCH, 41% WERE THE VBX, AND 9% WERE OTHER. A TOTAL OF 78 PATIENTS HAD CONCOMITANT PLACEMENT OR PREVIOUS PLACEMENT OF AN EVAR DEVICE FOR ANEURYSMAL AORTIC DISEASE. BILATERAL IBES WERE PLACED FOR BILATERAL IIAA IN 15 PATIENTS. THERE WERE SEVEN PATIENTS THAT HAD ENDOLEAK ASSOCIATED WITH IBE DEVICE AND FIVE PATIENTS REQUIRED INTERVENTION. ENDOLEAKS CONSISTED OF TWO TYPE IB REQUIRING DISTAL EXTENSION, ONE TYPE II REQUIRING COIL EMBOLIZATION, AND TWO TYPE III REQUIRING RELINING OF INTERNAL ILIAC STENT. PATENCY OF INTERNAL ILIAC LIMB WAS OBSERVED IN 83 PATIENTS. Manufacturer Narrative: PATEL, R., FATIMA, J., KANAMORI, L., ODERICH, G., SAQIB, N., TREFFALLS, R., LIMA, G., MENDES, B. (2025) A MULTI-INSTITUTIONAL COLLABORATION EVALUATING THE USE OF ILIAC BRANCH ENDOPROSTHESIS TO REPAIR INTERNAL ILIAC ARTERY ANEURYSMS. JOURNAL OF VASCULAR SURGERY, 81(4), E43. A.2. AGE: THE AVERAGE AGE OF THE SUBJECTS IS LISTED AS 73.8 YEARS OLD, THEREFORE AGE IS LISTED AS 74 YEARS. A.3A SEX: 92% OF THE SUBJECTS ARE MALE, THEREFORE SEX IS GIVEN AS MALE. D.1. BRAND NAME LISTED AS NOT OTHERWISE SPECIFIED: THE ABSTRACT LISTS ONLY THAT AN "AORTIC DEVICE" WAS INFECTED AND REQUIRED EXPLANTATION, LEADING TO DEATH. SINCE IT IS UNKNOWN WHAT BRAND OF AORTIC DEVICE WAS USED, WE LIST "NOT OTHERWISE SPECIFIED" E.1. DR. RAJ PATEL IS LISTED AS THE FIRST AUTHOR ON THE PAPER, AND SO IS LISTED HERE AS INITIAL REPORTER. HOWEVER, NO CONTACT INFORMATION WAS PROVIDED. MULTIPLE ATTEMPTS WERE MADE TO CONTACT ANOTHER AUTHOR OF THIS ABSTRACT WITH NO SUCCESS. THE CASE WILL BE CLOSED WITH AVAILABLE INFORMATION. H.6. TYPE OF INVESTIGATION CODE B22: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE CONDUCTED BECAUSE THE SERIAL/LOT NUMBER REMAINS UNKNOWN. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22464726&pc=NPT | 2015691-2025-05589 | 2025-06-17 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-09 04:00 | NPT | COMMANDER DELIVERY SYSTEM | Difficult or Delayed Positioning; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; High Blood Pressure/ Hypertension; Cardiac Tamponade; Cardiac Perforation | P140031 |  | 1 | Event Description: DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE PERFORMED BY OUR EDWARDS AFFILIATES IN POLAND USING A 26MM SAPIEN 3 ULTRA VALVE VIA TRANSFEMORAL ACCESS, THE DELIVERY SYSTEM WAS UNABLE TO CROSS THE NATIVE VALVE DESPITE TWO PRIOR BALLOON AORTIC VALVULOPLASTY'S (BAVS). THE DELIVERY SYSTEM AND VALVE WERE POSITIONED AT A HIGHLY HORIZONTAL ANGLE RELATIVE TO THE ANNULUS, AND THE GUIDEWIRE APPEARED TO BE IN AN UNUSUAL POSITION. A THIRD BAV WAS PERFORMED THROUGH CONTRALATERAL FEMORAL ACCESS, WHICH ENABLED SUCCESSFUL VALVE IMPLANTATION. HOWEVER, IMMEDIATELY FOLLOWING DEPLOYMENT, THE PATIENT EXPERIENCED A SUDDEN RISE IN PRESSURE FOLLOWED BY CARDIAC ARREST. THE PATIENT WAS INTUBATED, CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED, AND CATECHOLAMINES WERE ADMINISTERED. THE CARDIAC ARREST WAS REPORTEDLY CAUSED BY AN INJURY TO THE LEFT VENTRICLE, LOCATED BETWEEN THE LEFT ANTERIOR DESCENDING AND CIRCUMFLEX ARTERIES. MULTIPLE SOURCES OF BLEEDING WERE IDENTIFIED, AND CARDIAC TAMPONADE WAS OBSERVED. THE PROCEDURE WAS CONVERTED TO OPEN-CHEST SURGERY, DURING WHICH THE PERICARDIAL SAC WAS DRAINED AND BLEEDING SITES WERE ADDRESSED. THE PATIENT WAS PLACED ON EXTRACORPOREAL CIRCULATION BUT ULTIMATELY PASSED AWAY. THE DELIVERY SYSTEM WAS DISCARDED AT THE HOSPITAL. ACCORDING TO THE MEDICAL TEAM, THE ROOT CAUSE OF THE EVENT WAS LIKELY IMPROPER GUIDEWIRE POSITIONING, WHICH CONTRIBUTED TO THE DIFFICULTY IN CROSSING THE VALVE AND ULTIMATELY LED TO THE VENTRICULAR INJURY. IT IS SUSPECTED THAT THE COMPROMISED VENTRICULAR WALL MAY HAVE RUPTURED DUE TO THE ABRUPT INCREASE IN PRESSURE. Manufacturer Narrative: CORRECTION TO H6; IMPACT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSION. THE EVENT REPORTED IS ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. IMAGERY PROVIDED SHOWED THAT CALCIFICATION WAS PRESENT AT THE NATIVE ANNULUS, AND THE GUIDEWIRE APPEARED DIRECTED TOWARD THE VENTRICULAR WALL. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, DIFFICULTY OR INABILITY IN TRACKING THE DELIVERY SYSTEM TO THE LANDING ZONE, INCLUDING TRAVERSING THE AORTIC ARCH AND CROSSING THE NATIVE ANNULUS/BIOPROSTHESIS HAS NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE OF THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PATIENT FACTORS, POOR GUIDEWIRE MANAGEMENT, DAMAGED GUIDEWIRE, AND/OR EXCESSIVE MANIPULATION OF THE FLEX WHEEL. THE COMPLAINT FOR CROSSING DIFFICULTY HAS BEEN EMPIRICALLY CONFIRMED BASED ON THE INFORMATION AVAILABLE TO DATE. ANY UPDATES OR ADDITIONAL FINDINGS WILL BE SUBMITTED IN ACCORDANCE WITH FDA REPORTING REQUIREMENTS. AVAILABLE INFORMATION SUGGESTS PROCEDURAL FACTORS (GUIDEWIRE MANAGEMENT) AND/OR PATIENT FACTORS (CALCIFICATION) LIKELY CONTRIBUTED TO THE EVENT. REPORTED INFORMATION INDICATED THAT "AS PER MEDICAL OPINION THE PERCEIVED ROOT CAUSE OF THE EVENT WAS THAT THE GUIDEWIRE WAS NOT PROPERLY POSITIONED, THIS LED TO THE CROSSING DIFFICULTIES AND FINALLY TO THE VENTRICLE INJURY". IN ADDITION, 3MENSIO REVEALED THE PRESENCE OF CALCIFICATION IN PATIENT'S NATIVE ANNULUS. PROPER USE OF THE GUIDEWIRE IS IMPORTANT AS IT SERVES TO GUIDE THE DELIVERY SYSTEM DURING TRACKING AND CROSSING OF THE ANATOMY. AS SUCH, POOR GUIDEWIRE POSITIONING AND/OR LOSS OF GUIDEWIRE PLACEMENT MAY CAUSE THE DELIVERY SYSTEM TO INTERACT WITH PATIENT ANATOMY, LEADING TO DIFFICULTY TRACKING/CROSSING. FURTHERMORE, PATIENT FACTORS (I.E. CALCIFICATION, TORTUOSITY) MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERE WITH PASSAGE OF THE DELIVERY SYSTEM AND CAUSES DIFFICULTY DURING TRACKING/CROSSING. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: THE INVESTIGATION IS ONGOING. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR VENTRICULAR PERFORATION DURING A TVR PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION THAT A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THAT IMPROPER GUIDEWIRE POSITION CONTRIBUTED TO THE DIFFICULTY IN CROSSING AND VENTRICULAR INJURY. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22464997&pc=MHX | 9610816-2025-000581 | 2025-05-31 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-07-09 04:00 | MHX | INTELLIVUE MULTI-MEASUREMENT MODULE X3 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K171801 |  | 1 | Event Description: IT WAS REPORTED FOLLOWING A PATIENT INCIDENT, CLINICAL LEADERSHIP REQUESTED ASSISTANCE RETRIEVING AND REVIEWING THE CLINICAL AUDIT LOGS. THERE WAS A REQUEST FOR AN ON-SITE ENGINEER TO CONFIRM THE DEVICE WAS WORKING AS EXPECTED. IT WAS INDICATED THE PATIENT PASSED AWAY A FEW DAYS LATER BUT BIOMED WAS NOT SURE IF THE DEATH WAS RELATED TO THE INCIDENT. NO FURTHER DETAILS WERE PROVIDED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. Manufacturer Narrative: THE FOLLOWING FUNCTIONAL/DIAGNOSTIC TESTING AND COMMUNICATIONS WERE PERFORMED; THE PHILIPS REMOTE CLINICAL APPLICATION SPECIALIST (CAS) SPOKE WITH THE BIOMED WHO CONFIRMED THAT THE X3 MONITOR IS CURRENTLY IN USE AND THE DEVICE IS PERFORMING AS EXPECTED. THE CAS REQUESTED AN ONSITE VISIT TO RETRIEVE THE CLINICAL AUDIT, TECHNICAL LOGS, AND CONFIRM THE MONITOR IS PERFORMING AS EXPECTED. A PHILIPS TECHNICAL CONSULTANT WENT TO THE CUSTOMER'S SITE AND PULLED LOGS. THE WORK ORDER INDICATES THE SYSTEM MEETS SPECIFICATION FOR THE PERFORMED SERVICE AND IS RETURNED TO USE. GOOD FAITH EFFORTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, NO RESPONSE WAS RECEIVED FROM THE CUSTOMER CONTACT. ALTHOUGH DEVICE LOGS WERE AVAILABLE, NO DETAILS ABOUT THE PATIENT INCIDENT OR DEATH WERE PROVIDED AND NO DEVICE PROBLEM WAS ALLEGED; THEREFORE, THERE IS INSUFFICIENT INFORMATION FOR AN INVESTIGATION. BASED ON THE INFORMATION AVAILABLE THE CAUSE OF THE REPORTED PROBLEM WAS UNABLE TO BE DETERMINED. THE DEVICE REMAINS IN USE AT THE CUSTOMER'S SITE. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. Manufacturer Narrative: BASED ON THE LIMITED INFORMATION RECEIVED, IT CANNOT BE CONFIRMED WHETHER A DEVICE MALFUNCTION OCCURRED OR WHETHER THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22465063&pc=DSQ | 2916596-2025-04291 | 2025-05-22 04:00 | Death | THORATEC CORPORATION | 2025-07-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN HOSPICE. IT WAS SAID THE VENTRICULAR ASSIST DEVICE WAS DECOMMISSIONED AND FUNCTIONED AS INTENDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22465124&pc=LPB | 2029046-2025-02254 | 2025-06-12 04:00 | Death | BIOSENSE WEBSTER INC | 2025-07-09 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest; Unspecified Mental, Emotional or Behavioural Problem | P030031/S078 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION (AFIB) CARDIAC ABLATION, AND THE PATIENT EXPERIENCED COMA, MALIGNANT ARRHYTHMIA, AND DEATH. THE ATRIAL FIBRILLATION ABLATION SURGERY PROCEEDED NORMALLY, AND THE EQUIPMENT DID NOT EXPERIENCE ANY RELATED MALFUNCTIONS. HOWEVER, THE PATIENT FELL INTO A COMA WITH NO CONSCIOUSNESS, AND WITH PUPILS DILATED. SUBSEQUENTLY, MALIGNANT ARRHYTHMIA OCCURRED, AND IMMEDIATE MEASURES SUCH AS ELECTRICAL CARDIOVERSION AND MANUAL COMPRESSION WERE TAKEN FOR RESCUE. UNFORTUNATELY, THE RESCUE EFFORTS WERE UNSUCCESSFUL. INTERVENTION INCLUDED VENTILATOR, MANUAL CHEST COMPRESSION, ELECTRICAL DEFIBRILLATION. THE PHYSICIAN HAD NO IDEA ON WHAT THE CAUSE OF THE ADVERSE EVENT WAS BECAUSE THE SITUATION WAS SO URGENT THAT THERE WAS NO TIME FOR FURTHER EXAMINATION. FORCE VISUALIZATION FEATURES USED WERE GRAPH, DASHBOARD, VECTOR AND VISITAG. THE PARAMETERS FOR STABILITY WERE 2.5MM 3S. THERE WERE NO ERRORS REPORTED. Manufacturer Narrative: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION (AFIB) CARDIAC ABLATION, AND THE PATIENT EXPERIENCED COMA, MALIGNANT ARRHYTHMIA, AND DEATH. THE ATRIAL FIBRILLATION ABLATION SURGERY PROCEEDED NORMALLY, AND THE EQUIPMENT DID NOT EXPERIENCE ANY RELATED MALFUNCTIONS. HOWEVER, THE PATIENT FELL INTO A COMA WITH NO CONSCIOUSNESS, AND WITH PUPILS DILATED. SUBSEQUENTLY, MALIGNANT ARRHYTHMIA OCCURRED, AND IMMEDIATE MEASURES SUCH AS ELECTRICAL CARDIOVERSION AND MANUAL COMPRESSION WERE TAKEN FOR RESCUE. UNFORTUNATELY, THE RESCUE EFFORTS WERE UNSUCCESSFUL. INTERVENTION INCLUDED VENTILATOR, MANUAL CHEST COMPRESSION, ELECTRICAL DEFIBRILLATION. THE PHYSICIAN HAD NO IDEA ON WHAT THE CAUSE OF THE ADVERSE EVENT WAS BECAUSE THE SITUATION WAS SO URGENT THAT THERE WAS NO TIME FOR FURTHER EXAMINATION. FORCE VISUALIZATION FEATURES USED WERE GRAPH, DASHBOARD, VECTOR AND VISITAG. THE PARAMETERS FOR STABILITY WERE 2.5MM 3S. THERE WERE NO ERRORS REPORTED. DEVICE INVESTIGATION DETAILS: THE DEVICE WAS RETURNED TO JOHNSON & JOHNSON MEDTECH (J&J) FOR EVALUATION. A VISUAL INSPECTION AND REVISION OF ALL FEATURES WERE PERFORMED FOLLOWING J&J PROCEDURES. VISUAL ANALYSIS REVEALED NO DAMAGE OR ANOMALIES ON THE DEVICE. THE DEVICE FEATURES WERE REVIEWED, AND NO ISSUES WERE OBSERVED DURING THE PRODUCT INVESTIGATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER LOT 31581089L AND NO INTERNAL ACTION RELATED TO THE COMPLAINT WAS FOUND DURING THE REVIEW. NO MALFUNCTION WAS OBSERVED DURING THE PRODUCT ANALYSIS. OTHER ISSUES OR CIRCUMSTANCES MAY HAVE OCCURRED DURING THE USAGE OF THE DEVICE THAT COMPROMISED ITS PERFORMANCE. THE ROOT CAUSE OF THE ADVERSE EVENT REMAINS UNKNOWN. THE INSTRUCTION FOR USE (IFU) CONTAINS THE FOLLOWING WARNING AND PRECAUTIONS: WHEN USING THE CATHETER WITH CONVENTIONAL SYSTEMS (USING FLUOROSCOPY TO DETERMINE CATHETER TIP LOCATION), OR WITH THE CARTO¿ 3 SYSTEM, CAREFUL CATHETER MANIPULATION MUST BE PERFORMED IN ORDER TO AVOID CARDIAC DAMAGE, PERFORATION, OR TAMPONADE. DO NOT USE EXCESSIVE FORCE TO ADVANCE OR WITHDRAW THE CATHETER WHEN RESISTANCE IS ENCOUNTERED. THE FIRMNESS OF THE BRAIDED TIP DICTATES THAT CARE MUST BE TAKEN TO PREVENT PERFORATION OF THE HEART. AS PART OF JOHNSON & JOHNSON MEDTECH'S QUALITY PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. CORRECTION TO THE INITIAL REPORT: G 4. PMA/ 510(K) HAS BEEN UPDATED FROM P030031/S053 TO P030031/S078. THE ¿SUSPECTED MEDICAL DEVICE¿ REPORTED IN SECTION D OF THIS REPORT IS NOT MARKETED IN USA OR APPROVED BY THE FDA. HOWEVER, IT IS BEING REPORTED AS BIOSENSE WEBSTER CONSIDERS THIS AS A SIMILAR DEVICE TO THERMOCOOL® SMART TOUCH¿ BI-DIRECTIONAL NAVIGATION CATHETER APPROVED UNDER PMA # P030031/S078. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: THE BIOSENSE WEBSTER, INC. PRODUCT ANALYSIS LAB RECEIVED THE DEVICE FOR EVALUATION. THE ANALYSIS HAS BEGUN BUT IS NOT COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22465296&pc=KRH | 2025587-2025-04879 | 2025-06-04 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-09 04:00 | KRH | CG FUTURE BAND | Insufficient Information | Insufficient Information | K061127 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WHO HAD BEEN TREATED FOR MITRAL VALVE DISEASE WITH THE IMPLANTATION OF A CG FUTURE ANNULOPLASTY HEART RING IN THE MITRAL POSITION, PASSED AWAY. THE DEVICE WAS IMPLANTED ON (B)(6) 2025, AND THE DEATH OCCURRED ON (B)(6) 2025. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE PRODUCT COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22465570&pc=MHX | 1218950-2025-000385 | 2025-06-12 04:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-07-09 04:00 | MHX | PATIENT INFORMATION CENTER IX | Defective Alarm | Insufficient Information | K183387 |  | 1 | Event Description: IT WAS REPORTED THE ALARMS FOR APNEA, BRADYCARDIA, AND ASYSTOLE WERE NOT GENERATED. THE ICU PATIENT WAS FOUND PASSED AWAY IN THEIR BED AND RESUSCITATION ATTEMPTS WERE UNSUCCESSFUL. TESTING OF THE DEVICE ONSITE REVEALED NO TECHNICAL ANOMALIES. IT WAS CLARIFIED NONE OF THE ALARMS, AUDIBLE OR VISUAL, WERE GENERATED AT THE CENTRAL STATION. THE CONDITION OF THE PATIENT PRIOR AND THE CAUSE OF DEATH WERE UNKNOWN, BUT IT WAS INDICATED THE PATIENT WAS ELDERLY. Manufacturer Narrative: A PHILIPS FIELD SERVICE ENGINEER (FSE) WENT ONSITE AND COLLECTED THE LOGS OF THE PIC IX AND LOGS AND CONFIGURATION OF THE MP5. THE FSE DID NOT FIND ANY TECHNICAL DEFECTS WHILE ONSITE. INITIAL REVIEW OF THE LOGS INDICATES MANUAL DEACTIVATION OF THE PULSE ALARMS. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. SECTION E REPORTING INSTITUTION PHONE #: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22433909&pc=CAF | 3003399703-2025-22828 | 2024-11-16 05:00 | Death | AEROGEN LTD | 2025-07-08 04:00 | CAF | AEROGEN SOLO | Adverse Event Without Identified Device or Use Problem | Hypoxia | K070642 |  | 1 | Event Description: ON (B)(6) 2024, AN ADVERSE EVENT OCCURRED DURING A TRANSPLANT INVOLVING THE USE OF A PALL ULTIPOR 55 BREATHING CIRCUIT FILTER WITH AN AEROGEN SOLO NEBULISER. THE PATIENT, A 57-YEAR-OLD MALE, EXPERIENCED PROLONGED RESPIRATORY CHALLENGES DUE TO AN INCORRECT SETUP OF THE CIRCUIT, WHICH LIKELY CAUSED THE FILTER TO CLOG. THE FILTER AEROGEN SOLO NEBULISER WAS INTENDED TO DELIVER INHALED EPOPROSTENOL AFTER THE PATIENT WAS LIBERATED FROM CARDIOPULMONARY BYPASS. THE PATIENT WAS TAKEN OFF BYPASSAND CONNECTED TO A VENTILATOR FOR THE DELIVERY OF INHALED APOPROSTENOL. HOWEVER, WITHIN 2 HOURS, THE PATIENT'S TIDAL VOLUMES DECREASED, INDICATING DIFFICULTY IN VENTILATION AND OXYGENATION, LIKELY DUE TO THE CLOGGED FILTER OBSTRUCTING AIRFLOW AND EPOPROSTENOL DELIVERY FROM THE AEROGEN SOLO NEBULIZER. THIS LED TO AN ANOXIC BRAIN INJURY AND ULITMATELY THE PATIENT'S DEATH, WHICH OCCURED ON (B)(6) 2024. THE ADVERSE EVENT WAS IDENTIFIED AS POTENTIALLY LINKED TO THE SETUP OF THE FILTER WITH THE AEROGEN SOLO NEBULIZER AFTER DISCOVERING THAT THE FILTER ON THE VENTILATOR HAD CLOGGED, WHICH OBSTRUCTED THE DELIVERY OF OXYGEN AND EPOPROSTENOL.Event Description: ON NOVEMBER 16, 2024, AN ADVERSE EVENT OCCURRED DURING A TRANSPLANT INVOLVING THE USE OF A PALL ULTIPOR 55 BREATHING CIRCUIT FILTER WITH AN AEROGEN SOLO NEBULISER. THE PATIENT, A 57-YEAR OLD MALE, EXPERIENCED PROLONGED RESPIRATORY CHALLENGES DUE TO AN INCORRECT SETUP OF THE CIRCUIT, WHICH LIKELY CAUSED THE FILTER TO CLOG. THE FILTER AEROGEN SOLO NEBULISER WAS INTENDED TO DELIVER INHALED EPOPROSTENOL AFTER THE PATIENT WAS LIBERATED FROM CARDIOPULMONARY BYPASS. THE PATIENT WAS TAKEN OFF BYPASS AND CONNECTED TO A VENTILATOR FOR THE DELIVERY OF INHALED APOPROSTENOL. HOWEVER, WITHIN 2 HOURS, THE PATIENT'S TIDAL VOLUMES DECREASED, INDICATING DIFFICULTY IN VENTILATION AND OXYGENATION, LIKELY DUE TO THE CLOGGED FILTER OBSTRUCTING AIRFLOW AND EPOPROSTENOL DELIVERY FROM THE AEROGEN SOLO NEBULIZER. THIS LED TO AN ANOXIC BRAIN INJURY AND ULITMATELY THE PATIENTS DEATH, WHICH OCCURED ON (B)(6) 2024. THE ADVERSE EVENT WAS IDENTIFIED AS POTENTIALLY LINKED TO THE SETUP OF THE FILTER WITH THE AEROGEN SOLO NEBULIZER AFTER DISCOVERING THAT THE FILTER ON THE VENTILATOR HAD CLOGGED, WHICH OBSTRUCTED THE DELIVERY OF OXYGEN AND EPOPROSTENOL. Manufacturer Narrative: " ON NOVEMBER 16, 2024, AN ADVERSE EVENT OCCURRED DURING A TRANSPLANT INVOLVING THE USE OF A PALL ULTIPOR 55 BREATHING CIRCUIT FILTER WITH AN AEROGEN SOLO NEBULISER. THE PATIENT, A 57-YEAR OLD MALE, EXPERIENCED PROLONGED RESPIRATORY CHALLENGES DUE TO AN INCORRECT SETUP OF THE CIRCUIT, WHICH LIKELY CAUSED THE FILTER TO CLOG. THE FILTER AEROGEN SOLO NEBULISER WAS INTENDED TO DELIVER INHALED EPOPROSTENOL AFTER THE PATIENT WAS LIBERATED FROM CARDIOPULMONARY BYPASS. THE PATIENT WAS TAKEN OFF BYPASSAND CONNECTED TO A VENTILATOR FOR THE DELIVERY OF INHALED APOPROSTENOL. HOWEVER, WITHIN 2 HOURS, THE PATIENT'S TIDAL VOLUMES DECREASED, INDICATING DIFFICULTY IN VENTILATION AND OXYGENATION, LIKELY DUE TO THE CLOGGED FILTER OBSTRUCTING AIRFLOW AND EPOPROSTENOL DELIVERY FROM THE AEROGEN SOLO NEBULIZER. THIS LED TO AN ANOXIC BRAIN INJURY AND ULITMATELY THE PATIENTS DEATH, WHICH OCCURED ON (B)(6) 2024. THE ADVERSE EVENT WAS IDENTIFIED AS POTENTIALLY LINKED TO THE SETUP OF THE FILTER WITH THE AEROGEN SOLO NEBULIZER AFTER DISCOVERING THAT THE FILTER ON THE VENTILATOR HAD CLOGGED, WHICH OBSTRUCTED THE DELIVERY OF OXYGEN AND EPOPROSTENOL." THIS IS BEING INVESTIGATED AS A NON-RETURN. THE DEVICE REPORTED WAS NOT RETURNED TO AEROGEN FOR INVESTIGATION AS IT WAS DISCARDED BY THE CUSTOMER AND THERE WAS INSUFFICIENT EVIDENCE/INFORMATION SUPPLIED BY THE CUSTOMER TO DETERMINE A ROOT CAUSE. CLINICAL CONCLUSION BASED ON THE INFORMATION AVAILABLE THE DEATH REPORTED WAS LIKELY ATTRIBUTABLE SOLELY TO USER ERROR. THIS EVENT DID NOT OCCUR, AS CONFIRMED BY THE USER FACILITY, AS A RESULT OF DEVICE [21 CFR 803.3]: - FAILURE; - MALFUNCTION; - IMPROPER OR INADEQUATE DESIGN; - MANUFACTURE; OR - LABELLING. THE USER FACILITY REPRESENTATIVE CONFIRMED THAT "USER ERROR", IN THIS INSTANCE A MISTAKE MADE BY AN ATTENDING HEALTH CARE PROFESSIONAL IN THE CONFIGURATION IN WHICH THE PALL ULTIPOR 55 BREATHING CIRCUIT FILTER WAS PLACED IN RELATION TO THE AEROGEN® SOLO NEBULISER WAS THE CAUSE OF THIS EVENT12. THERE WAS NO PERFORMANCE ISSUE RELATED TO THE AEROGEN® SOLO NEBULISER REPORTED BY THE USER FACILITY. SUMMARY NO PERFORMANCE ISSUE WAS REPORTED WITH THE AEROGEN MEDICAL DEVICE INVOLVED. THE USER FACILITY CONFIRMED THAT AN "INCORRECT" CONFIGURATION OF THE MEDICAL DEVICES BY AN ATTENDING CLINICAL STAFF MEMBER WAS THE CAUSE OF THE EVENT. BASED ON THE INFORMATION AVAILABLE THE FOLLOWING IMDRF ANNEX A, E & F CODES MAY BE CONSIDERED. APPROPRIATE: A24 - ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM. E0726 - HYPOXIA. F22 - UNEXPECTED DETERIORATION. F23 - UNEXPECTED MEDICAL INTERVENTION. F02 - DEATH. THE FOLLOWING QUESTIONS (BOLD ITALICS) WERE PROVIDED TO THE USER FACILITY BY AEROGEN LTD. ON THE XXX. A WRITTEN RESPONSE (NORMAL TEXT) WAS PROVIDED TO QUESTION 1 ONLY. VERBAL RESPONSES (REPRODUCED AS FAITHFULLY AS POSSIBLE - NORMAL TEXT) WERE PROVIDED TO OTHER QUESTIONS IN THE FACE-TO-FACE INTERVIEW BY (B)(6) REPRESENTATIVES L PACKER ON THE 27TH JUNE 2025. DUE TO ONGOING LITIGATION (B)(6) DECLINE TO RESPOND TO SOME QUESTIONS ASKED. THESE ARE MARKED [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] BELOW. 1. COULD YOU DESCRIBE THE FORM OF RESPIRATORY SUPPORT (OR MECHANICAL VENTILATION) THE PATIENT WAS RECEIVING AT THE TIME, INCLUDING CONFIGURATION (ACTIVE OR PASSIVE HUMIDIFICATION, HUMIDIFIER USED, DUAL/SINGLE LIMB ETC.) AND INTERFACE TO THE PATIENT (E.G., ETT VS. TRACHEOSTOMY TUBE). MECHANICAL VENTILATION WAS ACHIEVED USING AN GE DATEX-OHMEDA AISYS CS2 ANESTHESIA SYSTEM, CONNECTED VIA A SINGLE LIMB ANESTHESIA CIRCUIT TO THE ETT OF THE PATIENT. THE PATIENT-SIDE HME FILTER WAS REMOVED FROM THE CIRCUIT. THERE WERE VENTILATOR-SIDE PLEATED HYDROPHOBIC MEMBRANE FILTERS (PALL ULTIPOR 55) PLACED BOTH IN THE INSPIRATORY AND EXPIRATORY PORT. THE AEROGEN SOLO NEBULIZER WAS PLACED IN THE INSPIRATORY PORT. THE INCORRECT SETUP WAS SUCH THAT THE AEROGEN SOLO NEBULIZER WAS PLACED IN THE INSPIRATORY PORT OF THE VENTILATOR FIRST, FOLLOWED BY THE PALL ULTIPOR FILTER, FOLLOWED BY THE ANESTHESIA CIRCUIT, LIKELY CAUSING THE HYDROPHOBIC FILTER TO CLOG WITH AERSOSOLIZED EPOPROSTANOL. 2. WAS THE PALL ULTIPOR 55 FILTER BEING EMPLOYED AS A FILTER (ON THE MACHINE SIDE PROTECTING THE EXPIRATORY PORT) OR AS A HME (AND POSITIONED PROXIMAL TO THE PATIENT)? A. IF THE PALL ULTIPOR 55 FILTER WAS PLACED ON THE EXPIRATORY LIMB, WAS ANOTHER HME EMPLOYED OR WAS ACTIVE HUMIDIFICATION USED? PALL ULTIPOR 55 BREATHING CIRCUIT FILTER WAS PLACED AT BOTH THE INSPIRATORY AND EXPIRATORY POT OF THE VENTILATOR. THIS IS NORMAL PRACTICE AT THIS INSTITUTION. 3. COULD YOU DESCRIBE WHERE THE AEROGEN® SOLO NEBULISER WAS PLACED IN THIS CIRCUIT IN RELATION TO THE PALL ULTIPOR 55 FILTER? N/A - CONSIDERED ADDRESSED IN PREVIOUS ANSWERS TO Q1 & Q2. 4. WAS THE PALL ULTIPOR 55 FILTER PLACED [?]UPSTREAM' OF THE AEROGEN® SOLO NEBULISER, I.E., PROXIMAL TO THE PATIENT WITH RESPECT TO THE POSITION OF NEBULISER POSITION? OR WAS THE AEROGEN® SOLO NEBULISER BETWEEN THE PALL ULTIPOR 55 FILTER AND THE PATIENT? YES. THE FILTER WAS PLACED BETWEEN THE NEBULISER AND THE PATIENT. 5. WHAT WAS THE INDICATION FOR AEROSOL DRUG DELIVERY USING THE AEROGEN® SOLO NEBULISER? REDUCTION OF PULMONARY VASCULAR RESISTANCE (PVR), RELIEF OF RIGHT VENTRICULAR AFTERLOAD. INSTITUTIONAL PRACTICE FOR PREVENTION OF PRIMARY GRAFT DYSFUNCTION (REFERENCED TO RCT CONDUCED AT (B)(6) IN 2022 - SEE BIBLIOGRAPHY). 6. WHEN WAS THE AEROGEN SOLO FIRST USED FOLLOWING THE PATIENT'S SEPARATION FROM BYPASS? 2 HOURS. 7. HOW LONG WAS THE NEBULISER IN USE DURING THE POST-OPERATIVE PERIOD? [CONSIDERED ADDRESSED BY RESPONSE TO Q6] 8. WERE THE STAFF INVOLVED FAMILIAR WITH AEROGEN'S IFU AND THE CORRECT SETUP OF THE DEVICE? YES, (B)(6) STAFF ARE VERY FAMILIAR WITH AEROGEN DEVICES, THE IFU AND USE AEROGEN NEBULISERS FOR IEPO AS STANDARD PRACTICE. 9. HAS THIS TYPE OF CONFIGURATION (AEROGEN SOLO WITH IN-LINE FILTER) BEEN USED BEFORE AT YOUR FACILITY, AND IF SO, HAVE ANY SIMILAR ISSUES OCCURRED? THIS SPECIFIC CONFIGURATION, AS OCCURRED IN THIS INCIDENT - NO. THE CORRECT CONFIGURATION, WITH A FILTER ON THE INSPIRATORY PORT & EXPIRATORY PORT, WITH AN AEROGEN NEBULISER PROXIMAL TO THE PATIENT IS STANDARD IN (B)(6). PAGE 11 OF 16 RELATED TO THE FILTER OBSTRUCTION: 10. IN THIS SPECIFIC CASE HOW LONG WAS THE PALL ULTIPOR 55 FILTER USED? AND A. AT WHAT FREQUENCY ARE THESE FILTERS TYPICALLY CHECKED ACCORDING TO YOUR INTERNAL SOPS, AND CHANGED/REPLACED? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 11. SPECIFICALLY, WAS THE FILTER SATURATED OR [?]CLOGGED'/'OBSTRUCTED' WITH PARTICULATE LIKE CRYSTALS? YES. THE FILTER WAS NOTED TO BE SATURATED BY THE HCPS. NO - THERE WAS NO PARTICULATE VISIBLE. 12. WAS THERE VISIBLE CONDENSATION OR SATURATION NOTED IN THE FILTER OR CIRCUIT? NO, THE SATURATION WAS [?]NON-OBVIOUS' TO THE ATTENDING CLINICAL TEAM. 13. AT WHAT POINT WAS A CLOGGED OR OBSTRUCTED FILTER SUSPECTED OR CONFIRMED? THE PROFOUND DROP IN TIDAL VOLUME WAS THE FIRST INDICATION TO THE CLINICAL STAFF OF A POTENTIAL ISSUE. IT WAS DISCOVERED SHORTLY AFTER THIS. 14. AT WHAT POINT WAS A CLOGGED OR OBSTRUCTED FILTER SUSPECTED OR CONFIRMED? CONSIDERED ADDRESSED BY RESPONSE TO Q13. RELATED TO THE AE: 15. WAS A PRE-USE CHECK OR INSPECTION OF THE BREATHING CIRCUIT AND FILTER CONDUCTED PRIOR TO INITIATION? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 16. WAS A PRE-USE CHECK OF THE VENTILATOR PERFORMED TO ENSURE THE PATIENT SETTINGS, ALARMS AND MONITORING DEVICES ARE ACCURATELY CONNECTED? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 17. THE REPORT REFERS TO A PROLONGED RESPIRATORY CHALLENGE[?]COULD YOU PROVIDE MORE DETAIL AROUND THIS? HOW LONG DID IT PERSIST AND WHAT SYMPTOMS WERE NOTICED/OBSERVED? AFTER THE DROP IN TIDAL VOLUME THE [?]RESPIRATORY CHALLENGE' OCCURRED FOR BETWEEN 15-20 MINS UNTIL THE SATURATED FILTER WAS DISCOVERED. WHEN THIS WAS REPLACED, AND CONFIGURATION CORRECT, THE PATIENT WAS VENTILATED SUCCESSFULLY. 18. WHAT WERE THE PATIENT'S OXYGENATION LEVELS (SPO2 OR SIMILAR) BEFORE THE AE. [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 19. RELATED TO THE "DECREASED TIDAL VOLUMES" AND "DIFFICULTY IN VENTILATION" WERE ANY OF THE FOLLOWING ALARMS RECORDED: A. HIGH PEEP, OR B. HIGH AIRWAY PRESSURE ALARM, OR C. LOW EXHALED TIDAL VOLUME ALARM, OR D. APNEA OR LOW MINUTE VENTILATION, OR E. FLOW SENSOR OR EXPIRATORY VALVE ERROR. NO - NO ALARMS SOUNDED TO ALERT THE CLINICAL TEAM. THE FIRST INDICATION OF AN ISSUE WAS THE DROP IN TIDAL VOLUME. 20. WERE KEY RESPIRATORY PARAMETERS SUCH AS TIDAL VOLUME, PEAK PRESSURES, RESISTANCE, OR ETCO CONTINUOUSLY MONITORED? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 21. WHAT CLINICAL MONITORING WAS IN PLACE (E.G., SPO[?], ETCO[?], FIO[?] TRENDS) [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 22. WHAT INTERVENTIONS WERE MADE IN RESPONSE TO THE OBSERVED DETERIORATION IN VENTILATION AND OXYGENATION? PAGE 12 OF 16 A BRONCHOSCOPY WAS ATTEMPTED TO DETERMINE IF AIRWAY OBSTRUCTION WAS PRESENT. UPON DETECTION THE FILTER WAS REMOVED - AT WHICH POINT THE PATIENT COULD BE VENTILATED SUCCESSFULLY. 23. WHEN WAS THE ONSET OF ANOXIC BRAIN INJURY FIRST SUSPECTED OR CONFIRMED? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 24. WERE OTHER POSSIBLE CAUSES OF HYPOXIA (E.G., AIRWAY OBSTRUCTION, PULMONARY EMBOLISM, BRONCHOSPASM) CONSIDERED OR RULED OUT? CONSIDERED ADDRESSED BY RESPONSE TO Q22 25. THE AE REPORT STATES THE AE WAS "LIKELY DUE TO A CLOGGED FILTER OBSTRUCTING AIRFLOW" AND REFERS TO A 2 HOUR TIME PERIOD OF WHICH PERIOD "DECREASED TIDAL VOLUME, DIFFICULTY IN VENTILATION AND OXYGENATION OCCURRED". WAS THE "CLOGGED" FILTER NOTED DURING THIS PERIOD OR AFTER? CONSIDERED ADDRESSED BY PREVIOUS ANSWERS. FINALLY: 26. HAS ANY CHANGE BEEN MADE TO LOCAL PRACTICE OR PROTOCOLS AS A RESULT OF THIS EVENT? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 27. IS THE AEROGEN DEVICE AVAILABLE TO BE SENT BACK FOR INVESTIGATION. NO - THE AEROGEN DEVICE WAS NOT RETAINED. CLINICAL STAFF REPORTED NO ISSUE WITH THE AEROGEN DEVICE, AND IT WAS NOT RETAINED. THE PALL ULTIPOR 55 BREATHING CIRCUIT FILTER WAS RETAINED. Manufacturer Narrative: COMMENTS FROM CALL WITH (B)(6): THEY STATED CLEARLY THAT THEIR STAFF PLACED THE AEROGEN DEVICE IN THE IMV CIRCUIT INCORRECTLY - WITH THE AEROGEN DEVICE DISTAL TO A FILTER (I.E., THE PATHWAY FOR THE AEROSOL TO THE PATIENT WAS OBSTRUCTED BY THE FILTER). THIS THEY STATED THIS WAS DUE TO AN ERROR BY A MEMBER OF THEIR STAFF. THE AEROGEN DEVICE WAS NOT RETAINED AS THEY FOUND NO FAULT WITH IT AND THEY DO NOT ATTRIBUTE THE ADVERSE EVENT TO THE AEROGEN DEVICE. FROM THEIR PERSPECTIVE THE USE OF THE AEROGEN DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE EVENT. QUESTIONS: A WRITTEN RESPONSE (NORMAL TEXT) WAS PROVIDED TO QUESTION 1 ONLY. VERBAL RESPONSES (REPRODUCED AS FAITHFULLY AS POSSIBLE - NORMAL TEXT) WERE PROVIDED TO OTHER QUESTIONS IN THE FACE-TO-FACE INTERVIEW BY (B)(6) REPRESENTATIVES L PACKER ON THE (B)(6) 2025. DUE TO ONGOING LITIGATION (B)(6) DECLINE TO RESPOND TO SOME QUESTIONS ASKED. THESE ARE MARKED [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] BELOW. COULD YOU DESCRIBE THE FORM OF RESPIRATORY SUPPORT (OR MECHANICAL VENTILATION) THE PATIENT WAS RECEIVING AT THE TIME, INCLUDING CONFIGURATION (ACTIVE OR PASSIVE HUMIDIFICATION, HUMIDIFIER USED, DUAL/SINGLE LIMB ETC.) AND INTERFACE TO THE PATIENT (E.G., ETT VS. TRACHEOSTOMY TUBE). MECHANICAL VENTILATION WAS ACHIEVED USING AN GE DATEX-OHMEDA AISYS CS2 ANESTHESIA SYSTEM, CONNECTED VIA A SINGLE LIMB ANESTHESIA CIRCUIT TO THE ETT OF THE PATIENT. THE PATIENT-SIDE HME FILTER WAS REMOVED FROM THE CIRCUIT. THERE WERE VENTILATOR-SIDE PLEATED HYDROPHOBIC MEMBRANE FILTERS (PALL ULTIPOR 55) PLACED BOTH IN THE INSPIRATORY AND EXPIRATORY PORT. THE AEROGEN SOLO NEBULIZER WAS PLACED IN THE INSPIRATORY PORT. THE INCORRECT SETUP WAS SUCH THAT THE AEROGEN SOLO NEBULIZER WAS PLACED IN THE INSPIRATORY PORT OF THE VENTILATOR FIRST, FOLLOWED BY THE PALL ULTIPOR FILTER, FOLLOWED BY THE ANESTHESIA CIRCUIT, LIKELY CAUSING THE HYDROPHOBIC FILTER TO CLOG WITH AERSOSOLIZED EPOPROSTANOL. WAS THE PALL ULTIPOR 55 FILTER BEING EMPLOYED AS A FILTER (ON THE MACHINE SIDE PROTECTING THE EXPIRATORY PORT) OR AS A HME (AND POSITIONED PROXIMAL TO THE PATIENT)? A. IF THE PALL ULTIPOR 55 FILTER WAS PLACED ON THE EXPIRATORY LIMB, WAS ANOTHER HME EMPLOYED OR WAS ACTIVE HUMIDIFICATION USED? PALL ULTIPOR 55 BREATHING CIRCUIT FILTER WAS PLACED AT BOTH THE INSPIRATORY AND EXPIRATORY POT OF THE VENTILATOR. THIS IS NORMAL PRACTICE AT THIS INSTITUTION. COULD YOU DESCRIBE WHERE THE AEROGEN® SOLO NEBULISER WAS PLACED IN THIS CIRCUIT IN RELATION TO THE PALL ULTIPOR 55 FILTER? N/A - CONSIDERED ADDRESSED IN PREVIOUS ANSWERS TO Q1 & Q2. WAS THE PALL ULTIPOR 55 FILTER PLACED UPSTREAM' OF THE AEROGEN® SOLO NEBULISER, I.E., PROXIMAL TO THE PATIENT WITH RESPECT TO THE POSITION OF NEBULISER POSITION? OR WAS THE AEROGEN® SOLO NEBULISER BETWEEN THE PALL ULTIPOR 55 FILTER AND THE PATIENT? YES. THE FILTER WAS PLACED BETWEEN THE NEBULISER AND THE PATIENT. WHAT WAS THE INDICATION FOR AEROSOL DRUG DELIVERY USING THE AEROGEN® SOLO NEBULISER? REDUCTION OF PULMONARY VASCULAR RESISTANCE (PVR), RELIEF OF RIGHT VENTRICULAR AFTERLOAD. INSTITUTIONAL PRACTICE FOR PREVENTION OF PRIMARY GRAFT DYSFUNCTION (REFERENCED TO RCT CONDUCED AT (B)(6) IN 2022 - SEE BIBLIOGRAPHY). WHEN WAS THE AEROGEN SOLO FIRST USED FOLLOWING THE PATIENT'S SEPARATION FROM BYPASS? 2 HOURS. HOW LONG WAS THE NEBULISER IN USE DURING THE POST-OPERATIVE PERIOD? [CONSIDERED ADDRESSED BY RESPONSE TO Q6] WERE THE STAFF INVOLVED FAMILIAR WITH AEROGEN'S IFU AND THE CORRECT SETUP OF THE DEVICE? YES, (B)(6) STAFF ARE VERY FAMILIAR WITH AEROGEN DEVICES, THE IFU AND USE AEROGEN NEBULISERS FOR IEPO AS STANDARD PRACTICE. HAS THIS TYPE OF CONFIGURATION (AEROGEN SOLO WITH IN-LINE FILTER) BEEN USED BEFORE AT YOUR FACILITY, AND IF SO, HAVE ANY SIMILAR ISSUES OCCURRED? THIS SPECIFIC CONFIGURATION, AS OCCURRED IN THIS INCIDENT - NO. THE CORRECT CONFIGURATION, WITH A FILTER ON THE INSPIRATORY PORT & EXPIRATORY PORT, WITH AN AEROGEN NEBULISER PROXIMAL TO THE PATIENT IS STANDARD IN (B)(6). RELATED TO THE FILTER OBSTRUCTION: IN THIS SPECIFIC CASE HOW LONG WAS THE PALL ULTIPOR 55 FILTER USED? AND A. AT WHAT FREQUENCY ARE THESE FILTERS TYPICALLY CHECKED ACCORDING TO YOUR INTERNAL SOPS, AND CHANGED/REPLACED? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] SPECIFICALLY, WAS THE FILTER SATURATED OR CLOGGED'/'OBSTRUCTED' WITH PARTICULATE LIKE CRYSTALS? YES. THE FILTER WAS NOTED TO BE SATURATED BY THE HCPS. NO - THERE WAS NO PARTICULATE VISIBLE. WAS THERE VISIBLE CONDENSATION OR SATURATION NOTED IN THE FILTER OR CIRCUIT? NO, THE SATURATION WAS NON-OBVIOUS' TO THE ATTENDING CLINICAL TEAM. AT WHAT POINT WAS A CLOGGED OR OBSTRUCTED FILTER SUSPECTED OR CONFIRMED? THE PROFOUND DROP IN TIDAL VOLUME WAS THE FIRST INDICATION TO THE CLINICAL STAFF OF A POTENTIAL ISSUE. IT WAS DISCOVERED SHORTLY AFTER THIS. AT WHAT POINT WAS A CLOGGED OR OBSTRUCTED FILTER SUSPECTED OR CONFIRMED? CONSIDERED ADDRESSED BY RESPONSE TO Q13. RELATED TO THE AE: WAS A PRE-USE CHECK OR INSPECTION OF THE BREATHING CIRCUIT AND FILTER CONDUCTED PRIOR TO INITIATION? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] WAS A PRE-USE CHECK OF THE VENTILATOR PERFORMED TO ENSURE THE PATIENT SETTINGS, ALARMS AND MONITORING DEVICES ARE ACCURATELY CONNECTED? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] THE REPORT REFERS TO A PROLONGED RESPIRATORY CHALLENGE COULD YOU PROVIDE MORE DETAIL AROUND THIS? HOW LONG DID IT PERSIST AND WHAT SYMPTOMS WERE NOTICED/OBSERVED? AFTER THE DROP IN TIDAL VOLUME THE RESPIRATORY CHALLENGE OCCURRED FOR BETWEEN 15-20 MINS UNTIL THE SATURATED FILTER WAS DISCOVERED. WHEN THIS WAS REPLACED, AND CONFIGURATION CORRECT, THE PATIENT WAS VENTILATED SUCCESSFULLY. WHAT WERE THE PATIENT'S OXYGENATION LEVELS (SPO2 OR SIMILAR) BEFORE THE AE. [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] RELATED TO THE "DECREASED TIDAL VOLUMES" AND "DIFFICULTY IN VENTILATION" WERE ANY OF THE FOLLOWING ALARMS RECORDED: A. HIGH PEEP, OR B. HIGH AIRWAY PRESSURE ALARM, OR C. LOW EXHALED TIDAL VOLUME ALARM, OR D. APNEA OR LOW MINUTE VENTILATION, OR E. FLOW SENSOR OR EXPIRATORY VALVE ERROR. NO - NO ALARMS SOUNDED TO ALERT THE CLINICAL TEAM. THE FIRST INDICATION OF AN ISSUE WAS THE DROP IN TIDAL VOLUME. WERE KEY RESPIRATORY PARAMETERS SUCH AS TIDAL VOLUME, PEAK PRESSURES, RESISTANCE, OR ETCO CONTINUOUSLY MONITORED? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] WHAT CLINICAL MONITORING WAS IN PLACE (E.G., SPO, ETCO, FIO TRENDS)? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] WHAT INTERVENTIONS WERE MADE IN RESPONSE TO THE OBSERVED DETERIORATION IN VENTILATION AND OXYGENATION? A BRONCHOSCOPY WAS ATTEMPTED TO DETERMINE IF AIRWAY OBSTRUCTION WAS PRESENT. UPON DETECTION THE FILTER WAS REMOVED - AT WHICH POINT THE PATIENT COULD BE VENTILATED SUCCESSFULLY. WHEN WAS THE ONSET OF ANOXIC BRAIN INJURY FIRST SUSPECTED OR CONFIRMED? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] WERE OTHER POSSIBLE CAUSES OF HYPOXIA (E.G., AIRWAY OBSTRUCTION, PULMONARY EMBOLISM, BRONCHOSPASM) CONSIDERED OR RULED OUT? CONSIDERED ADDRESSED BY RESPONSE TO Q22 THE AE REPORT STATES THE AE WAS "LIKELY DUE TO A CLOGGED FILTER OBSTRUCTING AIRFLOW" AND REFERS TO A 2 HOUR TIME PERIOD OF WHICH PERIOD "DECREASED TIDAL VOLUME, DIFFICULTY IN VENTILATION AND OXYGENATION OCCURRED". WAS THE "CLOGGED" FILTER NOTED DURING THIS PERIOD OR AFTER? CONSIDERED ADDRESSED BY PREVIOUS ANSWERS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22434521&pc=DSP | 2249723-2025-0002921 | 2025-06-23 04:00 | Death | DATASCOPE CORP. | 2025-07-08 04:00 | DSP | CARDIOSAVE HYBRID, 3.1 EDITION | Incorrect, Inadequate or Imprecise Result or Readings | Myocardial Infarction | K112372 |  | 1 | Event Description: IT WAS REPORTED THAT CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) PRESSURE WAVEFORM AND BLOOD PRESSURE VALUES WERE NO LONGER DISPLAYED DURING CLINICAL USE WITH SENSOR BALLOON. THE PATIENT DIED ON (B)(6) 2025. Manufacturer Narrative: A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22434885&pc=DXY | 9614453-2025-02601 | 2025-06-30 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-07-08 04:00 | DXY | THERA-I SR | Appropriate Term/Code Not Available | Atrial Fibrillation | P890003 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S SPOUSE; THE PATIENT EXPERIENCED AN ATRIAL FIBRILLATION (AF) EPISODE AND THE IMPLANTABLE PULSE GENERATOR (IPG) DID NOT "KICK IN". THE PATIENT IS DECEASED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22435487&pc=KNS | 2124215-2025-44063 | 2024-03-25 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-08 04:00 | KNS | REZUM | Adverse Event Without Identified Device or Use Problem | Urinary Retention; Urinary Tract Infection; Hematuria; Unspecified Tissue Injury |  |  | 1 | Event Description: IT WAS REPORTED IN THE JOHN WILEY & SONS AUSTRALIA THAT A RETROSPECTIVE STUDY WAS CONDUCTED TO EVALUATE THE SUCCESS RATE OF TRIAL WITHOUT CATHETER (TWOC) AND COMPARED PRE- AND POST-OPERATIVE (AT 6 MONTHS) UROFLOWMETRY RESULTS, INTERNATIONAL PROSTATE SYMPTOMS SCORE (IPSS), IPSS QUALITY OF LIFE SCORE (IPSS QOL), AND PROSTATE VOLUME. A TOTAL OF 63 PATIENTS WHO UNDERWENT REZUM PROCEDURES PERFORMED WERE ENROLLED (44 AND 19 PATIENTS HAD REFRACTORY URINARY RETENTION AND OBSTRUCTIVE UROPATHY, RESPECTIVELY) FROM JANUARY 2022 TO AUGUST 2022. THE FOLLOWING BOSTON SCIENTIFIC PRODUCTS WERE USED DURING THE PROCEDURES: REZUM DEVICE. ALL SURGICAL PROCEDURES WERE PERFORMED SUCCESSFULLY WITHOUT ANY INTRA-OPERATIVE COMPLICATIONS; HOWEVER, PRE-OPERATIVE PATIENT CONDITION WERE ACUTE RETENTION OF URINE FOR 2 PATIENTS, OBSTRUCTIVE UROPATHY ON 19 PATIENT AND 44 PATIENTS WITH REFRACTORY URINARY RETENTION. REGARDING POSTOPERATIVE COMPLICATIONS, TOTAL OF 11 PATIENTS HAD UNPLANNED READMISSIONS AFTER THE OPERATION, PATIENTS HAD EARLY UNPLANNED READMISSIONS (WITHIN 30 DAYS OF OPERATION) AND SIX PATIENTS HAD LATE UNPLANNED READMISSIONS (AFTER 30 DAYS OF OPERATION). IN THE EARLY UNPLANNED READMISSION GROUP, 2 PATIENTS HAD HEMATURIA AND 3 HAD CATHETER-ASSOCIATED PROBLEMS. ALL 5 WERE TREATED CONSERVATIVELY. IN THE LATE UNPLANNED READMISSION GROUP, 2 PATIENTS HAD HEMATURIA, AND 2 PATIENTS HAD URINARY TRACT INFECTIONS; 1 PATIENT WITH GROSS HEMATURIA REQUIRED BLADDER IRRIGATION. ANOTHER PATIENT DEVELOPED A URINARY TRACT INFECTION AND ACUTE URINARY RETENTION 13 DAYS AFTER HIS SECOND PROCEDURE; HE WAS TREATED WITH ANTIBIOTICS. ONE PATIENT DIED BEFORE HIS FIRST ATTEMPT AT TWOC (TRIAL WITHOUT CATHETER) AS A RESULT OF PNEUMONIA AND ONE PATIENT REFUSED FURTHER TWOC AFTER HIS SECOND ATTEMPT AND INSTEAD PREFERRED LONG-TERM CATHETERIZATION. IN CONCLUSION, THE STUDY SHOWED THAT PATIENT IPSS, QOL SCORE, AND MAXIMUM FLOW RATE WERE IMPROVED BY 40 TO 50 PERCENT WITH AN OVERALL SURGICAL RETREATMENT RATE OF 4.4 PERCENT IN 5 YEARS. FURTHER, MALE SEXUAL FUNCTION COULD BE PRESERVED. IN CONCLUSION, REZUM WAS SHOWN TO BE EFFECTIVE AND SAFE IN CATHETER-DEPENDENT CHINESE MEN WITH BPH IN MAKING THEM CATHETER FREE, IMPROVING THEIR VOIDING PARAMETERS, AND SYMPTOMS SCORE. THIS REPORT IS FOR THE DEATH INFORMATION PROVIDED AT THE LITERATURE ARTICLE SOURCE. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. CITATION: HENRY CHOW, CHI HO IP, YIU CHUNG LAM. (2024). TRANSURETHRAL WATER VAPOUR THERAPY (REZUM) FOR CATHETER-DEPENDENT MEN SECONDARY TO BENIGN PROSTATIC HYPERPLASIA: A RETROSPECTIVE STUDY IN A HONG KONG POPULATION. JOHN WILEY & SONS AUSTRALIA, LTD., 2024 (28), 87-92.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22435507&pc=OZD | 1220648-2025-30105 | 2025-06-19 04:00 | Death | ABIOMED, INC. | 2025-07-08 04:00 | OZD | IMPELLA CP | Positioning Failure; Restricted Flow rate | Cardiac Arrest; Tachycardia; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP POST PROCEDURE FOR ST-ELEVATION MYOCARDIAL INFARCTION AND HAD TO BE LEFT IN A DEEPER POSITION DUE TO HYPO-KINESIS OF LEFT VENTRICLE. THE SITE WAS APPROPRIATELY DRESSED AND GROIN BEST PRACTICES WERE USED. NO BLEEDING OR HEMATOMA AT THE GROIN SITE WERE NOTED. DURING SUPPORT, IT WAS NOTED THAT THE PATIENTS VITALS BEGAN TO CRASH AND CODE WAS CALLED. ADVANCED CARDIAC LIFE SUPPORT (ACLS) WAS PROVIDED FOR A TOTAL OF FIVE MINUTES WHICH OBTAINED RETURN OF SYSTEMIC CIRCULATION. AN ECHOCARDIOGRAM SHOWED DEEP PLACEMENT. HOWEVER, THE PHYSICIAN WAS UNABLE TO GET THE CATHETER IN A BETTER POSITION SO IT IS STILL CURRENTLY AROUND FIVE CENTIMETERS IN THE VENTRICLE. THE PATIENT'S FEET WERE NOTED TO BE COOL BILATERALLY SO BILATERAL LOWER EXTREMITY (BLE) ARTERIAL DOPPLER WAS COMPLETED. PER THE PRELIMINARY REPORT, BOTH LEGS HAD ADEQUATE FLOW. THE PATIENT WAS STILL MECHANICALLY VENTILATED AND SEDATED. THE PATIENT DID SUFFER A PULSELESS ELECTRICAL ACTIVITY ARREST WHILE TURNING THE PATIENT REQUIRING TWO ROUNDS OF CARDIOPULMONARY RESUSCITATION AND ACLS MEDICATIONS. IN ADDITION, IT WAS NOTED THAT THE ECHOCARDIOGRAM WAS COMPLETED AND THE IMPELLA CP WAS PULLED BACK TO A NEW LANDMARK. A LEFT VENTRICLE THROMBUS WAS NOTED ON ECHOCARDIOGRAM FROM THE PREVIOUS DAY. HOWEVER, IT WAS NOTED TO BE LARGER. THE PATIENT HAD A VENTRICULAR TACHYCARDIA ARREST AND THE FAMILY DECIDED TO MAKE THE PATIENT DO NOT RESUSCITATE. THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ Manufacturer Narrative: THE INVESTIGATION FOR VENTRICULAR TACHYCARDIA (VT,) POSITIONING ISSUES, LOW PUMP FLOW, AND THROMBUS HAS BEEN COMPLETED. CLINICAL DETAILS STATE THAT THE PATIENT SUFFERED A PEA ARREST OVERNIGHT. PATIENT HAD VTACH ARREST AND FAMILY DECIDED TO MAKE PATIENT DNR. THE ROOT CAUSE OF THE PEA ARREST/VTACH ARREST WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. THE CAUSE OF THE POSITIONING ISSUES WAS MOST LIKELY PATIENT CONDITION RELATED AS THE CLINICAL DETAILS STATED THE PUMP WAS LEFT DEEP DUE TO HYPOKINESIS OF LV. THE CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY PATIENT CONDITION RELATED AS CLINICAL DETAILS NOTED RV DYSFUNCTION, HIGH SVR, HYPOKINESIS OF THE LV, AND GIVING VOLUME TEMPORARILY RESOLVED SUCTION, WHICH CORRESPONDED WITH THE TRENDS SEEN IN THE DATA LOGS. THE ROOT CAUSE OF THE THROMBUS WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22435809&pc=PNJ | 9612164-2025-03353 | 2025-06-13 04:00 | Death | MEDTRONIC IRELAND | 2025-07-08 04:00 | PNJ | MICRA¿ AV2 | Positioning Problem | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE IPG EXHIBITED PLACEMENT DIFFICULTY. IT WAS FURTHER REPORTED THAT THE DELIVERY SYSTEM EXHIBITED DEFLECTION ISSUES. IT WAS FURTHER NOTED TWO WEEKS POST IMPLANT THAT THE PATIENT DIED. THE SYSTEM WAS ATTEMPTED NOT USED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT DEATH WAS DEEMED PROBABLE CARDIAC ARREST AND WAS NOT RELATED TO PROCEDURE. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE FULL LEAD WAS RETURNED AND ANALYZED. THE LUMEN OF THE DELIVERY SYSTEM WAS TORN. THE ANALYST NOTED THE FULL DELIVERY SYSTEM WAS RETURNED WITH THE DEVICE INTACT IN THE DEVICE CUP. THE STABILITY MEMBER WAS KINK/BUCKLED AT THE DISTAL END OF THE DELIVERY SYSTEM. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [(B)(6)]). D9: YES, RETURN DATE: 26-JUN-2025 H3: YES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MC2AVR1] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [SERIAL #(B)(6)). D9: YES, RETURN DATE: 26-JUN-2025. H3: YES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22436177&pc=NVZ | 3008973940-2025-04925 | 2025-06-23 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-07-08 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED FIVE DAYS POST IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. FOLLOW-UP HAS BEEN CARRIED OUT TO DETERMINE THE CAUSE OF DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22436725&pc=DQY | 2124215-2025-45001 | 2025-07-02 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-08 04:00 | DQY | WATCHMAN TRUSTEER? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Bradycardia; Cardiac Arrest; Low Blood Pressure/ Hypotension; Cardiac Tamponade; Pericardial Effusion | K240018 |  | 1 | Event Description: IT WAS REPORTED PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND A 35MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS SELECTED FOR USE. THE FIRST TRANSEPTAL PUNCTURE (TSP) WAS PERFORMED WITHOUT FLUOROSCOPIC GUIDANCE USING THE VERSACROSS CONNECT. IT WAS NOTED THAT THE TSP WAS NOT AT AN IDEAL POSITION. DESPITE THE SUBOPTIMAL LOCATION, THE PHYSICIAN PROCEEDED WITH ATTEMPTS TO DEPLOY THE WDS HOPING THE WAS TRUSTEER SHEATH WOULD HELP COMPENSATE FOR THE INADEQUATE TSP. MULTIPLE ATTEMPTS TO POSITION THE DEVICE CORRECTLY IN THE APPENDAGE WERE MADE, BUT THE POSITIONING WAS DETERMINED NOT ACCEPTABLE. THE PHYSICIAN RECAPTURED THE CLOSURE DEVICE AND REMOVED BOTH THE WDS AND WAS TRUSTEER SHEATH IN ORDER TO PERFORM A SECOND, BETTER-POSITIONED TSP. UPON REMOVAL OF THE DEVICES, A PERICARDIAL EFFUSION WAS NOTED. PERICARDIOCENTESIS WAS INITIATED, AND 2-3 SYRINGES (60CC EACH) OF BLOOD WERE DRAINED, TEMPORARILY STABILIZING THE PATIENT. A SECOND TSP WAS SUCCESSFULLY PERFORMED, AND THE TRUSTEER SHEATH AND 35MM DEVICE WERE RE-INTRODUCED AND THE CLOSURE DEVICE WAS REDEPLOYED. INITIALLY, THE TEAM BELIEVED THE CLOSURE DEVICE HELPED TO STOP THE EFFUSION, BUT THE EFFUSION REACCUMULATED. THE PHYSICIAN DECIDED TO RELEASE THE 35MM CLOSURE DEVICE WHICH WAS WELL POSITIONED IN THE APPENDAGE. ADDITIONAL DRAINAGE WAS THEN PERFORMED, BUT THE PATIENT WENT INTO CARDIAC ARREST. THE NURSES PERFORMED HEART MASSAGE. THE PATIENT EXPERIENCED SEVERE BRADYCARDIA, AND HYPOTENSION (ARTERIAL TENSION LOW 16 HIGH 34). A DEFIBRILLATOR (200 JOULESX1) WAS USED, AND LIFE CELL WAS USED TO GIVE ADDITIONAL BLOOD UNITS. A CODE WAS CALLED IN THE HOSPITAL TO ADVISE THE SURGERY TEAM. PROTAMINE WAS GIVEN TO THE PATIENT. THE PATIENT SEEMED TO STABILIZE SLIGHTLY, BUT THE EFFUSION PERSISTED. THE PHYSICIAN KEPT REMOVING 60CC SYRINGES OF BLOOD FROM THE PATIENT. THE PATIENT REMAINED IN THE INTERVENTIONAL CARDIOLOGY LAB UNDER ACTIVE RESUSCITATION AND WAS REPORTED TO HAVE DIED.Event Description: IT WAS REPORTED PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND A 35MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS SELECTED FOR USE. THE FIRST TRANSEPTAL PUNCTURE (TSP) WAS PERFORMED WITHOUT FLUOROSCOPIC GUIDANCE USING THE VERSACROSS CONNECT. IT WAS NOTED THAT THE TSP WAS NOT AT AN IDEAL POSITION. DESPITE THE SUBOPTIMAL LOCATION, THE PHYSICIAN PROCEEDED WITH ATTEMPTS TO DEPLOY THE WDS HOPING THE WAS TRUSTEER SHEATH WOULD HELP COMPENSATE FOR THE INADEQUATE TSP. MULTIPLE ATTEMPTS TO POSITION THE DEVICE CORRECTLY IN THE APPENDAGE WERE MADE, BUT THE POSITIONING WAS DETERMINED NOT ACCEPTABLE. THE PHYSICIAN RECAPTURED THE CLOSURE DEVICE AND REMOVED BOTH THE WDS AND WAS TRUSTEER SHEATH IN ORDER TO PERFORM A SECOND, BETTER-POSITIONED TSP. UPON REMOVAL OF THE DEVICES, A PERICARDIAL EFFUSION WAS NOTED. PERICARDIOCENTESIS WAS INITIATED, AND 2-3 SYRINGES (60CC EACH) OF BLOOD WERE DRAINED, TEMPORARILY STABILIZING THE PATIENT. A SECOND TSP WAS SUCCESSFULLY PERFORMED, AND THE TRUSTEER SHEATH AND 35MM DEVICE WERE RE-INTRODUCED AND THE CLOSURE DEVICE WAS REDEPLOYED. INITIALLY, THE TEAM BELIEVED THE CLOSURE DEVICE HELPED TO STOP THE EFFUSION, BUT THE EFFUSION REACCUMULATED. THE PHYSICIAN DECIDED TO RELEASE THE 35MM CLOSURE DEVICE WHICH WAS WELL POSITIONED IN THE APPENDAGE. ADDITIONAL DRAINAGE WAS THEN PERFORMED, BUT THE PATIENT WENT INTO CARDIAC ARREST. THE NURSES PERFORMED HEART MASSAGE. THE PATIENT EXPERIENCED SEVERE BRADYCARDIA, AND HYPOTENSION (ARTERIAL TENSION LOW 16 HIGH 34). A DEFIBRILLATOR (200 JOULESX1) WAS USED, AND LIFE CELL WAS USED TO GIVE ADDITIONAL BLOOD UNITS. A CODE WAS CALLED IN THE HOSPITAL TO ADVISE THE SURGERY TEAM. PROTAMINE WAS GIVEN TO THE PATIENT. THE PATIENT SEEMED TO STABILIZE SLIGHTLY, BUT THE EFFUSION PERSISTED. THE PHYSICIAN KEPT REMOVING 60CC SYRINGES OF BLOOD FROM THE PATIENT. THE PATIENT REMAINED IN THE INTERVENTIONAL CARDIOLOGY LAB UNDER ACTIVE RESUSCITATION AND WAS REPORTED TO HAVE DIED. IT WAS FURTHER REPORTED THAT THE PHYSICIAN SUSPECTED THAT THE INADEQUATE TRANSEPTAL LOCATION CAUSED THE WATCHMAN TRUSTEER ACCESS SYSTEM TO RUB AGAINST THE RIDGE AND POSSIBLY CAUSED A SMALL TEAR IN THE APPENDAGE. THERE WAS NO PERFORATION NOTED ON ANY IMAGING. THE CAUSE OF DEATH WAS DETERMINED TO BE TAMPONADE. Manufacturer Narrative: B5: DESCRIBE EVENT OR PROBLEM: UPDATED WITH ADDITIONAL INFORMATION RECEIVED. D1: BRAND NAME CORRECTED FROM WATCHMAN FLX PRO TO WATCHMAN TRUSTEER ACCESS SYSTEM. D4: LOT NUMBER UPDATED. H6: PATIENT CODE ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22436747&pc=OZD | 1220648-2025-30115 | 2025-06-16 04:00 | Death | ABIOMED, INC. | 2025-07-08 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH CARDIOMYOPATHY WAS ESCALATED FROM AN INTRA-AORTIC BALLOON PUMP (IABP) TO AN IMPELLA 5.5 DEVICE MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED BLEEDING AND ARRHYTHMIC EVENTS REQUIRING INTERVENTION. THE IMPELLA IMPLANT PROCEDURE TOOK PLACE IN AN OPERATING ROOM. THE PATIENT WAS INTUBATED, PREPPED, AND DRAPED. THE PHYSICIAN PERFORMED A CUTDOWN TO THE RIGHT AXILLARY ARTERY AND A10MM VASCUTEK GELWEAVE WAS ANASTOMOSED TO THE ARTERY. THE PEEL AWAY SHEATH WAS ADVANCED INTO THE GRAFT AND SECURED WITH A GRAFT LOCK. ACCESS TO THE LEFT VENTRICLE WAS OBTAINED AND THE IMPELLA 5.5 WAS INSERTED OVER THE GUIDEWIRE INTO THE LEFT VENTRICLE UNDER FLUOROSCOPY AND TRANSESOPHAGEAL ECHOCARDIOGRAM GUIDANCE. THE IMPELLA WAS STARTED AT PERFORMANCE LEVEL P-2 AND RAMPED UP TO P-7. THE PEEL AWAY SHEATH WAS REMOVED, GRAFT TRIMMED, AND THE BLUE SUTURE HUB WAS ADVANCED AND SECURED TO THE GRAFT. THE YELLOW PIN WAS REMOVED, AND THE BLUE SUTURE HUB WAS SECURED TO THE PATIENT. DRESSING WAS PLACED TO THE IMPELLA SITE AND THE IABP WAS REMOVED FROM THE RIGHT GROIN WITH FEMSTOP PLACEMENT. THE PATIENT WAS TRANSPORTED BACK TO THE CARDIOVASCULAR INTENSIVE CARE UNIT FOR CONTINUED MONITORING. IN THE EVENING, IT WAS NOTED THE IMPELLA SITE REMAINED CLEAR, DRY, AND INTACT, WITH NO EVIDENCE OF BLEEDING OR HEMATOMA. THE IMPELLA THREE-POINT FIXATION WAS IN PLACE. HOWEVER, THE FEMSTOP PLACED TO THE RIGHT GROIN WHERE THE IABP WAS BLEEDING. IT WAS NOTED HEPARIN DRIP WAS USED AND MAY HAVE CONTRIBUTED TO THE BLEEDING. MANUAL PRESSURE WAS HELD. ONE DAY LATER, THE DRESSING WAS CHANGED BY THE PRACTITIONER. WHILE THE DRESSING WAS CLEAR, DRY, AND INTACT, WITH NO EVIDENCE OF BLEEDING, A SMALL HEMATOMA WAS NOTED AT THE IMPELLA SITE. NOW FOUR DAYS POST START OF THE IMPELLA 5.5 SUPPORT, THE IMPELLA INSERTION SITE WAS NOTED TO HAVE A MINIMAL AMOUNT OF BLOODY DRAINAGE WITH STABLE HEMATOMA. NO INTERVENTION WAS PLANNED; THE CLINICAL REP CONFIRMED THAT THERE WAS NO KNOWN INTERVENTION FOR THE HEMATOMA. STAFF MONITORED; THE PHYSICIAN WAS AWARE. THE NEXT DAY, THE PATIENT WAS HAVING EPISODES OF VENTRICULAR TACHYCARDIA STORM, DESPITE THE AUTOMATED IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (AICD). THE PUMP SETTINGS WERE CHANGED, AND THE PATIENT WAS STARTED ON HEPARIN TO AGGRESSIVELY TITRATE FOR THERAPEUTIC RANGE, AS THE PATIENT HAD NOT BEEN PREVIOUSLY STARTED ON HEPARIN PRIOR TO HOSPITAL TRANSFER. OVERNIGHT, THE NEXT DAY, THE PATIENT HAD VENTRICULAR TACHYCARDIA AND WAS GIVEN AMIODARONE BOLUS. THE PATIENT CONTINUED TO HAVE PERSISTENT SMALL RUNS OF VENTRICULAR TACHYCARDIA. AN ECHOCARDIOGRAM WAS PERFORMED AT BEDSIDE AND CONFIRMED THE PUMP WAS IN STABLE POSITION. THE PATIENT'S BELLY WAS DISTENDED SO THE BOWEL REGIMEN WAS INCREASED. THE PATIENT WAS NOTED TO HAVE EMESIS OF OVER ONE LITER WITH INCREASED LACTIC AND DISTENDED BELLY. A COMPUTED TOMOGRAPHY DEMONSTRATED BOWEL ISCHEMIA. AFTER DISCUSSING CARE AND GOALS WITH THE FAMILY, THE PATIENT WAS PLACED ON DO NOT RESUSCITATE STATUS AND CHOSE TO WITHDRAW CARE. THE PATIENT ULTIMATELY EXPIRED. Manufacturer Narrative: THE PATIENT EXPIRY OCCURRED UPON WITHDRAWAL OF SUPPORT. HOWEVER, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA 5.5 DEVICE USED AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22437243&pc=DSP | 2248146-2025-0000357 | 2025-06-23 04:00 | Death | DATASCOPE CORP. | 2025-07-08 04:00 | DSP | TRANS-RAY PLUS 7.5FR 35CC IAB WITH ACCESSORIES | Incorrect, Inadequate or Imprecise Result or Readings | Myocardial Infarction | K122628 |  | 1 | Event Description: IT WAS REPORTED THAT DURING INTRA-AORTIC BALLOON (IAB) THERAPY, THE DEVICE WAS USED PER STANDARD PROCEDURE. HOWEVER, IN FULL AUTO MODE WITH ECG TRIGGER, THE TRP SENSOR PRESSURE WAVEFORM AND NUMERICAL VALUES WERE NOT DISPLAYED ON THE CSAVE MONITOR, DESPITE CONTINUED IAB CATHETER ASSISTANCE. RECONNECTING THE SENSOR CONNECTOR HAD NO EFFECT. THE TRANSDUCER WAS THEN CONNECTED TO THE Y-FITTING, ALLOWING ARTERIAL PRESSURE INPUT TO CONTINUE IABP THERAPY. THE PATIENT LATER PASSED AWAY ON (B)(6) 2025. SUBSEQUENT TESTING OF THE CARDIOSAVE UNIT WITH A TEST BALLOON AND SYSTEM TRAINER SHOWED NORMAL ARTERIAL PRESSURE WAVEFORM AND VALUES. Manufacturer Narrative: UPDATED FIELDS: B4, D8, D9, E1 (INITIAL REPORTER), G1 (CONTACT PERSON MFG SITE), G3, G6, H2, H3, H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, COMPONENT CODES AND INVESTIGATION CONCLUSIONS), H11. NO DECON OR VISUAL INSPECTION PERFORMED SINCE PRODUCT WAS NOT RETURNED AND COULD NOT BE EVALUATED. THEREFORE, WE WERE UNABLE TO CONFIRM OR DETERMINE A ROOT CAUSE FOR THE REPORTED FAILURE. THE FAILURE MODE IS ADDRESSED IN THE RISK FILE AND IS OPERATING WITHIN ITS RISK PROFILE. THE IFU ADDRESSES THE REPORTED FAILURE. THERE WERE NO NCMRS IDENTIFIED WHICH COULD CAUSE OR CONTRIBUTE TO THE REPORTED FAILURE. THE INVESTIGATION DOES NOT INDICATE THAT THE DEVICE WAS INADVERTENTLY RELEASED AS NON-CONFORMING OR AN ADULTERATED PRODUCT OR WAS A COUNTERFEIT. THE COMPLAINT HISTORY REVIEW DID NOT IDENTIFY AN ADVERSE TREND (INCREASE IN NUMBER OF COMPLAINTS OVER PAST THREE (3) MONTHS). BASED ON THE RATIONAL PROVIDED ABOVE, NO ESCALATION TO THE CAPA PROCESS IS REQUIRED. Manufacturer Narrative: A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22437743&pc=NIQ | 2024168-2025-06968 | 2025-06-21 04:00 | Death | ABBOTT VASCULAR INC. | 2025-07-08 04:00 | NIQ | XIENCE SIERRA¿ | Adverse Event Without Identified Device or Use Problem | Angina | P110019 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE MID LEFT ANTERIOR DESCENDING (LAD) CORONARY ARTERY WITH HEAVY CALCIFICATION, MODERATE TORTUOSITY AND 80% STENOSIS. THE LESION WAS ACCESSED VIA THE RIGHT FEMORAL ARTERY WITH A 6FR SHEATH AND A WIRE FROM ANOTHER MANUFACTURER. ORBITAL ATHERECTOMY WAS PERFORMED AND THERE WAS NO REFLOW OBSERVED DURING THE ATHERECTOMY RUNS. PRE-DILATATION WAS PERFORMED WITH A 3X12 MM NC BALLOON AT 12 ATMOSPHERES (ATM) AND THE 2.75X48 MM XIENCE XPEDITION AND 4.0X12 MM XIENCE SIERRA STENT WERE IMPLANTED FROM THE MID TO PROXIMAL LAD. THE LESION IN THE CIRCUMFLEX CORONARY ARTERY WAS PRE-DILATED WITH A 2.5X15 MM NC BALLOON A 12 ATM AND THE 2.5X38 MM XIENCE SIERRA STENT WAS IMPLANTED. THE STENT WAS POST-DILATED WITH 2.5X12 MM AND 3X8 MM BALLOONS. PLAIN OLD BALLOON ANGIOPLASTY WAS PERFORMED AT THE SECOND OBTUSE MARGINAL CORONARY ARTERY. FINAL RESULTS WERE GOOD WITH TIMI 3 FLOW AND NO DISSECTION OR THROMBUS WAS SEEN IN THE FINAL SHOT OF ANGIOGRAPHY. THE PATIENT WAS STABLE AND DISCHARGED FROM THE HOSPITAL ON (B)(6) 2025 WITH ANTIPLATELETS, STATINS AND OTHER CARDIAC SUPPORTIVE DRUGS. ON (B)(6) 2025 THE PATIENT EXPERIENCED SUDDEN CHEST PAIN AT HOME AND EXPIRED DUE TO AN UNKNOWN REASON. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22437945&pc=OZD | 1220648-2025-30123 | 2025-06-15 04:00 | Death | ABIOMED, INC. | 2025-07-08 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP AND EXPERIENCED BLEEDING AT THE GROIN SITE WHICH WAS RESOLVED WITH A FIGURE EIGHT DRESSING. THE PATIENT'S FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED THE FOLLOWING DAY. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF ACCESS SITE BLEEDING WAS MOST LIKELY USE ISSUE SINCE NON-ABIOMED PRODUCT WAS USED WHICH IS NOT RECOMMENDED PER INSTRUCTIONS FOR USE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER. ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) . ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22437957&pc=NPT | 2015691-2025-05538 | 2025-06-18 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-08 04:00 | NPT | COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem |  | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS AFFILIATES IN ITALY, DURING A TAVR PROCEDURE UTILIZING A 23MM SAPIEN 3 ULTRA VALVE VIA THE TRANSAXILLARY APPROACH, RESISTANCE WAS ENCOUNTERED WHILE ADVANCING THE DELIVERY SYSTEM THROUGH THE ESHEATH+. THE DECISION WAS MADE TO REMOVE THE DEVICES. A NEW 16F ESHEATH WAS THEN INTRODUCED, AND THE VALVE ADVANCED WITHOUT ISSUE. HOWEVER, FOLLOWING THE ALIGNMENT STEP, A TEAR IN THE AORTA WAS OBSERVED. THE PATIENT ULTIMATELY PASSED AWAY. PATIENT DEMOGRAPHICS WERE NOT PROVIDED. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. AORTIC DISSECTION OR RUPTURE MAY OCCUR WHEN MULTIPLE ATTEMPTS ARE MADE TO CROSS THE STENOTIC NATIVE VALVE, AND/OR WHEN EXCESSIVE FORCE IS USED. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). THE THV TRAINING MANUALS GUIDE TO FACILITATE SAFE CROSSING OF THE NATIVE OR FAILING VALVE, INCLUDING CAMERA PROJECTIONS, HANDLING DURING ADVANCEMENT, AND TROUBLESHOOTING TECHNIQUES IF DIFFICULTY IS ENCOUNTERED. AS STATED, EXCESSIVE FORCE SHOULD NOT BE USED WHEN THE DEVICE HAS DIFFICULTY CROSSING THE STENOTIC VALVE. ADDING TENSION TO THE WIRE, PULLING BACK THE SYSTEM TO RE-ORIENT THE VALVE, AS NEEDED, AND TORQUING THE FLEX CATHETER MAY HELP SOLVE THE PROBLEM. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION THAT A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS ARE INCONCLUSIVE AS RELEVANT PATIENT, AND PROCEDURAL FACTORS WERE NOT PROVIDED. HOWEVER, THE EVENT MAY BE RELATED TO THE MECHANISM DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22437959&pc=DSQ | 2916596-2025-04476 | 2025-06-02 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WENT INTO CARDIAC ARREST ON (B)(6) 2025 AND WAS TRANSPORTED TO THE (B)(6) HOSPITAL EMERGENCY DEPARTMENT (ED) WHERE CPR WAS ADMINISTERED BUT THE PATIENT WAS UNABLE TO BE REVIVED. CARDIOPULMONARY RESUSCITATION WAS PERFORMED, AND THE PATIENT WAS TRANSPORTED TO THE LOCAL EMERGENCY DPEARTMENT. THE PATIENT WAS UNABLE TO BE REVIVED. THE VENTRICULAR ASSIST DEVICE (VAD) WAS FUNCTIONING AS INTENDED. NO AUTOPSY WAS PERFORMED AND THE VAD WOULD NOT BE RETURNED. THERE WAS NO FURTHER INFORMATION AVAILABLE PER THE VAD COORDINATOR. Manufacturer Narrative: SECTION H6: INVESTIGATION CONCLUSIONS CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22438243&pc=GKT | 1722028-2025-00151 | 2025-06-13 04:00 | Death | TERUMO BCT | 2025-07-08 04:00 | GKT | RIKA | No Apparent Adverse Event; Insufficient Information | Insufficient Information; No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: THE CUSTOMER REPORTED A DONOR DEATH WITHIN 36 HOURS OF DONATING. PATIENT INFORMATION IS UNKNOWN AT THIS TIME. THIS REPORT IS BEING FILED DUE TO PATIENT DEATH, ALTHOUGH PER CURRENT INFORMATION THERE IS NO DETECTABLE MALFUNCTION WITH THE TERUMO BCT DEVICE OR ALLEGATION OF A MALFUNCTION.Event Description: THE CUSTOMER REPORTED A DONOR DEATH WITHIN 36 HOURS OF DONATING. PATIENT INFORMATION IS UNKNOWN AT THIS TIME. THIS REPORT IS BEING FILED DUE TO PATIENT DEATH, ALTHOUGH PER CURRENT INFORMATION THERE IS NO DETECTABLE MALFUNCTION WITH THE TERUMO BCT DEVICE OR ALLEGATION OF A MALFUNCTION. Manufacturer Narrative: THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION IN H.6 AND H.11. INVESTIGATION: THE MACHINE WAS CHECKED OUT AT THE CUSTOMER SITE. IT WAS DETERMINED THAT NO WORK WAS NEEDED AND THE DEVICE WAS RETURNED TO SERVICE. FURTHER EVALUATION OF THIS EVENT HAS DETERMINED THAT THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY, NOR IS THERE A LIKELY POTENTIAL FOR DEATH OR SERIOUS INJURY ASSOCIATED WITH THIS EVENT BASED ON ADDITIONAL INVESTIGATIONAL INFORMATION. IT WAS CONFIRMED THAT THE REPORTED OUTCOME OF THE DONOR DEATH IS DUE TO AN UNDISCLOSED HEART CONDITION. NO FURTHER REPORTING WILL BE PROVIDED AS THIS DOES NOT REPRESENT A REPORTABLE EVENT. Manufacturer Narrative: INVESTIGATION: THE MACHINE WAS CHECKED OUT AT THE CUSTOMER SITE. IT WAS DETERMINED THAT NO WORK WAS NEEDED AND THE DEVICE WAS RETURNED TO SERVICE. INVESTIGATION IS IN PROCESS, A FOLLOW-UP REPORT WILL BE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22438558&pc=DSQ | 2916596-2025-04492 | 2025-06-14 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR FAILURE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WAS SAID TO NOT BE RETURNING FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. D, IS ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22438791&pc=OZD | 1220648-2025-30131 | 2025-06-17 04:00 | Death | ABIOMED, INC. | 2025-07-08 04:00 | OZD | IMPELLA CP | Positioning Problem | Arrhythmia; Hemolysis | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD VENTRICULAR TACHYCARDIA REQUIRING 1 SHOCK AT 200J. SUPPORT CONTINUED. ADDITIONALLY, THE PATIENT'S LDH LEVELS WERE RISING THE PUMP WAS REPOSITIONED AND PULLED BACK. THE CLINICAL REP CONFIRMED THAT THE PATIENT HAD NO URINE OUTPUT SO HEMOLYSIS COULD NOT BE VISUALIZED AND CONFIRMED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22438839&pc=DSQ | 2916596-2025-04520 | 2025-06-19 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Infusion or Flow Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS READMITTED TO THE HOSPITAL ON (B)(6) 2025 FOR A SUDDEN DROP IN FLOW AND LOW FLOW ALARMS. AN ECHOCARDIOGRAM WAS PERFORMED AND SHOWED AN UNLOADED LEFT VENTRICLE AND ACUTE RIGHT HEART FAILURE. IT WAS NOTED THAT THE PATIENT PREVIOUSLY HAD LIMITED PUMP FUNCTION FROM THE RIGHT VENTRICLE, BUT THE LVAD IMPLANT WAS IN 2019. A NEW COMPUTED TOMOGRAPHY SCAN SHOWED NO SIGNS OF EXTRINSIC OUTFLOW GRAFT OBSTRUCTION (EOGO). LOG FILES WERE SUBMITTED FOR REVIEW. THE ISSUE WAS NOT A PUMP ISSUE, IT WAS MEMBRANE/STENOSIS ON THE AORTIC ANASTOMOSIS BETWEEN THE OUTFLOW GRAFT AND THE AORTA. THERE WERE NO SYMPTOMS REPORTED FROM THE PATIENT. THE PATIENT WAS INITIALLY TREATED WITH THROMBOLYSIS BY SUSPECTED PUMP THROMBOSIS. AFTER THAT THE PATIENT UNDERWENT IMPLANTATION FROM A VENO/ARTERY EXTRACORPOREAL LIFE SUPPORT (ECLS) SYSTEM TO STABILIZE THE HEMODYNAMICS. THE PUMP WAS EXCHANGED ON (B)(6) 2025. A NEW PUMP WAS PUT IN PLACE WITH A NEW OUTFLOW GRAFT WITH END-TO-END ANASTOMOSIS. THE SET SPEED FROM THE PUMP WAS 8000 RPM FOR 4LPM WITH TEMPORARY RIGHT VENTRICULAR (RV) SUPPORT. FURTHER ANALYSIS ON COMPUTED TOMOGRAPHY (CT) SCAN SHOWED STENOSIS OR ¿MEMBRANE¿ BETWEEN OLD OUTFLOW GRAFT AND ASCENDING AORTA, AND A NEW SURGICAL REVISION ON THE AORTIC ANASTOMOSIS ON (B)(6) 2025 WAS PERFORMED. AFTER THE SURGICAL APPROACH FROM THE ANASTOMOSIS THERE WERE NORMAL PARAMETERS ON THE LVAD.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO MULTIPLE ORGAN DYSFUNCTION SYNDROME. THERE WERE NO DEVICE OR THERAPY ISSUES, AND THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: SECTION G2 CORRECTION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22438996&pc=DSQ | 2916596-2025-04359 | 2025-06-15 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A STROKE AND PASSED AWAY AFTER WITHDRAWAL OF CARE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22439177&pc=OZD | 1220648-2025-30135 | 2025-06-17 04:00 | Death | ABIOMED, INC. | 2025-07-08 04:00 | OZD | IMPELLA 5.5 | Positioning Failure; Malposition of Device; Device Sensing Problem | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT AN IMPLANTED IMPELLA 5.5 WAS POINTED AT THE MITRAL VALVE. ATTEMPTS TO REPOSITION THE PUMP WERE UNSUCCESSFUL. ADDITIONALLY, THERE WERE PLACEMENT SIGNAL NOT RELIABLE ALARMS. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES AND OPTICAL SIGNAL ISSUE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. OPTICAL SIGNAL ISSUE: THE DATA LOG DID NOT RECORD ANY PLACEMENT SIGNAL NOT RELIABLE (PSNR) ALARMS ON THE DAY OF THE REPORTED ISSUE. THERE WERE NO ISSUES FOUND WITH THE 5S OPTICAL SIGNAL. THERE WAS NO ISSUE FOUND ON OPTICAL SIGNAL DURING THE REPORTED TIME OF PSNR EVENT. THE DEVICE FUNCTIONED/OPERATED TO SPECIFICATION. POSITIONING ISSUE: THE CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY USER-RELATED, AS THE PUMP WAS FOUND NEAR THE MITRAL APPARATUS DURING 18TH DAY OF SUPPORT, WITHOUT ANY REPORTED ISSUES WITH THE CATHETER ANCHOR ON THE PUMP. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITES STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22439792&pc=PNJ | 2017865-2025-91618 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-08 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO UNKNOWN CAUSES. FURTHER INFORMATION WAS REQUESTED BUT WAS NOT AVAILABLE. Manufacturer Narrative: CORRECTION: PLEASE RETRACT REPORT 2017865-2025-91618. INFORMATION RECEIVED FROM THE FIELD CONFIRMED THAT THE PATIENT'S DEATH WAS UNRELATED TO THEIR PACING SYSTEM. THERE WERE NO ALLEGATIONS OF DEVICE DEFICIENCY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440039&pc=DSQ | 2916596-2025-04218 | 2025-03-11 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE (MSOF). THE PATIENT EXPIRED UNDER REFRACTORY MULTI-ORGAN FAILURE WHICH WAS NOT DEVICE RELATED. THE OUTCOME WAS RELATED TO MULTIORGAN FAILURE. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) OPERATED AS EXPECTED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440105&pc=DSQ | 2916596-2025-04295 | 2021-10-13 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | HEARTMATE II LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Fall; Intracranial Hemorrhage | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO A SUBARACHNOID HEMORRHAGE CAUSED BY THE PATIENT ACCIDENTALLY STUMBLING AND FALLING ON THEIR HEAD. THERE WERE NO DEVICE RELATED ISSUES OR ALARMS PRIOR TO THE FALL. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6) , AND THE REPORTED BLEEDING AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LVAS IFU, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO RANGE, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: D4: UDI AND EXPIRATION DATE CORRECTED. H4: MANUFACTURE DATE CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440238&pc=DSQ | 2916596-2025-04292 | 2025-06-13 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. HOWEVER, THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED AND THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440290&pc=OZD | 1220648-2025-30140 | 2025-06-17 04:00 | Death | ABIOMED, INC. | 2025-07-08 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP HAD HEMOLYZED LABS AND DISCOLORED URINE. THE PATIENT ALSO HAD A THROMBUS IN THE LEFT VENTRICLE. TWO DAYS AFTER IMPLANTATION THE PATIENT CODED AND EXPIRED. Manufacturer Narrative: A4 AND A5 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440295&pc=DSQ | 2916596-2025-04266 | 2024-09-13 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR FAILURE FOLLOWED BY MULTI-SYSTEM ORGAN FAILURE. THE RIGHT HEART FAILURE WAS NOT CONSIDERED TO BE PRE-EXISTING; IT WAS CONSIDERED MILD IMPACT. THE PROGRESSION OF REFRACTORY MULTI-SYSTEM ORGAN FAILURE LED TO A PROGRESSIVE RIGHT VENTRICULAR FAILURE WHICH WAS RELATED TO THE DEATH. THE DEVICE WAS NOT RELATED TO THE MULTI SYSTEM ORGAN FAILURE OR THE RIGHT VENTRICULAR FAILURE. THE DEVICE OPERATED AS EXPECTED. THE DEATH WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE¿) STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THIS SECTION (UNDER ¿CAUTION!¿) ALSO EXPLAINS THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: H6: 3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440486&pc=MVK | 3008642652-2025-06606 | 2025-06-02 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-08 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 15:54:26 ON (B)(6) 2025. AT 06:46:36 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS TACHYCARDIA @ 100 BPM WITH NSVT. THE RHYTHM THEN DEGRADED TO VF. AT 06:47:12, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE DEVICE WAS SHUT DOWN AT 08:56:57 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440500&pc=MVK | 3008642652-2025-06607 | 2025-06-09 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-08 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 04:40:02 ON (B)(6) 2025. AT 12:04:42 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 12:06:20, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 12:23:15 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440512&pc=CBK | 2518422-2025-107171 | 2025-06-14 04:00 | Death | RESPIRONICS, INC. | 2025-07-08 04:00 | CBK | TRILOGY EVO, JAPAN | Adverse Event Without Identified Device or Use Problem | Respiratory Arrest | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT USING A FISHER AND PAYKEL VITERA MASK HAD DETACHED FROM THE CIRCUIT AT THE SWIVEL CONNECTOR AND A CIRCUIT DISCONNECT ALARM SOUNDED FROM THE TRILOGY EVO VENTILATOR. IT WAS REPORTED THAT THE PATIENT WAS FOUND LYING FACE DOWN IN RESPIRATORY ARREST. THE DOCTOR WAS CONTACTED, AND MANUAL VENTILATION WAS INITIATED, HOWEVER THE PATIENT EXPIRED. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT A PATIENT USING A FISHER AND PAYKEL VITERA MASK HAD DETACHED FROM THE CIRCUIT AT THE SWIVEL CONNECTOR AND A CIRCUIT DISCONNECT ALARM SOUNDED FROM THE TRILOGY EVO VENTILATOR. IT WAS REPORTED THAT THE PATIENT WAS FOUND LYING FACE DOWN IN RESPIRATORY ARREST. THE DOCTOR WAS CONTACTED, AND MANUAL VENTILATION WAS INITIATED, HOWEVER THE PATIENT EXPIRED. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". DURING THE EVALUATION OF THE DEVICE'S EVENT LOG NO ERROR INDICATING A DEVICE FAILURE FOR THE REPORTED DATE AND TIME WAS CONFIRMED. NO DEVICE FAILURE RELATED TO THE REPORTED ISSUE WAS CONFIRMED BY ANY OTHER INVESTIGATIONS EITHER. SINCE DIRT WAS CONFIRMED, INLINE PROXIMAL FILTER WAS REPLACED. SINCE USB EXTENSION CABLE WAS FOUND DAMAGED DUE TO HAVING BEEN PINCHED, IT WAS REPLACED. (THERE WAS NEITHER THERMAL DAMAGE NOR EXPOSED CONDUCTING WIRES.) PERFORMED FINAL TEST, BATTERY CHECK, VALVE CHECK, RUN IN TESTS AND CLEANING THEN CONFIRMED THE UNIT OPERATED PROPERLY. CONFIRMED THE SOFTWARE VERSION IS 1.06.10.00.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440560&pc=OZD | 1220648-2025-30144 | 2025-06-13 04:00 | Death | ABIOMED, INC. | 2025-07-08 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia; Cardiogenic Shock | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED LIMB ISCHEMIA AND EXPIRED FROM PROLONGED CARDIOGENIC SHOCK. Manufacturer Narrative: A3 AND A5 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440695&pc=FDS | 9610595-2025-13376 | 2025-06-01 04:00 | Death | AIZU OLYMPUS CO., LTD. | 2025-07-08 04:00 | FDS | EVIS LUCERA GASTROINTESTINAL VIDEOSCOPE | Adverse Event Without Identified Device or Use Problem | Necrosis |  |  | 1 | Event Description: IT WAS REPORTED THAT AN ALERT MEANT TO FILL TUBING WAS CONFIRMED, CAUSING CONCERN FOR HIGH BLOOD GLUCOSE LEVELS, WHICH WERE DISPLAYED AS "HIGH," THOUGH NO SPECIFIC VALUE WAS PROVIDED. THE LOWEST BLOOD GLUCOSE LEVEL RECORDED WAS 86 MG/DL. THE CUSTOMER TOOK CORRECTIVE MEASURES BY ADMINISTERING A CORRECTION BOLUS VIA THE PUMP AND CONSUMING CARBOHYDRATES TO MANAGE INSULIN EFFECTS. TECHNICAL SUPPORT EDUCATED THE CUSTOMER ON HANDLING ALERTS AND ACCIDENTAL INSULIN DELIVERY, REVIEWED LOGS TO VALIDATE THE FILL TUBING PROCESS, AND ADVISED FOLLOW-UP WITH HEALTHCARE PROFESSIONALS IF ISSUES CONTINUED.Event Description: NO ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER.Event Description: SUBSEQUENT FOLLOW-UP WITH THE CORRESPONDING AUTHOR CONFIRMED THAT THERE WAS NO MALFUNCTION OF THE DEVICE DURING ANY OF THE REPORTED PROCEDURES AND THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED ADVERSE EVENTS, INCLUDING DEATH. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE ADDITIONAL INFORMATION FROM THE AUTHOR CONFIRMING THAT THE ADVERSE EVENT AND DEATH THAT OCCURRED WERE NOT CAUSED BY THE DEVICE. NEW INFORMATION ADDED TO THE FOLLOWING FIELDS: B5. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE THE LEGAL MANUFACTURER'S INVESTIGATION. A DEFINITIVE ROOT CAUSE WAS NOT IDENTIFIED. BASED ON THE AVAILABLE INFORMATION, THE LEGAL MANUFACTURER WAS UNABLE TO DETERMINE THE PROBABLE CAUSE OF THE ADVERSE EVENTS SINCE THE DEVICE WAS NOT RETURNED FOR EVALUATION. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE. Manufacturer Narrative: THIS REPORT IS RELATED TO THE FOLLOWING LINKED PATIENT IDENTIFIERS: (B)(6). THE INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED OR IF ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440701&pc=FDS | 9610595-2025-13377 | 2025-06-01 04:00 | Death | AIZU OLYMPUS CO., LTD. | 2025-07-08 04:00 | FDS | EVIS LUCERA ELITE GASTROINTESTINAL VIDEOSCOPE | Adverse Event Without Identified Device or Use Problem | Necrosis |  |  | 1 | Event Description: OLYMPUS REVIEWED THE FOLLOWING LITERATURE TITLED ¿NECROSECTOMY AND ITS TIMING IN RELATION TO CLINICAL OUTCOMES OF EUS-GUIDED TREATMENT OF WALLED-OFF PANCREATIC NECROSIS: A MULTICENTER STUDY¿. LITERATURE SUMMARY: BACKGROUND AND AIMS: EUS-GUIDED TRANSMURAL DRAINAGE WITH ON-DEMAND ENDOSCOPIC NECROSECTOMY (EN) IS INCREASINGLY USED TO MANAGE WALLED-OFF NECROSIS (WON). IT HAS NOT BEEN FULLY ELUCIDATED HOW EN AND ITS TIMING ARE CORRELATED WITH TREATMENT OUTCOMES COMPARED WITH THE DRAINAGE-BASED APPROACH. METHODS: WITHIN A MULTI-INSTITUTIONAL COHORT OF 423 PATIENTS WITH PANCREATIC FLUID COLLECTIONS, INCLUDING 227 PATIENTS WITH WON, 153 PATIENTS WERE IDENTIFIED WHO RECEIVED THE STEP-UP TREATMENT AFTER THE INITIAL EUS-GUIDED DRAINAGE OF SYMPTOMATIC WON; THIS INCLUDED 102 EN PATIENTS AND 51 NON-EN (DRAINAGE) PATIENTS. USING THE COMPETING-RISKS MULTIVARIABLE PROPORTIONAL HAZARDS REGRESSION MODEL WITH ADJUSTMENT FOR POTENTIAL CONFOUNDERS, WE CALCULATED SUBDISTRIBUTION HAZARD RATIOS (SHRS) FOR CLINICAL TREATMENT SUCCESS (WON RESOLUTION) ACCORDING TO USE OF EN AND ITS TIMING. RESULTS: COMPARED WITH DRAINAGE ALONE, THE EN-BASED TREATMENT WAS ASSOCIATED WITH A SHORTER TIME TO CLINICAL SUCCESS WITH A MULTIVARIABLE SHR OF 1.66 (95% CONFIDENCE INTERVAL, 1.12-2.46). DESPITE A HIGHER RISK OF PROCEDURE-RELATED BLEEDING IN THE EN GROUP, THERE WERE NO DIFFERENCES IN THE RATES OF SEVERE ADVERSE EVENTS (7.8% VS 5.9% IN THE EN AND NON-EN GROUPS, RESPECTIVELY) OR MORTALITY (6.9% VS 9.8%). IN THE EN-TREATED PATIENTS, THE TIMING OF EN WAS NOT STATISTICALLY SIGNIFICANTLY ASSOCIATED WITH THE TIME TO CLINICAL SUCCESS (P TREND = .34). 7 PATIENT DEATHS WERE REPORTED IN THE STUDY. NO FURTHER DETAILS REGARDING THE PATIENTS' DEATHS WERE PROVIDED IN THE LITERATURE. CONCLUSIONS: AMONG PATIENTS RECEIVING EUS-GUIDED TREATMENT OF SYMPTOMATIC WON, THE USE OF EN IN ADDITION TO DRAINAGE PROCEDURES WAS ASSOCIATED WITH EARLIER DISEASE RESOLUTION. FURTHER RESEARCH IS DESIRED TO DETERMINE THE OPTIMAL TIMING OF INITIATING EN CONSIDERING A RISK¿BENEFIT BALANCE AND COST-EFFECTIVENESS. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION IDENTIFIED IN THE ARTICLE.Event Description: SUBSEQUENT FOLLOW-UP WITH THE CORRESPONDING AUTHOR CONFIRMED THAT THERE WAS NO MALFUNCTION OF THE DEVICE DURING ANY OF THE REPORTED PROCEDURES AND THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED ADVERSE EVENTS, INCLUDING DEATH. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE ADDITIONAL INFORMATION FROM THE AUTHOR AND TO PROVIDE ADDITIONAL INFORMATION BASED ON THE LEGAL MANUFACTURER'S FINAL INVESTIGATION. NEW INFORMATION ADDED TO THE FOLLOWING FIELDS: B5 AND H6. BASED ON THE RESULTS OF THE INVESTIGATION, THE RELATIONSHIP BETWEEN THE DEVICE AND THE ADVERSE EVENT CANNOT BE CONFIRMED. THERE WAS NO COMPLAINT REPORTED ON THE SUBJECT DEVICE. THERE IS NO EVIDENCE OF AN OLYMPUS DEVICE MALFUNCTION. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE. Manufacturer Narrative: THIS REPORT IS RELATED TO THE FOLLOWING LINKED PATIENT IDENTIFIERS: (B)(6). THE INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED OR IF ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440759&pc=DSQ | 2916596-2025-04296 | 2024-05-31 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT SUFFERED FROM REFRACTORY AND PROGRESSIVE MULTIPLE ORGAN FAILURE (MOF) LEADING TO DEATH. NO DEVICE RELATED ISSUES WERE MENTIONED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED AND THE DEVICE WOULD NOT RETURN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440779&pc=DSQ | 2916596-2025-04376 | 2025-05-05 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT SUFFERED FROM A REFRACTORY MULTIPLE ORGAN FAILURE (MOF) AND A SEPSIS DEVELOPED UNDER THESE CIRCUMSTANCES. THE PATIENT PASSED AWAY DUE TO SEPSIS UNDER MOF. IT WAS REPORTED THAT DRIVELINE OR DEVICE INFECTION WERE NOT A ROOT CAUSE AND THAT THE PATIENT WAS MULTI-MORBID UNDER PALLIATIVE CARE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETED. Manufacturer Narrative: E3 - OTHER OCCUPATION: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION, REV. B AND THE HEARTMATE 3 PATIENT HANDBOOK, REVISION C, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, SEPSIS, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. BOTH THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS IN REGARD TO PREVENTING INFECTION. THE IFU ALSO PROVIDES SUGGESTED RESPONSES IN THE EVENT OF INFECTION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440799&pc=DSQ | 2916596-2025-04371 | 2021-10-18 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Fall; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT STUMBLED AND FELL DOWN A DEEP STAIR. RESUSCITATION FAILED, AND THE PATIENT PASSED AWAY. DIRECT CAUSE OF DEATH WAS NOT PROVIDED. THERE WERE NO LVAD-RELATED EVENTS THAT MAY HAVE CONTRIBUTED TO THE FALL. THE DEVICE OPERATED AS EXPECTED. NO DEVICE ISSUES WERE MENTIONED. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: CORRECTED DATA: SECTION G2: REPORT SOURCE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440805&pc=DSQ | 2916596-2025-04277 | 2023-03-30 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440828&pc=NIQ | 2124215-2025-45099 | 2025-06-16 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-08 04:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Chest Pain |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT WAS ADMITTED FOR PRIMARY PERCUTANEOUS CORONARY INTERVENTION. THE TARGET LESION WAS LOCATED IN THE RIGHT CORONARY ARTERY. A 2.75 X 48 SYNERGY DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT. HOWEVER, POST DEPLOYMENT, CHEST PAIN WAS SUDDENLY STARTED, AND THE PATIENT WAS SENT TO CARDIAC CARE UNIT (CCU). FEW HOURS LATER, THE PATIENT EXPERIENCE CHEST PAIN AGAIN AND GOT CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS THEN PERFORMED BUT THE PATIENT DIED FROM CARDIAC ARREST. NO AUTOPSY WAS PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22440844&pc=DSQ | 2916596-2025-04220 | 2024-05-29 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR FAILURE (RVF) AND MULTI-SYSTEM ORGAN FAILURE. THE PATIENT SUFFERED FROM A REFRACTORY MULTIORGAN FAILURE (MOF) WITH A RVF IN THE COURSE WHICH WAS RELATED TO THE OUTCOME AT LEAST. LVAD WAS NOT RELATED TO RVF AND OPERATED AS EXPECTED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22441249&pc=DSQ | 2916596-2025-04298 | 2024-01-08 05:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT SUFFERED FROM END STAGE PROGRESSIVE MULTIPLE ORGAN FAILURE WHICH LED TO DEATH. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED AND WOULD NOT RETURN. IT WAS NOTED THAT THE MULTIPLE ORGAN FAILURE WAS DUE TO ADVANCED "HI" DISEASE PROGRESSION.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE MULTIPLE ORGAN FAILURE WAS DUE TO ADVANCED HEART INSUFFICIENCY/HEART FAILURE DISEASE PROGRESSION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE PUMP WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22441684&pc=NIK | 2017865-2025-91649 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-08 04:00 | NIK | GALLANT DR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND UNCONSCIOUS. THE PATIENT WAS TAKEN TO THE EMERGENCY ROOM AND EXPIRED SHORTLY AFTER DUE TO CARDIAC ARREST. THE PHYSICIAN DID NOT BELIEVE THAT A MALFUNCTION OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV) OR RIGHT ATRIAL (RA) LEADS, CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22441705&pc=QZI | 2124215-2025-45508 | 2025-06-13 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-08 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Convulsion/Seizure |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A PULSED FIELD ABLATION (PFA) PROCEDURE WAS COMPLETED SUCCESSFULLY WITH A FARAWAVE CATHETER AND THE PATIENT WAS DISCHARGED THE SAME DAY. TWO DAYS AFTER THE PROCEDURE, THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT (ED) STATING THAT WAS NOT FEELING WELL. A COMPUTED TOMOGRAPHY, CHEST X-RAY AND ECHOCARDIOGRAM WERE PERFORMED. WHILE IN THE ED, THE PATIENT SEIZED AND CODED, AND EVENTUALLY PASSED AWAY. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22441725&pc=NVN | 2017865-2025-91651 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-08 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND UNCONSCIOUS. THE PATIENT WAS TAKEN TO THE EMERGENCY ROOM AND EXPIRED SHORTLY AFTER DUE TO CARDIAC ARREST. THE PHYSICIAN DID NOT BELIEVE THAT A MALFUNCTION OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), RIGHT VENTRICULAR (RV) OR RIGHT ATRIAL (RA) LEADS, CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22442358&pc=QZI | 2124215-2025-45394 | 2025-07-02 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-08 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Pericardial Effusion |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED A PERICARDIAL EFFUSION, A FARAWAVE 31 MM DURING PROCEDURE TO TREAT AN ATRIAL FIBRILLATION (A FIB) WAS USED. TEE-GUIDED AND RADIOLOGICAL TRANSSEPTAL PUNCTURE, LEFT-SIDED PULMONARY VEINS ABLATED, RIGHT UPPER PULMONARY VEIN ABLATED. DURING THE ABLATION OF THE RIGHT UPPER PULMONARY VEIN, A DROP IN BLOOD PRESSURE WAS DETECTED. A CARDIAC ULTRASOUND WAS PERFORMED, AND A PERICARDIAL EFFUSION (PE) WAS DISCOVERED. CONSEQUENTLY, THE ABLATION WAS STOPPED, AND THE PERICARDIAL EFFUSION WAS PUNCTURED. THE PATIENT WAS HOSPITALIZED. THE CATHETER IS EXPECTED TO BE RETURNED.Event Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED A PERICARDIAL EFFUSION, A FARAWAVE 31 MM DURING PROCEDURE TO TREAT AN ATRIAL FIBRILLATION (A FIB) WAS USED. TEE-GUIDED AND RADIOLOGICAL TRANSSEPTAL PUNCTURE, LEFT-SIDED PULMONARY VEINS ABLATED, RIGHT UPPER PULMONARY VEIN ABLATED. DURING THE ABLATION OF THE RIGHT UPPER PULMONARY VEIN, A DROP IN BLOOD PRESSURE WAS DETECTED. A CARDIAC ULTRASOUND WAS PERFORMED, AND A PERICARDIAL EFFUSION (PE) WAS DISCOVERED. CONSEQUENTLY, THE ABLATION WAS STOPPED, AND THE PERICARDIAL EFFUSION WAS PUNCTURED. THE PATIENT WAS HOSPITALIZED. THE CATHETER IS EXPECTED TO BE RETURNED. IT WAS FURTHER REPORTED THAT FOLLOWING THE MEDICAL PROCEDURE IN WHICH THE PERICARDIAL EFFUSION OCCURRED, THE PATIENT DIED. ACCORDING TO THE TREATING PHYSICIAN, THERE WAS NO CAUSAL LINK BETWEEN THE COMPLICATION AND THE CATHETER USED. THE PHYSICIAN ATTRIBUTED THE EVENT TO MANUAL MANIPULATION DURING THE PROCEDURE, AND NOT TO ANY MALFUNCTION OR ISSUE RELATED TO THE DEVICE. BASED ON CURRENT INFORMATION AND THE PHYSICIAN'S ASSESSMENT, THERE IS NO INDICATION OF A DEVICE MALFUNCTION OR PRODUCT-RELATED ISSUE. THE CATHETER WAS USED AS INDICATED AND FUNCTIONED AS INTENDED. Manufacturer Narrative: THE DEVICE HAS NOT BEEN RECEIVED FOR ANALYSIS. UPON RECEIPT AND COMPLETION OF THE FAILURE ANALYSIS OF THE COMPLAINT DEVICE, IF THERE IS ANY FURTHER RELEVANT INFORMATION FROM THAT REVIEW, A SUPPLEMENTAL MEDWATCH WILL BE FILED. Manufacturer Narrative: THE DEVICE HAS NOT BEEN RECEIVED FOR ANALYSIS. UPON RECEIPT AND COMPLETION OF THE FAILURE ANALYSIS OF THE COMPLAINT DEVICE, IF THERE IS ANY FURTHER RELEVANT INFORMATION FROM THAT REVIEW, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22442658&pc=DSQ | 2916596-2025-04256 | 2022-07-18 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS DUE TO UNKNOWN ETIOLOGY. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WILL NOT BE EXPLANTED. Manufacturer Narrative: REPORT SUBTYPE: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22443007&pc=FJS | 1282497-2025-00485 | 2025-03-07 05:00 | Death | COVIDIEN LLC | 2025-07-08 04:00 | FJS | PERITONEAL CATHETER | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: PLI 10: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A PROSPECTIVE COHORT QUASI-EXPERIMENTAL STUDY ENROLLED A TOTAL OF 67 PATIENTS WITH ESRD (END-STAGE RENAL DISEASE) REQUIRING DIALYSIS WHO WERE RANDOMLY ASSIGNED TO EITHER PD (PERITONEAL DIALYSIS) OR HD (HEMODIALYSIS) TREATMENT REGIMENS. FROM THE POOL OF POTENTIAL PARTICIPANTS, 34 PATIENTS UNDERGOING HEMODIALYSIS AND 33 RECEIVING PERITONEAL DIALYSIS WERE IDENTIFIED. HOWEVER, SEVEN INDIVIDUALS (FOUR FROM THE HEMODIALYSIS GROUP AND THREE FROM THE PERITONEAL DIALYSIS GROUP) WERE EXCLUDED FROM THE STUDY, WHICH IS WHY THE FINAL COHORT CONSISTED OF 60 ESRD PATIENTS. ALL 60 PATIENTS HAD STAGE 5 CKD (CHRONIC KIDNEY DISEASE), AS EVIDENCED BY A CREATININE CLEARANCE OF LESS THAN 15 ML/MIN (MILLILITERS PER MINUTE), AND REQUIRED IMMEDIATE, UNSCHEDULED DIALYSIS THERAPY. COMORBIDITIES OBSERVED AMONG ALL PATIENTS ENROLLED IN THE STUDY INCLUDED DIABETES, HYPERTENSION, CARDIOVASCULAR DISEASE, CHRONIC HEART FAILURE (NEW YORK HEART ASSOCIATION [NYHA] STAGE III¿IV), CEREBROVASCULAR DISEASE, AUTOIMMUNE DISEASE, MALIGNANCY, HYPERVOLEMIA, AND UREMIA. NOTABLY, AMONG THE PD GROUP, 5 PATIENTS HAD PRE-EXISTING CHRONIC HEART FAILURE. THE PRIMARY OBJECTIVE WAS TO COMPARE MORTALITY RATES BETWEEN THE TWO GROUPS, WHILE SECONDARY OBJECTIVES INCLUDED COMPARING COMPLICATION RATES AND OVERALL HEALTHCARE COSTS. TWO-CUFF CURLED TENCKHOFF PD CATHETERS WERE INSERTED PERCUTANEOUSLY INTO 30 PATIENTS BY THE DIALYSIS UNIT NEPHROLOGISTS. TWO PD PATIENTS (6.7%) DIED. OVERALL MORTALITY RATES AFTER 3 MONTHS WERE 6.7% FOR PD. PLI 20: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A PROSPECTIVE COHORT QUASI-EXPERIMENTAL STUDY ENROLLED A TOTAL OF 67 PATIENTS WITH ESRD (END-STAGE RENAL DISEASE) REQUIRING DIALYSIS WHO WERE RANDOMLY ASSIGNED TO EITHER PD (PERITONEAL DIALYSIS) OR HD (HEMODIALYSIS) TREATMENT REGIMENS. FROM THE POOL OF POTENTIAL PARTICIPANTS, 34 PATIENTS UNDERGOING HEMODIALYSIS AND 33 RECEIVING PERITONEAL DIALYSIS WERE IDENTIFIED. HOWEVER, SEVEN INDIVIDUALS (FOUR FROM THE HEMODIALYSIS GROUP AND THREE FROM THE PERITONEAL DIALYSIS GROUP) WERE EXCLUDED FROM THE STUDY, WHICH IS WHY THE FINAL COHORT CONSISTED OF 60 ESRD PATIENTS. ALL 60 PATIENTS HAD STAGE 5 CKD (CHRONIC KIDNEY DISEASE), AS EVIDENCED BY A CREATININE CLEARANCE OF LESS THAN 15 ML/MIN (MILLILITERS PER MINUTE), AND REQUIRED IMMEDIATE, UNSCHEDULED DIALYSIS THERAPY. COMORBIDITIES OBSERVED AMONG ALL PATIENTS ENROLLED IN THE STUDY INCLUDED DIABETES, HYPERTENSION, CARDIOVASCULAR DISEASE, CHRONIC HEART FAILURE (NEW YORK HEART ASSOCIATION [NYHA] STAGE III¿IV), CEREBROVASCULAR DISEASE, AUTOIMMUNE DISEASE, MALIGNANCY, HYPERVOLEMIA, AND UREMIA. NOTABLY, AMONG THE PD GROUP, 5 PATIENTS HAD PRE-EXISTING CHRONIC HEART FAILURE. THE PRIMARY OBJECTIVE WAS TO COMPARE MORTALITY RATES BETWEEN THE TWO GROUPS, WHILE SECONDARY OBJECTIVES INCLUDED COMPARING COMPLICATION RATES AND OVERALL HEALTHCARE COSTS. TWO-CUFF CURLED TENCKHOFF PD CATHETERS WERE INSERTED PERCUTANEOUSLY INTO 30 PATIENTS BY THE DIALYSIS UNIT NEPHROLOGISTS. IN THE PD GROUP, INFECTIOUS COMPLICATIONS INCLUDED EXIT-SITE INFECTIONS IN 2 PATIENTS (6.7%) AND PERITONITIS IN 2 PATIENTS (6.7%), WHICH WAS SUCCESSFULLY TREATED WITH ANTIBIOTICS WITHOUT THE NEED FOR CATHETER REMOVAL. BACTEREMIA WAS REPORTED IN 1 PD PATIENT (3.3%). HOSPITALIZATION DAYS POST-DIALYSIS INITIATION AVERAGED 6.2 ± 3.3 FOR PD PATIENTS, WITH AN AVERAGE OF 4 RE-HOSPITALIZATIONS DURING THE OBSERVATION PERIOD. PLI 30: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A PROSPECTIVE COHORT QUASI-EXPERIMENTAL STUDY ENROLLED A TOTAL OF 67 PATIENTS WITH ESRD (END-STAGE RENAL DISEASE) REQUIRING DIALYSIS WHO WERE RANDOMLY ASSIGNED TO EITHER PD (PERITONEAL DIALYSIS) OR HD (HEMODIALYSIS) TREATMENT REGIMENS. FROM THE POOL OF POTENTIAL PARTICIPANTS, 34 PATIENTS UNDERGOING HEMODIALYSIS AND 33 RECEIVING PERITONEAL DIALYSIS WERE IDENTIFIED. HOWEVER, SEVEN INDIVIDUALS (FOUR FROM THE HEMODIALYSIS GROUP AND THREE FROM THE PERITONEAL DIALYSIS GROUP) WERE EXCLUDED FROM THE STUDY, WHICH IS WHY THE FINAL COHORT CONSISTED OF 60 ESRD PATIENTS. ALL 60 PATIENTS HAD STAGE 5 CKD (CHRONIC KIDNEY DISEASE), AS EVIDENCED BY A CREATININE CLEARANCE OF LESS THAN 15 ML/MIN (MILLILITERS PER MINUTE), AND REQUIRED IMMEDIATE, UNSCHEDULED DIALYSIS THERAPY. COMORBIDITIES OBSERVED AMONG ALL PATIENTS ENROLLED IN THE STUDY INCLUDED DIABETES, HYPERTENSION, CARDIOVASCULAR DISEASE, CHRONIC HEART FAILURE (NEW YORK HEART ASSOCIATION [NYHA] STAGE III¿IV), CEREBROVASCULAR DISEASE, AUTOIMMUNE DISEASE, MALIGNANCY, HYPERVOLEMIA, AND UREMIA. NOTABLY, AMONG THE PD GROUP, 5 PATIENTS HAD PRE-EXISTING CHRONIC HEART FAILURE. THE PRIMARY OBJECTIVE WAS TO COMPARE MORTALITY RATES BETWEEN THE TWO GROUPS, WHILE SECONDARY OBJECTIVES INCLUDED COMPARING COMPLICATION RATES AND OVERALL HEALTHCARE COSTS. TWO-CUFF CURLED TENCKHOFF PD CATHETERS WERE INSERTED PERCUTANEOUSLY INTO 30 PATIENTS BY THE DIALYSIS UNIT NEPHROLOGISTS. MECHANICAL COMPLICATIONS IN PD PATIENTS INCLUDED CATHETER TIP MIGRATION, DIALYSATE LEAKS, AND CATHETEROBSTRUCTION. IN THE PD GROUP, 4 PATIENTS (13.3%) EXPERIENCED CATHETER LEAKAGE, AND 4 PATIENTS (13.3%) EXPERIENCED CATHETER OBSTRUCTION AS THE PRIMARY MECHANICAL COMPLICATIONS AFTER 2 WEEKS OF THERAPY, BUT THESE WERE SUCCESSFULLY MANAGED WITHOUT TREATMENT INTERRUPTION. THERE WAS NO REPORTED PATIENT OUTCOME. PLI 40: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A PROSPECTIVE COHORT QUASI-EXPERIMENTAL STUDY ENROLLED A TOTAL OF 67 PATIENTS WITH ESRD (END-STAGE RENAL DISEASE) REQUIRING DIALYSIS WHO WERE RANDOMLY ASSIGNED TO EITHER PD (PERITONEAL DIALYSIS) OR HD (HEMODIALYSIS) TREATMENT REGIMENS. FROM THE POOL OF POTENTIAL PARTICIPANTS, 34 PATIENTS UNDERGOING HEMODIALYSIS AND 33 RECEIVING PERITONEAL DIALYSIS WERE IDENTIFIED. HOWEVER, SEVEN INDIVIDUALS (FOUR FROM THE HEMODIALYSIS GROUP AND THREE FROM THE PERITONEAL DIALYSIS GROUP) WERE EXCLUDED FROM THE STUDY, WHICH IS WHY THE FINAL COHORT CONSISTED OF 60 ESRD PATIENTS. ALL 60 PATIENTS HAD STAGE 5 CKD (CHRONIC KIDNEY DISEASE), AS EVIDENCED BY A CREATININE CLEARANCE OF LESS THAN 15 ML/MIN (MILLILITERS PER MINUTE), AND REQUIRED IMMEDIATE, UNSCHEDULED DIALYSIS THERAPY. COMORBIDITIES OBSERVED AMONG ALL PATIENTS ENROLLED IN THE STUDY INCLUDED DIABETES, HYPERTENSION, CARDIOVASCULAR DISEASE, CHRONIC HEART FAILURE (NEW YORK HEART ASSOCIATION [NYHA] STAGE III¿IV), CEREBROVASCULAR DISEASE, AUTOIMMUNE DISEASE, MALIGNANCY, HYPERVOLEMIA, AND UREMIA. NOTABLY, AMONG THE PD GROUP, 5 PATIENTS HAD PRE-EXISTING CHRONIC HEART FAILURE. THE PRIMARY OBJECTIVE WAS TO COMPARE MORTALITY RATES BETWEEN THE TWO GROUPS, WHILE SECONDARY OBJECTIVES INCLUDED COMPARING COMPLICATION RATES AND OVERALL HEALTHCARE COSTS. ALL CENTRAL VENOUS CATHETERS (CVCS) USED FOR HEMODIALYSIS WERE 11.5 FRENCH DOUBLE-LUMEN MAHURKAR CATHETERS. THREE HD PATIENTS (10.0%) DIED. PLI 50: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A PROSPECTIVE COHORT QUASI-EXPERIMENTAL STUDY ENROLLED A TOTAL OF 67 PATIENTS WITH ESRD (END-STAGE RENAL DISEASE) REQUIRING DIALYSIS WHO WERE RANDOMLY ASSIGNED TO EITHER PD (PERITONEAL DIALYSIS) OR HD (HEMODIALYSIS) TREATMENT REGIMENS. FROM THE POOL OF POTENTIAL PARTICIPANTS, 34 PATIENTS UNDERGOING HEMODIALYSIS AND 33 RECEIVING PERITONEAL DIALYSIS WERE IDENTIFIED. HOWEVER, SEVEN INDIVIDUALS (FOUR FROM THE HEMODIALYSIS GROUP AND THREE FROM THE PERITONEAL DIALYSIS GROUP) WERE EXCLUDED FROM THE STUDY, WHICH IS WHY THE FINAL COHORT CONSISTED OF 60 ESRD PATIENTS. ALL 60 PATIENTS HAD STAGE 5 CKD (CHRONIC KIDNEY DISEASE), AS EVIDENCED BY A CREATININE CLEARANCE OF LESS THAN 15 ML/MIN (MILLILITERS PER MINUTE), AND REQUIRED IMMEDIATE, UNSCHEDULED DIALYSIS THERAPY. COMORBIDITIES OBSERVED AMONG ALL PATIENTS ENROLLED IN THE STUDY INCLUDED DIABETES, HYPERTENSION, CARDIOVASCULAR DISEASE, CHRONIC HEART FAILURE (NEW YORK HEART ASSOCIATION [NYHA] STAGE III¿IV), CEREBROVASCULAR DISEASE, AUTOIMMUNE DISEASE, MALIGNANCY, HYPERVOLEMIA, AND UREMIA. NOTABLY, AMONG THE PD GROUP, 5 PATIENTS HAD PRE-EXISTING CHRONIC HEART FAILURE. THE PRIMARY OBJECTIVE WAS TO COMPARE MORTALITY RATES BETWEEN THE TWO GROUPS, WHILE SECONDARY OBJECTIVES INCLUDED COMPARING COMPLICATION RATES AND OVERALL HEALTHCARE COSTS. ALL CENTRAL VENOUS CATHETERS (CVCS) USED FOR HEMODIALYSIS WERE 11.5 FRENCH DOUBLE-LUMEN MAHURKAR CATHETERS. FOR HEMODIALYSIS (HD) PATIENTS, CATHETER-RELATED INFECTIONS, INCLUDING CVC-RELATED BACTEREMIA OR SEPSIS, WERE DIAGNOSED WHEN BLOOD CULTURES PERFORMED IN RESPONSE TO A SIGNIFICANT INCREASE IN C-REACTIVE PROTEIN (CRP) OR PATIENT DETERIORATION YIELDED POSITIVE RESULTS, OR WHEN BLOOD CULTURES DRAWN DIRECTLY FROM THE CATHETER WERE POSITIVE ALONG WITH CLINICAL AND LABORATORY EVIDENCE OF INFECTION AND SEPSIS. INFECTIOUS COMPLICATIONS INCLUDED EXIT-SITE INFECTIONS IN 6 PATIENTS (20%) AND BACTEREMIA IN 7 PATIENTS (23.3%) (P = 0.008). CATHETER INFECTION OCCURRED IN 6 PATIENTS (20.0%) (P <(><<)> 0.01), NECESSITATING CATHETER AND SITE REPLACEMENT IN 4 PATIENTS (13.3%). PATIENTS WERE HOSPITALIZED FOR AN AVERAGE OF 13.6 ± 4.1 DAYS POST-DIALYSIS INITIATION, WITH AN AVERAGE OF 5 RE-HOSPITALIZATIONS DURING THE OBSERVATION PERIOD. PLI 60: ACCORDING TO THE LITERATURE SOURCE OF STUDY, A PROSPECTIVE COHORT QUASI-EXPERIMENTAL STUDY ENROLLED A TOTAL OF 67 PATIENTS WITH ESRD (END-STAGE RENAL DISEASE) REQUIRING DIALYSIS WHO WERE RANDOMLY ASSIGNED TO EITHER PD (PERITONEAL DIALYSIS) OR HD (HEMODIALYSIS) TREATMENT REGIMENS. FROM THE POOL OF POTENTIAL PARTICIPANTS, 34 PATIENTS UNDERGOING HEMODIALYSIS AND 33 RECEIVING PERITONEAL DIALYSIS WERE IDENTIFIED. HOWEVER, SEVEN INDIVIDUALS (FOUR FROM THE HEMODIALYSIS GROUP AND THREE FROM THE PERITONEAL DIALYSIS GROUP) WERE EXCLUDED FROM THE STUDY, WHICH IS WHY THE FINAL COHORT CONSISTED OF 60 ESRD PATIENTS. ALL 60 PATIENTS HAD STAGE 5 CKD (CHRONIC KIDNEY DISEASE), AS EVIDENCED BY A CREATININE CLEARANCE OF LESS THAN 15 ML/MIN (MILLILITERS PER MINUTE), AND REQUIRED IMMEDIATE, UNSCHEDULED DIALYSIS THERAPY. COMORBIDITIES OBSERVED AMONG ALL PATIENTS ENROLLED IN THE STUDY INCLUDED DIABETES, HYPERTENSION, CARDIOVASCULAR DISEASE, CHRONIC HEART FAILURE (NEW YORK HEART ASSOCIATION [NYHA] STAGE III¿IV), CEREBROVASCULAR DISEASE, AUTOIMMUNE DISEASE, MALIGNANCY, HYPERVOLEMIA, AND UREMIA. NOTABLY, AMONG THE PD GROUP, 5 PATIENTS HAD PRE-EXISTING CHRONIC HEART FAILURE. THE PRIMARY OBJECTIVE WAS TO COMPARE MORTALITY RATES BETWEEN THE TWO GROUPS, WHILE SECONDARY OBJECTIVES INCLUDED COMPARING COMPLICATION RATES AND OVERALL HEALTHCARE COSTS. ALL CENTRAL VENOUS CATHETERS (CVCS) USED FOR HEMODIALYSIS WERE 11.5 FRENCH DOUBLE-LUMEN MAHURKAR CATHETERS. IN THE HD GROUP, CATHETER OBSTRUCTION (PARTIAL OR TOTAL) OCCURRED AS THE INITIAL COMPLICATION IN 3 PATIENTS (10%) AND WAS RESOLVED THROUGH CATHETER MANIPULATION. CATHETER LEAKAGE WAS REPORTED IN 5 PATIENTS (16.7%). THERE WAS NO REPORTED PATIENT OUTCOME. URGENT UNPLANNED PERITONEAL DIALYSIS VERSUS HEMODIALYSIS IN END-STAGE RENAL DISEASE: COSTS AND OUTCOMES: ABDULLAH ALHWIESH, 2025, THERAPEUTIC APHERESIS AND DIALYSIS. Manufacturer Narrative: ABDULLAH ALHWIESH, IBRAHIEM SAEED ABDUL-RAHMAN, MOHAMMED AHMED NASR EL-DIN, MOAZ ABDULGALIL, MOHAMMED ABDELMEGEED SAKR, KHADIJA M. A LSHEHABI, NADIA AL-AUDAH URGENT UNPLANNED PERITONEAL DIALYSIS VERSUS HEMODIALYSIS IN END-STAGE RENAL DISEASE: COSTS AND OUTCOMES THER APHER DIAL. 2025;29(3):437¿46. HTTPS://DOI.ORG/10. 1111/1744-9987.70007 D10 - CONCOMITANT PRODUCT: PERITONEAL CATHETER LOT UNKNOWN PERITONEAL CATHETER LOT UNKNOWN DIALYSIS UNKNOWN LOT UNKNOWN DIALYSIS UNKNOWN LOT UNKNOWN DIALYSIS UNKNOWN LOT UNKNOWN MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22443738&pc=MAE | 2135147-2025-03732 | 2017-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-08 04:00 | MAE | AMPLATZER PICCOLO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P020024 |  | 1 | Event Description: THE ARTICLE, "IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO EVALUATE THE EFFECT OF TRANSCATHETER PATENT DUCTUS ARTERIOSUS CLOSURE (TCDC) IN PREMATURE INFANTS ON PULMONARY EDEMA AND RESPIRATORY STATUS, AND MORE SPECIFICALLY TO CHARACTERIZE THE SUBSET OF PATIENTS WHO HAVE SIGNIFICANTLY IMPROVED CHEST X-RAY (CXR) IMMEDIATELY AFTER TCDC. DEVICES INCLUDED IN THE STUDY WERE AMPLATZER PICCOLO OCCLUDER, MEDTRONIC MICROVASCULAR PLUG, COVIDIEN MICROVASCULAR PLUG, AMPLATZER VASCULAR PLUG, AND AMPLATZER VASCULAR PLUG II. THE ARTICLE CONCLUDED CHEST X-RAY HAZINESS IMPROVED AFTER TCDC, WITH SMALLER INFANTS AND EARLIER CLOSURE HAVING MORE IMPROVEMENT. INFANTS WITH LUNG DISEASE HAD LESS NOTICEABLE IMPROVED EDEMA, INDICATING THE DIFFICULTY TO ASSESS THE HEMODYNAMIC SIGNIFICANCE OF THEIR PDA PRIOR TO CLOSURE. FURTHER STUDIES ARE NEEDED TO IDENTIFY WHICH NEONATES BENEFIT MOST FROM TCDC. [THE PRIMARY AND CORRESPONDING AUTHOR WAS AHMAD CHMAISSE, DIVISION OF CARDIOLOGY, RILEY HOSPITAL FOR CHILDREN, INDIANA UNIVERSITY SCHOOL OF MEDICINE, INDIANAPOLIS, IN, UNITED STATES, WITH CORRESPONDING EMAIL: ACHMAIS@IU.EDU]. THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2017 TO JUNE 2021. A TOTAL OF 68 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 59.5% RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE, WEIGHT, AND MAJORITY GENDER WAS NOT REPORTED. COMORBIDITIES INCLUDED PULMONARY HYPERTENSION, PULMONARY EDEMA, PRETERM BIRTH, AND PATENT DUCTUS ARTERIOSUS. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PULMONARY HYPERTENSION, PULMONARY EDEMA, PRETERM BIRTH, AND PATENT DUCTUS ARTERIOSUS. ONE OF THE COMPLICATION REPORTED WAS DEATH; THIS COMPLICATION IIS ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22444271&pc=KRD | 2135147-2025-03734 | 2017-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-08 04:00 | KRD | AMPLATZER VASCULAR PLUG | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K031810 |  | 1 | Event Description: THE ARTICLE, "IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO EVALUATE THE EFFECT OF TRANSCATHETER PATENT DUCTUS ARTERIOSUS CLOSURE (TCDC) IN PREMATURE INFANTS ON PULMONARY EDEMA AND RESPIRATORY STATUS, AND MORE SPECIFICALLY TO CHARACTERIZE THE SUBSET OF PATIENTS WHO HAVE SIGNIFICANTLY IMPROVED CHEST X-RAY (CXR) IMMEDIATELY AFTER TCDC. DEVICES INCLUDED IN THE STUDY WERE AMPLATZER PICCOLO OCCLUDER, MEDTRONIC MICROVASCULAR PLUG, COVIDIEN MICROVASCULAR PLUG, AMPLATZER VASCULAR PLUG, AND AMPLATZER VASCULAR PLUG II. THE ARTICLE CONCLUDED CHEST X-RAY HAZINESS IMPROVED AFTER TCDC, WITH SMALLER INFANTS AND EARLIER CLOSURE HAVING MORE IMPROVEMENT. INFANTS WITH LUNG DISEASE HAD LESS NOTICEABLE IMPROVED EDEMA, INDICATING THE DIFFICULTY TO ASSESS THE HEMODYNAMIC SIGNIFICANCE OF THEIR PDA PRIOR TO CLOSURE. FURTHER STUDIES ARE NEEDED TO IDENTIFY WHICH NEONATES BENEFIT MOST FROM TCDC. [THE PRIMARY AND CORRESPONDING AUTHOR WAS AHMAD CHMAISSE, DIVISION OF CARDIOLOGY, RILEY HOSPITAL FOR CHILDREN, INDIANA UNIVERSITY SCHOOL OF MEDICINE, INDIANAPOLIS, IN, UNITED STATES, WITH CORRESPONDING EMAIL: ACHMAIS@IU.EDU]. THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2017 TO JUNE 2021. A TOTAL OF 68 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 59.5% RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE, WEIGHT, AND MAJORITY GENDER WAS NOT REPORTED. COMORBIDITIES INCLUDED PULMONARY HYPERTENSION, PULMONARY EDEMA, PRETERM BIRTH, AND PATENT DUCTUS ARTERIOSUS. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2; DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PULMONARY HYPERTENSION, PULMONARY EDEMA, PRETERM BIRTH, AND PATENT DUCTUS ARTERIOSUS. ONE OF THE COMPLICATION REPORTED WAS DEATH; THIS COMPLICATION IS ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22444709&pc=KRD | 2135147-2025-03736 | 2017-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-08 04:00 | KRD | AMPLATZER VASCULAR PLUG II | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K071699 |  | 1 | Event Description: THE ARTICLE, "IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO EVALUATE THE EFFECT OF TRANSCATHETER PATENT DUCTUS ARTERIOSUS CLOSURE (TCDC) IN PREMATURE INFANTS ON PULMONARY EDEMA AND RESPIRATORY STATUS, AND MORE SPECIFICALLY TO CHARACTERIZE THE SUBSET OF PATIENTS WHO HAVE SIGNIFICANTLY IMPROVED CHEST X-RAY (CXR) IMMEDIATELY AFTER TCDC. DEVICES INCLUDED IN THE STUDY WERE AMPLATZER PICCOLO OCCLUDER, MEDTRONIC MICROVASCULAR PLUG, COVIDIEN MICROVASCULAR PLUG, AMPLATZER VASCULAR PLUG, AND AMPLATZER VASCULAR PLUG II. THE ARTICLE CONCLUDED CHEST X-RAY HAZINESS IMPROVED AFTER TCDC, WITH SMALLER INFANTS AND EARLIER CLOSURE HAVING MORE IMPROVEMENT. INFANTS WITH LUNG DISEASE HAD LESS NOTICEABLE IMPROVED EDEMA, INDICATING THE DIFFICULTY TO ASSESS THE HEMODYNAMIC SIGNIFICANCE OF THEIR PDA PRIOR TO CLOSURE. FURTHER STUDIES ARE NEEDED TO IDENTIFY WHICH NEONATES BENEFIT MOST FROM TCDC. [THE PRIMARY AND CORRESPONDING AUTHOR WAS AHMAD CHMAISSE, DIVISION OF CARDIOLOGY, RILEY HOSPITAL FOR CHILDREN, INDIANA UNIVERSITY SCHOOL OF MEDICINE, INDIANAPOLIS, IN, UNITED STATES, WITH CORRESPONDING EMAIL: ACHMAIS@IU.EDU]. THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2017 TO JUNE 2021. A TOTAL OF 68 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 59.5% RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE, WEIGHT, AND MAJORITY GENDER WAS NOT REPORTED. COMORBIDITIES INCLUDED PULMONARY HYPERTENSION, PULMONARY EDEMA, PRETERM BIRTH, AND PATENT DUCTUS ARTERIOSUS. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PULMONARY HYPERTENSION, PULMONARY EDEMA, PRETERM BIRTH, AND PATENT DUCTUS ARTERIOSUS. ONE OF THE COMPLICATION REPORTED WAS DEATH; THIS COMPLICATION IS ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B3: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22446640&pc=DSQ | 2916596-2025-04504 | 2025-06-17 04:00 | Death | THORATEC CORPORATION | 2025-07-08 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH A KNOWN HISTORY OF LOW FLOW ALARMS, DOWN TO AS LOW AT 1.8 LITERS PER MINUTE (LPM), WAS SEEN IN THE EMERGENCY DEPARTMENT AND DISCHARGED FOR GASTROENTERITIS. LABS WITH SUPRATHERAPEUTIC INTERNATIONAL NORMALIZE RATIO WERE SEEN AND THE PATIENT LEFT THE EMERGENCY DEPARTMENT STABLE. THE PATIENT WAS BROUGHT INTO THE EMERGENCY DEPARTMENT BY MEDICAL SERVICES AT APPROXIMATELY 22:50, TWO HOURS AFTER BEING DISCHARGED, FOR CARDIAC ARREST. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) HAD BEEN ALARMING FOR 20 MINUTES. RETURN OF SPONTANEOUS CIRCULATION (ROSC) WAS ACHIEVED WHILE IN THE FIELD. UPON ARRIVAL, THE PATIENT WAS IN PULSELESS ELECTRICAL ACTIVITY AND ADVANCED CARDIAC LIFE SUPPORT WAS RESTARTED. THE LVAD FLOW WAS 0.3 LPW. INTERROGATION WAS ABLE TO GO BACK TO 22:30, WHERE IT SHOWED THAT PATIENT HAD BEEN HAVING LOW FLOW ALARMS 1.8-2.4 LPM AND AT ABOUT 23:00, THE FLOWS BECAME <1.0. IN THE EMERGENCY DEPARTMENT, A FOCUSED ASSESSMENT WITH SONOGRAPHY AND A POINT OF CARE ULTRASOUND WERE PERFORMED AND REVEALED NO OBVIOUS BLEED. STAFF WERE ABLE TO ACHIEVE ROSC 4 TIMES, BUT ON 5TH TIME NO SUSTAINABLE RHYTHM WAS ACHIEVED AND TIME OF DEATH WAS CALLED. LOG FILES WERE OBTAINED BEFORE THE SYSTEM CONTROLLER WAS DISCONNECTED AND WERE SUBMITTED FOR REVIEW. THE EVENT LOG FILE CAPTURED LOW FLOW EVENTS ON (B)(6) 2025. THERE ARE ALSO SEVERAL LOW FLOW FLAGS WHICH DID NOT PERSIST LONG ENOUGH TO TRIGGER LOW FLOW ALARMS. THE PERIODIC LOG FILE CAPTURED INTERMITTENT LOW FLOW EVENTS BETWEEN (B)(6) 2025 AND (B)(6) 2025. THERE WERE NO ISSUES NOTED IN THE PUMP LOG FILES.Event Description: IT WAS ADDITIONALLY REPORTED THAT IT WAS UNKNOWN IF THE DEATH WAS CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION H6: HEALTH EFFECT - IMPACT CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. FURTHER, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW HAZARD ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THE ALARMS COULD NOT BE CONCLUSIVELY DETERMINED. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. THE CONTROLLER EVENT LOG FILES CONTAINED EVENTS FROM 17JUN2025 THROUGH 18JUN2025. SUSTAINED AND INTERMITTENT LOW FLOW HAZARD ALARMS WERE CAPTURED THROUGHOUT THE FILE WITH ESTIMATED FLOW REACHING AS LOW AS 0.3 LITERS PER MINUTE (LPM). A SPECIFIC CAUSE FOR THESE LOW FLOW EVENTS COULD NOT BE CONCLUSIVELY DETERMINED; HOWEVER, THE LOW FLOW EVENTS OCCURRING AFTER 22:50:00 ON 17JUN2025 APPEAR TO BE CONSISTENT WITH THE REPORTED CARDIAC ARREST. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR MLP-013432 WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. A ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE. THIS SECTION ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿HEARTMATE TOUCH COMMUNICATION SYSTEM¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LPM, A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. THIS SECTION ALSO OUTLINES SITUATIONS THAT CAN RESULT IN A LOW FLOW ALARM AND FURTHER EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. FOLLOWING SOFTWARE UPGRADES, THE HM3 LVAS POCKET GUIDES TO ALARMS FOR PATIENTS, REV. A, AND CLINICIANS, REV. A, EXPLAIN THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.0 LPM. SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING," AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿ OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22451386&pc=DYB | 2029046-2025-02220 | 2025-06-17 04:00 | Death | BIOSENSE WEBSTER INC | 2025-07-08 04:00 | DYB | VIZIGO | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiac Tamponade; Vascular Dissection | K170997 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT A CARDIAC ABLATION PROCEDURE WITH A VIZIGO AND A QDOT MICRO CATHETER. THE PATIENT EXPERIENCED ACCUMULATION OF BLOOD IN THE ABDOMEN TREATED WITH BLOOD TRANSFUSION FOR THE BLEED AND LATER THE PATIENT EXPERIENCED CARDIAC TAMPONADE AND EXPIRED TWO DAYS AFTER THE PROCEDURE. IT WAS INITIALLY REPORTED THAT AFTER THE PROCEDURE, AN ACCUMULATION OF BLOOD IN THE ABDOMEN WAS FOUND WHEN THE DRAPE WAS REMOVED. INTRACARDIAC WAS CHECKED BY ECHOCARDIOGRAPHY, AND NO PERICARDIAL EFFUSION WAS CONFIRMED. AFTER THE PATIENT LEFT THE PROCEDURE ROOM A BLOOD TRANSFUSION WAS PERFORMED. THE PHYSICIAN'S OPINION WAS THAT THERE WAS AN ISSUE WITH SOME ASPECT OF THE WORKFLOW DURING THE PUNCTURE BEFORE THE USE OF BWI PRODUCTS. ON 23-JUN-2025, ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT IN ADDITION TO THE ABDOMINAL BLOOD ACCUMULATION OBSERVED IMMEDIATELY AFTER THE PROCEDURE, CARDIAC TAMPONADE WAS CONFIRMED AT A LATER DATE AND THE PATIENT PASSED AWAY IN THE HOSPITAL 2 DAYS AFTER THE INDEX PROCEDURE. THE DETAILS WERE UNCLEAR, BUT THE PHYSICIAN SUSPECTED THAT THE ISSUE OCCURRED DURING THE PUNCTURE PRIOR TO THE USE OF BWI DEVICES. THE PHYSICIAN'S ASSESSMENT OF THE HEALTH HAZARD WAS SERIOUS. THE PHYSICIAN'S OPINIONS ON THE RELATIONSHIP BETWEEN THE EVENT AND THE PRODUCT WAS IMMEDIATELY AFTER THE PROCEDURE, INTRACARDIAC WAS CHECKED BY ECHOCARDIOGRAPHY, NO PERICARDIAL EFFUSION WAS CONFIRMED. NO ABNORMALITIES WERE OBSERVED WITH THE PRODUCT PRIOR TO OR DURING USE OF THE PRODUCTS. Manufacturer Narrative: ON 22-JUL-2025, ADDITIONAL INFORMATION ABOUT THE PATIENT AND EVENT WERE RECEIVED. IT WAS REPORTED THE PROCEDURAL ACTIVATED CLOTTING TIME (ACT) WAS 300 SECONDS. SEVERAL CATHETER EXCHANGES OCCURRED DURING THE PROCEDURE. THE ENERGY DELIVERED WAS RADIOFREQUENCY (RF). THERE WAS NO EVIDENCE OF STEAM POP. TRANSSEPTAL PUNCTURE WAS PERFORMED WITH UNKNOWN NEEDLE. THE VIZIGO SHEATH WAS NOT USED FOR TRANSSEPTAL ACCESS. NO ERROR MESSAGES OBSERVED ON BIOSENSE WEBSTER EQUIPMENT DURING THE PROCEDURE. THE CORRECT CATHETER SETTINGS WERE SELECTED ON THE GENERATOR. THERE WERE NO ISSUES WITH FLOW RATE CHANGES AT THE BEGINNING OF THE PROCEDURE. FORCE VISUALIZATION FEATURES USED INCLUDED VECTOR, REALTIME GRAPH. VISITAG MODULE PARAMETERS FOR STABILITY WERE RANGE: 3MM, TIME: 3SEC, FORCE OVER TIME (FOT): 25%, TAG SIZE: 2MM. NO ADDITIONAL FILTERS WRE USED. COLOR OPTIONS USED WERE TAG INDEX. THE PHYSICIAN'S OPINION ON THE CAUSE OF DEATH WAS CARDIAC TAMPONADE OR HEMORRHAGIC SHOCK BEFORE THE CARDIAC TAMPONADE WAS CONFIRMED. IT WAS REPORTED THAT DURING THE PROCEDURE, A DECREASE IN BLOOD PRESSURE WAS OBSERVED; HOWEVER, AT THAT TIME, NO PERICARDIAL EFFUSION WAS CONFIRMED ON ULTRASOUND. AFTER THE PROCEDURE, WHEN THE COVERING CLOTH WAS REMOVED, A SUBCUTANEOUS BLOOD ACCUMULATION IN THE ABDOMEN WAS NOTED, AND DRAINAGE AND TRANSFUSION WERE PERFORMED. THE PHYSICIAN SUSPECTED THAT VASCULAR DISSECTION DURING THE PUNCTURE MAY HAVE BEEN THE CAUSE. ON THE MORNING OF 2 DAYS AFTER THE PROCEDURE, CARDIAC TAMPONADE WAS FOUND, AND THE PATIENT SUBSEQUENTLY DIED. BASED ON THE ADDITIONAL INFORMATION RECEIVED, THE VIZIGO SHEATH IS THE MOST SUSPECTED DEVICE FOR VASCULAR HEMORRHAGE / DISSECTION SINCE THE VIZIGO COMES INTO DIRECT CONTACT WITH VASCULAR TISSUE. PLEASE CONSIDER THE CARDIAC TAMPONADE (E0605) AS REMOVED FROM FIELD H6. HEALTH EFFECT - CLINICAL CODE AND VASCULAR DISSECTION (E0515) AND HEMORRHAGE/BLEEDING (E0506) HAVE BEEN ADDED TO FIELD H6. HEALTH EFFECT - CLINICAL CODE. THE MOST SUSPECTED DEVICE FOR THE CARDIAC TAMPONADE CONTINUES TO BE QDOT MICRO ABLATION CATHETER AS IT DELIVERS ENERGY DIRECTLY TO CARDIAC TISSUE. BECAUSE THE CAUSE OF DEATH IS UNCLEAR WHETHER IT IS FROM HEMORRHAGE OR CARDIAC TAMPONADE, BOTH THE VIZIGO AND THE QDOT MICRO ABLATION CATHETER WILL CONTINUE TO CAPTURE THE DEATH. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REF. # (B)(4). Manufacturer Narrative: DEVICE EVALUATION DETAILS: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER LOT 60000518 AND NO INTERNAL ACTION RELATED TO THE COMPLAINT WAS FOUND DURING THE REVIEW. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REF.# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22451852&pc=NIQ | 2124215-2025-44582 | 2025-04-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-08 04:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IN (B)(6) 2020, THE SUBJECT PRESENTED WITH MYOCARDIAL INFARCTION AND UNSTABLE ANGINA AND WAS REFERRED FOR CARDIAC CATHETERIZATION. THE TARGET LESION WAS LOCATED IN THE MIDDLE RIGHT CORONARY ARTERY (RCA) WITH UNKNOWN% STENOSIS AND WAS 30 MM LONG, WITH A REFERENCE VESSEL DIAMETER OF 4.0 MM. THE TARGET LESION WAS TREATED WITH PRE-DILATATION AND FOLLOWED BY THE PLACEMENT OF A 4.00 MM X 32 MM SYNERGY STENT SYSTEM. FOLLOWING POST-DILATATION, THE RESIDUAL STENOSIS WAS NOTED TO BE 0%. FOUR DAYS LATER, THE SUBJECT WAS DISCHARGED ON CLOPIDOGREL AND OTHER ANTIPLATELET MEDICATION. IN (B)(6) 2025, THE SUBJECT DIED, AND THE PRIMARY CAUSE OF DEATH WAS UNKNOWN. THE SUBJECT WAS ON CLOPIDOGREL AND OTHER ANTIPLATELET MEDICATION. THERE WAS NO ACTION TAKEN TO TREAT THE EVENT AND IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: B2 DATE OF DEATH: ESTIMATED BASED ON THE EVENT DATE AS THE EXACT DATE WAS UNKNOWN. B3 DATE OF EVENT: ESTIMATED BASED ON THE EVENT DATE AS THE EXACT DATE WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22452497&pc=OAE | 2029046-2025-02221 | 2025-06-17 04:00 | Death | BIOSENSE WEBSTER INC | 2025-07-08 04:00 | OAE | QDOT MICRO | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade | P210027 |  | 1 | Event Description: IT WAS REPORTED A PATIENT UNDERWENT A CARDIAC ABLATION PROCEDURE WITH A VIZIGO AND A QDOT MICRO CATHETER. THE PATIENT EXPERIENCED ACCUMULATION OF BLOOD IN THE ABDOMEN TREATED WITH BLOOD TRANSFUSION FOR THE BLEED AND LATER THE PATIENT EXPERIENCED CARDIAC TAMPONADE AND EXPIRED TWO DAYS AFTER THE PROCEDURE. IT WAS INITIALLY REPORTED THAT AFTER THE PROCEDURE, AN ACCUMULATION OF BLOOD IN THE ABDOMEN WAS FOUND WHEN THE DRAPE WAS REMOVED. INTRACARDIAC WAS CHECKED BY ECHOCARDIOGRAPHY, AND NO PERICARDIAL EFFUSION WAS CONFIRMED. AFTER THE PATIENT LEFT THE PROCEDURE ROOM A BLOOD TRANSFUSION WAS PERFORMED. THE PHYSICIAN'S OPINION WAS THAT THERE WAS AN ISSUE WITH SOME ASPECT OF THE WORKFLOW DURING THE PUNCTURE BEFORE THE USE OF BWI PRODUCTS. ON (B)(6) 2025, ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT IN ADDITION TO THE ABDOMINAL BLOOD ACCUMULATION OBSERVED IMMEDIATELY AFTER THE PROCEDURE, CARDIAC TAMPONADE WAS CONFIRMED AT A LATER DATE AND THE PATIENT PASSED AWAY IN THE HOSPITAL 2 DAYS AFTER THE INDEX PROCEDURE. THE DETAILS WERE UNCLEAR, BUT THE PHYSICIAN SUSPECTED THAT THE ISSUE OCCURRED DURING THE PUNCTURE PRIOR TO THE USE OF BWI DEVICES. THE PHYSICIAN'S ASSESSMENT OF THE HEALTH HAZARD WAS SERIOUS. THE PHYSICIAN'S OPINIONS ON THE RELATIONSHIP BETWEEN THE EVENT AND THE PRODUCT WAS IMMEDIATELY AFTER THE PROCEDURE, INTRACARDIAC WAS CHECKED BY ECHOCARDIOGRAPHY, NO PERICARDIAL EFFUSION WAS CONFIRMED. NO ABNORMALITIES WERE OBSERVED WITH THE PRODUCT PRIOR TO OR DURING USE OF THE PRODUCTS. Manufacturer Narrative: ON 22-JUL-2025, ADDITIONAL INFORMATION ABOUT THE PATIENT AND EVENT WERE RECEIVED. IT WAS REPORTED THE PROCEDURAL ACTIVATED CLOTTING TIME (ACT) WAS 300 SECONDS. SEVERAL CATHETER EXCHANGES OCCURRED DURING THE PROCEDURE. THE ENERGY DELIVERED WAS RADIOFREQUENCY (RF). THERE WAS NO EVIDENCE OF STEAM POP. TRANSSEPTAL PUNCTURE WAS PERFORMED WITH UNKNOWN NEEDLE. THE VIZIGO SHEATH WAS NOT USED FOR TRANSSEPTAL ACCESS. NO ERROR MESSAGES OBSERVED ON BIOSENSE WEBSTER EQUIPMENT DURING THE PROCEDURE. THE CORRECT CATHETER SETTINGS WERE SELECTED ON THE GENERATOR. THERE WERE NO ISSUES WITH FLOW RATE CHANGES AT THE BEGINNING OF THE PROCEDURE. FORCE VISUALIZATION FEATURES USED INCLUDED VECTOR, REALTIME GRAPH. VISITAG MODULE PARAMETERS FOR STABILITY WERE RANGE: 3MM, TIME: 3SEC, FORCE OVER TIME (FOT): 25%, TAG SIZE: 2MM. NO ADDITIONAL FILTERS WRE USED. COLOR OPTIONS USED WERE TAG INDEX. THE PHYSICIAN'S OPINION ON THE CAUSE OF DEATH WAS CARDIAC TAMPONADE OR HEMORRHAGIC SHOCK BEFORE THE CARDIAC TAMPONADE WAS CONFIRMED. IT WAS REPORTED THAT DURING THE PROCEDURE, A DECREASE IN BLOOD PRESSURE WAS OBSERVED; HOWEVER, AT THAT TIME, NO PERICARDIAL EFFUSION WAS CONFIRMED ON ULTRASOUND. AFTER THE PROCEDURE, WHEN THE COVERING CLOTH WAS REMOVED, A SUBCUTANEOUS BLOOD ACCUMULATION IN THE ABDOMEN WAS NOTED, AND DRAINAGE AND TRANSFUSION WERE PERFORMED. THE PHYSICIAN SUSPECTED THAT VASCULAR DISSECTION DURING THE PUNCTURE MAY HAVE BEEN THE CAUSE. ON THE MORNING OF 2 DAYS AFTER THE PROCEDURE, CARDIAC TAMPONADE WAS FOUND, AND THE PATIENT SUBSEQUENTLY DIED. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REF. # (B)(4). Manufacturer Narrative: DEVICE EVALUATION DETAILS: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER 31560023L AND NO INTERNAL ACTION RELATED TO THE COMPLAINT WAS FOUND DURING THE REVIEW. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REF.# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22409824&pc=DWC | 1718850-2025-10002 | 2019-02-19 05:00 | Death | MUNICH | 2025-07-07 04:00 | DWC | HEATER-COOLER SYSTEM 3T | Microbial Contamination of Device | Bacterial Infection |  |  | 1 | Event Description: LIVANOVA RECEIVED REPORT THAT A MALE PATIENT UNDERWENT CARDIOTHORACIC SURGERY ON (B)(6) 2019. THE SURGICAL PROCEDURE WAS AN AORTIC ROOT REPLACEMENT AND REPLACEMENT OF ASCENDING AORTA. PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2025. HE WAS DIAGNOSED WITH DISSEMINATED M. CHIMAERA INFECTION ON (B)(6) 2025. HE WAS DISCHARGED FROM THE HOSPITAL ON (B)(6) 2025 AND READMITTED ON (B)(6) 2025 AND THEN DISCHARGED HOME WITH HOSPICE ON (B)(6) 2025. THE PATIENT ULTIMATELY DIED ON (B)(6) 2025. HE WAS 64 YEARS OLD WHEN HE DIED. THE DEATH WAS CAUSED BY COMPLICATIONS OF DISSEMINATED M. CHIMAERA INFECTION. Manufacturer Narrative: H.9. LIVANOVA DEUTSCHLAND IMPLEMENTED A FIELD SAFETY NOTICE FOR DISINFECTION AND CLEANING OF HEATER-COOLER DEVICES. THE Z NUMBER IS Z-2076/2081-2015. H11: LIVANOVA DEUTSCHLAND MANUFACTURES THE HEATER-COOLER SYSTEM 3T. THE INCIDENT OCCURRED IN (B)(6). LEGAL LAWSUIT HAS BEEN ISSUED AND NO FURTHER INFORMATION HAS BEEN MADE AVAILABLE. DEVICE HISTORY RECORD (DHR) REVIEW OF INVOLVED DEVICE SERIAL NUMBER HAS BEEN COMPLETED AND DID NOT IDENTIFY ANY DEVIATIONS OR NON-CONFORMITIES RELEVANT TO THE REPORTED ISSUE. THE DEVICE INVOLVED AND IN USE AT THE HOSPITAL AT THE TIME OF SURGERY ((B)(6) 2019) WAS NOT YET EQUIPPED WITH VACUUM AND SEALING KIT SINCE UPGRADE ACTIVITY WAS PERFORMED ON 25 FEBRUARY 2020.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22409923&pc=DTB | 2649622-2025-18158 | 2025-06-13 04:00 | Death | MPRI | 2025-07-07 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Device Dislodged or Dislocated | Perforation; Cardiac Tamponade; Pericardial Effusion | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT POST OPERATIVELY, THE RIGHT ATRIAL (RA) LEAD DISLODGED, POSSIBLY DUE TO PERFORATION, RESULTING IN PATIENT EXPER IENCING CARDIAC TAMPONADE AND PERICARDIAL EFFUSION. PERICARDIOCENTESIS WAS PERFORMED. PATIENT CODED ON THE PROCEDURE TABLE. FLUOROSCOPY IMAGING SPECIFIED THE MIGRATION OF THE RA LEAD INTO THE VENTRICLE. EVENTUALLY, IT WAS NOTED THAT THE PATIENT PASSED AWAY FOUR DA YS AFTER THE IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. CAUSE OF DEATH COULD NOT BE CONFIRMED IF IT WAS RELATED TO THE RA OR RIGHT VENTRICULAR (RV) LEAD. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22409924&pc=NVN | 2649622-2025-18160 | 2025-06-13 04:00 | Death | MPRI | 2025-07-07 04:00 | NVN | SELECTSECURE MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT POST OPERATIVELY, THE RIGHT ATRIAL (RA) LEAD DISLODGED POSSIBLY DUE TO PERFORATION RESULTING IN PATIENT EXPERIENCING CARDIAC TAMPONADE AND PERICARDIAL EFFUSION. PERICARDIOCENTESIS WAS PERFORMED. PATIENT CODED ON THE PROCEDURE TABLE. FLUOROSCOPY IMAGING SPECIFIED THE MIGRATION OF THE RA LEAD INTO THE VENTRICLE. EVENTUALLY, IT WAS NOTED THAT THE PATIENT PASSED AWAY FOUR DAYS AFTER THE IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. CAUSE OF DEATH COULD NOT BE CONFIRMED IF IT WAS RELATED TO THE RA OR RIGHT VENTRICULAR (RV) LEAD. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22410246&pc=QFG | 3013756811-2025-158877 | 2025-06-13 04:00 | Death | TANDEM DIABETES CARE | 2025-07-07 04:00 | QFG | T:SLIM X2 INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Diabetic Ketoacidosis | K203234 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. CAUSE OF DEATH WAS DUE TO HYPERGLYCEMIA RESULTING IN DIABETIC KETOACIDOSIS WITH A BLOOD GLUCOSE (BG) LEVEL OF 600 MG/DL. THE PUMP HISTORY WAS REVIEWED BY THE HEALTHCARE PROVIDER AND MULTIPLE HIGH BG ALERTS WERE OBSERVED TO HAVE OCCURRED BETWEEN (B)(6). THE CUSTOMER INITIALLY ARRIVED AT THE EMERGENCY ROOM ON (B)(6) AND WAS TRANSFERRED BY EMERGENCY SERVICES TO ANOTHER HOSPITAL LATER THAT SAME DAY. EMERGENCY SERVICES DISCARDED THE INFUSION SET USED AT THE TIME OF THE EVENT, AND NO INFORMATION WAS PROVIDED ON THE CONTINUOUS GLUCOSE MONITOR (CGM) SENSOR USED AT THE TIME OF THE EVENT. NO ADDITIONAL INFORMATION COULD BE CONFIRMED AT THE TIME OF THIS REPORT, AS PUMP DATA IS NOT AVAILABLE FOR REVIEW. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS RECEIVED. Manufacturer Narrative: D4 (SERIAL NUMBER), H4.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22410490&pc=OZD | 1220648-2025-30060 | 2025-06-17 04:00 | Death | ABIOMED, INC. | 2025-07-07 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT. PRIOR TO BEING PLACED ON THE IMPELLA CP, THE PATIENT ARRIVED WITH COMPLAINTS OF CHEST PAIN/PRESSURE ONGOING FOR 4-5 DAYS PRIOR. IT WAS PLANNED FOR THE PATIENT TO BE PLACED ON AN INTRA-AORTIC BALLOON BUMP (IABP), BUT THE PATIENT IMMEDIATELY PULSELESS ELECTRICAL ACTIVITY (PEA) ARRESTED. CODE WAS CALLED AND THE DECISION WAS MADE TO PLACE THE IMPELLA CP. DURING SUPPORT, THE PATIENT EXPERIENCED LIMB ISCHEMIA BILATERALLY. MOTTLING AND TEMPERATURE CHANGES WERE NOTED. THE DECISION WAS MADE TO SEND THE PATIENT FOR VASCULAR SURGERY TO EXPLANT THE DEVICE AND REPAIR. SUBSEQUENTLY, IT WAS REPORTED THAT THE PATIENT'S LABS WERE WORSENING. ADDITIONAL INFORMATION NOTED CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22410537&pc=OZD | 1220648-2025-30061 | 2025-06-17 04:00 | Death | ABIOMED, INC. | 2025-07-07 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED AN IMPELLA CP WAS PLACED TO SUPPORT A PATIENT ADMITTED IN WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK AND ARREST. THE PUMP WAS PLACED BUT THERE WERE PERSISTENT LOW FLOW/SUCTION EVENTS AND THE PATIENT'S LIMBS WERE ISCHEMIC BILATERALLY. THERE WAS BLEEDING FROM THE GROIN THAT PROMPTED TREATING WITH AN EXTRA SUTURE, FEMSTOP, AND TRANSFUSION OF TWO UNITS OF PACKED RED BLOOD CELLS. DESPITE ALL MEASURES, THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW, LIMB ISCHEMIA, AND ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. LOW PUMP FLOW: THE DATA LOG SHOWS SEVERAL SUCTION ALARMS THROUGHOUT THE CASE RUN, WITH AN OBSERVED DOWNWARD TREND IN PLACEMENT SIGNAL AND PUMP FLOW. PUMP FLOW WAS SEEN GRADUALLY DECREASING WHILE RUNNING ON THE SAME P-LEVEL. A MOTOR CURRENT SPIKE WAS REPORTED IN THE MIDDLE OF THE CASE WITHOUT IMPACTING MAJOR ISSUES IN PUMP PERFORMANCE. THE MOTOR CURRENT SPIKE WAS INDICATIVE OF BIOMATERIAL INTERACTION DURING THE CASE; THEREFORE, THE THROMBUS FAILURE MODE WAS ADDED AS AN INVESTIGATED FAILURE MODE. ACCORDING TO THE CLINICAL NOTES, THE PATIENT WAS ON PRESSURE SUPPORT, AND BLOOD PRODUCTS WERE GIVEN DUE TO BLOOD LOSS DURING SUPPORT. THE CAUSE OF THE LOW PUMP FLOW ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED, BASED ON DATA LOG REVIEW AND GIVEN CLINICAL DETAILS. LIMB ISCHEMIA - REVERSIBLE: THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. ACCESS SITE BLEEDING-MAJOR: THE CAUSE OF THE ACCESS SITE BREEDING ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED. THROMBUS: THE ROOT CAUSE OF THE THROMBUS WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22410636&pc=OZD | 1220648-2025-30062 | 2025-06-17 04:00 | Death | ABIOMED, INC. | 2025-07-07 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Tachycardia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 AND EXPERIENCED RECURRENT VENTRICULAR TACHYCARDIA AND VENTRICULAR FIBRILLATION DURING SUPPORT. THE PATIENT WAS TREATED WITH LIDOCAINE. AFTER NINE DAYS OF IMPELLA 5.5 SUPPORT, THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22411076&pc=DQY | 3006425876-2025-00580 | 2025-06-09 04:00 | Death | ARROW INTERNATIONAL LLC | 2025-07-07 04:00 | DQY | ARROW CVC SET: 4-LUMEN 8.5FR X 16CM | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | K862056 |  | 1 | Event Description: IT WAS REPORTED "AFTER REMOVAL OF THE CENTRAL VENOUS CATHETER (CVC) PLACED IN THE RIGHT SUBCLAVIAN VEIN, THE PATIENT BEGAN TO SHOW CHANGES IN CONSCIOUSNESS AND MYOCLONUS - PROBABLY GAS EMBOLISM. TREATMENT: HYPERBARIC CHAMBER TREATMENT, + VENTILATION INVASIVE MECHANICS + VASOPRESSOR. " THE PATIENT'S CURRENT CONDITION IS REPORTED AS "DECEASED". ADDITIONAL INFORMATION REPORTS "ROOT CAUSE ANALYSIS WAS CARRIED OUT ON THIS INCIDENT AND THERE WAS NO EVIDENCE OF PROBLEMS WITH THE DEVICE USED, HOWEVER, THE DISTRIBUTOR WAS MADE AWARE OF THE OCCURENCE."Event Description: IT WAS REPORTED "AFTER REMOVAL OF THE CENTRAL VENOUS CATHETER (CVC) PLACED IN THE RIGHT SUBCLAVIAN VEIN, THE PATIENT BEGAN TO SHOW CHANGES IN CONSCIOUSNESS AND MYOCLONUS - PROBABLY GAS EMBOLISM. TREATMENT: HYPERBARIC CHAMBER TREATMENT, + VENTILATION INVASIVE MECHANICS + VASOPRESSOR. " THE PATIENT'S CURRENT CONDITION IS REPORTED AS "DECEASED". ADDITIONAL INFORMATION REPORTS "ROOT CAUSE ANALYSIS WAS CARRIED OUT ON THIS INCIDENT AND THERE WAS NO EVIDENCE OF PROBLEMS WITH THE DEVICE USED, HOWEVER, THE DISTRIBUTOR WAS MADE AWARE OF THE OCCURENCE." Manufacturer Narrative: (B)(4) Manufacturer Narrative: QN#(B)(4). COMPLAINT VERIFICATION TESTING COULD NOT BE PERFORMED AS IT WAS REPORTED THAT THE SAMPLE IS NOT AVAILABLE FOR RETURN. A DEVICE HISTORY RECORD REVIEW WAS PERFORMED AND NO RELEVANT FINDINGS WERE IDENTIFIED. WITHOUT THE DEVICE TO EVALUATE THE COMPLAINT COULD NOT BE CONFIRMED AND THE PROBABLE CAUSE COULD NOT BE DETERMINED FROM THE AVAILABLE INFORMATION. TELEFLEX WILL CONTINUE TO MONITOR AND TREND FOR REPORTS OF THIS NATURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22411436&pc=DSQ | 2916596-2025-04156 | 2025-06-10 04:00 | Death | THORATEC CORPORATION | 2025-07-07 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Movement Disorder; Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE)/EMERGENT REDO STERNOTOMY, INCISION AND DRAINAGE OF STERNAL WOUND, REPAIR OF OUTFLOW GRAFT TRACT (STERNAL INFECTION), AND PARTIAL STERNOTOMY WERE PERFORMED. ON THE DAY OF SURGERY / POST OPERATIVE DAY 1, THE PATIENT HAD AN EMBOLIC STROKE. THEY EXPERIENCED LEFT HEMIPLEGIA AND RIGHT GAZE DEVIATION. A COMPUTED TOMOGRAPHY PERFUSION OF THE BRAIN WAS PERFORMED. THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS THE STROKE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND WAS STATED TO HAVE OPERATED AS EXPECTED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATION FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D, ARE CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THIS IFU LISTS POTENTIAL ADVERSE EVENTS INCLUDING INFECTION (LOCAL, DRIVELINE OR PUMP POCKET INFECTION), STROKE, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING INFECTION PREVENTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22412409&pc=PNJ | 9612164-2025-03320 | 2025-06-08 04:00 | Death | MEDTRONIC IRELAND | 2025-07-07 04:00 | PNJ | MICRA¿ VR2 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY FIVE WEEKS POST IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE PATIENT PASSED AWAY. FOLLOW-UP WAS CONDUCTED TO DETERMINE THE CAUSE OF DEATH. BUT NO INFORMATION COULD BE OBTAINED. PHYSICIAN ASSUMES DEATH NOT LINKED TO THE LEADLESS IPG. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22412948&pc=OZD | 1220648-2025-30081 | 2025-06-20 04:00 | Death | ABIOMED, INC. | 2025-07-07 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Device Sensing Problem | Multiple Organ Failure | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PLACEMENT SIGNAL UNRELIABLE ALARM APPEARED ON AN AUTOMATED IMPELLA CONTROLLER DURING IMPLANT. SUPPORT REMAINED UNAFFECTED AND THE PLACEMENT WAS MONITORED VIA MOTOR CURRENT AND IMAGING. AFTER TWO DAYS OF SUPPORT, THE PATIENT DEVELOPED MULTI ORGAN FAILURE AND CARE WAS WITHDRAWN. THE PATIENT SUBSEQUENTLY PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE FAILURE AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22413340&pc=DSQ | 2916596-2025-04452 | 2022-09-05 04:00 | Death | THORATEC CORPORATION | 2025-07-07 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT SUFFERED FROM A PROGRESSIVE MULTIPLE ORGAN FAILURE (MOF). THEY PASSED AWAY DUE TO MOF. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22413773&pc=DSQ | 2916596-2025-04255 | 2025-04-08 04:00 | Death | THORATEC CORPORATION | 2025-07-07 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6) AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED DUE TO PATIENT PRIVACY LAWS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED THROUGH THE MANUFACTURING EXECUTION SYSTEM (MES) AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22413788&pc=DSQ | 2916596-2025-04131 | 2025-06-13 04:00 | Death | THORATEC CORPORATION | 2025-07-07 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENTS; HOWEVER, NO ADDITIONAL INFORMATION WAS NOT PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22413789&pc=NIK | 2182208-2025-03268 | 2024-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-07-07 04:00 | NIK | MEDTRONIC CRT | Appropriate Term/Code Not Available | Cardiac Arrest |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING IMPLANTATION OF CARDIAC RESYNCHRONIZATION THERAPY (CRT) DEVICES AND LEFT VENTRICULAR (LV) LEAD PLACEMENT. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN, AND ALL SUDDEN UNKNOWN DEATHS WERE CATEGORIZED AS A SUDDEN CARDIAC DEATH. THERE WERE PATIENTS WHO EXPERIENCED REHOSPITALIZATIONS DUE TO HEART FAILURE. THERE IS NO ALLEGATION OF DEVICE OR LEAD-DEATH, OR HEART FAILURE RELATEDNESS INDICATED IN THE ARTICLE; HOWEVER, THE CAUSE OF DEATH AND HEART FAILURE-RELATEDNESS HAS BEEN REQUESTED, BUT NOT YET RECEIVED. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/64 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: OPTIMIZING IMPLANTATION OF CARDIAC RESYNCHRONIZATION THERAPY: A RANDOMIZED CONTROLLED TRIAL OF ELECTROPHYSIOLOGICAL OR ANATOMICAL LEFT VENTRICULAR LEAD PLACEMENT STRATEGY. RUSSIAN JOURNAL OF CARDIOLOGY. 2025; 30(2): 5720. DOI: 10.15829/1560-4071-2025-5720. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22413805&pc=DXY | 2182208-2025-03269 | 2024-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-07-07 04:00 | DXY | MEDTRONIC LEAD | Appropriate Term/Code Not Available | Cardiac Arrest |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING IMPLANTATION OF CARDIAC RESYNCHRONIZATION THERAPY (CRT) DEVICES AND LEFT VENTRICULAR (LV) LEAD PLACEMENT. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN, AND ALL SUDDEN UNKNOWN DEATHS WERE CATEGORIZED AS A SUDDEN CARDIAC DEATH. THERE WERE PATIENTS WHO EXPERIENCED REHOSPITALIZATIONS DUE TO HEART FAILURE. THERE IS NO ALLEGATION OF DEVICE OR LEAD-DEATH, OR HEART FAILURE RELATEDNESS INDICATED IN THE ARTICLE; HOWEVER, THE CAUSE OF DEATH AND HEART FAILURE-RELATEDNESS HAS BEEN REQUESTED, BUT NOT YET RECEIVED. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/64 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: OPTIMIZING IMPLANTATION OF CARDIAC RESYNCHRONIZATION THERAPY: A RANDOMIZED CONTROLLED TRIAL OF ELECTROPHYSIOLOGICAL OR ANATOMICAL LEFT VENTRICULAR LEAD PLACEMENT STRATEGY. RUSSIAN JOURNAL OF CARDIOLOGY. 2025; 30(2): 5720. DOI: 10.15829/1560-4071-2025-5720. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22414239&pc=OZD | 1220648-2025-30087 | 2025-06-20 04:00 | Death | ABIOMED, INC. | 2025-07-07 04:00 | OZD | IMPELLA CP | Malposition of Device | Asystole | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND HAD THE PUMP COME OUT OF POSITION. THE PATIENT WENT INTO ASYSTOLE. CARDIOPULMONARY RESUSCITATION WAS RENDERED BUT THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22415044&pc=MHX | 8030229-2025-05157 | 2025-06-10 04:00 | Death | NIHON KOHDEN CORPORATION | 2025-07-07 04:00 | MHX | PU-681RA | Appropriate Term/Code Not Available | Low Blood Pressure/ Hypotension | K102376 |  | 1 | Event Description: THE HOSPITAL'S RISK MANAGER AT THE HOSPITAL REPORTED THAT AFTER RETURNING FROM A PROCEDURE, THE PATIENT EXPIRED WHILE BEING MONITORED AT THE CENTRAL NURSE'S STATION (CNS). THIS OCCURRED ON (B)(6) 2025 AT 1200 HOURS. THE LOGS ARE CURRENTLY BEING PULLED FROM THE DEVICES. SHE WAS CONCERNED BECAUSE IN THE PAST, MONITOR TECHS HAD BEEN SILENCING THE ALARMS, AND THEY WANT TO VERIFY THAT THIS IS NOT THE CASE WITH THIS EVENT. SHE ALSO WANTED TO SEE WHAT HAPPENED WITH THE PATIENT. Manufacturer Narrative: THE HOSPITAL'S RISK MANAGER AT THE HOSPITAL REPORTED THAT AFTER RETURNING FROM A PROCEDURE, THE PATIENT EXPIRED WHILE BEING MONITORED AT THE CENTRAL NURSE'S STATION (CNS). THIS OCCURRED ON (B)(6) 2025 AT 1200 HOURS. THE LOGS ARE CURRENTLY BEING PULLED FROM THE DEVICES. SHE WAS CONCERNED BECAUSE IN THE PAST, MONITOR TECHS HAD BEEN SILENCING THE ALARMS, AND THEY WANT TO VERIFY THAT THIS IS NOT THE CASE WITH THIS EVENT. SHE ALSO WANTED TO SEE WHAT HAPPENED WITH THE PATIENT. NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. THE FOLLOWING FIELD CONTAINS LIMITED INFORMATION, AS ATTEMPTS TO OBTAIN THE INFORMATION WERE MADE, BUT NOT PROVIDED: D10 ATTEMPT # 1: 06/12/2025 EMAILED THE CUSTOMER VIA MICROSOFT OUTLOOK FOR ALL INFORMATION IN THE NI LIST ABOVE: NO REPLY WAS RECEIVED. ATTEMPT # 2: 06/25/2025 EMAILED THE CUSTOMER VIA MICROSOFT OUTLOOK FOR ALL INFORMATION IN THE NI LIST ABOVE: NO REPLY WAS RECEIVED. ATTEMPT # 3: 07/07/2025 EMAILED THE CUSTOMER VIA MICROSOFT OUTLOOK FOR ALL INFORMATION IN THE NI LIST ABOVE: NO REPLY WAS RECEIVED. ADDITIONAL DEVICE INFORMATION: D10 CONCOMITANT MEDICAL DEVICE: THE FOLLOWING DEVICES WERE USED IN CONJUNCTION WITH THE CNS: MULTIPLE PATIENT RECEIVER (ORG): MODEL #: NI. SERIAL #: NI. DEVICE MANUFACTURER DATA: NI. UNIQUE IDENTIFIER (UDI) #: NI. RETURNED TO NIHON KOHDEN: NA. ZM TRANSMITTER: MODEL #: ZM-531PA. SERIAL #: (B)(6). DEVICE MANUFACTURER DATA: 09/25/2020. UNIQUE IDENTIFIER (UDI) #: (B)(4). RETURNED TO NIHON KOHDEN: NA.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22418339&pc=QBJ | MW5172344 |  | Death | ABBOTT DIABETES CARE INC. | 2025-07-07 04:00 | QBJ | FREESTYLE LIBRE 2 | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: PATIENT PASSED BUT WAS UNABLE TO CONFIRM IF ON MEDICATION AT TIME OF PASSING. REPORTER DID NOT CONSENT TO FOLLOW UP. (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22418410&pc=QBJ | MW5172345 | 2025-06-30 04:00 | Death | ABBOTT DIABETES CARE INC. | 2025-07-07 04:00 | QBJ | FREE STYLE LIBRE 3 | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: PATIENT DECEASED. REPORTED BY HCP OFFICE. HCP OFFICE DID NOT CONSENT TO FOLLOW UP. NO FURTHER INFORMATION/CLARIFICATION AVAILABLE. (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22418662&pc=NKM | 2135147-2025-03695 | 2015-03-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-07 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "TRANSCATHETER MANAGEMENT OF LEFT-SIDED VALVULAR HEART DISEASE FOLLOWING HEART TRANSPLANTATION" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE SAFETY, EFFICACY, AND OUTCOMES OF TRANSCATHETER INTERVENTIONS IN THIS HIGH-RISK POPULATION. DEVICES MENTIONED INCLUDE MITRACLIP AND NAVITOR. THE ARTICLE CONCLUDED THAT TAVI AND M-TEER WERE FEASIBLE AND EFFICIENT IN IMPROVING ECHOCARDIOGRAPHIC ALLOGRAFT FUNCTION. ELECTIVE PROCEDURES WERE ASSOCIATED WITH A MEDIAN SURVIVAL OF OVER TWO YEARS. SURVIVAL OUTCOMES VARIED BASED ON PROCEDURE URGENCY. CLOSE MONITORING OF AS/MR SEEMS CRUCIAL IN HTX PATIENTS, WITH TIMELY INTERVENTION PRIOR DECOMPENSATION/SHOCK. FURTHER MULTICENTRE STUDIES ARE NEEDED TO ESTABLISH MANAGEMENT GUIDELINES FOR AS/MR IN HTX RECIPIENTS. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS HRISTIAN HINKOV AT DEUTSCHES HERZZENTRUM DER CHARITÉ INSTITUTION WITH CORRESPONDING EMAIL HRISTIAN.HINKOV@DHZC-CHARITE.DE]. THE TIME FRAME OF THE STUDY WAS MARCH 2015 TO APRIL 2024. A TOTAL OF 15 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 15 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. (B)(6), UNK MITRACLIP. PERI- AND POST-PROCEDURAL COMPLICATION INCLUDED DEATH, HEART FAILURE, SURGICAL INTERVENTION, PROLONGED HOSPITALIZATION. (B)(6), UNK NAVITOR. PERI- AND POST-PROCEDURAL COMPLICATION INCLUDED DEATH, SEPSIS, STROKE, ARRHYTHMIA, PERMANENT PACEMAKER, HEART BLOCK. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH LIMITED LIFE EXPECTANCY. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE, SURGICAL INTERVENTION, PROLONGED HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22418692&pc=DSQ | 2916596-2025-04357 | 2025-06-18 04:00 | Death | THORATEC CORPORATION | 2025-07-07 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. BASED ON A REVIEW OF AVAILABLE PATIENT'S RECORDS AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE ANY ADDITIONAL INFORMATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22418826&pc=DQY | 2124215-2025-44404 | 2025-06-30 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-07 04:00 | DQY | WATCHMAN FXD CURVE? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | K212228 |  | 1 | Event Description: IT WAS REPORTED A PERICARDIAL EFFUSION OCCURRED. A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE WAS BEING PERFORMED. PRIOR TO THE PROCEDURE, A TRACE PERICARDIAL EFFUSION WAS NOTED ON BASELINE TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). THE TRANS-SEPTAL PUNCTURE (TSP) WAS PERFORMED USING AN NRG RF NEEDLE AND A FIXED CURVE SHEATH UNDER INTRACARDIAC ECHOCARDIOGRAPHY (ICE) GUIDANCE. THE IMPLANTING PHYSICIAN EXPERIENCED DIFFICULTY OBTAINING A CLEAR ANTERIOR-POSTERIOR (AP) VIEW ON ICE BUT OBSERVED A GOOD TENTING OF THE SEPTUM IN THE INFERIOR/SUPERIOR VIEW. THE NEEDLE WAS INITIALLY ADVANCED ACROSS THE SEPTUM, THEN RETRACTED TO THE RIGHT ATRIUM. AFTER REVIEWING TEE IMAGING, THE PHYSICIAN PROCEEDED WITH SEPTAL CROSSING USING RF ENERGY. AS THE SHEATH WAS EXCHANGED FOR A WAS DOUBLE CURVE SHEATH AND THE PIGTAIL CATHETER WAS ADVANCED, THE TEE IMAGING PHYSICIAN NOTED THAT THE PREVIOUSLY OBSERVED TRACE EFFUSION HAD INCREASED IN SIZE. A PERICARDIOCENTESIS TRAY WAS PREPARED, AND CONTRAST IMAGING OF THE LEFT ATRIAL APPENDAGE (LAA) WAS PERFORMED. NO CONTRAST LEAK WAS OBSERVED FROM THE APPENDAGE OR INTO THE PERICARDIAL SPACE. PERICARDIOCENTESIS WAS COMPLETED AND THE PATIENT WAS TREATED WITH PROTAMINE, IV FLUIDS, AND HEMODYNAMIC SUPPORT PROVIDED BY THE CERTIFIED REGISTERED NURSE ANESTHETIST (CRNA). A CARDIOTHORACIC SURGEON WAS CALLED FOR CONSULTATION. THE PATIENT RECEIVED TWO UNITS OF PACKED RED BLOOD CELLS AND WAS TRANSFERRED TO THE OPERATING ROOM FOR SURGICAL MANAGEMENT. INTRAOPERATIVELY, A PERFORATION WAS IDENTIFIED IN THE LEFT ATRIAL WALL NEAR THE RIGHT INFERIOR PULMONARY VEIN, WHICH WAS DETERMINED TO BE THE SOURCE OF THE EFFUSION. THE PERFORATION WAS SURGICALLY REPAIRED, AND THE LAA WAS LIGATED. THE PATIENT TOLERATED THE PROCEDURE WELL, WITH NO FURTHER COMPLICATIONS, AND WAS TRANSFERRED TO RECOVERY IN STABLE CONDITION.Event Description: IT WAS REPORTED A PERICARDIAL EFFUSION OCCURRED. A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE WAS BEING PERFORMED. PRIOR TO THE PROCEDURE, A TRACE PERICARDIAL EFFUSION WAS NOTED ON BASELINE TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). THE TRANS-SEPTAL PUNCTURE (TSP) WAS PERFORMED USING AN NRG RF NEEDLE AND A FIXED CURVE SHEATH UNDER INTRACARDIAC ECHOCARDIOGRAPHY (ICE) GUIDANCE. THE IMPLANTING PHYSICIAN EXPERIENCED DIFFICULTY OBTAINING A CLEAR ANTERIOR-POSTERIOR (AP) VIEW ON ICE BUT OBSERVED A GOOD TENTING OF THE SEPTUM IN THE INFERIOR/SUPERIOR VIEW. THE NEEDLE WAS INITIALLY ADVANCED ACROSS THE SEPTUM, THEN RETRACTED TO THE RIGHT ATRIUM. AFTER REVIEWING TEE IMAGING, THE PHYSICIAN PROCEEDED WITH SEPTAL CROSSING USING RF ENERGY. AS THE SHEATH WAS EXCHANGED FOR A WAS DOUBLE CURVE SHEATH AND THE PIGTAIL CATHETER WAS ADVANCED, THE TEE IMAGING PHYSICIAN NOTED THAT THE PREVIOUSLY OBSERVED TRACE EFFUSION HAD INCREASED IN SIZE. A PERICARDIOCENTESIS TRAY WAS PREPARED, AND CONTRAST IMAGING OF THE LEFT ATRIAL APPENDAGE (LAA) WAS PERFORMED. NO CONTRAST LEAK WAS OBSERVED FROM THE APPENDAGE OR INTO THE PERICARDIAL SPACE. PERICARDIOCENTESIS WAS COMPLETED AND THE PATIENT WAS TREATED WITH PROTAMINE, IV FLUIDS, AND HEMODYNAMIC SUPPORT PROVIDED BY THE CERTIFIED REGISTERED NURSE ANESTHETIST (CRNA). A CARDIOTHORACIC SURGEON WAS CALLED FOR CONSULTATION. THE PATIENT RECEIVED TWO UNITS OF PACKED RED BLOOD CELLS AND WAS TRANSFERRED TO THE OPERATING ROOM FOR SURGICAL MANAGEMENT. INTRAOPERATIVELY, A PERFORATION WAS IDENTIFIED IN THE LEFT ATRIAL WALL NEAR THE RIGHT INFERIOR PULMONARY VEIN, WHICH WAS DETERMINED TO BE THE SOURCE OF THE EFFUSION. THE PERFORATION WAS SURGICALLY REPAIRED, AND THE LAA WAS LIGATED. THE PATIENT TOLERATED THE PROCEDURE WELL, WITH NO FURTHER COMPLICATIONS, AND WAS TRANSFERRED TO RECOVERY IN STABLE CONDITION. IT WAS FURTHER REPORTED THAT THE PATIENT WAS EXTUBATED ONE (1) DAY POST-OP, HOWEVER THE PATIENT BEGAN TO DECLINE. THREE (3) DAYS POST-OP THE PATIENT PASSED AWAY DUE TO SEPTIC SHOCK. Manufacturer Narrative: B5 DESCRIBE EVENT OR PROBLEM: UPDATED WITH ADDITIONAL INFORMATION RECEIVED. B2: DATE OF DEATH ESTIMATED AS (B)(6) 2025 AS THE PATIENT DIED 3 DAYS POST-OPERATION. H6: IMPACT CODE ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22419139&pc=DSQ | 2916596-2025-04257 | 2024-11-28 05:00 | Death | THORATEC CORPORATION | 2025-07-07 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO PULMONARY BLEEDING. THE PATIENT SUFFERED FROM PULMONARY BLEEDING WITH PULMONARY FAILURE IN THE COURSE WHICH LED TO THE DEATH AT THE END. NO DEVICE ISSUES WERE MENTIONED. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) OPERATED AS EXPECTED. THE DEATH WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22419917&pc=NPT | 2135147-2025-03703 | 2015-03-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-07 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Stroke/CVA; Sepsis; Heart Block | P190023 |  | 1 | Event Description: THE ARTICLE "TRANSCATHETER MANAGEMENT OF LEFT-SIDED VALVULAR HEART DISEASE FOLLOWING HEART TRANSPLANTATION" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE SAFETY, EFFICACY, AND OUTCOMES OF TRANSCATHETER INTERVENTIONS IN THIS HIGH-RISK POPULATION. DEVICES MENTIONED INCLUDE MITRACLIP AND NAVITOR. THE ARTICLE CONCLUDED THAT TAVI AND M-TEER WERE FEASIBLE AND EFFICIENT IN IMPROVING ECHOCARDIOGRAPHIC ALLOGRAFT FUNCTION. ELECTIVE PROCEDURES WERE ASSOCIATED WITH A MEDIAN SURVIVAL OF OVER TWO YEARS. SURVIVAL OUTCOMES VARIED BASED ON PROCEDURE URGENCY. CLOSE MONITORING OF AS/MR SEEMS CRUCIAL IN HTX PATIENTS, WITH TIMELY INTERVENTION PRIOR DECOMPENSATION/SHOCK. FURTHER MULTICENTRE STUDIES ARE NEEDED TO ESTABLISH MANAGEMENT GUIDELINES FOR AS/MR IN HTX RECIPIENTS. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS HRISTIAN HINKOV AT DEUTSCHES HERZZENTRUM DER CHARITÉ INSTITUTION WITH CORRESPONDING EMAIL HRISTIAN.HINKOV@DHZC-CHARITE.DE]. THE TIME FRAME OF THE STUDY WAS MARCH 2015 TO APRIL 2024. A TOTAL OF 15 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 15 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. (B)(6), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATION INCLUDED DEATH, HEART FAILURE, SURGICAL INTERVENTION, PROLONGED HOSPITALIZATION. (B)(6), UNK NAVITOR PERI- AND POST-PROCEDURAL COMPLICATION INCLUDED DEATH, SEPSIS, STROKE, ARRHYTHMIA, PERMANENT PACEMAKER, HEART BLOCK. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF NAVITOR DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH LIMITED LIFE EXPECTANCY. COMPLICATIONS REPORTED INCLUDED DEATH, SEPSIS, STROKE, ARRHYTHMIA, PERMANENT PACEMAKER, HEART BLOCK; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22420227&pc=MHX | 1218950-2025-000384 | 2023-07-26 04:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-07-07 04:00 | MHX | INTELLIVUE MX40 1.4GHZ | Device Alarm System | Insufficient Information | K113125 |  | 1 | Event Description: PHILIPS RECEIVED A COMPLAINT ON THE MX40 1.4 GHZ SMART HOPPING INDICATING A PATIENT CONNECTED TO TELEMETRY REMOVED THE LEADS, WHICH RESULTED IN THE GENERATION OF A BLUE INOP ALARM FOR 'LEADS OFF' INSTEAD OF RED, WHICH WAS EXPECTED BY THE USER. RESPIRATORY STAFF FOUND THE PATIENT IN PULSELESS ELECTRICAL ACTIVITY (PEA) WITH ECG LEADS OFF AND THE PATIENT HAD PASSED AWAY. IT WAS SUSPECTED WHEN TELEMETRY WENT OFFLINE FOR A BATTERY CHANGE, THE ALARM SETTINGS DEFAULTED BACK TO THE BEDSIDE SETTINGS; HOWEVER, THIS CANNOT BE CONFIRMED AT THIS TIME. Manufacturer Narrative: THE AUDIT LOGS WERE REVIEWED BY THE CLINICAL PRODUCT SPECIALIST (CPS) AND THE PRODUCT SPECIALIST ENGINEER (PSE) TO DETERMINE IF THE DEVICE FUNCTIONED AS EXPECTED. THEY EXPLAINED THAT PRIOR TO THE ECG ¿LEADS OFF¿ THAT GENERATED A CYAN COLOR ON (B)(6) 2023 AT 05:13:40, THE DEVICE WAS RECONNECTED TO THE PIC AT THE START OF MONITORING HENCE THE ¿NO DATA TELE¿ ENDED PRIOR TO THAT. IN THE INFORMATION FOR USE (IFU), IT STATES THAT THE ECG LEADS OFF WILL DISPLAY AS A CYAN INITIALLY UNTIL THE ECG SIGNAL IS OBTAINED. SO, IF THE DEVICE IS AT THE START OF MONITORING DUE TO PIC IX REBOOT OR IF IT IS OUT OF THE ACCESS POINT (AP) RANGE, THEN IT WILL GO CYAN AT THE START UNTIL THERE IS AN ECG SIGNAL. DUE TO THE DEVICE BEING RECONNECTED, THERE HAD BEEN NO ECG SIGNAL ACQUIRED; THEREFORE, THE INOP REMAINS CYAN IN COLOR. BASED ON THIS INFORMATION, IT WAS DETERMINED THAT NO PRODUCT MALFUNCTION OCCURRED. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS USER ERROR. THE REPORTED PROBLEM WAS NOT CONFIRMED THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22420304&pc=OZD | 1220648-2025-30098 | 2025-06-16 04:00 | Death | ABIOMED, INC. | 2025-07-07 04:00 | OZD | IMPELLA 5.5 | Positioning Failure | Multiple Organ Failure | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT AN IMPELLA 5.5 WAS IMPLANTED TOO DEEP. THE TEAM TRIED TO REPOSITION THE PUMP BUT WAS UNSUCCESSFUL AS THE TUOHY BORST WOULD NOT ALLOW PUMP MOVEMENT. THE PATIENT DIED FROM MULTI ORGAN FAILURE AFTER 12 DAYS OF SUPPORT. Manufacturer Narrative: ATHE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22420345&pc=DSQ | 2916596-2025-04235 | 2025-06-15 04:00 | Death | THORATEC CORPORATION | 2025-07-07 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND DECEASED ON THE MORNING OF (B)(6) 2025. THE CORONER SUSPECTED THE PATIENT HAD BEEN DECEASED FOR 4 DAYS. THE SYSTEM CONTROLLER WAS STILL ALARMING. LOG FILES WERE SUBMITTED FOR REVIEW AND CAPTURED PERSISTENT LOW FLOW HAZARD ALARMS BETWEEN (B)(6) 2025 AND (B)(6) 2025 UNTIL THE PUMP WAS DISCONNECTED FROM THE SYSTEM CONTROLLER AT 10:58 AM ON (B)(6) 2025. THE PERIODIC LOG CONTAINED A SINGLE LOW FLOW HAZARD RECORDED DURING THE PERIODIC INTERVAL ON (B)(6) 2025 AT 10:14 AM. THERE WERE NO OTHER UNUSUAL EVENTS. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22420350&pc=BZD | 2518422-2025-107141 | 2025-07-03 04:00 | Death | RESPIRONICS, INC. | 2025-07-07 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO DEATH OF A PATIENT. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. NO OTHER MEDICAL INFORMATION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22420588&pc=FKX | 0008030665-2025-01544 | 2025-06-26 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-07-07 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K222318 |  | 1 | Event Description: ON (B)(6) 2025 A HOSPITAL NURSE REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT HAD PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. THE PATIENT WAS AT THE END OF TREATMENT (STEP 8). THE NURSE DISCONNECTED THE PATIENT AND REALIZED THE PATIENT WAS NOT WAKING UP. THE NURSE TOOK THE PATIENT¿S VITAL SIGNS, AND HE SUBSEQUENTLY WENT INTO CARDIAC ARREST AND DIED. ADDITIONAL INFORMATION WAS OBTAINED THROUGH THE DIALYSIS UNIT NURSE. THE PATIENT WAS AN IN-PATIENT REHABILITATION (REHAB) DIALYSIS PATIENT. THE PATIENT WAS IN REHAB DUE TO A LEFT BELOW KNEE AMPUTATION (BKA) WHICH WAS UNRELATED TO DIALYSIS. THE DATE OF THE ADMISSION TO THE REHAB WAS UNKNOWN. THE NURSE STATED THAT ON (B)(6) 2025, THE HOSPITAL NURSE WAS DOING ROUNDS WHEN THEY WENT TO DISCONNECT THIS PATIENT FROM THE LIBERTY SELECT CYCLER AT THE END OF TREATMENT. THE NURSE WENT TO TAKE THE PATIENT¿S VITALS AND FOUND THE PATIENT¿S PULSE TO BE THREADY. THE PATIENT SUBSEQUENTLY WENT INTO CARDIAC ARREST. THE NURSE COULD NOT CONFIRM WHAT RESUSCITATIVE EFFORTS WERE ADMINISTERED. THE EFFORTS WERE NOT SUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DECEASED. THE CAUSE OF DEATH WAS CARDIAC ARREST. THE NURSE STATED THE DEATH WAS NOT RELATED TO PD THERAPY OR USE OF THE LIBERTY SELECT CYCLER OR OTHER FRESENIUS PRODUCT(S). NO FURTHER INFORMATION WAS PROVIDED.Event Description: ON (B)(6) 2025 A HOSPITAL NURSE REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT HAD PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. THE PATIENT WAS AT THE END OF TREATMENT (STEP 8). THE NURSE DISCONNECTED THE PATIENT AND REALIZED THE PATIENT WAS NOT WAKING UP. THE NURSE TOOK THE PATIENT¿S VITAL SIGNS, AND HE SUBSEQUENTLY WENT INTO CARDIAC ARREST AND DIED. ADDITIONAL INFORMATION WAS OBTAINED THROUGH THE DIALYSIS UNIT NURSE. THE PATIENT WAS AN IN-PATIENT REHABILITATION (REHAB) DIALYSIS PATIENT. THE PATIENT WAS IN REHAB DUE TO A LEFT BELOW KNEE AMPUTATION (BKA) WHICH WAS UNRELATED TO DIALYSIS. THE DATE OF THE ADMISSION TO THE REHAB WAS UNKNOWN. THE NURSE STATED THAT ON (B)(6) 2025, THE HOSPITAL NURSE WAS DOING ROUNDS WHEN THEY WENT TO DISCONNECT THIS PATIENT FROM THE LIBERTY SELECT CYCLER AT THE END OF TREATMENT. THE NURSE WENT TO TAKE THE PATIENT¿S VITALS AND FOUND THE PATIENT¿S PULSE TO BE THREADY. THE PATIENT SUBSEQUENTLY WENT INTO CARDIAC ARREST. THE NURSE COULD NOT CONFIRM WHAT RESUSCITATIVE EFFORTS WERE ADMINISTERED. THE EFFORTS WERE NOT SUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DECEASED. THE CAUSE OF DEATH WAS CARDIAC ARREST. THE NURSE STATED THE DEATH WAS NOT RELATED TO PD THERAPY OR USE OF THE LIBERTY SELECT CYCLER OR OTHER FRESENIUS PRODUCT(S). NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF CARDIAC ARREST AND DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND UTILIZATION OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE PATIENT WAS FOUND WITH A THREADY PULSE AND WENT INTO CARDIAC ARREST. ALTHOUGH IT IS UNKNOWN IF THE PATIENT HAD PRE-EXISTING CARDIAC DISEASE, CARDIAC DISEASE IS THE MAJOR CAUSE OF DEATH, ACCOUNTING FOR APPROXIMATELY 37 PERCENT OF ALL-CAUSE MORTALITY IN PATIENTS RECEIVING EITHER HEMODIALYSIS OR PERITONEAL DIALYSIS. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S CARDIAC ARREST AND DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE NOT VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT. THERE WERE VISUAL INDICATION OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NOT BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. SYSTEM AIR LEAK TEST PASSED. VALVE ACTUATION TEST PASSED. TEACH PUMPS TEST PASSED. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THERE WERE NO VISUAL DISCREPANCIES ENCOUNTERED DURING THE INTERNAL INSPECTION. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22420867&pc=OZD | 1220648-2025-30101 | 2025-06-16 04:00 | Death | ABIOMED, INC. | 2025-07-07 04:00 | OZD | IMPELLA 5.5 | Malposition of Device | Ventricular Fibrillation | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH AN IMPELLA 5.5 WAS AGITATED AND PULLED THE IMPELLA CATHETER TUBING OUT OF PLACE. THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION AND CARDIOPULMONARY RESUSCITATION WAS PROVIDED. AFTER 30 MINUTES THE PATIENT EXPIRED. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT AND POSITIONING ISSUES HAS BEEN COMPLETED. VT/VF: THE ROOT CAUSE OF THE VT/VF/AT/AF WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. POSITIONING ISSUES: DATA LOGS WERE REVIEWED, AND THERE WAS A CONCENTRATION OF POSITION IN VENTRICLE AND POSITION IN AORTA ALARMS LATE THE EVENING OF (B)(6) 2025. AT THIS TIME THERE WERE DISTINCT CHANGES IN MOTOR CURRENT (MC) AND PLACEMENT SIGNAL (PS) PULSATILITY ALIGNED WITH THE ALARMS. THE PUMP HAD BEEN FUNCTIONING NORMALLY UP TO THAT POINT. RETURNED PRODUCT WAS INVESTIGATED AND THE CATHETER ANCHOR STILL FUNCTIONED AS NORMAL. THERE WERE NO VISUAL DAMAGES/ABNORMALITIES NOTED ON THE PUMP BESIDES THE STERILE SLEEVE BEING SEPARATED FROM BOTH THE PROXIMAL AND DISTAL HUB. BASED ON THE CLINICAL DETAILS PROVIDED THAT THE PATIENT PULLED THE CATHETER OUT OF THE LEFT VENTRICLE (LV0, THE ROOT CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY PATIENT MOVEMENT. PRODUCT DAMAGE - STERILE SLEEVE DAMAGE: INVESTIGATION OF RETURNED PRODUCT FOUND THE STERILE SLEEVE TO BE SEPARATED FROM BOTH HUBS. THEREFORE, THE PRODUCE DAMAGE FAILURE MODE WAS ADDED AND INVESTIGATED. CLINICAL DETAILS REPORTED THAT THE PATIENT WAS AGITATED, PULLED ON THE IMPELLA CATHETER, CAUSING IT TO COME OUT OF THE LV. BASED ON THE CLINICAL DETAILS THAT THE PATIENT PULLED ON THE CATHETER, THE ROOT CAUSE OF THE PRODUCT DAMAGE (STERILE SLEEVE DAMAGE) WAS MOST LIKELY A USE ISSUE. B1 PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30101.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22423454&pc=NPT | 2025587-2025-04811 | 2025-07-03 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-07 04:00 | NPT | EVOLUT PRO PLUS DCS | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Stroke/CVA | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE INVOLVING THE SENTRANT #18 INTRODUCER SHEATH AND THE EVPROPLUS-29 VALVE, SIGNIFICANT FORCE WAS APPLIED TO INSERT THE INTRODUCER THROUGH THE FEMORAL ARTERY DUE TO RESISTANCE, BUT IT WAS EVENTUALLY ADVANCED SUCCESSFULLY. THE GUIDEWIRE WAS THEN NAVIGATED INTO THE LEFT VENTRICLE, AND THE VALVE DELIVERY SYSTEM WAS ADVANCED; HOWEVER, THE VALVE WAS NEVER DEPLOYED, AND NO PREDILATATION OR AGGRESSIVE ANNULAR MANIPULATION OCCURRED. BEFORE VALVE DEPLOYMENT, THE PATIENT EXPERIENCED A SUDDEN ONSET OF RESPIRATORY DISTRESS, WHICH RAPIDLY PROGRESSED TO ASYSTOLE AND CARDIAC ARREST. DESPITE RESUSCITATION EFFORTS, INCLUDING CARDIOPULMONARY RESUSCITATION, DRUG ADMINISTRATION, AND FULL SUPPORT FROM THE RESUSCITATION TEAM, THE PATIENT COULD NOT BE REVIVED. THERE WAS NO EVIDENCE OF DEVICE MALFUNCTION OR DIRECT INTERACTION WITH CARDIAC STRUCTURES, AND NO DIRECT IMPACT FROM THE VALVE SYSTEM ON THE PATIENT COULD BE CONFIRMED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE OF DEATH WAS UNKNOWN. IT COULD HAVE BEEN AN ISSUE WITH THE FEMORAL ACCESS AND THE FORCE APPLIED, OR A POSSIBLE STROKE LATER WAS CONSIDERED, HOWEVER IT IS UNKNOWN.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT CLARIFIED THAT A STROKE WAS NOT CONFIRMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT PRO PLUS VALVE; PRODUCT ID EVPROPLUS-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED B.5, D.4, H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. CORRECTED DATA: D4. FULL UDI ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22424977&pc=LWS | 2649622-2025-18281 | 2025-06-30 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-07 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Signal Artifact/Noise; Defibrillation/Stimulation Problem; Under-Sensing | Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING RESULTING IN A VERY DELAYED DETECTION IN TREATING VENTRICULAR FIBRILLATION (VF). THE EPISODE WAS BELOW THE RATE OF TREATABLE ZONES. NOISE WAS ALSO OBSERVED DUE TO CARDIOPULMONARY RESUSCITATION ON THE PATIENT. THE PATIENT WAS ULTIMATELY SHOCKED BY THE AMBULANCE FOR VF. THE LEAD REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING RESULTING IN A VERY DELAYED DETECTION IN TREATING VENTRICULAR FIBRILLATION (VF). THE EPISODE WAS BELOW THE RATE OF TREATABLE ZONES. NOISE WAS ALSO OBSERVED DUE TO CARDIOPULMONARY RESUSCITATION ON THE PATIENT. THE PATIENT WAS ULTIMATELY SHOCKED BY THE AMBULANCE FOR VF. THE LEAD REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. IT WAS FURTHER REPORTED THAT THE PATIENT DIED THE DAY FOLLOWING THE UNDERSENSING AND DELAYED DETECTION OF VENTRICULAR FIBRILLATION (VF). IT WAS NOTED THAT THE EVENT WAS DEEMED TO BE ASSOCIATED WITH THE PATIENT'S CONDITION AND THE DEATH WAS BELIEVED TO BE UNRELATED TO THE LEAD. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED RIGHT VENTRICULAR UNDERSENSING. THE DEVICE MEMORY INDICATED A DELAY IN VENTRICULAR TACHYARRHYTHMIA THERAPY. ANALYSIS OF THE DEVICE MEMORY OBSERVED NOISE ON THE ELECTROGRAM WAVEFORMS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: DTPA2D1 IMPLANTED: ON (B)(6) 2022. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22426970&pc=LWS | 2017865-2025-91522 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-07 04:00 | LWS | FORTIFY ASSURA VR ICD, US | Failure to Deliver Shock/Stimulation; Low impedance; Charging Problem | Arrhythmia | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PRESENTED WITH A SUSTAINED ARRYTHMIA NOTED. UPON EXAMINATION AN OVER-CURRENT DETECTION ISSUE, LOW DEFIBRILLATION IMPEDANCE, AND NO DEFIBRATION THERAPY WAS NOTED AND ALLEGED ON THE LEAD AND DEVICE. THE PATIENT WAS REPORTED TO HAVE PASSED AWAY. NO FURTHER INTERVENTION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22427504&pc=NVY | 2017865-2025-91523 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-07 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Failure to Deliver Shock/Stimulation; Low impedance; Charging Problem | Arrhythmia | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PRESENTED WITH A SUSTAINED ARRYTHMIA NOTED. UPON EXAMINATION AN OVER-CURRENT DETECTION ISSUE, LOW DEFIBRILLATION IMPEDANCE, AND NO DEFIBRATION THERAPY WAS NOTED AND ALLEGED ON THE LEAD AND DEVICE. THE PATIENT WAS REPORTED TO HAVE PASSED AWAY. NO FURTHER INTERVENTION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22408922&pc=BZD | 2518422-2025-107109 | 2025-03-21 04:00 | Death | RESPIRONICS, INC. | 2025-07-06 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION FROM THE PATIENT'S REPRESENTATIVE ALLEGING THAT THE PATIENT HAS PASSED AWAY DUE TO CANCER. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22403322&pc=FTL | 3011175548-2025-000007 | 2019-10-23 04:00 | Death | ATRIUM MEDICAL CORPORATION | 2025-07-04 04:00 | FTL | MESH, SURGICAL, POLYMERIC | Adverse Event Without Identified Device or Use Problem | Fistula; Neuropathy; Pain; Urinary Tract Infection; Quadriplegia |  |  | 1 | Event Description: MEDWATCH RECEIVED FROM PATIENT'S FAMILY MEMBER, ALLEGING COMPLICATIONS FROM AN UNKNOWN ATRIUM MESH PRODUCT. PATIENT UNDERWENT EMERGENCY SURGERY IN 2019 FOR A LIFE-THREATENING SEPTIC BOWEL BLOCKAGE, IN WHICH MESH WAS IMPLANTED. REPORTED COMPLICATIONS INCLUDE EXTREME NEUROPATHY, FUNCTIONAL QUADRIPLEGIA, LARGE ABDOMINAL WOUNDS WITH FISTULAS, NUMEROUS UTIS, AND PERSISTENT PAIN. THE HERNIA MESH PRODUCTS CAUSED SIGNIFICANT DAMAGE TO HIS INTERNAL ORGANS, LEADING TO MULTIPLE SURGERIES AND A CONTINUOUS DECLINE IN HIS HEALTH, AND SEVERE IMPACT TO QUALITY OF LIFE. THE PATIENT ULTIMATELY PASSED AWAY ON UNKNOWN DATE. Manufacturer Narrative: A FOLLOW-UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. Manufacturer Narrative: ADDITIONAL INFORMATION SECTION: H6. AN FDA MEDWATCH (MW5171293) WAS RECEIVED WHICH DID NOT IDENTIFY ANY SPECIFIC ATRIUM MESH PRODUCT WAS USED IN THE REPORTED EVENT. THE REVIEW OF THE UNFORTUNATE DETAILS DOES NOT INDICATE THAT AN ATRIUM HERNIA MESH WAS USED. THERE WAS NO MENTION OF ATRIUM MESH PART DESCRIPTIONS, PART NUMBERS/ CATALOG NUMBER OR PRODUCT LOT NUMBER. THE DETAILS DO MENTION BY NAME TWO OTHER MANUFACTURERS OF HERNIA MESH. THE DETAILS ALSO SUGGEST THAT THE MEDICAL RECORDS WERE REVIEWED VERY CLOSELY AND IF THERE WERE A MENTION OF THE TWO OTHER MANUFACTURERS AND NOT ATRIUM MEDICAL MESH. PRODUCT PART NUMBERS AND LOT NUMBERS AND DESCRIPTIONS ARE STANDARD USUALLY WITHIN MEDICAL RECORDS FROM CLINICAL PROCEDURES FOR TRACEABILITY. WITHOUT MORE DETAILED INFORMATION REGARDING THE SURGICAL MESH USED (THIRD PRODUCT) THE COMPLAINT CANNOT BE CONFIRMED. THERE IS NO EVIDENCE THAT ATRIUM HERNIA MESH WAS USED IN THE PROCEDURE BASED ON THE DETAILS PROVIDED. THE INSTRUCTION FOR USE FOR THE ATRIUM MEDICAL POLYPROPYLENE MESH AW010666 LISTS THE ADVERSE REACTION THAT MAY OCCUR WITH THE USE OF ANY SURGICAL MESH. "COMPLICATIONS THAT MAY OCCUR WITH THE USE OF ANY SURGICAL MESH INCLUDE, BUT ARE NOT LIMITED TO PAIN, INFLAMMATION, INFECTION, ALLERGIC REACTION, SEROMA, HEMATOMA, FISTULA FORMATION OR MECHANICAL DISRUPTION OF THE TISSUE AND/OR MESH MATERIAL AND POSSIBLY ADHESIONS WHEN PLACED IN DIRECT CONTACT WITH THE VISCERA AND OTHER ORGANS. THE PATIENT SHOULD BE ADVISED TO CONTACT THE PHYSICIAN SHOULD AN ADVERSE REACTION OCCUR." THERE WAS NO PRODUCT LOT NUMBER PROVIDED THEREFORE A DEVICE HISTORY RECORD REVIEW CANNOT BE CONDUCTED. AS THE PRODUCT WAS DISCONTINUED IN JULY 2019 THERE WAS NO CR/CAPA OR NCR QUERY CONDUCTED. BASED ON THE DETAILS OF THE COMPLAINT THE ROOT CAUSE IS IMPOSSIBLE TO DEFINE AS THERE IS NO EVIDENCE TO SUGGEST AN ATRIUM PRODUCT WAS USED IN THE PROCEDURE AND THEREFORE CAUSED ANY OF THE PATIENT¿S CONDITIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22403674&pc=DWC | 9611109-2025-90224 | 2019-02-06 05:00 | Death | MUNICH | 2025-07-04 04:00 | DWC | HEATER-COOLER SYSTEM 3T | Microbial Contamination of Device | Bacterial Infection |  |  | 1 | Event Description: LIVANOVA RECEIVED REPORT THAT A MALE PATIENT UNDERWENT CARDIOTHORACIC SURGERY ON (B)(6) 2019. THE SURGICAL PROCEDURE WAS AN AORTIC ROOT REPLACEMENT AND REPLACEMENT OF ASCENDING AORTA. PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2025. HE WAS DIAGNOSED WITH DISSEMINATED M. CHIMAERA INFECTION ON (B)(6) 2025. HE WAS DISCHARGED FROM THE HOSPITAL ON (B)(6) 2025 AND READMITTED ON (B)(6) 2025 AND THEN DISCHARGED HOME WITH HOSPICE ON (B)(6) 2025. THE PATIENT ULTIMATELY DIED ON (B)(6) 2025. HE WAS 64 YEARS OLD WHEN HE DIED. THE DEATH WAS CAUSED BY COMPLICATIONS OF DISSEMINATED M. CHIMAERA INFECTION. Manufacturer Narrative: H.9. LIVANOVA DEUTSCHLAND IMPLEMENTED A FIELD SAFETY NOTICE FOR DISINFECTION AND CLEANING OF HEATER-COOLER DEVICES. THE Z NUMBER IS Z-2076/2081-2015. H11: LIVANOVA DEUTSCHLAND MANUFACTURES THE HEATER-COOLER SYSTEM 3T. THE INCIDENT OCCURRED IN (B)(6). LEGAL LAWSUIT HAS BEEN ISSUED AND NO FURTHER INFORMATION HAS BEEN MADE AVAILABLE. DEVICE HISTORY RECORD (DHR) REVIEW OF INVOLVED DEVICE SERIAL NUMBER HAS BEEN COMPLETED AND DID NOT IDENTIFY ANY DEVIATIONS OR NON-CONFORMITIES RELEVANT TO THE REPORTED ISSUE. THE DEVICE INVOLVED AND IN USE AT THE HOSPITAL AT THE TIME OF SURGERY ((B)(6) 2019) WAS NOT YET EQUIPPED WITH VACUUM AND SEALING KIT SINCE UPGRADE ACTIVITY WAS PERFORMED ON (B)(6) 2020.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22403731&pc=DTQ | 9611109-2025-90225 | 2025-06-06 04:00 | Death | LIVANOVA DEUTSCHLAND | 2025-07-04 04:00 | DTQ | S5 CONSOLE | Failure to Auto Stop | Cardiac Arrest |  |  | 1 | Event Description: LIVANOVA RECEIVED A REPORT THAT THE ARTERIAL S5 ROLLER PUMP 150 DIDN'T STOP AT A LEVEL ALARM. A DETAILED DESCRIPTION OF THE INCIDENT WAS PROVIDED: THE PATIENT HAD BEEN ON CARDIOPULMONARY BYPASS FOR ABOUT 20 MINUTES. A NEW CARDIOPLEGIA WAS REQUESTED BY THE SURGEON, WHICH WAS PERFORMED. THE SURGEON ALSO REQUESTED THAT THE HEART BE FILLED, WHICH WAS PERFORMED AS WELL. THE VOLUME IN THE CARDIOTOMY DROPPED TO THE LEVEL SENSOR. THE SENSOR TRIGGERED A "LOW LEVEL" ALARM. THE USER ACKNOWLEDGED THE ALARM AND SAW THE LEVEL CONTINUE TO DROP BELOW THE SENSOR LEVEL, WHICH ACCORDING TO CUSTOMER WAS ABNORMAL, AND THE PUMP CONTINUED TO RUN AT FULL FLOW, WITH THE CARDIOTOMY ALMOST EMPTY. ARTERY AND VAIN WERE URGENTLY CLAMPED. AIR BUBBLES WERE PRESENT IN THE ARTERY UP TO THE BUBBLE SENSOR. TWO OTHER PERFUSIONISTS WERE CALLED FOR EMERGENCY ASSISTANCE. MEANWHILE, THE PATIENT WAS NOT RECEIVING PERFUSION. THE SURGEON IMMEDIATELY UNCLAMPED THE HEART AND PERFORMED INTERNAL CARDIAC MASSAGE. ON THE ANESTHETIC SIDE, RESUSCITATION WAS STARTED. A LOW FLOW RATE WAS OBSERVED FOR SEVERAL MINUTES. THE CONSUMABLE ON THE CARDIOTOMY, ALL TUBING AND OXYGENATOR WERE NOT DAMAGED AND WERE FUNCTIONAL. THANKS TO THE ACTION OF THE PERFUSIONISTS (RENEWED CRYSTALLOID DEBULKING, DISCONNECTION AND PURGING OF THE ARTERY BY THE SURGEON), THE CARDIOPULMONARY BYPASS COULD BE RESUMED 9 MINUTES LATER. ON JUNE 12, 2025, LIVANOVA WAS INFORMED THAT THE PATIENT HAS DIED, AND THE HEART LUNG MACHINE HAS BEEN SEGREGATED BY THE POLICE.Event Description: SEE INITIAL REPORT. Manufacturer Narrative: BASED ON TESTS CONDUCTED ON THE DEVICE AND LOG FILE ANALYSIS, THE CAUSE OF THE REPORTED INCIDENT IS TRACED TO THE USER ACTIVATING THE SELECTIVE OVERRIDE ON THE LEVEL MONITORING FUNCTION, THEREFORE PREVENTING THE PUMP TO STOP AT A LEVEL ALARM. A DETAILED EXPLANATION OF THE EXPECTED BEHAVIOUR OF THE DEVICE AT A LEVEL ALARM AND INVESTIGATION FINDINGS IS PROVIDED BELOW. TECHNICAL INSPECTION AND FUNCTIONAL TESTS RESULTS: BASED ON THE TESTS PERFORMED ON THE MACHINE BY A LIVANOVA SERVICE REPRESENTATIVE, ALL SENSORS (LEVEL, BUBBLE, AND PRESSURE) WERE TESTED AND WERE FOUND TO BE WORKING AS PER SPECIFICATIONS. THE MONITORING FUNCTION CONTROL ON THE PUMP WAS TESTED AS WELL AND THE PUMP FUNCTIONED AS INTENDED (THE PUMP STOPS WHEN ALARMS ARE TRIGGERED). LOG FILE ANALYSIS RESULTS: THE ANALYSIS OF THE LOG FILE REVEALED THAT RIGHT BEFORE THE EVENT REPORTED BY THE CUSTOMER OCCURRED, A LEVEL ALARM CORRECTLY TRIGGERED A PUMP STOP. AFTER THAT, THE LEVEL ALARM TONE WAS SWITCHED OFF AND BACK ON BY THE USER. THEN, THE OVERRIDE OF THE LEVEL MONITORING FUNCTION (SELECTIVE OVERRIDE) WAS ACTIVATED BY THE USER (AT 10:32:13 A.M.) DISABLING THE LEVEL CONTROL ON THE PUMP. IN CASE THE OVERRIDE IS ACTIVATED, THE PUMP IS EXPECTED NOT TO STOP WHEN A LEVEL ALARM IS ISSUED AND THIS CONDITION LASTS 5 MINUTES. AFTER THE ACTIVATION OF THE OVERRIDE ON THE LEVEL MONITORING AND CONTROL, PRESSURE AND BUBBLE ALARMS WERE ISSUED AND THE PUMP CORRECTLY STOPPED AS EXPECTED, SINCE THE OVERRIDE FUNCTION WAS APPLIED BY THE USER ONLY ON THE LEVEL MONITORING FUNCTION. AT 10:33:02 A.M. THE BUBBLE SENSOR MONITORING FUNCTION WAS SWITCHED OFF. THE SELECTIVE OVERRIDE OF THE LEVEL MONITORING AND CONTROL WAS AUTOMATICALLY DEACTIVATED AFTER 5 MINUTES (AT 10:37:12 A.M.). THE CONCLUSION OF DEVICE LOG ANALYSIS IS THAT THE MACHINE WORKED AND BEHAVED AS PER SPECIFICATIONS AND DID NOT MALFUNCTION. THIS IS IN ACCORDANCE WITH THE FINDINGS OF THE TESTS PERFORMED BY THE LIVANOVA SERVICE REPRESENTATIVE. INVESTIGATION CONCLUSION THE DEVICE OPERATED AS EXPECTED, TRIGGERING THE LEVEL ALARM WHEN REQUIRED. THE USER VOLUNTARILY DEACTIVATED THE LEVEL ALARM AND CONTINUED OPERATION FOR FIVE MINUTES WITHOUT LEVEL MONITORING. A SIGNIFICANT DROP IN THE RESERVOIR LEVEL MAY FACILITATE AIR TO ENTER THE OXYGENATOR. SUBSEQUENTLY, PRESSURE AND BUBBLE ALARMS WERE ISSUED, AND THE PUMP WAS CORRECTLY STOPPED AS EXPECTED. EVENTUALLY, ALSO THE BUBBLE SENSOR MONITORING FUNCTION WAS SWITCHED OFF. BASED ON THE INFORMATION CURRENTLY AVAILABLE, WE CAN REASONABLY EXCLUDE A DEVICE MALFUNCTION AS THE CAUSE OR CONTRIBUTING FACTOR TO THE CLINICAL EVENT. Manufacturer Narrative: H11: LIVANOVA DEUTSCHLAND MANUFACTURES THE S5 CONSOLE. THE INCIDENT OCCURRED IN NATES, FRANCE. THE MODEL NUMBER OF THE ROLLER PUMP INVOLVED IS CODE 10-80-00 AND THE SERIAL NUMBER IS (B)(6). BASED ON THE TESTS PERFORMED ON THE MACHINE BY A LIVANOVA SERVICE REPRESENTATIVE, ALL SENSORS (LEVEL, BUBBLE, AND PRESSURE) WERE TESTED AND WERE FOUND TO BE WORKING AS PER SPECIFICATIONS. THE MONITORING FUNCTION CONTROL ON THE PUMP WAS TESTED AS WELL AND THE PUMP FUNCTIONED AS INTENDED (THE PUMP STOPS WHEN ALARMS ARE TRIGGERED). THE ANALYSIS OF THE LOG FILE REVEALED THAT RIGHT BEFORE THE EVENT REPORTED BY THE CUSTOMER OCCURRED, A LEVEL ALARM CORRECTLY TRIGGERED A PUMP STOP. AFTER THAT, THE LEVEL ALARM TONE WAS SWITCHED OFF AND BACK ON BY THE USER. THEN, THE OVERRIDE OF THE LEVEL MONITORING FUNCTION (SELECTIVE OVERRIDE) WAS ACTIVATED BY THE USER (AT 10:32:13 A.M.) DISABLING THE LEVEL CONTROL ON THE PUMP. IN CASE THE OVERRIDE IS ACTIVATED, THE PUMP IS EXPECTED NOT TO STOP WHEN A LEVEL ALARM IS ISSUED AND THIS CONDITION LASTS 5 MINUTES. AFTER THE ACTIVATION OF THE OVERRIDE ON THE LEVEL MONITORING AND CONTROL, PRESSURE AND BUBBLE ALARMS WERE ISSUED AND THE PUMP CORRECTLY STOPPED AS EXPECTED, SINCE THE OVERRIDE FUNCTION WAS APPLIED BY THE USER ONLY ON THE LEVEL MONITORING FUNCTION. AT 10:33:02 A.M. THE BUBBLE SENSOR MONITORING FUNCTION WAS SWITCHED OFF. THE SELECTIVE OVERRIDE OF THE LEVEL MONITORING AND CONTROL WAS AUTOMATICALLY DEACTIVATED AFTER 5 MINUTES (AT 10:37:12 A.M.). THE CONCLUSION OF DEVICE LOG ANALYSIS IS THAT THE MACHINE WORKED AND BEHAVED AS PER SPECIFICATIONS AND DID NOT MALFUNCTION. THIS IS IN ACCORDANCE WITH THE FINDINGS OF THE TESTS PERFORMED BY THE LIVANOVA SERVICE REPRESENTATIVE. MEDICAL ASSESSMENT HELD WITH LIVANOVA CLINICAL EXPERTS ALSO CONFIRMS USER ERROR AND HIGHLIGHTED THAT SINCE THE BUBBLE SENSOR STOPPED THE PUMP, AIR LIKELY DID NOT ENTER THE PATIENT AND IS LIKELY THAT THE 9 MINUTES OF LACK OF PERFUSION LED TO PATIENT HARM. HOWEVER, IT IS CLEAR CONSEQUENCE OF INCORRECT MANAGEMENT OF THE SYSTEM BY THE USER. BASED ON TESTS CONDUCTED ON THE DEVICE AND LOG FILE ANALYSIS, THE CAUSE OF THE REPORTED INCIDENT IS TRACED TO THE USER ACTIVATING THE SELECTIVE OVERRIDE ON THE LEVEL MONITORING FUNCTION, THEREFORE PREVENTING THE PUMP TO STOP AT A LEVEL ALARM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22392002&pc=DTE | 2182208-2025-03199 | 2024-06-02 04:00 | Death | MEDTRONIC, INC. | 2025-07-03 04:00 | DTE | DUAL CHAMBER TEMPORARY PACEMAKER | Insufficient Information | Unspecified Heart Problem | K150246 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS DECEASED POST CARDIAC SURGERY. IT WAS UNKNOWN WHETHER THE EPG CONTRIBUTED TO THE EVENT. THE DATE OF PASSING WAS UNKNOWN AND NO FURTHER DETAILS ARE AVAILABLE.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS DECEASED ON (B)(6) 2024. RESUSCITATION PROTOCOLS THAT INCLUDE ADVANCED CARDIAC LIFE SUPPORT (ACLS) WAS TAKEN AT THE TIME OF THE EVENT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: B3. DATE OF EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22393654&pc=FTL | 1213643-2025-00554 | 2019-10-23 04:00 | Death | DAVOL INC., SUB. C.R. BARD, INC. -1213643 | 2025-07-03 04:00 | FTL | BARD FLAT MESH | Patient-Device Incompatibility | Fistula; Unspecified Infection; Bowel Perforation | PREAMENDMENT |  | 1 | Event Description: PER MEDWATCH REPORT MW5171291: "MY NAME IS (B)(6) THE SON OF (B)(6) WHO DIED AS A RESULT OF DEFECTIVE HERNIA MESH PRODUCTS, SPECIFICALLY BARD DOVAL AND COVIDIEN AS WELL AS A THIRD PRODUCT I BELIEVE. I WAS MY FATHER'S CARE PROVIDER FOR THE LAST THREE YEARS OF HIS LIFE AS HIS CONDITION BECAME SEVERE AND DEBILITATING. IN 2019 HE WENT IN FOR EMERGENCY SURGERY DUE TO A LIFE-THREATENING BOWEL BLOCKAGE THAT BECAME SEPTIC. IT TOOK 11 DAYS FOR HIM TO WAKE UP AND HE NEVER WALKED AGAIN FOR THE LAST TWO YEARS OF HIS LIFE. WHEN I MET WITH THE SURGEON DR. (B)(6) HE STATED "THIS WAS ONE OF THE MOST DIFFICULT SURGERIES HE EVER PERFORMED". BEAR IN MIND DR. (B)(6) HAD BEEN PERFORMING SURGERIES FOR 35 YEARS TO PUT THAT STATEMENT IN PERSPECTIVE. THE PRODUCTS USED IN MY FATHER'S HERNIA SURGERIES DESTROYED HIS INTERNAL ORGANS TO THE POINT THEY HAD TO BE REMOVED AND RESECTED. HIS ORGANS WERE DEAD, DYING, PERFORATED, AND SO ON. EVERYTHING I SPEAK ABOUT IN THIS EMAIL IS FULLY DOCUMENTED IN HIS MEDICAL RECORDS. I PERSONALLY SPENT 3 YEARS READING THROUGH EVERY PAGE AND GOING OVER THEM WITH A FINE TOOTH COMB. I SOLD MY HOME AND LEFT MY CAREER AS A MORTGAGE COMPLIANCE OFFICER TO LIVE WITH MY FATHER THE LAST 2-2.5 YEARS OF HIS LIFE TO TAKE CARE OF HIM 24/7 AS MY MOTHER, HIS WIFE PASSED AWAY IN 2011. AFTER HIS LAST HERNIA MESH REVISION SURGERY IN 2019 DUE TO A BOWEL BLOCKAGE HE NEVER RECOVERED. HE SUFFERED EXTREME NEUROPATHY, FUNCTIONAL QUADRIPLEGIA, LARGE ABDOMINAL WOUNDS WITH FISTULAS THAT I HAD TO PACK EVERY DAY, NUMEROUS UTI'S FROM HIS CATHETER, ETC... IN THE WEEKS FOLLOWING HIS SURGERY HIS SURGEON HAD TO PULL MESH THROUGH HIS FISTULAS DURING AN OFFICE VISIT DUE TO THE EXTREME PAIN HE WAS IN. I SPENT TWO YEARS LISTENING TO HIM CRYING OUT IN PAIN AND SOBBING ON A DAILY BASIS, AND UNLESS YOU'VE WATCHED A PARENT SUFFER LIKE THAT YOU HAVE NO IDEA WHAT IT DOES TO YOU. HIS WOUNDS NEVER HEALED DUE TO THE MESH INFECTION. I HAD TO CALL 911 OVER 20 TIMES IN THE LAST YEAR OF HIS LIFE AND HE HAD TO ENDURE THE PAIN OF TRAVELING TO THE HOSPITAL AND THE PAIN WAS EXCRUCIATING. HIS QUALITY OF LIFE WAS GONE COMPLETELY. MY FATHER DIED A VERY SLOW, AGONIZING, DEPRESSIVE, DEMORALIZING, TORTUOUS DEATH, TO PUT IT MILDLY. THESE ARE FACTS. TRUST ME, I HAD THE TRAUMATIZING TASK OF BEING THERE 24/7 THROUGHOUT THE ENTIRE 2 PLUS YEAR ORDEAL. MAKE NO MISTAKE HE SUFFERED FROM EVERY HERNIA MESH COMPLICATION YOU CAN FIND ON GOOGLE PLUS THE EXTREME NEUROPATHY, FUNCTION QUADRIPLEGIA, AND DEATH. HE COULD NOT WALK, FEED HIMSELF, BATHE, USE THE TOILET, SCRATCH HIS NOSE, PET HIS DOG, OR ANYTHING ELSE FOR THAT MATTER. NOBODY SHOULD EVER HAVE TO GO THROUGH THE HELL MY FATHER, MYSELF, AND MY FAMILY WENT THROUGH DUE TO THESE MEDICAL PRODUCTS. IT WAS A TRAGEDY OF EPIC PROPORTIONS FOR ALL OF US TO SAY THE VERY LEAST. REFERENCE REPORTS MW5171292, MW5171293." Manufacturer Narrative: THE INFORMATION OBTAINED IS LIMITED TO THE RECEIVED MEDWATCH REPORT. MULTIPLE ATTEMPTS HAVE BEEN MADE TO TRY AND OBTAIN ADDITIONAL INFORMATION WITH NO RESPONSE RECEIVED. AS REPORTED, PATIENT HAD A COMPLEX MEDICAL HISTORY WITH MULTIPLE IMPLANTS AND SURGICAL PROCEDURES. BASED ON THE INFORMATION PROVIDED, NO CONCLUSION CAN BE MADE AS TO THE DEGREE TO WHICH THE BARD/BD MESH IMPLANT MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT¿S ALLEGED CLINICAL COURSE. INFECTION AND FISTULA ARE KNOWN INHERENT RISKS OF HERNIA REPAIR SURGERY AND ARE LABELED AS POSSIBLE COMPLICATIONS IN THE INSTRUCTIONS-FOR-USE, SUPPLIED WITH THE DEVICE. REGARDING INFECTION THE WARNINGS SECTION STATES "IF AN INFECTION DEVELOPS, TREAT THE INFECTION AGGRESSIVELY. CONSIDERATION SHOULD BE GIVEN REGARDING THE NEED TO REMOVE THE MESH. AN UNRESOLVED INFECTION MAY REQUIRE REMOVAL OF THE MESH." NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. NOTE, THE ACTUAL DATE OF DEATH WAS NOT PROVIDED IN THE REPORT SENT TO BD. WE HAVE DOCUMENTED THIS REQUIRED DATE FIELD WITH A BEST ESTIMATE. NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22393740&pc=LWS | 2649622-2025-17983 | 2025-06-07 04:00 | Death | MPRI | 2025-07-03 04:00 | LWS | SPRINT QUATTRO SECURE MRI SURESCAN | Under-Sensing | Cardiac Arrest; Tachycardia; Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS RIGHT ATRIAL (RA) LEAD UNDERSENSING ON STORED EPISODES. IT WAS ALSO NOTED THERE WAS INTERMITTENT UNDERSENSING ON THE RIGHT VENTRICULAR (RV) LEAD. THE PATIENT HAD A CARDIAC ARREST, UNSTABLE VENTRICULAR TACHYCARDIA (VT), VENTRICULAR FIBRILLATION (VF), AND PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: 407652 LEAD IMPLANTED (B)(6) 2007. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22394144&pc=BZD | 2518422-2025-106747 | 2025-06-04 04:00 | Death | RESPIRONICS, INC. | 2025-07-03 04:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF REP DREAMSTATION AUTO CPAP, DOM - RECRT. PATIENT'S SPOUSE CONFIRMING THAT THE AFFECTED DEVICE WAS RECEIVED BACK AND MENTIONED THAT THE PATIENT PASSED AWAY 2 YEARS AGO AND INQUIRE ABOUT ELS. CALLER IS MADE AWARE THAT A COURT APPROVED THIRD-PARTY SETTLEMENT ADMINISTRATOR AND IS ADVISED THAT TO COMMUNICATE TO ANGELION GROUP REGARDING SETTLEMENT CLAIMS. A CONTACT NUMBER IS PROVIDED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22394191&pc=NKE | 9614453-2025-02565 | 2025-05-31 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-07-03 04:00 | NKE | SOLARA QUAD CRT-P MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P010015 |  | 1 | Event Description: IT WAS REPORTED THAT 10 DAYS POST IMPLANTABLE PULSE GENERATOR (IPG) TO CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) UPGRADE, THE PATIENT PASSED AWAY. ADDITIONAL INFORMATION WAS REQUESTED BUT WAS NOT AVAILABLE; HOWEVER, THE PATIENT'S OBITUARY INDICATED THE PATIENT PASSED AWAY AFTER A LONG ILLNESS. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 459888 LEAD IMPLANTED (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22394194&pc=OJX | 2649622-2025-17996 | 2025-05-31 04:00 | Death | MPRI | 2025-07-03 04:00 | OJX | ATTAIN PERFORMA S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT 10 DAYS POST IMPLANTABLE PULSE GENERATOR (IPG) TO CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) UPGRADE, THE PATIENT PASSED AWAY. ADDITIONAL INFORMATION WAS REQUESTED BUT WAS NOT AVAILABLE; HOWEVER, THE PATIENT'S OBITUARY INDICATED THE PATIENT PASSED AWAY AFTER A LONG ILLNESS. Manufacturer Narrative: CORRECTION: D6B. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. THE ANALYST NOTED A PROXIMAL SEGMENT OF THE LEAD 6.3 CM LONG WAS RETURNED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID W4TR03 CRT-P IMPLANTED (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22394250&pc=CCW | 3010244187-2025-00569 | 2025-05-29 04:00 | Death | MEMPHIS RTG CONTRACT | 2025-07-03 04:00 | CCW | MCGRATH | Failure to Run on Battery | Asystole | EXEMPT |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, TWO VIDEO LARYNGOSCOPE BATTERIES SHOWED INACCURATE REMAINING MINUTES AND THEN FLASHED BATTERY EMPTY SIGN. PATIENT WAS DECEASED DUE TO INJURIES, AND NOT ASSOCIATED WITH THE BATTERY FAILURES. Manufacturer Narrative: D10 CONCOMITANT PRODUCT: MCGRATH 3.6V BATTERY 340-000-000 LOT: H23051604. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22394280&pc=CCW | 3010244187-2025-00570 | 2025-05-29 04:00 | Death | MEMPHIS RTG CONTRACT | 2025-07-03 04:00 | CCW | MCGRATH | Failure to Run on Battery | Asystole | EXEMPT |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE, TWO VIDEO LARYNGOSCOPE BATTERIES SHOWED INACCURATE REMAINING MINUTES AND THEN FLASHED BATTERY EMPTY SIGN. PATIENT WAS DECEASED DUE TO INJURIES, AND NOT ASSOCIATED WITH THE BATTERY FAILURES. Manufacturer Narrative: D10 CONCOMITANT PRODUCT: MCGRATH 3.6V BATTERY 340-000-000 LOT: H23090304. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22394313&pc=LWQ | 1649833-2025-00024 | 2025-05-23 04:00 | Death | ARTIVION, INC. ¿ AUSTIN | 2025-07-03 04:00 | LWQ | ON-X PROSTHETIC AORTIC VALVE WITH CONFORM-X SEWING RING AND EXTENDED HOLDER (ONX | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P000037 |  | 1 | Event Description: ACCORDING TO THE INITIAL NOTIFICATION ON (B)(6) 2025, "THE PATIENT UNDERWENT SURGERY FOR A MITRAL VALVE REPLACEMENT. A NEW DOCTOR, WHO HAS ONLY RECENTLY STARTED WORKING AT THE HOSPITAL, PERFORMED THE OPERATION. THE SURGEON IS FAMILIAR WITH THE ONX VALVE FROM HIS PREVIOUS WORK AT ANOTHER CLINIC. THE CONSIGNMENT STOCK AT THE HOSPITAL ONLY CONTAINS ONX VALVES FOR THE AORTIC POSITION, DIFFERENT SIZES OF ONXACE. DESPITE THE REQUEST FOR AN ONX MITRAL VALVE, THE OR NURSING STAFF HANDED THE SURGEON AN ONX AORTIC VALVE. NEITHER THE PERSON HANDING OVER THE VALVE, NOR THE STERILE OR NURSE, NOR THE SURGEON NOTICED THAT IT WAS THE WRONG PRODUCT. ACCORDING TO THE REPORTING DOCTOR, IT WAS A HUMAN ERROR ON OF THE ENTIRE SURGICAL TEAM INVOLVED AND THE CASE IS BEING INVESTIGATED INTERNALLY AT THE HOSPITAL NOW. IMAGES OF BOTH THE OUTER PACKAGING AND THE INNER PACKAGING OF THE ONX VALVES WERE REQUESTED FOR THIS PURPOSE." Manufacturer Narrative: THIS INVESTIGATION IS CURRENTLY ONGOING. ANY ADDITIONAL INFORMATION WILL BE PROVIDED IN THE FOLLOW-UP REPORT. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22394411&pc=QJZ | 3012172416-2025-00059 | 2025-04-27 04:00 | Death | XENIOS AG | 2025-07-03 04:00 | QJZ | XLUNG KIT 230 | Insufficient Information | Hemolysis | K191407 |  | 1 | Event Description: A USER FACILITY REPORTED, ON (B)(6) 2025, FROM 1:30 P.M. TO 10:00 P.M., THE DEVICE FUNCTIONED PROPERLY ON A PATIENT UNDER MEDICALLY INDUCED COMA. AROUND 10:00 P.M., A MALFUNCTION OCCURRED ACCOMPANIED BY SEVERAL ALERT MESSAGES: CLOT DETECTED IN THE CIRCUIT, AIR BUBBLE DETECTED, AND PUMP HEAD STOPPED. A DE-BUBBLING PROCEDURE WAS INITIATED BY THE USERS, BUT THE SAME ERROR MESSAGES REAPPEARED. ANTICOAGULATION PROTOCOL IN PLACE: A TOTAL OF (B)(4) IU OF SODIUM HEPARIN WAS USED TO FLUSH THE CIRCUIT, AND THE PATIENT WAS RECEIVING A CONTINUOUS INFUSION OF HEPARIN AT (B)(4) IU/DAY AS PART OF THE ANTICOAGULATION PROTOCOL. THE PATIENT EXPIRED. THE FACILITY MENTIONED THAT THERE IS NO DIRECT CAUSAL LINK ESTABLISHED, ALTHOUGH THE INTENSIVE CARE UNIT TEAM INDICATED THAT THE EVENT CONTRIBUTED TO THE OUTCOME, AS THE PATIENT EXPERIENCED HEMOLYSIS DUE TO THE EVENT. THE FACILITY WAS NOT ABLE TO PROVIDE KIT INFORMATION. THE FACILITY DID NOT MENTION ANY PART OF THE CIRCUIT THAT COULD HAVE LED TO THE EVENT. THE COMPLAINT SAMPLE WAS NOT AVAILABLE FOR PRODUCT EVALUATION. UPON FOLLOW-UP FROM THE XENIOS COMPLAINT MANAGER, IT WAS REPORTED THIS PATIENT WAS PLACED ON ECMO SUPPORT ON AN UNDETERMINED DATE FOR AN UNREPORTED CAUSE. THE XLUNG KIT 230 AND CONSOLE WERE EXCHANGED FOLLOWING THIS EVENT. THE PATIENT EXPIRED DURING ECMO SUPPORT THE FOLLOWING DAY DUE TO AN UNREPORTED CAUSE. IT WAS CONFIRMED THAT THE PATIENT¿S DEATH WAS NOT A DIRECT RESULT OF A MALFUNCTION OR DEFICIENCY OF ANY XENIOS PRODUCT(S) OR DEVICE(S). Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22394433&pc=NVZ | 3008973940-2025-04792 | 2025-06-03 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-07-03 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED ONE WEEK POST IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM, AS PER THE HEALTHCARE PROFESSIONAL, THE PATIENT'S P-WAVES DECREASED AND THE RIGHT ATRIAL (RA) LEAD THRESHOLD HAS INCREASED SINCE IMPLANT. THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. (B)(6) 2025: IT WAS FURTHER REPORTED THE PATIENT PASSED AWAY APPROXIMATELY TWO WEEKS POST IMPLANT OF THE IPG SYSTEM. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22394457&pc=DTB | 2649622-2025-18005 | 2025-06-03 04:00 | Death | MPRI | 2025-07-03 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED ONE WEEK POST IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM, AS PER THE HEALTHCARE PROFESSIONAL, THE PATIENT'S P-WAVES DECREASED AND THE RIGHT ATRIAL (RA) LEAD THRESHOLD HAS INCREASED SINCE IMPLANT. THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. (B)(6) 2025: IT WAS FURTHER REPORTED THE PATIENT PASSED AWAY APPROXIMATELY TWO WEEKS POST IMPLANT OF THE IPG SYSTEM. ADDITIONAL INFOR MATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22394776&pc=MFA | 3007284006-2025-00122 | 2025-06-09 04:00 | Death | SPECTRANETICS CORPORATION | 2025-07-03 04:00 | MFA | GLIDELIGHT LASER SHEATH | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation; Cardiac Perforation; Pericardial Effusion | P960042 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV), A RIGHT ATRIAL (RA), AND LEFT VENTRICULAR (LV) LEADS DUE TO CIED SYSTEM/POCKET INFECTION. SPECTRANETICS LLD LEAD LOCKING DEVICES (LLDS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 16F GLIDELIGHT LASER SHEATH ON THE RA LEAD, STEADY PROGRESS WAS MADE, AND THE LEAD WAS REMOVED. THEN, WHEN REMOVING THE GLIDELIGHT FROM THE BODY, THE PATIENT'S BLOOD PRESSURE DROPPED, AND A PERICARDIAL EFFUSION WAS DETECTED VIA TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). RESCUE EFFORTS BEGAN, INCLUDING RESCUE BALLOON, BYPASS, AND STERNOTOMY. A PERFORATION IN THE SVC/RA JUNCTION WAS DISCOVERED; HOWEVER, WAS UNABLE TO BE REPAIRED. THE RV AND LV LEADS REMAINED, AND THE PATIENT DID NOT SURVIVE THE PROCEDURE. THIS REPORT CAPTURES THE 16F GLIDELIGHT IN USE WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION, BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: D4/H4) THE LOT NUMBER WAS NOT PROVIDED; THUS, THE FOLLOWING INFORMATION IS UNKNOWN: DEVICE SERIAL NUMBER, UNIQUE ID, EXPIRATION DATE, AND MANUFACTURE DATE. H3) THE DEVICE WAS DISCARDED; THUS, NO INVESTIGATION COULD BE COMPLETED. H6) GREAT VESSEL, CARDIAC PERFORATION, AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE GLIDELIGHT DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22395044&pc=OZD | 1220648-2025-30041 | 2025-06-19 04:00 | Death | ABIOMED, INC. | 2025-07-03 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED MANY SUCTION EVENTS. AN ECHOCARDIOGRAM SHOWED THE IMPELLA PLACEMENT WAS GOOD. IN ADDITION, THERE WAS BLEEDING FROM THE ACCESS SITE. FOUR UNITS OF PACKED RED BLOOD CELLS WERE TRANSFUSED AND A FEMSTOP WAS PLACED. ULTIMATELY, THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW AND VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. VASCULAR DAMAGE - PERIPHERAL VESSEL: PATIENT NOTED TO HAVE COAGULOPATHY. THE CAUSE OF THE VASCULAR DAMAGE - PERIPHERAL VESSEL WAS LIKELY PATIENT CONDITION RELATED AS PATIENT HAD COAGULATION DISORDER PER CLINICAL DETAILS. LOW PUMP FLOW: THE PUMP WAS NOT RETURNED. DATA LOG REVIEW SHOWED SUCTION ALARMS OCCURRING WITH LOWER THAN EXPECTED FLOW RATES, AT INSTANCES OF P-7, FLOWS FELL BELOW THE EXPECTED 2.9L/MIN, AS LOW AS 1.9L/MIN. P-LEVELS WERE REDUCED AND GRADUALLY BROUGHT BACK UP. THE CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY PATIENT CONDITION RELATED DUE TO THE NOTED BLEEDING OCCURRING WITH SUCTION EVENTS THAT RESOLVED WITH REDUCING THE P-LEVEL AND GIVING VOLUME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22395058&pc=DRE | 3007284006-2025-00123 | 2025-06-10 04:00 | Death | SPECTRANETICS CORPORATION | 2025-07-03 04:00 | DRE | TIGHTRAIL SUB-C ROTATING DILATOR SHEATH | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K161333 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV), RIGHT ATRIAL (RA), AND TWO LEFT VENTRICULAR (LV) LEADS DUE TO BACTEREMIA. SPECTRANETICS LLD EZ LEAD LOCKING DEVICES (LLD EZS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. USING MULTIPLE SPECTRANETICS DEVICES (16FR GLIDELIGHT LASER SHEATHS, LARGE VISISHEATH DILATOR SHEATHS, A 11F TIGHTRAIL MINI DILATOR SHEATH, 13FR TIGHTRAIL SUB-C DILATOR SHEATHS, 13FR TIGHTRAIL DILATOR SHEATHS (LONG), 14FR GLIDELIGHT, AND 11FR TIGHTRAIL SUB-C DILATOR SHEATH) ON EACH LEAD, THE LEADS STARTED TO BREAK DOWN AND FALL APART, ATTRIBUTED TO THE LEADS BEING EXTREMELY CALCIFIED TOGETHER. AFTER SWITCHING BETWEEN THE LEADS, AND TRYING SEVERAL DIFFERENT TOOLS, THE BLEEDING FROM THE POCKET INCREASED, AND THE PATIENT¿S BLOOD PRESSURE DROPPED. AT THIS POINT, THE LEAD EXTRACTION WAS ABORTED, AND RESCUE EFFORTS BEGAN INCLUDING SUTURING THE POCKET, AND BLOOD PRODUCTS TO CONTROL THE BLOOD LOSS. THE PATIENT STABILIZED; HOWEVER, PASSED AWAY THE FOLLOWING MORNING. PER THE PHYSICIAN, THERE WAS NO INJURY IDENTIFIED AND CAUSE OF DEATH WAS UNCLEAR, BUT MAY BE DUE TO THE BLOOD LOSS DURING THE PROCEDURE. THIS REPORT CAPTURES THE 11F TIGHTRAIL SUB-C DEVICE LAST USED DURING THE PROCEDURE, WHICH RESULTED IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: A2) DATE OF BIRTH - NOT AVAILABLE FROM FACILITY. A4) PATIENT WEIGHT - NOT AVAILABLE FROM FACILITY. B7) OTHER RELEVANT HISTORY - NOT AVAILABLE FROM FACILITY. D4/H4) THE LOT NUMBER WAS NOT PROVIDED; THUS, THE FOLLOWING INFORMATION IS UNKNOWN: UNIQUE ID, EXPIRATION DATE, AND MANUFACTURE DATE. H3) THE TIGHTRAIL WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) DEATH IS A KNOWN RISK OF COMPLICATION WITH USE OF THE TIGHTRAIL DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22395087&pc=OZD | 1220648-2025-30042 | 2025-06-25 04:00 | Death | ABIOMED, INC. | 2025-07-03 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT WAS IMPLANTED WITH A IMPELLA CP FOR SUPPORT OF ACUTE MYOCARDIAL INFARCTION. IT WAS REPORTED THAT DURING CARDIOPULMONARY RESUSCITATION, THE PATIENT EXPERIENCED PULMONARY BLEEDING AND HEMOTHORAX. THE BLEEDING COULD NOT BE CONTROLLED AND HAD PROGRESSED TO MEDIASTINAL SHIFT. IN ADDITION, SUCTION OCCURRED ON THE IMPELLA CP DUE TO INSUFFICIENT INTRAVASCULAR VOLUME. IT WAS REPORTED THAT SYSTEMIC CIRCULATION COULD NOT BE MAINTAINED DUE TO THE BLEEDING, AND BOTH THE EXTRACORPOREAL MEMBRANE OXYGENATION AND IMPELLA CP WERE UNABLE TO PROVIDE SUFFICIENT FLOW, LEADING TO THE PATIENT'S DEATH FROM CIRCULATORY FAILURE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22395094&pc=NPT | 2025587-2025-04763 | 2025-06-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-03 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 9 YEARS AND 5 MONTHS AFTER THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC PULMONARY VALVE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNSPECIFIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22395399&pc=DRE | 3007284006-2025-00124 | 2025-06-18 04:00 | Death | SPECTRANETICS CORPORATION | 2025-07-03 04:00 | DRE | TIGHTRAIL ROTATING DILATOR SHEATH | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation; Cardiac Perforation | K150360 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV), RIGHT ATRIAL (RA), AND LEFT VENTRICULAR (LV) LEAD DUE TO BACTEREMIA. SPECTRANETICS LLD EZ LEAD LOCKING DEVICES (LLD EZS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 12F GLIDELIGHT LASER SHEATH, 14F GLIDELIGHT, AND A SPECTRANETICS 13F TIGHTRAIL ROTATING DILATOR SHEATH, THE RA AND THE LV LEADS WERE REMOVED. THEN, USING A 16F GLIDELIGHT ON THE RV LEAD, ADVANCEMENT WAS MADE TO THE INFERIOR VENA CAVA (IVC)/RA JUNCTION, WHEN RESISTANCE WAS MET. SWITCHING TO THE 13F TIGHTRAIL, FURTHER PROGRESS WAS UNSUCCESSFUL, AND AFTER BACKING OFF AND RE-ENGAGING THE LESION, THE TIGHTRAIL MOVED FORWARD, AND THE LEAD WAS REMOVED. HOWEVER, THE PATIENT'S BLOOD PRESSURE DROPPED, AND RESCUE EFFORTS BEGAN, INCLUDING RESCUE BALLOON, CPR, ADMINISTRATION OF BLOOD PRODUCTS AND FLUIDS, DEFIBRILLATOR, AND STERNOTOMY. A SUPERIOR VENA CAVA (SVC)/RA JUNCTION PERFORATION WAS DISCOVERED AND REPAIRED. DESPITE REPAIR AND EXTENSIVE LIFESAVING EFFORTS, THE PATIENT WAS DECLARED BRAIN DEAD DUE TO BLOOD LOSS. THE PATIENT DID NOT SURVIVE. THIS REPORT CAPTURES THE 13F TIGHTRAIL IN USE IN THE AREA WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION, BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: A2) PATIENT DATE OF BIRTH - NOT AVAILABLE FROM FACILITY. A3B) PATIENT GENDER - NOT AVAILABLE FROM FACILITY. A5) PATIENT ETHNICITY - NOT AVAILABLE FROM FACILITY. B6/B7) PATIENT INFORMATION REGARDING RELEVANT TESTS/LABORATORY DATA OR MEDICAL HISTORY - NOT AVAILABLE FROM FACILITY. D4/H4) THE LOT NUMBER WAS NOT PROVIDED, THUS THE FOLLOWING INFORMATION IS UNKNOWN: UNIQUE ID, EXPIRATION DATE, AND MANUFACTURE DATE. H3) THE TIGHTRAIL WAS NOT RETURNED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) GREAT VESSEL PERFORATION, CARDIAC PERFORATION, AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE TIGHTRAIL DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22395410&pc=NIM | 2124215-2025-42986 | 2025-06-03 04:00 | Death | SILK ROAD MEDICAL | 2025-07-03 04:00 | NIM | ENROUTE TRANSCAROTID STENT | Adverse Event Without Identified Device or Use Problem | Muscle Weakness; Hemorrhagic Stroke |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A BRAIN BLEED WITHIN 24 HOURS OF THE TCAR PROCEDURE AND SUBSEQUENTLY PASSED AWAY. THE PATIENT WAS PRESENTED AS SYMPTOMATIC WITH A VERY TIGHT STENOSIS FOR A HIGH SURGICAL RISK FOR THE RIGHT-SIDED TCAR PROCEDURE. THE TCAR PROCEDURE WAS COMPLETED. FOLLOWING THE PROCEDURE, ACCORDING TO THE NURSE PRACTITIONER, THE PATIENT RECEIVED LOVENOX POST-PROCEDURE. THE PATIENT WAS ALERT AND FOLLOWING COMMANDS BUT EXHIBITED SIGNS OF LEFT ARM WEAKNESS, WHICH WAS PRESENT PRIOR TO THE PROCEDURE, ACCORDING TO THE PHYSICIAN. DURING THE NIGHT, THE PATIENT DEVELOPED A BRAIN BLEED, AND THE FOLLOWING MORNING, THE FAMILY OPTED TO REPLACE THE CARE CONSENT WITH COMFORT CARE. A CT SCAN AND NEUROLOGY CONSULTATION WERE ALSO CONDUCTED. THE PATIENT ULTIMATELY EXPIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22395602&pc=BZD | 2518422-2025-107044 | 2025-06-09 04:00 | Death | RESPIRONICS, INC. | 2025-07-03 04:00 | BZD | DREAMSTATION AUTO CPAP W/HUM/CELL, DOM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED THAT A DREAMSTATION AUTO CPAP W/HUM/CELL, DOM DEVICE PATIENT HAS PASSED AWAY. IT WAS NOT SUGGESTED THAT THE DEVICE CAUSED THE DEATH. THE PATIENT'S SPOUSE CONTACTED THE MANUFACTURER REGARDING RECEIVING A CHECK FOR THE ECONOMIC SETTLEMENT. WHILE NO ASSERTIONS HAVE BEEN MADE ABOUT THE DEVICE CAUSING THE PATIENT'S DEATH, FURTHER DETAILS ARE NECESSARY. A REPORT WILL BE SUBMITTED TO ALL RELEVANT AGENCIES, AND IF MORE INFORMATION BECOMES AVAILABLE, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22395797&pc=NIK | 9614453-2025-02570 | 2025-06-11 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-07-03 04:00 | NIK | COBALT¿ HF QUAD CRT-D MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Aspiration/Inhalation; Sepsis | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED SEPSIS AND POSSIBLE ASPIRATION. THE PATIENT WAS REPORTED TO BE DECEASED. THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) FOR HIGH-RATE NON-SUSTAINED EPISODES AND SENSING INTEGRITY COUNTER (SIC). THE RIGHT VENTRICULAR (RV) LEAD ALSO EXHIBITED LOSS OF CAPTURE DURING A ANTI-TACHYCARDIA PACING (ATP) TREATED EPISODE. IT WAS UNABLE TO BE DETERMINED IF THE LIA WAS A PERFORMANCE ISSUE OR DUE TO AN AGONAL RHYTHM. ADDITIONAL INFORMATION RELATED TO THE CAUSE AND CIRCUMSTANCES OF DEATH WERE REQUESTED AND NOT RECEIVED. Manufacturer Narrative: CONTINUATION OF D10: 479888 LEAD IMPLANT DATE (B)(6) 2023. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22395928&pc=DTB | 2649622-2025-18039 | 2025-06-11 04:00 | Death | MPRI | 2025-07-03 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Aspiration/Inhalation; Sepsis | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED SEPSIS AND POSSIBLE ASPIRATION. THE PATIENT WAS REPORTED TO BE DECEASED. THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) FOR HIGH-RATE NON-SUSTAINED EPISODES AND SENSING INTEGRITY COUNTER (SIC). THE RIGHT VENTRICULAR (RV) LEAD ALSO EXHIBITED LOSS OF CAPTURE DURING A ANTI-TACHYCARDIA PACING (ATP) TREATED EPISODE. IT WAS UNABLE TO BE DETERMINED IF THE LIA WAS A PERFORMANCE ISSUE OR DUE TO AN AGONAL RHYTHM. ADDITIONAL INFORMATION RELATED TO THE CAUSE AND CIRCUMSTANCES OF DEATH WERE REQUESTED AND NOT RECEIVED. Manufacturer Narrative: CONTINUATION OF D10: 6944-65 LEAD IMPLANT DATE (B)(6) 2001. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22395970&pc=LWS | 2649622-2025-18041 | 2025-06-11 04:00 | Death | MPRI | 2025-07-03 04:00 | LWS | SPRINT QUATTRO LEAD | Failure to Capture; Over-Sensing | Aspiration/Inhalation; Sepsis | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED SEPSIS AND POSSIBLE ASPIRATION. THE PATIENT WAS REPORTED TO BE DECEASED. THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) FOR HIGH-RATE NON-SUSTAINED EPISODES AND SENSING INTEGRITY COUNTER (SIC). THE RIGHT VENTRICULAR (RV) LEAD ALSO EXHIBITED LOSS OF CAPTURE DURING AN ANTI-TACHYCARDIA PACING (ATP) TREATED EPISODE. IT WAS UNABLE TO BE DETERMINED IF THE LIA WAS A PERFORMANCE ISSUE OR DUE TO AN AGONAL RHYTHM. ADDITIONAL INFORMATION RELATED TO THE CAUSE AND CIRCUMSTANCES OF DEATH WERE REQUESTED AND NOT RECEIVED. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANT DATE (B)(6) 2001. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22396001&pc=OJX | 2649622-2025-18043 | 2025-06-11 04:00 | Death | MPRI | 2025-07-03 04:00 | OJX | ATTAIN STABILITY QUAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Aspiration/Inhalation; Sepsis | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED SEPSIS AND POSSIBLE ASPIRATION. THE PATIENT WAS REPORTED TO BE DECEASED. THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) FOR HIGH-RATE NON-SUSTAINED EPISODES AND SENSING INTEGRITY COUNTER (SIC). THE RIGHT VENTRICULAR (RV) LEAD ALSO EXHIBITED LOSS OF CAPTURE DURING AN ANTI-TACHYCARDIA PACING (ATP) TREATED EPISODE. IT WAS UNABLE TO BE DETERMINED IF THE LIA WAS A PERFORMANCE ISSUE OR DUE TO AN AGONAL RHYTHM. ADDITIONAL INFORMATION RELATED TO THE CAUSE AND CIRCUMSTANCES OF DEATH WERE REQUESTED AND NOT RECEIVED. Manufacturer Narrative: CONTINUATION OF D10: DTPB2Q1 CRT-D IMPLANT DATE (B)(6) 2023. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22396016&pc=NPT | 2025587-2025-04771 | 2025-07-01 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-03 04:00 | NPT | EVOLUT PRO PLUS DCS | Difficult or Delayed Positioning; Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Perforation; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE VIA LEFT CAROTID ACCESS USING A SELF-EXPANDING VALVE, AN ELDERLY PATIENT WITH A HISTORY OF PERMANENT PACEMAKER IMPLANTATION, ATRIAL FIBRILLATION, AND A FRAIL CLINICAL CONDITION UNDERWENT RECAPTURE MANEUVERS TO OPTIMIZE VALVE POSITIONING. DURING VALVE POSITIONING AND RECAPTURES, IT WAS OBSERVED THAT THE STIFF GUIDEWIRE WAS DEEPLY ADVANCED. BASED ON THE PROCEDURAL SEQUENCE AND FLUOROSCOPIC IMAGES, IT WAS SUSPECTED THAT THE GUIDEWIRE MAY HAVE CONTRIBUTED TO A RIGHT ATRIAL AND/OR LEFT VENTRICULAR PERFORATION, WHICH COULD HAVE RESULTED IN A PERICARDIAL EFFUSION. CARDIAC ARREST OCCURRED, AND A PERICARDIAL EFFUSION AND POSSIBLE RIGHT ATRIAL PERFORATION WERE SUSPECTED. THE PATIENT WAS TRANSFERRED TO THE OPERATING ROOM FOR CARDIAC SURGERY, AND A PERICARDIAL PATCH REPAIR WAS BEING CONSIDERED. DUE TO THE PATIENT¿S ADVANCED AGE AND FRAILTY, THE CARDIAC SURGERY TEAM AND THE FAMILY DECIDED NOT TO PROCEED WITH THE INTERVENTION. THE PATIENT PASSED AWAY.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE VIA LEFT CAROTID ACCESS USING A SELF-EXPANDING VALVE, AN ELDERLY PATIENT WITH A HISTORY OF PERMANENT PACEMAKER IMPLANTATION, ATRIAL FIBRILLATION, AND A FRAIL CLINICAL CONDITION UNDERWENT RECAPTURE MANEUVERS TO OPTIMIZE VALVE POSITIONING. DURING VALVE POSITIONING AND RECAPTURES, IT WAS OBSERVED THAT THE STIFF GUIDEWIRE WAS DEEPLY ADVANCED. BASED ON THE PROCEDURAL SEQUENCE AND FLUOROSCOPIC IMAGES, IT WAS SUSPECTED THAT THE GUIDEWIRE MAY HAVE CONTRIBUTED TO A RIGHT ATRIAL AND/OR LEFT VENTRICULAR PERFORATION, WHICH COULD HAVE RESULTED IN A PERICARDIAL EFFUSION. CARDIAC ARREST OCCURRED, AND A PERICARDIAL EFFUSION AND POSSIBLE RIGHT ATRIAL PERFORATION WERE SUSPECTED. THE PATIENT WAS TRANSFERRED TO THE OPERATING ROOM FOR CARDIAC SURGERY, AND A PERICARDIAL PATCH REPAIR WAS BEING CONSIDERED. DUE TO THE PATIENT¿S ADVANCED AGE AND FRAILTY, THE CARDIAC SURGERY TEAM AND THE FAMILY DECIDED NOT TO PROCEED WITH THE INTERVENTION. THE PATIENT PASSED AWAY. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE MEDICAL TEAM, THE CAUSE OF DEATH WAS A VENTRICULAR PERFORATION. THE PHYSICIAN EXCLUDED THE VALVE AND DCS AS CONTRIBUTING CAUSES TO THE DEATH. THE VALVE WAS RECAPTURED TWICE. A NON-MEDTRONIC (COOK) GUIDEWIRE WAS USED DURING THE PROCEDURE.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE VIA LEFT CAROTID ACCESS USING A SELF-EXPANDING VALVE, AN ELDERLY PATIENT WITH A HISTORY OF PERMANENT PACEMAKER IMPLANTATION, ATRIAL FIBRILLATION, AND A FRAIL CLINICAL CONDITION UNDERWENT RECAPTURE MANEUVERS TO OPTIMIZE VALVE POSITIONING. DURING VALVE POSITIONING AND RECAPTURES, IT WAS OBSERVED THAT THE STIFF GUIDEWIRE WAS DEEPLY ADVANCED. BASED ON THE PROCEDURAL SEQUENCE AND FLUOROSCOPIC IMAGES, IT WAS SUSPECTED THAT THE GUIDEWIRE MAY HAVE CONTRIBUTED TO A RIGHT ATRIAL AND/OR LEFT VENTRICULAR PERFORATION, WHICH COULD HAVE RESULTED IN A PERICARDIAL EFFUSION. CARDIAC ARREST OCCURRED, AND A PERICARDIAL EFFUSION AND POSSIBLE RIGHT ATRIAL PERFORATION WERE SUSPECTED. THE PATIENT WAS TRANSFERRED TO THE OPERATING ROOM FOR CARDIAC SURGERY, AND A PERICARDIAL PATCH REPAIR WAS BEING CONSIDERED. DUE TO THE PATIENT¿S ADVANCED AGE AND FRAILTY, THE CARDIAC SURGERY TEAM AND THE FAMILY DECIDED NOT TO PROCEED WITH THE INTERVENTION. THE PATIENT PASSED AWAY. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE MEDICAL TEAM, THE CAUSE OF DEATH WAS A VENTRICULAR PERFORATION. THE PHYSICIAN EXCLUDED THE VALVE AND DCS AS CONTRIBUTING CAUSES TO THE DEATH. THE VALVE WAS RECAPTURED TWICE. A NON-MEDTRONIC (COOK) GUIDEWIRE WAS USED DURING THE PROCEDURE. ADDITIONAL INFORMATION WAS RECEIVED THAT AFTER THE POSITIONING MANEUVERS; THE VALVE WAS SUCCESSFULLY IMPLANTED BUT THE PATIENT SUBSE QUEENLY DEVELOPED HEMODYNAMIC INSTABILITY. A DEFINITIVE MEDICAL DIAGNOSIS WAS NOT ESTABLISHED, BUT PRESUMED DUE TO TECHNICAL FACTORS RELATED TO THE PROCEDURE, PROCEDURAL SEQUENCE, AND THE COMPLEXITY OF THE CASE INCLUDING THE DEEPLY ADVANCED POSITION OF THE GUIDEWIRE. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: E VPROPLUS-26, SERIAL/LOT #: (B)(6), UBD: 21-AUG-2026, UDI#: (B)(4) SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW: THERE WERE FOUR IMAGES PROVIDED FOR REVIEW OF THE EVENT. THERE WAS NO FLUOROSCOPIC LOAD INSPECTION OR COMPUTED TOMOGRAPHY (CT) SCAN PROVIDED FOR REVIEW. THERE WAS AN EXECUTIVE SUMMARY PROVIDED FOR ANATOMICAL CONSIDERATIONS. IT WAS REPORTED THAT A 26MM VALVE WAS IMPLANTED VIA THE LEFT CAROTID ACCESS. PER INSTRUCTIONS FOR USE (IFU) CRITERIA, THIS WAS THE APPROPRIATELY SIZED VALVE FOR THE PATIENT¿S ANATOMY. PER THE EXECUTIVE SUMMARY, THE RIGHT FEMORAL ARTERY (RFA) AND LEFT FEMORAL ARTERY (LFA) VESSELS ARE NOT LARGE ENOUGH TO ACCOMMODATE THE DELIVERY CATHETER SYSTEM (DCS) FOR A 26MM VALVE. EVIDENCE IS CONSISTENT WITH A POSSIBLE BOVINE ARCH. IT WAS REPORTED THAT DURING VALVE POSITIONING AND RECAPTURES, IT WAS OBSERVED THAT THE STIFF GUIDEWIRE WAS DEEPLY ADVANCED. BASED ON THE PROCEDURAL SEQUENCE AND FLUOROSCOPIC IMAGES, IT WAS SUSPECTED THAT THE GUIDEWIRE MAY HAVE CONTRIBUTED TO A RIGHT ATRIAL AND/OR LEFT VENTRICULAR PERFORATION, WHICH COULD HAVE RESULTED IN A PERICARDIAL EFFUSION. WITH THE IMAGING, IT IS POSSIBLE THAT THERE IS A PERICARDIAL EFFUSION PRESENT WHEN THE VALVE WAS DEPLOYED TO THE POINT OF NO RECAPTURE AND AN IMAGE SHOWS THE VALVE AFTER FULL RELEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B2, B4, D4, H4, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22396057&pc=QFG | 3013756811-2025-114055 | 2025-06-03 04:00 | Death | TANDEM DIABETES CARE | 2025-07-03 04:00 | QFG | T:SLIM X2 INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Hyperglycemia | K201214 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON AN UNSPECIFIED DATE. PER THE CONTACT, THE CAUSE OF DEATH WAS UNKNOWN. REPORTEDLY, THE CUSTOMER EXPERIENCED A HIGH BLOOD GLUCOSE LEVEL OF OVER 1500 MG/DL AND WAS WEARING THE PUMP AT THE TIME OF DEATH. MULTIPLE ATTEMPTS WERE MADE BY TANDEM TECHNICAL SUPPORT TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, NO ADDITIONAL INFORMATION COULD BE OBTAINED. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22396118&pc=NPT | 2015691-2025-05439 | 2025-03-18 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-03 04:00 | NPT | SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Endocarditis; Insufficient Information | P140031 |  | 1 | Event Description: AS REPORTED BY EDWARDS IMPLANT PATIENT REGISTRY (IPR), APPROXIMATELY 5 YEARS 7 MONTHS AND 5 DAYS POST TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WITH A 29 MM SAPIEN 3 VALVE, THE VALVE WAS EXPLANTED, AND A NEW 25 MM SURGICAL VALVE WAS IMPLANTED. THE CAUSE OF THE VALVE EXPLANT IS UNKNOWN AT THIS TIME. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED TO CORRECT THE REPORTABILITY OF THE EVENT PREVIOUSLY REPORTED. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN CORRECTED AND UPDATED: B2 (OUTCOMES ATTRIBUTED TO ADVERSE EVENTS), B7 (OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS), H1 (TYPE OF REPORTABLE EVENT), H6 (HEALTH EFFECT - IMPACT CODE, HEALTH EFFECT - CLINICAL CODE, DEVICE CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) AND H11 (ADDITIONAL NARRATIVE/DATA). ADDITIONAL INFORMATION WAS RECEIVED VIA MEDICAL RECORDS REVEALING THE PATIENT HAD DEVELOPED ENDOCARDITIS APPROXIMATELY 5 YEARS, 6 MONTHS AND 17 DAYS POST TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WITH A 29 MM SAPIEN 3 VALVE. THE PATIENT WAS DIAGNOSED WITH VANCOMYCIN-RESISTANT ENTEROCOCCUS (VRE) BACTEREMIA. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) CONFIRMED THE PRESENCE OF MITRAL VALVE ENDOCARDITIS. THE PATIENT SUBSEQUENTLY UNDERWENT MITRAL VALVE REPLACEMENT (MVR) AND AORTIC VALVE REPLACEMENT (AVR) PROCEDURE WITH EXPLANT OF THE 29 MM SAPIEN 3 VALVE. APPROXIMATELY 16 DAYS POST PROCEDURE, THE PATIENT PASSED AWAY WITH ENDOCARDITIS OF THE MITRAL VALVE LISTED AS A CAUSE OF DEATH. AT THE TIME OF THIS REPORT, THE INFORMATION AVAILABLE DOES NOT SUGGEST THE 29 MM SAPIEN 3 VALVE MALFUNCTIONED, OR THAT THE USE OR MISS-USE OF THE DEVICE CAUSED OR CONTRIBUTED TO THE VALVE EXPLANT AND SUBSEQUENT PATIENT DEATH. THERE WAS NO ALLEGATION OF DEFICIENCIES RELATED TO THE IDENTITY, QUALITY, DURABILITY, RELIABILITY, SAFETY, LABELING EFFECTIVENESS, OR PERFORMANCE OF THIS EDWARDS SAPIEN 3 VALVE. THEREFORE, THE VALVE EXPLANT EVENT IS NO LONGER CONSIDERED FDA REPORTABLE. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22396180&pc=OZD | 1220648-2025-30052 | 2025-06-15 04:00 | Death | ABIOMED, INC. | 2025-07-03 04:00 | OZD | IMPELLA CP | Crack; Restricted Flow rate; Malposition of Device | Hematoma; Hemolysis; Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT. DURING THE IMPLANT, THE PEEL AWAY SHEATH CRACKED. THE PATIENT WAS BLEEDING. THE PEEL AWAY SHEATH WAS REMOVED AND THE REPOSITIONING SHEATH WAS IMPLANTED. THE BLEEDING STOPPED. ADDITIONALLY, THE PATIENT WAS FOUND TO HAVE HEMOLYSIS. THE PUMP POSITION WAS SHALLOW BUT THE DOCTOR DID NOT WANT TO REPOSITION SO SUPPORT WAS CONTINUED. THE PATIENT ALSO HAD LIMB ISCHEMIA. A FEMORAL-FEMORAL BYPASS WAS PLACED. FURTHERMORE, THE PUMP WAS GOING TO BE EXPLANTED AND REPLACED WITH AN IMPELLA 5.5 DEVICE. UPON ARRIVAL TO THE INTENSIVE CARE UNIT (ICU), THE PATIENT HAD A HEMATOMA TO THE RIGHT GROIN. SHORTLY AFTER ARRIVAL, THE PATIENT BEGAN TO DECOMPENSATE SIGNIFICANTLY REQUIRING RAPID TRANSFUSION AND MASSIVE TRANSFUSION PROTOCOL. AT THIS TIME, IT WAS NOTED PATIENT HAD A GROWING HEMATOMA TO THE LEFT GROIN. AN EMERGENT SURGICAL CUT DOWN WAS PERFORMED IN THE ICU TO CONTROL THE BLEEDING. HEMOSTASIS WAS ACHIEVED ON THE LEFT GROIN. THE DOCTOR SURGICALLY REPAIRED THE RIGHT GROIN AT THIS TIME AS WELL. THE PATIENT MAINTAINED FLOWS ON IMPELLA CP WITH MINIMAL SUCTION ALARMS DURING RESUSCITATION. THE DOCTOR STATED THIS ADVERSE EVENT WAS DUE TO HEMOSTASIS DEVICES USED TO CLOSE THE EXTERNAL FEMORAL-FEMORAL BYPASS AND CLOSURE OF THE IMPELLA CP INSERTION SITE. THE PATIENT EXPIRED. Manufacturer Narrative: ABIOMED'S MEDICAL SAFETY OFFICER REVIEWED THE EVENT AND DETERMINED THE PATIENT WAS ESCALATED TO AND EXPIRED ON THE IMPELLA 5.5 BUT WAS A SERIOUS INJURY FOR THE ATRIAL FIBRILLATION. THE IMPELLA CP WAS RELATED ADVERSE EVENTS (E.G., LIMB ISCHEMIA, ACCESS BLEEDING REQUIRING BLOOD TRANSFUSION, ETC.) ARE MORE LIKELY ASSOCIATED WITH THE DEMISE OUTCOME. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THIS IS ONE OF TWO MEDICAL DEVICE REPORTS ASSOCIATED WITH THE PATIENT DEATH. REFER TO INITIAL MEDICAL DEVICE REPORT FOR THE IMPELLA 5.5 SERIAL NUMBER (B)(6) WITH DATE OF EVENT 17-JUN-2025 AND IDENTIFIER ENDING IN (B)(6). THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES, INTRODUCER DAMAGE ¿ MECHANICAL, ACCESS SITE BLEEDING, VASCULAR DAMAGE - PERIPHERAL VESSEL, HEMOLYSIS, AND LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. POSITIONING ISSUES: CLINICAL DETAILS REPORT THAT THE PATIENT BEGAN TO EXPERIENCE HEMOLYSIS SYMPTOMS THE MORNING OF (B)(6) 2025. AT THIS TIME, POSITION WAS CHECKED WITH IMAGING, AND IT WAS NOTED THAT THE PUMP WAS SHALLOW. THE PHYSICIAN DECIDED NOT TO REPOSITION. OF NOTE, THE PATIENT WAS REPORTED TO HAVE A SMALL LEFT VENTRICLE (LV). NO CLINICAL DETAILS WERE PROVIDED TO SUGGEST HOW THE PUMP HAD BECOME SHALLOW. DATA LOGS WERE REVIEWED AND THERE WAS ONLY 1 POSITION IN AORTA ALARM AT THE END OF THE CASE. THIS WAS PRESUMABLY DURING PUMP EXPLANT. PRODUCT WAS NOT RETURNED FOR INVESTIGATION. SINCE INSUFFICIENT CLINICAL DETAILS WERE PROVIDED REGARDING HOW THE PUMP CAME OUT OF POSITION, THE ROOT CAUSE OF THE POSITIONING ISSUES COULD NOT BE DETERMINED. INTRODUCER DAMAGE - MECHANICAL: CLINICAL DETAILS REPORT THAT DURING THE INSERTION PROCEDURE, THE INTRODUCER SHEATH CRACKED. THERE WERE NO VESSEL CONDITIONS REPORTED NOR EXCESSIVE FORCE USED. OF NOTE, HOWEVER, IT WAS REPORTED THAT THE INSERTION ANGLE WAS 45 DEGREES. THE INTRODUCER SHEATH WAS PULLED OUT OF THE BODY AND PEELED AWAY, THEN THE REPOSITIONING SHEATH WAS ADVANCED TO RESOLVE THE BLEEDING. PRODUCT WAS NOT RETURNED FOR INVESTIGATION. BASED ON THE CLINICAL DETAIL PROVIDED THAT THE INSERTION ANGLE WAS STEEP, THE ROOT CAUSE OF THE INTRODUCER DAMAGE WAS DETERMINED TO MOST LIKELY BE A USE ISSUE. ACCESS SITE BLEED: CLINICAL DETAILS REPORT THAT DURING ACCESS, THE INTRODUCER SHEATH WAS FOUND TO BE CRACKED. THE ACCESS SITE WAS BLEEDING. THE PEEL AWAY SHEATH WAS REMOVED, THE REPOSITIONING UNIT WAS ADVANCED, AND THE BLEEDING SUBSIDED. IT WAS REPORTED THAT THE INTRODUCER WAS INSERTED AT A 45 DEGREE ANGLE. PRODUCT WAS NOT RETURNED FOR INVESTIGATION. CLINICAL DETAILS PROVIDED THAT THE INTRODUCER SHEATH WAS INSERTED AT A 45 DEGREE ANGLE AND CRACKED. THE BLEEDING RESOLVED AFTER ADVANCING THE REPOSITIONING UNIT. BASED ON THIS INFORMATION, THE ROOT CAUSE OF THE ACCESS SITE BLEED WAS DETERMINED TO MOST LIKELY BE A USE ISSUE. VASCULAR DAMAGE - PERIPHERAL VESSEL: CLINICAL DETAILS REPORT THAT SHORTLY AFTER THE IMPELLA CP WAS EXPLANTED AND THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5, THE CP ACCESS SITE AND THE LEFT FEMORAL ARTERY (LFA) (OTHER SITE RELATED TO FEM-FEM BYPASS) DEVELOPED HEMATOMAS. ADDITIONALLY, BOTH VESSELS REQUIRED SURGICAL REPAIR. THE PHYSICIAN BELIEVED THAT THE ADVERSE EVENT WAS DUE TO THE HEMOSTASIS DEVICES USED TO CLOSE THE EXTERNAL FEM-FEM BYPASS AND CP INSERTION SITES. HOWEVER, NO FURTHER DETAILS WERE PROVIDED, PARTICULARLY IF EXCESSIVE FORCE WAS USED OR IF THE PATIENT'S CONDITION MAY HAVE PLAYED A ROLE. SINCE INSUFFICIENT CLINICAL DETAILS WERE PROVIDED, THE ROOT CAUSE OF THE VASCULAR DAMAGE COULD NOT BE DETERMINED. HEMOLYSIS: DATA LOG REVIEW DID NOT CONFIRM ANY POSITIONING ISSUES. THERE WAS 1 POSITION IN AORTA ALARM AT THE END OF THE CASE, PRESUMABLY DURING PUMP EXPLANT. THERE WERE NO ABNORMAL MOTOR CURRENT (MC) SPIKES DURING THE CASE. THAT BEING SAID, THE PLACEMENT SIGNAL (PS) AND MC SHOWED VARIABILITY THROUGHOUT THE CASE. THERE WAS ONE INSTANCE EARLY IN THE CASE WHERE THERE WAS A SIGNIFICANT INCREASE IN THE PS/MAP, AND IT CORRELATED WITH A DROP IN MC MINIMUM VALUES. THAT BEING SAID, MC MAXIMUM VALUES REMAINED STABLE AND NO SUCTION ALARMS TRIGGERED. THIS WAS AN ACUTE EVENT THAT RESOLVED QUICKLY. LATER IN THE EVENING OF 15/JUN/2025, THERE WAS A TREND DOWN IN PS THAT ALIGNED WITH A TREND DOWN IN MC. THE MC RECOVERED, BUT PS CONTINUED TO BE VARIABLE. THEREFORE, IT IS UNCLEAR IF THE PATIENT'S CONDITION COULD HAVE CAUSED THIS DROP IN MEAN FLOWS. PRODUCT WAS NOT RETURNED FOR INVESTIGATION. SINCE CLINICAL DETAILS REPORTED SEVERAL POTENTIAL FACTORS THAT COULD PLAY A ROLE IN HEMOLYSIS, DATA LOG REVIEW WAS INCONCLUSIVE, AND PRODUCT WAS NOT RETURNED FOR INVESTIGATION, THE ROOT CAUSE OF THE HEMOLYSIS COULD NOT BE DETERMINED. LIMB ISCHEMIA: THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. G1 REPORTING CONTACT EMAIL REVISED ON MANUFACTURER DEVICE REPORT 1220648-2025-30052 IN ACCORDANCE WITH UPDATED PROCEDURES. H6 COMPONENT CODE 4742 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30052.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22396320&pc=FRN | 1416980-2025-03628 | 2025-05-31 04:00 | Death | BAXTER HEALTHCARE CORPORATION | 2025-07-03 04:00 | FRN | NOVUM IQ | Insufficient Flow or Under Infusion | Cardiac Arrest | K211122 |  | 1 | Event Description: IT WAS REPORTED THAT A NOVUM IQ LARGE VOLUME PUMP UNDER-INFUSED AND THE PATIENT SUBSEQUENTLY DIED. THE PATIENT REPORTEDLY WAS INITIALLY TRANSPORTED TO THE HOSPITAL VIA MED FLIGHT AFTER REPORTEDLY SUFFERING BURNS TO 95% OF THEIR BODY. THE CUSTOMER REPORTED THAT THE PATIENT WAS INTUBATED EN ROUTE. THE PATIENT WAS REPORTED TO HAVE ARRIVED AT THE HOSPITAL IN CRITICAL CONDITION AND WAS ¿RUSHED¿ TO THE OPERATING ROOM FOR AN ESCHAROTOMY DUE TO SWELLING. THE PATIENT REPORTEDLY WASN¿T RECEIVING ENOUGH OXYGEN FROM THE VENTILATOR. AT AN UNSPECIFIED TIME IN THE HOSPITAL THE NOVUM PUMP REPORTEDLY WAS PROGRAMMED TO RUN 1000ML OF NORMAL SALINE AT 999ML/HOUR. WHEN THE PUMP REPORTEDLY ALARMED, THE INFUSION COMPLETED AND THE VOLUME TO BE INFUSED WAS REPORTED TO HAVE READ ¿ZERO¿ ON THE SCREEN. THE MEDICAL STAFF REPORTEDLY OBSERVED THAT THE NORMAL SALINE BAG STILL HAD APPROXIMATELY 250-300ML REMAINING IN THE BAG. THAT NIGHT THE FAMILY WAS INFORMED OF THE PATIENT¿S CONDITION, AND IT WAS REPORTED THAT ¿LIMITED CARE¿ WAS REQUESTED BE GIVEN TO THE PATIENT. THE PATIENT REPORTEDLY EXPIRED THE FOLLOWING DAY AT 00:55. AN AUTOPSY WAS NOT PERFORMED. THE REPORTER STATED: "NO OFFICIAL DEATH RECORD WAS AVAILABLE; HOWEVER, ¿THE LIKELY CAUSE OF DEATH WOULD HAVE BEEN CARDIAC ARREST ULTIMATELY, SECONDARY TO 95% BURNS.¿ THE HOSPITAL BIOMEDICAL DEPARTMENT REPORTEDLY TESTED THE DEVICE THE SAME DAY, AND THE TECHNICIAN COULD NOT DUPLICATE THE REPORTED ISSUE; THEREFORE, THE PUMP WAS PUT BACK INTO SERVICE. NO FURTHER INFORMATION WAS AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: H11: THE EVENT HISTORY LOG SHOWED SEVERAL VOLUMES TO BE INFUSED (VTBI), PROGRAMED AT 990 ML, 800 ML, 200 ML, 950 ML, 100 ML, 1000 ML, 100 ML, AND 1000 ML. ON THE FIRST BAG, THERE WERE SIX DOWNSTREAM OCCLUSIONS (DSO). DSO ALARMS MAY INDICATE THERE WERE PARTIAL/BORDERLINE OCCLUSION CONDITIONS, WHICH COULD CAUSE UNDER-INFUSION. THE DEVICE WAS TESTED AT THE USER FACILITY AND NO ISSUES WERE OBSERVED THEREFORE THE PUMP WAS PUT BACK INTO SERVICE. THE DEVICE WAS NOT RETURNED; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22396736&pc=DSP | 2248146-2025-0000345 | 2025-04-29 04:00 | Death | DATASCOPE CORP. | 2025-07-03 04:00 | DSP | LINEAR 7.5FR. 34CC IAB (WITH APA) | Burst Container or Vessel; Optical Problem | Insufficient Information | K041281 |  | 1 | Event Description: IT WAS REPORTED FIRST THAT DURING INTRA-AORTIC BALLOON (IAB) THERAPY, A FAILURE OCCURRED IN THE FIBER OPTIC SENSOR MODULE, RESULTING IN THE BALLOON RUPTURING INSIDE THE PATIENT. ALTHOUGH NO COMPLICATIONS DIRECTLY RELATED TO THE BALLOON RUPTURE WERE REPORTED, THE PATIENT SUBSEQUENTLY DIED. Manufacturer Narrative: UPDATED FIELDS: B4, D9, G1 (CONTACT PERSON MFG SITE), G3, G6, H2, H3, H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, COMPONENT CODES AND INVESTIGATION CONCLUSIONS), H11. NO DECON OR VISUAL INSPECTION PERFORMED SINCE PRODUCT WAS NOT RETURNED AND COULD NOT BE EVALUATED. THEREFORE, WE WERE UNABLE TO CONFIRM OR DETERMINE A ROOT CAUSE FOR THE REPORTED FAILURE. Manufacturer Narrative: DUE TO CHARACTER LIMITATION IN E1, FULL INITIAL REPORTER NAME: (B)(6). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22398068&pc=NVY | 3002807561-2025-00014 | 2025-05-31 04:00 | Death | MEDTRONIC B.V. | 2025-07-03 04:00 | NVY | OMNIASECURE¿ MRI SURESCAN¿ LEAD | Over-Sensing; Pacing Problem; Under-Sensing | Ventricular Fibrillation; Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT A REVIEW OF A REMOTE TRANSMISSION SHOWED MULTIPLE SHORT V-V INTERVALS ON THE RIGHT VENTRICULAR (RV) LEAD. ADDITIONALLY, THE PRESENTING RHYTHM SHOWED A POSSIBLE PAUSE AND UNDERSENSING NOTED POTENTIALLY RESULTING IN VENTRICULAR FIBRILLATION (VF) IT WAS REPORTED THAT A REVIEW OF A REMOTE TRANSMISSION SHOWED MULTIPLE SHORT V-V INTERVALS ON THE RIGHT VENTRICULAR (RV) LEAD. ADDITIONALLY, THE PRESENTING RHYTHM SHOWED A POSSIBLE PAUSE AND UNDERSENSING NOTED POTENTIALLY RESULTING IN VENTRICULAR FIBRILLATION (VF) NOT BEING DETECTED OR TREATED. UPON FURTHER REVIEW OF THE EPISODES IT WAS NOTED THAT THE RHYTHM APPEARED TO BE AN AGONAL RHYTHM, PULSELESS ELECTRICAL ACTIVITY. THE PATIENT PASSED AWAY. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY.Event Description: IT WAS FURTHER REPORTED THAT A LEAD INTEGRITY ALERT (LIA) WAS TRIGGERED AND PHYSIOLOGIC OVERSENSING WAS OBSERVED. ADDITIONALLY, IT WAS SUSPECTED THAT THE PATIENT'S DEATH WAS RELATED TO VENTRICULAR FIBRILLATION (VF) AND SUDDEN CARDIAC DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22398859&pc=QZK | 2017233-2025-06380 | 2025-06-03 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-07-03 04:00 | QZK | GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Internal Organ Perforation; Rupture | P230023 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED FROM THE AAA24-01 CLINICAL STUDY MANAGER: ON (B)(6) 2025, THE PATIENT UNDERWENT ENDOVASCULAR TREATMENT OF A THORACOABDOMINAL AORTIC ANEURYSM USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) DEVICE SYSTEM. THE MAXIMUM DIAMETER OF THE AORTIC DISEASE WAS 67MM. IT WAS REPORTED THAT, DURING THE PROCEDURE, THE PATIENT EXPERIENCED A SPLENIC ARTERY RUPTURE. THE PHYSICIAN ATTEMPTED TO REPAIR THE RUPTURE, BUT ULTIMATELY A SPLENECTOMY WAS PERFORMED. THE PATIENT TOLERATED THE PROCEDURE. REPORTEDLY, THE PATIENT EXPERIENCED COMPLICATIONS POST PROCEDURE, INCLUDING LIVER SHOCK, AND CONTINUED TO DECLINE. THE PATIENT WAS PLACED ON COMFORT CARE. ON (B)(6) 2025, THE PATIENT EXPIRED. DURING THE AUTOPSY ON (B)(6) 2025, THE SITE EXPLANTED THE TAMBE DEVICE SYSTEM. THE EXPLANTED DEVICE SYSTEM WAS RETURNED. IT WAS NOTED THAT THE DEVICE EXPLANT WAS NOT REPORTEDLY RELATED TO AN ALLEGATION AGAINST THE DEVICE SYSTEM, BUT THAT THE EXPLANT WAS PERFORMED DURING THE ALREADY PLANNED AUTOPSY PROCEDURE. Manufacturer Narrative: ADDITIONAL DEVICES INCLUDED ON THIS REPORT ARE AS FOLLOWS: CATALOG #BXB085902A/ SERIAL #(B)(6)/ UDI #(B)(4). CATALOG #BXB095902A/ SERIAL #(B)(6)/ UDI #((B)(4). H.6. INVESTIGATION FINDINGS FOR ANALYSIS OF PRODUCTION RECORDS: CODE C19 - THE REVIEW OF THE MANUFACTURING PAPERWORK VERIFIED THAT THE LOTS INVOLVED IN THIS EVENT MET ALL PRE-RELEASE SPECIFICATIONS. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22399009&pc=PRL | 2017233-2025-06381 | 2025-06-03 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-07-03 04:00 | PRL | GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Internal Organ Perforation; Rupture | P160021 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED FROM THE (B)(4) CLINICAL STUDY MANAGER: ON (B)(6) 2025, THE PATIENT UNDERWENT ENDOVASCULAR TREATMENT OF A THORACOABDOMINAL AORTIC ANEURYSM USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) DEVICE SYSTEM. THE MAXIMUM DIAMETER OF THE AORTIC DISEASE WAS 67MM. TWO GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS DEVICES WERE IMPLANTED IN THE PATIENT'S CELIAC ARTERY. IT WAS REPORTED THAT, DURING THE PROCEDURE, THE PATIENT EXPERIENCED A SPLENIC ARTERY RUPTURE. THE PHYSICIAN ATTEMPTED TO REPAIR THE RUPTURE, BUT ULTIMATELY A SPLENECTOMY WAS PERFORMED. THE PATIENT TOLERATED THE PROCEDURE. REPORTEDLY, THE PATIENT EXPERIENCED COMPLICATIONS POST PROCEDURE, INCLUDING LIVER SHOCK, AND CONTINUED TO DECLINE. THE PATIENT WAS PLACED ON COMFORT CARE. ON (B)(6) 2025, THE PATIENT EXPIRED. DURING THE AUTOPSY ON (B)(6) 2025, THE SITE EXPLANTED THE TAMBE DEVICE SYSTEM. THE EXPLANTED DEVICE SYSTEM WAS RETURNED. IT WAS NOTED THAT THE DEVICE EXPLANT WAS NOT REPORTEDLY RELATED TO AN ALLEGATION AGAINST THE DEVICE SYSTEM, BUT THAT THE EXPLANT WAS PERFORMED DURING THE ALREADY PLANNED AUTOPSY PROCEDURE. Manufacturer Narrative: ADDITIONAL DEVICES INCLUDED ON THIS REPORT ARE AS FOLLOWS: CATALOG # BXB095902A/ SERIAL #( B)(6)/ UDI #(B)(4) AS A COMPONENT OF THE SAME SYSTEM. CATALOG # ATAA43720160A/ SERIAL # (B)(6)/ UDI # (B)(4). W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22399208&pc=BZD | 2518422-2025-107086 | 2025-07-01 04:00 | Death | RESPIRONICS, INC. | 2025-07-03 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED; THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22401087&pc=MNS | 2518422-2025-107090 | 2025-07-01 04:00 | Death | RESPIRONICS, INC. | 2025-07-03 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DEATH OF A PATIENT. THE MANUFACTURER RECEIVED A CALL FROM THE PATIENT'S CLINIC STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22401284&pc=NPU | 2015691-2025-05466 | 2025-06-16 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-03 04:00 | NPU | EDWARDS COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Perforation | P140031 |  | 1 | Event Description: AS REPORTED BY OUR AFFILIATES IN BRAZIL, IT WAS A CASE OF AN IMPLANT OF A 29MM SAPIEN 3 VALVE INSIDE A NON-EDWARDS SURGICAL VALVE, IN MITRAL POSITION BY TRANSFEMORAL APPROACH. A LUNDERQUIST DOUBLE CURVE (DC) GUIDEWIRE WAS USED DURING THE PROCEDURE. FOLLOWING THE VALVE DEPLOYMENT, LOW BLOOD PRESSURE WAS OBSERVED, AND INVESTIGATION WAS STARTED TO DISCOVER WHAT HAPPENED. AFTER A FEW MINUTES, A VENTRICULAR PERFORATION WAS IDENTIFIED, LIKELY OCCURRING DURING THE VALVE IMPLANT. CONSEQUENTLY, IT WAS DECIDED TO CONVERT TO OPEN-HEART SURGERY AND TRANSFER THE PATIENT TO THE SURGICAL CENTER. THE PATIENT PASSED AWAY DURING THE SURGERY. AS PER MEDICAL OPINION, THE PERCEIVED ROOT CAUSE OF THIS EVENT WAS VENTRICULAR PERFORATION CAUSED BY THE GUIDEWIRE. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR VENTRICULAR PERFORATION DURING A TVR PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THE VENTRICULAR PERFORATION WAS CAUSED BY THE GUIDEWIRE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22380378&pc=OZO | 2032227-2025-202442 | 2025-05-12 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-02 04:00 | OZO | 630G INSULIN PUMP MMT-1715KL 630G | Adverse Event Without Identified Device or Use Problem | Hypoglycemia; Insufficient Information | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2025, AND THE CAUSE OF DEATH WAS ACUTE MYOCARDIAL INFARCTION (HEART ATTACK). THE CUSTOMER REPORTED HYPOGLYCEMIA WAS TREATED WITH HOSPITALIZATION. THE EVENT INVOLVED PRODUCTS MMT-864A, MMT-332A, AND MMT-1715KL. TROUBLESHOOTING WAS PERFORMED AND FOUND THAT THE DOCUMENT OF ANY HEALTH ISSUES OR ILLNESSES THAT MAY HAVE CONTRIBUTED OR LED UP TO PASSING WAS THE DIABETES AND HEART PROBLEMS. THE CUSTOMER STATED THAT THE PUMP WAS NOT WORN AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-864A AND MMT-332A. MMT-1715KL WAS REQUESTED, AND THE CUSTOMER'S RESPONSE WAS THAT THE DEVICE WILL BE RETURNED. THE CUSTOMER WILL DISCONTINUE THE USE OF THE DEVICE. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22380800&pc=BZD | 2518422-2025-106981 | 2025-06-16 04:00 | Death | RESPIRONICS, INC. | 2025-07-02 04:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF A PATIENT'S DEATH FROM A DAUGHTER-IN-LAW STATING THAT THE PATIENT WENT INTO A NURSING HOME AND USED THE DEVICE EVERY NIGHT AND THAT SOMEBODY SPILLED WATER INTO THE DEVICE. SHE SAID THE PATIENT PASSED AWAY A FEW DAYS AGO NOT DUE TO THE DEVICE; NOTICED THE STAFF TO HAVE SWITCHED OUT THE DEVICE; ON IT SAYS DONATED AND TOOK OUT THE SD CARD. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22381734&pc=OZD | 1220648-2025-29988 | 2025-06-13 04:00 | Death | ABIOMED, INC. | 2025-07-02 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Insufficient Information | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 TO SUPPORT A PATIENT ADMITTED IN FOR COROANRY ARTERY BYPASS GRAFT AND FOUND TO HAVE POSTCARDIOTOMY CARDIOGENIC SHOCK. THE PUMP WAS SUPPORTING DESPITE KNOWN HEPARIN-INDUCED THROMBOCYTOPENIA (HIT) POSITIVE STATUS WHEN AFTER THREE DAYS THE HIGHPURGE PRESSURE WAS TREATED WITH THE ADDITION OF LYSIS, DESPITE THE HIT STATUS. LATER, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA PUMP WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST. SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER: ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22382442&pc=OZD | 1220648-2025-30000 | 2025-06-18 04:00 | Death | ABIOMED, INC. | 2025-07-02 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND AFTER THE IMPLANT, A LEFT VENTRICLE PERFORATION WAS NOTED. THE PUMP WAS EXPLANTED. THE PATIENT WAS STABLE AFTER THE EXPLANT. HOWEVER, THE PATIENT ENDED UP EXPIRING. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. IMPELLA CP WITH SMARTASSIST SYSTEM FOR USE DURING CARDIOGENIC SHOCK SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22382630&pc=OZD | 1220648-2025-30002 | 2025-06-18 04:00 | Death | ABIOMED, INC. | 2025-07-02 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Tachycardia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT ON IMPELLA CP SUPPORT HAD A COMPUTED TOMOGRAPHY (CT) OF THE BRAIN THAT SHOWED MULTIPLE CEREBROVASCULAR ACCIDENTS. CT WAS DONE AGAIN SHOWING EVOLVING STROKES BUT NOT WORSENING. NEUROLOGIST CONCLUDED THAT THE STROKES WERE NOT THE CULPRIT FOR THE PATIENT ABSENCE OF NEUROLOGY STATUS. ADDITIONALLY, THE PATIENT EXPERIENCED RUN OF VENTRICULAR TACHYCARDIA REQUIRING ONE ROUND OF CARDIOVERSION. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION ARREST AND THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22382739&pc=NPV | 2025587-2025-04720 | 2025-06-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-02 04:00 | NPV | HARMONY TRANSCATHETER PULMONARY VALVE | Insufficient Information; Patient Device Interaction Problem | Arrhythmia; Hemorrhage/Bleeding; Pulmonary Edema; Ventricular Fibrillation; Thrombosis/Thrombus | P200046 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER THE IMPLANTATION OF THE HARMONY TRANSCATHETER PULMONARY VALVE, CARDIAC DYSFUNCTION AND A CHANGE OF THE ANTERIOR WALL PORTION OF THE HEART OCCURRED. CATHETERIZATION REVEALED OCCLUSION OF THE LEFT CORONARY ARTERY, AND A STENT IMPLANTATION WAS PERFORMED AS A RESULT OF THE OCCLUSION. THE PATIENT WAS HOSPITALIZED IN THE INTENSIVE CARE UNIT AND SUBSEQUENTLY DIED.Event Description: IT WAS REPORTED THAT AFTER THE IMPLANTATION OF THE HARMONY TRANSCATHETER PULMONARY VALVE, CARDIAC DYSFUNCTION AND A CHANGE OF THE ANTERIOR WALL PORTION OF THE HEART OCCURRED. CATHETERIZATION REVEALED OCCLUSION OF THE LEFT CORONARY ARTERY, AND A STENT IMPLANTATION WAS PERFORMED AS A RESULT OF THE OCCLUSION. THE PATIENT WAS HOSPITALIZED IN THE INTENSIVE CARE UNIT AND SUBSEQUENTLY DIED. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE HARMONY VALVE IMPLANT PROCEDURE WAS FREE OF COMPLICATIONS. HOWEVER, AFTER THE PATIENT WAS EXTUBATED, PULMONARY EDEMA WAS NOTED; DIURETIC AND SUBSEQUENT INOTROPIC MEDICATION WAS ADMINISTERED. A COMPUTED TOMOGRAPHY WAS PERFORMED OF THE CHEST AND SHOWED A THICKENING OF PROBABLE BLOOD ORIGIN AT THE LUNG LEVEL, "NOT REPLENISHED". THERE WAS NO ACTIVE BLEED AT THE LEVEL OF THE ACCESSES. THE PATIENT PROGRESSIVELY WORSENED AND WAS ADMITTED TO THE POST-OPERATIVE INTENSIVE CARE UNIT. A BLEED AT THE RIGHT FEMORAL ACCESS WAS ADDRESSED AND PROMPTLY ARRESTED WITH MANUAL COMPRESSION, COMPRESSION BANDAGE, AND THE PATIENT WAS ADMINISTERED A CONCENTRATED BLOOD TRANSFUSION. THE PATIENT WAS EVALUATED BY A VASCULAR SURGEON WITH ULTRASOUND AND BRONCHOSCOPY, NO FURTHER BLEEDING WAS NOTED. IN THE MORNING HOURS, POST-OPERATIVE DAY ONE, THE PATIENT EXPERIENCED ARRHYTHMIAS INCLUDING VENTRICULAR FIBRILLATION; DIRECT CURRENT ELECTRICAL SHOCK WAS DELIVERED. A CORONAROGRAPHY WAS PERFORMED AND SHOWED COMPRESSION OF THE COMMON TRUNCUS OSTIA WHICH REQUIRED A STENT IMPLANT. THE PATIENT EXPERIENCED A SIGNIFICANT AGGRAVATION OF CONDITION WITH EXITUS IN THE MORNING HOURS ON POST-OPERATIVE DAY TWO.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH HAD MULTIFACTORIAL ETIOLOGY INCLUDING PULMONARY EDEMA FROM PULMONARY OVERFLOW, ANEMIA FROM BRONCHIAL BRANCH PERFORATION, LOW FLOW RATE, COMMON TRUNK STENOSIS WITH MYOCARDIAL ISCHEMIA AND METABOLIC ACIDOSIS. THE VALVE COULD NOT BE EXCLUDED AS A CONTRIBUTING CAUSE OF DEATH AS IT CAUSED THE COMPRESSION OF THE CORONARY COMMON TRUNK. THE DELIVERY CATHETER SYSTEM (DCS) COULD NOT BE EXCLUDED AS A CONTRIBUTING CAUSE OF DEATH AS IT FLOWED ON THE GUIDEWIRE THAT PERFORATED THE PULMONARY BRANCH. THE DCS DID NOT DISLODGE CALCIUM OR PLAQUE THAT OCCLUDED THE CORONARY ARTERY. THE PATIENT'S UNDERLYING MEDICAL HISTORY CONTRIBUTED TO DEATH DUE TO THE PRESENCE OF THE KNOWN ECTASIA OF THE ASCENDING AORTA, MAJOR DILATATION OF THE RIGHT VENTRICLE, AND LOW EJECTION FRACTION OF THE RIGHT VENTRICLE. NO THROMBOSIS IN THE CORONARY PRIOR TO THE PATIENT DEATH, ONLY NARROWING OF THE OSTIALIS COMMON TRUNCUS.Event Description: IT WAS REPORTED THAT AFTER THE IMPLANTATION OF THE HARMONY TRANSCATHETER PULMONARY VALVE, CARDIAC DYSFUNCTION AND A CHANGE OF THE ANTERIOR WALL PORTION OF THE HEART OCCURRED. CATHETERIZATION REVEALED OCCLUSION OF THE LEFT CORONARY ARTERY, AND A STENT IMPLANTATION WAS PERFORMED AS A RESULT OF THE OCCLUSION. THE PATIENT WAS HOSPITALIZED IN THE INTENSIVE CARE UNIT AND SUBSEQUENTLY DIED. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE HARMONY VALVE IMPLANT PROCEDURE WAS FREE OF COMPLICATIONS. HOWEVER, AFTER THE PATIENT WAS EXTUBATED, PULMONARY EDEMA WAS NOTED; DIURETIC AND SUBSEQUENT INOTROPIC MEDICATION WAS ADMINISTERED. A COMPUTED TOMOGRAPHY WAS PERFORMED OF THE CHEST AND SHOWED A THICKENING OF PROBABLE BLOOD ORIGIN AT THE LUNG LEVEL, "NOT REPLENISHED". THERE WAS NO ACTIVE BLEED AT THE LEVEL OF THE ACCESSES. THE PATIENT PROGRESSIVELY WORSENED AND WAS ADMITTED TO THE POST-OPERATIVE INTENSIVE CARE UNIT. A BLEED AT THE RIGHT FEMORAL ACCESS WAS ADDRESSED AND PROMPTLY ARRESTED WITH MANUAL COMPRESSION, COMPRESSION BANDAGE, AND THE PATIENT WAS ADMINISTERED A CONCENTRATED BLOOD TRANSFUSION. THE PATIENT WAS EVALUATED BY A VASCULAR SURGEON WITH ULTRASOUND AND BRONCHOSCOPY, NOFURTHER BLEEDING WAS NOTED. IN THE MORNING HOURS, POST-OPERATIVE DAY ONE, THE PATIENT EXPERIENCED ARRHYTHMIAS INCLUDING VENTRICULAR FIBRILLATION; DIRECT CURRENT ELECTRICAL SHOCK WAS DELIVERED. A CORONAROGRAPHY WAS PERFORMED AND SHOWED COMPRESSION OF THE COMMON TRUNCUS OSTIA WHICH REQUIRED A STENT IMPLANT. THE PATIENT EXPERIENCED A SIGNIFICANT AGGRAVATION OF CONDITION WITH EXITUS IN THE MORNING HOURS ON POST-OPERATIVE DAY TWO. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH HAD MULTIFACTORIAL ETIOLOGY INCLUDING PULMONARY EDEMA FROM PULMONARY OVERFLOW, ANEMIA FROM BRONCHIAL BRANCH PERFORATION, LOW FLOW RATE, COMMON TRUNK STENOSIS WITH MYOCARDIAL ISCHEMIA AND METABOLIC ACIDOSIS. THE VALVE COULD NOT BE EXCLUDED AS A CONTRIBUTING CAUSE OF DEATH AS IT CAUSED THE COMPRESSION OF THE CORONARY COMMON TRUNK. THE DELIVERY CATHETER SYSTEM (DCS) COULD NOT BE EXCLUDED AS A CONTRIBUTING CAUSE OF DEATH AS IT FLOWED ON THE GUIDEWIRE THAT PERFORATED THE PULMONARY BRANCH. THE DCS DID NOT DISLODGE CALCIUM OR PLAQUE THAT OCCLUDED THE CORONARY ARTERY. THE PATIENT'S UNDERLYING MEDICAL HISTORY CONTRIBUTED TO DEATH DUE TO THE PRESENCE OF THE KNOWN ECTASIA OF THE ASCENDING AORTA, MAJOR DILATATION OF THE RIGHT VENTRICLE, AND LOW EJECTION FRACTION OF THE RIGHT VENTRICLE. NO THROMBOSIS IN THE CORONARY PRIOR TO THE PATIENT DEATH, ONLY NARROWING OF THE OSTIALIS COMMON TRUNCUS. Manufacturer Narrative: ADDITIONAL INFORMATION: THE SYSTEM REPORTED DEVICE SPECIFICS WERE RECEIVED: PRODUCT ID: HARMONY-DCS; PRODUCT LOT NUMBER: 0012699637; PRODUCT TYPE: 0195-HEART VALVES; MANUFACTURED DATE: 2025-03-04; USE BEFORE DATE: 2026-03-04; FULL UDI #: (B)(4). NON-IMPLANTABLE DEVICE CORRECTED DATA: THE ABOVE SYSTEM REPORTED DEVICE WAS INCORRECTLY NOTED AS A CONCOMITANT PRODUCT IN THE INITIAL MEDWATCH REPORT. LET THIS REPORT ACCURATELY REFLECT THE DEVICE AS A SYSTEM REPORTED DEVICE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. H6. ADDITIONAL ANNEX E CODES WERE ADDED. ADDITIONAL CODES. ANNEX F CODES WERE ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5 SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID HARMONY-DCS; SERIAL/LOT NUMBER: N/A; USE BY DATE: N/A; UDI NUMBER: N/A H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID HARMONY-DCS; PRODUCT LOT NUMBER (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22382852&pc=NIK | 2182208-2025-03191 | 2024-04-17 04:00 | Death | MEDTRONIC, INC. | 2025-07-02 04:00 | NIK | MEDTRONIC CRT | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE CLINICAL IMPACT OF AN ADAPTIVE CARDIAC RESYNCHRONIZATION THERAPY (ACRT) ALGORITHM. PATIENTS WERE IMPLANTED WITH A CARDIAC RESYNCHRONIZATION THERAPY (CRT) DEVICE EITHER WITH OR WITHOUT THE ALGORITHM. THE AUTHORS DESCRIBED PATIENT DEATHS IN BOTH GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE. THERE WERE PATIENTS WHO EXPERIENCED HEART FAILURE (HF) HOSPITALIZATIONS WHICH INCLUDED UNEXPLAINED SYNCOPE. THE PATIENTS WITH THE USE OF THE ACRT ALGORITHM HAD LESS ALL-CAUSE DEATH AND/OR HOSPITALIZATION RATES FOR HF THAN THE GROUP WITHOUT. THERE IS NO ALLEGATION OF DEVICE-DEATH OR HF HOSPITALIZATION RELATEDNESS INDICATED IN THE ARTICLE. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: EFFECT OF ADAPTIVE CARDIAC RESYNCHRONIZATION THERAPY IN HEART FAILURE PATIENTS WITH PACEMAKER DEPENDENCY. JOURNAL OF INTERVENTIONAL CARDIAC ELECTROPHYSIOLOGY. 2024. 67:1571¿1577. DOI: 10.1007/S10840-024-01776-0BY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22383346&pc=NPT | 2025587-2025-04725 | 2025-04-04 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-02 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information; Patient Device Interaction Problem; Device Stenosis | Myocardial Infarction; Cardiogenic Shock; Low Cardiac Output; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT AN UNSPECIFIED AMOUNT OF TIME AFTER THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, AORTIC STENOSIS WAS NOTED. APPROXIMATELY 2 YEARS AND 5 MONTHS AFTER THE PROCEDURE, A MYOCARDIAL INFARCTION WAS NOTED AND A STENT WAS PLACED. APPROXIMATELY 2 WEEKS LATER, THE PATIENT WAS HOSPITALIZED DUE TO AN ACUTE SYSTOLIC HEART FAILURE EXACERBATION. SUBSEQUENTLY, A TEMPORARY HEART PUMP WAS PLACED. A RIGHT HEART CATHETERIZATION WAS PERFORMED WHICH REVEALED REDUCED CARDIAC OUTPUT. DURING THE HOSPITALIZATION, THE PATIENT ENTERED CARDIOGENIC SHOCK. THE PATIENT WAS ADMITTED TO THE CARDIOVASCULAR INTENSIVE CARE UNIT AND STARTED ON VASOPRESSOR MEDICATION. PER THE PHYSICIAN, THE PATIENT REQUIRED AT LEAST THREE DIFFERENT VASOPRESSORS FOR A PROLONGED DURATION. THE HEART PUMP AND VASOPRESSOR THERAPY WAS CONTINUED FOR SEVERAL DAYS WITH NO SIGNIFICANT IMPROVEMENT. AS A RESULT, THE PATIENT'S FAMILY DECIDED TO PURSUE PALLIATIVE CARE FOLLOWED BY COMFORT CARES. APPROXIMATELY 19 DAYS AFTER PLACEMENT OF THE HEART PUMP, THE PATIENT PASSED AWAY.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT CORRECTED PREVIOUSLY REPORTED INFORMATION INDICATING THAT THE AORTIC STENOSIS AND MYOCARDIAL INFARCTION FIRST OCCURRED APPROXIMATELY 2 YEARS AND 7 MONTHS FOLLOWING THE IMPLANT OF THE VALVE WHEN THE PATIENT WAS INITIALLY ADMITTED INTO THE HOSPITAL. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS LIKELY DUE TO THE MYOCARDIAL INFARCTION, AND THE VALVE COULD NOT BE RULED OUT AS A CONTRIBUTING FACTOR TO THE DEATH. Manufacturer Narrative: B3: DATE IS APPROXIMATE. YEAR IS CONFIRMED VALID. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. CORRECTED DATA: B3. D6A. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22383727&pc=NKM | 2135147-2025-03595 | 2015-08-01 04:00 | Death | ABBOTT MEDICAL | 2025-07-02 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Off-Label Use | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "VENTRICULAR INTERACTION AND CLINICAL OUTCOMES IN PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE ASSOCIATION BETWEEN VENTRICULAR INTERACTION AND CLINICAL OUTCOMES IN PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR (TTVR). DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP, PASCAL, CARDIOBAND. THE ARTICLE CONCLUDED THAT VENTRICULAR INTERACTION, ASSESSED BY LV ECCENTRICITY INDEX, WAS ASSOCIATED WITH CLINICAL OUTCOMES AFTER TTVR, AND ITS POSTPROCEDURAL IMPROVEMENT WAS LINKED TO BETTER OUTCOMES. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS DR (B)(6) AT (B)(6) HOSPITAL, WITH CORRESPONDING EMAIL TA.CHI.TSU.TETSU@GMAIL.COM. THE TIME FRAME OF THE STUDY WAS AUGUST 2015 TO AUGUST 2022. A TOTAL OF 304 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 208 (76%) RECEIVED AN ABBOTT DEVICE. AVERAGE AGE WAS 79 AND MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED FUNCTIONAL AND DEGENERATIVE TRICUSPID REGURGITATION, DIABETES, HYPERTENSION, COPD, CORONARY ARTERY DISEASE, PRIOR MYOCARDIAL INFARCTION, PRIOR CARDIAC SURGERY, ATRIAL FIBRILLATION, HEART FAILURE, CIED, HEART FAILURE MEDICATION, MITRAL REGURGITATION. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED TRICLIP: DEATH, HEART FAILURE, HOSPITALIZATION, UNCHANGED TR, RECURRENT TR MITRACLIP: DEATH, HEART FAILURE, HOSPITALIZATION, UNCHANGED TR, RECURRENT TR, OFF-LABEL USE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING FUNCTIONAL AND DEGENERATIVE TRICUSPID REGURGITATION, DIABETES, HYPERTENSION, COPD, CORONARY ARTERY DISEASE, PRIOR MYOCARDIAL INFARCTION, PRIOR CARDIAC SURGERY, ATRIAL FIBRILLATION, HEART FAILURE, CIED, HEART FAILURE MEDICATION, MITRAL REGURGITATION. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE, HOSPITALIZATION, UNCHANGED TR, RECURRENT TR, OFF-LABEL USE; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. THE REPORTED OFF-LABEL USE WAS ASSOCIATED WITH USING A MITRACLIP ON THE TRICUSPID VALVE. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: VENTRICULAR INTERACTION AND CLINICAL OUTCOMES IN PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22383775&pc=DSQ | 2916596-2025-04167 | 2025-05-01 04:00 | Death | THORATEC CORPORATION | 2025-07-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Septic Shock; Multiple Organ Failure; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED FOR ONGOING DRIVELINE INFECTION MANAGEMENT, WHICH PROGRESSED TO SEPTIC SHOCK AND MULTISYSTEM ORGAN FAILURE, WHICH LEAD TO DEATH. THE DEATH WAS CONSIDERED DEVICE RELATED DUE TO THE DRIVELINE EXIT SITE. THE DEVICE OPERATED AS EXPECTED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED DRIVELINE INFECTION; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, PUMP POCKET), SEPSIS, MULTI ORGAN FAILURE AND DYSFUNCTION (INCLUDING RIGHT HEART FAILURE, RESPIRATORY FAILURE, HEPATIC DYSFUNCTION, AND RENAL DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22383783&pc=DSQ | 2916596-2025-04109 | 2025-06-13 04:00 | Death | THORATEC CORPORATION | 2025-07-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Tachycardia; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A HISTORY OF VENTRICULAR TACHYCARDIA (VT) PRIOR TO THE LEFT VENTRICULAR ASSIST DEVICE (LVAD). ON (B)(6) 2025, THE PATIENT SUFFERED VT ABLATION AND ACUTELY DECOMPENSATED INTO RIGHT VENTRICULAR FAILURE (RVF). ON (B)(6) 2025, THE PATIENT WAS TERMINALLY WEANED. THE LVAD HAD FUNCTIONED AS INTENDED. NO EQUIPMENT WOULD BE RETURNED. Manufacturer Narrative: SECTION H6 CORRECTION: HEALTH EFFECT - IMPACT CODE 4624 - SURGICAL INTERVENTION ADDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE LVAD WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE IFU LISTS CARDIAC ARRHYTHMIA, RIGHT HEART FAILURE, AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE IFU ALSO LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THE IFU DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6 STATES: "RIGHT HEART FAILURE CAN OCCUR FOLLOWING IMPLANTATION OF THE PUMP. RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH ELEVATED PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP." NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22384027&pc=QZI | 2124215-2025-43904 | 2025-06-03 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-02 04:00 | QZI | FARAWAVE NAV | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest |  |  | 1 | Event Description: IT WAS REPORTED IN AN DISRUPT-AF STUDY, THE PATIENT PASSED AWAY FROM CARDIAC ARREST APPROXIMATELY 3 MONTHS AFTER HAVING A PULSED FIELD ABLATION (PFA) PROCEDURE. IT WAS NOT KNOWN WHAT CAUSED THE CARDIAC ARREST. IT WAS FURTHER REPORTED THE PFA PROCEDURE WAS PERFORMED TO TREAT PAROXYSMAL ATRIAL FIBRILLATION. THE PATIENT MEDICAL HISTORY CONDITIONS ATTRIBUTABLE TO THE PATIENTS DEATH IS ATRIAL FIBRILLATION, BLADDER CANCER, CAD, MI, NSTEMI, AND STROKE. THE PHYSICIAN WAS CONTACTED FOR ADDITIONAL INFORMATION. IT WAS FURTHER REPORTED THE PATIENT WAS SEEN FOR A ONE-MONTH FOLLOW-UP A FEW DAYS BEFORE THE EVENT, WITH NO OVERT COMPLICATIONS FROM THE PFA PROCEDURE. THE ECG DEMONSTRATED SINUS RHYTHM WITH POTENTIAL QT PROLONGATION. ADDITIONALLY, THE PATIENT HAD AN UNREMARKABLE EKG AT LEAST TWO MONTHS AFTER HER ABLATION, SHOWING NO SIGNS OF HEART BLOCK, LONG QTC, OR ANY INDICATION OF A RELATIONSHIP TO THE ABLATION PROCEDURE. THE PHYSICIAN DOES NOT BELIEVE THE EVENT IS RELATED TO THE PFA PROCEDURE. THE CARDIAC ARREST THE PATIENT EXPERIENCED APPEARS TO HAVE BEEN DUE TO PEA. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22384073&pc=NPT | 2025587-2025-04732 | 2024-10-31 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-02 04:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Fibrosis; Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 6 YEARS AND 5 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC AORTIC THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. FURTHER INFORMATION WAS NOT AVAILABLE.Event Description: ADDITIONAL INFORMATION RECEIVED INDICATED PULMONARY FIBROSIS LED TO THE PATIENT¿S DEATH. THE PHYSICIAN INDICATED THE EVENT WAS UNRELATED TO THE VALVE. AT THIS TIME, THERE WAS NO REPORT THE VALVE AS A CONTRIBUTING FACTOR TO THE PATIENT¿S DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22384207&pc=MNS | 2518422-2025-107024 | 2025-07-01 04:00 | Death | RESPIRONICS, INC. | 2025-07-02 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED BECAUSE THE PATIENT HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22384227&pc=NPT | 2135147-2025-03607 | 2007-06-01 04:00 | Death | ABBOTT MEDICAL | 2025-07-02 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P190023 |  | 1 | Event Description: THE ARTICLE, "COMPUTED TOMOGRAPHY SCAN PREDICTORS AND PROGNOSTIC IMPACT OF COMBINED PULMONARY HYPERTENSION IN PATIENTS WITH AORTIC VALVE STENOSIS UNDERGOING TAVI", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO EVALUATE THE CORRELATION BETWEEN COMPUTED TOMOGRAPHY (CT) DERIVED MAIN PULMONARY ARTERY (MPA), RIGHT PA (RPA), LEFT PA (LPA) DIAMETERS AND MPA/ASCENDING AORTA (AA) RATIO AND PULMONARY HYPERTENSION (PH) SUBTYPES DEFINED AT RIGHT HEART CATHETERIZATION (RHC) AND TO EVALUATE THEIR PROGNOSTIC IMPACT. DEVICES INCLUDED IN THE STUDY WERE SAPIEN/XT/3/3 ULTRA (EDWARDS LIFESCIENCES), AND MYVAL (MERIL LIFE SCIENCES), LOTUS (BOSTON SCIENTIFIC), COREVALVE/EVOLUT R/EVOLUT PRO (MEDTRONIC), ACURATE/NEO (BOSTON SCIENTIFIC), PORTICO (ABBOT VASCULAR, SANTA CLARA, CALIFORNIA, USA), NAVITOR (ABBOTT) AND JENAVALVE (JENAVALVE). THE ARTICLE CONCLUDED THAT INCREASE IN MPA/AA CT-DERIVED RATIO IS PREDICTIVE OF COMBINED PH, HIGHLIGHTING PATIENTS WHO COULD BENEFIT FROM A RHC IN TERM OF CARDIOVASCULAR (CV) STRATIFICATION BEFORE TAVI. [THE PRIMARY AND CORRESPONDING AUTHOR WAS (B)(6), DEPARTMENT OF CARDIAC, THORACIC AND VASCULAR SCIENCES AND PUBLIC HEALTH, (B)(6), WITH CORRESPONDING EMAIL: GIUSEPPE.TARANTINI.1@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS FROM JUNE 2007 TO DECEMBER 2022. A TOTAL OF 638 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH IT WAS NOT SPECIFIED HOW MANY RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 81 YEARS (NO PH), 81 YEARS (IPCPH), AND 80 YEARS (COPH). THE MAJORITY GENDER COULD NOT BE DETERMINED. COMORBIDITIES INCLUDED ARTERIAL HYPERTENSION, DIABETES MELLITUS, DYSLIPIDEMIA, SMOKING HISTORY, CORONARY ARTERY DISEASE, CHRONIC KIDNEY DISEASE, AND PULMONARY HYPERTENSION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF COMPUTED TOMOGRAPHY SCAN PREDICTORS AND PROGNOSTIC IMPACT OF COMBINED PULMONARY HYPERTENSION IN PATIENTS WITH AORTIC VALVE STENOSIS UNDERGOING TAVI WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING ARTERIAL HYPERTENSION, DIABETES MELLITUS, DYSLIPIDEMIA, SMOKING HISTORY, CORONARY ARTERY DISEASE, CHRONIC KIDNEY DISEASE, AND PULMONARY HYPERTENSION. ONE OF THE COMPLICATION REPORTED WAS DEATH; THIS COMPLICATION IS ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: COMPUTED TOMOGRAPHY SCAN PREDICTORS AND PROGNOSTIC IMPACT OF COMBINED PULMONARY HYPERTENSION IN PATIENTS WITH AORTIC VALVE STENOSIS UNDERGOING TAVI. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: COMPUTED TOMOGRAPHY SCAN PREDICTORS AND PROGNOSTIC IMPACT OF COMBINED PULMONARY HYPERTENSION IN PATIENTS WITH AORTIC VALVE STENOSIS UNDERGOING TAVI.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22384423&pc=OZD | 1220648-2025-30014 | 2025-06-11 04:00 | Death | ABIOMED, INC. | 2025-07-02 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Anemia; Arrhythmia; Cardiac Arrest; Hemolysis; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 AND THE PATIENT CODED. THERE WERE CONCERNS FOR LOW HEMOGLOBIN/HEMATOCRIT LEVELS. THE PATIENT REQUIRED PACKED RED BLOOD CELLS. NO SOURCE OF BLEEDING WAS IDENTIFIED. THE PATIENT WENT INTO ARRYTHMIA AND CARDIOPULMONARY RESUSCITATION WAS PERFORMED. ULTIMATELY, THE PATIENT PASSED AWAY. THE MEDICAL DOCTOR DID NOT FEEL THIS TO BE IMPELLA RELATED.Event Description: IT WAS REPORTED MULTIPLE BLOOD PRODUCTS GIVEN FOR BLEEDING. THE SOURCE OF THE BLEEDING WAS NEVER DETERMINED. ON THE LAST DAY OF SUPPORT THE PATIENT WAS NOTED TO HAVE HEMOLYSIS AND WAS TRANSFUSED WITH 2 UNITS OF RED BLOOD CELLS (RBC) BLOOD PRODUCT. PATIENT EXPIRED. CAUSE OF DEATH NOTED AS UNKNOWN; BLEEDING SUSPECTED AS LABS SHOWED A HEMATOCRIT WAS 5. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL, VF/VT/AF/AT, AND HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING: SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ B2 INTERVENTION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30014. B5 ADDITIONAL INFORMATION THAT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30014. H6 HEALTH EFFECT - CLINICAL CODES 1888 AND 1886 WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30014. H10 INSTRUCTIONS FOR USE WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30014. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22384732&pc=OZD | 1220648-2025-30015 | 2025-06-12 04:00 | Death | ABIOMED, INC. | 2025-07-02 04:00 | OZD | IMPELLA 5.5 | Failure to Advance | Cardiac Perforation; Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AFTER MULTIPLE FAILED ATTEMPTS TO ADVANCE AN IMPELLA 5.5 FOR SUPPORT ESCALATION, THE PROCEDURE WAS ABORTED. THE PHYSICIAN WAS UNABLE TO ADVANCE THE IMPELLA DUE TO THE PATIENT¿S ANATOMY. IMMEDIATELY FOLLOWING THE DELIVERY FAILURE, THE PATIENT WAS PLACED ON ECPELLA SUPPORT WITH AN EXISTING IMPELLA IMPLANT. HOWEVER, THE PATIENT DECOMPENSATED AND PASSED AWAY WITHIN THE HOUR. IT IS UNKNOWN IF THE FAILURE TO ESCALATE SUPPORT CONTRIBUTED TO THE PATIENT'S PASSING SHORTLY THEREAFTER.Event Description: A TRANSESOPHAGEAL ECHOCARDIOGRAM SHOWED THAT THE PATIENT HAD A LARGE PERICARDIAL EFFUSION. THE PHYSICIAN OPENED THE CHEST AND DRAINED EFFUSION AND MULTIPLE BLOOD PRODUCTS WERE GIVEN. HOWEVER, THE PATIENT DECOMPENSATED AND PASSED AWAY WITHIN THE HOUR. Manufacturer Narrative: THE IMPELLA PUMP WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES AND VENTRICLE PERFORATION HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS RETURNED FOR EVALUATION. FAILURE MODE PRODUCT DAMAGE (CATHETER KINK) ADDED AFTER PRODUCT EVALUATION SINCE A CATHETER KINK WAS FOUND. GUIDEWIRE WAS DELIVERED SUCCESSFULLY BUT UNABLE TO ADVANCE IMPELLA 5.5 AFTER MULTIPLE ATTEMPTS DUE TO PATIENT ANATOMY. DELIVERY WAS ABORTED AND PATIENT WAS CANNULATED FOR EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) INSTEAD. IMPELLA 5.5 WAS RETURNED AND A CATHETER KINK WAS FOUND BELOW THE MOTOR HOUSING. DELIVERY ISSUE - THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED SINCE THE PUMP WAS UNABLE TO PASS THROUGH DIFFICULT ANATOMY. CATHETER KINK - THE ROOT CAUSE OF THE PRODUCT DAMAGE (CATHETER KINK) WAS MOST LIKELY PATIENT CONDITION RELATED SINCE THE PUMP WAS UNABLE TO PASS THROUGH DIFFICULT ANATOMY. VENTRICLE PERFORATION - THE ROOT CAUSE OF THE VENTRICLE PERFORATION WAS NOT DETERMINED DUE TO INSUFFICIENT CLINICAL INFORMATION. IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ B5 ADDITIONAL INFORMATION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30015. H6 HEALTH EFFECT - CLINICAL CODE 4580 WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30015. H6 HEALTH EFFECT - CLINICAL CODE 213 AND HEALTH EFFECT - IMPACT CODES 4643 AND 4624 WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30015. H10 INSTRUCTIONS FOR USE WERE INCOMPLETE ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30015.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22384764&pc=NPT | 2025587-2025-04734 | 2023-11-20 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-02 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Non specific EKG/ECG Changes; Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 YEAR AND 7 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO AN UNKNOWN REASON.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH WAS CARDIAC ARREST. ADDITIONALLY, THE VALVE COULD BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE DEATH. IT WAS NOTED THAT THE DEATH COULD HAVE BEEN RELATED TO THE PATIENT'S UNDERLYING MEDICAL HISTORY BECAUSE THE PATIENT HAD A PREVIOUS PERMANENT PACEMAKER.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PERMANENT PACEMAKER WAS IMPLANTED APPROXIMATELY 13 DAYS FOLLOWING THE INITIAL IMPLANT OF THE VALVE DUE TO AN UNSPECIFIED CONDUCTION DISTURBANCE. Manufacturer Narrative: UPDATED DATA: B2, B5, H6. CORRECTED DATA: B3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22385242&pc=NPT | 2025587-2025-04740 | 2023-06-08 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-02 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE, FOLLOWING DEPLOYMENT, THE PATIENT WAS HYPOTENSIVE. MEDICATION WAS ADMINISTERED AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED TO CIRCULATE THE MEDICATION. THE PATIENT WAS MOVED TO THE INTENSIVE CARE UNIT (ICU) AND SUBSEQUENTLY STARTED DECOMPENSATING. IT WAS NOTED THAT THE PATIENT HAD A LEFT VENTRICLE PERFORATION THAT HAD RESULTED IN PERICARDIAL EFFUSION. IT WAS REPORTED THAT THE GUIDEWIRE HAD CAUSED THE PERFORATION. THE TYPE OF GUIDEWIRE WAS UNKNOWN. SURGICAL INTERVENTION WAS PERFORMED TO REPAIR THE LEFT VENTRICLE. FOLLOWING SURGERY, THE PATIENT RETURNED TO THE ICU. LATER THAT DAY, THE PATIENT EXPIRED. THE CAUSE OF DEATH WAS NOT REPORTED. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: GWBC30 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0193-GUIDEWIRE; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID D-EVOLUTFX-2329 (UNKNOW SERIAL/LOT); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED: G3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22385394&pc=OZD | 1220648-2025-30018 | 2025-05-14 04:00 | Death | ABIOMED, INC. | 2025-07-02 04:00 | OZD | IMPELLA 5.5 | Deformation Due to Compressive Stress; Device Sensing Problem | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND DURING SUPPORT, THE PLACEMENT SIGNAL BECAME INTERMITTENT LATER NOTED TO BE FROM A KINK IN THE PUMP CATHETER. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. THE PATIENT PRESENTED TO THE HOSPITAL NOTED WITH "MORE" HEART FAILURE SYMPTOMS AND WAS IMPLANTED WITH THE IMPELLA 5.5 DEVICE FOR MCS AND BRIDGE TO TRANSPLANT. APPROXIMATELY FIVE DAYS AFTER START OF THE IMPELLA MCS, THE PLACEMENT SIGNAL BECAME INTERMITTENT, ERRATIC. THE STAFF WAS ADVISED THAT THE DEVICE COULD STILL FUNCTION WITHOUT THE PLACEMENT SIGNAL AND WALKED THROUGH HOW TO DISABLE THE PLACEMENT SIGNAL IF THE ALARM COULD NOT BE RESOLVED. ADDITIONALLY, THE STAFF WAS INFORMED OF HOW TO USE HEMODYNAMICS, ECHO MONITORING AND TRENDS TO MONITOR THE SUPPORT STATUS. SOME TIME THEREAFTER, THE ATTENDING NOTICED A KINK IN THE IMPELLA CATHETER AT THE DRESSING SITE. THE KINK WAS FIXED, AND THE PLACEMENT SIGNAL NORMALIZED. THE STATUS OF THE PATIENT FOLLOWING THE EVENT WAS INDICATED AS UNKNOWN. HOWEVER, IMPELLA MCS CONTINUED. APPROXIMATELY NINE DAYS LATER IT WAS NOTED THE PATIENT WAS TAKEN TO THE ELECTROPHYSIOLOGY LAB FOR ARTERIAL LEAD PLACEMENT. THE PATIENT WAS NOTED TO HAVE REMAINED WITH RESPIRATORY ISSUES REQUIRING DAILY BRONCHOSCOPY ONCE LUNGS WERE STABLE. LATER THAT AFTERNOON, IT WAS REPORTED THE PATIENT ARRESTED FOR APPROXIMATELY TWO HOURS. THE FAMILY WAS CALLED TO BEDSIDE, THE DECISION WAS MADE FOR COMFORT MEASURES ONLY, AND THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF OPTICAL SIGNAL ISSUE AND PRODUCT DAMAGE (CANNULA KINK) HAS BEEN COMPLETED. NO PRODUCT OR DATA LOGS WERE RETURNED FOR EVALUATION. OPTICAL SIGNAL ISSUE: CLINICAL DETAIL NOTED THAT INTERMITTENT "PLACEMENT SIGNAL NOT RELIABLE" ALARMS WERE NOTED DURING SUPPORT WITH ERRATIC PS WAVEFORMS. ALL OTHER WAVEFORMS WITHIN NORMAL LIMITS. ISSUE RESOLVED IN THE MORNING. ADDITIONALLY, AO PLACEMENT SIGNAL WAS READING HIGH AORTIC PRESSURES WHICH WAS OUT OF PROPORTION FROM THE CLINICAL A-LINE. A KINK IN THE IMPELLA CATHETER WAS OBSERVED AT THE SITE DRESSING. KINK WAS FIXED AND THE AO PLACEMENT SIGNAL NORMALIZED. THE ROOT CAUSE OF THE OPTICAL SIGNAL ISSUE COULD NOT BE DEFINITIVELY DETERMINED AS NO PRODUCT OR LOGS WERE RETURNED FOR EVALUATION. PRODUCT DAMAGE (CANNULA KINK): CLINICAL DETAILS NOTED THAT A KINK IN THE CATHETER AT THE SITE DRESSING WAS OBSERVED DURING SUPPORT. KINK WAS ABLE TO BE FIXED. THE ROOT CAUSE OF THE PRODUCT DAMAGE WAS MOST LIKELY A CATHETER KINK PER CLINICAL DETAILS, HOWEVER, THE CAUSE OF THE CATHETER KINK WAS UNABLE TO BE DETERMINED. H6 MEDICAL DEVICE PROBLEM CODE 2917 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30018. H6 TYPE OF INVESTIGATION 4112 AND INVESTIGATION FINDINGS 3211 WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30018. H10 INVESTIGATION RESULTS WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30018. Manufacturer Narrative: IT IS UNKNOWN IF THERE WAS ANY IMPACT TO THE FLOW AS A RESULT OF THE KINKED CATHETER. THUS, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR TO THE OVERALL OUTCOME. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITES STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22385615&pc=DSQ | 2916596-2025-04464 | 2025-06-15 04:00 | Death | THORATEC CORPORATION | 2025-07-02 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Dyspnea; Fatigue; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure; Gastrointestinal Hemorrhage; Swelling/ Edema | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN HOSPICE CARE DUE TO WITHDRAWAL FROM CARE AND THE PUMP WAS TURNED OFF. THE PATIENT WAS IN TERMINAL RIGHT HEART FAILURE.Event Description: IT WAS LATER REPORTED THAT THE PATIENT'S RIGHT HEART FAILURE DID NOT EXIST PRE-LEFT VENTRICULAR ASSIST DEVICE IMPLANT AND THAT IT WAS UNKNOWN IF THE DEVICE CAUSED OR CONTRIBUTED TO THE FAILURE. THE PATIENT EXPERIENCED SYMPTOMS OF SWELLING IN LEGS, SHORTNESS OF BREATH, FATIGUE, LOSS OF APPETITE, AND LOW FLOW ALARMS. DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND WAS SAID TO HAVE BEEN OPERATING AS EXPECTED. IT WAS ALSO SAID THAT THE PATIENT HAD A HISTORY OF THROMBUS IN OUTFLOW CANNULA AND WAS OFF ANTICOAGULATION DUE TO RECURRENT GASTROINTESTINAL BLEEDS. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22385720&pc=LLI | 22385720 | 2025-06-09 04:00 | Death | INVACARE | 2025-07-02 04:00 | LLI | INVACARE | Adverse Event Without Identified Device or Use Problem | Physical Entrapment |  |  | 1 | Event Description: PT WAS FOUND DECEASED AND LYING ON THE FLOOR, WITH HIS ARM CAUGHT IN THE BED RAIL. PT'S HEAD WAS ON THE FLOOR WITH MODERATE BLOOD ON THE FLOOR. LOWER BODY SUSPENDED IN AIR. MEDICAL EXAMINER CONTACTED AND DECLINED CASE. DELAY IN OBTAINING DEVICE INFO FROM THE EXTERNAL COMPANY THAT PROVIDED BED AND RAILS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22385725&pc=LLI | 22385725 | 2025-06-09 04:00 | Death | INVACARE | 2025-07-02 04:00 | LLI | INVACARE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Physical Entrapment |  |  | 1 | Event Description: PT WAS FOUND DECEASED AND LYING ON THE FLOOR, WITH HIS ARM CAUGHT IN THE BED RAIL. PT'S HEAD WAS ON THE FLOOR WITH MODERATE BLOOD ON THE FLOOR. LOWER BODY SUSPENDED IN AIR. MEDICAL EXAMINER CONTACTED AND DECLINED CASE. DELAY IN OBTAINING DEVICE INFO FROM THE EXTERNAL COMPANY THAT PROVIDED BED AND RAILS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22385928&pc=DRE | 2124215-2025-44042 | 2025-06-11 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-07-02 04:00 | DRE | VERSACROSS CONNECT LAAC ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Air Embolism; Bradycardia; Cardiac Arrest | K220414 |  | 1 | Event Description: IT WAS REPORTED A PATIENT DEATH OCCURRED. THE PROCEDURE WAS CANCELLED. DURING A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A VERSACROSS CONNECT LAAC WAS SELECTED FOR USE. THE PHYSICIAN PERFORMED A SUCCESSFUL TRANSSEPTAL PUNCTURE (TSP) USING WATCHMAN DOUBLE CURVE SHEATH (WAS). WHEN REMOVING THE DILATOR, IT WAS NOTED THAT THERE WAS NO BLEED. THE END OF THE SHEATH WAS LIFTED UP AND HEARD A VACUUM/SUCKING SOUND, QUICKLY PLACED THE DILATOR BACK DOWN ON THE TABLE, LIFTED END OF THE SHEATH UP AGAIN AND HEARD THE SAME SUCKING SOUND AGAIN, QUICKLY LOWERED THE SHEATH DOWN. THIS IS WHEN IT WAS NOTED AIR BUBBLES ON THE INTRACARDIAC ECHOCARDIOGRAPHY (ICE) IMAGES AND THE PATIENT'S HEART RATED DROPPED. TRANSCUTANEOUS PACING WAS INITIATED, PATIENT WAS INTUBATED, FLUIDS GIVEN AND CODE STARTED ONCE PATIENT CONTINUED TO DECLINE. CATH TEAM WAS CALLED IN. THE PATIENT WAS UNABLE TO BE SUCCESSFULLY RESUSCITATED. THE CODE ENDED AND PATIENT PRONOUNCED DEAD APPROXIMATELY AT 9:30 AM. IN THE PHYSICIAN OPINION, HE WAS UNCERTAIN IF THE PATIENT POSSIBLY HAD A VERY LOW OR NEGATIVE LEFT ATRIUM PRESSURE OR POSSIBLY TAKING A BREATH DURING MAC ANESTHESIA CAUSING NEGATIVE PRESSURE RESULTING IN NO BLEED-BACK FROM THE SHEATH AFTER THE DILATOR WAS REMOVED. THE PHYSICIANS WERE CERTAIN THE SHEATH/DILATOR WERE FLUSHED PROPERLY PRIOR TO INSERTION. SHEATH WAS INSPECTED AFTER EVENT AND APPEARS TO BE INTACT AND WORKING PROPERLY. THE DILATOR IS EXPECTED TO BE RETURNED FOR ANALYSIS. AIR ENTRY BELIEVED TO BE AT THE PORT OF THE DOUBLE CURVE SHEATH RESULTING IN AIR IN THE LEFT SIDE OF THE HEART. THE VALVE WAS CLOSED AT THE END OF THE SHEATH. ASPIRATION FROM THE SIDE PORT. FLUID AND ANESTHESIA SUPPORT. AIR INGRESS WHEN REMOVED DILATOR TO LEFT SIDE OF THE HEART. PATIENT STATUS DECLINED, CODE INITIATED, THEN DEATH. AIR INGRESS TO LEFT SIDE OF HEART CAUSING CARDIAC ARREST. NO DIFFICULTY WITH THE TSP WORKFLOW THAT WAS APPARENT DURING THE PROCEDURE. THE PHYSICIAN DID NOTE THAT THE FEMORAL VEIN ANATOMY WAS TORTUOUS BUT DID NOT APPEAR TO HAVE DIFFICULTY ADVANCING THE SHEATH. THE PHYSICIAN DID NOT MENTION ANYTHING ABOUT A DEFECTIVE PRODUCT OR THAT THE SHEATH OR VERSACROSS CONNECT WAS THE ISSUE. IN THE PHYSICIAN OPINION, IT WAS MORE OF A LOW OR NEGATIVE LEFT ATRIAL PRESSURE AND THE USE OF MAC THAT MAY HAVE CONTRIBUTED TO THE PROBLEM. Manufacturer Narrative: THE DEVICE HAS NOT BEEN RECEIVED FOR ANALYSIS. UPON RECEIPT AND COMPLETION OF THE FAILURE ANALYSIS OF THE COMPLAINT DEVICE, IF THERE IS ANY FURTHER RELEVANT INFORMATION FROM THAT REVIEW, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22385949&pc=OZD | 1220648-2025-30023 | 2025-05-14 04:00 | Death | ABIOMED, INC. | 2025-07-02 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Failure to Pump; Device Sensing Problem | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND WAS TRANSFERRED TO A DIFFERENT HOSPITAL. UPON AIC (CONSOLE) TO AIC TRANSFER, AN "IMPELLA DEFECTIVE¿ ALARM APPEARED. A NEW AIC USED AND ¿PLACEMENT SIGNAL NOT RELIABLE¿ ALARMED. DURING THIS TIME, THE IMPELLA WAS SHUT OFF THREE TIMES AND RESTARTED. IT WAS FOUND THAT THE IMPELLA CORD NOT COMPLETELY PLUGGED IN AS A SMALL GAP WAS VISIBLE BETWEEN CONSOLE PLUG IN AND THE IMPELLA CORD. IMPELLA CORD UNPLUGGED TO ASSESS PRONGS, WHICH WERE ALL STRAIGHT THEN REPLUGGED IN, STILL NOT FITTING FLUSH INTO PLUG IN. THE DOCTOR WAS OK WITH DISABLING PLACEMENT SIGNAL (PS) UNRELIABLE ALARM. THE ALARMS WERE MUTED. SUPPORT CONTINUED. HOWEVER, IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITES STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22386227&pc=GEX | 2124215-2025-43846 | 2024-11-11 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-02 04:00 | GEX | UNKNOWN | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism |  |  | 1 | Event Description: A RETROSPECTIVE OBSERVATIONAL STUDY PUBLISHED IN THE WORLD JOURNAL OF UROLOGY WAS CONDUCTED AT THE UNIVERSITY HOSPITAL CENTER OF BESANCON, FRANCE, COMPARING POSTOPERATIVE COMPLICATIONS AFTER LASER PHOTO VAPORIZATION OF THE PROSTATE (PVP) IN ANTICOAGULATED VERSUS NON-ANTICOAGULATED PATIENTS, INCLUDING ALL INDIVIDUALS WHO UNDERWENT PVP BETWEEN 2013 AND 2022, FOR A TOTAL OF 811 PATIENTS AND A FINAL MATCHED COHORT OF 245 AFTER PROPENSITY SCORE MATCHING (95 WITH ANTICOAGULANT THERAPY AND 150 WITHOUT). ALL PROCEDURES WERE PERFORMED USING THE GREENLIGHT 180 W XPS LASER SYSTEM EQUIPPED WITH A MOXY FIBER, FOLLOWING INTERNATIONAL USER GUIDE RECOMMENDATIONS: 80 W POWER WAS INITIALLY APPLIED AT THE BLADDER NECK, INCREASED UP TO 180 W FOR VAPORIZATION OF THE LATERAL AND MEDIAN LOBES (STARTING AT THE PROSTATE BASE), AND RETURNING TO 80 W AT THE APEX, ALWAYS WITHOUT EXCEEDING THE VERUMONTANUM, MAINTAINING A 3-5 MM DISTANCE BETWEEN PROSTATE AND FIBER, AND GUIDED BY INTRAOPERATIVE PROSTATIC ENDORECTAL ULTRASOUND; THE PROCEDURE WAS COMPLETED WHEN THE CAPSULE FIBERS BECAME VISIBLE, AND A THREE-WAY 22 CH BLADDER CATHETER WAS LEFT IN PLACE WITH IRRIGATION UNTIL THE NEXT MORNING. WITHIN THREE MONTHS POST-SURGERY, 86 PATIENTS (35%) EXPERIENCED COMPLICATIONS, THE MAIN ONES BEING SIGNIFICANT HEMATURIA (38 CASES, 15%), ACUTE URINARY RETENTION (54 CASES, 22%), AND URINARY SEPSIS (24 CASES, 9.8%), WITH 92% OF ALL EVENTS CLASSIFIED AS MILD OR MODERATE. THE ANTICOAGULATED GROUP SHOWED A HIGHER OVERALL COMPLICATION RATE (43% VS. 30%) AND, SPECIFICALLY, A HIGHER INCIDENCE OF HEMATURIA (26% VS. 8.7%). REGARDING SEVERE COMPLICATIONS, ONE ANTICOAGULATED PATIENT REQUIRED TRANSFUSION DUE TO POSTOPERATIVE RETZIUS HEMATOMA EVACUATION, ANOTHER DEVELOPED POSTOPERATIVE MYOCARDIAL INFARCTION REQUIRING STENT PLACEMENT, AND A THIRD EXPERIENCED SEPTIC SHOCK NECESSITATING RESUSCITATION MEASURES; IN THE NON-ANTICOAGULATED GROUP, TWO PATIENTS REQUIRED SURGERY FOR POSTOPERATIVE HEMATURIA, AND ONE DIED FROM POSTOPERATIVE PULMONARY EMBOLISM. IN CONCLUSION, THIS STUDY DEMONSTRATES THAT LASER PHOTO VAPORIZATION OF THE PROSTATE PERFORMED WITH THE GREENLIGHT 180 W XPS SYSTEM AND MOXY FIBER IS A SAFE PROCEDURE EVEN FOR FRAIL AND ANTICOAGULATED PATIENTS, WITH A LOW RISK OF SEVERE COMPLICATIONS AND TRANSFUSION, SEVERITY AND HOSPITAL STAY COMPARABLE TO EXISTING LITERATURE, BUT WITH A HIGHER INCIDENCE OF HEMATURIA IN ANTICOAGULATED INDIVIDUALS, MOST OF WHICH WAS MANAGED MEDICALLY. THE RELATIONSHIP BETWEEN THE PROCEDURE OR DEVICE AND THE DEATH WAS NOT PROVIDED.
THIS REPORT REFERS TO THE DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22386444&pc=MCX | 3004742232-2025-00191 | 2025-06-11 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-07-02 04:00 | MCX | DIAMONDBACK CORONARY | Unintended System Motion; Failure to Advance | Ischemia; Perforation of Vessels | P130005 |  | 1 | Event Description: IT WAS REPORTED THAT THE DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) AND VIPERWIRE ADVANCE CORONARY GUIDE WIRE WITH FLEX TIP WERE USED FOR TREATMENT OF HEAVILY CALCIFIED, HEAVILY TORTUOUS, 98-99% STENOSED LESION IN PROXIMAL TO MID RIGHT CORONARY ARTERY (RCA). PRIMARY WIRING WITH THE VIPERWIRE WAS ATTEMPTED BUT IT COULD NOT GET THROUGH. THE VIPERWIRE WAS EXCHANGED FOR A FEW DIFFERENT UNSPECIFIED WIRES AND FINALLY CROSSED WITH A NON-ABBOTT GUIDE WIRE. THE OAD WAS THEN ADVANCED BUT COULD NOT CROSS THE LESION IN THE RCA DURING SEVERAL ATTEMPTS DUE TO PATIENT ANATOMY. THE WIRE WAS STARTING TO COME BACK WHILE TRYING TO ADVANCE THE CROWN. OAD SPINNING WAS STOPPED AND THE TIP BUSHING WAS ATTEMPTED TO BE ADVANCED FURTHER INTO THE LESION WITH THE OAD OFF TO ADD SUPPORT TO ADVANCING THE VIPERWIRE FURTHER INTO THE RCA BUT THE WIRE COULD NOT BE ADVANCED DUE TO LACK OF GUIDE SUPPORT. AROUND THIS TIME, IT WAS REALIZED THE TIP BUSHING OF THE OAD APPEARED TO BE OUTSIDE THE VESSEL WALL. ANGIOGRAPHIC IMAGE SHOWED A LARGE VESSEL PERFORATION THAT WAS BLEEDING INTO THE PERICARDIUM. THE OAD WAS THEN REMOVED, AND ALL WIRE POSITIONS WERE LOST DURING OAD REMOVAL. THE PHYSICIAN WAS ABLE TO WIRE BACK TO THE POINT OF THE PERFORATION AND OCCLUDED THE RCA WITH ABBOTT TREK BALLOON AT THE OSTIUM BUT COULD NOT RECROSS THE LESION TO PLACE A COVERED STENT INTO PLACE. SURGERY WAS THE ONLY OPTION TO REPAIR THE VESSEL AND BYPASS IT. THE PROCEDURE WAS ABORTED DUE TO PERFORATION. THE PATIENT WAS SENT TO OPERATING ROOM FOR EMERGENCY BYPASS GRAFT TO THE RCA. THE PATIENT WAS PLACED ON VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) AT THE END OF THE SURGERY. THE PATIENT WAS IN CRITICAL CONDITION. ON (B)(6) 2025 THE PATIENT REMAINED IN HOSPITAL ON ECMO. ON (B)(6) 2025, THE PATIENT EXPIRED. THE PRIMARY CAUSE OF DEATH WAS COMPLICATION IN THE CARDIAC CATHETERIZATION LABORATORY DUE TO PERFORATION. THE SECONDARY CAUSE OF DEATH WAS COMPLICATION FROM ECMO. IN THE OPINION OF THE PHYSICIAN, THE OAD CAUSED OR CONTRIBUTED TO THE PERFORATION AS IT WENT THROUGH THE VESSEL WALL DUE TO SIGNIFICANT TORTUOSITY. DURING THE SECOND TREATMENT, THE PHYSICIAN FELT THE OAD CROWN JUMP AND SAW THE VIPERWIRE BUCKLE OUT; THERE WAS A LOT OF WIRE BIAS. ELECTROCARDIOGRAM ST CHANGES WERE OBSERVED.Event Description: IT WAS REPORTED THAT THE DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) AND VIPERWIRE ADVANCE CORONARY GUIDE WIRE WITH FLEX TIP WERE USED FOR TREATMENT OF HEAVILY CALCIFIED, HEAVILY TORTUOUS, 98-99% STENOSED LESION IN PROXIMAL TO MID RIGHT CORONARY ARTERY (RCA). PRIMARY WIRING WITH THE VIPERWIRE WAS ATTEMPTED BUT IT COULD NOT GET THROUGH. THE VIPERWIRE WAS EXCHANGED FOR A FEW DIFFERENT UNSPECIFIED WIRES AND FINALLY CROSSED WITH A NON-ABBOTT GUIDE WIRE. THE OAD WAS THEN ADVANCED BUT COULD NOT CROSS THE LESION IN THE RCA DURING SEVERAL ATTEMPTS DUE TO PATIENT ANATOMY. THE WIRE WAS STARTING TO COME BACK WHILE TRYING TO ADVANCE THE CROWN. OAD SPINNING WAS STOPPED AND THE TIP BUSHING WAS ATTEMPTED TO BE ADVANCED FURTHER INTO THE LESION WITH THE OAD OFF TO ADD SUPPORT TO ADVANCING THE VIPERWIRE FURTHER INTO THE RCA BUT THE WIRE COULD NOT BE ADVANCED DUE TO LACK OF GUIDE SUPPORT. AROUND THIS TIME, IT WAS REALIZED THE TIP BUSHING OF THE OAD APPEARED TO BE OUTSIDE THE VESSEL WALL. ANGIOGRAPHIC IMAGE SHOWED A LARGE VESSEL PERFORATION THAT WAS BLEEDING INTO THE PERICARDIUM. THE OAD WAS THEN REMOVED, AND ALL WIRE POSITIONS WERE LOST DURING OAD REMOVAL. THE PHYSICIAN WAS ABLE TO WIRE BACK TO THE POINT OF THE PERFORATION AND OCCLUDED THE RCA WITH ABBOTT TREK BALLOON AT THE OSTIUM BUT COULD NOT RECROSS THE LESION TO PLACE A COVERED STENT INTO PLACE. SURGERY WAS THE ONLY OPTION TO REPAIR THE VESSEL AND BYPASS IT. THE PROCEDURE WAS ABORTED DUE TO PERFORATION. THE PATIENT WAS SENT TO OPERATING ROOM FOR EMERGENCY BYPASS GRAFT TO THE RCA. THE PATIENT WAS PLACED ON VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) AT THE END OF THE SURGERY. THE PATIENT WAS IN CRITICAL CONDITION. ON (B)(6) 2025 THE PATIENT REMAINED IN HOSPITAL ON ECMO. ON (B)(6) 2025, THE PATIENT EXPIRED. THE PRIMARY CAUSE OF EXPIRATION IS COMPLICATION IN THE CARDIAC CATHETERIZATION LABORATORY DUE TO PERFORATION. THE SECONDARY CAUSE OF EXPIRATION IS COMPLICATION FROM ECMO. IN THE OPINION OF THE PHYSICIAN, THE OAD CAUSED OR CONTRIBUTED TO THE PERFORATION AS IT WENT THROUGH THE VESSEL WALL BECAUSE OF SIGNIFICANT TORTUOSITY. DURING THE SECOND TREATMENT, THE PHYSICIAN FELT THE OAD CROWN JUMP AND SAW THE VIPERWIRE BUCKLE OUT; THERE WAS A LOT OF WIRE BIAS. ELECTROCARDIOGRAM ST CHANGES WERE OBSERVED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. Manufacturer Narrative: CORRECTION: B2. ADDITIONAL INFORMATION: H6 (CODES 4118, 3233, AND 11 NO LONGER APPLY), H10. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED FAILURE TO ADVANCE, CROWN JUMP, PATIENT DEATH, HOSPITALIZATION, PERFORATION OF VESSELS, ISCHEMIA, DELAY TO TREATMENT, SURGICAL INTERVENTION, AND UNEXPECTED MEDICAL INTERVENTION APPEARS TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE IT IS POSSIBLE THAT THE REPORTED FAILURE TO ADVANCE, CROWN JUMP, PATIENT DEATH, HOSPITALIZATION, PERFORATION OF VESSELS, ISCHEMIA, DELAY TO TREATMENT, SURGICAL INTERVENTION, AND UNEXPECTED MEDICAL INTERVENTION ARE A RESULT OF THE PATIENT¿S DISEASE SEVERITY AND/OR USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22386974&pc=HAW | 1723170-2025-02589 | 2025-05-09 04:00 | Death | MEDTRONIC NAVIGATION, INC | 2025-07-02 04:00 | HAW | MEDTRONIC NAVIGATION | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: DOI: 10.3171/2025.1.PEDS24387. SUMMARY: OBJECTIVE STEREOTACTIC LASER ABLATION (SLA) OR LASER INTERSTITIAL THERMAL THERAPY (LITT) IS AN EMERGING ALTERNATIVE TO CONVENTIONAL SURGERY FOR CNS TUMORS. FURTHER CHARACTERIZATION OF ITS EFFECTIVENESS AND SAFETY IN THE PEDIATRIC POPULATION IS NEEDED. METHODS A REVIEW WAS CONDUCTED ACCORDING TO PRISMA GUIDELINES USING THE PUBMED/MEDLINE, SCOPUS, EMBASE, GOOGLE SCHOLAR, AND SCIENCE DIRECT DATABASES TO INVESTIGATE THE EFFECTIVENESS AND SAFETY IN THE USE OF LITT TO TREAT PEDIATRIC CNS TUMORS IN CHILDREN. RESULTS A TOTAL OF 24 ARTICLES MET THE INCLUSION CRITERIA. FROM THE 4 RETROSPECTIVE STUDIES COMPRISING 97 PATIENTS (MEAN AGE 11.4 YEARS, 55.0% MALES IN THE POOLED PREVALENCE), THE MOST COMMON TUMOR TYPES WERE PILOCYTIC ASTROCYTOMA (31.5%) AND SUBEPENDYMAL GIANT CELL ASTROCYTOMA (SEGA; 68.5%). TUMORS WERE PRIMARILY LOCATED IN THE FRONTAL LOBE (29.7%) AND THALAMUS (24.4%). POSTOPERATIVE COMPLICATIONS INCLUDED TRANSIENT NEUROLOGICAL DEFICITS IN 12.1% AND PERMANENT DEFICITS IN 6.0% OF PATIENTS. AT A MEAN FOLLOW-UP OF 43.9 MONTHS, MASS REDUCTION WAS OBSERVED IN 68.8% OF PATIENTS AND OVERALL DISEASE IMPROVEMENT IN 91.9% OF PATIENTS, AND THE MORTALITY RATE WAS 2.6%. FROM THE 20 CASE REPORTS/SERIES INVOLVING 67 PATIENTS (MEAN AGE 10.8 ± 4.7 YEARS, 52.2% MALES), THE MEAN TUMOR SIZE WAS 15.7 ± 8.7 CM3. PREDOMINANT TUMOR SUBTYPES WERE PILOCYTIC ASTROCYTOMA (29.9%) AND SEGA (16.4%). TUMORS WERE LOCATED IN ELOQUENT AREAS IN 61.3% OF CASES, NOTABLY THE THALAMUS (24.2%) AND VENTRICULAR SYSTEM (24.2%). PRIOR TREATMENTS INCLUDED SURGERY (78.8%), CHEMOTHERAPY (51.5%), AND IMMUNOTHERAPY (27.3%). KEY LITT PARAMETERS WERE DURATION (7.2 ± 8.8 MINUTES), DOSE (10.2 ± 2.4 W), AND EXTENT OF TUMOR VOLUME DECREASE (68.3% ± 30.4%). THE MOST COMMONLY USED LITT SYSTEM WAS VISUALASE (95.7%). POSTOPERATIVE COMPLICATIONS WERE REPORTED IN 26.9% OF PATIENTS, INCLUDING TRANSIENT NEUROLOGICAL DEFICITS (55.6%), PERILESIONAL EDEMA (22.2%), AND HYDROCEPHALUS (22.2%). HOSPITAL STAYS WERE = 3 DAYS IN 92.7% OF PATIENTS. TUMOR SIZE REDUCTION WAS ACHIEVED IN 86.7% OF PATIENTS, AND 78.9% EXPERIENCED DISEASE IMPROVEMENT. COMPARATIVE ANALYSIS SHOWED THAT A GREATER EXTENT OF ABLATION WAS ASSOCIATED WITH A REDUCED NEED FOR SECONDARY SURGERY (P = 0.038, OR 0.94) AND IMPROVED DISEASE OUTCOMES (P = 0.023, OR 1.05). LONGER LITT DURATION WAS SIGNIFICANTLY ASSOCIATED WITH POSTOPERATIVE COMPLICATIONS (P = 0.050). CONCLUSIONS LITT APPEARS EFFECTIVE IN REDUCING TUMOR SIZE AND IMPROVING DISEASE OUTCOMES IN PEDIATRIC CNS TUMORS. THE LONG-TERM EFFECTIVENESS OF LITT IN PEDIATRIC BRAIN TUMORS REQUIRES FURTHER RANDOMIZED PROSPECTIVE INVESTIGATION. REPORTED EVENT(S): THIS ARTICLE DOCUMENTED REVIEW OF 24 ARTICLES THAT DOCUMENT 20 CASE REPORTS / CASE SERIES AND 4 RETROSPECTIVE STUDIES. THESE REPORTS WERE COMPRISED OF 97 PATIENTS (54.99% MALE, 45.01% FEMALE) WITH A MEAN AGE OF 11.43 YEARS. THE MORTALITY RATE WAS 2.55%. Manufacturer Narrative: H3, H6: NO PRODUCTS WERE RETURNED TO MEDTRONIC FOR ANALYSIS. CODES B17, C20, AND D15 ARE APPLICABLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22386976&pc=GEX | 1723170-2025-02590 | 2025-05-09 04:00 | Death | MEDTRONIC NAVIGATION, INC | 2025-07-02 04:00 | GEX | MEDTRONIC NAVIGATION | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K081656 |  | 1 | Event Description: DOI: 10.3171/2025.1.PEDS24387. SUMMARY: OBJECTIVE STEREOTACTIC LASER ABLATION (SLA) OR LASER INTERSTITIAL THERMAL THERAPY (LITT) IS AN EMERGING ALTERNATIVE TO CONVENTIONAL SURGERY FOR CNS TUMORS. FURTHER CHARACTERIZATION OF ITS EFFECTIVENESS AND SAFETY IN THE PEDIATRIC POPULATION IS NEEDED. METHODS A REVIEW WAS CONDUCTED ACCORDING TO PRISMA GUIDELINES USING THE PUBMED/MEDLINE, SCOPUS, EMBASE, GOOGLE SCHOLAR, AND SCIENCE DIRECT DATABASES TO INVESTIGATE THE EFFECTIVENESS AND SAFETY IN THE USE OF LITT TO TREAT PEDIATRIC CNS TUMORS IN CHILDREN. RESULTS A TOTAL OF 24 ARTICLES MET THE INCLUSION CRITERIA. FROM THE 4 RETROSPECTIVE STUDIES COMPRISING 97 PATIENTS (MEAN AGE 11.4 YEARS, 55.0% MALES IN THE POOLED PREVALENCE), THE MOST COMMON TUMOR TYPES WERE PILOCYTIC ASTROCYTOMA (31.5%) AND SUBEPENDYMAL GIANT CELL ASTROCYTOMA (SEGA; 68.5%). TUMORS WERE PRIMARILY LOCATED IN THE FRONTAL LOBE (29.7%) AND THALAMUS (24.4%). POSTOPERATIVE COMPLICATIONS INCLUDED TRANSIENT NEUROLOGICAL DEFICITS IN 12.1% AND PERMANENT DEFICITS IN 6.0% OF PATIENTS. AT A MEAN FOLLOW-UP OF 43.9 MONTHS, MASS REDUCTION WAS OBSERVED IN 68.8% OF PATIENTS AND OVERALL DISEASE IMPROVEMENT IN 91.9% OF PATIENTS, AND THE MORTALITY RATE WAS 2.6%. FROM THE 20 CASE REPORTS/SERIES INVOLVING 67 PATIENTS (MEAN AGE 10.8 ± 4.7 YEARS, 52.2% MALES), THE MEAN TUMOR SIZE WAS 15.7 ± 8.7 CM3. PREDOMINANT TUMOR SUBTYPES WERE PILOCYTIC ASTROCYTOMA (29.9%) AND SEGA (16.4%). TUMORS WERE LOCATED IN ELOQUENT AREAS IN 61.3% OF CASES, NOTABLY THE THALAMUS (24.2%) AND VENTRICULAR SYSTEM (24.2%). PRIOR TREATMENTS INCLUDED SURGERY (78.8%), CHEMOTHERAPY (51.5%), AND IMMUNOTHERAPY (27.3%). KEY LITT PARAMETERS WERE DURATION (7.2 ± 8.8 MINUTES), DOSE (10.2 ± 2.4 W), AND EXTENT OF TUMOR VOLUME DECREASE (68.3% ± 30.4%). THE MOST COMMONLY USED LITT SYSTEM WAS VISUALASE (95.7%). POSTOPERATIVE COMPLICATIONS WERE REPORTED IN 26.9% OF PATIENTS, INCLUDING TRANSIENT NEUROLOGICAL DEFICITS (55.6%), PERILESIONAL EDEMA (22.2%), AND HYDROCEPHALUS (22.2%). HOSPITAL STAYS WERE = 3 DAYS IN 92.7% OF PATIENTS. TUMOR SIZE REDUCTION WAS ACHIEVED IN 86.7% OF PATIENTS, AND 78.9% EXPERIENCED DISEASE IMPROVEMENT. COMPARATIVE ANALYSIS SHOWED THAT A GREATER EXTENT OF ABLATION WAS ASSOCIATED WITH A REDUCED NEED FOR SECONDARY SURGERY (P = 0.038, OR 0.94) AND IMPROVED DISEASE OUTCOMES (P = 0.023, OR 1.05). LONGER LITT DURATION WAS SIGNIFICANTLY ASSOCIATED WITH POSTOPERATIVE COMPLICATIONS (P = 0.050). CONCLUSIONS LITT APPEARS EFFECTIVE IN REDUCING TUMOR SIZE AND IMPROVING DISEASE OUTCOMES IN PEDIATRIC CNS TUMORS. THE LONG-TERM EFFECTIVENESS OF LITT IN PEDIATRIC BRAIN TUMORS REQUIRES FURTHER RANDOMIZED PROSPECTIVE INVESTIGATION. REPORTED EVENT(S): THIS ARTICLE DOCUMENTED REVIEW OF 24 ARTICLES THAT DOCUMENT 20 CASE REPORTS / CASE SERIES AND 4 RETROSPECTIVE STUDIES. THESE REPORTS WERE COMPRISED OF 97 PATIENTS (54.99% MALE, 45.01% FEMALE) WITH A MEAN AGE OF 11.43 YEARS. THE MORTALITY RATE WAS 2.55%. Manufacturer Narrative: H3, H6: NO PRODUCTS WERE RETURNED TO MEDTRONIC FOR ANALYSIS. CODES B17, C20, AND D15 ARE APPLICABLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22387533&pc=NIU | 2124215-2025-44117 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-02 04:00 | NIU | ELUVIATM DRUG-ELUTING VASCULAR STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | RWD2400413 | 25 | Event Description: THIS REPORT SUMMARIZES 25 REPORTED INCIDENTS FOR PATIENT DEATHS WITHIN 30 DAYS POST INDEX PROCEDURE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. THE ELUVIA DRUG-ELUTING VASCULAR STENT SYSTEM WAS SELECTED FOR USE. PATIENT DEATHS WERE REPORTED WITHIN 30 DAYS FOLLOWING THE INDEX PROCEDURE. 72.5% OF PATIENTS WERE OVER THE AGE OF 65 YEARS. 39.7% OF PATIENTS WERE FEMALE. TYPE OF PROCEDURE: TREATMENT OF PERIPHERAL ARTERIAL DISEASE (PAD). BOSTON SCIENTIFIC PERFORMED A RETROSPECTIVE DATA ANALYSIS USING THE AGGREGATE ELECTRONIC HEALTH RECORD DATABASE FROM TRUVETA. THIS ANALYSIS AIMED TO EVALUATE ADVERSE EVENT RATES RELATED TO THE ELUVIA? DRUG-ELUTING VASCULAR STENT SYSTEM DEVICE IN REAL-WORLD USE. THE IMPLANT PROCEDURES WERE PERFORMED FROM FEBRUARY 2011 TO JUNE 2024. PATIENT EVENTS WERE IDENTIFIED AS EVENT TERMS WITH RATES. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT. DATA OBTAINED FROM TRUVETA FOR THIS STUDY IS DE-IDENTIFIED BEFORE BEING ACCESSED BY BOSTON SCIENTIFIC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO OUR ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE STUDY DATA DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. WITHIN THIS ANALYSIS, THE TOTAL PATIENT POPULATION SIZE WAS 440 INDIVIDUALS TREATED WITH ELUVIA AND 2785 INDIVIDUALS TREATED WITH BARE METAL STENTS.Event Description: THIS REPORT SUMMARIZES 25 REPORTED INCIDENTS FOR PATIENT DEATHS WITHIN 30 DAYS POST INDEX PROCEDURE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. THE ELUVIA DRUG-ELUTING VASCULAR STENT SYSTEM WAS SELECTED FOR USE. PATIENT DEATHS WERE REPORTED WITHIN 30 DAYS FOLLOWING THE INDEX PROCEDURE. 72.5% OF PATIENTS WERE OVER THE AGE OF 65 YEARS. 39.7% OF PATIENTS WERE FEMALE. TYPE OF PROCEDURE: TREATMENT OF PERIPHERAL ARTERIAL DISEASE (PAD). BOSTON SCIENTIFIC PERFORMED A RETROSPECTIVE DATA ANALYSIS USING THE AGGREGATE ELECTRONIC HEALTH RECORD DATABASE FROM TRUVETA. THIS ANALYSIS AIMED TO EVALUATE ADVERSE EVENT RATES RELATED TO THE ELUVIA? DRUG-ELUTING VASCULAR STENT SYSTEM DEVICE IN REAL-WORLD USE. THE IMPLANT PROCEDURES WERE PERFORMED FROM (B)(6) 2011 TO (B)(6) 2024. PATIENT EVENTS WERE IDENTIFIED AS EVENT TERMS WITH RATES. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT. DATA OBTAINED FROM TRUVETA FOR THIS STUDY IS DE-IDENTIFIED BEFORE BEING ACCESSED BY BOSTON SCIENTIFIC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO OUR ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE STUDY DATA DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. WITHIN THIS ANALYSIS, THE TOTAL PATIENT POPULATION SIZE WAS 440 INDIVIDUALS TREATED WITH ELUVIA AND 2785 INDIVIDUALS TREATED WITH BARE METAL STENTS.Event Description: Unique Complaint ID Number,Event Date,Date Entered,Manufacturer Aware Date,Brand Name,Generic Name,Model Number,Lot Number,Catalog Number,Serial Number,UDI Number,PMA / 510K Number,Date Returned to Manufacturer,Type of Reportable Event,Event Description,Manufacturer Narrative,Medical History,Patient Age,Patient Gender,Patient Weight,Date Implanted,Date Explanted,Health Effect Clinical Code,Health Effect Impact Code,Device Problem Code,Device Component Code,Investigation Type Code,Investigation Findings Code,Investigation Conclusion Code,Remedial Action Type,Latest Line Item Version
<br>19985499,,7/2/2025,4/4/2025,ELUVIATM Drug-Eluting Vascular Stent System,,,,,,,,,D,"This report summarizes 25 reported incidents for patient deaths within 30 days post index procedure. Note that multiple event terms may apply to a single patient but are captured separately under reporting guidelines. Device relationship to the events reported is not provided. <br><br>The ELUVIA Drug-Eluting Vascular Stent System was selected for use. Patient deaths were reported within 30 days following the index procedure. <br><br>72.5% of patients were over the age of 65 years. <br>39.7% of patients were female<br><br>Type of Procedure: Treatment of peripheral arterial disease (PAD).<br><br>Boston Scientific performed a retrospective data analysis using the aggregate electronic health record database from Truveta. This analysis aimed to evaluate adverse event rates related to the ELUVIA¿ Drug-Eluting Vascular Stent System device in real-world use. The implant procedures were performed from February 2011 to June 2024. Patient events were identified as event terms with rates. Multiple event terms may apply to a single patient. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC. Within this analysis, the total patient population size was 440 individuals treated with Eluvia and 2785 individuals treated with bare metal stents.","Time Frame of Event Date Collection: Patients in Truveta health systems who received an ELUVIA device between February 5, 2011 and June 7, 2024. <br>Summary of Adverse Events: 30 day mortality; 30 day, 90 day, 12 month and 24 month amputation (both all amputation and same limb amputation), and 12 and 24 month target vessel revascularization (TVR). <br>Total number of ELUVIA patients: 440<br>Device Evaluation: Under the terms and conditions of the study, anonymized data was provided. No products were returned as part of the study. It cannot be determined if these events have been previously reported or if the devices were returned for analysis as part of a previously reported spontaneous complaint. However, the events reported were anticipated in nature as defined by the ""Potential Adverse Events"" list in the FDA-approved Instructions for Use (IFU). <br>Contextual Analysis: <br>Limitations: this was an exploratory analysis of real world data using Truveta¿s electronic health records data including feasibility assessment of a natural language process model (NLPM) in early infancy. Only 10% of the healthcare systems providing data contributed to the NLPM reintervention data. Only patients with unique device identifier (UDI) codes were included in this analysis and therefore, it is likely that not all uses of ELUVIA devices are captured. While Truveta data as a whole is nationally representative when compared to the US Census, the population included for specific studies may not be representative. Further, data on patients who receive care outside of Truveta health systems would not be captured in Truveta data.<br>The 12-month freedom from amputation rate (same limb) within the Truveta analysis of 94.6% aligns with expectations derived from published literature which range from 92.9%-100% [1-3]. <br>Similarly, the 12-month TVR rate of 3.4% observed in this analysis is lower than rates reported in the literature (6.9%) [4]. The lower observed rates are due to the limited contributing data, missing and/or incomplete post index procedure intervention progress notes, and language learning model still in development. Importantly, TVR does not imply that the same lesion was treated during the reintervention. <br>The 30-day all-cause mortality rate of 5.7% is higher than rates reported in the literature (0-2.8%) [4-7], however, this is likely due to the exclusion of patients with a documented life expectancy of less than 12 months in addition to other inclusion/ exclusion criteria that may lead to a healthier population. <br>While health analytics platforms hold great promise for advancing healthcare, there are important limitations that must be considered when interpreting these data. Electronic health record data may be missing or inconsistent due to the varied sources, timeliness of data entry, methods of collection, differences in nomenclature between EHR platforms, and documentation. Coding practices can complicate the process of integrating and harmonizing data across multiple sources. Only a small percentage of health systems were actually providing UDI data at the time of this data query. Finally, we performed this analysis early in the realm of electronic health record data. Systems have improved, data has been enriched, submission of health system data has increased, and granularity of the data has evolved.<br>REFERENCES<br>1. Kim, J., et al., Korean Multicenter Registry of ELUVIA Stent for Femoropopliteal Artery Disease: K-ELUVIA Registry. Korean Circ J, 2024. 54(9): p. 565-576.<br>2. Wittig, T., et al., Randomized Trial Comparing a Stent-Avoiding With a Stent-Preferred Strategy in Complex Femoropopliteal Lesions. JACC Cardiovasc Interv, 2024. 17(9): p. 1134-1144.<br>3. Soga, Y., et al., High-Dose Drug-Coated Balloon Versus Polymer-Based Drug-Eluting Stent for Femoropopliteal Artery Disease Treatment. J Endovasc Ther, 2024: p. 15266028241267759.<br>4. Gray, W.A., et al., A polymer-coated, paclitaxel-eluting stent (Eluvia) versus a polymer-free, paclitaxel-coated stent (Zilver PTX) for endovascular femoropopliteal intervention (IMPERIAL): a randomised, non-inferiority trial. Lancet, 2018. 392(10157): p. 1541-1551.<br>5. Muller-Hulsbeck, S., et al., Twelve-Month Results From the MAJESTIC Trial of the Eluvia Paclitaxel-Eluting Stent for Treatment of Obstructive Femoropopliteal Disease. J Endovasc Ther, 2016. 23(5): p. 701-7.<br>6. Giannopoulos, S., et al., Concomitant Drug-Coated Balloon Angioplasty With Bail-Out Use of Eluvia Drug-Eluting Stent: Is There Any Downside to a Double Dose of Paclitaxel? J Invasive Cardiol, 2022. 34(6): p. E469-E476.<br>7. Lansink, W. Real world all comers data from the REGAL registry ¿ 2-year results [PowerPoint Slides] Leizpig Interventional Course 2023. 6 June 2023. LINC_2023_programme.pdf",peripheral arterial disease (PAD) ,72.5% of patients were over the age of 65 years. ,,,,,E2403,F02,A24,G07001,B17,C19,D12,No,1
<br>19985499,,7/2/2025,4/4/2025,ELUVIATM Drug-Eluting Vascular Stent System,,,,,,,,,D,"This report summarizes 25 reported incidents for patient deaths within 30 days post index procedure. Note that multiple event terms may apply to a single patient but are captured separately under reporting guidelines. Device relationship to the events reported is not provided. <br><br>The ELUVIA Drug-Eluting Vascular Stent System was selected for use. Patient deaths were reported within 30 days following the index procedure. <br><br>72.5% of patients were over the age of 65 years. <br>39.7% of patients were female<br><br>Type of Procedure: Treatment of peripheral arterial disease (PAD).<br><br>Boston Scientific performed a retrospective data analysis using the aggregate electronic health record database from Truveta. This analysis aimed to evaluate adverse event rates related to the ELUVIA¿ Drug-Eluting Vascular Stent System device in real-world use. The implant procedures were performed from February 2011 to June 2024. Patient events were identified as event terms with rates. Multiple event terms may apply to a single patient. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC. Within this analysis, the total patient population size was 440 individuals treated with Eluvia and 2785 individuals treated with bare metal stents.","Time Frame of Event Date Collection: Patients in Truveta health systems who received an ELUVIA device between February 5, 2011 and June 7, 2024. <br>Summary of Adverse Events: 30 day mortality; 30 day, 90 day, 12 month and 24 month amputation (both all amputation and same limb amputation), and 12 and 24 month target vessel revascularization (TVR). <br>Total number of ELUVIA patients: 440<br>Device Evaluation: Under the terms and conditions of the study, anonymized data was provided. No products were returned as part of the study. It cannot be determined if these events have been previously reported or if the devices were returned for analysis as part of a previously reported spontaneous complaint. However, the events reported were anticipated in nature as defined by the ""Potential Adverse Events"" list in the FDA-approved Instructions for Use (IFU). <br>Contextual Analysis: <br>Limitations: this was an exploratory analysis of real world data using Truveta¿s electronic health records data including feasibility assessment of a natural language process model (NLPM) in early infancy. Only 10% of the healthcare systems providing data contributed to the NLPM reintervention data. Only patients with unique device identifier (UDI) codes were included in this analysis and therefore, it is likely that not all uses of ELUVIA devices are captured. While Truveta data as a whole is nationally representative when compared to the US Census, the population included for specific studies may not be representative. Further, data on patients who receive care outside of Truveta health systems would not be captured in Truveta data.<br>The 12-month freedom from amputation rate (same limb) within the Truveta analysis of 94.6% aligns with expectations derived from published literature which range from 92.9%-100% [1-3]. <br>Similarly, the 12-month TVR rate of 3.4% observed in this analysis is lower than rates reported in the literature (6.9%) [4]. The lower observed rates are due to the limited contributing data, missing and/or incomplete post index procedure intervention progress notes, and language learning model still in development. Importantly, TVR does not imply that the same lesion was treated during the reintervention. <br>The 30-day all-cause mortality rate of 5.7% is higher than rates reported in the literature (0-2.8%) [4-7], however, this is likely due to the exclusion of patients with a documented life expectancy of less than 12 months in addition to other inclusion/ exclusion criteria that may lead to a healthier population. <br>While health analytics platforms hold great promise for advancing healthcare, there are important limitations that must be considered when interpreting these data. Electronic health record data may be missing or inconsistent due to the varied sources, timeliness of data entry, methods of collection, differences in nomenclature between EHR platforms, and documentation. Coding practices can complicate the process of integrating and harmonizing data across multiple sources. Only a small percentage of health systems were actually providing UDI data at the time of this data query. Finally, we performed this analysis early in the realm of electronic health record data. Systems have improved, data has been enriched, submission of health system data has increased, and granularity of the data has evolved.<br>REFERENCES<br>1. Kim, J., et al., Korean Multicenter Registry of ELUVIA Stent for Femoropopliteal Artery Disease: K-ELUVIA Registry. Korean Circ J, 2024. 54(9): p. 565-576.<br>2. Wittig, T., et al., Randomized Trial Comparing a Stent-Avoiding With a Stent-Preferred Strategy in Complex Femoropopliteal Lesions. JACC Cardiovasc Interv, 2024. 17(9): p. 1134-1144.<br>3. Soga, Y., et al., High-Dose Drug-Coated Balloon Versus Polymer-Based Drug-Eluting Stent for Femoropopliteal Artery Disease Treatment. J Endovasc Ther, 2024: p. 15266028241267759.<br>4. Gray, W.A., et al., A polymer-coated, paclitaxel-eluting stent (Eluvia) versus a polymer-free, paclitaxel-coated stent (Zilver PTX) for endovascular femoropopliteal intervention (IMPERIAL): a randomised, non-inferiority trial. Lancet, 2018. 392(10157): p. 1541-1551.<br>5. Muller-Hulsbeck, S., et al., Twelve-Month Results From the MAJESTIC Trial of the Eluvia Paclitaxel-Eluting Stent for Treatment of Obstructive Femoropopliteal Disease. J Endovasc Ther, 2016. 23(5): p. 701-7.<br>6. Giannopoulos, S., et al., Concomitant Drug-Coated Balloon Angioplasty With Bail-Out Use of Eluvia Drug-Eluting Stent: Is There Any Downside to a Double Dose of Paclitaxel? J Invasive Cardiol, 2022. 34(6): p. E469-E476.<br>7. Lansink, W. Real world all comers data from the REGAL registry ¿ 2-year results [PowerPoint Slides] Leizpig Interventional Course 2023. 6 June 2023. LINC_2023_programme.pdf",peripheral arterial disease (PAD) ,72.5% of patients were over the age of 65 years. ,,,,,E2403,F02,A24,G07001,B17,C19,D12,No,1
<br>19985499,,7/2/2025,4/4/2025,ELUVIATM Drug-Eluting Vascular Stent System,,,,,,,,,D,"This report summarizes 25 reported incidents for patient deaths within 30 days post index procedure. Note that multiple event terms may apply to a single patient but are captured separately under reporting guidelines. Device relationship to the events reported is not provided. <br><br>The ELUVIA Drug-Eluting Vascular Stent System was selected for use. Patient deaths were reported within 30 days following the index procedure. <br><br>72.5% of patients were over the age of 65 years. <br>39.7% of patients were female<br><br>Type of Procedure: Treatment of peripheral arterial disease (PAD).<br><br>Boston Scientific performed a retrospective data analysis using the aggregate electronic health record database from Truveta. This analysis aimed to evaluate adverse event rates related to the ELUVIA¿ Drug-Eluting Vascular Stent System device in real-world use. The implant procedures were performed from February 2011 to June 2024. Patient events were identified as event terms with rates. Multiple event terms may apply to a single patient. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC. Within this analysis, the total patient population size was 440 individuals treated with Eluvia and 2785 individuals treated with bare metal stents.","Time Frame of Event Date Collection: Patients in Truveta health systems who received an ELUVIA device between February 5, 2011 and June 7, 2024. <br>Summary of Adverse Events: 30 day mortality; 30 day, 90 day, 12 month and 24 month amputation (both all amputation and same limb amputation), and 12 and 24 month target vessel revascularization (TVR). <br>Total number of ELUVIA patients: 440<br>Device Evaluation: Under the terms and conditions of the study, anonymized data was provided. No products were returned as part of the study. It cannot be determined if these events have been previously reported or if the devices were returned for analysis as part of a previously reported spontaneous complaint. However, the events reported were anticipated in nature as defined by the ""Potential Adverse Events"" list in the FDA-approved Instructions for Use (IFU). <br>Contextual Analysis: <br>Limitations: this was an exploratory analysis of real world data using Truveta¿s electronic health records data including feasibility assessment of a natural language process model (NLPM) in early infancy. Only 10% of the healthcare systems providing data contributed to the NLPM reintervention data. Only patients with unique device identifier (UDI) codes were included in this analysis and therefore, it is likely that not all uses of ELUVIA devices are captured. While Truveta data as a whole is nationally representative when compared to the US Census, the population included for specific studies may not be representative. Further, data on patients who receive care outside of Truveta health systems would not be captured in Truveta data.<br>The 12-month freedom from amputation rate (same limb) within the Truveta analysis of 94.6% aligns with expectations derived from published literature which range from 92.9%-100% [1-3]. <br>Similarly, the 12-month TVR rate of 3.4% observed in this analysis is lower than rates reported in the literature (6.9%) [4]. The lower observed rates are due to the limited contributing data, missing and/or incomplete post index procedure intervention progress notes, and language learning model still in development. Importantly, TVR does not imply that the same lesion was treated during the reintervention. <br>The 30-day all-cause mortality rate of 5.7% is higher than rates reported in the literature (0-2.8%) [4-7], however, this is likely due to the exclusion of patients with a documented life expectancy of less than 12 months in addition to other inclusion/ exclusion criteria that may lead to a healthier population. <br>While health analytics platforms hold great promise for advancing healthcare, there are important limitations that must be considered when interpreting these data. Electronic health record data may be missing or inconsistent due to the varied sources, timeliness of data entry, methods of collection, differences in nomenclature between EHR platforms, and documentation. Coding practices can complicate the process of integrating and harmonizing data across multiple sources. Only a small percentage of health systems were actually providing UDI data at the time of this data query. Finally, we performed this analysis early in the realm of electronic health record data. Systems have improved, data has been enriched, submission of health system data has increased, and granularity of the data has evolved.<br>REFERENCES<br>1. Kim, J., et al., Korean Multicenter Registry of ELUVIA Stent for Femoropopliteal Artery Disease: K-ELUVIA Registry. Korean Circ J, 2024. 54(9): p. 565-576.<br>2. Wittig, T., et al., Randomized Trial Comparing a Stent-Avoiding With a Stent-Preferred Strategy in Complex Femoropopliteal Lesions. JACC Cardiovasc Interv, 2024. 17(9): p. 1134-1144.<br>3. Soga, Y., et al., High-Dose Drug-Coated Balloon Versus Polymer-Based Drug-Eluting Stent for Femoropopliteal Artery Disease Treatment. J Endovasc Ther, 2024: p. 15266028241267759.<br>4. Gray, W.A., et al., A polymer-coated, paclitaxel-eluting stent (Eluvia) versus a polymer-free, paclitaxel-coated stent (Zilver PTX) for endovascular femoropopliteal intervention (IMPERIAL): a randomised, non-inferiority trial. Lancet, 2018. 392(10157): p. 1541-1551.<br>5. Muller-Hulsbeck, S., et al., Twelve-Month Results From the MAJESTIC Trial of the Eluvia Paclitaxel-Eluting Stent for Treatment of Obstructive Femoropopliteal Disease. J Endovasc Ther, 2016. 23(5): p. 701-7.<br>6. Giannopoulos, S., et al., Concomitant Drug-Coated Balloon Angioplasty With Bail-Out Use of Eluvia Drug-Eluting Stent: Is There Any Downside to a Double Dose of Paclitaxel? J Invasive Cardiol, 2022. 34(6): p. E469-E476.<br>7. Lansink, W. Real world all comers data from the REGAL registry ¿ 2-year results [PowerPoint Slides] Leizpig Interventional Course 2023. 6 June 2023. LINC_2023_programme.pdf",peripheral arterial disease (PAD) ,72.5% of patients were over the age of 65 years. ,,,,,E2403,F02,A24,G07001,B17,C19,D12,No,1
<br>19985499,,7/2/2025,4/4/2025,ELUVIATM Drug-Eluting Vascular Stent System,,,,,,,,,D,"This report summarizes 25 reported incidents for patient deaths within 30 days post index procedure. Note that multiple event terms may apply to a single patient but are captured separately under reporting guidelines. Device relationship to the events reported is not provided. <br><br>The ELUVIA Drug-Eluting Vascular Stent System was selected for use. Patient deaths were reported within 30 days following the index procedure. <br><br>72.5% of patients were over the age of 65 years. <br>39.7% of patients were female<br><br>Type of Procedure: Treatment of peripheral arterial disease (PAD).<br><br>Boston Scientific performed a retrospective data analysis using the aggregate electronic health record database from Truveta. This analysis aimed to evaluate adverse event rates related to the ELUVIA¿ Drug-Eluting Vascular Stent System device in real-world use. The implant procedures were performed from February 2011 to June 2024. Patient events were identified as event terms with rates. Multiple event terms may apply to a single patient. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC. Within this analysis, the total patient population size was 440 individuals treated with Eluvia and 2785 individuals treated with bare metal stents.","Time Frame of Event Date Collection: Patients in Truveta health systems who received an ELUVIA device between February 5, 2011 and June 7, 2024. <br>Summary of Adverse Events: 30 day mortality; 30 day, 90 day, 12 month and 24 month amputation (both all amputation and same limb amputation), and 12 and 24 month target vessel revascularization (TVR). <br>Total number of ELUVIA patients: 440<br>Device Evaluation: Under the terms and conditions of the study, anonymized data was provided. No products were returned as part of the study. It cannot be determined if these events have been previously reported or if the devices were returned for analysis as part of a previously reported spontaneous complaint. However, the events reported were anticipated in nature as defined by the ""Potential Adverse Events"" list in the FDA-approved Instructions for Use (IFU). <br>Contextual Analysis: <br>Limitations: this was an exploratory analysis of real world data using Truveta¿s electronic health records data including feasibility assessment of a natural language process model (NLPM) in early infancy. Only 10% of the healthcare systems providing data contributed to the NLPM reintervention data. Only patients with unique device identifier (UDI) codes were included in this analysis and therefore, it is likely that not all uses of ELUVIA devices are captured. While Truveta data as a whole is nationally representative when compared to the US Census, the population included for specific studies may not be representative. Further, data on patients who receive care outside of Truveta health systems would not be captured in Truveta data.<br>The 12-month freedom from amputation rate (same limb) within the Truveta analysis of 94.6% aligns with expectations derived from published literature which range from 92.9%-100% [1-3]. <br>Similarly, the 12-month TVR rate of 3.4% observed in this analysis is lower than rates reported in the literature (6.9%) [4]. The lower observed rates are due to the limited contributing data, missing and/or incomplete post index procedure intervention progress notes, and language learning model still in development. Importantly, TVR does not imply that the same lesion was treated during the reintervention. <br>The 30-day all-cause mortality rate of 5.7% is higher than rates reported in the literature (0-2.8%) [4-7], however, this is likely due to the exclusion of patients with a documented life expectancy of less than 12 months in addition to other inclusion/ exclusion criteria that may lead to a healthier population. <br>While health analytics platforms hold great promise for advancing healthcare, there are important limitations that must be considered when interpreting these data. Electronic health record data may be missing or inconsistent due to the varied sources, timeliness of data entry, methods of collection, differences in nomenclature between EHR platforms, and documentation. Coding practices can complicate the process of integrating and harmonizing data across multiple sources. Only a small percentage of health systems were actually providing UDI data at the time of this data query. Finally, we performed this analysis early in the realm of electronic health record data. Systems have improved, data has been enriched, submission of health system data has increased, and granularity of the data has evolved.<br>REFERENCES<br>1. Kim, J., et al., Korean Multicenter Registry of ELUVIA Stent for Femoropopliteal Artery Disease: K-ELUVIA Registry. Korean Circ J, 2024. 54(9): p. 565-576.<br>2. Wittig, T., et al., Randomized Trial Comparing a Stent-Avoiding With a Stent-Preferred Strategy in Complex Femoropopliteal Lesions. JACC Cardiovasc Interv, 2024. 17(9): p. 1134-1144.<br>3. Soga, Y., et al., High-Dose Drug-Coated Balloon Versus Polymer-Based Drug-Eluting Stent for Femoropopliteal Artery Disease Treatment. J Endovasc Ther, 2024: p. 15266028241267759.<br>4. Gray, W.A., et al., A polymer-coated, paclitaxel-eluting stent (Eluvia) versus a polymer-free, paclitaxel-coated stent (Zilver PTX) for endovascular femoropopliteal intervention (IMPERIAL): a randomised, non-inferiority trial. Lancet, 2018. 392(10157): p. 1541-1551.<br>5. Muller-Hulsbeck, S., et al., Twelve-Month Results From the MAJESTIC Trial of the Eluvia Paclitaxel-Eluting Stent for Treatment of Obstructive Femoropopliteal Disease. J Endovasc Ther, 2016. 23(5): p. 701-7.<br>6. Giannopoulos, S., et al., Concomitant Drug-Coated Balloon Angioplasty With Bail-Out Use of Eluvia Drug-Eluting Stent: Is There Any Downside to a Double Dose of Paclitaxel? J Invasive Cardiol, 2022. 34(6): p. E469-E476.<br>7. Lansink, W. Real world all comers data from the REGAL registry ¿ 2-year results [PowerPoint Slides] Leizpig Interventional Course 2023. 6 June 2023. LINC_2023_programme.pdf",peripheral arterial disease (PAD) ,72.5% of patients were over the age of 65 years. ,,,,,E2403,F02,A24,G07001,B17,C19,D12,No,1
<br>19985499,,7/2/2025,4/4/2025,ELUVIATM Drug-Eluting Vascular Stent System,,,,,,,,,D,"This report summarizes 25 reported incidents for patient deaths within 30 days post index procedure. Note that multiple event terms may apply to a single patient but are captured separately under reporting guidelines. Device relationship to the events reported is not provided. <br><br>The ELUVIA Drug-Eluting Vascular Stent System was selected for use. Patient deaths were reported within 30 days following the index procedure. <br><br>72.5% of patients were over the age of 65 years. <br>39.7% of patients were female<br><br>Type of Procedure: Treatment of peripheral arterial disease (PAD).<br><br>Boston Scientific performed a retrospective data analysis using the aggregate electronic health record database from Truveta. This analysis aimed to evaluate adverse event rates related to the ELUVIA¿ Drug-Eluting Vascular Stent System device in real-world use. The implant procedures were performed from February 2011 to June 2024. Patient events were identified as event terms with rates. Multiple event terms may apply to a single patient. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC. Within this analysis, the total patient population size was 440 individuals treated with Eluvia and 2785 individuals treated with bare metal stents.","Time Frame of Event Date Collection: Patients in Truveta health systems who received an ELUVIA device between February 5, 2011 and June 7, 2024. <br>Summary of Adverse Events: 30 day mortality; 30 day, 90 day, 12 month and 24 month amputation (both all amputation and same limb amputation), and 12 and 24 month target vessel revascularization (TVR). <br>Total number of ELUVIA patients: 440<br>Device Evaluation: Under the terms and conditions of the study, anonymized data was provided. No products were returned as part of the study. It cannot be determined if these events have been previously reported or if the devices were returned for analysis as part of a previousl... Edited for length. See MAUDE Adverse Event Report for complete event narratives.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22387546&pc=NPT | 2015691-2025-05396 | 2025-06-10 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-02 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem |  | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES AFFILIATE IN ITALY, THE PATIENT UNDERWENT A TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WITH A 20 MM SAPIEN 3 ULTRA VALVE. AFTER VALVE DEPLOYMENT, THE PATIENT PASSED AWAY DUE TO AN ACUTE CORONARY OCCLUSION. THE IMPLANT DATE, EVENT DATE, AND DATE OF DEATH ARE UNKNOWN AS NO ADDITIONAL DETAILS HAVE BEEN PROVIDED. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ALTHOUGH A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED, THE EVENT MAY BE RELATED TO THE MECHANISMS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR CORRECTION AND INCLUDES ADDITIONAL INFORMATION PROVIDED AND FROM THE UPDATED INVESTIGATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN CORRECTED/UPDATED: B2 (OUTCOME ATTRIBUTED TO ADVERSE EVENT), B3 (DATE OF EVENT), B7 (OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS), D6 (IMPLANT DATE), H6 (INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS) AND H11 (ADDITIONAL NARRATIVE/DATA). ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT ATTEMPTS WERE MADE TO PERFORM PERCUTANEOUS CORONARY INTERVENTION (PCI); HOWEVER, CORONARY CANNULATION WAS UNSUCCESSFUL DUE TO CALCIUM. IT WAS BELIEVED THAT CALCIFICATION LOCATED OVER THE LEFT CORONARY SINUS OCCLUDED THE LEFT CORONARY ARTERY (LCA). PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT FACTORS (CALCIFICATION) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22387581&pc=NPT | 2015691-2025-05395 | 2025-05-09 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-02 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEATH VALVE | Insufficient Information | Insufficient Information | P140031 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN FRANCE. AS REPORTED, THIS WAS A CASE OF AN IMPLANT OF A 29MM SAPIEN 3 IN THE MITRAL POSITION BY TRANSEPTAL APPROACH. THE PATIENT WAS VERY FRAGILE AND HAD A MITRAL VALVE AREA OF AROUND 850MM. THE FIRST VALVE IMPLANTED RESULTED IN A TOO ATRIAL POSITION. THEREFORE, IT WAS DECIDED TO IMPLANT A SECOND VALVE. THE SECOND VALVE WAS NOT WORKING PROPERLY ACCORDING TO IMAGES, AND IT WAS DECIDED TO IMPLANT A THIRD VALVE. UNFORTUNATELY, FEW HOURS AFTER PROCEDURE THE PATIENT PASSED AWAY. THE ROOT CAUSE OF THE EVENT MAY HAVE BEEN DUE TO THE LARGE SIZE OF THE MITRAL VALVE AROUND 850+MM2. THE PATIENT WAS VERY FRAGILE, AND THIS WAS A ''RESCUE PROCEDURE'' FOR PATIENT. Manufacturer Narrative: THE DEVICE WAS USED IN OFF-LABEL IMPLANTATION IN THE NATIVE MITRAL POSITION. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO NATIVE MITRAL PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE. PER THE INSTRUCTIONS FOR USE (IFU), NON-EMERGENT REOPERATION, EMERGENCY CARDIAC SURGERY, REOPERATION, DEVICE MIGRATION OR MALPOSITION REQUIRING INTERVENTION AND DEVICE EXPLANT ARE LISTED AS POTENTIAL RISKS ASSOCIATED WITH THE DEVICE AND TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT LONG-TERM DURABILITY HAS NOT BEEN ESTABLISHED FOR THE VALVE. REGULAR MEDICAL FOLLOW-UP IS ADVISED TO EVALUATE VALVE PERFORMANCE. ACCELERATED DETERIORATION OF THE VALVE DUE TO CALCIFIC DEGENERATION MAY OCCUR IN CHILDREN, ADOLESCENTS, OR YOUNG ADULTS AND IN PATIENTS WITH AN ALTERED CALCIUM METABOLISM. VALVE RECIPIENTS SHOULD BE MAINTAINED ON ANTICOAGULANT/ANTIPLATELET THERAPY, EXCEPT WHEN CONTRAINDICATED, AS DETERMINED BY THEIR PHYSICIAN. THIS DEVICE HAS NOT BEEN TESTED FOR USE WITHOUT ANTICOAGULATION. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. VALVE REINTERVENTION MAY BE DUE TO VALVE DYSFUNCTION (E.G., SEVERE, OR CLINICALLY SIGNIFICANT PARAVALVULAR LEAK, CENTRAL LEAK, SIGNIFICANT MALPOSITION, EARLY/LATE ENDOCARDITIS, THROMBOSIS, STRUCTURAL VALVE DETERIORATION OR DEGENERATION) AND/OR OTHER REASONS UNRELATED TO THE EDWARDS DEVICES (E.G., TREATMENT OF OTHER PATIENT COMORBIDITIES). DURING THE MANUFACTURING PROCESS, ALL SAPIEN VALVES (ALL MODELS) ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. IN THIS CASE, DESPITE MULTIPLE INVESTIGATIONAL ATTEMPTS, IT WAS NOT POSSIBLE TO OBTAIN ADDITIONAL DETAILS REGARDING THIS EVENT. DUE TO LIMITED INFORMATION, A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AT THIS TIME. HOWEVER, IT IS LIKELY THAT PATIENT AND PROCEDURE RELATED FACTORS (LARGE ANNULAR AREA AND OFF LABEL NATIVE MITRAL POSITION) CAUSED OR CONTRIBUTED TO THIS EVENT. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE INFORMATION WILL BE REVIEWED, AND THE FILE UPDATED ACCORDINGLY. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22387598&pc=LWS | 3008973940-2025-04764 | 2025-07-01 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-07-02 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Signal Artifact/Noise; Fracture; Over-Sensing; Pacing Problem | Non specific EKG/ECG Changes; Dizziness; Pericardial Effusion; Presyncope | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) DUE TO HIGH RATE NON-SUSTAINED EPISODES AND SENSING INTEGRITY COUNTER (SIC). IN ADDITION, THE LEAD EXHIBITED NOISE AND T-WAVE OVERSENSING (TWOS) WHICH CAUSED LOSS OF PACING WITH SIGNIFICANT PAUSES UP TO FIVE SECONDS. THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS LIGHT-HEADED DUE TO THE PACING INHIBITION. IT WAS NOTED THE DEVICE WAS REPROGRAMMED TO PACE ONLY OFF OF THE LEFT VENTRICULAR (LV) LEAD PENDING THE RIGHT VENTRICULAR (RV) LEAD EXTRACTION PROCEDURE. IT WAS REPORTED THAT THE RV LEAD EXTRACTION WAS ATTEMPTED, BUT COMPLICATIONS OF PERICARDIAL EFFUSION OCCURRED DURING THE PROCEDURE. IT WAS NOTED THE PATIENT DIED AS A RESULT OF THE COMPLICATION.Event Description: IT WAS FURTHER REPORTED THAT THE RV LEAD EXHIBITED OVERSENSING AND PAUSES WITH ISOMETRIC MOVEMENTS. THE LEAD HAD A POSSIBLE FRACTURE. THE PATIENT EXHIBITED PRESYNCOPE AND WAS HOSPITALIZED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE RV LEAD EXHIBITED OVERSENSING AND PAUSES WITH ISOMETRIC MOVEMENTS. THE PATIENT EXHIBITED PRESYNCOPE AND WAS HOSPITALIZED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED AND ANALYZED. THE ANALYSIS INDICATED THAT THE LV1/DISTAL CONDUCTOR WAS FRACTURED IN-VIVO IN THE ANCHORING SLEEVE AREA. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: A4, B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 1688T LEAD, 1056K LEAD, IMPLANTED ON (B)(6) 2007. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22390751&pc=BZD | 2518422-2025-107047 | 2025-06-17 04:00 | Death | RESPIRONICS, INC. | 2025-07-02 04:00 | BZD | REP DREAMSTATION AUTO BIPAP, DOM-RECRT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER INQUIRING ABOUT A DREAMSTATION DEVICE RECEIVED FROM THEIR DME. THE MANUFACTURER WAS INFORMED THAT THE PREVIOUS USER OF THE DEVICE WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22390988&pc=NPT | 2015691-2025-05405 | 2025-06-09 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-02 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem |  | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES JAPAN AFFILIATE, THE PATIENT UNDERWENT A TRANSFEMORAL TAVR (TRANSCATHETER AORTIC VALVE REPLACEMENT) PROCEDURE WITH A 26MM SAPIEN 3 ULTRA (S3U) VALVE IN THE NATIVE AORTIC POSITION. POST S3U VALVE DEPLOYMENT, PROLONGED HYPOTENSION WAS OBSERVED. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) DID NOT CONFIRM EFFUSION OR DISSECTION, BUT REDUCED WALL MOTION WAS NOTED. A LEFT MAIN TRUNK (LMT) OBSTRUCTION WAS SUSPECTED DUE TO THE NATIVE LEAFLET BEING LOCATED NEAR THE UPWARD LMT; PER MEDICAL OPINION, THE S3U VALVE LIKELY PUSHED THE SINOTUBULAR JUNCTION (STJ) AND CAUSED THE OBSTRUCTION. A CARDIAC MASSAGE WAS INITIATED, AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) OF THE LEFT FEMORAL VENOUS DRAINAGE/RIGHT FEMORAL ARTERIAL RETURN WAS ESTABLISHED SIMULTANEOUSLY. AORTOGRAPHY CONFIRMED ZERO FLOW IN THE LMT. JUST AS A PERCUTANEOUS CORONARY INTERVENTION (PCI) WAS ABOUT TO BE EXECUTED, TEE REVEALED A DISSECTION AT THE ASCENDING AORTA, WHICH ALSO EXTENDED TO THE DESCENDING AORTA; PER MEDICAL OPINION, THE ECMO ACCESS CAUSED THE RETROGRADE DISSECTION. A THORACOTOMY WAS EXECUTED ALONG WITH SURGICAL AORTIC VALVE REPLACEMENT AND ASCENDING AORTA REPLACEMENT. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). APPROXIMATELY 1 HOUR LATER, THE PATIENT PASSED AWAY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PROCEDURAL FACTORS LIKELY CONTRIBUTED TO THE CORONARY ARTERY OBSTRUCTION, INCLUDING THE RETROGRADE DISSECTION. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22391440&pc=PNJ | 2017865-2025-91331 | 2025-06-13 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-02 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY SEVERAL DAYS POST-IMPLANT PROCEDURE. THE CAUSE OF DEATH WAS UNKNOWN; HOWEVER, THE PHYSICIAN DID NOT ALLEGE THAT THE PATIENT'S DEATH WAS RELATED TO THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22391538&pc=GCJ | 2955842-2025-28782 | 2025-06-11 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-07-02 04:00 | GCJ | WECK VISTA | Insufficient Information | Unspecified Tissue Injury; Insufficient Information | K152663 |  | 1 | Event Description: IT WAS REPORTED THAT DURING INITIAL PORT PLACEMENT FOR A DA VINCI-ASSISTED OVARIAN CYSTECTOMY SURGICAL PROCEDURE, AN INJURY OCCURRED TO AN UNSPECIFIED VITAL STRUCTURE LEADING TO THE PATIENT'S DEATH. THE INTUITIVE CLINICAL SALES REPRESENTATIVE WAS INFORMED BY OR STAFF THAT AFTER PORT PLACEMENT, A DA VINCI ENDOSCOPE WAS INTRODUCED AND THE INJURY IN THE PATIENT'S ABDOMEN WAS OBSERVED. THE SYSTEM HAD BEEN DOCKED FOR THREE MINUTES PRIOR TO THIS OBSERVATION. THE INJURY REQUIRED CONVERSION TO AN OPEN PROCEDURE. IT WAS REPORTED THAT TRAUMA AND VASCULAR SURGEONS WERE ENLISTED TO ASSIST WITH THE OPEN PROCEDURE; HOWEVER, NO SPECIFIC INFORMATION REGARDING THE SURGICAL INTERVENTION WAS PROVIDED. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) WHERE AND ULTIMATELY EXPIRED ON AN UNSPECIFIED DATE. THERE WAS NO ALLEGATION THAT A MALFUNCTION OF A DA VINCI SYSTEM, INSTRUMENT OR ACCESSORY OCCURRED. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION HAVE BEEN MADE; HOWEVER, NO FURTHER INFORMATION HAS BEEN PROVIDED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: THE FOLLOWING PRODUCTS WERE PREPPED FOR USE AND INSTALLED ON THE ROBOT: 0 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, PROGRASP FORCEPS, FENESTRATED BIPOLAR FORCEPS. HOWEVER, ONLY THE MONOPOLAR CURVED SCISSORS AND THE FENESTRATED BIPOLAR FORCEPS INSTRUMENTS HAD LOGGED TIME IN USE OF 3 MINUTES AND 2 MINUTES RESPECTIVELY. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AS A RESULT OF A PORT PLACEMENT INJURY AT THE BEGINNING OF THE CASE. THE EXACT TYPE OF INJURY IS NOT PROVIDED. ALTHOUGH PORT PLACEMENT INJURIES ARE OFTEN SURGICAL TECHNICAL ERRORS AND NO ALLEGATION HAS BEEN MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT, FURTHER DETAILS OF THE EVENT ARE NOT PROVIDED. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION B2 PATIENT DATE OF DEATH: THE DATE OF DEATH PROVIDED IS ESTIMATED. THE EVENT OCCURRED ON (B)(6) 2025, AND THE INTUITIVE CSR WAS NOTIFIED ON (B)(6) 2025 THAT THE PATIENT HAD EXPIRED "AT THE END OF LAST WEEK." Manufacturer Narrative: ADDITIONAL INFORMATION SECTION B5: THE SURGEON REPORTED THAT THIS EVENT WAS UNRELATED TO THE DA VINCI SYSTEM OR INSTRUMENTS, AND WAS ADVISED BY THE HOSPITAL ADMINISTRATION TO NOT DISCLOSE ALL THE DETAILS OF THE EVENT AS IT WAS NOT ROBOTIC RELATED. SECTION H11: A REVIEW OF THE SYSTEM LOGS FOUND THAT NO ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE FIVE SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM SHOWED THAT THE SYSTEM FUNCTIONED NORMALLY, NO INTRAOPERATIVE EVENTS OR ISSUES WERE FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22391750&pc=GCJ | 2955842-2025-28786 | 2025-06-04 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-07-02 04:00 | GCJ | WECK VISTA | Insufficient Information | Hemorrhage/Bleeding; Great Vessel Perforation | K152663 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED OVARIAN CYSTECTOMY AND ENDOMETRIOSIS RESECTION PROCEDURE, THE PATIENT'S AORTA WAS INADVERTENTLY PERFORATED DURING THE INITIAL TROCAR INSERTION, PRIOR TO DOCKING THE ROBOTIC SYSTEM. ACCESS INTO THE PATIENT'S ABDOMEN WAS PERFORMED USING A THIRD-PARTY 5MM LAPAROSCOPIC ENDOSCOPE AND A DA VINCI 8MM CANNULA WITH AN 8MM BLADELESS OPTICAL OBTURATOR. BLEEDING WAS IDENTIFIED WITHIN MINUTES, AND THE PROCEDURE WAS CONVERTED TO LAPAROTOMY TO ATTEMPT HEMOSTASIS. THE PATIENT WAS PLACED ON A MASS TRANSFUSION PROTOCOL. A VASCULAR SURGEON WAS CALLED TO ASSIST; HOWEVER, THE PATIENT EXPIRED ON THE OPERATING TABLE BEFORE THE VASCULAR SURGEON WAS ABLE TO INTERVENE. Manufacturer Narrative: THE REPORTED EVENT OCCURRED PRIOR TO DOCKING THE DA VINCI SYSTEM; THEREFORE, NO DA VINCI SYSTEM OR INSTRUMENT LOGS ARE AVAILABLE FOR REVIEW. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AS A RESULT OF A TROCAR INJURY TO THE AORTA WHILE USING A BLADELESS OPTICAL OBTURATOR AND A 3RD PARTY SCOPE. ALTHOUGH NO ALLEGATION IS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT, HOW THE INJURY OCCURRED IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22368661&pc=OZP | 2032227-2025-201157 | 2025-05-25 04:00 | Death | MEDTRONIC MINIMED | 2025-07-01 04:00 | OZP | UNKNOWN NGP | Adverse Event Without Identified Device or Use Problem | Hyperglycemia |  |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025 AT HOME. THE CAUSE OF DEATH WAS HYPERGLYCEMIA. BLOOD GLUCOSE VALUE WAS 500 MG/DL. THE INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: UNOMEDICAL, UNK_NGP. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING. IT WAS UNKNOWN WHETHER THE CUSTOMER USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT, AND IT WAS UNKNOWN WHETHER THE SMARTGUARD/AUTO MODE OF THE INSULIN PUMP WAS USED OR NOT. UNOMEDICAL AND UNK-NGP WERE REQUESTED, AND CUSTOMER RESPONSE WAS THE DEVICE'S WILL BE RETURNED. Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22369842&pc=FNL | 1419652-2025-00033 | 2025-04-04 04:00 | Death | ARJOHUNTLEIGH POLSKA SP. Z O.O. | 2025-07-01 04:00 | FNL | MINUET 2 | Entrapment of Device | Strangulation |  |  | 1 | Event Description: THE PATIENT WAS TRYING TO GET OUT OF BED AND BECAME ENTRAPPED BETWEEN THE HOSPITAL BED AND BED LEVER, IT IS REPORTED AT THIS TIME RESULTED IN SUFFOCATION, LEADING TO DEATH. ON (B)(6) 2025 ARJO WAS INFORMED THAT THE AFFECTED ARJO BED USED FOR AT THAT TIME WAS A MINUET 2 BED; THE DEVICE WAS ISOLATED BY A CUSTOMER. CURRENTLY, NO DEVICE EVALUATION WAS PERFORMED; THE CUSTOMER DID NOT PROVIDE A DEVICE SERIAL NUMBER OR LOCATION DESPITE ATTEMPTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22369843&pc=FNL | 3007420694-2025-00157 | 2025-04-04 04:00 | Death | ARJOHUNTLEIGH POLSKA SP. Z O.O. | 2025-07-01 04:00 | FNL | MINUET 2 | Entrapment of Device | Strangulation |  |  | 1 | Event Description: THE PATIENT WAS TRYING TO GET OUT OF BED AND BECAME ENTRAPPED BETWEEN THE HOSPITAL BED AND BED LEVER, IT IS REPORTED AT THIS TIME RESULTED IN SUFFOCATION, LEADING TO DEATH. ON 04 JUN 2025 ARJO WAS INFORMED THAT THE AFFECTED ARJO BED USED FOR AT THAT TIME WAS A MINUET 2 BED; THE DEVICE WAS ISOLATED BY A CUSTOMER. CURRENTLY, NO DEVICE EVALUATION WAS PERFORMED; THE CUSTOMER DID NOT PROVIDE A DEVICE SERIAL NUMBER OR LOCATION DESPITE ATTEMPTS. Manufacturer Narrative: THE INVESTIGATION IS STILL ONGOING. FURTHER INFORMATION WILL BE PROVIDED IN THE FINAL REPORT. Manufacturer Narrative: THE INVESTIGATION IS ONGOING. FURTHER INFORMATION WILL BE PROVIDED IN THE FINAL REPORT. H3 OTHER TEXT: CUSTOMER DID NOT PROVIDE ANY DEVICE DETAIL SUCH SERIAL NUMBER, LOCATION OF THE DEVICE, THE EVALUATION WAS NOT PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22370180&pc=DSQ | 2916596-2025-04095 | 2025-05-12 04:00 | Death | THORATEC CORPORATION | 2025-07-01 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Dyspnea; Fatigue; Renal Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DEALT WITH DEPRESSION SINCE THEIR SPOUSE PASSED AWAY. THE PATIENT NOTIFIED THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) TEAM THAT THEY NO LONGER WANTED TO HEAR FROM HEART FAILURE (HF)/VENTRICULAR ASSIST DEVICE (VAD) COORDINATORS IN (B)(6) 2025. THE PATIENT HAD REFUSED ALL CONTACT AND REPORTEDLY STOPPED TAKING ALL MEDICATIONS AT THAT TIME. THEY RE-PRESENTED TO THE HOSPITAL ON (B)(6) 2025 WITH COMPLAINS OF SHORTNESS OF BREATH (SOB) AND FATIGUE. LABS AT THAT TIME WAS NOTABLE FOR CREATININE LEVEL OF 18, BLOOD UREA NITROGEN (BUN) LEVEL OF 163, AND BRAIN NATRIURETIC PEPTIDE (BNP) LEVEL OF 64,000. THE PATIENT SUBSEQUENTLY REQUIRED HEMODIALYSIS (HD). THE PATIENT REQUESTED TO BE TRANSITIONED TO HOSPICE BECAUSE THEY DID NOT WANT TO CONTINUE DIALYSIS IN AN OUTPATIENT SETTING AND WANTED TO RETURN TO HOME DUE TO POOR QUALITY OF LIFE. THE RENAL DYSFUNCTION WAS NOT CONSIDERED TO BE DEVICE RELATED. THE KIDNEY FAILURE WAS DUE TO THE PATIENT SELF-STOPPING ALL MEDICATIONS AND WANTING TO MOVE TOWARD HOSPICE. THE PATIENT EVENTUALLY PASSED AWAY ON (B)(6) 2025. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE HAD OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RENAL DYSFUNCTION, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22370501&pc=FRN | 2016493-2025-91927 | 2025-05-13 04:00 | Death | CAREFUSION SD | 2025-07-01 04:00 | FRN | ALARIS SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K133532 |  | 1 | Event Description: IT WAS REPORTED THE CUSTOMER IS REQUESTING A LOG REVIEW FOLLOWING AN EVENT WHERE THE PATIENT HAS DIED. ALTHOUGH DUE DILIGENCE ATTEMPTS HAVE BEEN MADE, AT THIS TIME, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: THE ACTUAL DATE OF DEATH IS UNKNOWN. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22370631&pc=BZD | 2518422-2025-047820 | 2021-10-18 04:00 | Death | RESPIRONICS, INC. | 2025-07-01 04:00 | BZD | DREAMSTATION AUTO CPAP | Insufficient Information | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION IT IS OBSERVED THAT THERE WAS NO VISIBLE FOAM PARTICLES FOUND. THE DEVICE WAS SCRAPPED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22370817&pc=BZD | 2518422-2025-106953 | 2025-06-26 04:00 | Death | RESPIRONICS, INC. | 2025-07-01 04:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM | Insufficient Information | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION RELATED TO A REP DREAMSTATION AUTO CPAP. THE REPORTER ALLEGES THE PATIENT RECEIVED A DEVICE WITH A DIFFERENT SERIAL NUMBER THAN THEIR ASSIGNED DEVICE. THE ALLEGATION INCLUDES THAT THE PATIENT IS DEAD. THERE IS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22371145&pc=DTB | 2649622-2025-17741 | 2025-06-10 04:00 | Death | MPRI | 2025-07-01 04:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS VERY SICK PRIOR TO THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM IMPLANT. POST-OPERATIVELY, THE PATIENT PASSED AWAY.Event Description: ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND IS NOT AVAILABLE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22371207&pc=LWS | 2649622-2025-17743 | 2025-06-10 04:00 | Death | MPRI | 2025-07-01 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS VERY SICK PRIOR TO THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM IMPLANT. POST-OPERATIVELY, THE PATIENT PASSED AWAY.Event Description: ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND IS NOT AVAILABLE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22371223&pc=LWS | 9614453-2025-02527 | 2025-06-10 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-07-01 04:00 | LWS | COBALT¿ XT DR MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS VERY SICK PRIOR TO THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM IMPLANT. POST-OPERATIVELY, THE PATIENT PASSED AWAY.Event Description: ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND IS NOT AVAILABLE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22371303&pc=LNQ | 3027386225-2025-00081 | 2025-05-12 04:00 | Death | ENTERRA MEDICAL, INC. | 2025-07-01 04:00 | LNQ | ENTERRA II IPG | No Apparent Adverse Event | Insufficient Information | H990014 |  | 1 | Event Description: PRODUCT RETURNED TO MEDTRONIC FOR ANALYSIS: REPORTED IN MEDTRONIC DATABASE. RECEIVED FROM (B)(6) IN (B)(6). CODED AS C-500 NON-COMPLAINT AT MEDTRONIC AND CLOSED. ORIGINALLY IMPLANTED ON (B)(6) 2022.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22371445&pc=DSQ | 2916596-2025-03984 | 2025-06-05 04:00 | Death | THORATEC CORPORATION | 2025-07-01 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hypervolemia | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED WITH AN EXCESS AMOUNT OF LOW FLOW ALARMS OVERNIGHT WITH NO OVERT SYMPTOMS. THE LOG FILE CAPTURED LOW FLOW EVENTS ON (B)(6) 2025. THERE DID NOT APPEAR TO BE ANY TYPE OF EXTERNAL EQUIPMENT CAUSING THE EVENTS. THE PATIENT HAD BEEN ON CONTINUOUS INFUSION OF LASIX FOR DIURESIS DUE TO ELEVATED JUGULAR VEIN DISTENTION (JVD). WHEN THE ALARMS STARTED, THE LASIX WAS HELD. THE PATIENT'S MEAN ARTERIAL PRESSURE (MAP) WAS 60-70S, TEMPERATURE 96.7 - 98 DEGREES FAHRENHEIT, PULSE 73-83, RESPIRATORY RATE 16-27, PERIPHERAL OXYGEN SATURATION OF 93%-100%, AND ARTERIAL LINE BLOOD PRESSURE (81-100)/ (52-66). THE COMPLETE BLOOD COUNT (CBC) WAS STABLE AND CONSISTENT WITH PRIOR VALUES. ELECTROLYTES WERE REPLACED AS NEEDED (PRN) BUT ALSO REMAIN IN SAME RANGE. BRAIN NATRIURETIC PEPTIDE (BNP) WAS ELEVATED FROM BASELINE CORRESPONDING WITH VOLUME ASSESSMENTS BY PHYSICIANS. ON (B)(6) 2025, THE PATIENT WAS TRANSITIONING TO COMFORT CARE AND MADE DO NOT RESUSCITATE. THE PATIENT PASSED AWAY DUE TO CARDIOPULMONARY ARREST. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22371463&pc=MVK | 3008642652-2025-06358 | 2025-06-17 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-01 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED EIGHT APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 17:05:12 ON (B)(6) 2025. AT 18:28:04, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 70 BPM DEGRADING TO VF. AT 18:28:54, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SEVERE BRADYCARDIA/SINUS BRADYCARDIA FROM 10-30 BPM WITH PVC¿S. AT 18:29:28, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF THE TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @50 BPM WITH NSVT. AT 18:32:50, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF THE TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 80 BPM WITH PAC¿S AND PVC¿S AT 18:35:17, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF THE TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 80 BPM. AT 18:36:59, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF THE TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 80 BPM WITH PAC¿S. AT 18:38:20, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF THE TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 70 BPM WITH PAC¿S. AT 18:40:54, THE PATIENT RECEIVED THE SEVENTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF THE TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 70 BPM. AT 18:42:57, THE PATIENT RECEIVED THE EIGHTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF THE TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 70 BPM WITH NSVT. THE RHYTHM THEN DEGRADED TO VF WITH MOTION ARTIFACT AND ELECTRODE LEAD FALL OFF. THE DEVICE WAS SHUT DOWN AT 19:21:40 ON (B)(6) 2025. Manufacturer Narrative: THE MONITOR/ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION. THE REVIEW OF THE DATA INDICATED THAT THE DEVICE POWERED ON NORMALLY AND WAS ABLE TO ACQUIRE THE PATIENT¿S ECG SIGNAL ON THE LAST DAY OF USE CAPTURED IN THE DATA DOWNLOAD. NO DEFICIENCIES ALLEGED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22371562&pc=LNQ | 3027386225-2025-00084 | 2025-05-12 04:00 | Death | ENTERRA MEDICAL, INC. | 2025-07-01 04:00 | LNQ | ENTERRA II IPG | No Apparent Adverse Event | Insufficient Information | H990014 |  | 1 | Event Description: PRODUCT RETURNED TO MEDTRONIC FOR ANALYSIS - REPORTED IN MEDTRONIC DATABASE. RECEIVED FROM (B)(6). CODED AS C-500 NON-COMPLAINT AT MEDTRONIC AND CLOSED. ORIGINALLY IMPLANTED ON (B)(6) 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22371880&pc=OZD | 1220648-2025-29942 | 2025-06-12 04:00 | Death | ABIOMED, INC. | 2025-07-01 04:00 | OZD | IMPELLA CP | Failure to Advance; Patient-Device Incompatibility | Cardiac Arrest | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT DURING IMPLANTATION OF AN IMPELLA CP, THE PEEL AWAY SHEATH COULD NOT BE ADVANCED DUE TO THE PATIENT'S ANATOMY. THE PATIENT EXPERIENCED CARDIAC ARREST AND REQUIRED EPINEPHRINE, ATROPINE, AND BICARBONATE. THE PHYSICIAN RESUMED IMPLANTATION USING A 11 FRENCH TERUMO SHEATH AND AN INTRA-AORTIC BALLOON PUMP WAS IMPLANTED INSTEAD. CARDIOPULMONARY RESUSCITATION CONTINUED BUT RETURN OF SPONTANEOUS CIRCULATION WAS NOT OBTAINED AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION DIFFICULTY INSERTING/REMOVING INTRODUCER AND CARDIAC ARREST HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE CAUSE OF THE DIFFICULTY INSERTING INTRODUCER ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED, BASED ON GIVEN CLINICAL DETAILS. THE ROOT CAUSE OF THE CARDIAC ARREST WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. B.2 REQUIRED INTERVENTION WAS INADVERTENTLY OMITTED FROM THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29942.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22372225&pc=NVN | 2182208-2025-03172 | 2024-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-07-01 04:00 | NVN | SELECTSECURE MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LONG-TERM CLINICAL OUTCOMES WITH LEFT BUNDLE BRANCH PACING (LBBP), LEFT VENTRICULAR SEPTAL PACING (LVSP), AND BIVENTRICULAR PACING (BIVP) FOR CARDIAC RESYNCHRONIZATION THERAPY. THE AUTHORS DESCRIBED PATIENT DEATHS IN ALL THREE GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DESCRIBED AS ALL-CAUSE WITH CARDIAC AND NON-CARDIAC DEATHS. THERE WERE PATIENTS WHO EXPERIENCED WORSENING HEART FAILURE SYMPTOMS WHICH REQUIRED ORAL OR INTRAVENOUS (IV) DIURETICS WITH HOSPITALIZATIONS IN ALL GROUPS, PHRENIC NERVE STIMULATION AND ONE PNEUMOTHORAX WHICH ONLY OCCURRED IN THE BIVP GROUP, AND INFECTIONS WHICH OCCURRED IN THE BIVP AND LVSP GROUPS. THERE WERE LEADS IN THE BIVP AND LVSP GROUPS WHICH EXHIBITED DISLODGMENTS AND INCREASED CAPTURE THRESHOLDS. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/63 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: COMPARISONS OF LONG-TERM CLINICAL OUTCOMES WITH LEFT BUNDLE BRANCH PACING, LEFT VENTRICULAR SEPTAL PACING, AND BIVENTRICULAR PACING FOR CARDIAC RESYNCHRONIZATION THERAPY. HEART RHYTHM. 2024; 21:1342¿1353. DOI: 10.1016/J.HRTHM.2024.03.007 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22372399&pc=NPT | 2025587-2025-04669 | 2025-02-25 05:00 | Death | MEDTRONIC IRELAND | 2025-07-01 04:00 | NPT | EVOLUT FX DCS | Patient Device Interaction Problem | Disseminated Intravascular Coagulation (DIC); Hemorrhage/Bleeding; Multiple Organ Failure; Embolism/Embolus; Thrombosis/Thrombus | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A "MASSIVE" HEMORRHAGE WAS NOTED OF SEVERITY THREE OR HIGHER ON THE BLEEDING ACADEMIC RESEARCH CONSORTIUM SCALE. THE PATIENT REQUIRED URGENT TREATMENT. A TRANSFEMORAL APPROACH WAS USED DUE TO SEVERE FRAILTY AND POOR VASCULAR CONDITION. SUBSEQUENTLY, A CHOLESTEROL EMBOLISM OCCURRED LEADING TO DISSEMINATED INTRAVASCULAR COAGULATION AND MULTIPLE ORGAN FAILURE. TWO DAYS AFTER THE PROCEDURE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS CHOLESTEROL EMBOLISM. PER THE PHYSICIAN, THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND THE DEVICE. THERE WAS A CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE PROCEDURE.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A "MASSIVE" HEMORRHAGE WAS NOTED OF SEVERITY THREE OR HIGHER ON THE BLEEDING ACADEMIC RESEARCH CONSORTIUM SCALE. THE PATIENT REQUIRED URGENT TREATMENT. A TRANSFEMORAL APPROACH WAS USED DUE TO SEVERE FRAILTY AND POOR VASCULAR CONDITION. SUBSEQUENTLY, A CHOLESTEROL EMBOLISM OCCURRED LEADING TO DISSEMINATED INTRAVASCULAR COAGULATION AND MULTIPLE ORGAN FAILURE. TWO DAYS AFTER THE PROCEDURE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS CHOLESTEROL EMBOLISM. PER THE PHYSICIAN, THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND THE DEVICE. THERE WAS A CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE PROCEDURE. THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE DEATH.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A "MASSIVE" HEMORRHAGE WAS NOTED OF SEVERITY THREE OR HIGHER ON THE BLEEDING ACADEMIC RESEARCH CONSORTIUM SCALE. THE PATIENT REQUIRED URGENT TREATMENT. A TRANSFEMORAL APPROACH WAS USED DUE TO SEVERE FRAILTY AND POOR VASCULAR CONDITION. SUBSEQUENTLY, A CHOLESTEROL EMBOLISM OCCURRED LEADING TO DISSEMINATED INTRAVASCULAR COAGULATION AND MULTIPLE ORGAN FAILURE. TWO DAYS AFTER THE PROCEDURE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS CHOLESTEROL EMBOLISM. PER THE PHYSICIAN, THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND THE DEVICE. THERE WAS A CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE PROCEDURE. THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT ACCESS WAS OBTAINED ON THE LEFT SIDE. THE VALVE, THE GUIDEWIRE, AND THE DELIVERY CATHETER SYSTEM (DCS) ALL DID NOT CONTRIBUTE TO THE VASCULAR INJURY. IT IS POSSIBLE THAT THE NON-MEDTRONIC SHEATH CONTRIBUTED TO THE VASCULAR INJURY. IT IS SUSPECTED THAT EITHER THE PERIPHERAL BALLOON OR THE SHEATH WAS THE CAUSE OF THE HEMORRHAGE AT THE LEFT COMMON ILIAC ARTERY. OF NOTE, DISSEMINATED INTRAVASCULAR COAGULATION REFERS TO ORGAN FAILURE CAUSED BY BLEEDING AND THROMBOSIS. NEITHER THE VALVE NOR THE DCS CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT FX VALVE; PRODUCT ID EVOLUTFX-26 (SERIAL: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED DATA: B5. ADDED SECOND PARAGRAPH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. ADDED THIRD PARAGRAPH. D3. D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22372442&pc=LNQ | 3027386225-2025-00074 | 2025-01-29 05:00 | Death | ENTERRA MEDICAL, INC. | 2025-07-01 04:00 | LNQ | ENTERRA II IPG | Insufficient Information | Insufficient Information | H990014 |  | 1 | Event Description: PATIENT DEATH. PRODUCT RETURNED TO MEDTRONIC FOR ANALYSIS - REPORTED IN MEDTRONIC DATABASE. CHECKED IN AT MDT (B)(6) 2025. MDT RECORD ID: (B)(4). AS REPORTED EVENT DESCRIPTION: PATIENT DEATH. CAUSE OF DEATH: PENDING FURTHER STUDIES. (LEADS RECEIVED CONNECTED TO DEVICE AND SEGMENTED AT PROXIMAL END). REPORTED IN DATABASE ONLY; WAITING FOR DEVICE ANALYSIS FROM MDT. DEVICE EXPLANTED AS A RESULT OF PATIENT DEATH. NO INDICATION THAT PATIENT DEATH WAS DUE TO DEVICE. NO FURTHER ACTION TO BE TAKEN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22372672&pc=OZP | 2032227-2025-202074 | 2025-03-21 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-01 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; High Blood Pressure/ Hypertension; Hypoglycemia; Myocardial Infarction | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2025, AND THE CAUSE OF DEATH WAS CORONARY ARTERY DISEASE, TYPE 1 DIABETES, HYPOGLYCEMIA, HYPERTENSION, AND HYPERCHOLESTEROLEMIA. THE EVENT INVOLVED PRODUCTS UNK_RESERVOIR, UNOMEDICAL, AND MMT-1884. TROUBLESHOOTING WAS PERFORMED AND FOUND THAT THE DOCUMENT OF ANY HEALTH ISSUES OR ILLNESSES THAT MAY HAVE CONTRIBUTED TO OR LED UP TO PASSING WAS SPECIFIED. THE CUSTOMER WAS STATED THAT THE PUMP WAS WORN AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR UNK_RESERVOIR, UNOMEDICAL, AND MMT-1884.Event Description: UPDATE SUMMARY: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2025, AND THE CAUSE OF DEATH WAS CORONARY ARTERY DISEASE, HYPERTENSION, AND HYPERCHOLESTEROLEMIA. THE EVENT INVOLVED PRODUCTS UNK_RESERVOIR, UNOMEDICAL, AND MMT-1884. TROUBLESHOOTING WAS PERFORMED AND FOUND THAT THE DOCUMENT OF ANY HEALTH ISSUES OR ILLNESSES THAT MAY HAVE CONTRIBUTED TO OR LED UP TO PASSING WAS SPECIFIED. THE CUSTOMER WAS STATED THAT THE PUMP WAS WORN AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR UNK_RESERVOIR, UNOMEDICAL, AND MMT-1884. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22372739&pc=DRM | 0003015876-2025-01321 | 2025-06-10 04:00 | Death | JOLIFE AB - 3005445717 | 2025-07-01 04:00 | DRM | LUCAS 3 CHEST COMPRESSION SYSTEM | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Unspecified Respiratory Problem | K161768 |  | 1 | Event Description: THE CUSTOMER CONTACTED STRYKER TO REPORT THAT THEIR DEVICE WAS USED ON A PATIENT IN CARDIAC ARREST. DURING USE OF THE LUCAS DEVICE, BLEEDING WAS OBSERVED FROM THE TRACHEAL INTUBATION TUBE AND THE NURSE REPORTED A SUSPECTED LUNG INJURY. THE PATIENT ASSOCIATED WITH THE REPORTED EVENT DID NOT SURVIVE. Manufacturer Narrative: STRYKER CONTINUES TO INVESTIGATE THE REPORTED FAILURE AND WILL SUBMIT A SUPPLEMENTAL REPORT ON THIS EVENT TO THE FDA AS PROVIDED BY 21 CFR 803.56. STRYKER PERFORMED A CLINICAL REVIEW OF THE REPORTED EVENT AND DETERMINED THAT THE DEVICE MAY HAVE CONTRIBUTED TO THE PATIENT OUTCOME. STRYKER CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. STRYKER REQUESTS PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. THE CUSTOMER PROVIDED STRYKER WITH THE AVAILABLE PATIENT INFORMATION. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22372741&pc=MKJ | 0003015876-2025-01320 | 2025-05-30 04:00 | Death | PHYSIO-CONTROL, INC. - 3015876 | 2025-07-01 04:00 | MKJ | LIFEPAK 15 MONITOR/DEFIBRILLATOR | Failure to Deliver Energy | Ventricular Fibrillation | P160026 |  | 1 | Event Description: THE CUSTOMER CONTACTED STRYKER TO REPORT THAT THEIR DEVICE WAS NOT DELIVERING A SHOCK. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THE PATIENT INVOLVED IN THE REPORTED EVENT IS DECEASED. Manufacturer Narrative: STRYKER CONTINUES TO INVESTIGATE THE REPORTED FAILURE AND WILL SUBMIT A SUPPLEMENTAL REPORT ON THIS EVENT TO THE FDA AS PROVIDED BY 21 CFR 803.56. STRYKER PERFORMED AN INITIAL CLINICAL REVIEW OF THE REPORTED EVENT AND DETERMINED THE DEVICE CONTRIBUTION TO THE PATIENT OUTCOME IS UNKNOWN. DUE TO CHARACTER LIMITATIONS SECTION A1, PATIENT IDENTIFIER, WAS LEFT BLANK. THE PATIENT IDENTIFIER IS (B)(6). STRYKER CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. STRYKER REQUESTS PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. THE CUSTOMER PROVIDED STRYKER WITH THE AVAILABLE PATIENT INFORMATION. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22373083&pc=MVK | 3008642652-2025-06368 | 2025-04-02 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-01 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED ONE APPROPRIATE TREATMENT. THE PATIENT ALSO RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 13:32:17 ON (B)(6) 2025. AT 03:18:58 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 80 BPM DEGRADING TO VT @ 200 BPM WITH MOTION ARTIFACT. AT 03:21:06, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VT @ 200 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 50 BPM WITH PVC¿S. AT 03:21:38, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. NSVT AND OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS NSVT. THE POST SHOCK RHYTHM WAS SINUS BRADYCARDIA AT 30 BPM. AT 03:22:13, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. NSVT AND OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION THE RHYTHM AT THE TIME OF TREATMENT WAS SINUS BRADYCARDIA @ 20 BPM WITH PVC¿S. POST SHOCK RHYTHM WAS SEVERE BRADYCARDIA @ 10 BPM. THE PATIENT WAS IN SINUS TACHYCARDIA @ 130 BPM WITH PVC¿S AT THE TIME OF THE DEVICE SHUTDOWN AT 03:49:14 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22373340&pc=OZD | 1220648-2025-29956 | 2025-06-15 04:00 | Death | ABIOMED, INC. | 2025-07-01 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT(PT) PRESENTED TO EMERGENCY ROOM AND CODED A TOTAL OF 5 TIMES. ONCE RETURN OF SPONTANEOUS CIRCULATION (ROSC)WAS ACHIEVED DR. (B)(6) DECIDED FOR IMPELLA CP PLACEMENT IF ADEQUATE VASCULATURE AND PERCUTANEOUS CORONARY INTERVENTION (PCI). IT WAS NOTED THAT ONCE THE PATIENT WAS MOVED OVER TO THE TABLE AND THE PATIENT CODED AGAIN. CARDIOPULMONARY RESUSCITATION PERFORMED WHILE PATIENT WAS PREPPED FOR IMPELLA PLACEMENT. ROSC ACHIEVED AND IMPELLA INSERTED VIA THE RIGHT FEMORAL ARTERY UTILIZING BEST PRACTICES WITHOUT DIFFICULTY. FLOWS OF 3.4 L/MIN OBTAINED. PCI OF LAD PERFORMED WITH 3 STENTS BEING DEPLOYED. PCI OF CIRC WAS THEN DONE WITH 1 STENT BEING DEPLOYED. DURING SUPPORT IMPELLA ALARMED SUCTION AND UNABLE TO FLOW ABOVE 1.0 L/MIN AND WAVEFORMS BECAME DISSOCIATED. AN ECHOCARDIOGRAM (ECHO) PERFORMED AGAIN SHOWING IMPELLA IN THE PROPER POSITION FREE IN THE MID-VENTRICULAR SPACE. DISCUSSION HAD WITH PHYSICIANS ABOUT POSSIBLE BIO-MATERIAL INGESTION. DECISION MADE TO EXPLANT THE PUMP. WHILE PREPARING FOR TRANSFER AN ECHO WAS PERFORMED TO CHECK FOR THROMBUS IN THE LEFT VENTRICLE AND NONE WAS VISUALIZED. THE IMPELLA WAS EXPLANTED WITHOUT DIFFICULTY. IMPELLA CP WAS EXAMINED AND BIO-MATERIAL WAS FOUND IN THE INLET. DECISION MADE BY DR. TO NOT IMPLANT ANOTHER CP DUE TO POSSIBILITY OF MORE BIOMATERIAL INGESTION. ADDITIONAL INFORMATION RECEIVED: THE PUMP IS NOT AVAILABLE FOR RETURN; IT WAS DISPOSED OF BY THE STAFF. THE PATIENT PASSED AWAY THE NEXT DAY. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW AND THROMBUS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. LOW PUMP FLOWS: DATA LOGS WERE REVIEWED AND THE DROP IN FLOWS WITH SUCTION WAS CONFIRMED. WHILE RUNNING AT P-9, PUMP FLOWS SUDDENLY DROPPED FROM 3.65 L/MIN TO 1.02 L/MIN, AND 8 SECONDS LATER, THE SUCTION ALARM TRIGGERED. THE SUCTION ALARM TRIGGERED BECAUSE MOTOR CURRENT (MC) MAX VALUES WERE VERY LOW, INDICATING A HIGH DIFFERENTIAL PRESSURE EVEN DURING SYSTOLE, WHEN THE VALVE IS OPEN. DURING THE SUBSEQUENT LOW FLOWS, THE PLACEMENT SIGNAL (PS) REMAINED PULSATILE AND AORTIC WITH VALUES AROUND 119/63 MMHG, SUGGESTING THAT THE PATIENT'S LEFT VENTRICLE (LV) WAS CONTRACTING. THIS ALL ALIGNS WITH THE CLINICAL REPORT OF BIOMATERIAL BEING VISUALIZED IN THE INFLOW CAGE UPON EXPLANT AND NO CANNULA KINK. SINCE BIOMATERIAL WAS IN THE INFLOW CAGE, IT HAD NOT MIGRATED DOWN TO THE IMPELLER, AND THEREFORE, NO MOTOR CURRENT SPIKE IS TO BE EXPECTED. THE BIOMATERIAL LIKELY OCCLUDED THE CANNULA FROM TYPICAL FLOW RATES AT P-9. PRODUCT WAS NOT RETURNED FOR INVESTIGATION. BASED ON THE CLINICAL DETAILS PROVIDED THAT BIOMATERIAL WAS VISUALIZED IN THE INFLOW CAGE AND DATA LOGS SHOWED A SUDDEN DROP IN PUMP FLOWS WHILE PS REMAINED STABLE, THE ROOT CAUSE OF THE LOW PUMP FLOWS WAS DETERMINED TO MOST LIKELY BE BIOMATERIAL. THROMBUS: THE ROOT CAUSE OF THE THROMBUS WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION, SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22373536&pc=NPT | 2025587-2025-04667 | 2024-08-10 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-01 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Sepsis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH AND 3 WEEKS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY DUE TO SEPSIS. PER THE PHYSICIAN, NEITHER THE DEVICE NOR THE PROCEDURE CONTRIBUTED TO THE CAUSE OF DEATH. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22373631&pc=DXY | 2182208-2025-03173 | 2024-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-07-01 04:00 | DXY | MEDTRONIC LEAD | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LONG-TERM CLINICAL OUTCOMES WITH LEFT BUNDLE BRANCH PACING (LBBP), LEFT VENTRICULAR SEPTAL PACING (LVSP), AND BIVENTRICULAR PACING (BIVP) FOR CARDIAC RESYNCHRONIZATION THERAPY. THE AUTHORS DESCRIBED PATIENT DEATHS IN ALL THREE GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DESCRIBED AS ALL-CAUSE WITH CARDIAC AND NON-CARDIAC DEATHS. THERE WERE PATIENTS WHO EXPERIENCED WORSENING HEART FAILURE SYMPTOMS WHICH REQUIRED ORAL OR INTRAVENOUS (IV) DIURETICS WITH HOSPITALIZATIONS IN ALL GROUPS, PHRENIC NERVE STIMULATION AND ONE PNEUMOTHORAX WHICH ONLY OCCURRED IN THE BIVP GROUP, AND INFECTIONS WHICH OCCURRED IN THE BIVP AND LVSP GROUPS. THERE WERE LEADS IN THE BIVP AND LVSP GROUPS WHICH EXHIBITED DISLODGMENTS AND INCREASED CAPTURE THRESHOLDS. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/63 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: COMPARISONS OF LONG-TERM CLINICAL OUTCOMES WITH LEFT BUNDLE BRANCH PACING, LEFT VENTRICULAR SEPTAL PACING, AND BIVENTRICULAR PACING FOR CARDIAC RESYNCHRONIZATION THERAPY. HEART RHYTHM. 2024; 21:1342¿1353. DOI: 10.1016/J.HRTHM.2024.03.007 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22373789&pc=HGM | 9610816-2025-000527 | 2025-04-13 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-07-01 04:00 | HGM | AVALON FETAL MONITOR FM50 | Wireless Communication Problem | Insufficient Information | K140535 |  | 1 | Event Description: IT WAS REPORTED THE STAFF ATTEMPTED TO CONNECT ULTRASOUND AND TOCO TRANSDUCERS TO THE MOTHER BUT COULD NOT GET THE TRANSDUCERS TO CONNECT. THE MOTHER, WHO WAS PREGNANT WITH TWINS, WAS IN THE LABOR AND DELIVERY ROOM AND ALL DEVICES WERE WORKING CORRECTLY. THE MOTHER WAS TRANSPORTED TO THE OR TO DELIVER THE FIRST BABY BUT THEN TRANSPORTED BACK TO L&D TO DELIVER THE SECOND BABY, WHO WAS IN A GOOD POSITION FOR A NATURAL BIRTH. WHEN ATTEMPTING TO CONNECT THE MOTHER TO WIRELESS TRANSDUCERS, THE TRANSDUCERS WOULD NOT CONNECT TO THE FETAL MONITOR. AFTER TROUBLESHOOTING FOR 8 MINUTES, A WIRED TRANSDUCER WAS OBTAINED; HOWEVER, AROUND MINUTE 10, THE BABY HAD ALREADY BEEN BORN. THIRTY MINUTES LATER, THE SECOND BABY PASSED AWAY DUE TO NATURAL CAUSES; HOWEVER, THE CAUSE OF DEATH REMAINS UNKNOWN. THE MONITOR WAS VISUALLY INSPECTED BY THE FIELD SERVICE ENGINEER, REVEALING NO ANOMALIES. THE ULTRASOUND AND TOCO TRANSDUCERS WERE LATCHED IMMEDIATELY UPON DOCKING AND THE ENGINEER WAS UNABLE TO REPLICATE THE COMPLAINT. Manufacturer Narrative: A PHILIPS FIELD SERVICE ENGINEER (FSE) VISITED THE CUSTOMER SITE. THE ULTRASOUND (US) AND TOCO TRANSDUCERS INVOLVED WERE UNKNOWN AND COULD NOT BE PROVIDED. THE STAFF PROCURED A US AND TOCO FOR TESTING PURPOSE. THE US AND TOCO TRANSDUCERS LATCHED IMMEDIATELY UPON DOCKING AND READINGS WERE REPLICATED ON FM50. THE FSE WAS UNABLE TO REPRODUCE ERRORS OR THE CUSTOMER ALLEGATION. A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) REVIEWED THE SAMPLE 15 MINUTE TRACE PROVIDED (SEE ATTACHMENT). THE PSE INDICATED THAT THE INFORMATION PROVIDED DID NOT APPEAR TO BE RELATED TO THE DEVICES INVOLVED IN THE COMPLAINT; HOWEVER, THE STRIP MAKES IT CLEAR THAT THERE IS A LOT OF "NOISE" (WIRELESS INTERFERENCE) IN THE CUSTOMER'S ENVIRONMENT. THE PSE INDICATED THAT THE CUSTOMER SHOULD VERIFY THAT THE RSSI LEVELS DO NOT EXCEED 100 FOR OPTIMAL PERFORMANCE OF THE DEVICE. THE PSE ALSO INDICATED IT WOULD BE UNUSUAL FOR BOTH THE TOCO AND ULTRASOUND TRANSDUCERS TO FAIL, WHICH POINTS TO A POTENTIALLY 'NOISY' ENVIRONMENT; HOWEVER, THE CAUSE COULD NOT BE DEFINITIVELY DETERMINED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. AFTER FURTHER INVESTIGATION RELATED REPORT NUMBER(S) 9610816-2025-000560 & 9610816-2025-000561 HAVE BEEN DETERMINED TO BE NON-REPORTABLE. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. THE ACTUAL TRANSDUCERS USED DURING THE EVENT WERE UNABLE TO BE LOCATED AS IT IS PRESUMED, THEY ARE PART OF A POOL OF DEVICES. THE STRIPS PROVIDED DO NOT APPEAR TO BE THE RELATED TO THE DEVICES INVOLVED IN THE COMPLAINT; HOWEVER, THE STRIPS MAKE IT CLEAR THERE IS A LOT OF NOISE IN THE CUSTOMER'S ENVIRONMENT. IT WAS INDICATED THAT THE CUSTOMER SHOULD VERIFY THE RSSI LEVELS DO NOT EXCEED 100 FOR OPTIMAL PERFORMANCE OF THE DEVICE. THE PSE INDICATED IT WOULD BE UNUSUAL FOR BOTH THE TOCO AND ULTRASOUND TRANSDUCERS TO FAIL, WHICH POINTS TO A POTENTIALLY 'NOISY' ENVIRONMENT; HOWEVER, THE CAUSE COULD NOT BE DEFINITIVELY DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22373909&pc=NVZ | 9614453-2025-02533 | 2025-06-01 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-07-01 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; No Clinical Signs, Symptoms or Conditions | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY POST MAGNETIC RESONANCE IMAGING (MRI). THE PATIENT RECEIVED AN MRI OF THE FOOT DUE TO A SYSTEMIC INFECTION AND IT WAS NOTED THAT POST MRI THE RIGHT VENTRICULAR (RV) HIS BUNDLE LEAD EXHIBITED HIGH AND RISING THRESHOLDS AND LOSS OF CAPTURE. THE PHYSICIAN ALLEGES THE LEAD MALFUNCTIONED DUE TO THE MRI. THE PATIENT PASSED AWAY IN THE INTENSIVE CARE UNIT (ICU) AND AN IMPLANTABLE PULSE GENERATOR (IPG) CHECK NOTED ONE EPISODE OF NON-SUSTAINED VENTRICULAR TACHYCARDIA. THE CAUSE OF DEATH WAS REQUESTED BUT IS NOT AVAILABLE. Manufacturer Narrative: CORRECTED H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: EVFXPLUS-29 TRANSCATHETER VALVE IMPLANTED ON (B)(6) 2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22373911&pc=NVN | 2649622-2025-17774 | 2025-06-01 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-01 04:00 | NVN | SELECTSECURE MRI SURESCAN | Failure to Capture; Capturing Problem; High Capture Threshold | Unspecified Infection | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY POST MAGNETIC RESONANCE IMAGING (MRI). THE PATIENT RECEIVED AN MRI OF THE FOOT DUE TO A SYSTEMIC INFECTION AND IT WAS NOTED THAT POST MRI THE RIGHT VENTRICULAR (RV) HIS BUNDLE LEAD EXHIBITED HIGH AND RISING THRESHOLDS AND LOSS OF CAPTURE. THE PHYSICIAN ALLEGES THE LEAD MALFUNCTIONED DUE TO THE MRI. THE PATIENT PASSED AWAY IN THE INTENSIVE CARE UNIT (ICU) AND AN IMPLANTABLE PULSE GENERATOR (IPG) CHECK NOTED ONE EPISODE OF NON-SUSTAINED VENTRICULAR TACHYCARDIA. THE CAUSE OF DEATH WAS REQUESTED BUT IS NOT AVAILABLE. Manufacturer Narrative: CORRECTED H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: EVFXPLUS-29 TRANSCATHETER VALVE IMPLANTED (B)(6) 2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22373912&pc=NRY | 3011370111-2025-00064 | 2024-05-18 04:00 | Death | NEURAVI LTD. | 2025-07-01 04:00 | NRY | EMBOTRAP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Intracranial Hemorrhage | K173452 |  | 1 | Event Description: AS REPORTED VIA THE EXCELLENT STUDY (B)(6), A 70-YEAR-OLD MALE (SUBJECT (B)(6)) WITH A MEDICAL HISTORY OF CONGESTIVE HEART FAILURE, ATRIAL FIBRILLATION, PREVIOUS ISCHEMIC STROKE (¿UN-UNK-2019¿), AND HYPERTENSION, PRESENTED WITH A WITNESSED STROKE ON (B)(6) 2024 AT 01:30. THE PATIENT WAS PRESENTED TO THE TREATING HOSPITAL ON THE SAME DAY AT 08:30. INTRAVENOUS TISSUE PLASMINOGEN ACTIVATOR (TPA) WAS NOT ADMINISTERED AT THE TIME OF STROKE PRESENTATION. THE SUSPECTED ORIGIN OF THE EMBOLISM WAS ¿ARTERY TO ARTERY (TANDEM LESION WITH PROXIMAL ICA STENOSIS AND INTRACRANIAL OCCLUSION).¿ THE PATIENT¿S BASELINE NIHSS SCORE WAS 12 AND A MRS SCORE OF ¿1-NO SIGNIFICANT DISABILITY. ABLE TO CARRY OUT ALL USUAL DUTIES AND ACTIVITIES.¿ THE PATIENT UNDERWENT AN ENDOVASCULAR MECHANICAL THROMBECTOMY ON (B)(6) 2024 USING AN EMBOTRAP II 5X33 REVASC. DEV. (ET007533/ 23E251AV) FOR AN OCCLUSION AT THE M1 SEGMENT OF THE LEFT MIDDLE CEREBRAL ARTERY (MCA). THE PRE-PASS MTICI SCORE WAS 0. THE 1ST PASS RESULTED IN A MTICI SCORE OF 1, WITH CLOT RETRIEVAL IN THE STENT RETRIEVER. THE 2ND PASS RESULTED IN A MTICI SCORE OF 2A, WITH CLOT RETRIEVAL IN THE STENT RETRIEVER. DURING THE PROCEDURE, A GUIDEWIRE WAS USED, BUT THE BRAND WAS NOT SPECIFIED. A 0.021 PROWLER SELECT PLUS (606S255X/ 31135553) MICROCATHETER, A 0.066 NAVIEN INTERMEDIATE CATHETER, AND A 0.079 PENUMBRA NEURON MAX LONG SHEATH WERE ALSO USED. THERE WERE NO INTRAOPERATIVE STUDY DEVICE DEFICIENCIES. THE PATIENT¿S 24-HOUR POST-PROCEDURE NIHSS SCORE WAS 40. ON (B)(6) 2024, THE PATIENT EXPERIENCED THE ADVERSE EVENT OF ¿ISCHEMIC STROKE WITH POST-INFARCTION HEMORRHAGE,¿ WHICH BECAME KNOWN TO THE SITE AND SPONSOR ON 21-MAY-2024. THE PRINCIPAL INVESTIGATOR (PI) ASSESSED THIS EVENT AS A SERIOUS, SEVERE ADVERSE EVENT, THAT WAS UNRELATED TO THE STUDY DEVICE, UNRELATED TO THE LARGE BORE CATHETER, UNRELATED TO THE CEREGLIDE71, AND UNRELATED TO THE SURGICAL PROCEDURE. AS TO WHETHER THE EVENT WAS ANTICIPATED, THE ANSWER FIELD IS RECORDED AS ¿N/A (DOES NOT MEET ABOVE CRITERIA)¿. THE EVENT WAS MEDICINALLY TREATED. THE OUTCOME WAS FATAL, AS THE PATIENT EXPIRED ON (B)(6) 2024. ADDITIONAL INFORMATION WAS RECEIVED ON 02-JUN-2024. SUMMARY: THE INFORMATION PROVIDED CONFIRMED THE EVENT OF ¿ISCHEMIC STROKE WITH POST-INFARCTION HEMORRHAGE¿ DID NOT OCCUR DURING THE PROCEDURE. THERE WAS NO VESSEL TRAUMA ASSOCIATED WITH THE EVENT. THE EVENT WAS ATTRIBUTED TO PROGRESSION OF INDEX STROKE. ADDITIONAL/MODIFIED INFORMATION WAS RECEIVED ON 19-JUN-2024. SUMMARY: THE ADVERSE EVENT TERM "ISCHEMIC STROKE WITH POST-INFARCTION HEMORRHAGE" WAS UPDATED TO "ISCHEMIC STROKE." ADDITIONAL INFORMATION WAS RECEIVED ON 20-JUN-2024. SUMMARY: ON (B)(6) 2024, THE PATIENT EXPERIENCED THE ADVERSE EVENT OF ¿POST-INFARCTION HEMORRHAGE,¿ WHICH BECAME KNOWN TO THE SITE ON 21-MAY-2024 AND TO THE SPONSOR ON 20-JUN-2024. THE PRINCIPAL INVESTIGATOR (PI) ASSESSED THIS EVENT AS A SERIOUS, SEVERE ADVERSE EVENT, THAT WAS UNRELATED TO THE STUDY DEVICE, UNRELATED TO THE LARGE BORE CATHETER, UNRELATED TO THE CEREGLIDE71, AND UNRELATED TO THE SURGICAL PROCEDURE. AS TO WHETHER THE EVENT WAS ANTICIPATED, THE ANSWER FIELD IS RECORDED AS ¿N/A (DOES NOT MEET ABOVE CRITERIA)¿. THE EVENT WAS MEDICINALLY TREATED. THE OUTCOME IS RECORDED AS ¿FATAL¿ WITH AN END DATE OF (B)(6) 2024. LISTED. ADDITIONAL/MODIFIED INFORMATION WAS RECEIVED ON 27-JUN-2024. SUMMARY: THE ADVERSE EVENT TERM "ISCHEMIC STROKE" WAS UPDATED TO "PROGRESSION OF INDEX ISCHEMIC STROKE" ADDITIONAL/MODIFIED INFORMATION WAS RECEIVED ON 25-SEP-2024. SUMMARY: THE ADVERSE EVENT TERM: "PROGRESSION OF INDEX ISCHEMIC STROKE" WAS UPDATED TO "AGGRAVATED CEREBRAL INFARCTION." REGARDING THE EVENT OF ¿POST-INFARCTION HEMORRHAGE,¿ THE OUTCOME OF "FATAL" WAS UPDATED TO "RECOVERING/RESOLVING," THE ANSWER FIELD FOR ¿IS THE ADVERSE EVENT SERIOUS?¿ WAS UPDATED FROM "YES" TO "NO," AND THE END DATE OF "(B)(6) 2024" WAS UPDATED TO "[BLANK]." ADDITIONAL/MODIFIED INFORMATION WAS NOTED ON THE CLINICAL DATABASE AT THE TIME OF THIS REVIEW, 26-SEP-2024. SUMMARY: THE SUSPECTED ORIGIN OF THE EMBOLISM WAS UPDATED FROM ¿ARTERY TO ARTERY (TANDEM LESION WITH PROXIMAL ICA STENOSIS AND INTRACRANIAL OCCLUSION)¿ TO ¿CARDIOEMBOLIC.¿ ADDITIONAL INTERVENTION PERFORMED AFTER THE FIRST THROMBECTOMY PASS WAS NOTED AS ¿PROXIMAL CAROTID ATHEROSCLEROTIC LESION STENTING.¿ AN ADDITIONAL OCCLUSION SITE WAS NOTED; THE ADDITIONAL OCCLUSION SITE WAS THE LEFT POSTERIOR COMMUNICATING ARTERY (PCOMM), WHICH WAS TARGETED (IN ADDITION TO THE M1 SITE) FOR BOTH THE 1ST AND 2ND PASS. AN ADDITIONAL RETRIEVAL ATTEMPT DURING THE PROCEDURE WAS NOTED; THE 3RD PASS WAS MADE AT THE PCOMM SITE USING A DIRECT CONTACT ASPIRATION DEVICE ALONE, WHICH RESULTED IN A MTICI SCORE OF 2B, WITH CLOT RETRIEVAL IN THE ASPIRATE. THE INTERMEDIATE CATHETER USED WAS UPDATED FROM ¿A 0.066 NAVIEN INTERMEDIATE CATHETER¿ TO A ¿0.055 YIN SHE INTERMEDIATE CATHETER.¿ ON (B)(6) 2024, THE PATIENT WAS DISCHARGE TO ANOTHER HOSPITAL WITH A NIHSS SCORE OF 40 AND A MRS SCORE OF ¿5-SEVERE DISABILITY. BEDRIDDEN, INCONTINENT, AND REQUIRING CONSTANT NURSING CARE AND ATTENTION.¿ THE OUTCOME FOR THE EVENT OF ¿AGGRAVATED CEREBRAL INFARCTION¿ WAS UPDATED TO ¿FATAL.¿ ADDITIONAL/MODIFIED INFORMATION WAS RECEIVED ON 28-NOV-2024. SUMMARY: THE PATIENT¿S EXPIRATION DATE WAS UPDATED FROM "(B)(6) 2024" TO "(B)(6) 2024." ADDITIONAL/MODIFIED INFORMATION WAS RECEIVED ON 13-FEB-2025. SUMMARY: REGARDING THE EVENT OF ¿POST-INFARCTION HEMORRHAGE,¿ THE SEVERITY WAS UPDATED FROM "SEVERE" TO "MODERATE." ADDITIONAL/MODIFIED INFORMATION WAS RECEIVED ON 24-JUN-2025. SUMMARY: ON 19-JUN-2025, A CLINICAL EVENT COMMITTEE (CEC) ADJUDICATION FOR THE ADVERSE EVENT OF ¿AGGRAVATED CEREBRAL INFARCTION¿ WAS RECEIVED. THE CEC ADJUDICATED EVENT TERM WAS UPDATED TO ¿PROGRESSION OF INDEX STROKE.¿ THE RELATIONSHIP OF EVENT TO THE EMBOTRAP STUDY DEVICE AND TO THE STUDY PROCEDURE WERE UPDATED TO ¿POSSIBLY RELATED.¿ THE RELATIONSHIP OF EVENT TO THE EMBOVAC LARGE BORE CATHETER AND CEREGLIDE 71 INTERMEDIATE CATHETER REMAINS ASSESSED AS ¿NOT RELATED.¿ IT WAS FURTHER COMMENTED, ¿LARGE BASELINE STROKE.¿ PER THE ADDITIONAL/MODIFIED INFORMATION NOTED ON THE CLINICAL DATABASE ON 26-SEP-2024; IT WAS DISCLOSED THREE RETRIEVAL ATTEMPTS WERE MADE DURING THE PROCEDURE. THE USE OF THE EMBOTRAP III DEVICE WAS TERMINATED BEFORE THE 3 ALLOWED RETRIEVAL ATTEMPTS PER THE INSTRUCTIONS FOR USE (IFU) AND AN ALTERNATE DEVICE WAS UTILIZED TO CONTINUE THE PROCEDURE BASED ON THE DISCRETION OF THE TREATING PHYSICIAN. NO NEW PATIENT CONSEQUENCES HAVE OCCURRED BECAUSE OF THE USED DEVICE. THE DEVICE PERFORMED AS INTENDED WITH NO ALLEGED QUALITY ISSUE IN =2 PASSES. THEREFORE, THE CODES HAVE BEEN UPDATED APPROPRIATELY; THE EVENTS WILL BE CODED BUT NOT REPORTED AND THE FILE WILL BE CONVERTED TO A PRODUCT COMPLAINT. THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. Manufacturer Narrative: MANUFACTURER REF#: (B)(4). SECTION A1. PATIENT IDENTIFIER: (B)(6). THE DEVICE WAS DISCARDED; THEREFORE, NO FURTHER INVESTIGATION CAN BE PERFORMED. A REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH LOT 23E251AV PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESS THAT CAN BE RELATED TO THE REPORTED COMPLAINT. PER THE ADDITIONAL/MODIFIED INFORMATION RECEIVED ON 24-JUN-2025, REGARDING THE EVENT OF ¿AGGRAVATED CEREBRAL INFARCTION/PROGRESSION OF INDEX STROKE,¿ WHICH RESULTED IN DEATH; CEREBRAL INFARCTION IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE USE OF THE EMBOTRAP III STUDY DEVICE AND IS MENTIONED IN THE INSTRUCTIONS FOR USE (IFU) AS SUCH. THE PRINCIPAL INVESTIGATOR ASSESSED THE EVENT OF ¿AGGRAVATED CEREBRAL INFARCTION/PROGRESSION OF INDEX STROKE¿ AS NOT RELATED TO THE USED DEVICE NOR PROCEDURE; HOWEVER, THE CLINICAL EVENT COMMITTEE (CEC) ADJUDICATION DEEMED THE EVENT AS ¿POSSIBLY RELATED¿ TO THE USED EMBOTRAP III STUDY DEVICE AND TO THE STUDY PROCEDURE. BASED ON THIS ASSESSMENT, THE CORRELATING RELATIONSHIP BETWEEN THE EVENT TO THE USED EMBOTRAP III STUDY DEVICE CANNOT BE RULED OUT ENTIRELY. BASED ON THE ASSESSMENT OF THE CLINICAL EVENT COMMITTEE (CEC), THIS EVENT DOES MEET US FDA REPORTING CRITERIA UNDER 21 CFR 803 WITH A CLASSIFICATION OF ¿DEATH,¿ WITH AN AWARENESS DATE OF 24-JUN-2025. THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. AS PART OF THE POST MARKET SURVEILLANCE PROGRAM, INFORMATION FROM THIS COMPLAINT IS TRENDED TO IDENTIFY STATISTICAL SIGNALS FOR CONSIDERATION OF FURTHER CORRECTION ACTION. SINCE THERE WAS NO EVIDENCE TO SUGGEST THE EVENT WAS RELATED TO A MANUFACTURING OR DESIGN ISSUE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS BLANK; THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803 (AND/OR PART 4, AS APPLICABLE). THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY CERENOVUS, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, CERENOVUS, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22373932&pc=DTB | 2649622-2025-17775 | 2025-06-01 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-01 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY POST MAGNETIC RESONANCE IMAGING (MRI). THE PATIENT RECEIVED AN MRI OF THE FOOT DUE TO A SYSTEMIC INFECTION AND IT WAS NOTED THAT POST MRI THE RIGHT VENTRICULAR (RV) HIS BUNDLE LEAD EXHIBITED HIGH AND RISING THRESHOLDS AND LOSS OF CAPTURE. THE PHYSICIAN ALLEGES THE LEAD MALFUNCTIONED DUE TO THE MRI. THE PATIENT PASSED AWAY IN THE INTENSIVE CARE UNIT (ICU) AND AN IMPLANTABLE PULSE GENERATOR (IPG) CHECK NOTED ONE EPISODE OF NON-SUSTAINED VENTRICULAR TACHYCARDIA. THE CAUSE OF DEATH WAS REQUESTED BUT IS NOT AVAILABLE. Manufacturer Narrative: CONTINUATION OF D10: EVFXPLUS-29 TRANSCATHETER VALVE IMPLANTED (B)(6) 2024 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22374069&pc=DYE | 2015691-2025-05326 | 2025-05-30 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-01 04:00 | DYE | CARPENTIER-EDWARDS PERIMOUNT MAGNA MITRAL EASE PERICARDIAL BIOPROSTHESIS | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Rupture; Insufficient Information | P860057 |  | 1 | Event Description: IT WAS REPORTED A 31MM 7300TFX MITRAL VALVE WAS EXPLANTED AT IMPLANT DUE TO SIZING ISSUES. THE EXPLANTED DEVICE WAS REPLACED WITH A 25MM 7300TFX MITRAL VALVE. PATIENT EXPIRED DUE TO HOLE IN ATRIUM.Event Description: IT WAS REPORTED A 31MM 7300TFX MITRAL VALVE WAS EXPLANTED AT IMPLANT DUE TO ATRIAL RUPTURE. THE EXPLANTED DEVICE WAS REPLACED WITH A 25MM 7300TFX MITRAL VALVE. PATIENT EXPIRED DUE TO ATRIAL RUPTURE. PER MEDICAL RECORDS, PATIENT WAS SUFFERING FROM SEVERE MAC AND STENOSIS WITH SEVERE MULTIVESSEL CORONARY ARTERY DISEASE. INTRAOPERATIVELY, SURGEON EXTENSIVELY DISSECTED THE HEART AND SONDERGAARD'S GROOVE. LEFT ATRIOTOMY WAS THEN PERFORMED. UPON INSPECTION, SURGEON NOTES DUE TO HORRIFIC MAC AND PRIOR TAVR, THERE WAS DIFFICULTY EXPOSING OF THE MITRAL VALVE AND WAS HIGH RISK FOR AV RUPTURE. SURGEON DEBRIDED THE ANTERIOR LEAFLET AND MINIMALLY DEBRIDED THE POSTERIOR LEAFLET TISSUE. SURGEON WAS UNABLE TO EASILY PASS THROUGH MITRAL ANNULUS, THEREFORE, A 31MM 7300TFX VALVE WAS PLACED SUPRA-ANNULARLY ON THE ATRIUM TISSUE. COMING OFF CPB, BLEEDING WAS NOTED AND ATRIAL RUPTURE WAS SUSPECTED. THE VALVE WAS EXPLANTED AND REPLACED WITH A 25MM 7300 WHICH WAS PLACED WITHIN THE TRUE ANNULUS AND THE ATRIAL TISSUE RUPTURE WAS REPAIRED. UPON COMING OFF CPB, EXCESSIVE BLEEDING WAS NOTED IN THE REGION OF THE LEFT ATRIAL DOME. THIS WAS UNABLE TO BE REPAIRED SUCCESSFULLY. PATIENT EXPIRED IN THE OR. Manufacturer Narrative: H11: ADDITIONAL MANUFACTURER NARRATIVE: THE DEVICE HISTORY RECORD (DHR) REVIEW WAS COMPLETED AND THIS DEVICE PASSED ALL MANUFACTURING AND STERILIZATION INSPECTIONS PRIOR TO RELEASE FOR DISTRIBUTION. Manufacturer Narrative: ADDED INFORMATION TO H6. THE COMPLAINT IS ABLE TO BE CONFIRMED VIA MEDICAL RECORDS. THERE IS NO EVIDENCE TO SUGGEST AN EDWARDS MANUFACTURING DEFECT. IT WAS REPORTED A 31MM 7300TFX MITRAL VALVE WAS EXPLANTED AT IMPLANT DUE TO ATRIAL RUPTURE. THE EXPLANTED DEVICE WAS REPLACED WITH A 25MM 7300TFX MITRAL VALVE. PATIENT EXPIRED DUE TO ATRIAL RUPTURE. THIS IS AN 84-YEAR-OLD FEMALE PATIENT WITH HISTORY OF TAVR, SEVERE MULTIVESSEL CAD, PAFIB. THE SUBJECT DEVICE WAS NOT RETURNED FOR EVALUATION AS THERE WAS NO ALLEGATION TOWARDS DEVICE MALFUNCTION OR CONTRIBUTION TO THE PATIENT'S DEATH. MEDICAL RECORDS WERE PROVIDED. PER THE MEDICAL RECORDS, PATIENT WAS SUFFERING FROM SEVERE MAC (MITRAL ANNULAR CALCIFICATION) AND STENOSIS WITH SEVERE MULTIVESSEL CORONARY ARTERY DISEASE. INTRAOPERATIVELY, SURGEON EXTENSIVELY DISSECTED THE HEART AND SONDERGAARD GROOVE. LEFT ATRIOTOMY WAS THEN PERFORMED. UPON INSPECTION, SURGEON NOTES DUE TO HORRIFIC MAC AND PRIOR TAVR, THERE WAS DIFFICULTY EXPOSING OF THE MITRAL VALVE AND WAS HIGH RISK FOR AV RUPTURE. SURGEON DEBRIDED THE ANTERIOR LEAFLET AND MINIMALLY DEBRIDED THE POSTERIOR LEAFLET TISSUE. SURGEON WAS UNABLE TO EASILY PASS THROUGH MITRAL ANNULUS, THEREFORE, A 31MM 7300TFX VALVE WAS PLACED SUPRA-ANNULARLY ON THE ATRIUM TISSUE. COMING OFF CPB, BLEEDING WAS NOTED AND ATRIAL RUPTURE WAS SUSPECTED. THE VALVE WAS EXPLANTED AND REPLACED WITH A 25MM 7300TFX THE RUPTURE WAS REPAIRED. UPON COMING OFF CPB, EXCESSIVE BLEEDING WAS NOTED IN THE REGION OF THE LEFT ATRIAL DOME. THIS WAS UNABLE TO BE REPAIRED SUCCESSFULLY. PATIENT EXPIRED IN THE OPERATING ROOM. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED, AND NO RELEVANT NON-CONFORMANCES WERE IDENTIFIED. THE INSTRUCTIONS FOR USE (IFU) HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARD TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THE REPORTED TYPE OF EVENT IS INCLUDED IN THE IFU. BASED ON THE INFORMATION AVAILABLE, THE COMPLAINT WAS ABLE TO BE CONFIRMED VIA MEDICAL RECORDS. PER THE MEDICAL RECORDS, THE PATIENT SEVERE MITRAL ANNULAR CALCIFICATION AND MULTIVESSEL CORONARY ARTERY DISEASE INCREASED THE PATIENT'S RISK OF COMPLICATIONS DURING THE SURGERY. ADDITIONALLY, THE PATIENT PRIOR TAVR MITRAL VALVE AND NATIVE ANATOMY MADE IT MORE DIFFICULT FOR THE SURGEON TO REACH THE MITRAL VALVE, INCREASING RISK FOR AV RUPTURE. THE SURGEON REPORTED SIZING ISSUES AS A RESULT OF DIFFICULTY REACHING THE PATIENT MITRAL ANNULUS. THEREFORE, THE MOST LIKELY CAUSE OF THE REPORTED ATRIAL RUPTURE IS PATIENT AND PROCEDURAL FACTORS. AN EDWARDS DEFECT HAS NOT BEEN CONFIRMED. Manufacturer Narrative: EDWARDS WILL CONTINUE TO REVIEW AND MONITOR ALL EVENTS. TRENDS ARE MONITORED ON A THE INVESTIGATION IS STILL IN PROGRESS; THEREFORE, A CONCLUSION HAS YET TO BE ESTABLISHED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY UPON INVESTIGATION COMPLETION. THE INVESTIGATION IS STILL IN PROGRESS; THEREFORE, A CONCLUSION HAS YET TO BE ESTABLISHED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY UPON INVESTIGATION COMPLETION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22374243&pc=OZD | 1220648-2025-29958 | 2025-06-15 04:00 | Death | ABIOMED, INC. | 2025-07-01 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Ischemia; Sepsis | P140003 |  | 1 | Event Description: A COMPLAINANT REPORTED THAT A PATIENT WITH ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK WAS PLACED ON IMPELLA CP FOR HEMODYNAMIC SUPPORT. DURING SUPPORT IT WAS REPORTED THAT THE PATIENT HAD EXPERIENCED LEG ISCHEMIA THAT BEGAN DURING/AFTER INTRODUCER INSERTION. A SENIOR DOCTOR CHECKED POSITION AND LEG PERFUSION(FEM.-FEM BYPASS). IT WAS NOTED THAT THE PUMP RAN WITHOUT PROBLEMS ON P-6, HOWEVER HIGHER P-LEVEL WAS NOT POSSIBLE DUE TO SUCTION ALARMS. DISCUSSIONS OF ESCALATION TO BIPELLA WERE HAD HOWEVER, NOT AN OPTION DUE TO PATIENT OVERALL BAD CONDITIONS AND BAD PROGNOSIS. ADDITIONALLY THE PATIENT HAD A EARLY ONSET OF SEPSIS, AS A RESULT OF THAT EXTRAVASAL VOLUME SHIFT AND SUCTION ALARMS. ULTIMATELY CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22375004&pc=JPA | 2250033-2025-00001 | 2025-06-04 04:00 | Death | ACCRIVA DIAGNOSTICS, INC. | 2025-07-01 04:00 | JPA | HEMOCHRON WHOLE BLOOD MICROCOAGULATION SYSTEM - HEMOCHRON SIGNATURE ELITE INSTR | Insufficient Information | Insufficient Information | K193041 |  | 1 | Event Description: HEALTHCARE PROFESSIONAL REPORTED THAT A HEMOCHRON SIGNATURE ELITE ANALYZER WAS USED ON SEVERAL PATIENT CASES WHERE THEY OBSERVED UNEXPECTED RESULTS. THESE CASES INVOLVED PATIENTS WHO WENT INTO ARREST. ONE PATIENT COULD NOT BE RESUSCITATED, RESULTING IN DEATH. NO ADDITIONAL DETAILS OR PATIENT INFORMATION ARE AVAILABLE AT THIS TIME.Event Description: HEALTHCARE PROFESSIONAL REPORTED THAT A HEMOCHRON SIGNATURE ELITE ANALYZER WAS USED ON SEVERAL PATIENT CASES WHERE THEY OBSERVED UNEXPECTED RESULTS. THESE CASES INVOLVED PATIENTS WHO WENT INTO ARREST. ONE PATIENT COULD NOT BE RESUSCITATED, RESULTING IN DEATH. THE NUMBER OF PATIENTS INVOLVED WAS NOT SPECIFIED. THE PATIENT WHO PASSED AWAY WAS UNDERGOING A CORONARY INTERVENTION PROCEDURE. NO ERROR CODES WERE SEEN FROM THE INSTRUMENT, AND ELECTRONIC QUALITY CONTROL (EQC) AND LIQUID QUALITY CONTROL (LQC) TESTING PASSED. NO OTHER MEDICAL OR LABORATORY INFORMATION WAS PROVIDED. Manufacturer Narrative: THIS MDR REFERENCES ACCRIVA DIAGNOSTICS' COMPLAINT NUMBER (B)(4) FOR THE HEMOCHRON WHOLE BLOOD MICROCOAGULATION SYSTEM - HEMOCHRON SIGNATURE ELITE INSTRUMENT (B)(6). THE INSTRUMENT HAS NO EXPIRATION DATE. ANALYSIS OF PRODUCTION RECORDS DID NOT PROVIDE ANY INFORMATION RELEVANT TO THIS COMPLAINT. THE DHR FOR ELITE INSTRUMENT (B)(6) WAS REVIEWED AND ANALYZED FOR POTENTIAL NON-CONFORMITIES, OR ANOMALIES, WHICH COULD HAVE LED TO THIS COMPLAINT ALLEGATION. REVIEW AND ANALYSIS INDICATED NO NON-CONFORMITIES OR ANOMALIES THAT CAN BE A FACTOR OR A POTENTIAL ROOT CAUSE TO THE AFOREMENTIONED ALLEGATION. THE PRODUCT INITIAL RELEASE TESTING WAS FOUND TO MEET ALL SPECIFICATIONS. ACCRIVA DIAGNOSTICS, INC. HAS REQUESTED FURTHER INFORMATION FROM THE REPORTER AND IS AWAITING A RESPONSE. Manufacturer Narrative: FOLLOW-UP 1: FOLLOWING THE INITIAL MDR REPORT, HEMOCHRON SIGNATURE ELITE INSTRUMENT SE12657 WAS RETURNED TO ACCRIVA FOR EVALUATION. THE INSTRUMENT WAS TESTED AND MET ALL TEST CRITERIA. NO ERRORS OCCURRED DURING TESTING, AND THE COMPLAINT COULD NOT BE REPRODUCED. IN ADDITION, A REVIEW OF IN-HOUSE TESTING PERFORMED ON RETAINS OF THE ASSOCIATED CUVETTE LOTS SHOWED THAT THE CUVETTES PERFORMED AS EXPECTED BASED ON ESTABLISHED SPECIFICATIONS: CUVETTE LOT A5JLR015 TESTED UNDER LSR 2025-CSS-058, AND CUVETTE LOT A5JLR011 TESTED UNDER LSR 2025-CSS-050. BASED ON THE INVESTIGATION FINDINGS AND THE INFORMATION AVAILABLE, THE ADVERSE EVENT PROBLEM CODES HAVE BEEN UPDATED AS FOLLOWS: HEALTH EFFECT CLINICAL CODE (E): 4580 - INSUFFICIENT INFORMATION. HEALTH EFFECT IMPACT CODE (F): 1802 - DEATH. MEDICAL DEVICE PROBLEM CODE (A): 3190 - INSUFFICIENT INFORMATION. COMPONENT CODE (G): 593 - ANALYZER. TYPE OF INVESTIGATION (B): 10 - TESTING OF ACTUAL/SUSPECTED DEVICE; 3331 - ANALYSIS OF PRODUCTION RECORDS. INVESTIGATION FINDINGS (C): 213 - NO DEVICE PROBLEM FOUND. INVESTIGATION CONCLUSIONS (D): 4315 - CAUSE NOT ESTABLISHED. THE NUMBER OF PATIENTS AFFECTED REMAINS UNCONFIRMED AS THE INITIAL REPORTER HAS NOT RESPONDED TO ACCRIVA'S REQUEST FOR THIS INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22375075&pc=HCG | 3008881809-2025-00369 | 2025-04-21 04:00 | Death | STRYKER NEUROVASCULAR CORK | 2025-07-01 04:00 | HCG | TARGET HELICAL ULTRA 2.5MM X 6CM | Fracture | No Clinical Signs, Symptoms or Conditions | K242243 |  | 1 | Event Description: IT WAS REPORTED THAT DURING INCIDENTAL CEREBRAL ANEURYSM PROCEDURE, THE OPERATOR TRIED TO EXTRACT THE SUBJECT COIL ONCE THE RELEASE PROCESS HAD STARTED. THE SUBJECT COIL COULD NOT BE REMOVED FROM THE PATIENT'S ANATOMY, WHICH CAUSED PATIENT'S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22375267&pc=NHL | 3006630150-2025-05031 | 2025-05-03 04:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-07-01 04:00 | NHL | VERCISE GENUS? | Adverse Event Without Identified Device or Use Problem | Fall | P150031 |  | 1 | Event Description: IT WAS REPORTED THAT THE DEEP BRAIN STIMULATION (DBS) PATIENT PASSED AWAY AFTER EXPERIENCING A FALL. THIS EVENT WAS REPORTED AS NOT RELATED TO THE PROCEDURE. THE RELATIONSHIP TO THE DEVICE HARDWARE OR DEVICE STIMULATION WAS NOT REPORTED AND IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22375345&pc=DSQ | 2916596-2025-04082 | 2025-05-30 04:00 | Death | THORATEC CORPORATION | 2025-07-01 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED; HOWEVER, IT WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. DUE TO PRIVACY REASONS THE HOSPITAL WOULD NOT PROVIDE THE CAUSE OF DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER MLP-051174, AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO INFORMATION REGARDING THE PATIENT¿S EXPIRATION WILL BE PROVIDED DUE TO PRIVACY REASONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR MLP-051174 WERE REVIEWED THROUGH THE MANUFACTURING EXECUTION SYSTEM (MES) AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22375385&pc=OZD | 1220648-2025-29968 | 2025-06-13 04:00 | Death | ABIOMED, INC. | 2025-07-01 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Stroke/CVA; Hematoma | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT IMPLANTED WITH AN IMPELLA 5.5 SUPPORT EXPERIENCED SINUS TACHYCARDIA REQUIRING DEFIBRILLATION. ADDITIONALLY, A STROKE CODE WAS CALLED THOUGH A CT WAS PERFORMED AND SHOWED NO SIGNS OF BLEEDING. IT WAS FURTHER REPORTED THAT THE PATIENT HAD HEMATOMA NEAR THE IMPELLA ACCESS SITE AND THEREFORE, HEPARIN WAS WITHHELD. THE PATIENT'S FAMILY DECIDED TO WITHDRAW CARE, AND THE PATIENT EXPIRED. Manufacturer Narrative: A4 AND A5, RACE, ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22375556&pc=MVK | 3008642652-2025-06390 | 2025-05-21 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-01 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED 8 APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 18:12:06 ON (B)(6) 2025. AT 12:48:11 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 12:48:47, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM. AT 12:50:54, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VT @ 250 BPM. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 60 BPM WITH PAC¿S. AT 12:52:34, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF THE TREATMENT WAS VT @ 250 BPM. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 60 BPM WITH PAC¿S. AT 12:54:46, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF THE TREATMENT WAS VT @ 250 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 50 BPM WITH PAC¿S. AT 12:56:23, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF THE TREATMENT WAS VT @ 250 BPM. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 70 BPM WITH PACS/PVCS. AT 12:58:45, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF THE TREATMENT WAS VT @ 260 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 20 BPM. AT 12:59:19, THE PATIENT RECEIVED THE SEVENTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF THE TREATMENT WAS VT @ 260 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA FROM 10-30 BPM WITH HB/UNCONDUCTED P WAVES AND PVCS. AT 13:04:40, THE PATIENT RECEIVED THE EIGHTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF THE TREATMENT WAS VT @ 260 BPM WITH MOTION ARTIFACT AND ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS SINUS RHYTHM @70 BPM WITH PACS. AT 13:06:07 AND 13:07:00, TWO ARRHYTHMIAS WERE DETECTED. ECG SHOWS VT @ 260 BPM WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. THE DEVICE PROPERLY DETECTED VT. CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF DELAYED AND PREVENTED THE LIFEVEST FROM TREATING THE PATIENT UNTIL THE ELECTRODE BELT WAS DISCONNECTED. THE ELECTRODE BELT WAS DISCONNECTED AT 13:07:30 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22375626&pc=NPS | 2135147-2025-03579 | 2015-08-01 04:00 | Death | ABBOTT MEDICAL | 2025-07-01 04:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "VENTRICULAR INTERACTION AND CLINICAL OUTCOMES IN PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE ASSOCIATION BETWEEN VENTRICULAR INTERACTION AND CLINICAL OUTCOMES IN PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR (TTVR). DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP, PASCAL, CARDIOBAND. THE ARTICLE CONCLUDED THAT VENTRICULAR INTERACTION, ASSESSED BY LV ECCENTRICITY INDEX, WAS ASSOCIATED WITH CLINICAL OUTCOMES AFTER TTVR, AND ITS POSTPROCEDURAL IMPROVEMENT WAS LINKED TO BETTER OUTCOMES. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS DR TETSU TANAKA AT UNIVERSITY HOSPITAL BONN, BONN, GERMANY WITH CORRESPONDING EMAIL TA.CHI.TSU.TETSU@GMAIL.COM. THE TIME FRAME OF THE STUDY WAS AUGUST 2015 TO AUGUST 2022. A TOTAL OF 304 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 208 (76%) RECEIVED AN ABBOTT DEVICE. AVERAGE AGE WAS 79 AND MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED FUNCTIONAL AND DEGENERATIVE TRICUSPID REGURGITATION, DIABETES, HYPERTENSION, COPD, CORONARY ARTERY DISEASE, PRIOR MYOCARDIAL INFARCTION, PRIOR CARDIAC SURGERY, ATRIAL FIBRILLATION, HEART FAILURE, CIED, HEART FAILURE MEDICATION, MITRAL REGURGITATION. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED. TRICLIP: DEATH, HEART FAILURE, HOSPITALIZATION, UNCHANGED TR, RECURRENT TR. MITRACLIP: DEATH, HEART FAILURE, HOSPITALIZATION, UNCHANGED TR, RECURRENT TR, OFF-LABEL USE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING FUNCTIONAL AND DEGENERATIVE TRICUSPID REGURGITATION, DIABETES, HYPERTENSION, COPD, CORONARY ARTERY DISEASE, PRIOR MYOCARDIAL INFARCTION, PRIOR CARDIAC SURGERY, ATRIAL FIBRILLATION, HEART FAILURE, CIED, HEART FAILURE MEDICATION, MITRAL REGURGITATION. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE, HOSPITALIZATION, UNCHANGED TR, RECURRENT TR, THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: "VENTRICULAR INTERACTION AND CLINICAL OUTCOMES IN PATIENTS UNDERGOING TRANSCATHETER TRICUSPID VALVE REPAIR".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22375690&pc=NGV | 2135147-2025-03582 | 2019-05-01 04:00 | Death | ABBOTT MEDICAL | 2025-07-01 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P200049 |  | 1 | Event Description: THE ARTICLE "CONCOMITANT TRANSCATHETER EDGE-TO-EDGE REPAIR AND LEFT ATRIAL APPENDAGE OCCLUSION" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE SAFETY AND FEASIBILITY OF CONCOMITANT MITRAL TRANSCATHETER EDGE TO EDGE REPAIR (M-TEER) AND LEFT ATRIAL APPENDAGE OCCLUSION (LAAO). DEVICES MENTIONED INCLUDE MITRACLIP, AMPLATZER AMULET, AMPLATZER TORQVUE 45X45 DELIVERY SYSTEM. THE ARTICLE CONCLUDED THAT CONCOMITANT LAAO WITH M-TEER IS AN ATTRACTIVE OPTION FOR PATIENTS WITH INDICATIONS FOR BOTH PROCEDURES TO SIMPLIFY RECOVERY TIME, REDUCE COST AND MITIGATE THE CUMULATIVE RISKS ASSOCIATED WITH VASCULAR ACCESS AND TRANSSEPTAL PUNCTURE WHEN PERFORMED ON SEPARATE OCCASIONS. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS OMAR ABDEL-RAZEK AT UNIVERSITY OF OTTAWA WITH CORRESPONDING EMAIL *OABDEL-RAZEK@OTTAWAHEART.CA;*. THE TIME FRAME OF THE STUDY WAS MAY 2019 AND SEPTEMBER 2024. A TOTAL OF 15 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. AVERAGE WAS 80 YEARS AND MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDE PRIOR MYOCARDIAL INFARCTION, DIABETES, HYPERTENSION, DYSLIPIDEMIA, SMOKING, PRIOR STROKE, PERIPHERAL ARTERIAL DISEASE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, DEGENERATIVE AND FUNCTIONAL MITRAL REGURGITATION, ATRIAL FIBRILLATION, HEART FAILURE. (B)(6) - COMPLICATIONS INCLUDE DEATH, MAJOR BLEEDING, PERICARDIAL EFFUSION, PERICARDIOCENTESIS, HOSPITALIZATION, RECURRENT MITRAL REGURGITATION. (B)(6) - COMPLICATIONS INCLUDE DEATH, MAJOR BLEEDING, PERICARDIAL EFFUSION, PERICARDIOCENTESIS, HOSPITALIZATION. (B)(6) - ATRIAL SEPTAL DEFECT WITH INTERVENTION, PERICARDIAL EFFUSION, PERICARDIOCENTESIS. (B)(6)- ATRIAL SEPTAL DEFECT WITH INTERVENTION, PERICARDIAL EFFUSION, PERICARDIOCENTESIS. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF AMULET PROCEDURES WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PRIOR MYOCARDIAL INFARCTION, DIABETES, HYPERTENSION, DYSLIPIDEMIA, SMOKING, PRIOR STROKE, PERIPHERAL ARTERIAL DISEASE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, DEGENERATIVE AND FUNCTIONAL MITRAL REGURGITATION, ATRIAL FIBRILLATION, HEART FAILURE. COMPLICATIONS REPORTED INCLUDED COMPLICATIONS INCLUDE DEATH, MAJOR BLEEDING, PERICARDIAL EFFUSION, PERICARDIOCENTESIS, HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2 DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: "CONCOMITANT TRANSCATHETER EDGE-TO-EDGE REPAIR AND LEFT ATRIAL APPENDAGE OCCLUSION".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22376030&pc=PYX | 1220648-2025-29975 | 2025-06-14 04:00 | Death | ABIOMED, INC. | 2025-07-01 04:00 | PYX | IMPELLA RP | Fluid/Blood Leak | Multiple Organ Failure | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT (PT) WAS ADMITTED FOR UNEXPLAINED ANEMIA, WEIGHT LOSS, AND ABDOMINAL PAIN. CARDIAC EVAL/TTE REVEALED VEGETATION ON THE AORTIC VALVE LEAFLETS. EF 45%. ADDITIONALLY, PT WAS DIAGNOSED WITH ACUTE CHOLECYSTITIS. ABX TREATMENT STARTED AND PT OPTIMIZED FOR SURGERY. THE PATIENT UNDERWENT BENTALL. SURGERY WAS COMPLICATED BY COMPLEX REMOVAL OF VALVE AND FRIABLE TISSUE. DECISION WAS MADE FOR URGENT PLACEMENT OF IMPELLA 5.5 AS THEY WERE UNABLE TO WEAN FROM BYPASS. ON ARRIVAL IN THE OR, THE IMPELLA 5.5 WAS INSERTED AND INITIATED AT P2. UPON ASSESSMENT OF HEMODYNAMICS AND HEART FUNCTION, PHYSICIAN MADE THE DECISION TO PROCEED WITH INSERTION OF RP FLEX. DURING SUPPORT, A NURSE CALLED ABOUT PLACEMENT SIGNAL ALARM WITH THE RP. ALARM EXPLAINED AND AUDIO WAS SILENCED. ADDITIONALLY, THERE WAS A NOTIFICATION WITH THE RP OF PURGE PRESSURE LOW. THE PURGE FLOW 30, PP IN 300'S. HOWEVER, WAS INTERMITTENT AFTER BAG CHANGE. NO LEAKS NOTED. IT WAS NOTED THAT THE PATIENT ON BIPELLA SUPPORT WAS CRITICALLY ILL WITH POOR PROGNOSIS, ABSENT OF ANY NEUROLOGICAL FUNCTION, ON MULTIPLE HIGH-DOSE PRESSORS/INOTROPES WITH ACUTE-ONSET MULTIPLE ORGAN DYSFUNCTION SYNDROME. FAMILY MADE DECISION TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA RP SYSTEM WITH AUTOMATED IMPELLA CONTROLLER AND IMPELLA RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER INSTRUCTIONS FOR USE AND CLINICAL REFERENCE MANUAL. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ARRHYTHMIA, ATRIAL FIBRILLATION, BLEEDING, CARDIAC TAMPONADE, CARDIOGENIC SHOCK, DEATH, DEVICE MALFUNCTION, HEMOLYSIS, HEPATIC FAILURE, INSERTION SITE INFECTION, PERFORATION, PHLEGMASIA CERULEA DOLENS (A SEVERE FORM OF DEEP VENOUS THROMBOSIS), PULMONARY VALVE INSUFFICIENCY, RESPIRATORY DYSFUNCTION, SEPSIS, THROMBOCYTOPENIA, THROMBOTIC VASCULAR (NON-CENTRAL NERVOUS SYSTEM COMPLICATION, TRICUSPID VALVE INJURY, CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION), VENOUS THROMBOSIS, VENTRICULAR FIBRILLATION AND/OR TACHYCARDIA."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22376494&pc=OZD | 1220648-2025-29982 | 2025-06-14 04:00 | Death | ABIOMED, INC. | 2025-07-01 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT, THE PATIENT DEVELOPED BLEEDING AT THE ACCESS SITE, WHICH REQUIRED MEDICAL INTERVENTION. THE PATIENT HAD A HISTORY OF ST-ELEVATION MYOCARDIAL INFARCTION (STEMI) AND HAD PRESENTED WITH CHEST PAIN, NAUSEA, AND VOMITING. THE DECISION WAS MADE TO IMPLANT AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. A MATTRESS SUTURE WAS PLACED BY THE PHYSICIAN TO ADDRESS MINOR GROIN SITE OOZING. FOLLOWING CLINICAL DETERIORATION, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS, ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER, ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS, ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION)." Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF ACCESS SITE BLEEDING IS DETERMINED TO BE USE ISSUE BECAUSE HEMOSTASIS WAS ACHIEVED BY PLACING MATTRESS SUTURES AT THE SITE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22376533&pc=LWS | 2649622-2025-17800 | 2025-06-25 04:00 | Death | MPRI | 2025-07-01 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing; Under-Sensing | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AMYLOIDOSIS, WHICH ULTIMATELY LED TO HEART FAILURE. A REVIEW OF THE DEVICE DATA BY QUALIFIED PERSONNEL REVEALED THAT THE EPISODE APPEARED TO HAVE TERMINATED DUE TO UNDERSENSING OF VERY FINE, LOW AMPLITUDE VENTRICULAR FIBRILLATION (VF) THAT FELL OUT OF THE DEVICE'S DETECTION ZONE. IT WAS FURTHER NOTED THAT A LEAD INTEGRITY ALERT (LIA) TRIGGERED ON THE RIGHT VENTRICULAR (RV) LEAD DUE TO HIGH-RATE NON-SUSTAINED EPISODES AND SHORT V-V INTERVALS, AND THERE WERE MULTIPLE FAILED SHOCKS OBSERVED DURING THE EPISODE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22376636&pc=DSQ | 2916596-2025-04120 | 2025-06-13 04:00 | Death | THORATEC CORPORATION | 2025-07-01 04:00 | DSQ | HEARTMATE II LEFT VENTRICULAR ASSIST DEVICE | Adverse Event Without Identified Device or Use Problem | Fall; Intracranial Hemorrhage; Hemorrhagic Stroke | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO HEMORRHAGIC STROKE FROM INTRACRANIAL BLEEDING. THIS WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE PATIENT WAS ADMITTED TO THE EMERGENCY ROOM DUE TO HEAD TRAUMA SUSTAINED FROM A FALL AND PASSED AWAY WITHIN A FEW HOURS OF ADMISSION DUE TO MASSIVE HEMORRHAGIC STROKE. THIS WAS NOT PUMP RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. THE CAUSE OF THE FALL WAS NOT DETERMINED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (HMII LVAS) INSTRUCTIONS FOR USE (IFU), REV. B, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND STROKE, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, UNDER ¿ANTICOAGULATION¿, PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22377176&pc=DTB | 2649622-2025-17808 | 2025-06-09 04:00 | Death | MPRI | 2025-07-01 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Erythema; Fever; Pain; Cardiac Tamponade | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY FOUR MONTHS POST CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) CHANGEOUT PROCEDURE THE PATIENT DEVELOPED A POCKET INFECTION WITH NOTED FEVER, PAIN AND REDNESS. THE CRT-D SYSTEM WAS TO BE EXPLANTED HOWEVER DURING THE EXTRACTION PROCESS OF THE RIGHT ATRIAL (RA) LEAD THE PATIENT EXPERIENCED PERICARDIAL TAMPONADE AND DIED. Manufacturer Narrative: CONTINUATION OF D10: DTPA2QQ CRT-D IMPLANTED (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22377313&pc=DSQ | 2916596-2025-04059 | 2025-05-24 04:00 | Death | THORATEC CORPORATION | 2025-07-01 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Tricuspid Valve Insufficiency/ Regurgitation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO COMPLICATIONS POST OPERATIVELY. IT WAS NOTED THAT THE PATIENT SUFFERED FROM MODERATE TRICUSPID VALVE (TV) ISSUES/REGURGITATION PRIOR TO IMPLANT. THE PATIENT HAD THEIR TRICUSPID VALVE REPLACED POST-OP. IT WAS UNABLE TO BE DETERMINED IF LVAD THERAPY CONTRIBUTED TOO OR WORSENED THE TV ISSUE. THE HEALTHCARE PROVIDERS THOUGHT THEY MAY HAVE MADE THE WRONG DECISION NOT REPLACING THE VALVE DURING THE LVAD IMPLANT. THE PATIENT SUFFERED FROM VASOPLEGIA, AND EVENTUALLY PASSED AWAY DUE TO WORSENING HEMODYNAMIC SITUATION. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND THE DEVICE WOULD NOT RETURN. Manufacturer Narrative: A. PATIENT INFORMATION WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION D5 - OPERATOR OF DEVICE: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 5, ¿SURGICAL PROCEDURES¿, EXPLAINS THAT MODERATE TO SEVERE AORTIC INSUFFICIENCY MUST BE CORRECTED AT THE TIME OF DEVICE IMPLANT. THIS SECTION ALSO STATES THAT IF THE AORTIC INSUFFICIENCY IS NOT ADDRESSED, THE DEVICE WILL NOT BE ABLE TO PROVIDE THE INTENDED FLOW. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22378523&pc=NPT | 2025587-2025-04698 | 2024-08-31 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-01 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS AND 16 DAYS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO CARDIOPULMONARY ARREST. IT WAS NOTED THAT LIFE SAVING TREATMENT WAS PERFORMED BUT THE PATIENT ULTIMATELY DIED.Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS AND 16 DAYS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO CARDIOPULMONARY ARREST. IT WAS NOTED THAT LIFE SAVING TREATMENT WAS PERFORMED BUT THE PATIENT ULTIMATELY DIED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE CARDIOPULMONARY RESUSCITATION WAS PERFORMED PRIOR TO THE PATIENT'S DEATH. PER THE PHYSICIAN, THE DELIVERY CATHETER SYSTEM CAN BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE CAUSE OF DEATH. Manufacturer Narrative: UPDATED DATA: B5. ADDED SECOND PARAGRAPH. D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22378544&pc=NPT | 2025587-2025-04699 | 2025-06-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-01 04:00 | NPT | EVOLUT FX DCS | Failure to Advance; Patient Device Interaction Problem | Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE TRANSCATHETER AORTIC VALVE EVFXPLUS-23, A PATIENT WITH A HISTORY OF PREVIOUS COARCTATION SURGERY AND A FRAGILE AORTA EXPERIENCED SIGNIFICANT PROCEDURAL DIFFICULTIES. THE SYSTEM WOULD NOT TRACK AROUND THE AORTIC ARCH DESPITE VARIOUS TECHNIQUES, INCLUDING THE USE OF A STIFFER WIRE AND BALLOON, AND ULTIMATELY COULD NOT BE ADVANCED. A DIFFERENT MANUFACTURER'S VALVE (23 SAPIEN) WAS IMPLANTED INSTEAD. COMPUTED TOMOGRAPHY WAS PERFORMED AND IDENTIFIED A TEAR IN THE AORTA AT THE LEFT COMMON CAROTID. THE DISSECTION WAS LIKELY CAUSED BY PUSHING THE DEVICE AROUND THE ARCH. THE PATIENT DIED AS A RESULT OF AORTIC DISSECTION.Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE TRANSCATHETER AORTIC VALVE EVFXPLUS-23, A PATIENT WITH A HISTORY OF PREVIOUS COARCTATION SURGERY AND A FRAGILE AORTA EXPERIENCED SIGNIFICANT PROCEDURAL DIFFICULTIES. THE SYSTEM WOULD NOT TRACK AROUND THE AORTIC ARCH DESPITE VARIOUS TECHNIQUES, INCLUDING THE USE OF A STIFFER WIRE AND BALLOON, AND ULTIMATELY COULD NOT BE ADVANCED. A DIFFERENT MANUFACTURER'S VALVE (23 SAPIEN) WAS IMPLANTED INSTEAD. COMPUTED TOMOGRAPHY WAS PERFORMED AND IDENTIFIED A TEAR IN THE AORTA AT THE LEFT COMMON CAROTID. THE DISSECTION WAS LIKELY CAUSED BY PUSHING THE DEVICE AROUND THE ARCH. THE PATIENT DIED AS A RESULT OF AORTIC DISSECTION. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE AORTIC DISSECTION OCCURRED DURING DEVICE ADVANCEMENT; HOWEVER, A SPECIFIC DEVICE WAS NOT MENTIONED. NO TREATMENT WAS PROVIDED TO ADDRESS THE AORTIC DISSECTION. ACCORDING TO THE PHYSICIAN, THERE WAS NOTHING IDENTIFIED ON THE 3MENSIO IMAGING TO INDICATE THE PATIENT'S ANATOMY CONTRIBUTED TO THE DISSECTION; HOWEVER, IT WAS NOTED THE TWO PREVIOUS CARDIAC SURGERIES COULD ACCOUNT FOR TISSUE FRAGILITY. IT WAS FURTHER NOTED THE PATIENT'S UNDERLYING MEDICAL HISTORY DID NOT CONTRIBUTE TO THE PATIENT'S DEATH. THE MEDTRONIC VALVE COULD BE EXCLUDED AS A CONTRIBUTING FACTOR AS IT WAS NEVER DEPLOYED. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012768942); PRODUCT TYPE: 0195-HEART VALVES. SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT FX PLUS VALVE; PRODUCT ID: EVFXPLUS-23 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. D10. THE INITIAL MEDWATCH WAS SUBMITTED WITH SECTION H11 NOTING THE VALVE LOADED IN THE SUSPECT DCS WAS A SYSTEM REPORTED DEVICE; HOWEVER, ADDITIONAL INFORMATION WAS RECEIVED THAT CLEARLY EXCLUDES THE VALVE AS A SYSTEM REPORTED DEVICE. INSTEAD, LET THIS REPORT REFLECT THE VALVE IS A CONCOMITANT DEVICE LOADED IN THE DCS WHEN THE DCS WAS ADVANCED THROUGH THE AORTA AND CONTRIBUTED TO THE AORTIC DISSECTION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22378763&pc=NPT | 2025587-2025-04703 | 2025-06-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-01 04:00 | NPT | EVOLUT FX PLUS VALVE | Adverse Event Without Identified Device or Use Problem; Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A TRANSCATHETER AORTIC VALVE REPLACEMENT IMPLANT WITH A TRANSCATHETER AORTIC VALVE EVFXPLUS-26, A PATIENT WITH A MEDICAL HISTORY OF RENAL IMPAIRMENT, SEVERE AORTIC VALVE STENOSIS, AND CARDIOMYOPATHY ARRIVED AT THE CRITICAL CARE UNIT. THE PATIENT EXPERIENCED A SUDDEN ONSET OF NAUSEA AND BECOMING UNWELL SUDDENLY. THIS LED TO A CARDIAC ARREST, REQUIRING RESUSCITATION. DURING RESUSCITATION, THE PATIENT WAS RETURNED TO THE CATHETER ROOM FOR X-RAY. THE PATIENT SUBSEQUENTLY DIED, WITH THE CAUSE OF DEATH UNKNOWN.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT AN AUTOPSY WAS PERFORMED BUT NO CAUSE OF DEATH WAS FOUND SO THE CAUSE OF DEATH REMAINS UNKNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22379275&pc=BTT | 9611451-2025-00586 |  | Death | FISHER & PAYKEL HEALTHCARE LTD | 2025-07-01 04:00 | BTT | FISHER & PAYKEL HEALTHCARE | Adverse Event Without Identified Device or Use Problem | Aspiration/Inhalation; Cardiac Arrest | K110019 |  | 1 | Event Description: A HEALTHCARE FACILITY IN (B)(6) REPORTED VIA A FISHER & PAYKEL (F&P) HEALTHCARE FIELD REPRESENTATIVE, AN EVENT INVOLVING A FISHER & PAYKEL HEALTHCARE MR850 RESPIRATORY HUMIDIFIER. IT WAS REPORTED THAT THERE WAS CONDENSATION, CAUSING BUILD UP NEAR THE CONNECTION OF THE BREATHING CIRCUIT WITH THE ENDOTRACHEAL TUBE (ETT), AND WATER TO ENTER THE PATIENT'S ETT. THE HEALTHCARE FACILITY REPORTED THAT APPROXIMATELY 50 MINUTES AFTER THE WATER ENTERED THE ETT, THE PATIENT EXPERIENCED "FLUID OVERLOAD AND CARDIAC ARREST". IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY DECEASED. F&P HEALTHCARE HAS REQUESTED THE RETURN OF THE SUBJECT MR850 RESPIRATORY HUMIDIFIER FOR EVALUATION AND IS CURRENTLY IN THE PROCESS OF OBTAINING FURTHER INFORMATION ABOUT THE INCIDENT, INCLUDING DETAILS ON THE BRAND AND MODEL OF THE BREATHING CIRCUIT IN USE AT THE TIME OF THE REPORTED EVENT. Manufacturer Narrative: (B)(4). FISHER AND PAYKEL (F&P) HEALTHCARE IS CURRENTLY IN THE PROCESS OF OBTAINING FURTHER INFORMATION REGARDING THE REPORTED EVENT. WE WILL PROVIDE A FOLLOW-UP REPORT UPON COMPLETION OF OUR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22379525&pc=NAY | 2955842-2025-28450 | 2025-05-11 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-07-01 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Unspecified Tissue Injury | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI-ASSISTED LARGE PARAESOPHAGEAL HERNIA REPAIR, THE PATIENT SUSTAINED AN ULCER PERFORATION AND ULTIMATELY EXPIRED. THE INITIAL ROBOTIC PROCEDURE WAS COMPLETED SUCCESSFULLY WITHOUT ANY INTRAOPERATIVE COMPLICATIONS, AND THERE WERE NO ERRORS OR COMPLICATIONS RELATED TO THE DA VINCI SYSTEM. THE SURGEON REPORTED THAT THE PATIENT WAS RECOVERING AS EXPECTED, WITH PLANS FOR DISCHARGE FROM THE HOSPITAL TWO WEEKS POST-PROCEDURE. HOWEVER, THE PATIENT EXPERIENCED A PERFORATED SUPERFICIAL ULCER AND SUBSEQUENTLY PASSED AWAY. THE SURGEON BELIEVES THAT THE PERFORATION AND SUBSEQUENT COMPLICATIONS WERE NOT RELATED TO THE SURGICAL PROCEDURE, BUT WERE A RESULT OF THE PATIENT'S HISTORY OF ACUTE-ON-CHRONIC ULCERS. Manufacturer Narrative: REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS WERE LOGGED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS OCCURRED. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, TIP-UP FENESTRATED GRASPER, FENESTRATED BIPOLAR FORCEPS, PROGRASP FORCEPS, MEGA SUTURECUT NEEDLE DRIVER, VESSEL SEALER EXTEND. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THESE PRODUCTS. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE INSTRUMENTS CONFIRMED THAT THERE WERE NO RELATED NON-CONFORMANCES DOCUMENTED. A REVIEW OF THE VESSEL SEALER EXTEND LOGS FOUND THE INSTRUMENT WAS INSTALLED ONCE AND PASSED HOMING. THE CUT COMPLETE AND SEAL EVENTS WERE COMPLETED WITH NO RELEVANT ERRORS NOTED. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT IN THIS REPORT DEVELOPED A PERFORATION DIRECTLY RELATED TO AN ULCER TWO WEEKS AFTER THE INITIAL ROBOTIC PROCEDURE FOR A PARAESOPHAGEAL HERNIA REPAIR. THE PATIENT HAD A HISTORY OF CHRONIC ULCER DISEASE. THIS COMPLICATION EVENTUALLY LED TO THE PATIENT¿S DEATH. NO ALLEGATION IS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION B2 DATE OF DEATH: THE EXACT DATE OF DEATH WAS NOT PROVIDED. AN ESTIMATED DATE OF DEATH IS ENTERED THAT IS 2 WEEKS POST-PROCEDURE, PER REPORT THAT PATIENT WAS TO BE DISCHARGED TWO WEEKS POST-PROCEDURE, HOWEVER, EXPERIENCED A PERFORATED ULCER AND SUBSEQUENTLY PASSED AWAY. .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22379526&pc=EOQ | 2955842-2025-27960 | 2025-06-04 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-07-01 04:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Insufficient Information | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT CODED DURING AN ION TRANSBRONCHIAL LUNG BIOPSY OF THE OF THE LEFT UPPER LOBE (LUL). THE PROCEDURE WAS ABORTED. THE PHYSICIAN INFORMED THE INTUITIVE ENDOLUMINAL TERRITORY ASSOCIATE (ETA) AFTER THE CASE THAT THE PATIENT DID NOT RECOVER AND SUBSEQUENTLY PASSED AWAY. NO ADDITIONAL INFORMATION WAS PROVIDED; THERE WAS NO REPORT OF ANY MALFUNCTIONS WITH THE ION SYSTEM OR ANY INSTRUMENTS AND ACCESSORIES. A REVIEW OF THE PROCEDURE VIDEO ON SITE BY THE ETA FOUND NO ISSUES OR MALFUNCTIONS OF THE ION SYSTEM DURING THE PROCEDURE. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE PHYSICIAN HAVE BEEN MADE; HOWEVER, NO NEW INFORMATION WAS PROVIDED. Manufacturer Narrative: SYSTEM LOGS ARE NOT AVAILABLE FOR REVIEW. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT OF UNKNOWN AGE UNDERWENT AN ION LUNG BIOPSY FROM A LESION LOCATED IN THE LEFT UPPER LOBE. THE PATIENT CODED AND ULTIMATELY DIED. THERE WAS NO REPORTED MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. BASED ON THE AVAILABLE DATA THE DEATH WAS POSSIBLY PROCEDURE RELATED BUT NOT DEVICE RELATED. ATTEMPTS AT OBTAINING FURTHER DATA HAVE BEEN UNSUCCESSFUL. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. --FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. --FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. --KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023 --BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22622812&pc=GEI | 1717344-2025-00943 | 2025-03-27 04:00 | Death | COVIDIEN | 2025-07-28 04:00 | GEI | UNKNOWN LIGASURE INSTRUMENT | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY ANALYZED INTRA-OPERATIVE FACTORS AND OUTCOMES OF ELDERLY PATIENTS TREATED FOR HEP ATOCELLULAR CARCINOMA (HCC) BASED UPON AGE AND SURGICAL MODALITY BETWEEN JANUARY 2018 AND JUNE 2024. IT WAS NOTED THAT A DURING LIVER RESECTION, SMALL DUCTS AND LIVER PARENCHYMA DISTANT FROM THE HEPATIC PEDICLE AND MAIN TRUNK OF THE HEPATIC VEIN WERE COAGULATED AND DIVIDED USING A LIGASURE OR COMPETITOR PRODUCT. THERE WERE 252 PATIENTS INCLUDED IN THE STUDY AND COMPLICATIONS INCLUDED DEATH. ARTICLE: PROGNOSTIC ANALYSIS OF ELDERLY PATIENTS WITH HEPATOCELLULAR CARCINOMA: AN EXPLORATION AND MACHINE LEARNING MODEL PREDICTION BASED ON AGE STRATIFICATION AND SURGICAL APPROACH AUTHOR: CHIYU CAI PUBLISHED DATE: 14 APRIL 2025 PUBLICATION: JOURNAL OF HEPATOCELLULAR CARCINOMA I. Manufacturer Narrative: D10 CONCOMITANT PRODUCTS: UNKNOWN LIGASUR, UNKNOWN LIGASURE INSTRUMENT (LOT#:UNKNOWN) CHIYU CAI. "PROGNOSTIC ANALYSIS OF ELDERLY PATIENTS WITH HEPATOCELLULAR CARCINOMA: AN EXPLORATION AND MACHINE LEARNING MODEL PREDICTION BASED ON AGE STRATIFICATION AND SURGICAL APPROACH". 2025, JOURNAL OF HEPATOCELLULAR CARCINOMA MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22627979&pc=DSQ | 2916596-2025-04957 | 2025-07-16 04:00 | Death | THORATEC CORPORATION | 2025-07-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Renal Failure; Heart Failure/Congestive Heart Failure; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. THE PATIENT'S HISTORY OF HEMODIALYSIS WAS NOT A RESULT OF HEARTMATE 3 AND WAS PRIOR TO IMPLANT. THE PATIENT'S RIGHT HEART FAILURE WAS ALSO PRIOR TO IMPLANT.Event Description: THE CAUSE OF DEATH WAS SEPSIS FROM THE LINE FOR DIALYSIS. THIS WAS NOT DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. THE DEVICE WOULD NOT BE RETURNED. THE DEVICE DID NOT WORSEN THE PRE-EXISTING RIGHT HEART FAILURE AND KIDNEY DYSFUNCTION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THERE WERE NO DEVICE-RELATED ISSUES WITH (B)(6) REPORTED OR IDENTIFIED THROUGH THIS EVALUATION. IT WAS DETERMINED THAT THE REPORTED INFORMATION DID NOT MEET THE REQUIREMENTS OF A COMPLAINT; HOWEVER, THIS RECORD WILL REMAIN A COMPLAINT AS AN INITIAL MDR (MEDICAL DEVICE REPORT) HAS ALREADY BEEN SUBMITTED TO THE UNITED STATES FOOD AND DRUG ADMINISTRATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. NO SPECIFIC DEVICE-RELATED ISSUES WERE REPORTED; HOWEVER, THE HEARTMATE 3 LVAS IFU PROVIDES A LIST OF ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM IN SECTION 1, ¿INTRODUCTION¿. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22628081&pc=NPT | 2015691-2025-06080 | 2025-07-09 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-28 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation | P140031 |  | 1 | Event Description: AS REPORTED, IT WAS A CASE OF AN IMPLANT OF A 26MM SAPIEN 3 ULTRA VALVE, IN AORTIC POSITION BY TRANSFEMORAL APPROACH. THE ESHEATH+ WAS INTRODUCED THROUGH RIGHT FEMORAL ACCESS, AND THE SENTINEL CEREBRAL PROTECTION SYSTEM WAS SUCCESSFULLY DEPLOYED VIA RIGHT RADIAL ACCESS. A REFERENCE PIGTAIL CATHETER WAS NAVIGATED TO INITIATE VALVE CROSSING. FOLLOWING SUCCESSFUL CROSSING OF THE NATIVE AORTIC VALVE, A PRE-SHAPED GUIDEWIRE WAS POSITIONED IN THE LEFT VENTRICLE. THE COMMANDER DELIVERY SYSTEM, WITH THE VALVE, WAS ADVANCED THROUGH THE ESHEATH+ WITHOUT ISSUE. THE VALVE WAS ALIGNED BETWEEN RADIOGRAPHIC MARKERS, AND THE REFERENCE PIGTAIL WAS USED TO CONFIRM POSITIONING. RAPID PACING WAS INITIATED, AND THE BALLOON WAS INFLATED TO DEPLOY THE VALVE. THE VALVE WAS IMPLANTED IN A GOOD POSITION. AFTER VALVE DEPLOYMENT, THE PATIENT DEVELOPED BRADYCARDIA, FOLLOWED BY HYPOTENSION AND PROGRESSION TO ASYSTOLE. IMAGING VIA FLUOROSCOPY AND ECHOCARDIOGRAPHY REVEALED CARDIAC TAMPONADE. CARDIOPULMONARY RESUSCITATION (CPR) AND PERICARDIOCENTESIS WERE INITIATED. AUTOTRANSFUSION OF ASPIRATED PERICARDIAL BLOOD WAS PERFORMED, RESULTING IN INITIAL SIGNS OF HEMODYNAMIC RECOVERY. CONTROL ANGIOGRAPHY RULED OUT PARAVALVULAR LEAK, ANNULAR RUPTURE, AND VENTRICULAR PERFORATION. IMAGING FINDINGS SUGGESTED AN AORTIC INJURY ABOVE THE NON-CORONARY SINUS AS THE LIKELY SOURCE OF TAMPONADE. SURGICAL EXPLORATION WAS UNDERTAKEN TO IDENTIFY AND REPAIR THE INJURY. DURING THE PROCEDURE, THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION. DESPITE DEFIBRILLATION AND DIRECT CARDIAC MASSAGE, THE PATIENT REMAINED IN ASYSTOLE AND WAS SUBSEQUENTLY DECLARED DECEASED. SURGICAL FINDINGS CONFIRMED A POSTERIOR PERFORATION NEAR THE SINO-TUBULAR JUNCTION (STJ). PER MEDICAL ASSESSMENT, THE PERCEIVED ROOT CAUSE OF THE EVENT WAS STJ PERFORATION ASSOCIATED WITH VALVE DEPLOYMENT. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS INDICATE PER MEDICAL OPINION, THE PERCEIVED ROOT CAUSE OF THE EVENT WAS STJ PERFORATION ASSOCIATED WITH VALVE DEPLOYMENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22628857&pc=OZD | 1220648-2025-30743 | 2025-07-15 04:00 | Death | ABIOMED, INC. | 2025-07-28 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT HAD A POSSIBLE CEREBROVASCULAR ACCIDENT. THE STAFF WERE UNSURE WHAT TYPE OF STROKE IT WAS. CARE WAS WITHDRAWN AND PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22629173&pc=NGV | 2124215-2025-49565 | 2025-06-03 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-28 04:00 | NGV | WATCHMAN LAA CLOSURE DEVICE & DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  |  | 1 | Event Description: REPORTED VIA JOURNAL ARTICLE. IT WAS REPORTED THAT PATIENTS EXPERIENCED A CEREBRAL VASCULAR ACCIDENT. THIS STUDY AIMED TO COMPARE THE EFFICACY OF LEFT ATRIAL APPENDAGE CLOSURE PERFORMED BY EXCISION (LAAC-EE) VS. OCCLUSION (LAAC-PO) FOR STROKE PREVENTION IN PATIENTS WITH RECURRENT ATRIAL FIBRILLATION UNDERGOING RADIOFREQUENCY ABLATION. 160 CONSECUTIVE PATIENTS (109 UNDERGOING LAAC-EE AND 51 UNDERGOING LAAC-PO) WERE EVALUATED. RESULTS: A TOTAL OF 16 PATIENTS PASSED AWAY DURING THE DURATION OF THE STUDY, 7 FROM THE LAAC-EE GROUP AND 9 FROM LAAC-PO GROUP. DEATHS FROM UNKNOWN REASONS, WHERE FAMILY MEMBERS DID NOT KNOW THE CAUSE OF DEATH OR DID NOT WANT TO SPEAK ABOUT IT, WERE OBSERVED IN 7 OF THE 16 PATIENTS. DEATH FROM CEREBRAL INFARCTION WAS OBSERVED IN 3 PATIENTS AND 3 PATIENTS DIED FROM LUNG CANCER. 1 PATIENT PASSED AWAY DUE TO HAVING CONTRACTED COVID-19, ANOTHER 1 PATIENT PASSED AWAY FROM BILE DUCT CANCER AND ANOTHER 1 FROM BREAST CANCER. THERE WERE 13 PATIENTS WHICH HAD EPISODES OF POSTOPERATIVE CEREBRAL INFARCTION TIA. 2 PATIENTS SUFFERED CEREBRAL HEMORRHAGES, AND 1 PATIENT SUFFERED FROM LIMB EMBOLISM. AFTER 5 YEARS OF FOLLOW-UP, STROKE OCCURRED IN 64 PATIENTS, FROM WHICH 32 WERE IN THE EXCISION GROUP AND 32 IN THE OCCLUSION GROUP. Manufacturer Narrative: B2: DATE OF DEATH UNKNOWN. B3 DATE OF EVENT: ARTICLE PUBLISH DATE USED AS EVENT DATE IS UNKNOWN. D4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI)# AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. D6A IMPLANT DATE WAS REPORTED AS BETWEEN 2015 AND 2019. LITERATURE CITATION: ABIBE GOIA, A., XU, M., & QIUMING, H. (2025). LONG-TERM OUTCOMES OF LEFT ATRIAL APPENDAGE CLOSURE TECHNIQUES ON STROKE PREVENTION OF RECURRENT ATRIAL FIBRILLATION PATIENTS: EPICARDIAL EXCISION VERSUS PERCUTANEOUS OCCLUSION. FRONTIERS IN CARDIOVASCULAR MEDICINE, 12, 1601303.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22629695&pc=MIH | 9612164-2025-03658 | 2025-07-04 04:00 | Death | MEDTRONIC IRELAND | 2025-07-28 04:00 | MIH | ENDURANT IIS BIFURCATED STENT GRAFT | Difficult to Remove; Activation, Positioning or Separation Problem | Hemorrhage/Bleeding | P100021 |  | 1 | Event Description: AN ENDURANT II BIFURCATE STENT GRAFT WAS INTENDED TO BE IMPLANTED DURING THE ENDOVASCULAR TREATMENT OF A 60MM AAA. IT WAS REPORTED THAT DURING THE INDEX PROCEDURE, WHILE OPENING OF THE PROXIMAL SUPRARENAL RING, TWO STENT STRUTS BECAME STUCK IN THE SOCKET OF THE DELIVERY SYSTEM AND DID NOT OPEN. IT WAS CONFIRMED THE DELIVERY SYSTEM APPEAR NORMAL PRIOR TO USE. BAIL-OUT TECHNIQUES WERE APPLIED ACCORDING TO INSTRUCTIONS, RESULTING IN ONE STRUT DETACHING, BUT ONE STRUT REMAINED ATTACHED TO THE DELIVERY SYSTEM AND COULD NOT BE DETACHED AND RELEASED DESPITE REPEATED ATTEMPTS. THE PHYSICIAN INITIALLY ATTEMPTED TO RELEASE THE SUPRARENAL STENT AS DESCRIBED IN THE INSTRUCTIONS FOR USE (IFU) BY USING A HEMOSTAT TO DEPRESS THE TABS ON THE BACK-END WHEEL, REMOVING THE WHEEL, AND PUSHING THE T-TUBE. DESPITE COMPLETING THE TECHNIQUE, THE SUPRARENAL STENTS DID NOT RELEASE. THE PROCESS WAS REPEATED WITH THE SAME RESULT. ANGIOGRAPHY SUGGESTED THAT THE SUPRARENAL STENT SPIKES WERE STUCK IN THE SPINDLE. A COMPLIANT BALLOON WAS DEPLOYED IN AN ATTEMPT TO SEPARATE THE SPIKES, RESULTING IN THE RELEASE OF TWO, WHILE THE REMAINING SPIKES REMAINED ATTACHED. ADDITIONAL ATTEMPTS USING BOTH COMPLIANT AND NON-COMPLIANT BALLOONS DID NOT ACHIEVE COMPLETE RELEASE. THE PHYSICIAN ATTEMPTED AN ALTERNATIVE APPROACH BY PASSING A SNARE THROUGH THE UPPER ACCESS POINT. A SNARE WAS PASSED TO HOOK THE PROXIMAL COMPONENT OF THE DELIVERY SYSTEM AND TRACTION WAS APPLIED IN THE OPPOSITE DIRECTION TO RELEASE THE REMAINING SUPRARENAL STENTS. THIS TECHNIQUE WAS ALSO UNSUCCESSFUL. PER THE PHYSICIAN THE CAUSE OF THE EVENT IS UNDETERMINED. DUE TO THE PATIENT'S COMORBIDITIES, ENDOVASCULAR REPAIR WAS INITIALLY RECOMMENDED OVER OPEN SURGERY. HOWEVER, THE INABILITY TO REMOVE THE DELIVERY SYSTEM AND IMPLANT THE STENT GRAFT REQUIRED AN OPEN SURGICAL CONVERSION. THE PATIENT EXPIRED DURING THE OPEN PROCEDURE DUE TO BLOOD LOSS. Manufacturer Narrative: IT WAS REPORTED THAT THE INFRARENAL NECK ANGULATION AT THE LOCATION OF THE SUPRARENAL STENTS WAS 15 DEGREES. IT WAS CLARIFIED THAT TWO SUPRARENAL STENT STRUTS BECAME STUCK IN THE SPINDLE. THE BACK-END WHEEL WAS ROTATED AS FAR AS IT COULD GO TO RELEASE THE SUPRARENAL STENT BEFORE A BAILOUT WAS ATTEMPTED. THIS PROCESS WAS DOUBLE-CHECKED BY THE PHYSICIAN. DURING THE USE OF BAILOUT TECHNIQUES, ONE STENT GRAFT WAS SUCCESSFULLY RELEASED FROM THE SPINDLE, WHILE ONE STENT REMAINED ATTACHED TO IT. THE REMAINING STENT WAS RELEASED FROM THE DELIVERY SYSTEM VIA OPEN SURGERY USING SURGICAL TWEEZERS AND FORCEPS. THE OPERATING TEAM THAT PERFORMED THE OPEN SURGERY REPORTED THAT THE DELIVERY SYSTEM APPEARED INTACT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22629755&pc=NPT | 9617601-2025-00703 | 2025-07-24 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-07-28 04:00 | NPT | EVOLUT PRO PLUS VALVE | Difficult or Delayed Positioning; Activation Failure; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT VIA TRANSFEMORAL ACCESS INTO AN EXISTING 21 MM NON-MEDTRONIC SURGICAL VALVE (LABCOR), NO PRE-IMPLANT BALLOON DILATION WAS PERFORMED. THE VALVE WAS RECAPTURED TWICE DURING IMPLANTATION. AFTER RELEASE, A POST-IMPLANT BALLOON DILATION WAS PERFORMED WITH A NON-MEDTRONIC (ATLAS GOLD) NON-COMPLIANT BALLOON TO OPTIMIZE THE OUTCOME. A FINAL ARTERIOGRAPHY REVEALED PATENT CORONARY ARTERIES WITH NO SIGNS OF CORONARY INVOLVEMENT. WHILE THE PATIENT WAS BEING TRANSFERRED FROM THE CATHETERIZATION LAB TABLE, SUDDEN HYPOTENSION WAS OBSERVED. DUE TO THE HEMODYNAMIC INSTABILITY, THE PATIENT WAS REPOSITIONED ON THE TABLE, AND REPEAT ARTERIOGRAPHY REVEALED AN OCCLUSION OF THE LEFT CORONARY ARTERY (LCA). AN ATTEMPT WAS MADE TO RECANALIZE THE LCA AND ANGIOPLASTY WAS PERFORMED. DURING THIS PERIOD, WHICH LASTED APPROXIMATELY 40 MINUTES, THE PATIENT DEVELOPED CARDIOPULMONARY ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED AND CONTINUED FOR APPROXIMATELY ONE HOUR WITHOUT SUCCESS. THE PATIENT DIED AT THE END OF CPR.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE VALVE WAS RECAPTURED TWICE AND A POST-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED DUE TO UNDER-EXPANSION OF THE VALVE. PER THE PHYSICIAN, THE VALVE, DELIVERY CATHETER SYSTEM (DCS), AND GUIDEWIRE DID NOT CAUSE OR CONTRIBUTE TO THE OCCLUSION AND SUBSEQUENT DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PHYSICIAN THOUGHT THAT SEQUESTRATION (WHEN THE CUSPS OF THE NATIVE AORTIC VALVE OR A PREVIOUS PROSTHESIS ARE DISPLACED BY THE NEW VALVE) OF THE SINUSES OF VALSALVA MAY HAVE OCCURRED WHICH MAY HAVE LEAD TO ACUTE CORONARY OBSTRUCTION. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: PRO PLUS DELIVERY CATHETER SYSTEM (DCS), PRODUCT ID: D-EVPROP23-29, LOT: 0012549996, USE BY DATE: 20-NOV-2026, UDI: (B)(4). OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: L-EVPROP23-29 (LOT: 0012483160); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID: GWBC30 (LOT: 92402819); PRODUCT TYPE: 0193-GUIDEWIRE; PRODUCT ID: SENSH1828W (LOT: P2F25C0023); PRODUCT TYPE: 0199-INTRODUCER; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22630028&pc=DTB | 2649622-2025-20519 | 2025-04-28 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-28 04:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Tachycardia; Ventricular Fibrillation | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED IN A CAR ACCIDENT WHILE RIGHT VENTRICULAR (RV) LEAD WAS CONNECTED TO THE PATIENT DURING THE ACCIDENT. POST MORTEM INTERROGATION SHOWED A 8 SEC POLYMORPHIC VENTRICULAR TACHYCARDIA (VT)/VENTRICULAR FIBRILLATION (VF) EPISODE CORRESPONDING IN TIME WITH THE ACCIDENT. CUSTOMER BELIEVED INITIAL ARRHYTHMIA CAUSED THE ACCIDENT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE FULL LEAD WAS RETURNED AND ANALYZED. THE ANALYSIS INDICATED THAT THE OUTER INSULATION OF THE LEAD DEVELOPED A BREACH DUE TO METAL ION OXIDATION WHILE IN VIVO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22630052&pc=BZD | 2518422-2025-108270 | 2025-07-10 04:00 | Death | RESPIRONICS, INC. | 2025-07-28 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K113068 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO AN ALLEGATION OF A DEATH OF A PATIENT. IT IS NOT YET CLEAR IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH OF THE PATIENT. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22630389&pc=DSQ | 3007042319-2025-01523 | 2025-07-22 04:00 | Death | HEARTWARE, INC. | 2025-07-28 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT EXPIRED. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE. LOG FILE ANALYSIS WAS NOT PERFORMED SINCE LOG FILES WERE NOT AVAILABLE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH IS UNKNOWN. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF SYNCOPE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22630746&pc=DSQ | 3007042319-2025-01525 | 2024-12-21 05:00 | Death | HEARTWARE, INC. | 2025-07-28 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT PASSED AWAY IN HOSPICE CARE. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS COULD NOT BE PERFORMED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS NOT PROVIDED. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS FOR THIS PATIENT, IT WAS NOTED THAT THE PATIENT HAD A HISTORY OF THROMBUS AND CARDIAC ARRHYTHMIA. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22630840&pc=OZD | 1220648-2025-30759 | 2024-09-27 04:00 | Death | ABIOMED, INC. | 2025-07-28 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Renal Failure; Thrombocytopenia | P140003 |  | 1 | Event Description: ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REPORTED THAT A PATIENT WAS IMPLANTED WITH IMPELLA 5.5 FOR HIGH-RISK SURGERY ON (B)(6) 2024. IT WAS REPORTED THAT THE "PATIENT ENDURED THROMBOCYTOPENIA WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA DEVICE USED: ¿PATIENT DEVELOPED WORSENING BLOOD LOSS ANEMIA AFTER IMPELLA IMPLANT. NADIR HB: 7.2 ON (B)(6) POST-IMPELLA IMPLANT. TRANSFUSED WITH A TOTAL OF 2 PACKED RED BLOOD CELLS (PRBC), AND 4 PLATELETS DURING IMPELLA SUPPORT FROM ON (B)(6). PLAVIX, LOVENOX ON HOLD. RESUME AS ABLE." ADDITIONALLY, THE PATIENT DEVELOPED WORSENING OF ACUTE KIDNEY INJURY (AKI) WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA DEVICE USED: ¿AKI REQUIRING CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT). CREATININE TRENDED UPWARDS PRIOR TO IMPELLA IMPLANT. INITIALLY, CRRT RECOMMENDED BUT FAMILY REFUSED DIALYSIS. THE SUPPORT WAS FOR JUST OVER 3 DAYS WHEN THE PATIENT HAD CARE WITHDRAWN AND PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA PUMP WAS NOT RETURNED FOR INVESTIGATION. SHOULD NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. INSTRUCTIONS FOR USE: IMPELLA 5.5 FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22630843&pc=OZD | 1220648-2025-30760 | 2025-07-16 04:00 | Death | ABIOMED, INC. | 2025-07-28 04:00 | OZD | IMPELLA CP | Pumping Stopped | Hemorrhage/Bleeding; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AFTER AN AUTOMATED IMPELLA CONTROLLER (AIC) TO AIC TRANSFER, AN IMPLANTED IMPELLA CP PUMP FAILED TO RESUME SUPPORT. DESPITE ATTEMPTING ON FOUR DIFFERENT AICS, THE CONSOLES WOULD NOT RECOGNIZE THE IMPELLA CP DEVICE. THE PUMP WAS SUBSEQUENTLY REMOVED AND PLANS WERE MADE TO ESCALATE TO IMPELLA 5.5 SUPPORT, BUT THE PATIENT IS CURRENTLY TOO UNSTABLE FOR SURGERY. THE PATIENT WAS REPORTED TO BE IN CRITICAL CONDITION.Event Description: ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE PUMP WAS EXPLANTED AND THE PATIENT WAS COAGULOPATHIC AND TOO UNSTABLE TO GO INTO SURGERY FOR IMPELLA 5.5. THE PATIENT WAS HAVING BLEEDING ISSUES PRIOR TO THE CP IMPLANT DUE TO COMPLICATIONS OF CARDIOGENIC SHOCK. THE PATIENT EXPIRED 2-3 DAYS LATER FROM A SUBDURAL HEMORRHAGE. IT IS UNKNOWN IF THE IMPELLA CP CAUSED OR CONTRIBUTED TO THE SUBDURAL HEMORRHAGE. Manufacturer Narrative: THE IMPELLA PUMP WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿ACHIEVING AN ACT = 250 SECONDS PRIOR TO REMOVING THE DILATOR WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. MONITOR MOTOR CURRENT AND CONSIDER REPLACING IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: TROUBLESHOOTING THE PURGE SYSTEM: ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED: ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: H6 HEALTH IMPACT CLINICAL CODE AND IMPACT CODE UPDATED WITH ADDITIONAL INFORMATION. A4 PATIENT WEIGHT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22630994&pc=OZD | 1220648-2025-30765 | 2025-07-16 04:00 | Death | ABIOMED, INC | 2025-07-28 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE US COMPLAINANT HAD A PATIENT ARREST OUTSIDE IN COMMUNITY AND GIVEN CARDIOPULMONARY RESUSCITATION (CPR). THE PATIENT WAS DEFIBRILLATED AND INTUBATED AND RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED. THE PATIENT WAS TAKEN THE CARDIAC CATHETERIZATION LAB AND PERCUTANEOUS CORONARY INTERVENTION (PCI) WAS PERFORMED AND AN INTRA-AORTIC BALLOON PUMP (IABP) WAS PLACED. UNABLE TO EVALUATE NEURO STATUS. PUPILS PINPOINT AND FIXED. THE PATIENT CONTINUED TO DECOMPENSATE AND THE DECISION WAS MADE TO ESCALATE THE IABP TO IMPELLA CP. THE PATIENT WAS TRANSFERRED TO THE TERTIARY CENTER ON SUPPORT. WHILE ON SUPPORT THE PATIENT SUFFERED A GASTROINTESTINAL BLEED AND HAD BLOOD PRODUCTS INFUSED. THE FAMILY MADE CHOICE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA PUMP WAS NOT RETURNED FOR INVESTIGATION. SHOULD NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. B2: DATE OF DEATH UNKNOWN. D6B: EXPLANT DATE UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22631397&pc=OZD | 1220648-2025-30768 | 2025-07-13 04:00 | Death | ABIOMED, INC. | 2025-07-28 04:00 | OZD | IMPELLA CP | Positioning Failure; Increase in Pressure; Detachment of Device or Device Component | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE IMPELLA CP HAD PURGE SYSTEM BLOCKED/PURGE PRESSURE HIGH ALARMS. NO KINKS OR BLOCKAGES WERE OBSERVED. THE PURGE CASSETTE CHANGED WITH NO IMPROVEMENT. TISSUE PLASMINOGEN ACTIVATOR PROTOCOL WAS DONE. ADDITIONALLY, THE CARE TEAM REPOSITIONED IMPELLA CP FOR POSITIONAL ALARMS, AND TUOHY-BORST WAS BROKE. THE CATHETER WAS SECURED BY USING A STAT LOCK AT 97CM FROM 90CM. IT WAS FURTHER REPORTED THAT THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22631403&pc=OZD | 1220648-2025-30769 | 2025-07-15 04:00 | Death | ABIOMED, INC. | 2025-07-28 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT FOLLOWING IMPELLA CP IMPLANT, A MODERATE AMOUNT OF BLEEDING WAS SEEN FROM THE RIGHT GROIN / IMPELLA SHEATH. IN ADDITION, THE PATIENT DEVELOPED HEMOLYSIS. FEMOSTOP WAS APPLIED AND AN ADDITIONAL SUTURE WAS ADDED TO ACHIEVE HEMOSTASIS. BLOOD PRODUCTS WERE ALSO GIVEN TO TREAT THE BLEED AND HEMOLYSIS. THE PATIENT'S FAMILY MADE THE DECISION TO FORGO ESCALATION OF SUPPORT AND THE PATIENT WAS MADE COMFORT CARE. THE PATIENT PASSED AWAY FOLLOWING WITHDRAWAL OF CARE. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: ACCESS BLEED: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ HEMOLYSIS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22631486&pc=NGV | 2124215-2025-49988 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-28 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR CARDIAC ARREST. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22631552&pc=NPU | 2015691-2025-06101 | 2024-10-23 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-28 04:00 | NPU | EDWARDS SAPIEN 3 TRANSCATHETER HEATH VALVE | Migration or Expulsion of Device; Malposition of Device | Foreign Body Embolism | P140031 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN COLOMBIA. THROUGH REVIEW OF MEDICAL ARTICLE " FROM REPAIR TO REPLACEMENT RING DETACHMENT AND SECONDARY ANNULAR INJURY DURING TRANSCATHETER MITRAL VALVE-IN-RING PROCEDURE", CORRESPONDING AUTHOR DR. ANDRES FELIPE OCHOA-DIAZ ET AL., THE FOLLOWING EVENT WAS IDENTIFIED AS PERTAINING TO AN EDWARDS DEVICE. AS REPORTED, THIS WAS A CASE OF AN IMPLANT OF A 57-YEAR-OLD FEMALE PATIENT THAT PRESENTED WITH SHORTNESS OF BREATH AND ACUTE DECOMPENSATED HEART FAILURE. SIX MONTHS PRIOR, THE PATIENT UNDERWENT CORONARY ARTERY BYPASS GRAFTING AND SURGICAL MITRAL VALVE REPAIR WITH ANNULOPLASTY. THE PATIENT HAD AN IMPLANTED A NON-EDWARDS ANNULOPLASTY RING IN MITRAL POSITION. A TRANSEPTAL PUNCTURE AND MITRAL BALLOON VALVULOPLASTY WERE PERFORMED SUCCESSFULLY. THE VALVE WAS POSITIONED IN THE ANNULAR LOCATION BUT DURING DEPLOYMENT IT WAS OBSERVED THAT THE RING WAS PARTIALLY DEHISCENCE, PROBABLY CAUSED BY THE BALLOON VALVULOPLASTY OR COULD HAVE PREEXISTED AND GONE UNDETECTED IN THE PREPROCEDURAL IMAGING. THE RING WAS FINALLY DETHATCHED AND RESULTED IN MALPOSITION. THIS LED TO A MALPOSITION OF THE EDWARDS VALVE FOLLOWED BY A MIGRATION AND FINALLY THE VALVE EMBOLIZED TO THE LEFT ATRIUM WITH TORRENTIAL MITRAL REGURGITATION AND HEMODYNAMIC COLLAPSE OF THE PATIENT. THE PROCEDURE WAS CONVERTED TO SURGICAL, THE VALVE WAS REMOVED AND IT WAS OBSERVED AN ANNULAR TEAR WITH 90% DEHISCENCE OF THE PROSTHETIC RING, THAT WAS SOLVED. UNFORTUNATELY, THE PATIENT PASSED AWAY FORTY-EIGHT HOURS AFTER PROCEDURE DUE TO LEFT VENTRICULAR DYSFUNCTION AND VASOPLEGIC AND CARDIOGENIC SHOCK.Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN COLOMBIA. EDWARDS RECEIVED THE MEDICAL ARTICLE, ''FROM REPAIR TO REPLACEMENT RING DETACHMENT AND SECONDARY ANNULAR INJURY DURING TRANSCATHETER MITRAL VALVE-IN-RING PROCEDURE'', CORRESPONDING AUTHOR DR. ANDRES FELIPE OCHOA-DIAZ ET AL., THE FOLLOWING EVENT WAS IDENTIFIED AS PERTAINING TO AN EDWARDS DEVICE. AS REPORTED, THIS WAS A CASE OF AN IMPLANT OF A 57-YEAR-OLD FEMALE PATIENT THAT PRESENTED SHORTNESS OF BREATH AND ACUTE DECOMPENSATED HEART FAILURE. SIX MONTHS BEFORE, UNDERWENT CORONARY ARTERY BYPASS GRAFTING AND SURGICAL MITRAL VALVE REPAIR WITH ANNULOPLASTY. THE PATIENT HAD IMPLANTED A NON-EDWARDS RING IN MITRAL POSITION. A TRANSEPTAL PUNCTURE AND MITRAL BALLOON VALVULOPLASTY WERE DONE SUCCESSFULLY. THE VALVE WAS POSITIONED IN THE ANNULAR LOCATION BUT DURING DEPLOYMENT IT WAS OBSERVED THAT THE RING WAS PARTIALLY DEHISCENCE, PROBABLY CAUSED BY THE BALLOON VALVULOPLASTY OR COULD HAVE PREEXISTED AND GONE UNDETECTED IN THE PREPROCEDURAL IMAGING. THE RING WAS FINALLY DETACHED AND RESULTED IN MALPOSITION. THIS LED TO A MALPOSITION OF THE EDWARDS VALVE FOLLOWED BY A TORRENTIAL MITRAL REGURGITATION AND A MIGRATION OF THE VALVE TO THE LEFT ATRIUM WHICH LED THE PATIENT TO COLLAPSE HEMODYNAMICALLY. THE PROCEDURE WAS CONVERTED TO SURGICAL, THE VALVE WAS REMOVED AND IT WAS OBSERVED AN ANNULAR TEAR WITH 90% DEHISCENCE OF THE PROSTHETIC RING, THAT WAS SOLVED. UNFORTUNATELY, THE PATIENT PASSED AWAY FORTY-EIGHT HOURS AFTER PROCEDURE DUE TO LEFT VENTRICULAR DYSFUNCTION AND VASOPLEGIC AND CARDIOGENIC SHOCK. Manufacturer Narrative: SUPPLEMENTAL REPORT SUBMITTED TO INCLUDE ADDITIONAL INFORMATION RECEIVED THROUGH FOLLOW-UP. UPDATED D4 SERIAL NUMBER, EXPIRATION DATE, AND UDI. UPDATE D6 AND H4. Manufacturer Narrative: SUPPLEMENTAL REPORT SUBMITTED TO CORRECT B2, B5, AND H1. UPDATED H6. THE DEVICE WAS NOT RETURNED FOR EVALUATION AS IT REMAINED IMPLANTED. THEREFORE, A NO PRODUCT RETURN ENGINEERING EVALUATION WAS PERFORMED. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE COMPLAINT OF VALVE UNKNOWN REGURGITATION, WAS CONFIRMED BASED ON THE PROVIDED ARTICLE. DUE TO UNAVAILABILITY OF VALVE SERIAL NUMBER, DHR REVIEW WAS UNABLE TO BE PERFORMED TO DETERMINE IF A MANUFACTURING NON-CONFORMANCE CONTRIBUTED TO THE COMPLAINT. A REVIEW OF THE IFU REVEALED NO DEFICIENCIES. DURING THE MANUFACTURING PROCESS, ALL SAPIEN 3 VALVES ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE COMPLAINT EVENT. IT SHOULD BE NOTED THAT THE THV WAS IMPLANTED IN MITRAL POSITION INSIDE A PRE-EXISTING NON-EDWARDS RING. THE SAPIEN 3 (S3) VALVE WAS INDICATED FOR USE IN PATIENTS WITH SYMPTOMATIC HEART DISEASE DUE TO A FAILING AORTIC BIOPROSTHETIC VALVE OR A FAILING MITRAL SURGICAL BIOPROSTHETIC VALVE (STENOSED, INSUFFICIENT, OR COMBINED). THEREFORE, THIS WAS AN OFF-LABEL OPERATION. AS REPORTED, ''THE VALVE WAS POSITIONED IN THE ANNULAR LOCATION BUT DURING DEPLOYMENT IT WAS OBSERVED THAT THE RING WAS PARTIALLY DEHISCENCE [...] THE RING WAS FINALLY DETACHED AND RESULTED IN MALPOSITION. THIS LED TO A MALPOSITION OF THE EDWARDS VALVE FOLLOWED BY A TORRENTIAL MITRAL REGURGITATION AND A MIGRATION OF THE VALVE TO THE LEFT ATRIUM''. PER THE INSTRUCTIONS FOR USE (IFU), VALVE REGURGITATION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE TAVR PROCEDURE. IN THIS CASE, THE DISLOCATION OF THE PRE-EXISTING RING DURING THE THV VALVE DEPLOYMENT LED TO A MALPOSITIONED THV AND THE PROSTHETIC RING WAS TRAPPED BETWEEN THE TRANSCATHETER HEART VALVE AND THE NATIVE MITRAL ANNULUS. THE OFF-LABEL USE IN COMBINATION WITH THV MALPOSITION AND RING TRAPPED BETWEEN THV AND ANNULUS MAY CAUSE A LACK OF APPROPRIATE SEALING OF THE THV TO THE TARGET SITE OR CONTRIBUTED TO IMPROPER LEAFLET COAPTATION. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PROCEDURAL FACTORS (OFF LABEL OPERATION, THV MALPOSITION) MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. NO DEVICE OR LABELING PROBLEM WAS IDENTIFIED DURING THE EVALUATION. THEREFORE, NO FURTHER CORRECTIVE OR PREVENTIVE ACTION IS REQUIRED. PER THE INSTRUCTIONS FOR USE (IFU), VALVE MALPOSITION REQUIRING INTERVENTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO MALPOSITION, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, PRESERVED EJECTION FRACTION, SIGNIFICANT LANDING ZONE CALCIFICATION, LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE VALVE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR MALPOSITION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. THE DEVICE WAS USED IN OFF-LABEL IMPLANTATION IN THE MITRAL POSITION. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO VALVE-IN-RING IN MITRAL PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THAT PROCEDURAL FACTORS (DETACHMENT OF THE PRE-EXISTING RING) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. PER THE INSTRUCTIONS FOR USE (IFU), VALVE MIGRATION REQUIRING INTERVENTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH TRANSCATHETER VALVE REPLACEMENT (TVR). ACCORDING TO THE LITERATURE REVIEW, VALVE MIGRATION RESULTS WHEN FORCES ACTING ON THE TRANSCATHETER HEART VALVE (THV) OVERCOME THE STRENGTH OF ATTACHMENT OF THE VALVE TO THE ANNULUS. STENT VALVES ARE SUBJECTED TO ANTEGRADE EJECTION FORCES DURING SYSTOLE. LESS-THAN-SEVERE AND NON-UNIFORMLY DISTRIBUTED CALCIFICATION OF THE LEAFLETS, INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION CAN CONTRIBUTE TO VALVE MIGRATION. ADDITIONALLY, RESIDUAL OVERHANGING LEAFLETS CAN EXERT DOWNWARDS FORCE DURING DIASTOLE, CAUSING MIGRATION OF THE THV TOWARDS THE VENTRICLE. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR MIGRATION (I.E., MINIMAL LEAFLET CALCIFICATION), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. THE DEVICE WAS USED IN OFF-LABEL IMPLANTATION IN THE MITRAL POSITION. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO VALVE-IN-RING IN MITRAL PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE (DETACHMENT OF THE PRE-EXISTING RING) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. Manufacturer Narrative: THE INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE COMPLETED UPON COMPLETION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22631689&pc=NGV | 2124215-2025-50076 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-28 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR DEVICE EXPLANTATION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22631789&pc=LKN | 1722028-2025-00169 | 2025-07-04 04:00 | Death | TERUMO BCT | 2025-07-28 04:00 | LKN | SPECTRA OPTIA | Insufficient Information | Bradycardia; Cardiac Arrest; Dizziness; Convulsion, Clonic; Convulsion, Tonic; Convulsion/Seizure | K183081 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT APPROXIMATELY 6-8 MINUTES AFTER STARTING A THERAPEUTIC PLASMA EXCHANGE (TPE) FOR A RENAL TRANSPLANT PATIENT, THE PATIENT REPORTED FEELING DIZZY AND EXPERIENCED OCULAR RETROVERSION WITH GENERALIZED TONIC-CLONIC MOVEMENTS. THE OBJECTIVE OF THIS PROCEDURE WAS DESENSITIZE. THE ICU NURSE AND PHYSICIANS WERE NOTIFIED, THE PROCEDURE WAS PAUSED, AND THE SEIZURE MANAGEMENT WAS INITIATED. THE PATIENT BEGAN TO PRESENT RESPIRATORY DEPRESSION, THE ICU STAFF ARRIVED, AND POSITIVE VENTILATION WAS PERFORMED. THE PATIENT SUBSEQUENTLY DEVELOPED BRADYCARDIA AND CARDIAC ARREST WITH NO RESPONSE TO RESUSCITATION MANEUVERS. AFTER 45 MINUTES OF CPR THE PATIENT WAS DECLARED DECEASED. PER THE CUSTOMER, ON THE DAY PRIOR (B)(6) 2025 THE FIRST PROCEDURE WAS PERFORMED USING A CENTRAL VENOUS CATHETER AS ACCESS WHICH FUNCTIONED CORRECTLY AND THE PATIENT TOLERATED IT WELL. THE PLASMA APHERESIS WAS PERFORMED AFTER THE TRANSPLANT WITHOUT ANY INCONVENIENCE WITH A PERSONALIZED PRIMING. THE COLLECTION SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. THIS REPORT IS BEING FILED DUE TO PATIENT DEATH, ALTHOUGH PER CURRENT INFORMATION THERE IS NO DETECTABLE MALFUNCTION WITH THE TERUMO BCT DEVICE OR ALLEGATION OF A MALFUNCTION.Event Description: THE CUSTOMER REPORTED THAT APPROXIMATELY 6-8 MINUTES AFTER STARTING A THERAPEUTIC PLASMA EXCHANGE (TPE) FOR A RENAL TRANSPLANT PATIENT. THE PATIENT REPORTED FEELING DIZZY AND EXPERIENCED OCULAR RETROVERSION WITH GENERALIZED TONIC-CLONIC MOVEMENTS. THE OBJECTIVE OF THIS PROCEDURE WAS DESENSITIZE. THE ICU NURSE AND PHYSICIANS WERE NOTIFIED, THE PROCEDURE WAS PAUSED, AND THE SEIZURE MANAGEMENT WAS INITIATED. THE PATIENT BEGAN TO PRESENT RESPIRATORY DEPRESSION, THE ICU STAFF ARRIVED, AND POSITIVE VENTILATION WAS PERFORMED. THE PATIENT SUBSEQUENTLY DEVELOPED BRADYCARDIA AND CARDIAC ARREST WITH NO RESPONSE TO RESUSCITATION MANEUVERS. AFTER 45 MINUTES OF CPR, THE PATIENT WAS DECLARED DECEASED. PER THE CUSTOMER, ON THE DAY PRIOR (ON (B)(6) 2025), THE FIRST PROCEDURE WAS PERFORMED USING A CENTRAL VENOUS CATHETER AS ACCESS WHICH FUNCTIONED CORRECTLY AND THE PATIENT TOLERATED IT WELL. THE PLASMA APHERESIS WAS PERFORMED AFTER THE TRANSPLANT WITHOUT ANY INCONVENIENCE WITH A PERSONALIZED PRIMING. THE COLLECTION SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. THIS REPORT IS BEING FILED DUE TO PATIENT DEATH, ALTHOUGH PER CURRENT INFORMATION THERE IS NO DETECTABLE MALFUNCTION WITH THE TERUMO BCT DEVICE OR ALLEGATION OF A MALFUNCTION. Manufacturer Narrative: THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION IN H.6 AND H.11. INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. REVIEW OF THE RUN DATA FILE SHOWED THAT THE SYSTEM TO BE OPERATING AS INTENDED WITH NO SIGNALS OR ALARMS TO INDICATE UNEXPECTED DEVICE BEHAVIOR. SIGNALS IN THE RUN DATA FILE DO NOT SHOW ANY ISSUES WITH THE DEVICE. ALL SAFETY SYSTEMS REMAINED IN PLACE AND THE CORRECT ALARMS WERE RAISED. THE DEVICE WAS FOUND TO BE PERFORMING AS INTENDED. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) FOR THIS LOT SHOWED NO IRREGULARITIES DURING MANUFACTURING THAT WERE RELEVANT TO THIS ISSUE. INVESTIGATION IS IN PROCESS, A FOLLOW-UP REPORT WILL BE PROVIDED. Manufacturer Narrative: INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. REVIEW OF THE RUN DATA FILE SHOWED THAT THE SYSTEM TO BE OPERATING AS INTENDED WITH NO SIGNALS OR ALARMS TO INDICATE UNEXPECTED DEVICE BEHAVIOR. SIGNALS IN THE RUN DATA FILE DO NOT SHOW ANY ISSUES WITH THE DEVICE. ALL SAFETY SYSTEMS REMAINED IN PLACE AND THE CORRECT ALARMS WERE RAISED. THE DEVICE WAS FOUND TO BE PERFORMING AS INTENDED. INVESTIGATION IS IN PROCESS, A FOLLOW-UP REPORT WILL BE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22631800&pc=DQX | 2025587-2025-05297 | 2025-07-24 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-28 04:00 | DQX | CONFIDA BRECKER CURVE GUIDEWIRE | Difficult or Delayed Positioning; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension | K132623 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT VIA TRANSFEMORAL ACCESS INTO AN EXISTING 21 MM NON-MEDTRONIC SURGICAL VALVE (LABCOR), NO PRE-IMPLANT BALLOON DILATION WAS PERFORMED. THE VALVE WAS RECAPTURED TWICE DURING IMPLANTATION. AFTER RELEASE, A POST-IMPLANT BALLOON DILATION WAS PERFORMED WITH A NON-MEDTRONIC (ATLAS GOLD) NON-COMPLIANT BALLOON TO OPTIMIZE THE OUTCOME. A FINAL ARTERIOGRAPHY REVEALED PATENT CORONARY ARTERIES WITH NO SIGNS OF CORONARY INVOLVEMENT. WHILE THE PATIENT WAS BEING TRANSFERRED FROM THE CATHETERIZATION LAB TABLE, SUDDEN HYPOTENSION WAS OBSERVED. DUE TO THE HEMODYNAMIC INSTABILITY, THE PATIENT WAS REPOSITIONED ON THE TABLE, AND REPEAT ARTERIOGRAPHY REVEALED AN OCCLUSION OF THE LEFT CORONARY ARTERY (LCA). AN ATTEMPT WAS MADE TO RECANALIZE THE LCA AND ANGIOPLASTY WAS PERFORMED. DURING THIS PERIOD, WHICH LASTED APPROXIMATELY 40 MINUTES, THE PATIENT DEVELOPED CARDIOPULMONARY ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED AND CONTINUED FOR APPROXIMATELY ONE HOUR WITHOUT SUCCESS. THE PATIENT DIED AT THE END OF CPR.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE VALVE WAS RECAPTURED TWICE AND A POST-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED DUE TO UNDER-EXPANSION OF THE VALVE. PER THE PHYSICIAN, THE VALVE, DELIVERY CATHETER SYSTEM (DCS), AND GUIDEWIRE DID NOT CAUSE OR CONTRIBUTE TO THE OCCLUSION AND SUBSEQUENT DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PHYSICIAN THOUGHT THAT SEQUESTRATION (WHEN THE CUSPS OF THE NATIVE AORTIC VALVE OR A PREVIOUS PROSTHESIS ARE DISPLACED BY THE NEW VALVE) OF THE SINUSES OF VALSALVA MAY HAVE OCCURRED WHICH MAY HAVE LEAD TO ACUTE CORONARY OBSTRUCTION. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID EVPROPLUS-23 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025 EXPLANT DATE N/A PRODUCT ID D-EVPROP23-29 (LOT: 0012549996); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID L-EVPROP23-29 (LOT: 0012483160); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID SENSH1828W (LOT: P2F25C0023); PRODUCT TYPE: 0199-INTRODUCER; IMPLANT DATE N/A; EXPLANT DATE N/A SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22631804&pc=OZD | 1220648-2025-30772 | 2025-07-12 04:00 | Death | ABIOMED, INC. | 2025-07-28 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AFTER ONE DAY OF SUPPORT, THE PATIENT DEVELOPED HEMOLYSIS. THE STAFF WAS AWARE OF THE SITUATION AND PLANNED TO MONITOR THE CONDITION VIA FLUID STATUS, LAB DRAWS, AND POSITIONING CHECKS. SUPPORT CONTINUED FOR ONE MORE DAY, AT WHICH POINT THE FAMILY CHOSE TO PLACE THE PATIENT ON COMFORT CARE. THE PATIENT SUBSEQUENTLY EXPIRED FOLLOWING THAT DECISION. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER. ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER. DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22631852&pc=NGV | 2124215-2025-50067 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-28 04:00 | NGV | WATCHMAN FLX? PRO | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR IMPLANT EMBOLIZATION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025 AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22631868&pc=NGV | 2124215-2025-50090 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-28 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Thromboembolism |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR THROMBOEMBOLISM. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 229 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22631969&pc=DSQ | 2916596-2025-04985 | 2025-07-19 04:00 | Death | THORATEC CORPORATION | 2025-07-28 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT NEVER RECOVERED AFTER IMPLANTATION. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION", LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22632356&pc=BSZ | 9611500-2025-00415 | 2025-06-25 04:00 | Death | DRÄGERWERK AG & CO. KGAA | 2025-07-28 04:00 | BSZ | PERSEUS A500 | No Flow; Device Handling Problem | Cardiac Arrest | K133886 |  | 1 | Event Description: IT WAS REPORTED THAT THE ANAESTHETIC MACHINE IN THEATER 4 COULD NOT FILL UP THE BREATHING BAG WHEN THEY STARTED THE CASE. THE PATIENT REPORTEDLY WAS RESUSCITATED BUT PASSED AWAY.Event Description: IT WAS REPORTED THAT THE ANAESTHETIC MACHINE IN THEATER 4 COULD NOT FILL UP THE BREATHING BAG WHEN THEY STARTED THE CASE. THE PATIENT REPORTEDLY WAS RESUSCITATED BUT PASSED AWAY. Manufacturer Narrative: THE INVESTIGATION IS STILL ON-GOING. THE RESULTS WILL BE PROVIDED WITH A FOLLOW-UP REPORT. Manufacturer Narrative: THE INVESTIGATION WAS PERFORMED BASED ON THE DEVICE LOG FILE. THE CASE IN QUESTION WAS STARTED AT 16:03 USING THE EXTERNAL FRESH GAS OUTLET WITH A FRESH GAS FLOW OF 1.0 L/MIN AND 100 VOL% OXYGEN. IN THE FIRST 6 MINUTES AN ETCO2 BETWEEN 20 AND 36 MMHG WAS MEASURED, THE INSPIRATORY OXYGEN CONCENTRATION WAS - COMPARED TO THE FRESH GAS FLOW SETTING - LOW WITH VALUES BETWEEN 21 AND 38 VOL%. THE DEVICE REPEATEDLY ALARMED FIO2 LOW. NO PRESSURE BUILD-UP WAS RECORDED DURING THE WHOLE PROCEDURE. NO ETCO2 WAS DETECTED IN THE FOLLOWING 10 MINUTES. AT AROUND 16:15, THE INSPIRATORY O2 CONCENTRATION ROSE SHARPLY FROM 21 TO 52 AND THEN UP TO 79 VOL%. BETWEEN 16:19 AND 16:22, AN ETCO2 BETWEEN 8 AND 21 MMHG WAS TEMPORARILY MEASURED, BUT SUBSEQUENTLY FELL BACK TO 0. FURTHERMORE, THE INSPIRATORY OXYGEN CONCENTRATION ALSO FELL TO VALUES BETWEEN 21 AND 25 VOL% - DESPITE AN INTERIM INCREASE IN FRESH GAS FLOW TO 5.5 L/MIN (16:18) AND REPEATED ACTIVATION OF EMERGENCY O2 DOSING. AT 16:35 THE UNIT WAS PLACED IN STANDBY. THE ANALYSIS OF THE DEVICE LOGFILE HAS GIVEN NO HINTS FOR A MALFUNCTION OF THE DEVICE. ACCORDING TO THE RECORDS, THE DEVICE WAS USED IN EXT. CGO MODE. IN THIS CASE THE SET FRESH GAS FLOW IS PROVIDED VIA THE EXTERNAL FRESH-GAS OUTLET. IT IS CONCEIVABLE THAT THE NON-REBREATHING SYSTEM WAS NOT (CORRECTLY) CONNECTED, THE PATIENT WAS CONNECTED TO THE INTEGRATED BREATHING SYSTEM, OR THE WRONG BREATHING BAG (BAG OF THE INTEGRATED BREATHING SYSTEM) WAS USED TO MANUALLY VENTILATE THE PATIENT. THE CORRECT CONNECTION AND USE OF A NON-REBREATHING SYSTEM IS DESCRIBED IN THE INSTRUCTIONS FOR USE OF THE ANESTHESIA MACHINE. IT CONTAINS THE FOLLOWING WARNING: "INSUFFICIENT GAS SUPPLY TO THE PATIENT: NON-REBREATHING SYSTEMS ARE ONLY SUITABLE AND INTENDED FOR MANUAL VENTILATION OR SPONTANEOUS BREATHING AND MAY ONLY BE CONNECTED TO THE EXTERNAL FRESH-GAS OUTLET. WHEN USING A NON-REBREATHING SYSTEM, ENSURE AN ADEQUATE GAS MONITORING." FINALLY, AS NO FURTHER INFORMATION REGARDING THE VENTILATION SETUP WAS PROVIDED (ALTHOUGH REQUESTED) THE EXACT ROOT CAUSE COULDN¿T BE DETERMINED. THE DEVICE ISSUED ADEQUATE ALARMS TO KEEP THE USER INFORMED OF THE CURRENT SITUATION AT ALL TIMES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22632730&pc=GDO | 3011137372-2025-00295 | 2025-05-09 04:00 | Death | TELEFLEX MEDICAL | 2025-07-28 04:00 | GDO | WECK HZ APPLIER SMALL 8" CVD | Separation Failure | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT "THE INTENSIVE CARE UNIT INFORMS US THAT THE PATIENT IS DRAINING HEAVILY THROUGH THE PLEUR-EVAC AND MUST UNDERGO A REVIEW, WHICH INVOLVES PLACING SURGICAL DRAPES, REMOVING SKIN STITCHES, FASCIA, AND STERNUM WIRES. ABUNDANT CLOTS ARE EVIDENT IN THE MEDIASTINUM, WITH ACTIVE ARTERIAL BLEEDING FROM A BLEEDING COLLATERAL OF THE INTERNAL MAMMARY ARTERY. HEMOSTASIS IS PERFORMED WITH A NEW CLIP, CONTROLLING THE BLEEDING. CLOTS ARE EVACUATED FROM THE MEDIASTINUM, AND BRIDGES, ANASTOMOSES, SUTURES, STERNAL WALL, AND MEDIASTINUM ARE EXAMINED WITHOUT FINDING OTHER SITES OF BLEEDING". THE PATIENT STATUS IS REPORTED AS "DECEASED". NO FURTHER INFORMATION IS AVAILABLE.Event Description: IT WAS REPORTED THAT "THE INTENSIVE CARE UNIT INFORMS US THAT THE PATIENT IS DRAINING HEAVILY THROUGH THE PLEUR-EVAC AND MUST UNDERGO A REVIEW, WHICH INVOLVES PLACING SURGICAL DRAPES, REMOVING SKIN STITCHES, FASCIA, AND STERNUM WIRES. ABUNDANT CLOTS ARE EVIDENT IN THE MEDIASTINUM, WITH ACTIVE ARTERIAL BLEEDING FROM A BLEEDING COLLATERAL OF THE INTERNAL MAMMARY ARTERY. HEMOSTASIS IS PERFORMED WITH A NEW CLIP, CONTROLLING THE BLEEDING. CLOTS ARE EVACUATED FROM THE MEDIASTINUM, AND BRIDGES, ANASTOMOSES, SUTURES, STERNAL WALL, AND MEDIASTINUM ARE EXAMINED WITHOUT FINDING OTHER SITES OF BLEEDING". THE PATIENT STATUS IS REPORTED AS "DECEASED". NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). THE SAMPLE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER REPORTED: "THE DHR FOR THE RETURNED INSTRUMENTS WAS REVIEWED AND FOUND COMPLETE WITHOUT ANY IRREGULARITIES. THIS INSTRUMENT WAS PRODUCED AT THE TECOMET, INC., KENOSHA, WI, FACILITY AS PART OF A COMBINED 99 PC. LOT IN JANUARY OF 2018 FROM LOT NUMBER: 06C1747927. IT WAS FOUND THAT THIS ORDER WAS MADE FROM THE CORRECT MATERIALS AND COMPONENTS, AND ALL APPROVED PROCESSES WERE FOLLOWED. IT CAN THEN BE STATED THAT THE ALLEGED NON-CONFORMANCE THAT INCITED THIS COMPLAINT WAS NOT DUE TO AN ERROR IN TECOMET - KENOSHA'S MANUFACTURING PROCESSES. EVALUATION OF THE RETURNED INSTRUMENT AS RECEIVED SHOWS THAT THE OPEN JAW TIP SET GAP IS SLIGHTLY UNDERSIZED, THUS MAKING IT MORE DIFFICULT TO LOAD A CLIP. WE ARE ABLE TO VALIDATE THIS COMPLAINT FOR THE RETURNED INSTRUMENT. FURTHER EVALUATION SHOWS THAT THE BEND ON THE SPRING MECHANISM HAS BEEN ALTERED, THUS MAKING IT EASIER TO CLOSE THE APPLIER WHILE LOADING A CLIP, WHICH IN TURN CHANGED THE TIP SET GAP, MAKING IT SLIGHTLY UNDERSIZED. AFTER THE INITIAL EVALUATION, THE SPRING WAS RE-BENT TO THE PROPER BEND, AND NOW THIS INSTRUMENT IS ABLE TO PICK UP, RETAIN, CLOSE, AND RELEASE MULTIPLE CLIPS WITHOUT ANY DIFFICULTY. WE ARE UNABLE TO DETERMINE WHAT CAUSED THE SPRING BEND TO CHANGE ON THIS DEVICE, THUS MAKING IT DIFFICULT TO LOAD A CLIP, BUT MISHANDLING OF THIS DEVICE AT THE END USER'S FACILITY IS SUSPECTED. ALL INSTRUMENTS FROM THIS LOT WERE 100% VISUALLY INSPECTED AND FUNCTION TESTED BEFORE SHIPMENT TO THE CUSTOMER, AS THIS IS A STANDARDIZED PROCEDURE AT THIS FACILITY FOR THIS PRODUCT LINE. DUE TO THESE FINDINGS, NO FURTHER ACTIONS WILL BE TAKEN IN RESPONSE TO THIS COMPLAINT, AND THIS RECORD WILL BE DEEMED CLOSED." TELEFLEX WILL CONTINUE TO MONITOR AND TREND ON COMPLAINTS OF THIS NATURE. Manufacturer Narrative: QN #: (B)(4). N/A. OTHER REMARKS: N/A. CORRECTED DATA: N/A.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22632774&pc=DSQ | 2916596-2025-04984 | 2025-07-20 04:00 | Death | THORATEC CORPORATION | 2025-07-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR FAILURE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS RIGHT HEART FAILURE AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22632895&pc=MHX | 9610816-2025-000684 | 2025-07-04 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-07-28 04:00 | MHX | INTELLIVUE MX800 PATIENT MONITOR | Defective Alarm | Insufficient Information | K150310 |  | 1 | Event Description: IT WAS REPORTED THAT ASYSTOLE WAS NOT SIGNALED BY A RED ALARM. IT WAS INDICATED THAT THERE WAS CLINICAL DETERIORATION OF THE PATIENT LEADING TO DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. Manufacturer Narrative: E1: REPORTING INSTITUTION PHONE #: (B)(6). E1: REPORTER PHONE #: (B)(6). PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22632917&pc=NGV | 2124215-2025-50086 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-28 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR UNSPECIFIED INFECTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22633028&pc=NGV | 2124215-2025-50583 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-28 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR DEATH CARDIOVASCULAR PROCEDURE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 4 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22633710&pc=DSQ | 3007042319-2025-01529 | 2025-07-21 04:00 | Death | HEARTWARE, INC. | 2025-07-28 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Unspecified Nervous System Problem; Respiratory Arrest; Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT PASSED AWAY, AND THE CAUSE OF DEATH IS UNKNOWN.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY ROOM WITH NEUROLOGICAL DEFICITS. THE PATIENT'S INTERNATIONAL NORMALIZED RATIO (INR) AT THE TIME OF ADMISSION WAS 19. THE PATIENT EXPERIENCED A SUBDURAL HEMATOMA AND SUBSEQUENTLY EXPIRED OF RESPIRATORY ARREST. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION. UPDATED FIELDS: B5. DESC EVT PROBLEM H6 PATIENT CODES INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THE COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT DIED AFTER EXPERIENCING RESPIRATORY ARREST. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, RESPIRATORY DYSFUNCTION, NEUROLOGICAL DYSFUNCTION, AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22634649&pc=NPT | 2025587-2025-05300 | 2024-08-04 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Dyspnea; Unspecified Infection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT 30 DAYS AFTER THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT WAS HOSPITALIZED DUE TO DYSPNEA. FOLLOWING HOSPITALIZATION, THE PATIENT DEVELOPED PNEUMONIA AND DIED. THE CAUSE OF DEATH WAS REPORTED TO BE PULMONARY IN ORIGIN. PER THE PHYSICIAN, THERE WAS NO RELATIONSHIP BETWEEN THE DEATH AND DEVICE OR IMPLANT PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: D3, D4, H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22635841&pc=NGV | 2124215-2025-50068 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-28 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 4 | Event Description: THIS REPORT SUMMARIZES 4 DEATH EVENTS FOR ADDITIONAL DEVICE REQUIRED. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 60 TO 88 YEARS, PATIENT SEX: 3 FEMALE; 1 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 6 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22635907&pc=DSQ | 2916596-2025-04965 | 2025-07-16 04:00 | Death | THORATEC CORPORATION | 2025-07-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Shock; Respiratory Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT ORIGINALLY DECOMPENSATED ON (B)(6) 2025 WHICH REQUIRED REINTUBATION. THEY CONTINUED TO HAVE INCREASING PRESSOR REQUIREMENTS AND WAS GIVEN METHYLENE BLUE AND INITIATED ON CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) ON (B)(6) 2025. THE PATIENT'S PRESSOR REQUIREMENTS CONTINUED TO INCREASE, AND THEY DEVELOPED SHOCK AND RESPIRATORY FAILURE. THE DECISION WAS MADE TO TAKE THE PATIENT TO THE OPERATING ROOM (OR) FOR PROTEKDUO RIGHT VENTRICULAR ASSIST DEVICE (RVAD) WITH OXYGENATOR. UNFORTUNATELY, THE PATIENT CONTINUED TO DECOMPENSATE DURING THE PROCEDURE, SO RVAD WAS ABORTED, AND CARE WAS ULTIMATELY WITHDRAWN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22636299&pc=OZD | 1220648-2025-30792 | 2025-07-02 04:00 | Death | ABIOMED, INC. | 2025-07-28 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY IMPLANTED WITH AN IMPELLA CP THEN FOLLOWED BY AN IMPELLA RP FLEX DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). THE PATIENT EXPERIENCED LIMB ISCHEMIA, BLEEDING AND SUBSEQUENTLY EXPIRED. THE PATIENT PRESENTED WITH CHEST PAIN AND WAS EMERGENTLY TAKEN TO THE CATHETERIZATION LAB. THE PATIENT SUBSEQUENTLY CODED AND WAS SHOCKED EIGHT TIMES BEFORE PROCEEDING WITH THE IMPELLA CP DEVICE INSERTED VIA THE RIGHT FEMORAL ARTERY DURING CODE BLUE PROCEDURE. FLOWS WERE INITIATED AND EVENTUALLY THE PATIENT STABILIZED WITH RETURN OF SPONTANEOUS CIRCULATION. THE CULPRIT WAS A 99% OCCLUDED PROXIMAL RIGHT CORONARY ARTERY WHICH WAS TREATED WITH ANGIOPLASTY AND DRUG-ELUTING STENTING. THE DISEASE TO THE LEFT ANTERIOR DESCENDING AND DIAGONAL ARTERIES WERE PLANNED TO BE TREATED ON A DIFFERENT DAY. A RIGHT HEART CATHETERIZATION WAS PERFORMED, AND THE PHYSICIAN PROCEEDED WITH AN IMPELLA RP PLACEMENT VIA THE RIGHT FEMORAL VEIN INTO LEFT PULMONARY ARTERY WITHOUT ISSUE. RIGHT-SIDED SUPPORT WAS STARTED WITH FLOWS OF 2.5-3.0 L/MIN ACHIEVED. THE PATIENT WAS "EXTREMELY" FLUID RESPONSIVE AND NO MARKED IMPROVEMENT WITH RP FLEX PLACEMENT. ECHOCARDIOGRAM IN THE PROCEDURE AREA SHOWED LEFT VENTRICULAR FUNCTION OF APPROXIMATELY 55% AND RIGHT VENTRICULAR FUNCTION SLIGHTLY DIMINISHED, HOWEVER, NOT DILATED. THE PATIENT WAS SENT TO UNIT ON BIPELLA (IMPELLA CP AND RP FLEX) SUPPORT. APPROXIMATELY ONE DAY INTO BIPELLA MCS, SYSTEMIC HEPARIN DRIP WAS ORDERED WITH GOAL OF ACTIVATED CLOTTING TIME ABOVE 200 SECONDS. THE PATIENT HAD DECREASED PULSES IN BOTH EXTREMITIES WITH MORE NOTICEABLE SYMPTOMS ON THE IMPELLA SIDE. A FEM-FEM BYPASS WAS PERFORMED TO MANAGE THE ISCHEMIA. THE PATIENT CONTINUED TO HAVE WORSENING HEMODYNAMICS. IN ADDITION, ABDOMINAL BLEEDING WAS OBSERVED, FOR WHICH MULTIPLE BLOOD PRODUCTS WERE GIVEN. A BEDSIDE ABDOMINAL TAP REVEALED BLOOD IN ABDOMEN THAT WAS FIRM. THE PATIENT CONTINUED TO DETERIORATE FROM THE UNDERLYING CONDITION AND THE DECISION WAS MADE TO WITHDRAW CARE. Manufacturer Narrative: THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE OUTCOME GIVEN THE DEVICE RELATED LIMB ISCHEMIA EVENT REQUIRING SURGICAL INTERVENTION IN ADDITION TO THE HEPARIN FOR IMPELLA SUPPORT AND ITS UNKNOWN IMPACT TO THE PATIENT'S ABDOMINAL BLEED. THIS IS ONE OF TWO DEVICES ASSOCIATED WITH THE EVENT. REFER TO MEDICAL DEVICE REPORT FOR THE IMPELLA RP FLEX SN (B)(6), DATE OF EVENT 03-JUL-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). THE IMPELLA CP DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22636383&pc=QZI | 2182208-2025-03728 | 2025-04-29 04:00 | Death | MEDTRONIC, INC. | 2025-07-28 04:00 | QZI | AFFERA | Adverse Event Without Identified Device or Use Problem | Pericardial Effusion | P240013 |  | 1 | Event Description: IT WAS REPORTED THAT A COMPLEX VENTRICULAR TACHYCARDIA (VT) ABLATION WAS PERFORMED ON A PATIENT WITH LIFE-THREATENING, REFRACTORY VENTRICULAR TACHYCARDIA (VT). AFTER EXTENSIVE ABLATION FROM BOTH ENDOCARDIAL AND EPICARDIAL APPROACHES, VT REMAINED INDUCIBLE. THE PROCEDURE WAS CONCLUDED DUE TO THE COMPLEXITY, PROCEDURAL DURATION, AND LIMITED ADDITIONAL THERAPEUTIC TARGETS. POST-ABLATION, INTRACARDIAC ECHOCARDIOGRAPHY (ICE) IMAGING SHOWED A SMALL CIRCUMFERENTIAL PERICARDIAL EFFUSION. HEMOSTASIS WAS ACHIEVED USING ARTERIAL AND VENOUS CLOSURE DEVICES, AND A PERICARDIAL DRAIN WAS LEFT IN PLACE. THE PATIENT WAS TRANSFERRED BACK TO THE CARDIAC INTENSIVE CARE UNIT (CICU) UNDER ECMO (EXTRACORPOREAL MEMBRANE OXYGENATION) AND BVM (BAG-VALVE-MASK) SUPPORT FOR CONTINUED MANAGEMENT UNDER CRITICAL CONDITIONS. DURING THE PATIENT¿S TEN-DAY HOSPITALIZATION, REPEAT ECHOCARDIOGRAPHY UNDER FULL SUPPORT REVEALED A SEVERELY REDUCED LEFT VENTRICULAR EJECTION FRACTION - LVEF (15¿20%) WITH NO MEANINGFUL CARDIAC RECOVERY. ADDITIONALLY, GIVEN PERSISTENT MULTI-ORGAN FAILURE, FAILURE TO WEAN FROM MECHANICAL SUPPORT, AND LIMITED NEUROLOGIC PROGNOSIS, A MULTIDISCIPLINARY TEAM, TOGETHER WITH THE PATIENT¿S FAMILY, ELECTED TO WITHDRAW LIFE-SUSTAINING THERAPY AND THE PATIENT DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22636475&pc=OMP | 3009897021-2025-00032 | 2025-06-19 04:00 | Death | KINETIC CONCEPTS, INC | 2025-07-28 04:00 | OMP | ACTIV.A.C.¿ THERAPY SYSTEM | Improper or Incorrect Procedure or Method | Cardiac Arrest; Hemorrhage/Bleeding | K201571 |  | 1 | Event Description: ON (B)(6) 2025, THE FOLLOWING INFORMATION WAS PROVIDED BY THE TISSUE VIABILITY NURSE: ON OR ABOUT (B)(6) 2025, THE PATIENT UNDERWENT SURGICAL DEBRIDEMENT AND PLACEMENT OF THE ACTIV.A.C.¿ THERAPY UNIT TO A LACERATION TO THE RIGHT LOWER LEG; THE WOUND WAS A SUPERFICIAL AND WOULD NOT BE INDICATED FOR V.A.C.® THERAPY. THE DEVICE WAS TURNED OFF OVERNIGHT DUE TO AN UNKNOWN REASON, AND THE V.A.C.® DRESSING WAS LEFT IN SITU. THE PUMP WAS RESTARTED IN THE MORNING; AFTER RESTARTING V.A.C.® THERAPY, THERE WAS A PERIOD OF BLEEDING. THE NURSE INITIALLY DECIDED TO REMOVE THE DEVICE BUT WAS ADVISED BY THE MEDICAL TEAM THAT ¿THE BLOOD WAS OLD, AND THE PATIENT WAS BEING MONITORED¿; V.A.C.® THERAPY CONTINUED. DESPITE A SECOND BLEEDING EPISODE, V.A.C.® THERAPY WAS NOT STOPPED. THE PATIENT EXPERIENCED A CARDIAC ARREST ATTRIBUTED TO HYPOVOLEMIA AND ANEMIA; THE PATIENT EXPIRED SEVERAL DAYS LATER. ON (B)(6) 2025, THE FOLLOWING INFORMATION WAS PROVIDED BY THE REPRESENTATIVE: THE HEALTHCARE PROVIDER BELIEVES THAT V.A.C.® THERAPY MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT¿S DEATH. THE EXACT DATE OF THE ALLEGED EVENT IS UNKNOWN BUT OCCURRED SOMETIME BETWEEN 19-JUN-2025 AND 24-JUN-2025. ON 21-JUL-2025, A DEVICE EVALUATION WAS COMPLETED BY SOLVENTUM QUALITY ENGINEERING. ON 14-JUN-2024, THE DEVICE WAS TESTED PER QUALITY CONTROL PROCEDURE BY A SOLVENTUM SERVICE CENTER, AND THE UNIT PASSED THE QUALITY CONTROL CHECKS AND MET SPECIFICATIONS. ON (B)(6) 2025, THE DEVICE WAS PLACED WITH THE PATIENT. ON 25-JUN-2025, THE DEVICE WAS TESTED PER QUALITY CONTROL PROCEDURE BY A SOLVENTUM SERVICE CENTER AND THE UNIT PASSED THE QUALITY CONTROL CHECKS AND MET SPECIFICATIONS. INSPECTION AND TESTING OF THE DEVICE DID NOT REVEAL ANY EVIDENCE OF AN OPERATIONAL MALFUNCTION WITH THE UNIT. Manufacturer Narrative: BASED ON THE INFORMATION PROVIDED, IT CANNOT BE DETERMINED THAT THE ALLEGED CARDIAC ARREST RESULTING IN THE PATIENT'S DEATH IS RELATED TO THE ACTIV.A.C.¿ THERAPY SYSTEM. THE PATIENT RECENTLY UNDERWENT SURGICAL DEBRIDEMENT, AND THE WOUND WAS NOTED TO BE SUPERFICIAL. THE DEVICE PASSED QUALITY CONTROL CHECKS BEFORE AND AFTER PATIENT PLACEMENT. THE TISSUE VIABILITY NURSE INDICATED THE PATIENT¿S WOUND RECENTLY UNDERWENT SURGICAL DEBRIDEMENT, AND WOULD NOT BE INDICATED FOR V.A.C.® THERAPY BECAUSE IT WAS SUPERFICIAL. ADDITIONALLY, THE PATIENT EXPERIENCED MULTIPLE BLEEDING EPISODES, AND THE ACTIV.A.C.¿ THERAPY SYSTEM WAS LEFT ON AGAINST THE NURSE¿S ADVICE. THEREFORE, THIS EVENT IS BEING REPORTED DUE TO POTENTIAL USE ERROR. DEVICE LABELING, AVAILABLE IN PRINT, STATES: WITH OR WITHOUT USING V.A.C.® THERAPY, CERTAIN PATIENTS ARE AT HIGH RISK OF BLEEDING COMPLICATIONS. THE FOLLOWING TYPES OF PATIENTS ARE AT INCREASED RISK OF BLEEDING, WHICH, IF UNCONTROLLED, COULD BE POTENTIALLY FATAL. PATIENTS WHO HAVE WEAKENED OR FRIABLE BLOOD VESSELS OR ORGANS IN OR AROUND THE WOUND AS A RESULT OF, BUT NOT LIMITED TO: SUTURING OF THE BLOOD VESSEL (NATIVE ANASTOMOSIS OR GRAFTS)/ORGAN INFECTION. TRAUMA, RADIATION, PATIENTS WITHOUT ADEQUATE WOUND HEMOSTASIS, PATIENTS WHO HAVE BEEN ADMINISTERED ANTICOAGULANTS OR PLATELET AGGREGATION INHIBITORS, PATIENTS WHO DO NOT HAVE ADEQUATE TISSUE COVERAGE OVER VASCULAR STRUCTURES. IF V.A.C.® THERAPY IS PRESCRIBED FOR PATIENTS WHO HAVE AN INCREASED RISK OF BLEEDING COMPLICATIONS, THEY SHOULD BE TREATED AND MONITORED IN A CARE SETTING DEEMED APPROPRIATE BY THE TREATING PHYSICIAN. IF ACTIVE BLEEDING DEVELOPS SUDDENLY OR IN LARGE AMOUNTS DURING V.A.C.® THERAPY, OR IF FRANK (BRIGHT RED) BLOOD IS SEEN IN THE TUBING OR IN THE CANISTER, IMMEDIATELY STOP V.A.C.® THERAPY, LEAVE DRESSING IN PLACE, TAKE MEASURES TO STOP THE BLEEDING AND SEEK IMMEDIATE MEDICAL ASSISTANCE. THE V.A.C.® THERAPY UNITS AND DRESSINGS SHOULD NOT BE USED TO PREVENT, MINIMIZE OR STOP VASCULAR BLEEDING. PROTECT VESSELS AND ORGANS: ALL EXPOSED OR SUPERFICIAL VESSELS AND ORGANS IN OR AROUND THE WOUND MUST BE COMPLETELY COVERED AND PROTECTED PRIOR TO THE ADMINISTRATION OF V.A.C.® THERAPY. WARNINGS PROTECT VESSELS AND ORGANS: ALL EXPOSED OR SUPERFICIAL VESSELS AND ORGANS IN OR AROUND THE WOUND MUST BE COMPLETELY COVERED AND PROTECTED PRIOR TO THE ADMINISTRATION OF V.A.C.® THERAPY. ALWAYS ENSURE THAT V.A.C.® FOAM DRESSINGS DO NOT COME IN CONTACT WITH VESSELS OR ORGANS. USE A THICK LAYER OF NATURAL TISSUE SHOULD PROVIDE THE MOST EFFECTIVE PROTECTION. IF A THICK LAYER OF NATURAL TISSUE IS NOT AVAILABLE OR IS NOT SURGICALLY POSSIBLE, BIOENGINEERED TISSUE OR MULTIPLE LAYERS OF NON-ADHERENT DRESSING MATERIAL MAY BE CONSIDERED AS AN ALTERNATIVE, IF DEEMED BY THE TREATING PHYSICIAN TO PROVIDE A COMPLETE PROTECTIVE BARRIER. IF USING NON-ADHERENT MATERIALS, ENSURE THEY ARE SECURED IN A MANNER THAT WILL MAINTAIN THEIR PROTECTIVE POSITION THROUGHOUT THERAPY. CONSIDERATION SHOULD ALSO BE GIVEN TO THE NEGATIVE PRESSURE SETTING AND THERAPY MODE WHEN USED WHEN INITIATING THERAPY. CAUTION SHOULD BE TAKEN WHEN TREATING LARGE WOUNDS THAT MAY CONTAIN HIDDEN VESSELS WHICH MAY NOT BE READILY APPARENT. THE PATIENT SHOULD BE CLOSELY MONITORED FOR BLEEDING IN A CARE SETTING DEEMED APPROPRIATE BY THE TREATING PHYSICIAN. DISCLAIMER: THIS INFORMATION IS SUBMITTED PURSUANT TO 21 CFR 803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A KINETIC CONCEPTS, INC. PRODUCT MALFUNCTIONED, IS DEFECTIVE OR HAS CAUSED SERIOUS INJURY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22636506&pc=LWP | 2124215-2025-50609 | 2025-07-02 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-28 04:00 | LWP | ACCOLADE? MRI EL DR | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Bradycardia; Respiratory Arrest | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE, AFTER THE RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEADS WERE IN PLACE, THE PATIENT'S HEART RATE WAS OBSERVED TO BE 30-40 BPM AND RESPIRATORY ARREST WAS OBSERVED. THE PACEMAKER WAS CONNECTED TO THE LEADS AND PACING WAS INITIATED AT 60 BPM WITH MAXIMUM OUTPUTS. IT WAS NOTED THE RV LEAD WAS PLACED IN THE LEFT BUNDLE BRANCH (LBB) SO A PERFORATION WAS NOT SUSPECTED. CHEST COMPRESSIONS WERE PERFORMED, AND A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) SHOWED SOME BLOOD IN THE PERICARDIUM, WHICH MAY HAVE BEEN DUE TO THE CHEST COMPRESSIONS. IT WAS THOUGHT A PULMONARY EMBOLISM HAD OCCURRED. THE PATIENT DIED. THE DEVICE AND LEADS WERE NOT REMOVED POST-MORTEM. IT WAS LATER REPORTED THAT NO AUTOPSY WAS PERFORMED AND THE PHYSICIAN DID NOT BELIEVE THE IMPLANT OF THE LEADS CONTRIBUTED TO THE PATIENT'S DEATH.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE, AFTER THE RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEADS WERE IN PLACE, THE PATIENT'S HEART RATE WAS OBSERVED TO BE 30-40 BPM AND RESPIRATORY ARREST WAS OBSERVED. THE PACEMAKER WAS CONNECTED TO THE LEADS AND PACING WAS INITIATED AT 60 BPM WITH MAXIMUM OUTPUTS. IT WAS NOTED THE RV LEAD WAS PLACED IN THE LEFT BUNDLE BRANCH (LBB) SO A PERFORATION WAS NOT SUSPECTED. CHEST COMPRESSIONS WERE PERFORMED, AND A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) SHOWED SOME BLOOD IN THE PERICARDIUM, WHICH MAY HAVE BEEN DUE TO THE CHEST COMPRESSIONS. IT WAS THOUGHT A PULMONARY EMBOLISM HAD OCCURRED. THE PATIENT DIED. THE DEVICE AND LEADS WERE NOT REMOVED POST-MORTEM. IT WAS LATER REPORTED THAT NO AUTOPSY WAS PERFORMED AND THE PHYSICIAN DID NOT BELIEVE THE IMPLANT OF THE LEADS CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THIS REPORT IS FILED TO CORRECT FIELDS B2 DATE OF DEATH AND B3 DATE OF EVENT. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22636539&pc=MVK | 3008642652-2025-07238 | 2025-06-22 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-28 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 09:24:10 ON (B)(6) 2025. AT 16:32:06, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 16:33:46, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 16:35:23 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22636588&pc=MVK | 3008642652-2025-07240 | 2025-03-13 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-28 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 06:51:14 ON (B)(6) 2025. AT 18:42:11, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS IDIOVENTRICULAR RHYTHM @ 60 BPM DEGRADING TO VF. AT 18:42:48, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH PACER SPIKES AND INTERMITTENT CARDIAC ACTIVITY. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE ELECTRODE BELT WAS DISCONNECTED AT 18:52:45 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22636768&pc=NGV | 2124215-2025-50089 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-28 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Unspecified Tissue Injury |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR HEART VALVE DAMAGE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025 AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22636916&pc=NPT | 2025587-2025-05307 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-28 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Unspecified Infection | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS INFECTION. THE TIME TO EVENT WAS TWO HUNDRED SEVENTY DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 65. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22637327&pc=LPB | 2029046-2025-02461 | 2025-05-15 04:00 | Death | BIOSENSE WEBSTER INC | 2025-07-28 04:00 | LPB | NGEN RF GENERATOR | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Fistula | P030031 |  | 1 | Event Description: IT WAS REPORTED 3 TO 4 WEEKS AFTER PULMONARY VEIN ISOLATION (PVI) ABLATION PROCEDURE WITH A THERMOCOOL SMARTTOUCH SF CATHETER THE PATIENT EXPERIENCED STROKE AND EXPIRED. THE PATIENT EXPERIENCED DEATH APPROXIMATELY 3-4 WEEKS POST PVI ABLATION AFTER RETURNING TO THE HOSPITAL WITH A STROKE. A COMPUTED TOMOGRAPHY (CT) SCAN WAS COMPLETED TO RULE OUT A FISTULA, BUT NOTHING WAS FOUND. PATIENT PASSED AWAY OVER THE WEEKEND, APPROXIMATELY ON (B)(6). LIMITED INFORMATION AVAILABLE, BUT THE PATIENT WAS THEN FOUND TO REMAIN IN ATRIAL FIBRILLATION THE FOLLOWING DAY AND WAS TREATED FOR REDO PVI ABLATION WITH RADIOFREQUENCY WITHIN 48 HOURS OF THE FIRST PFA ABLATION. THE FAMILY REFUSED AUTOPSY. ON 2-JUL-2024, ADDITIONAL INFORMATION WAS RECEIVED INDICATING THE PHYSICIAN¿S OPINION FOR CAUSE OF DEATH WAS SUSPECTED FISTULA, BUT OVERALL, NOT CONFIRMED. THE INITIAL CASE DATE WAS ON (B)(6), FOR PULMONARY VEIN ISOLATION (PVI) WITH VARIPULSE USING 30 ML IRRIGATION DURING ABLATION. UNDERWENT PVI WITH VARIPULSE AND POST ABLATION, AT THE HOTEL THAT NIGHT, WENT INTO PERSISTENT AF. RETURNED TO THE HOSPITAL NEXT DAY, REPEATED UNSUCCESSFUL CARDIOVERSIONS. THREE DAYS AFTER THE INITIAL ABLATION PROCEDURE, THE PATIENT WAS BROUGHT BACK FOR REDO ABLATION WITH RADIOFREQUENCY (RF). RIGHT POSTERIOR CARINA, LEFT INFERIOR PULMONARY VEIN (LIPV) FOUND TO BE RECONNECTED. THE PHYSICIANS RE-ISOLATED THE VEINS AND POSTERIOR BOX WITH RADIOFREQUENCY ABLATION. THE PATIENT WAS CARDIOVERTED AT END OF PROCEDURE, AND THE PATIENT LEFT IN SINUS RHYTHM. TWO WEEKS LATER THE PHYSICIAN FOLLOWS UP, AND THE PATIENT WAS DOING WELL. THREE WEEKS LATER THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT WITH STROKE, AND ONE SYMPTOM WAS APPARENTLY ESOPHAGEAL BLEEDING. THE PATIENT DIED IN THE HOSPITAL, AND THE FAMILY DENIED AUTOPSY. THE NEUROLOGICAL IMPAIRMENT EVENT OCCURRED WAS CEREBROVASCULAR ACCIDENT (CVA) /STROKE AND TRANSIENT ISCHEMIC ATTACK (TIA). A CT SCAN IMAGING WAS PERFORMED TO CONFIRM THE STROKE. NO EVIDENCE OF CHAR DURING THE PROCEDURE. NO EVIDENCE OF BLOOD THROMBUS/CLOT DURING THE PROCEDURE. THE CORRECT CATHETER SETTINGS WERE SELECTED ON THE GENERATOR. NO ISSUES WITH FLOW RATE CHANGE AT THE START OF ABLATION. THE EVENT DESCRIPTION INDICATED THAT THE INITIAL ABLATION PROCEDURE WAS PERFORMED ON (B)(6), BUT NOT THE ADVERSE EVENT. THE DEVICES USED WERE THERMOCOOL SMARTTOUCH SF CATHETER AND NGEN GENERATOR. BASED ON THE INFORMATION AVAILABLE, THE EVENT HAS BEEN DETERMINED TO BE MDR REPORTABLE AGAINST THE THERMOCOOL SMARTTOUCH SF CATHETER AND THE NGEN RF GENERATOR. Manufacturer Narrative: THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. NOTE: PER FDA REQUEST, MDR SUBMISSIONS FOR THE NGEN RF GENERATOR ARE TO BE REPORTED WITH PMA DETAILS OF THE CATHETER USED ALONG WITH THE NGEN RF GENERATOR. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REF. #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22637328&pc=DSQ | 2916596-2025-04794 | 2025-07-08 04:00 | Death | THORATEC CORPORATION | 2025-07-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Tachycardia; Ventricular Fibrillation; Multiple Organ Failure; Aspiration Pneumonitis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. THE MULTI-SYSTEM ORGAN FAILURE WAS THOUGHT TO HAVE BEEN A NATURAL PROGRESSION OF THE PATIENT'S UNDERLYING DISEASE PROCESS FOR WHICH THEY WERE RECEIVING LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY. THE DEATH WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE PATIENT PRESENTED WITH SYMPTOMS OF LETHARGY, HYPOTENSION, AND HYPOXIA. THE PATIENT DETERIORATED SIGNIFICANTLY OVERNIGHT WITH PROGRESSIVE MULTI-SYSTEM ORGAN FAILURE, RECURRENT VENTRICULAR TACHYCARDIA/VENTRICULAR FIBRILLATION, AND VERY LARGE ASPIRATION PNEUMONIA. THE PATIENT SPONTANEOUSLY CONVERTED DURING THE FIRST EPISODE OF ARRHYTHMIA AND THEN REQUIRED SHOCK BY IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR (ICD) FOR THE SECOND EPISODE. THE ICD WAS IMPLANTED PRIOR TO THE PUMP IMPLANT. THE PATIENT DEVELOPED SILENT ASPIRATION POST INTUBATION FOR IMPLANT. THEY WERE ON DYSPHAGIA 3 DIET ONCE EXTUBATED. THE SITE STATED IT WAS POSSIBLE THE ASPIRATION WAS DUE TO NAUSEA AND VOMITING. THE PATIENT DEVELOPED VENTRICULAR TACHYCARDIA/FIBRILLATION IN 2021 PRIOR TO LVAD IMPLANT. THE ARRHYTHMIA WAS NOT THOUGHT TO HAVE BEEN DEVICE OR THERAPY RELATED. THE LVAD FUNCTIONED AS INTENDED WITHOUT ANY ALARMS OR ISSUES. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, ARRHYTHMIA, INFECTION, AND DEATH, WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THE IFU ALSO LISTS ARRHYTHMIA AND INFECTION AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22637340&pc=NPT | 2025587-2025-05306 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-28 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Unspecified Infection | P130021 | E2014038 | 2 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE INFECTIONS. THE AVERAGE TIME TO EVENT WAS NINETY-EIGHT DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 83 TO 90 AND OVER. THERE WAS 1 MALE AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038. QUARTERLY REPORTING PERIOD: Q2 2025. TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 2. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22637377&pc=NPT | 2025587-2025-05309 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-28 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Renal Failure | P130021 | E2014038 | 2 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE RENAL. THE AVERAGE TIME TO EVENT WAS TWENTY-TWO DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 80 TO 86. THERE WAS 1 MALE AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 2 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22637379&pc=NPT | 2025587-2025-05308 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Renal Failure | P130021 | E2014038 | 11 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE RENAL. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED FORTY-FIVE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 67 TO 90 AND OVER. THERE WERE 5 MALES AND 6 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 11 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22637447&pc=NPT | 2025587-2025-05310 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Unspecified Nervous System Problem | P130021 | E2014038 | 11 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE NEUROLOGIC. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED SEVENTY-EIGHT DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 70 TO (B)(6) AND OVER. THERE WERE 4 MALES AND 7 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 11 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22637574&pc=NPT | 2015691-2025-06122 | 2025-07-10 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-28 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Rupture; Cardiac Tamponade; Pericardial Effusion | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES AFFILIATE IN CANADA, REGARDING SUCCESSFUL IMPLANT OF A 26MM SAPIEN 3 ULTRA RESILIA IN AORTIC POSITION BY TRANSFEMORAL APPROACH. APPROXIMATELY 3 HOURS AFTER VALVE IMPLANT, PATIENT HAD A PEA ARREST AND RESUSCITATION WAS ACHIEVED. TTE SHOWED SIGNIFICANT PERICARDIAL EFFUSION LIKELY SECONDARY TO AORTIC RUPTURE POST PROCEDURE, WHICH CAUSED A CARDIAC TAMPONADE. THE PATIENT WAS NOT A SURGICAL CANDIDATE, AND A PERICARDIAL DRAIN WAS PLACED. HOWEVER, THERE WAS REACCUMULATION OF BLOOD IMMEDIATELY POST EVACUATION. THE TREATMENT WAS DISCONTINUED AS PERICARDIAL EFFUSION/TAMPONADE ENDUED, AND PATIENT PASSED AWAY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT FACTORS (HEAVILY CALCIFIED ANNULUS) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22637617&pc=NPT | 2025587-2025-05311 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-28 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Unspecified Nervous System Problem | P130021 | E2014038 | 3 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE NEUROLOGIC. THE AVERAGE TIME TO EVENT WAS TWENTY-FOUR DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 78 TO 81. THERE WAS 1 MALE AND 2 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038. QUARTERLY REPORTING PERIOD: Q2 2025. TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 3. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22637621&pc=NPT | 2025587-2025-05312 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-28 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Unspecified Nervous System Problem | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS NEUROLOGIC. THE TIME TO EVENT WAS SEVENTEEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 60. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22638526&pc=DSQ | 2916596-2025-05028 | 2025-07-19 04:00 | Death | THORATEC CORPORATION | 2025-07-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22638994&pc=MGB | 3004939290-2025-00614 | 2025-06-09 04:00 | Death | CORDIS US CORP | 2025-07-28 04:00 | MGB | MYNXGRIP | Adverse Event Without Identified Device or Use Problem | Hematoma; Unspecified Vascular Problem | P040044 |  | 1 | Event Description: AS REPORTED, AFTER INADVERTENTLY CANNULATING THE RIGHT COMMON FEMORAL VEIN (CFV), A 6/7F MYNXGRIP VASCULAR CLOSURE DEVICE (VCD) WAS USED TO SEAL THE VEIN. POST-PROCEDURE, THE PATIENT DEVELOPED A HEMATOMA. A FEMORAL COMPRESSION DEVICE WAS APPLIED, BUT THE HEMATOMA PERSISTED. SURGICAL INTERVENTION WAS REQUIRED; THE SURGEON DRAINED THE HEMATOMA AND DISCOVERED A TEAR IN THE VEIN, WHICH WAS THEN REPAIRED. UNFORTUNATELY, THE PATIENT LATER EXPIRED. A PHYSICIAN WAS PERFORMING A PERCUTANEOUS CORONARY INTERVENTION (PCI) ON A PATIENT WITH MULTIPLE COMORBIDITIES. DURING AN ATTEMPT TO ACCESS THE RIGHT COMMON FEMORAL ARTERY (CFA), THE PHYSICIAN INADVERTENTLY CANNULATED THE RIGHT COMMON FEMORAL VEIN (CFV). SUBSEQUENTLY, THE PHYSICIAN ACCESSED THE LEFT COMMON FEMORAL ARTERY (CFA) AND PROCEEDED WITH THE INTERVENTION. THE HOSPITAL IS CURRENTLY CONDUCTING A ROOT CAUSE ANALYSIS OF THE EVENT. ALTHOUGH THEY DO NOT ATTRIBUTE THE ADVERSE OUTCOME TO THE USE OF THE MYNX GRIP DEVICE, THEY HAVE NOTIFIED US DUE TO THE ONGOING INVESTIGATION. ADDITIONAL INFORMATION WAS REQUESTED BUT NOT PROVIDED. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: ADDITIONAL INFORMATION IS PENDING AND WILL BE SUBMITTED WITHIN 30 DAYS UPON RECEIPT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22639120&pc=LWS | 3004209178-2025-12795 | 2025-07-13 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-28 04:00 | LWS | EVERA MRI XT DR SURESCAN | Insufficient Information | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S FAMILY MEMBER THAT THE PATIENT IS DECEASED, AND THEY REQUESTED TO HAVE THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) INTERROGATED. THE FAMILY MEMBER ALSO STATED THAT THEY HAD SPOKEN TO THE PATIENT'S PHYSICIAN, AND "WE BOTH WONDERED WHY THE DEVICE DID NOT GO OFF". NO FURTHER INFORMATION COULD BE OBTAINED THROUGH FOLLOW-UP WITH THE CLINIC. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22639286&pc=NGV | 2124215-2025-50078 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-28 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR SURGICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 16 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22639393&pc=MVK | 3008642652-2025-07239 | 2025-05-30 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-28 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED SIX APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 19:43:02 ON (B)(6) 2025. AT 03:42:14 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 60 BPM WITH NSVT TRANSITIONING TO VT @ 260 BPM. AT 03:42:56, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS BEAT TRANSITIONING BACK TO VT @ 280 BPM. AT 03:43:31, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 80 BPM DEGRADING TO VF. AT 03:43:52, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 03:44:23, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 03:44:53, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF WITH MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. AT 04:23:50, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT TRANSITIONING TO VF. AT 04:24:57, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH TACTILE ARTIFACT AND ELECTRODE LEAD FALLOFF. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE ELECTRODE BELT WAS DISCONNECTED AT 04:48:12 ON (B)(6) 2025. Manufacturer Narrative: THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22639523&pc=NPT | 2025587-2025-05305 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-28 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Unspecified Infection | P130021 | E2014038 | 26 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE INFECTIONS. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED NINETY-THREE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 58 TO 90 AND OVER. THERE WERE 12 MALES AND 14 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 26 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22639602&pc=QZI | 2124215-2025-50569 | 2025-07-03 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-28 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Cardiac Tamponade |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. THE FARAWAVE CATHETER WAS INSERTED INTO THE FARADRIVE SHEATH AND ADVANCED A FEW CENTIMETERS. THE PHYSICIAN THEN ASPIRATED THE FARADRIVE SHEATH FOR AIR MANAGEMENT. DURING THIS TIME, THE FARASTAR GENERATOR PREPARATION BUTTON WAS PRESSED TO PREPARE THE CATHETER. THE FARAWAVE CATHETER AND BOSTON SCIENTIFIC STARTER ROSEN WIRE WERE ADVANCED INTO THE LEFT ATRIUM. AT THIS MOMENT, THE PHYSICIAN MADE THE PROCEDURAL ROOM WARE THAT THE FARAWAVE CATHETER WAS NOT BEING VISUALIZED ON THE MAPPING SYSTEM. CONVERSATION OCCURRED ABOUT NEEDING TO TROUBLESHOOT THE VISUALIZATION ISSUE. THE PHYSICIAN THEN ADJUSTED HIS SHEATH AND INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER TO HELP VISUALIZE THE FARAWAVE CATHETER WITH ICE. AT THIS POINT, THE MAPPING TEAM SUGGESTED A REBOOT OF THE MAPPING SYSTEM WITH THE IDEA OF ADDRESSING THE VISUALIZATION ERROR THAT OCCURRED. WHILE THE TEAM WAS WORKING ON THE REBOOT, THE PHYSICIAN REQUESTED AN ACT. AT THAT TIME, THE PHYSICIAN NOTICED THE PATIENT BECOMING HYPOTENSIVE AND ADDRESSED THE PROCEDURAL ROOM AND STAFF WITH THIS FINDING. THE PHYSICIAN REQUESTED ICE TO INSPECT FOR POSSIBLE PERFORATION BUT WAS UNABLE TO DO SO BECAUSE THE MAPPING SYSTEM WAS DOWN IN A REBOOT PHASE WHICH ALSO TURNED OFF THE ICE MONITOR. THE LAB STAFF ENTERS THE ROOM AND BEGINS TO ASSIST THE PHYSICIAN TO ASSESS THE PATIENT. SHORTLY AFTER, CARDIAC CHEST COMPRESSION BEGAN ON THE PATIENT, BUT THE PATIENT DID NOT MAKE IT.Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. THE FARAWAVE CATHETER WAS INSERTED INTO THE FARADRIVE SHEATH AND ADVANCED A FEW CENTIMETERS. THE PHYSICIAN THEN ASPIRATED THE FARADRIVE SHEATH FOR AIR MANAGEMENT. DURING THIS TIME, THE FARASTAR GENERATOR PREPARATION BUTTON WAS PRESSED TO PREPARE THE CATHETER. THE FARAWAVE CATHETER AND BOSTON SCIENTIFIC STARTER ROSEN WIRE WERE ADVANCED INTO THE LEFT ATRIUM. AT THIS MOMENT, THE PHYSICIAN MADE THE PROCEDURAL ROOM WARE THAT THE FARAWAVE CATHETER WAS NOT BEING VISUALIZED ON THE MAPPING SYSTEM. CONVERSATION OCCURRED ABOUT NEEDING TO TROUBLESHOOT THE VISUALIZATION ISSUE. THE PHYSICIAN THEN ADJUSTED HIS SHEATH AND INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER TO HELP VISUALIZE THE FARAWAVE CATHETER WITH ICE. AT THIS POINT, THE MAPPING TEAM SUGGESTED A REBOOT OF THE MAPPING SYSTEM WITH THE IDEA OF ADDRESSING THE VISUALIZATION ERROR THAT OCCURRED. WHILE THE TEAM WAS WORKING ON THE REBOOT, THE PHYSICIAN REQUESTED AN ACT. AT THAT TIME, THE PHYSICIAN NOTICED THE PATIENT BECOMING HYPOTENSIVE AND ADDRESSED THE PROCEDURAL ROOM AND STAFF WITH THIS FINDING. THE PHYSICIAN REQUESTED ICE TO INSPECT FOR POSSIBLE PERFORATION BUT WAS UNABLE TO DO SO BECAUSE THE MAPPING SYSTEM WAS DOWN IN A REBOOT PHASE WHICH ALSO TURNED OFF THE ICE MONITOR. THE LAB STAFF ENTERS THE ROOM AND BEGINS TO ASSIST THE PHYSICIAN TO ASSESS THE PATIENT. SHORTLY AFTER, CARDIAC CHEST COMPRESSION BEGAN ON THE PATIENT, BUT THE PATIENT DID NOT MAKE IT. IT WAS FURTHER REPORTED THAT WHEN THE PATIENT WAS FOUND TO BE IN TAMPONADE, A PERICARDIOCENTESIS WAS PERFORMED TO TRY AND SOLVE THE ISSUE. BLOOD AND PLASMA PRODUCTS WERE GIVEN TO THE PATIENT AS WELL BEFORE CALLING FOR CARDIOTHORACIC SURGERY. BUT THE PATIENT PASSED AWAY BEFORE THE SURGERY. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. UPDATED FIELD B5 WITH ADDITIONAL INFORMATION OBTAINED. H6 IMPACT CODE UPDATED, ADDED F2302: BLOOD TRANSFUSION CODE AND F2203: IMAGING REQUIRED. UPDATED FIELD H10 WITH NEW MEDWATCH REPORT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22639899&pc=DSQ | 2916596-2025-04924 | 2025-07-13 04:00 | Death | THORATEC CORPORATION | 2025-07-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Disconnection; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. ON THE MORNING OF (B)(6) 2025 THE PATIENT REPORTED VIA THE HOSPITALS 24-HOUR EMERGENCY LINE SEVERAL REPLACE EMERGENCY BACKUP BATTERY (EBB) AND LOW FLOW ALARMS. THE HOSPITAL ATTEMPTED TO CALL THE PATIENT BACK WITH NO RESPONSE, AFTER SEVERAL ATTEMPTS THE HOSPITAL REACHED OUT TO THE PATIENTS FAMILY FOR A WELLNESS CHECK. THE PATIENT WAS FOUND BY POLICE PASSED AWAY WITH BATTERIES AND DRIVELINE CONNECTED AND THE SYSTEM CONTROLLER ALARMING. THE PATIENTS FRIEND THEN DISCONNECTED THE PATIENT'S SYSTEM CONTROLLER AND BATTERIES UNTIL THE SYSTEM CONTROLLER SHUT OFF. LOG FILE REVIEW REVEALED EBB FAULTS AS WELL AS INTERMITTENT LOW FLOW HAZARD ALARMS ON (B)(6) 2025 BETWEEN 08:37 AM AND 09:49 AM. AT 10:56 AM THE SYSTEM CONTROLLER WAS DISCONNECTED FROM THE PUMP, AND NOT RECONNECTED. THE PERIODIC LOG FILE CAPTURED EBB FAULTS ON (B)(6) 2025 AT 05:35 AM. ADDITIONALLY A SET OF DRIVELINE DISCONNECTS ALARMS WERE NOTED BETWEEN 06:35 AM AND 07:35 AM. THE DRIVELINE DISCONNECT ALARMS OCCURRED WHILE THE PATIENT WAS CONNECTED TO BATTERY POWER.Event Description: THE CAUSE OF DEATH WAS NOT DETERMINED, AND IT WAS NOT KNOWN WHY THE DRIVELINE WAS DISCONNECTED THEN RECONNECTED. THE SITE WAS UNSURE IF THE PATIENT PASSING AWAY WAS DEVICE OR THERAPY RELATED, OR IF THE DEVICE OPERATING AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILES CONFIRMED PUMP STOP EVENTS ASSOCIATED WITH A DISCONNECTION OF THE DRIVELINE; HOWEVER, A SPECIFIC CAUSE FOR THE EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SUBMITTED LOG FILES ALSO CONFIRMED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THESE EVENTS COULD NOT CONCLUSIVELY BE DETERMINED. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED. THE SUBMITTED SYSTEM CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM 13JUL2025. THE FILE CAPTURED LOW FLOW ALARMS ON (B)(6) 2025. ADDITIONALLY, THE DRIVELINE WAS DISCONNECTED ON (B)(6) 2025 THAT RESULTED IN A PUMP STOP. THE PUMP REMAINED DISCONNECTED FOR THE REMAINDER OF THE LOG FILE. THE PUMP APPEARED TO BE OPERATING AS INTENDED WHILE THE DRIVELINE WAS CONNECTED. IT WAS LATER COMMUNICATED THAT THE CAUSE OF DEATH WAS NOT DETERMINED, AND IT WAS NOT KNOWN WHY THE DRIVELINE WAS DISCONNECTED THEN RECONNECTED. THE SITE WAS UNSURE IF THE PATIENT OUTCOME WAS DEVICE OR THERAPY RELATED, OR IF THE DEVICE OPERATING AS EXPECTED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THIS SECTION ALSO ADDRESSES ALL PUMP PARAMETERS. SECTION 2 OF THE IFU, "SYSTEM OPERATIONS" (UNDER "SYSTEM CONTROLLER WARNINGS AND CAUTIONS"), AND SECTION 2 OF THE PATIENT HANDBOOK, ¿HOW YOUR HEART PUMP WORKS¿, INSTRUCT THE USER TO CHECK THE SYSTEM CONTROLLER DRIVELINE CONNECTOR TO CONFIRM THAT THE DRIVELINE IS SECURELY INSERTED IN THE SOCKET. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, THE PUMP STOPS. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, PROMPTLY RECONNECT IT TO RESUME PUMP OPERATION. THE PATIENT HANDBOOK FURTHER EXPLAINS THAT THE PUMP CANNOT RUN WITHOUT POWER. SECTION 2 OF THE IFU ALSO CONTAINS A SECTION TITLED "CONNECTING THE DRIVELINE TO THE SYSTEM CONTROLLER", WHICH PROVIDES INSTRUCTIONS ON CONNECTING/DISCONNECTING THE DRIVELINE TO/FROM THE SYSTEM CONTROLLER AND MAKING SURE THAT IT IS FULLY AND PROPERLY INSERTED INTO THE SYSTEM CONTROLLER SOCKET. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS, AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 6, UNDER "EDUCATING AND TRAINING PATIENTS, FAMILIES, AND CAREGIVERS", EXPLAINS THAT DURING THE PATIENT SELECTION, PREIMPLANT, AND POSTOPERATIVE PERIOD, THE PATIENT MUST RECEIVE INSTRUCTIONS REGARDING THE OPERATION AND CARE OF EVERY SYSTEM COMPONENT (INCLUDING THE DRIVELINE AND WHAT TO DO IN AN EMERGENCY). SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. THE PATIENT HANDBOOK ALSO INSTRUCTS THE USER TO CALL THEIR HOSPITAL CONTACT IF THEY THINK THAT, FOR ANY REASON, ANY PORTION OF THEIR EQUIPMENT IS NOT FUNCTIONING AS USUAL, IS BROKEN, OR THEY ARE UNCOMFORTABLE WITH THE OPERATION OF THE EQUIPMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22621528&pc=NIQ | 2024168-2025-07764 | 2018-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-07-26 04:00 | NIQ | XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED IN A SUMMARY ARTICLE THE LONG TERM CLINICAL EFFECTS OF USE OF THE XIENCE V VERSUS USE OF A NON-ABBOTT DRUG ELUTING STENT (DES) IN PATIENTS FROM 2008-2020. THE STUDY FOUND THAT DURING FOLLOW-UPS THAT OCCURRED WITHIN 5 YEARS POST STENT IMPLANTATION THE XIENCE V DES DEMONSTRATED SIMILAR SAFETY AND EFFICACY AS THE NON-ABBOTT EPS. DURING POST PROCEDURE FOLLOW-UP SAME PATIENTS EXPERIENCED STENT THROMBOSIS, TARGET VESSEL REVASCULARIZATION, MYOCARDIAL INFARCTION AND DEATH. ADDITIONAL INFORMATION CAN BE FOUND IN THE SUMMARY ARTICLE, "LONG-TERM CLINICAL OUTCOMES OF BIODEGRADABLE-VERSUS DURABLE-POLYMER-COATED EVEROLIMUS-ELUTING STENTS IN REAL-WORLD POST-MARKETING STUDY". NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT DEATH AND THE RELATIONSHIP TO PRODUCT, IF ANY, CANNOT BE DETERMINED. THE INFORMATION RECEIVED IN THE COMPLAINT WAS OBTAINED THROUGH A PROACTIVE LITERATURE SEARCH. SPECIFIC MODEL AND LOT NUMBERS WERE NOT AVAILABLE AND CAUSALITY BETWEEN THE DOCUMENTED DEVICE USAGE AND THE PATIENT EFFECTS COULD NOT BE ESTABLISHED. FACTORS THAT MAY CONTRIBUTE TO THE PERFORMANCE OUTCOMES LISTED IN THE ARTICLE INCLUDE, BUT ARE NOT LIMITED TO, DEVICE TO PATIENT INTERACTION, PATIENT ANATOMICAL CONDITIONS, DEVICE TO USER INTERACTION AND MANUFACTURING OR MATERIAL ISSUES. DUE TO THE LIMITED INFORMATION AVAILABLE, THE EXACT CAUSE CANNOT BE DETERMINED. THE REPORTED PATIENT EFFECTS OF DEATH, MYOCARDIAL INFARCTION, STENOSIS, THROMBOSIS/ THROMBUS, HEMORRHAGE/BLOOD LOSS/BLEEDING, RENAL FAILURE, STROKE/CVA AND TRANSIENT ISCHEMIC ATTACK ARE LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. B3, B6, AND D6:ESTIMATED DATES D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE ¿LONG-TERM CLINICAL OUTCOMES OF BIODEGRADABLE- VERSUS DURABLE-POLYMER-COATED EVEROLIMUS-ELUTING STENTS IN REAL-WORLD POST-MARKETING STUDY¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22621945&pc=NIK | 9614453-2025-02950 | 2025-06-27 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-07-26 04:00 | NIK | COBALT¿ XT HF CRT-D MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER THE PATIENT UNDERWENT A CARDIAC PROCEDURE, THE PATIENT EXPERIENCED A PINS AND NEEDLES SENSATION AND DISCOMFORT BEHIND THE SITE OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD). APPROXIMATELY FIVE DAYS LATER, THE PATIENT UNDERWENT A AN UPGRADE PROCEDURE TO A CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WHICH INCLUDED THE IMPLANT OF AN LEFT VENTRICULAR (LV) LEAD AND EXPLANT OF THE ICD. ON THE DAY OF THE PROCEDURE, THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: CONTINUATION OF D10: 383069 LEAD IMPLANTED: (B)(6) 2025, 5076-45 LEAD IMPLANTED: (B)(6) 2016. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22606421&pc=OZP | 2032227-2025-218325 | 2025-06-11 04:00 | Death | MEDTRONIC MINIMED | 2025-07-25 04:00 | OZP | UNKNOWN NGP | Device Alarm System; Excess Flow or Over-Infusion; Insufficient Flow or Under Infusion; Battery Problem | Malaise; Diabetic Ketoacidosis; Confusion/ Disorientation |  |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER HAS PASSED AWAY ON (B)(6) 2025. IT WAS UNKNOWN THAT THE CUSTOMER WAS ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. IT WAS ALSO REPORTED THAT THE INSULIN PUMP OVER INJECTED INSULIN AND THE CAUSE OF DEATH WAS DIABETIC KETOACIDOSIS. ALSO, 12 HOURS PRIOR TO THE DEATH CUSTOMER WAS NOT FEELING WELL AND DISORIENTED. CUSTOMER'S SISTER REPORTED THAT SHE BELIEVES MALFUNCTION OF THE DEVICE AS THE CAUSE OF DEATH AS THE DEVICE HAS MULTIPLE ISSUES WITH THE BATTERY, OVER OR UNDER INJECTING INSULIN, NO ALARMS WHEN THE READINGS WERE NOT WITHIN THE NORMAL RANGE. THE CUSTOMER¿S BLOOD GLUCOSE WAS UNKNOWN. THE EVENT INVOLVED PRODUCT(S) MMT-332A, UNOMEDICAL, UNK_NGP. THE CUSTOMER WAS WEARING THE INSULIN PUMP AT THE TIME OF DEATH. IT WAS UNKNOWN THAT THE CUSTOMER WAS USING SENSORS. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A, UNOMEDICAL, UNK_NGP. Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22606424&pc=OYC | 2032227-2025-218326 | 2007-02-21 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-25 04:00 | OYC | PARADIGM REAL-TIME INSULIN INFUSION PUMP MMT-722NALL | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150019 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY. THE EVENT INVOLVED PRODUCT(S) MMT-722NAL. TROUBLESHOOTING WAS PERFORMED AND FOUND THE REASON FOR PASSING WAS HEART COMPLICATION. THE DATE OF THE EVENT WAS 21 FEBRUARY 2007. THE INSULIN PUMP WAS NOT WORN AT THE TIME OF PASSING. NO HARM REQUIRING MEDICAL INTERVENTION WAS REPORTED. MMT-722NAL WAS RETURNED FOR ANALYSIS. Manufacturer Narrative: THE CUSTOMER PASSED AWAY ON THE EVENT DATE LISTED ON SMARTSOLVE AS 21-FEB-2007. HOWEVER, THE CUSTOMER'S NOTES LISTED, PASSED AWAY ON (B)(6) 2007. ALL OPERATING CURRENTS ARE WITHIN SPECIFICATION. OFF NO POWER ALARM FUNCTIONS PROPERLY. THE PUMP PASSED THE DISPLACEMENT TEST, REWIND, SELF TEST AND A21 ERROR TEST. THE PUMP ALARMED A33 DURING THE PRIME/A33 TEST AT 4.0 LBS. UNABLE TO PERFORM THE BASIC OCCLUSION TEST, OCCLUSION TEST, EXCESSIVE NO DELIVERY TEST AND THE DAT DUE TO A33 ALARM. APPLY 1K OHM RESISTANCE ACROSS PIN 1 AND PIN 3 ON THE FSR CONNECTOR. VERIFY THAT SEATING OCCURRED AND THEN DELIVER A 2.0 UNIT BOLUS. VERIFY THAT A ¿NO DELIVERY¿ ALARM OCCURRED. THE PUMP SEATS IMMEDIATELY AND DECLARES A ¿NO DELIVERY¿ ALARM IMMEDIATELY (THE SYSTEM IS RESPONDING AS EXPECTED). A33 DURING THE PRIME/A33 TEST WAS DUE TO DUE TO FAULTY FSR (GOLD). THE FOLLOWING WERE NOTED DURING VISUAL INSPECTION: MINOR SCRATCHED DISPLAY WINDOW, CRACKED CASE AT DISPLAY WINDOW CORNER, SCRATCHED RESERVOIR TUBE WINDOW, MUSSING END CAP STICKER AND CRACKED RESERVOIR TUBE LIP. THE TEST P-CAP AND RESERVOIR DOES LOCK IN PLACE IN THE RESERVOIR COMPARTMENT. HISTORY DOWNLOAD WAS SUCCESSFUL USING THDS AND CARELINK UPLOAD WAS SUCCESSFUL. THE PUMP WAS RECEIVED WITHOUT A BATTERY. NO DAMAGE NOTED ON THE ORIGINAL BATTERY CAP. UNABLE TO VERIFY DAILY INSULIN TOTAL, BASAL INSULIN DELIVERED TOTAL AND BOLUS INSULIN DELIVERED TOTAL FOR EVENT DATES OF 21-FEB-2007 AND 02/06/2007 DUE TO NO DATA AVAILABLE IN THE PUMP HISTORY FILE. UNABLE TO PERFORM THE HISTORY REVIEW 1 WEEK PRIOR TO THE EVENT DATES OF 21-FEB-2007 AND 02/06/2007 IN THE PUMP HISTORY FILE DUE TO NO DATA AVAILABLE. PUMP WAS CUT OPEN TO PERFORM VISUAL INSPECTION AND FOUND MOISTURE DAMAGE ON THE ELECTRONIC ASSEMBLY, CORRODED TRANSDUCER AND CORRODED MOTOR HOME SWITCH. UNABLE TO COMPLETE THE FUNCTIONAL TESTING DUE TO A33 ALARM DURING THE PRIME/A33 TEST. A33 CONFIRMED. RETAINER RING DAMAGE CONFIRMED (FOUND CRACKED RESERVOIR TUBE LIP DURING VISUAL INSPECTION). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22607298&pc=OZP | 2032227-2025-218770 | 2025-06-11 04:00 | Death | MEDTRONIC MINIMED | 2025-07-25 04:00 | OZP | UNKNOWN NGP | Excess Flow or Over-Infusion; Insufficient Flow or Under Infusion; Battery Problem; Insufficient Information | Malaise; Diabetic Ketoacidosis; Confusion/ Disorientation |  |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY WITH ON 11TH JUNE'25. THE EVENT INVOLVED PRODUCT(S) UNK_NGP, MMT-332A & UNOMEDICAL. THE CUSTOMER'S SISTER REPORTED THAT THEY BELIEVED THE INSULIN PUMP WAS EITHER UNDER & OVER INJECTING INSULIN, HAD MULTIPLE BATTERY ISSUES, NO ALARMS WHEN THE READINGS WERE NOT WITHIN THE NORMAL RANGE. THE CUSTOMER'S SISTER FELT THE INSULIN PUMP MALFUNCTIONED CAUSING THE PATIENT TO BECOME DISORIENTED AND WASN'T FEELING WELL, RESULTING IN THEIR DEATH DUE TO DIABETES KETOACIDOSIS. IT WAS UNKNOWN IF THE INSULIN PUMP WAS USED WITHIN 48 HOURS OF THE REPORTED EVENT OR IF THE AUTOMODE/SMART GUARD FEATURE WAS ACTIVE. NO FURTHER DETAILS WERE PROVIDED. THE PRODUCT(S) UNK_NGP, MMT-332A & UNOMEDICAL WILL NOT BE RETURNED FOR ANALYSIS. Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22607940&pc=OWB | 3003768277-2025-006921 | 2025-06-27 04:00 | Death | PHILIPS MEDICAL SYSTEMS NEDERLAND B.V. | 2025-07-25 04:00 | OWB | ALLURA XPER FD20/10 & ALLURA XPER FD20/20 | Positioning Failure | Insufficient Information | K033737 |  | 1 | Event Description: IT WAS REPORTED TO PHILIPS THAT DURING A CEREBRAL ANGIOGRAPHY, THE USER EXPERIENCED AN ISSUE WITH THE MOVEMENT OF THE SYSTEM'S C-ARM. THE REPORT INDICATED THAT THERE WAS A PATIENT DEATH; HOWEVER, NO FURTHER INFORMATION REGARDING THE ALLEGED DEATH HAS BEEN PROVIDED TO DATE. AN INVESTIGATION INTO THIS COMPLAINT HAS BEEN INITIATED BY PHILIPS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22608381&pc=NIK | 9614453-2025-02933 | 2025-06-07 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-07-25 04:00 | NIK | COBALT¿ HF CRT-D MRI SURESCAN¿ | Failure to Deliver Shock/Stimulation; Device Sensing Problem | Tachycardia; Loss of consciousness | P010031 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S REPRESENTATIVE THAT THE PATIENT IS DECEASED AND STATED THAT ON THE DAY OF THE PATIENT'S DEATH, THE PATIENT COLLAPSED AND THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) FAILED TO DELIVER THERAPY. THE PATIENT'S REPRESENTATIVE STATED THEY WERE CONTACTED BY THE PATIENT'S PHYSICIAN AFTER THE PATIENT PASSED AWAY AND WERE INFORMED THAT ON THE DATE OF DEATH, THE DEVICE DETECTED AN EPISODE OF UNSUSTAINED TACHYCARDIA WHERE NO THERAPIES WERE DELIVERED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22608383&pc=OZD | 1220648-2025-30713 | 2025-07-09 04:00 | Death | ABIOMED, INC. | 2025-07-25 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT DEVELOPED A VANCOMYCIN-RESISTANT ENTEROCOCCI INFECTION AND YEAST INFECTION IN THE PATIENT'S BLOOD. PROVIDERS STATED THIS IS FROM THE STERNOTOMY SITE. DUE TO INFECTION, THE PLAN IS TO WEAN THE PATIENT FROM IMPELLA SUPPORT. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS, ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22608389&pc=NGV | 2124215-2025-50084 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-25 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 11 | Event Description: THIS REPORT SUMMARIZES 11 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 66 TO (B)(6) YEARS PATIENT SEX: 5 FEMALES; 6 MALES. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 78 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22608391&pc=NGV | 2124215-2025-50069 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-25 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 3 | Event Description: THIS REPORT SUMMARIZES 3 DEATH EVENTS FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 82 TO 88 YEARS. PATIENT SEX: 1 FEMALE; 2 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 2 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22609006&pc=PNJ | 9612164-2025-03641 | 2025-02-04 05:00 | Death | MEDTRONIC IRELAND | 2025-07-25 04:00 | PNJ | MICRA AV | Appropriate Term/Code Not Available | Insufficient Information | P150033 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING TREATMENT OPTIONS FOR ATRIOVENTRICULAR BLOCK (AVB) WITH A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE AT A TWO-YEAR FOLLOW-UP. THERE WERE PATIENTS WHO EXPERIENCED HEART FAILURE HOSPITALIZATIONS AND ONE PATIENT WITH A VASCULAR ACCESS HEMATOMA WHICH WAS UNEVENTFUL AND DID NOT REQUIRE INTERVENTION. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/74 YEARS OLD. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: ATRIOVENTRICULAR BLOCK TREATMENT: PACING SITE, AV SYNCHRONY, OR BOTH? JOURNAL OF CLINICAL MEDICINE. 2025. 14, 980. DOI: 10.3390/JCM14030980. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22609330&pc=OZD | 1220648-2025-30720 | 2025-07-11 04:00 | Death | ABIOMED, INC. | 2025-07-25 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND AFTER THE IMPLANT, DISTAL PULSES WERE UNABLE TO BE FOUND. A FEMOROFEMORAL BYPASS WAS PLACED AND PULSES WERE STILL UNABLE TO BE FOUND. SUPPORT WAS CONTINUED. ADDITIONALLY, THE PATIENT HAD RED URINE WITH HEMOLYZING LABS. THE P LEVEL WAS LOWERED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER. ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION. ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22609382&pc=NPT | 9617601-2025-00686 | 2025-06-27 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-07-25 04:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Gradient Increase; Patient Device Interaction Problem; Thickening of Material | Endocarditis; Low Blood Pressure/ Hypotension; Thrombosis/Thrombus | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 4 YEARS AND 10 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, A CHANGE IN MEAN GRADIENT WAS OBSERVED IN THE PATIENT'S AORTIC VALVE. SUBSEQUENTLY, A POST-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED TO THE VALVE; HOWEVER, THE PATIENT LOST PRESSURE AFTER BALLOON INFLATION. CARDIOPULMONARY RESUSCITATION (CPR) WAS ATTEMPTED BUT DID NOT RESOLVE THE PRESSURE LOSS, AND THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, THE VALVE AND PROCEDURE DID NOT CAUSE OR CONTRIBUTE TO THE DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE MEAN GRADIENT AFTER THE VALVE WAS IMPLANTED WAS 5 MMHG. THE PATIENT'S MEAN GRADIENT WAS 50 MILLIMETERS OF MERCURY (MM HG) WHEN THEY PRESENTED TO THE HOSPITAL. ADDITIONALLY, IT WAS NOTED THAT THE PATIENT HAD ENDOCARDITIS THAT COULD HAVE CONTRIBUTED TO THE ELEVATED GRADIENT. THE IMPLANT DEPTH WAS 4 MM ON THE NON-CORONARY CUSP (NCC) AND 4 MM ON THE LEFT CORONARY CUSP (LCC). IT WAS FURTHER REPORTED THAT LEAFLET THICKENING AND A THROMBUS WERE POSSIBLY OBSERVED ON THE BIOPROSTHETIC VALVE. THE VALVE WAS INITIALLY IMPLANTED IN THE PATIENT'S NATIVE ANNULUS. PER THE PHYSICIAN, THE POST-IMPLANT BALLOON AORTIC VALVULOPLASTY POSSIBLY CONTRIBUTED TO THE DEATH AS IT LEAD TO THE LOSS OF PRESSURE. ADDITIONALLY, IT WAS THOUGHT THAT THE PATIENT'S PRE-EXISTING COMORBIDITIES CONTRIBUTED TO THE DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22609430&pc=MKJ | 3004123209-2025-00202 | 2021-07-04 04:00 | Death | HEARTSINE TECHNOLOGIES LTD | 2025-07-25 04:00 | MKJ | PACKAGE,500P,PP01,ES,500-BAS-AS-10 | Battery Problem | Cardiac Arrest | P160008 |  | 1 | Event Description: THE CUSTOMER CONTACTED HEARTSINE TO REPORT THAT THEIR DEVICE HAS AN EXPIRED BATTERY AND WAS USED ON A CARDIAC ARREST PATIENT. THE PATIENT DID NOT SURVIVE THE EVENT. A CLINICAL REVIEW WILL BE COMPLETED TO DETERMINE DEVICE CONTRIBUTION. Manufacturer Narrative: HEARTSINE HAS REQUESTED RETURN OF THE DEVICE FOR INVESTIGATION. THE CUSTOMER HAS STATED THEY WILL NOT BE SENDING THE DEVICE TO HEARTSINE DUE TO LOCAL LAWS. PICTURES OF THE DEVICE FILES WERE SENT TO HEARTSINE AND A CLINICAL REVIEW WAS COMPLETED. FIFTEEN SHOCKS WERE DELIVERED WITH THE HEARTSINE SAM PAD. THE ELECTRONIC FILE DOWNLOADED FROM THE AED IS REQUIRED TO PROVIDE A FULL CLINICAL AND USABILITY REVIEW. HEARTSINE CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. HEARTSINE REQUESTS PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. THE CUSTOMER PROVIDED HEARTSINE WITH THE AVAILABLE PATIENT INFORMATION. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK. THE SAM 500P DEVICE IS NOT AVAILABLE FOR DISTRIBUTION IN THE UNITED STATES BUT IS SIMILAR TO THE SAM 350P AND 450P WHICH IS DISTRIBUTED IN THE UNITED STATES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22609977&pc=NPT | 2015691-2025-06054 | 2025-06-30 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-07-25 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem |  | P140031 |  | 1 | Event Description: AS REPORTED BY OUR FIELD CLINICAL SPECIALIST, AFTER DEPLOYMENT OF A 26 MM SAPIEN 3 ULTRA RESILIA (S3UR) VALVE DURING A TRANSFEMORAL TAVR PROCEDURE, THE PATIENT DEVELOPED VENTRICULAR FIBRILLATION, AND DIRECT CURRENT (DC) CARDIOVERSION WAS PERFORMED. AFTER DC, THE PATIENT WENT INTO CARDIAC ARREST, AND CHEST COMPRESSIONS WERE INITIATED. A RETURN OF SPONTANEOUS CIRCULATION (ROSC) WAS ACHIEVED WITHIN SECONDS, BUT THE SYSTOLIC BLOOD PRESSURE (SBP) SURGED TO 300 MMHG. THE PATIENT THEN EXPERIENCED ANOTHER CARDIAC ARREST, AND CHEST COMPRESSIONS WERE RESUMED. PERCUTANEOUS CARDIOPULMONARY SUPPORT (PCPS) WAS ESTABLISHED, AND TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) REVEALED FINDINGS CONSISTENT WITH ASCENDING AORTIC DISSECTION AND CARDIAC TAMPONADE. PERICARDIOCENTESIS WAS ATTEMPTED, BUT HEMOSTASIS COULD NOT BE ACHIEVED, LEADING TO A DECISION TO PERFORM OPEN-CHEST HEMOSTASIS. ALTHOUGH BLEEDING DECREASED WITH OPEN-CHEST INTERVENTION, THE FAMILY WAS INFORMED, AND THEY DECLINED ADDITIONAL PROCEDURES SUCH AS ASCENDING AORTIC REPLACEMENT. THEREFORE, THE CHEST WAS CLOSED WITHOUT FURTHER INTERVENTION, AND THE PATIENT WAS RETURNED TO THE ICU ON PCPS. THE FALSE LUMEN OF THE ASCENDING AORTIC DISSECTION RUPTURED, CAUSING BLEEDING AND CARDIAC TAMPONADE. THE PATIENT PASSED AWAY AT MIDNIGHT ON THE DAY OF THE TAVR PROCEDURE. THE PHYSICIAN BELIEVED THAT THE ADMINISTRATION OF MEDICATION CAUSED THE BLOOD PRESSURE TO RISE TO 300 MMHG, AND THAT THE ASCENDING AORTIC DISSECTION AND RUPTURE MAY HAVE BEEN TRIGGERED BY CHEST COMPRESSIONS. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. AORTIC DISSECTION OR RUPTURE MAY OCCUR WHEN MULTIPLE ATTEMPTS ARE MADE TO CROSS THE STENOTIC NATIVE VALVE, AND/OR WHEN EXCESSIVE FORCE IS USED. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). THE THV TRAINING MANUALS GUIDE TO FACILITATE SAFE CROSSING OF THE NATIVE OR FAILING VALVE, INCLUDING CAMERA PROJECTIONS, HANDLING DURING ADVANCEMENT, AND TROUBLESHOOTING TECHNIQUES IF DIFFICULTY IS ENCOUNTERED. AS STATED, EXCESSIVE FORCE SHOULD NOT BE USED WHEN THE DEVICE HAS DIFFICULTY CROSSING THE STENOTIC VALVE. ADDING TENSION TO THE WIRE, PULLING BACK THE SYSTEM TO RE-ORIENT THE VALVE, AS NEEDED, AND TORQUING THE FLEX CATHETER MAY HELP SOLVE THE PROBLEM. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PROCEDURAL FACTORS (CHEST COMPRESSIONS) MAY HAVE CAUSED OR CONTRIBUTED TO THE AORTIC DISSECTION EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22610186&pc=BZD | 2518422-2025-108242 | 2025-07-08 04:00 | Death | RESPIRONICS, INC. | 2025-07-25 04:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THIS COMPLAINT HAS BEEN REVIEWED AND IT HAS BEEN DETERMINED THAT THE CUSTOMER HAS ALLEGED PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22610296&pc=OZD | 1220648-2025-30724 | 2025-06-29 04:00 | Death | ABIOMED, INC. | 2025-07-25 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE INSERTION OF IMPELLA 5.5 DEVICE TO A PATIENT UNDERGOING A HIGH-RISK PERCUTANEOUS CORONARY INTERVENTION (PCI) WAS UNSUCCESSFUL. THE DEVICE WAS UNABLE TO PASS THROUGH THE AXILLARY ARTERY, PROCEED DOWN TO THE LEFT VENTRICLE AFTER ATTEMPTS AND DIFFERENT TECHNIQUES. CONSEQUENTLY, THE IMPELLA 5.5 IMPLANT WAS ABORTED, AND THE PHYSICIAN ATTEMPT WITH AN IMPELLA CP WHICH WAS SUCCESSFULLY PLACED. AFTER SUCCESSFUL INSERTION OF THE IMPELLA CP, THE PHYSICIAN PROCEEDED WITH PCI OF THE LEFT MAIN INTERNAL MAMMARY ARTERY TO THE LEFT ANTERIOR DESCENDING ARTERY WITH PERCUTANEOUS CORONARY ANGIOPLASTY AND STENTS (X2) PLACEMENT. THE PATIENT WAS TRANSFERRED TO THE UNIT INTUBATED AND SEDATED NOTED TO BE IN STABLE CONDITION. TWO DAYS AFTER START OF THE IMPELLA CP MECHANICAL CIRCULATORY SUPPORT, THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA BUT RESOLVED. THE NEXT DAY, THE PATIENT WAS TAKEN FOR A RIGHT HEART CATHETERIZATION AND NEW PULMONARY ARTERY CATHETER PLACEMENT. THE PLAN WAS TO WEAN AND EXPLANT THE IMPELLA CP. HOWEVER, THREE DAYS LATER, THE PATIENT REMAINED INTUBATED AND SEDATED, FAMILY DECIDED TO WITHDRAW CARE, AND THE PATIENT WOULD SUBSEQUENTLY EXPIRE. Manufacturer Narrative: ALTHOUGH CARE WAS WITHDRAWN, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR GIVEN THE ARRHYTHMIC EVENT THAT OCCURRED THREE DAYS PRIOR TO THE PATIENT¿S DEMISE. THE UNTOWARD IMPACT, IF ANY, IS UNKNOWN. THIS IS ONE OF TWO DEVICES RELATED TO THE PATIENT'S IMPLANT EXPERIENCE. REFER TO MANUFACTURING REPORT NUMBER 1220648-2025-30712 FOR THE REPORT ON THE IMPELLA 5.5 DEVICE THAT WAS UNABLE TO BE DELIVERED. THE IMPELLA CP DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22610323&pc=DSQ | 3004582654-2025-00042 | 2025-07-07 04:00 | Death | BERLIN HEART GMBH | 2025-07-25 04:00 | DSQ | EXCOR® | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160035 |  | 1 | Event Description: BERLIN HEART GMBH CLINICAL AFFAIRS WAS INFORMED BY THE ITALIAN DISTRIBUTOR ON 2025-07-08 THAT A PATIENT WITH EXCOR BLOOD PUMP PU VALVES; 10 ML IN/OUT; Ø 6 MM; SUFFERED CEREBRAL HEMORRHAGE AND WAS DECEASED ON (B)(6) 2025. THE PATIENT'S DEATH WAS NOT CAUSED BY A THROMBUS IN THE EXCOR PUMP. IT WAS DUE TO A SERIES OF COMPLICATIONS / EVENTS. Manufacturer Narrative: THE LVAD EXCOR BLOOD PUMP PU VALVES; 10 ML IN/OUT; Ø 6 MM; SN (B)(6) WAS IN USE ON THE PATIENT FROM (B)(6) 2025 UNTIL IT WAS REMOVED FROM THE PATIENT ON (B)(6) 2025. THE EVENT OCCURRED ON (B)(6) 2025, WHICH IS (13 DAYS) AFTER THE PUMP WAS PLACED ON THE PATIENT. WE HAVE REVIEWED THE PRODUCTION RECORDS OF THE EXCOR BLOOD PUMP AND THE PUMP WAS PRODUCED ACCORDING TO OUR SPECIFICATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22610637&pc=QLG | 3016525500-2025-00058 | 2025-03-19 04:00 | Death | BIGFOOT BIOMEDICAL, INC | 2025-07-25 04:00 | QLG | BIGFOOT UNITY DIABETES MANAGEMENT SYSTEM | Application Program Freezes, Becomes Nonfunctional | No Clinical Signs, Symptoms or Conditions | K202145 |  | 1 | Event Description: A CUSTOMER REPORTED RECEIVING DEVICE ERROR FROM A BIGFOOT UNITY DEVICE. THIS IS A KNOWN ISSUE FOR THE SERIAL NUMBER ASSOCIATED WITH THIS COMPLAINT, ADDRESSED BY THE FIELD ACTION FA-001 (FDA RECALL 3016525500/05162025/C). THERE WAS NO REPORT OF DEATH, SERIOUS INJURY, OR MISTREATMENT ASSOCIATED WITH THIS EVENT.Event Description: A CUSTOMER REPORTED RECEIVING DEVICE ERROR FROM A BIGFOOT UNITY DEVICE. THIS IS A KNOWN ISSUE FOR THE SERIAL NUMBER ASSOCIATED WITH THIS COMPLAINT, ADDRESSED BY THE FIELD ACTION FA-001 (FDA RECALL 3016525500/05162025/C). THERE WAS NO REPORT OF DEATH, SERIOUS INJURY, OR MISTREATMENT ASSOCIATED WITH THIS EVENT. Manufacturer Narrative: FOLLOWING THE SUBMISSION OF A CORRECTION/REMOVAL FOR CERTAIN BIGFOOT UNITY DEVICES, IT HAS BEEN DETERMINED THAT THIS COMPLAINT IS A REPORTABLE MALFUNCTION. Manufacturer Narrative: THIS SERVES AS A CORRECTION REPORT. SECTION D4 (PRIMARY UDI NUMBER) WAS INCORRECTLY DOCUMENTED IN THE PREVIOUS REPORT. CORRECTION HAS BEEN MADE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22610717&pc=LIT | 3007635982-2025-00164 | 2025-06-27 04:00 | Death | CORDIS CORPORATION | 2025-07-25 04:00 | LIT | SABERX RADIANZ | Burst Container or Vessel | No Clinical Signs, Symptoms or Conditions | K201377 |  | 1 | Event Description: AS REPORTED, AFTER PLACING A SMART RADIANZ (UNKNOWN SIZE), POST-DILATION WAS PERFORMED USING A 9X40 SABERX RADIANZ PERCUTANEOUS TRANSLUMINAL ANGIOPLASTY (PTA) BALLOON. THERE WAS NO ISSUE OR MALFUNCTION WITH THE SMART STENT THAT WAS IMPLANTED. WITH A RATED BURST PRESSURE (RBP) OF 10 ATM, THE TRAINEE DOCTOR MISTAKENLY APPLIED 20 ATM. THE BALLOON RUPTURED AND THE VESSEL ALSO RUPTURED. THE BLOOD PRESSURE DROPPED. A NON-CORDIS COVERED STENT WAS PLACED FOR HEMOSTASIS. HOWEVER, THE PATIENT DIED THE NEXT DAY. THE PATIENT WAS UNDERGOING A TRANS-RADIAL INTERVENTION (TRI) OF THE ILIAC ARTERY. THE PRODUCT WAS STORED PROPERLY ACCORDING TO THE INSTRUCTIONS FOR USE (IFU). THERE WAS NO DIFFICULTY REMOVING THE PRODUCT FROM THE HOOP. THERE WAS NO DIFFICULTY REMOVING THE PROTECTIVE BALLOON COVER. THERE WAS NO DIFFICULTY REMOVING THE STYLET OR ANY OF THE STERILE PACKAGING COMPONENTS. NO KINKS OR OTHER DAMAGES WERE NOTED PRIOR TO INSERTING THE PRODUCT INTO THE PATIENT. THE DEVICE WAS PREPPED PER THE IFU. THE DEVICE PREPPED NORMALLY, MAINTAINING NEGATIVE PRESSURE. THE LESION WAS MINIMALLY CALCIFIED. THE VESSEL TORTUOSITY WAS MODERATE. THE PERCENTAGE OF STENOSIS WAS 90%. THE CATHETER WAS NEVER IN AN ACUTE BEND. THE BALLOON CATHETER DID NOT KINK WHILE BEING USED. THE CONTRAST TO SALINE RATIO WAS 50:50. THE DEVICE WILL BE RETURNED FOR EVALUATION. Manufacturer Narrative: THIS DEVICE IS AVAILABLE FOR ANALYSIS, BUT THE ENGINEERING REPORT IS NOT YET AVAILABLE. HOWEVER, IT WILL BE SUBMITTED WITHIN 30 DAYS UPON RECEIPT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22610772&pc=NPT | 9617601-2025-00688 | 2025-07-02 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-07-25 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Cardiac Arrest; Chest Pain; Thrombosis/Thrombus | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 8 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT EXPERIENCED CHEST PAIN DURING HOSPITALIZATION FOR GASTROINTESTINAL TREATMENT. THE PATIENT WENT INTO CARDIAC ARREST DUE TO ACUTE CORONARY SYNDROME (ACS). A STENT WAS PLACED WITH PERCUTANEOUS CORONARY INTERVENTION (PCI) AND CORONARY ARTERY TREATMENT WAS PERFORMED. HOWEVER, THE PATIENT DIED DUE TO ACS. IT WAS NOTED THAT A POST-OPERATIVE COMPUTED TOMOGRAPHY (CT) SCAN REVEALED A THROMBUS WHICH MAYBE HAVE BEEN A CONTRIBUTING FACTOR. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. B2. YEAR VALID, MONTH AND DATE APPROXIMATE. B3. YEAR VALID, MONTH AND DATE APPROXIMATE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22611303&pc=DRM | MW5173298 | 2025-07-19 04:00 | Death | ZOLL CIRCULATION, INC | 2025-07-25 04:00 | DRM | AUTO-PULSE RESUSCITATION SYSTEM | Defective Device; Device Displays Incorrect Message | Asystole |  |  | 1 | Event Description: THE PATIENT PRESENTED IN ASYSTOLE, PULSELESS AN APNEIC WITH CARDIOPULMONARY RESUSCITATION (CPR) IN PROGRESS. THE AUTO-PULSE WAS APPLIED AND MALFUNCTIONED SEVERAL TIMES GIVING A CODE 17 AND WOULD NOT OPERATE PROPERLY. HAD TO RESTART DEVICE SEVERAL TIMES TO OPERATE. PULSE WAS NEVER REGAINED AND PATIENT WAS PRONOUNCED ON SCENE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22611340&pc=NPS | 2135147-2025-04124 | 2016-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-25 04:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "RIGHT VENTRICULAR¿PULMONARY ARTERY COUPLING IN TRICUSPID REGURGITATION: PROGNOSTIC VALUE AND IMPACT OF TREATMENT STRATEGY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE PROGNOSTIC IMPLICATIONS OF RVPAC IN A EUROPEAN REGISTRY OF PATIENTS WITH TRICUSPID REGURGITATION UNDERGOING EITHER T-TEER OR MEDICAL MANAGEMENT. DEVICES MENTIONED INCLUDE TRICLIP AND NON-ABBOTT DEVICE. THE ARTICLE CONCLUDED THAT POORER RVPAC REFLECTS HIGHER BASELINE RISK AND MORTALITY, REGARDLESS OF TREATMENT. T-TEER IS ASSOCIATED WITH BETTER SURVIVAL ACROSS A RANGE OF RVPAC VALUES, INCLUDING THOSE LESS THAN PREVIOUSLY SUGGESTED THRESHOLDS. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS KARL-PHILIPP ROMMEL AT UNIVERSITÄTSMEDIZIN JOHANNES GUTENBERG-UNIVERSITY INSTITUTION WITH CORRESPONDING EMAIL **. THE TIME FRAME OF THE STUDY WAS JANUARY 2016 TO DECEMBER 2022. A TOTAL OF 946 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION OR FLUTTER, CORONARY ARTERY DISEASE, LEFT VENTRICULAR EJECTION FRACTION. (B)(6), UNK TRICLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22612808&pc=QMG | 3015053858-2025-00093 | 2025-06-25 04:00 | Death | SHOCKWAVE MEDICAL, INC. | 2025-07-25 04:00 | QMG | SHOCKWAVE C2 CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels; Cardiac Tamponade | P200039 |  | 1 | Event Description: THE PATIENT, AN 80 YEAR OLD FEMALE, HAD TOTAL OCCLUSION IN THE RIGHT CORONARY ARTERY (RCA), HIGHLY CALCIFIED DIFFUSE STENOSIS IN THE LEFT MAIN TRUNK (LMT) TO THE LEFT ANTERIOR DESCENDING ARTERY (LAD), IN THE LAD DISTAL, AND IN THE OSTIUM (OS) OF THE CIRCUMFLEX ARTERY (CX). THE PATIENT WAS INTUBATED PRIOR TO THE PROCEDURE DUE TO BAD RESPIRATORY STATUS, AND THE PATIENT HAD A BLOOD PRESSURE READING BELOW 100. ABIOMED IMPELLA WAS INSERTED PRIOR TO THE TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE DUE TO THE EXTREME EMERGENCY SITUATION. AN ATTEMPT WAS MADE TO TREAT THE HIGHLY CALCIFIED LESION IN THE LEFT CIRCUMFLEX ARTERY (LCX) USING THE BOSTON SCIENTIFIC ROTABLATOR 1.5 MM USING A 6 FR EBU 3.75 SHEATH, BUT IT COULD NOT CROSS THE MAIN LESION DIAMETER (MLD). HOWEVER, A BOSTON SCIENTIFIC ROTABLATOR 1.25 MM WAS ABLE TO CROSS THE LESION. ANOTHER 7FR EBU 3.75 SHEATH WAS IN PLACE FOR AN ADDITIONAL GUIDEWIRE, LIKELY TO MINIMIZE THE NEED TO SWAP SHEATHS THEY ALREADY HAD IT IN PLACE WITH A BACK UP DOUBLE-GUIDEWIRE (GW). FOLLOWING THIS, INTRAVASCULAR ULTRASOUND (IVUS) WAS PERFORMED TO ASSESS THE CONDITION OF THE VESSEL. A CALCIFICATION 3+ WAS OBSERVED DURING THE IVUS ASSESSMENT. TO AVOID DISRUPTING HEMODYNAMIC STATUS THROUGH ADDITIONAL ATHERECTOMY, A SHOCKWAVE C2 CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER WAS UTILIZED FOR DILATION. UNFORTUNATELY, VESSEL PERFORATION OCCURRED AT THE TREATMENT SITE DURING THIS PROCEDURE. A 2.0 MM BALLOON CATHETER WAS SUCCESSFULLY EMPLOYED TO ACHIEVE HEMOSTASIS AFTER THE PERFORATION AND AN ABBOTT XIENCE DRUG-ELUTING STENT (DES) MEASURING 2.5 X 38 MM WAS PLACED. DESPITE SUCCESSFUL HEMOSTASIS, CARDIAC TAMPONADE DEVELOPED, REQUIRING DRAINAGE TO ADDRESS THE COMPLICATION. AFTER SUCCESSFUL MANAGEMENT OF THE COMPLICATIONS AND ACHIEVING HEMOSTASIS, THE PATIENT WAS MOVED BACK TO THE HOSPITAL ROOM FOR FURTHER MONITORING AND CARE. AFTER THE PROCEDURE, AN IVUS EXAMINATION PERFORMED PRIOR TO THE IVL REVEALED THAT THE THREE LAYERS OF THE VESSEL WERE PARTIALLY DISRUPTED, INDICATING THAT THERE WAS DAMAGE CAUSED BY THE ROTABLATOR ABLATION, WHICH ULTIMATELY RESULTED IN A RUPTURE DURING THE IVL TREATMENT. THE PATIENT REMAINED HOSPITALIZED FOR THE TREATMENT OF CARDIAC ARREST. AS OF THE DATE OF AWARENESS, THE PATIENT HAD BEEN EXTUBATED, INDICATING AN IMPROVEMENT IN HER RESPIRATORY CONDITION. THE PHYSICIAN BELIEVES THAT THE ADVERSE EVENTS WERE PRIMARILY CAUSED BY VESSEL DAMAGE RESULTING FROM THE ATHERECTOMY PERFORMED PRIOR TO THE IVL. THE PHYSICIAN DOES NOT BELIEVE THAT THE SHOCKWAVE C2 CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER CONTRIBUTED TO OR CAUSED THESE ADVERSE EVENTS. THE DEVICE WAS USED IN ACCORDANCE WITH ITS INSTRUCTIONS FOR USE (IFU) AND CONTRAINDICATIONS, AND THERE WAS NO DEVICE MALFUNCTION. THE PATIENT'S CURRENT STATUS, INCLUDING ANY WORSENING SYMPTOMS OR RECOVERY DETAILS, IS YET TO BE ANNOUNCED. PER ADDITIONAL INFORMATION RECEIVED ON JULY 9, 2025, SHOCKWAVE MEDICAL JAPAN WAS INFORMED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO DISSEMINATED INTRAVASCULAR COAGULATION (DIC). Manufacturer Narrative: THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION CANNOT BE PERFORMED. BASED ON THE REPORTED INFORMATION, NO INFORMATION WAS PROVIDED SUGGESTING THE IVL DEVICE CONTRIBUTED TO THE VESSEL PERFORATION, CARDIAC TAMPONADE, AND PATIENT DEATH. THERE WAS NO REPORTED IVL CATHETER ISSUE. THE PHYSICIAN BELIEVES THAT THE ADVERSE EVENTS WERE PRIMARILY CAUSED BY VESSEL DAMAGE RESULTING FROM THE ATHERECTOMY PERFORMED PRIOR TO THE IVL . THE PHYSICIAN DOES NOT BELIEVE THAT THE SHOCKWAVE C2 CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER CONTRIBUTED TO OR CAUSED THESE ADVERSE EVENTS. SHOCKWAVE MEDICAL HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO BEING DISTRIBUTED. A REVIEW OF THE MANUFACTURING AND TEST DOCUMENTATION FOR SUBJECT LOT DOES NOT REVEAL ANY ISSUES WITH THE MANUFACTURING OF THE DEVICE. THE DEVICE PASSED ALL OF SHOCKWAVE MEDICAL, INC. ACCEPTANCE CRITERIA PRIOR TO SHIPPING. COMPLAINT TRENDS WILL CONTINUE TO BE MONITORED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22612989&pc=DSQ | 2916596-2025-04889 | 2025-07-16 04:00 | Death | THORATEC CORPORATION | 2025-07-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD BEEN ON HOSPICE CARE SINCE (B)(6) 2025 DUE TO A CHRONIC DRIVELINE INFECTION. THE PATIENT ULTIMATELY PASSED AWAY ON (B)(6) 2025. NO AUTOPSY TO BE COMPLETED.Event Description: THE ONSET OF THE PATIENT'S CHRONIC DRIVELINE INFECTION WAS REPORTED UNDER MFR #: 2916596-2023-06861. Manufacturer Narrative: SECTION H6: HEALTH EFFECT - CLINICAL CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22613179&pc=MTD | 2029046-2025-02444 | 2025-07-01 04:00 | Death | BIOSENSE WEBSTER INC | 2025-07-25 04:00 | MTD | OCTARAY MAPPING CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade | K193237 |  | 1 | Event Description: IT WAS REPORTED THAT DURING CARDIAC ABLATION PROCEDURE WHICH INCLUDED THE USE OF AN OCTARAY MAPPING CATHETER, THE PATIENT EXPERIENCED TAMPONADE THAT REQUIRED DRAINAGE. THE PATIENT DIED. CARDIAC TAMPONADE OCCURRED DURING THE PROCEDURE. THE OCTARAY WAS THE ONLY DEVICE USED DURING THE CASE; HOWEVER, THE PHYSICIAN THOUGHT THAT IT WAS DUE TO THE TRANSSEPTAL. "QUICK MINUTES" AFTER THE MAPPING STARTED, THE BLOOD PRESSURE HAD FALLEN AND THEY DID AN ECHOCARDIOGRAPHY AND THEY SAW A MAJOR TAMPONADE. HE DIRECTLY STOPPED THE MAPPING AND TREATED THE TAMPONADE. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT WAS THE PUNCTURE OF THE TRANSSEPTAL AND THAT THERE WAS NO IMPLICATION WITH BIOSENSE WEBSTER (BWI) PRODUCTS. THE PUNCTURE WAS DONE WITH THE ETO AND BRK 71CM. HOWEVER, THE EVENT WAS DISCOVERED AT THE START OF MAPPING. NO ABLATION CATHETER WAS USED. THERE WAS NO ERROR MESSAGE OBSERVED ON BWI EQUIPMENT DURING THE PROCEDURE. INTERVENTION WAS PROVIDED WHICH INCLUDED THE PHYSICIAN DRAINED THE PATIENT; HOWEVER, THE PATIENT PASSED AWAY. Manufacturer Narrative: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER 31642001L AND NO INTERNAL ACTION RELATED TO THE COMPLAINT WAS FOUND DURING THE REVIEW. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22613624&pc=NPT | 2025587-2025-05289 | 2025-01-07 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-25 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 7 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22613835&pc=MNT | 2518422-2025-048357 | 2025-07-21 04:00 | Death | RESPIRONICS, INC. | 2025-07-25 04:00 | MNT | RESPIRONICS | No Audible Alarm | Insufficient Information | K102985 |  | 1 | Event Description: PHILIPS RECEIVED A COMPLAINT ON THE V60 VENTILATOR INDICATING THAT A PATIENT PASSED WHILE ON UNIT. IT WAS REPORTED THAT THE DEVICE WAS DISCONNECTED FROM THE PATIENT AND THE PATIENT PASSED 5:20 AM. THE CUSTOMER STATED THAT IT WAS REPORTED THAT THE TUBE SET WAS REMOVED FROM MASK, BUT THE DEVICE DID NOT ALARM. THE CUSTOMER IS REQUESTING AN ON-SITE SERVICE TO HAVE THE DEVICE EVALUATED. BASED UPON THE INFORMATION PROVIDED AND CURRENTLY AVAILABLE, THE DEVICE CAUSE AND/OR CONTRIBUTION TO THE REPORTED EVENT OF DEATH WHILE ON THE DEVICE CANNOT CURRENTLY BE CONFIRMED, NOR REFUTED, BASED ON THE EVIDENCE PRESENT. FURTHER GOOD FAITH EFFORTS AND INFORMATION GATHERING WILL BE PERFORMED TO DISPOSITION DEVICE CAUSE AND/OR CONTRIBUTION TO THE PATIENT DEATH. Manufacturer Narrative: THE CUSTOMER ALLEGED THAT THE TUBE WAS REMOVED FROM THE PATIENT MASK AND THE V60 DID NOT ALARM AS EXPECTED. THE CUSTOMER REQUESTED AN ONSITE INSPECTION OF THE DEVICE. ON 04AUG2025, A GOOD FAITH EFFORT (GFE) RESPONSE WAS RECEIVED FROM THE MARKET CARDIOPULMONARY MANAGER AT THE CUSTOMER SITE. THE CUSTOMER STATED THAT THE PATIENT WAS ALERT AND ORIENTED WITHIN 30-45 MINUTES PRIOR TO THE "CODE BLUE" BEING CALLED. THE PATIENT WAS NOT ON CENTRAL MONITORING, BUT THEY WERE ON A BEDSIDE OXIMETER WITH ALERTS AND ALARMS SET. THE PATIENT PRE-EXISTING CONDITIONS INCLUDED ENCEPHALOPATHY NOS (NOT OTHERWISE SPECIFIED) AND ACUTE AND CHRONIC RESPIRATORY FAILURE WITH HYPOXIA AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD). THE HOSPITAL STAFF REPORTED THAT THE DEVICE HAD BEEN FUNCTIONING PROPERLY, AS THE DISCONNECT ALARM HAD ALARMED MULTIPLE TIMES THROUGHOUT THE NIGHT. A FIELD SERVICE ENGINEER (FSE) WENT TO THE CUSTOMER SITE ON (B)(6) 2025 TO EVALUATE THE DEVICE. THE FSE BOOTED THE VENTILATOR IN NORMAL MODE, AND DURING THE BOOT UP THE FSE DID HEAR AUDIBLE ALARM TONES FROM THE DEVICE. THE FSE RECORDED THE BOOTH SEQUENCE AND NOTED THAT NO ERRORS WERE PRODUCED OR OBSERVED. THE DEVICE DIAGNOSTIC REPORT (DRPT) WAS DOWNLOADED TO BE REVIEWED FOR ERRORS, DIAGNOSTIC CODES, CHECK VENT, AND VENT-INOP ALARMS. THE FSE STATED THAT THEY DID NOT FIND ANY CODES OR ERRORS INDICATING A FAILURE OF THE DEVICE. THE FSE NOTED THAT THE DEVICE BEGAN PROVIDING SERVICE AT ABOUT 10:34PM ON (B)(6) 2025 AND RAN CONTINUOUSLY UNTIL THE MORNING OF (B)(6) 2025 WITH MULTIPLE PATIENT DISCONNECT ALARMS BEING LOGGED THROUGHOUT THE EVENING AND EARLY MORNING. AT APPROXIMATELY 5:58AM, THE EVENT LOG SHOWED THAT THE VENTILATOR VOLUME WAS REDUCED FROM A LEVEL OF 5 TO A LEVEL OF 1, WITH 1 BEING THE LOWEST VOLUME SETTING POSSIBLE. THE FSE VERIFIED THAT THE ALARM VOLUME WAS STILL AT LEVEL 1 DURING THEIR EVALUATION OF THE DEVICE ON 04AUG2025. THE FSE CONTINUED TO TEST THE UNIT, WITH THE DEVICE PASSING A COMPLETE PERFORMANCE VERIFICATION TEST (PVT). THE FSE STATED THAT THEY FOCUSED ON THE ALARM TESTS, AND THE UNIT PASSED EVERY TIME. ALL SETTINGS AND OTHER INVESTIGATION PROCEDURES WERE STATED TO HAVE BEEN DOCUMENTED BY THE FSE. THE SYSTEM MET SPECIFICATIONS AND WAS RETURNED TO THE CUSTOMER READY FOR USE. NO ISSUE WAS OBSERVED OR COMMUNICATED BY THE FSE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22614075&pc=DSQ | 2916596-2025-04929 | 2025-07-11 04:00 | Death | THORATEC CORPORATION | 2025-07-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AFTER MULTIPLE ATTEMPTS TO WEAN THEM OFF OF VASOPRESSORS AND INOTROPES. ADDITIONAL INFORMATION WAS NOT PROVIDED.Event Description: AS OF (B)(6) 2025, THE SITE STATE THAT THE CAUSE OF DEATH WAS CARDIOGENIC/VASODILATORY SHOCK. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND WAS CONFIRMED TO HAVE OPERATED AS EXPECTED. THE DEVICE WOULD NOT BE RETURNED. Manufacturer Narrative: CORRECTED DATA: SECTION B3: DATE OF DEATH CORRECTED. SECTION H6: HEALTH EFFECT - IMPACT CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE ACCOUNT INDICATED THAT HEARTMATE 3 LVAS WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION", LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: B5 - NARRATIVE INFORMATION UPDATED. H6 - ADVERSE EVENT PROBLEM CODES UPDATED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22615264&pc=NVN | 2017865-2025-95844 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-25 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR, RIGHT VENTRICULAR LEAD AND RIGHT ATRIAL LEAD, HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD WAS RETURNED IN TWO PIECES. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES, HOWEVER AN INTERNAL SHORT WAS NOTED BETWEEN CONDUCTOR PATHS AT THE CONNECTOR PORTION OF THE LEAD CONSISTENT WITH PROCEDURAL DAMAGE. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22615283&pc=NVY | 2017865-2025-95845 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-25 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR, RIGHT VENTRICULAR LEAD AND RIGHT ATRIAL LEAD, HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD WAS RETURNED IN THREE PIECES. ELECTRICAL TESTING WAS NORMAL. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22615297&pc=NGV | 2135147-2025-04128 | 2025-06-26 04:00 | Death | ABBOTT MEDICAL | 2025-07-25 04:00 | NGV | AMPLATZER AMULET | Migration or Expulsion of Device | Arrhythmia; Obstruction/Occlusion | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 28MM AMPLATZER AMULET WAS SUCCESSFULLY IMPLANTED IN THE LAA. THE PATIENT HAD CHALLENGING ANATOMY AND THE DEVICE NEEDED TO BE REPOSITIONED SEVERAL TIMES. A GENTLE TUG TEST WAS NOTED TO HAVE BEEN PERFORMED. LATER, IT WAS NOTED THAT THE DEVICE HAD EMBOLIZED INTO THE LEFT VENTRICLE RESULTING IN A PARTIAL BLOCKAGE OF BLOOD FLOW. THE DEVICE WAS RETRIEVED VIA SNARE AFTER SIX HOURS TO RESOLVE THE EVENT. THE PHYSICIAN ALLEGES THAT THE POSITION AND ORIENTATION OF THE DEVICE WAS THE CAUSE OF THE EMBOLIZATION. THE PATIENT WAS TAKEN TO THE ICU TO BE MONITORED. THE PATIENT HAVE HAD ARRHYTHMIA AND WAS PARTIALLY HEMODYNAMIC UNSTABLE. ON THE SAME DAY THE PATIENT PASSED AWAY IN THE ICU DUE TO MULTIPLE ORGAN FAILURE. SINCE THE PATIENT HAD A COMPLICATION DURING THE PROCEDURE, THIS IS ASSUMED TO BE THE CAUSE OF DEATH. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22615574&pc=NVN | 2649622-2025-20433 | 2025-06-27 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-25 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER THE PATIENT UNDERWENT A CARDIAC PROCEDURE, THE PATIENT EXPERIENCED A PINS AND NEEDLES SENSATION AND DISCOMFORT BEHIND THE SITE OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD). APPROXIMATELY FIVE DAYS LATER, THE PATIENT UNDERWENT A AN UPGRADE PROCEDURE TO A CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM WHICH INCLUDED THE IMPLANT OF AN LEFT VENTRICULAR (LV) LEAD AND EXPLANT OF THE ICD. ON THE DAY OF THE PROCEDURE, THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: CONTINUATION OF D10: DTPA2D1 CRTD IMPLANTED: (B)(6) 2025 694758 LEAD IMPLANTED: (B)(6) 2012. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22616235&pc=NVZ | 2017865-2025-95756 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-25 04:00 | NVZ | ELLIPSE DR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR, RIGHT VENTRICULAR LEAD AND RIGHT ATRIAL LEAD, HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR WAS RETURNED DUE TO PATIENT DEATH. THE DEVICE WAS FOUND TO COMMUNICATE APPROPRIATELY WITH A MERLIN PROGRAMMER, AND THE DEVICE IMAGE WAS ABLE TO BE DOWNLOADED SUCCESSFULLY. THE VISUAL SCREEN WAS ACCEPTABLE AND ELECTRICAL TESTING WAS NORMAL. NO OTHER ANOMALIES WERE FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22618691&pc=PNJ | 2017865-2025-95971 | 2025-07-07 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-25 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PRESENTED FOR LEADLESS PACEMAKER IMPLANT PROCEDURE. DURING PROCEDURE, IT WAS OBSERVED THAT THE RIGHT VENTRICLE (RV) WAS PERFORATED. THE PERFUSION TEAM WAS URGENTLY CALLED TO MANAGE THE RESULTING PERICARDIAL EFFUSION. THE PATIENT UNDERWENT EMERGENCY THORACOTOMY, BUT UNFORTUNATELY, DID NOT SURVIVE THE EVENT AND PASSED AWAY. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22618736&pc=PNJ | 2017865-2025-95972 | 2025-07-07 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-25 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT ATRIUM | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Cardiac Perforation; Pericardial Effusion; No Clinical Signs, Symptoms or Conditions | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PRESENTED FOR LEADLESS PACEMAKER IMPLANT PROCEDURE. DURING PROCEDURE, IT WAS OBSERVED THAT THE RIGHT VENTRICLE (RV) WAS PERFORATED. THE PERFUSION TEAM WAS URGENTLY CALLED TO MANAGE THE RESULTING PERICARDIAL EFFUSION. THE PATIENT UNDERWENT EMERGENCY THORACOTOMY, BUT UNFORTUNATELY, DID NOT SURVIVE THE EVENT AND PASSED AWAY. Manufacturer Narrative: CORRECTION: PLEASE RETRACT THIS REPORT 2017865-2025-95972. UPON REVIEW, THE RIGHT ATRIAL LEADLESS PACEMAKER (RALP) SHOULD NOT HAVE BEEN SUBMITTED AS A MEDICAL DEVICE REPORT (MDR) AS RIGHT VENTRICLE WAS PERFORATED AND NO MALFUNCTION ALLEGED ON RALP DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22618757&pc=PNJ | 2017865-2025-95973 | 2025-07-07 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-25 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PRESENTED FOR LEADLESS PACEMAKER IMPLANT PROCEDURE. DURING PROCEDURE, IT WAS OBSERVED THAT THE RIGHT VENTRICLE (RV) WAS PERFORATED. THE PERFUSION TEAM WAS URGENTLY CALLED TO MANAGE THE RESULTING PERICARDIAL EFFUSION. THE PATIENT UNDERWENT EMERGENCY THORACOTOMY, BUT UNFORTUNATELY, DID NOT SURVIVE THE EVENT AND PASSED AWAY. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22620835&pc=MIH | 9612164-2025-03640 | 2024-07-05 04:00 | Death | MEDTRONIC IRELAND | 2025-07-25 04:00 | MIH | VALIANT STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100040 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ¿ DEVELOPMENT AND VALIDATION OF A USER FRIENDLY MORPHOLOGY GRADING SYSTEM (PATENT) PREDICTING AORTIC REMODELLING AFTER THORACIC ENDOVASCULAR AORTIC REPAIR IN HIGH RISK UNCOMPLICATED TYPE B AORTIC DISSECTION¿ THE TIME FRAME OF THIS STUDY WAS OVER AN ELEVEN YEAR PERIOD. VALIANT STENT GRAFTS WERE IMPLANTED IN THE ENDOVASCULAR TREATMENT OF TYPE B AORTIC DISSECTIONS. 351PATIENTS WERE INCLUDED. PATIENT MORTALITY WAS REPORTED OF WHICH 19 PATIENT DEATHS WERE DEEMED AORTIC RELATED. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS Manufacturer Narrative: MEDTRONIC RECEIVED THE FOLLOWING INFORMATION FROM A JOURNAL ARTICLE ENTITLED; ¿DEVELOPMENT AND VALIDATION OF A USER FRIENDLY MORPHOLOGY GRADING SYSTEM (PATENT) PREDICTING AORTIC REMODELLING AFTER THORACIC ENDOVASCULAR AORTIC REPAIR IN HIGH RISK UNCOMPLICATED TYPE B AORTIC DISSECTION¿ SHEN Y, WANG J, ZHAO J, HUANG B, WENG C, WANG T EUR J VASC ENDOVASC SURG (2024) 68, 579E587 HTTPS://DOI.ORG/10.1016/J.EJVS.2024.07.004 A.2 MEDIAN A.3 AVERAGE MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22621033&pc=NPT | 2135147-2025-04111 | 2022-05-01 04:00 | Death | ABBOTT MEDICAL | 2025-07-25 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P190023 |  | 1 | Event Description: THE ARTICLE, "SAPIEN3 ULTRA RESILIA VERSUS NAVITOR VALVE SYSTEM IN PATIENTS WITH A SMALL AORTIC ANNULUS: INSIGHTS FROM THE OCEAN TAVI REGISTRY", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, MULTICENTER STUDY ON THE SHORT-TERM CLINICAL OUTCOMES AND TRANSCATHETER HEART VALVES (THV) HEMODYNAMIC PERFORMANCE OF SAPIEN3 ULTRA RESILIA (S3UR) AND NAVITOR IN PATIENTS WITH SMALL ANNULUS, DEFINED AS AN ANNULUS AREA OF = 430 MM2. DEVICES INCLUDED IN THE STUDY WERE SAPIEN 3 ULTRA RESILIA AND NAVITOR. THE ARTICLE CONCLUDED THAT TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) USING S3UR RESULTED IN LOWER RATES OF ISCHEMIC STROKE, PACEMAKER IMPLANTATION, AND PVL COMPARED WITH THAT USING NAVITOR. THE INCIDENCE OF MODERATE OR SEVERE PPM WAS SIMILAR IN BOTH GROUPS; HOWEVER, THE NAVITOR TAVR GROUP SHOWED SUPERIOR HEMODYNAMIC VALVE PERFORMANCE, COMPARED WITH S3UR TAVR GROUP. [THE PRIMARY AND CORRESPONDING AUTHOR WAS MASANORI YAMAMOTO, DEPARTMENT OF CARDIOLOGY, TOYOHASHI HEART CENTER 21-1 GOBUTGOI, OYAMACHO, TOYOHASHI, AICHI, 441=8530, JAPAN WITH CORRESPONDING E-MAIL: MASA-NORI@NMS.AC.JP] THE TIME FRAME OF THE STUDY WAS FROM MAY 2022 TO MARCH 2024. A TOTAL OF 1224 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 518 (42.3%) RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 85 YEARS AND THE MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED HEART BLOCK, BICUSPID AORTIC VALVES, PRIOR MYOCARDIAL INFARCTION, AORTIC STENOSIS, AND SMALL AORTIC ANNULUS. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF SAPIEN3 ULTRA RESILIA VERSUS NAVITOR VALVE SYSTEM IN PATIENTS WITH A SMALL AORTIC ANNULUS: INSIGHTS FROM THE OCEAN TAVI REGISTRY WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HEART BLOCK, BICUSPID AORTIC VALVES, PRIOR MYOCARDIAL INFARCTION, AORTIC STENOSIS, AND SMALL AORTIC ANNULUS. SOME OF THE COMPLICATIONS REPORTED WERE RENAL FAILURE, STROKE, ATRIAL FIBRILLATION, HEMORRHAGE AND DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. FROM MEDICAL REVIEW: AVR USING S3UR RESULTED IN LOWER RATES OF ISCHEMIC STROKE, PACEMAKER IMPLANTATION, AND PVL COMPARED WITH THAT USING NAVITOR. THE INCIDENCE OF MODERATE OR SEVERE PPM WAS SIMILAR IN BOTH GROUPS; HOWEVER, THE NAVITOR TAVR GROUP SHOWED SUPERIOR HEMODYNAMIC VALVE PERFORMANCE, COMPARED WITH S3UR TAVR GROUP. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: SAPIEN3 ULTRA RESILIA VERSUS NAVITOR VALVE SYSTEM IN PATIENTS WITH A SMALL AORTIC ANNULUS: INSIGHTS FROM THE OCEAN-TAVI REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22621075&pc=NIQ | 2024168-2025-07754 | 2025-06-30 04:00 | Death | ABBOTT VASCULAR INC. | 2025-07-25 04:00 | NIQ | XIENCE ALPINE | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P110019 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROCEDURE WAS PERFORMED TO TREAT A MODERATELY CALCIFIED AND TORTUOUS LESION IN THE LEFT ANTERIOR DESCENDING (LAD). DURING THE PROCEDURE A 3X18MM AND 2.75X38MM XIENCE ALPINE DRUG ELUTING STENTS (DES) WERE SUCCESSFULLY IMPLANTED AT THE TARGET LESION WITH NO ADVERSE PATIENT EFFECTS NOR CLINICAL DELAY BEING NOTED. HOWEVER, 24 HOURS AFTER THE PROCEDURE THE PATIENT WAS NOTED TO EXPIRE. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE PATIENT PRESENTED WITH UNSTABLE ANGINA PRIOR TO THE PROCEDURE. FOLLOWING STENT IMPLANTATION, ANGIOGRAPHY WAS PERFORMED, AND A THROMBUS WAS OBSERVED. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE REPORTED PATIENT EFFECTS OF DEATH AND THROMBOSIS ARE LISTED IN THE XIENCE ALPINE EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECTS OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22584892&pc=KFM | 8010762-2025-0000314 | 2025-07-22 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-07-24 04:00 | KFM | HEART LUNG MACHINE | Electrical /Electronic Property Problem; Pumping Stopped | Cardiac Arrest | K991864 |  | 1 | Event Description: THE EVENT OCCURRED IN JAPAN. IT WAS REPORTED THAT THE ERROR MESSAGE ¿REFERENC ERROR¿ OCCURRED ON THE ROTAFLOW CONSOLE AND THE PUMP STOPPED DUE TO THE "REFERENC ERROR" DURING PATIENT USE. THE CUSTOMER USED THE EMERGENCY DRIVE AND REPLACED THE AFFECTED ROTAFLOW CONSOLE WITH ANOTHER ROTAFLOW CONSOLE. THE PUMP HAD BEEN USED ON A PATIENT WHO WAS IN A BAD CONDITION, AND THE PATIENT LATER PASSED AWAY, BUT THE PUMP HAD ONLY BEEN STOPPED FOR A SHORT TIME AND HIS BLOOD PRESSURE WAS STABLE, SO THERE IS NO CAUSAL RELATIONSHIP TO THIS ERROR ACCORDING TO THE CUSTOMER. DUE TO THE REPORTED DEATH OF THE PATIENT A REPORT IS REQUIRED. COMPLAINT ID: (B)(4).Event Description: COMPLAINT ID: (B)(4). Manufacturer Narrative: THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE JAPANESE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, ROTAFLOW CONSOLE¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K991864. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR ROTAFLOW CONSOLE WITH CATALOG NUMBER 701051711. Manufacturer Narrative: THE EVENT OCCURRED IN JAPAN. IT WAS REPORTED THAT THE ERROR MESSAGE ¿REFERENCE ERROR¿ OCCURRED ON THE ROTAFLOW CONSOLE AND THE PUMP STOPPED DUE TO THE "REFERENCE ERROR" DURING PATIENT USE. THE CUSTOMER USED THE EMERGENCY DRIVE AND REPLACED THE AFFECTED ROTAFLOW CONSOLE WITH ANOTHER ROTAFLOW CONSOLE. THE PUMP HAD BEEN USED ON A PATIENT WHO WAS IN A BAD CONDITION, AND THE PATIENT LATER PASSED AWAY, BUT THE PUMP HAD ONLY BEEN STOPPED FOR A SHORT TIME AND HIS BLOOD PRESSURE WAS STABLE, SO THERE IS NO CAUSAL RELATIONSHIP TO THIS ERROR ACCORDING TO THE CUSTOMER. FURTHER INFORMATION HAS BEEN PROVIDED BY THE SSU (SALES AND SERVICE UNIT) DATED ON 2025-07-22 AS FOLLOWS: A DIALYSIS PATIENT WAS URGENTLY HOSPITALIZED FOR CHEST PAIN ON (B)(6) 2025. CAG WAS PERFORMED ON (B)(6) 2025, BUT NO LESION WAS FOUND, AND THE PATIENT WAS MOVED DIRECTLY TO THE ICU. CHEST PAIN DEVELOPED AGAIN LATE AT NIGHT ON (B)(6) 2025, THE PATIENT BECAME UNCONSCIOUS, AND SUFFERED A CARDIAC ARREST, SO ECMO WAS STARTED AT 5:40 AM ON (B)(6) 2025 (USING ROTAFLOW). THERE WAS NO CHANGE IN THE PATIENT CONDITION BEFORE OR AFTER DEVICE REPLACEMENT. THE EXACT DATE OF THE DEATH WAS ON (B)(6) 2025. DUE TO THE REPORTED PUMP STOP DURING TREATMENT A REPORT IS REQUIRED. THE AFFECTED ROTAFLOW CONSOLE WITH S/N: (B)(6) WAS INVESTIGATED BY A GETINGE FIELD SERVICE TECHNICIAN AND THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO PARTS HAS BEEN REPLACED. THE DEVICE WAS TESTED AND PASSED ALL FUNCTIONAL AND SAFETY TESTS ACCORDING TO THE SPECIFICATIONS. A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-08-07 WITH FOLLOWING CONCLUSION: "THE REPORTED EVENT INVOLVED A ¿REFERENCE¿ ERROR ON A ROTAFLOW PUMP, WHICH RESULTED IN THE UNEXPECTED STOPPAGE OF THE PUMP DURING ONGOING V-A ECMO THERAPY IN A CRITICALLY ILL 69-YEAR-OLD PATIENT WITH DIABETES AND CARDIAC ARREST. THE CLINICAL TEAM RESPONDED PROMPTLY BY INITIATING CHEST COMPRESSIONS AND ACTIVATING THE EMERGENCY DRIVE WITHIN APPROXIMATELY THREE MINUTES, FOLLOWED BY REPLACEMENT OF THE DEVICE. THE TOTAL TIME TO RESUME FULL CIRCULATORY SUPPORT WAS REPORTED TO BE AROUND TEN MINUTES. DURING THIS PERIOD, THE PATIENT¿S BLOOD PRESSURE REMAINED STABLE, AND NO DETERIORATION IN THE CLINICAL CONDITION WAS OBSERVED IMMEDIATELY BEFORE OR AFTER THE INCIDENT. THE DISPLAYED ERROR, AS REPORTED BY THE USER, INCLUDED ¿ERR OR!!¿ IN THE SPEED AND FLOW DISPLAY FIELDS, ALONG WITH THE MESSAGE ¿[REFERENCE]¿ IN THE STATUS AREA. ACCORDING TO THE INSTRUCTIONS FOR USE, THIS INDICATES THAT THE INTERNAL VOLTAGE REFERENCE WAS OUTSIDE THE VALID RANGE, WHICH RESULTS IN BOTH AN ACOUSTIC ALARM AND IMMEDIATE PUMP STOP. POST-EVENT ANALYSIS INVOLVED A FULL RUNNING TEST AND INTERNAL DIAGNOSTICS, BUT THE ISSUE COULD NOT BE REPRODUCED, AND NO HARDWARE ABNORMALITIES WERE FOUND. AS SUCH, A DEFINITIVE ROOT CAUSE COULD NOT BE IDENTIFIED. REASONABLE AND POSSIBLE CAUSES MAY INCLUDE A TRANSIENT ELECTRONIC OR VOLTAGE INSTABILITY, INTERNAL SIGNAL ERROR, OR MECHANICAL OR ELECTRICAL INFLUENCE SUCH AS A BRIEF SHOCK OR POWER FLUCTUATION. HOWEVER, AS THE REPORTED INCIDENT COULD NOT BE REPRODUCED BY THE FIELD SERVICE TECHNICIAN OF GETINGE, WHICH MAY INDICATE THAT THE ROOT CAUSE MAY BE FOUND WITHIN A BRIEF POWER FLUCTUATION. THE CLINICAL TEAM NOTED A ¿CRACK¿ SOUND WAS HEARD JUST PRIOR TO THE EVENT, AND THE DEVICE WAS REPORTEDLY POWERED VIA A TABLE TAP, WHICH MAY HAVE CONTRIBUTED TO A BRIEF INSTABILITY - THOUGH THIS REMAINS SPECULATIVE. NO USER ERROR HAS BEEN IDENTIFIED IN RELATION TO SETUP OR OPERATION. THE CLINICAL TEAM FOLLOWED APPROPRIATE EMERGENCY PROCEDURES, AND NO DEVIATIONS FROM STANDARD USE WERE REPORTED. WHILE THE EVENT DOES CONSTITUTE A TEMPORARY DEVICE MALFUNCTION, THERE IS NO EVIDENCE OF PERSISTENT PERFORMANCE DEGRADATION, AND THE ISSUE WAS RESOLVED WITH REPLACEMENT OF THE PUMP. AS STATED BY THE COMPLAINT NARRATIVE: ¿THE PUMP HAD BEEN USED ON A PATIENT WHO WAS IN A BAD CONDITION, AND THE PATIENT LATER DIED, BUT THE PUMP HAD ONLY BEEN STOPPED FOR A SHORT TIME AND HIS BLOOD PRESSURE WAS STABLE, SO THERE IS NO CAUSAL RELATIONSHIP TO THIS ERROR.¿ THEREFORE, THE CLINICAL OUTCOME IS BELIEVED TO BE PRIMARILY RELATED TO THE PATIENT¿S MEDICAL STATE RATHER THAN THE DEVICE INCIDENT. IN CONCLUSION, THE DEVICE BEHAVED IN ACCORDANCE WITH ITS DESIGN RESPONSE TO A CRITICAL INTERNAL ERROR AS DESCRIBED IN THE IFU. WHILE THE ROOT CAUSE OF THE VOLTAGE REFERENCE FAULT COULD NOT BE CONCLUSIVELY DETERMINED, ALL AVAILABLE DATA INDICATE THAT THE INCIDENT WAS ISOLATED, AND NO DIRECT LINK TO PATIENT HARM HAS BEEN ESTABLISHED." THE REVIEW OF THE NON-CONFORMITIES WAS PERFORMED ON 2025-07-22 AND DURING THE PERIOD OF 2012-10-31 TO 2025-07-22 DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED PRODUCT AND FAILURE. THERE IS NO INDICATION THAT MANUFACTURING ISSUES OCCURRED DURING THIS TIME, THUS PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE AFFECTED ROTAFLOW DEVICE WAS PRODUCED IN 2012-10-31. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22585840&pc=PYX | 1220648-2025-30656 | 2025-07-07 04:00 | Death | ABIOMED, INC. | 2025-07-24 04:00 | PYX | AUTOMATED IMPELLA CONTROLLER | Failure to Read Input Signal | Insufficient Information | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT WHILE SETTING UP FOR A NEW CASE, AN AUTOMATED IMPELLA CONTROLLER (AIC) FAILED TO RECOGNIZE AN INSTALLED PURGE DISC. A CASE START WAS ATTEMPTED MULTIPLE TIMES, BUT THE ISSUE PERSISTED AND THE DECISION WAS MADE TO REPLACE THE CONSOLE. DESPITE A SUCCESSFUL SUBSEQUENT IMPLANT / START-UP OF SUPPORT, THE PATIENT CONTINUED TO DETERIORATE FROM THEIR UNDERLYING CONDITION AND PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION HAS BEEN COMPLETED. IMC LOGS REVEALED THAT THE CONSOLE ISSUED PURGE DISC NOT DETECTED ALARM (EVENT SET #402) SEVERAL TIMES. THE CONSOLE WAS RETURNED AND EXAMINED, A BROKEN PURGE DISC FLAG WAS IDENTIFIED. THE PURGE DISC FLAG WAS OBSERVED TO BE BROKEN AND WAS THE ROOT CAUSE OF THE INABILITY TO DETECT THE PURGE DISC. THE POSSIBLE ROOT CAUSES OF A FRACTURE OF THE PURGE FLAG WHICH IS NOT TIED TO A MANUFACTURING OR DESIGN DEFECT, TYPICALLY CAUSED BY FLUID INGRESS INTO THE PURGE PRESSURE SENSOR ASSEMBLY. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22585885&pc=OZD | 1220648-2025-30657 | 2025-07-08 04:00 | Death | ABIOMED, INC. | 2025-07-24 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE PATIENT SUFFERED A GASTROINTESTINAL BLEED DURING IMPELLA 5.5 SUPPORT. HEPARIN WAS STOPPED AND BLOOD PRODUCTS WERE GIVEN TO MANAGE THE CONDITION. SUPPORT CONTINUED UNINTERRUPTED. TWO DAYS LATER, A CT SCAN REVEALED THAT THE PATIENT HAD MULTIPLE BRAIN BLEEDS. THE PATIENT EXPIRED ON SUPPORT TWO DAYS LATER. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: POTENTIAL ADVERSE EVENTS (UNITES STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22585895&pc=NGV | 2124215-2025-48885 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR SURGICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 387 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22586021&pc=DSP | 2248146-2025-0000420 | 2025-07-09 04:00 | Death | DATASCOPE CORP. | 2025-07-24 04:00 | DSP | MEGA 7.5FR. 40CC IAB | Migration or Expulsion of Device | Myocardial Infarction | K120868 |  | 1 | Event Description: IT WAS REPORTED DURING USE, THROUGH THE EMERGENCY SUPPORT LINE, THAT THE PATIENT HAD LOW BLOOD PRESSURE WITH HIGH AUGMENTATION. AN X-RAY IMAGE WAS REVIEWED WHICH SHOWED THE IAB MARKER AT THE 8TH INTERCOSTAL SPACE, INDICATING THE BALLOON WAS POSITIONED LOW. THE USER WAS ADVISED TO NOTIFY THE PROVIDER. APPROXIMATELY ONE HOUR LATER, THE USER CALLED BACK AND SHARED ANOTHER X-RAY SCREENSHOT, WHICH SHOWED THAT THE BALLOON POSITION REMAINED RELATIVELY UNCHANGED AND WAS NOTED TO BE ALMOST AT THE HUB. IT WAS RECOMMENDED TO RETURN THE PATIENT TO THE CATH LAB FOR FURTHER EVALUATION UNDER FLUOROSCOPY, AS THE RADIOPAQUE TIP SHOULD BE POSITIONED BETWEEN THE 2ND AND 3RD INTERCOSTAL SPACE. THE PATIENT EXPIRED. Manufacturer Narrative: A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22586147&pc=OZD | 1220648-2025-30662 | 2025-07-09 04:00 | Death | ABIOMED, INC. | 2025-07-24 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Perforation; Pericardial Effusion | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD AN IMPELLA RP FLEX AND INTRA-AORTIC BALLOON PUMP (IABP) PATIENTS CONDITION DETERIORATING SO THE TEAM MADE THE DECISION TO REPLACE THE IABP TO AN IMPELLA CP. THE PATIENT WAS BROUGHT TO THE CATHRTERIZATION LAB FOR THE EXCHANGE. UPON ADVANCING THE IMPELLA INTO THE ASCENDING AORTA AND PRIOR TO CROSSING THE AORTIC VALVE, THE PATIENT BECAME UNSTABLE, MEAN ARTIERIAL PRESSURE DROPPED INTO THE 40'S AND THE PATIENT BECAME BRADYCARDIC. MULTIPLE ROUNDS OF EPINEPHRINE AND BICARBONATE WERE GIVEN. DURING THE LOSS IN BLOOD PRESSURE, THE MEDICAL DOCTOR (MD) ATTEMPTED TO ADVANCE THE IMPELLA INTO THE LEFT VENTRICLE (LV) TO INITIATE SUPPORT. THE PATIENT¿S DESCENDING AND ASCENDING AORTA WERE NOTED TO BE VERY TORTUROUS AND THE MD WAS UNABLE TO ADVANCED THE IMPELLA INTO THE LV FAR ENOUGH TO INITIATE SUPPORT DUE TO THE IMPELLA CATHETER BUCKLING IN THE DESCENDING AORTA. THE PATIENT CONTINUED TO BE UNSTABLE. THE DECISION WAS MADE TO USE A SINGLE ACCESS TO PLACE THE 7 FRENCH SHEATH AND TO USE A JR4 CATHETER WITH A STIFF WIRE TO HELP STRAIGHTEN THE TORTUOUS ANATOMY. WHEN THE JR4 AND WIRE WERE IN THE ASCENDING AORTA, THE MD WAS ABLE TO EASILY ADVANCE THE IMPELLA INTO THE LV WITHOUT RESISTANCE OR DIFFICULTY. THE JR4 AND WIRES WERE REMOVED AND SUPPORT WAS INITIATED ON AUTO WITH FLOWS 0.2-0.4 LITERS PER MINUTE WITH MAP IN THE 40S. MULTIPLE ROUNDS OF EPINEPHRINE AND BICARBONATE WERE CONTINUED TO BE GIVEN. CONTINUOUS SUCTION ALARMS OCCURRING ON AUTO AND THE PATIENT WAS TAKEN OUT OF AUTO MODE TO P8. THERE CONTINUED TO BE SUCTION ALARMS SO THE P-LEVEL WAS THEN REDUCED TO P6. SUCTION WAS INTERMITTENTLY RESOLVED UNTIL THE PATIENT'S HEART RATE AND MEAN ARTERIAL PRESSURE DROPPED AGAIN. THEY WER UNABLE TO STABILIZE THE PATIENT WITH CONTINUOUS SUCTION ALARMS ON THE IMPELLA CP. THE MD REQUESTED A STAT ECHOCARDIOGRAM TO ASSESS THE IMPELLA AND FOR POSSIBLE EFFUSION. UPON ARRIVAL, THE ECHOCARDIOGRAM REVEALED THE PIGTAIL OF THE IMPELLA CP THROUGH THE LEFT VENTRICLE WALL AND A LARGE PERICARDIAL EFFUSION WAS PRESENT. THE DECISION WAS MADE TO PLACE A PERICARDIAL DRAIN. INITIALLY, THE MD DRAINED 200 CUBIC CENTIMETERS OF BLOOD AND THEN HOOKED TO A VACUTAINER. THE VACUTAINER CONTINUED TO FILL UP AND THE DECISION WAS MADE TO HOOK A PERICARDIAL DRAIN VIA A 3-WAY STOPCOCK DIRECTLY TO THE PATIENT¿S CENTRAL LINE TO BE ABLE TO GIVE BACK BLOOD THE MD WAS PULLING OUT OF DRAIN. EMERGENT BLOOD WAS ORDERED AND ALSO GIVEN. THE EFFUSION WOULD NOT STOP DRAINING AND AFTER PULLING AND GIVING BACK OVER A LITER OF BLOOD, IT WAS DETERMINED BY THE TEAM THAT IT WAS NOT GOING TO RESOLVE. THE TEAM SPOKE WITH FAMILY AND THE DECISION WAS MADE TO NOT ESCALATE CARE FROM THAT POINT ON. THE PATIENT WAS MADE DO NOT ATTEMPT RESUSCITATION AND WAS TAKEN BACK TO INTENSIVE CARE UNIT (ICU) WITH THE CATHETERIZATION LAB TEAM AND ICU NURSES. UPON ARRIVAL TO UNIT, THE PATIENT EXPIRED ON IMPELLA RP FLEX AND IMPELLA CP SUPPORT SHORTLY AFTER. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. IMPELLA CP WITH SMARTASSIST SYSTEM WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22586544&pc=DSQ | 2916596-2025-04772 | 2025-07-11 04:00 | Death | THORATEC CORPORATION | 2025-07-24 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS DUE TO END OF THERAPY. PUMP WAS WORKING AS EXPECTED. NO PUMP RELATED ISSUES WERE REPORTED. THE PATIENT WAS AT HOME IN UNSTABLE CONDITIONS, BUT THE PATIENT DENIED RE-ADMISSION TO THE HOSPITAL. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS IFU, REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22586673&pc=MKJ | 1220908-2025-02772 | 2025-07-16 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-07-24 04:00 | MKJ | R SERIES | Failure to Deliver Energy; Unable to Obtain Readings; Failure of Device to Self-Test | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE DISPLAYED A "ECG DISABLED" MESSAGE. THE DEVICE ALSO DISPLAYED A RED "X" INDICATOR AND FAILED CODE READINESS TEST. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL CANADA FOR EVALUATION. THE CUSTOMER'S REPORT WAS OBSERVED DURING REVIEW OF THE DEVICE DATA LOGS. HOWEVER, THE DEVICE WAS PUT THROUGH EXTENSIVE FUNCTIONAL TESTING WITHOUT DUPLICATING THE REPORT. THE ACTIVITY LOG INDICATED THE DEVICE FAILED A POWER-ON SELF-TEST DUE TO AN ECG ERROR LOGGED BY THE DEVICE. THIS FAULT DISABLED ECG FUNCTIONALITY AND CHANGED THE READINESS INDICATOR TO A RED 'X', AS REPORTED BY THE OPERATOR. WE WERE UNABLE TO DUPLICATE A FAULT, BUT THE ANALOG BOARD WILL BE REPLACED AS A PRECAUTION. THE DEVICE WILL BE RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS FOR REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22586770&pc=OZD | 1220648-2025-30667 | 2025-07-08 04:00 | Death | ABIOMED, INC. | 2025-07-24 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND EXPERIENCED LIMB ISCHEMIA. THE TEAM PLACED A REPERFUSION SHEATH TO RESTORE FLOW AND PULSES. THE PUMP WAS SUPPORTING FOR 46 HOURS WHEN CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22586820&pc=NKM | 2135147-2025-04038 | 2025-06-27 04:00 | Death | ABBOTT MEDICAL | 2025-07-24 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100009 |  | 1 | Event Description: IN A SOCIAL MEDIA POST ON (B)(6) 2025, THE OUTCOMES OF SUMMARIZED MITRACLIP PROCEDURES WAS REPORTED FROM AN UNKNOWN LITERATURE ARTICLE. THE OUTCOMES AT TWO YEAR FOLLOW-UP ARE DESCRIBED AS ALL-CAUSE MORTALITY, HEART FAILURE HOSPITALIZATION, AND RECURRENT MITRAL REGURGITATION. THE SAMPLE SIZE OF PATIENT RECEIVING TRANSCATHETER EDGE TO EDGE REPAIR (TEER) WAS 602. THE AUTHOR OF THE SOCIAL MEDIA POST DOES NOT MAKE ANY ALLEGATIONS AGAINST THE MITRACLIP DEVICE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE IS NOT RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22587026&pc=LJT | 3006260740-2025-05169 | 2019-09-07 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-07-24 04:00 | LJT | M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Unspecified Infection; Sepsis | K873213 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITIGATION PROCESS THAT APPROXIMATELY TWO YEARS, SIX MONTHS AND TWENTY-FIVE DAYS LATER PORT PLACEMENT PROCEDURE, THE PATIENT DEVELOPED INFECTION, BACTEREMIA, SEPSIS AND ORGAN DYSFUNCTION. THE PORT WAS REMOVED AND WAS REPLACED WITH A NEW DEVICE. LATER THE PATIENT PASSED AWAY AND CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS ON (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H11: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22587089&pc=NGV | 2124215-2025-49006 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Endocarditis |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR ENDOCARDITIS. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 372 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22587219&pc=OZD | 1220648-2025-30668 | 2025-07-14 04:00 | Death | ABIOMED, INC. | 2025-07-24 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, A THROMBUS WAS NOTED IN THE LEFT ATRIUM AND THE LEFT VENTRICLE. SUPPORT WAS CONTINUED. ADDITIONALLY, OOZING WAS NOTED. THE LOCATION OF THE OOZING IS UNKNOWN AND THE INTERVENTION FOR THE OOZING IS UNKNOWN. THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION. SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22587386&pc=DYB | 3007284313-2025-04136 | 2025-06-27 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-07-24 04:00 | DYB | GORE® DRYSEAL FLEX INTRODUCER SHEATH | No Apparent Adverse Event | Obstruction/Occlusion | K160254 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, A PATIENT WAS TREATED FOR A PSEUDOANEURYSM IN THE ASCENDING AORTA USING A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (CTAG) AND A GORE® DRYSEAL FLEX INTRODUCER SHEATH (DSF). PRIOR TO THE CTAG BEING ADVANCED, A THROMBUS WAS DISLODGED FROM THE PROXIMAL THORACIC AORTA. THE THROMBUS TRAVELLED INTO THE SMA. AFTER DEPLOYMENT OF THE CTAG, THE PHYSICIAN ATTEMPTED TO PERFORM A THROMBECTOMY, AND ENDED UP DISSECTING THE SMA. THE PHYSICIAN CONVERTED TO OPEN REPAIR, AND REPAIRED THE SMA. AFTER CLOSING UP THE PATIENT AND TAKING THEM OUT OF THE OPERATING ROOM, THE PATIENT CODED. CHEST COMPRESSIONS WERE STARTED. THE PATIENT THEN DIED OF A LACERATED LIVER. IT IS UNCLEAR WHETHER THE THROMBUS WAS DISLODGED BY A GUIDEWIRE OR THE DSF. Manufacturer Narrative: H.6. INVESTIGATION FINDINGS CODE C20: THE DEVICE WAS DISCARDED AT THE TREATING FACILITY AND WAS THEREFORE NOT AVAILABLE FOR ANALYSIS. H.6. TYPE OF INVESTIGATION CODE B22: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE CONDUCTED BECAUSE THE SERIAL/LOT NUMBER REMAINS UNKNOWN. H.6. INVESTIGATION CONCLUSIONS CODE D12: IT SHOULD BE NOTED THAT, PER THE GORE® DRYSEAL FLEX INTRODUCER SHEATH INSTRUCTIONS FOR USE, COMPLICATIONS ASSOCIATED WITH THE USE OF THE GORE® DRYSEAL FLEX INTRODUCER SHEATH THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO, EMBOLIZATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22587402&pc=NGV | 2124215-2025-49005 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR UNSPECIFIED INFECTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 372 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22587600&pc=NGV | 2124215-2025-48908 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 9 | Event Description: THIS REPORT SUMMARIZES 9 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 73 TO (B)(6) YEARS. PATIENT SEX: 2 FEMALE; 7 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 586 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22587633&pc=KFM | 3013876692-2025-00059 | 2025-07-22 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-07-24 04:00 | KFM | HEART LUNG MACHINE | Electrical /Electronic Property Problem; Pumping Stopped | Cardiac Arrest | K991864 |  | 1 | Event Description: THE EVENT OCCURRED IN JAPAN. IT WAS REPORTED THAT THE ERROR MESSAGE ¿REFERENC ERROR¿ OCCURRED ON THE ROTAFLOW CONSOLE AND THE PUMP STOPPED DUE TO THE "REFERENC ERROR" DURING PATIENT USE. THE CUSTOMER USED THE EMERGENCY DRIVE AND REPLACED THE AFFECTED ROTAFLOW CONSOLE WITH ANOTHER ROTAFLOW CONSOLE. THE PUMP HAD BEEN USED ON A PATIENT WHO WAS IN A BAD CONDITION, AND THE PATIENT LATER PASSED AWAY, BUT THE PUMP HAD ONLY BEEN STOPPED FOR A SHORT TIME AND HIS BLOOD PRESSURE WAS STABLE, SO THERE IS NO CAUSAL RELATIONSHIP TO THIS ERROR ACCORDING TO THE CUSTOMER. DUE TO THE REPORTED DEATH OF THE PATIENT A REPORT IS REQUIRED. COMPLAINT ID: (B)(4).Event Description: COMPLAINT ID: (B)(4). Manufacturer Narrative: THE EVENT OCCURRED IN JAPAN. IT WAS REPORTED THAT THE ERROR MESSAGE ¿REFERENCE ERROR¿ OCCURRED ON THE ROTA FLOW CONSOLE AND THE PUMP STOPPED DUE TO THE "REFERENCE ERROR" DURING PATIENT USE. THE CUSTOMER USED THE EMERGENCY DRIVE AND REPLACED THE AFFECTED ROTA FLOW CONSOLE WITH ANOTHER ROTA FLOW CONSOLE. THE PUMP HAD BEEN USED ON A PATIENT WHO WAS IN A BAD CONDITION, AND THE PATIENT LATER PASSED AWAY, BUT THE PUMP HAD ONLY BEEN STOPPED FOR A SHORT TIME AND HIS BLOOD PRESSURE WAS STABLE, SO THERE IS NO CAUSAL RELATIONSHIP TO THIS ERROR ACCORDING TO THE CUSTOMER. FURTHER INFORMATION HAS BEEN PROVIDED BY THE SSU (SALES AND SERVICE UNIT) DATED ON 2025-07-22 AS FOLLOWS: A DIALYSIS PATIENT WAS URGENTLY HOSPITALIZED FOR CHEST PAIN ON (B)(6) 2025. CAG WAS PERFORMED ON (B)(6) 2025, BUT NO LESION WAS FOUND, AND THE PATIENT WAS MOVED DIRECTLY TO THE ICU. CHEST PAIN DEVELOPED AGAIN LATE AT NIGHT ON (B)(6) 2025, THE PATIENT BECAME UNCONSCIOUS, AND SUFFERED A CARDIAC ARREST, SO ECMO WAS STARTED AT 5:40 AM ON (B)(6) 2025 (USING ROTA FLOW). THERE WAS NO CHANGE IN THE PATIENT CONDITION BEFORE OR AFTER DEVICE REPLACEMENT. THE EXACT DATE OF THE DEATH WAS ON (B)(6) 2025. DUE TO THE REPORTED PUMP STOP DURING TREATMENT A REPORT IS REQUIRED. THE AFFECTED ROTA FLOW CONSOLE WITH S/N: (B)(6) WAS INVESTIGATED BY A GETINGE FIELD SERVICE TECHNICIAN AND THE REPORTED FAILURE COULD NOT BE REPRODUCED. NO PARTS HAS BEEN REPLACED. THE DEVICE WAS TESTED AND PASSED ALL FUNCTIONAL AND SAFETY TESTS ACCORDING TO THE SPECIFICATIONS. A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON (B)(6) 2025 WITH FOLLOWING CONCLUSION: "THE REPORTED EVENT INVOLVED A ¿REFERENCE¿ ERROR ON A ROTA FLOW PUMP, WHICH RESULTED IN THE UNEXPECTED STOPPAGE OF THE PUMP DURING ONGOING V-A ECMO THERAPY IN A CRITICALLY ILL 69-YEAR-OLD PATIENT WITH DIABETES AND CARDIAC ARREST. THE CLINICAL TEAM RESPONDED PROMPTLY BY INITIATING CHEST COMPRESSIONS AND ACTIVATING THE EMERGENCY DRIVE WITHIN APPROXIMATELY THREE MINUTES, FOLLOWED BY REPLACEMENT OF THE DEVICE. THE TOTAL TIME TO RESUME FULL CIRCULATORY SUPPORT WAS REPORTED TO BE AROUND TEN MINUTES. DURING THIS PERIOD, THE PATIENT¿S BLOOD PRESSURE REMAINED STABLE, AND NO DETERIORATION IN THE CLINICAL CONDITION WAS OBSERVED IMMEDIATELY BEFORE OR AFTER THE INCIDENT. THE DISPLAYED ERROR, AS REPORTED BY THE USER, INCLUDED ¿ERR OR!!¿ IN THE SPEED AND FLOW DISPLAY FIELDS, ALONG WITH THE MESSAGE ¿[REFERENCE]¿ IN THE STATUS AREA. ACCORDING TO THE INSTRUCTIONS FOR USE, THIS INDICATES THAT THE INTERNAL VOLTAGE REFERENCE WAS OUTSIDE THE VALID RANGE, WHICH RESULTS IN BOTH AN ACOUSTIC ALARM AND IMMEDIATE PUMP STOP. POST-EVENT ANALYSIS INVOLVED A FULL RUNNING TEST AND INTERNAL DIAGNOSTICS, BUT THE ISSUE COULD NOT BE REPRODUCED, AND NO HARDWARE ABNORMALITIES WERE FOUND. AS SUCH, A DEFINITIVE ROOT CAUSE COULD NOT BE IDENTIFIED. REASONABLE AND POSSIBLE CAUSES MAY INCLUDE A TRANSIENT ELECTRONIC OR VOLTAGE INSTABILITY, INTERNAL SIGNAL ERROR, OR MECHANICAL OR ELECTRICAL INFLUENCE SUCH AS A BRIEF SHOCK OR POWER FLUCTUATION. HOWEVER, AS THE REPORTED INCIDENT COULD NOT BE REPRODUCED BY THE FIELD SERVICE TECHNICIAN OF GETINGE, WHICH MAY INDICATE THAT THE ROOT CAUSE MAY BE FOUND WITHIN A BRIEF POWER FLUCTUATION. THE CLINICAL TEAM NOTED A ¿CRACK¿ SOUND WAS HEARD JUST PRIOR TO THE EVENT, AND THE DEVICE WAS REPORTEDLY POWERED VIA A TABLE TAP, WHICH MAY HAVE CONTRIBUTED TO A BRIEF INSTABILITY - THOUGH THIS REMAINS SPECULATIVE. NO USER ERROR HAS BEEN IDENTIFIED IN RELATION TO SETUP OR OPERATION. THE CLINICAL TEAM FOLLOWED APPROPRIATE EMERGENCY PROCEDURES, AND NO DEVIATIONS FROM STANDARD USE WERE REPORTED. WHILE THE EVENT DOES CONSTITUTE A TEMPORARY DEVICE MALFUNCTION, THERE IS NO EVIDENCE OF PERSISTENT PERFORMANCE DEGRADATION, AND THE ISSUE WAS RESOLVED WITH REPLACEMENT OF THE PUMP. AS STATED BY THE COMPLAINT NARRATIVE: ¿THE PUMP HAD BEEN USED ON A PATIENT WHO WAS IN A BAD CONDITION, AND THE PATIENT LATER DIED, BUT THE PUMP HAD ONLY BEEN STOPPED FOR A SHORT TIME AND HIS BLOOD PRESSURE WAS STABLE, SO THERE IS NO CAUSAL RELATIONSHIP TO THIS ERROR.¿ THEREFORE, THE CLINICAL OUTCOME IS BELIEVED TO BE PRIMARILY RELATED TO THE PATIENT¿S MEDICAL STATE RATHER THAN THE DEVICE INCIDENT. IN CONCLUSION, THE DEVICE BEHAVED IN ACCORDANCE WITH ITS DESIGN RESPONSE TO A CRITICAL INTERNAL ERROR AS DESCRIBED IN THE IFU. WHILE THE ROOT CAUSE OF THE VOLTAGE REFERENCE FAULT COULD NOT BE CONCLUSIVELY DETERMINED, ALL AVAILABLE DATA INDICATE THAT THE INCIDENT WAS ISOLATED, AND NO DIRECT LINK TO PATIENT HARM HAS BEEN ESTABLISHED." THE REVIEW OF THE NON-CONFORMITIES WAS PERFORMED ON 2025-07-22 AND DURING THE PERIOD OF 2012-10-31 TO 2025-07-22 DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED PRODUCT AND FAILURE. THERE IS NO INDICATION THAT MANUFACTURING ISSUES OCCURRED DURING THIS TIME, THUS PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE AFFECTED ROTA FLOW DEVICE WAS PRODUCED IN 2012-10-31. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. Manufacturer Narrative: THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE JAPANESE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, ROTAFLOW CONSOLE¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K991864. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR ROTAFLOW CONSOLE WITH CATALOG NUMBER 701051711.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22587668&pc=NGV | 2124215-2025-48924 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR A VERY LATE DEVICE THROMBOSIS (MORE THAN 1 YEAR POST IMPLANT). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 488 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22587670&pc=NGV | 2124215-2025-48909 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Endocarditis |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR ENDOCARDITIS. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 387 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22587671&pc=NGV | 2124215-2025-49007 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 6 | Event Description: THIS REPORT SUMMARIZES 6 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 61 TO 87 YEARS. PATIENT SEX: 3 FEMALE; 3 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 558 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22587725&pc=NGV | 2124215-2025-48911 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENTS FOR UNSPECIFIED INFECTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025 AVERAGE TIME TO EVENT: 387 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22587861&pc=LJT | 3006260740-2025-05171 | 2025-06-27 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-07-24 04:00 | LJT | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K063377 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITIGATION PROCESS THAT APPROXIMATELY ONE YEAR AND TEN MONTHS POST PORT PLACEMENT PROCEDURE, THE PORT WAS REMOVED FROM THE PATIENT. LATER THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS 01-JAN-1900 FOR THE MDR SUBMISSION REQUIREMENT. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22587885&pc=NPT | 2015691-2025-06011 | 2025-07-03 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-24 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM, 29MM | Difficult to Remove; Difficult to Advance | Perforation | P140031 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS FIELD CLINICAL SPECIALIST, DURING A TAVR PROCEDURE, THE PROCEDURE BEGAN WITH SUCCESSFUL INSERTION OF THE SHEATH, FOLLOWED BY ADVANCEMENT OF THE 29MM SAPIEN 3 ULTRA RESILIA VALVE AND COMMANDER DELIVERY SYSTEM. HOWEVER, DUE TO ANATOMICAL COMPLEXITY AT THE DISTAL END OF THE SHEATH INCLUDING A KNOWN DESCENDING AORTIC ANEURYSM, SEVERE VESSEL TORTUOSITY, AND SEVERE CALCIFICATION, THE VALVE AND PUSHER COULD NOT BE NAVIGATED THROUGH THE VESSEL. DURING ATTEMPTS TO TRAVERSE THE AFFECTED SEGMENT, AN AORTIC PERFORATION OCCURRED. THE DELIVERY SYSTEM COULD NOT BE WITHDRAWN THROUGH THE ESHEATH, AS THE VALVE HAD MIGRATED TOWARD THE TAPERED TIP AND WAS NO LONGER ALIGNED WITH THE PUSHER. THE PATIENT HAD A KNOWN DESCENDING AORTIC ANEURYSM, SEVERE ANNULAR AND VESSEL CALCIFICATION, AND SEVERE TORTUOSITY. EMERGENT SURGICAL INTERVENTION WAS PERFORMED, INCLUDING A SMALL CUT-DOWN, BALLOON TAMPONADE, AND PLACEMENT OF COVERED STENTS. THE SYSTEM WAS REMOVED AS A UNIT. DESPITE THESE EFFORTS, THE PATIENT DID NOT SURVIVE. THE EXACT CAUSE OF DEATH WAS NOT FORMALLY LABELED BUT WAS ATTRIBUTED TO THE VASCULAR COMPLICATION IN THE DESCENDING/ABDOMINAL AORTA. THE SHEATH WAS NOT INSERTED AT A STEEP ANGLE, AND THE VESSEL HAD BEEN PRE-DILATED TO 16 FRENCH. THE DELIVERY SYSTEM AND VALVE WERE PREPARED AND CRIMPED ACCORDING TO THE INSTRUCTIONS FOR USE (IFU). NO DEVICE DEFICIENCIES OR DAMAGES WERE REPORTED, AND NO USE ERROR WAS IDENTIFIED. THE DEVICES WERE DISCARDED AND ARE NOT AVAILABLE FOR RETURN. NO PROCEDURAL IMAGES, VIDEOS, OR 3MENSIO REPORT ARE AVAILABLE. THE PERCEIVED ROOT CAUSE OF THE EVENT WAS ATTRIBUTED TO ANATOMICAL CHALLENGES, INCLUDING THE DESCENDING AORTIC ANEURYSM, SEVERE VESSEL TORTUOSITY, SEVERE CALCIFICATION, AND SMALL LUMINAL DIAMETER. THE PATIENT'S LUMINAL DIAMETER WAS REPORTED TO BE APPROXIMATELY 10 MM. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22587909&pc=LJT | 3006260740-2025-05170 | 2018-06-20 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-07-24 04:00 | LJT | VACCESS CT LOW-PROFILE POWER-INJECTABLE IMPLANTABLE PORT, CHRONOFLEX, 8F | Adverse Event Without Identified Device or Use Problem | Sepsis | K063377 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITIGATION PROCESS THAT APPROXIMATELY NINE MONTHS AND FOURTEEN DAYS LATER PORT PLACEMENT PROCEDURE, THE PATIENT DEVELOPED SEVERE SEPSIS. LATER THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22588119&pc=PNB | 3005099803-2025-03277 | 2023-08-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | PNB | WALLFLEX BILIARY | Device Stenosis | Inflammation |  |  | 1 | Event Description: BOSTON SCIENTIFIC CORPORATION BECAME AWARE OF AN EVENT INVOLVING AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEMS AND WALLFLEX BILIARY RX FULLY COVERED STENT SYSTEMS THROUGH THE ARTICLE TITLED, "OPTIMIZING EUS-GUIDED CHOLEDOCHODUODENOSTOMY WITH LUMEN-APPOSING METAL STENTS FOR PRIMARY DRAINAGE OF MALIGNANT DISTAL BILIARY OBSTRUCTION (SCORPION-IIP): A PROSPECTIVE PILOT STUDY", BY (B)(6) ET AL. PER THE ARTICLE, BETWEEN AUGUST 2023 TO APRIL 2024, 27 PATIENTS SUFFERING FROM MALIGNANT DISTAL BILE DUCT OBSTRUCTION UNDERWENT EUS-GUIDED CHOLEDOCHODUODENOSTOMY (EUS-CDS) USING AXIOS AS LUMEN-APPOSING METAL STENTS (LAMS) WITH A PLACEMENT OF A WALLFLEX FULLY COVERED SELF-EXPANDING METAL STENTS THROUGH IT. ONE MORTALITY OCCURRED DURING THE TEST PERIOD OF THESE PROCEDURES POTENTIALLY RELATED TO INGROWTH IN THE WALLFLEX DEVICE AND CHOLECYSTITIS. THE PATIENT ALSO HAD CONCOMITANT RENAL FAILURE PLEASE SEE THE REFERENCED ARTICLE FOR FULL DETAILS. Manufacturer Narrative: BLOCK B2: THE EXACT DATE OF THE PATIENTS' DEATH WAS NOT REPORTED. THE RETROSPECTIVE STUDY START DATE OF (B)(6) 2023, IS USED FOR THE ESTIMATED DATE OF DEATH. BLOCK B3: THE EXACT DATE OF EVENT WAS NOT REPORTED. THE RETROSPECTIVE STUDY START DATE OF AUGUST 1, 2023, IS USED FOR THE ESTIMATED DATE OF EVENT. BLOCK D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WERE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK G2: LITERATURE SOURCE: JESKA A. FRITZSCHE; PAUL FOCKENS; MARC G. BESSELINK; OLIVER R. BUSCH; FREEK DAAMS; MATTHEUS C. B. WIELENGA; JOHANNA W. WILMINK; ROGIER P. VOERMANS; ROY L. J. VAN WANROOIJ. "OPTIMIZING EUS-GUIDED CHOLEDOCHODUODENOSTOMY WITH LUMEN-APPOSING METAL STENTS FOR PRIMARY DRAINAGE OF MALIGNANT DISTAL BILIARY OBSTRUCTION (SCORPION-IIP): A PROSPECTIVE PILOT STUDY." GASTROINTESTINAL ENDOSCOPY 2025; VOLUME 101, NO. 5: 1009-1016. BLOCK H6: IMDRF PATIENT CODE E2326 CAPTURES THE REPORTABLE EVENT OF CHOLECYSTITIS. IMDRF IMPACT CODE F02 CAPTURES THE PATIENT'S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22588314&pc=OZD | 1220648-2025-30680 | 2025-07-10 04:00 | Death | ABIOMED, INC. | 2025-07-24 04:00 | OZD | IMPELLA CP | Pumping Stopped | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THIRTEEN DAYS INTO IMPELLA CP SUPPORT, THE PUMP STOPPED UNEXPECTEDLY. ATTEMPTS TO RESTART THE PUMP WERE UNSUCCESSFUL AND THE DEVICE WAS SUBSEQUENTLY REMOVED. THE PATIENT WAS ESCALATED TO IMPELLA 5.5 SUPPORT FOLLOWING IMPELLA CP PUMP REMOVAL. PATIENT IS STABLE POST PUMP REPLACEMENT.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT EXPIRED ON THE SUBSEQUENT DEVICE, AN IMPELLA 5.5 (SERIAL NUMBER (B)(6)) AFTER EXPERIENCING A LARGE HEMORRHAGIC STROKE DURING IMPELLA 5.5. SUPPORT. THE PATIENT¿S FAMILY DECIDED TO WITHDRAW CARE. THE PATIENT EXPIRED. THIS COMPLAINT INVOLVES 2 PUMPS: PUMP 1: IMPELLA CP SERIAL NUMBER (B)(6). PUMP 2: IMPELLA 5.5, SERIAL NUMBER (B)(6). THIS MDR SPECIFICALLY ADDRESSES IMPELLA CP WITH SERIAL NUMBER 562192 WHICH IS THE SUBJECT OF THIS REPORT. A SEPARATE MEDWATCH REPORT WILL BE SUBMITTED FOR THE OTHER DEVICE ASSOCIATED WITH THIS COMPLAINT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: REREVIEW OF THE EVENT OCCURRED AND NOTED THE DEMISE OUTCOME IS ASSOCIATED WITH IMPELLA CP AS THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP AS AN ASSOCIATED FACTOR GIVEN THE PUMP STOP/LOSS OF SUPPORT. REFER TO MANUFACTURER REPORT NUMBER 1220648-2025-45845 FOR THE REPORT ON THE IMPELLA 5.5.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22588355&pc=OZD | 1220648-2025-30679 | 2025-07-10 04:00 | Death | ABIOMED, INC. | 2025-07-24 04:00 | OZD | IMPELLA CP | Fluid/Blood Leak; Pumping Stopped | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT EIGHTEEN DAYS INTO IMPELLA CP SUPPORT, A PURGE PRESSURE LOW ALARM APPEARED. NO VISIBLE LEAKS WERE OBSERVED. THREE HOURS LATER, A CLOT INGESTION RESULTED IN A SPIKE IN MOTOR CURRENT AND AN EVENTUAL PUMP STOP. ATTEMPTS TO RESTART THE PUMP WERE UNSUCCESSFUL AND SWAPPING THE AUTOMATED IMPELLA CONTROLLER DID NOT RESOLVE THE ISSUE. THE PATIENT REMAINED SUPPORTED VIA ECMO, BUT LATER PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22588571&pc=MKJ | 1220908-2025-02582 | 2025-06-09 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-07-24 04:00 | MKJ | X SERIES | No Display/Image | Cardiac Arrest | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE POWERED ON WITHOUT DISPLAY. EVENTUALLY THE SCREEN DID RETURN AND THE EVENT CONTINUED WITHOUT INCIDENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION. DURING THE INVESTIGATION IT WAS NOTED THAT THE UNIT SHUT DOWN DURING OPERATION. THE DEVICE WAS IN USE FOR 2 HOURS ON BATTERY POWER WITHOUT ISSUE. THE DEVICE THEN PROMPTED LOW BATTERY MESSAGES AND GAVE A SHUTDOWN WARNING 15 MINUTES LATER. THE DEVICE FUNCTIONED AS EXPECTED WHEN EVENTUALLY REPOWERED BY AUXILIARY. THE DEVICE TESTING AND LOG REVIEW SUPPORTS THE DEVICE IS FUNCTIONING AS EXPECTED. THE DEVICE WAS FULLY EVALUATED, INCLUDING TESTING WITH A LOW BATTERY, DISCHARGING, AND STRESSING WITH NO ANOMALIES OR FAULTS IDENTIFIED. THE DEVICE WILL BE RECERTIFIED FOR CLINICAL USE AND RETURNED TO THE CUSTOMER. NO TREND IDENTIFIED AGAINST SIMILAR REPORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22588865&pc=OZD | 1220648-2025-30686 | 2025-07-10 04:00 | Death | ABIOMED, INC. | 2025-07-24 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ON IMPELLA 5.5 SUPPORT DEVELOPED GASTROINTESTINAL BLEEDING REQUIRING TRANSFUSION OF MULTIPLE BLOOD PRODUCTS. EVENTUALLY, PATIENT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22589260&pc=NGV | 2124215-2025-49711 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN FXD CURVE? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR BLOOD LOSS. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22589272&pc=NGV | 2124215-2025-49712 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN FXD CURVE? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Retroperitoneal Hemorrhage |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR RETROPERITONEAL HEMORRHAGE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 1 DAY. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22589347&pc=NGV | 2124215-2025-49879 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN TRUSTEER? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR VESSEL TRAUMA. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22589526&pc=NIK | 9614453-2025-02927 | 2025-07-01 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-07-24 04:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P010031 |  | 1 | Event Description: THE PATIENT DIED APPROXIMATELY TWO WEEKS POST DEVICE REPLACEMENT. THE CAUSE OF DEATH WAS PROVIDED AS CARDIOVASCULAR DISEASE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22589582&pc=NGV | 2124215-2025-49713 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN FXD CURVE? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR VESSEL TRAUMA. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025 AVERAGE TIME TO EVENT: 4 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22589614&pc=NGV | 2124215-2025-49875 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN TRUSTEER? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR BLOOD LOSS. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22589801&pc=NGV | 2124215-2025-49710 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN FXD CURVE? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR UNEXPECTED MEDICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 4 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22590381&pc=NGV | 2124215-2025-49877 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-24 04:00 | NGV | WATCHMAN TRUSTEER? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Hematoma |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR HEMATOMA. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22591009&pc=DSQ | 2916596-2025-04875 | 2025-05-26 04:00 | Death | THORATEC CORPORATION | 2025-07-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. IT WAS REPORTED THAT THERE WAS NO PUMP ISSUES AND THAT THIS WAS A NATURAL DEATH. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22591230&pc=OZD | 1220648-2025-30700 | 2025-07-14 04:00 | Death | ABIOMED, INC. | 2025-07-24 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Device Sensing Problem | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT PERSISTENT SUCTION ISSUES COINCIDING WITH A LOW PLACEMENT SIGNAL ALARM APPEARED DURING IMPELLA CP SUPPORT. THE PATIENT'S LEFT VENTRICLE LOOKED UNDERFILLED, BUT DESPITE GIVING VOLUME AND BLOOD PRODUCTS, THE SUCTION REMAINED. THE DECISION WAS ULTIMATELY MADE TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22591257&pc=DSQ | 2916596-2025-04877 | 2023-03-19 04:00 | Death | THORATEC CORPORATION | 2025-07-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIORGAN FAILURE. IT WAS REPORTED THAT THERE WERE NO PUMP ISSUES. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE PUMP WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22591348&pc=DSQ | 2916596-2025-04876 | 2025-02-03 05:00 | Death | THORATEC CORPORATION | 2025-07-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIORGAN FAILURE WHILE ON HOSPICE, WITH NO PUMP ISSUES. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTI ORGAN FAILURE (INCLUDING RESPIRATORY FAILURE, RIGHT HEART FAILURE, HEPATIC DYSFUNCTION, AND RENAL DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22591540&pc=DSQ | 2916596-2025-04945 | 2025-07-13 04:00 | Death | THORATEC CORPORATION | 2025-07-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. ADDITIONAL INFORMATION WAS NOT PROVIDED. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WILL NOT BE EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTI ORGAN FAILURE (INCLUDING RESPIRATORY FAILURE, RIGHT HEART FAILURE, HEPATIC DYSFUNCTION, AND RENAL DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22591580&pc=NPT | 2025587-2025-05259 | 2025-06-01 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-24 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Patient Device Interaction Problem | Vascular Dissection | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING A COMPARISON OF THE 29-MM SAPIEN 3 WITH THE 34-MM EVOLUT-R IN PATIENTS WITH A LARGE AORTIC ANNULUS. THE STUDY POPULATION INCLUDED 172 PATIENTS WHO WERE PREDOMINANTLY MALE WITH A MEAN AGE OF 83 YEARS OLD. MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION; 62 PATIENTS WERE IMPLANTED WITH A MEDTRONIC EVOLUT R BIOPROSTHETIC VALVE. THREE PERI-PROCEDURAL DEATHS AND FIVE INTRA-HOSPITAL DEATHS OCCURRED IN THE STUDY POPULATION. ONE PROCEDURAL DEATH WAS WITH AN EVOLUT R PATIENT DUE TO AN AORTIC DISSECTION. NO FURTHER DETAILS WERE PROVIDED ON THIS DEATH. REGARDING THE REMAINING DEATHS, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. AMONG ALL EVOLUT R PATIENTS, CLINICAL OBSERVATIONS INCLUDED: STROKE, MAJOR VASCULAR OR BLEEDING COMPLICATION, MODERATE TO SEVERE AORTIC REGURGITATION, AND ARRHYTHMIA REQUIRING PERMANENT PACEMAKER IMPLANT. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: B3: DATE IS APPROXIMATE. MONTH AND YEAR ARE CONFIRMED VALID. CITATION: CANITROT ET AL. AN OBSERVATIONAL SINGLE-CENTER STUDY: COMPARISON OF THE 29-MM SAPIEN 3 WITH THE 34-MM EVOLUT-R IN PATIENTS WITH A LARGE AORTIC ANNULUS. JOURNAL OF INTERVENTIONAL CARDIOLOGY. JUNE 2025 (1). DOI: 10.1155/JOIC/5992132. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22591671&pc=DSQ | 2916596-2025-04907 | 2025-07-17 04:00 | Death | THORATEC CORPORATION | 2025-07-24 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE VENTRICULAR ASSIST DEVICE (VAD) COORDINATOR DID NOT MENTION ANY DEVICE RELATED CAUSES LEADING TO PATIENT DEATH. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WILL NOT BE EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22591695&pc=GDW | 3005075853-2025-05868 | 2025-07-08 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-07-24 04:00 | GDW | ECHELON 3000 60MM STANDARD | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Unspecified Tissue Injury | K213633 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN OPEN LEFT NEPHRECTOMY, THEY FIRED ACROSS THE RENAL AND SOMEHOW, THEY GOT INTO THE AORTA. THE PATIENT STARTS BLEEDING AND WAS STABLE OFF THE TABLE THEN WAS TRANSFERRED TO ICU BUT THE PATIENT PASSED AWAY. THERE'S NO ISSUE WITH THE DEVICE BUT THIS IS AN ADVERSE EVENT. Manufacturer Narrative: (B)(4). DATE SENT: 8/13/2025. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: I DO NOT KNOW THE INTERACTION BETWEEN THE AORTA AND THE STAPLER. LEFT RADICAL NEPHRECTOMY HE WILL USE THE 45 BUT IF HE CANT GET ACROSS THE RENAL VESSELS WITH THE 45 HE WILL GO TO THE 60 I DON'T KNOW IF IT WAS SKELETONIZED DON'T KNOW THE SIZE OF THE RENAL ARTERY HE SAID WHEN HE GOT THE KIDNEY REMOVED FROM THE BODY HE STARTED HAVING BLEEDING FROM THE AORTA. HE DOES NOT KNOW HOW IT REALLY HAPPENED. HE SAID HE FELT GOOD THAT HE GOT EVERYTHING FREED UP, AND ONCE THE KIDNEY WAS REMOVED BLOOD STARTED COMING FROM THE AORTA. DO NOT KNOW BLOOD LOSS IN ML THE PATIENT DID HAVE A TRANSFUSION THERE WERE CLAMPS NEAR THE AREA WERE STAPLER WAS FIRED HE HASN'T SAID YES OR NO IF HE THINKS THE STAPLER CONTRIBUTED TO THE BLEEDING THE DEVICE IS NOT AVAILABLE FOR ANALYSIS BECAUSE IT STAPLED AND CUT LIKE IT WAS SUPPOSED TO, TO FREE UP THE KIDNEY FOR REMOVAL. THEY DID NOT KEEP IT BECAUSE IT PERFORMED LIKE IT SUPPOSED TO. UNKNOWN ON THE AUTOPSY REPORT SURGEON HAS NOT REQUESTED TO SPEAK TO ETHICON MEDICAL OR ENGINEER. Manufacturer Narrative: (B)(4). DATE SENT: 7/24/2025. D4: BATCH # UNK. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE CLARIFY THE INTERACTION BETWEEN THE AORTA AND STAPLER? WHAT WAS THE INDICATION FOR THE PROCEDURE? DOES THE SURGEON ROUTINELY TAKE THE RENAL VESSELS WITH THE ECH60S? WAS THIS THE ENTIRE RENAL HILUM OR A SKELETONIZED RENAL ARTERY/VEIN? WHAT WAS THE APPROX. SIZE OF RENAL ARTERY? WHAT IS MEANT BY "THEY GOT INTO THE AORTA"? WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? WAS A TRANSFUSION REQUIRED? ANY CLIPS, CLAMPS, OR GRASPERS NEAR THE AREA WHERE THE DEVICE WAS BEING FIRED? DOES THE SURGEON BELIEVE THE ALLEGED DEVICE MALFUNCTION MAY HAVE CAUSED OR CONTRIBUTED TO THE SURGICAL OUTCOME? IF SO, WHY? IS THE DEVICE AVAILABLE FOR ANALYSIS? IS THERE AN AUTOPSY REPORT AVAILABLE? IS THE SURGEON INTERESTED IN SPEAKING WITH ETHICON MEDICAL AND ENGINEERING STAFF? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22591696&pc=GDW | 3005075853-2025-05867 | 2025-06-01 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-07-24 04:00 | GDW | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Hemorrhage/Bleeding; Embolism/Embolus | K141952 |  | 1 | Event Description: IT WAS REPORTED THAT POST-OP OF A LIVER RESECTION THAT THREE TO FOUR WEEKS AFTER THE PROCEDURE THE PATIENT PASSED. WHILE TRYING TO TAKE OUT THE PEDICLE BILE WAS COMING OUT, IN ORDER TO REINFORCE THE STAPLE LINE SUTURE WAS USED TO STOP THE BILE LEAKAGE. TWO DAYS AFTER THE CASE THEY HAD TO EMBOLIZE THE PATIENT. FIVE DAYS LATER THERE WAS SOME SORT OF VENOUS BLEEDING. Manufacturer Narrative: (B)(4). DATE SENT: 7/24/2025. B3: EXACT EVENT DATE UNK, ENTERED 6/1/2025 AS ONLY THE MONTH WAS PROVIDED. D4: BATCH # UNK. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: THE SURGEON HAD A PATIENT EXPIRE AFTER USING A PVS DEVICE. THE PATIENT PASSED ON (B)(6) 2025. THE INITIAL LIVER RESECTION PROCEDURE OCCURRED 3-4 WEEKS PRIOR. DURING THE INITIAL PROCEDURE THE SURGEON EXPERIENCED A BILE LEAK FROM THE STAPLE LINE AND REINFORCED WITH SUTURE. DAYS LATER THE SURGEON HAD TO EMBOLIZE THE PATIENT AND THE PATIENT HAD FURTHER COMPLICATIONS THAT LEAD TO THE EXPIRATION. HOW LONG WAS THE PROCEDURE? UNK. LIVER RESECTION TYPE? UNK. ANATOMICAL STRUCTURE ION OR DIVIDED? UNK. HAS THE REP BEEN IN LIVER CASES PRIOR? DOES THE SURGEON ISOLATE THE BILE DUCTS OR OFTEN INCORPORATE LIVER TISSUE? NO, THEY TRY NOT TO GET ANY LIVER TISSUE IN THE PVS DEVICE AND ISOLATE THE VESSEL. WERE ANY OTHER INSTRUMENTS USED IN THE DISSECTION? HARMONIC 1100 AND CLIPS. HAS AN AUTOPSY BEEN COMPLETED? UNK. IS THE SURGEON INTERESTED IN AN EXTERNAL RRT? UNK. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22591807&pc=LWS | 2649622-2025-20256 | 2025-07-21 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-24 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Dyspnea; Ventricular Fibrillation; Heart Failure/Congestive Heart Failure | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE, THE RIGHT VENTRICULAR (RV) DEFIBRILLATOR LEAD WAS PLACED IN THE PATIENT. THE PATIENT THEN EXPERIENCED SEVERE SHORTNESS OF BREATH AND SUDDEN ACUTE HEART FAILURE, AND SO THE SURGEON IMMEDIATELY TERMINATED THE PROCEDURE. AFTER INTUBATION AND RESCUE, THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). THE NEXT DAY, THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION (VF) AND DIED AFTER ADDITIONAL RESCUE ATTEMPTS FAILED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22591836&pc=NPT | 2025587-2025-05263 | 2025-06-30 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-24 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information; Central Regurgitation | Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 YEARS AND 4 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE VALVE FAILED DUE TO MODE RATE-SEVERE AORTIC INSUFFICIENCY. IT WAS FURTHER REPORTED THAT FOLLOWING THIS, THE PATIENT DIED DUE TO AN UNKNOWN REASON.Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 YEARS AND 4 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE VALVE FAILED DUE TO MODE RATE-SEVERE AORTIC INSUFFICIENCY. IT WAS FURTHER REPORTED THAT FOLLOWING THIS, THE PATIENT DIED DUE TO AN UNKNOWN REASON. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT NO TREATMENT OR INTERVENTION WAS PERFORMED FOR THE FAILED VALVE. ACCORDING TO THE PHYSICIAN, THE CAUSE OF DEATH WAS DUE TO MULTIPLE FACTORS. THE PHYSICIAN EXPLAINED THAT THE PATIENT WAS DISCHARGED FROM THE HOSPITAL TO BE READMITTED FOR TAVR (TRANSCATHETER AORTIC VALVE REPLACEMENT), BUT DURING THE DISCHARGE, THE PATIENT WENT TO THE EMERGENCY DEPARTMENT WITH AN ABDOMINAL BLEED. THE PHYSICIAN THEN SPECIFIED THAT THE PATIENT HAD A SPLENIC BLEED AND COILS WERE PLACED. ADDITIONALLY, THE PHYSICIAN INDICATED THAT THE VALVE CANNOT BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 (SECOND PARAGRAPH). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22591844&pc=NPT | 2025587-2025-05262 | 2025-03-01 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-24 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Viral Infection; Decreased Appetite | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 7 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT WAS ADMITTED TO THE HOSPITAL DUE TO COVID-19. SUBSEQUENTLY, THE PATIENT'S APPETITE DECREASED AND THE PATIENT DIED. THE CAUSE OF THE DEATH WAS UNKNOWN. PER THE PHYSICIAN, THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEVICE OR IMPLANT PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22591853&pc=OZD | 1220648-2025-30704 | 2025-07-14 04:00 | Death | ABIOMED, INC. | 2025-07-24 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia; Cardiac Perforation | P140003 |  | 1 | Event Description: US COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, IT WAS FOUND THAT THE ONLY FLOW TO THE LEFT LEG WAS FROM PROFUNDA. SWAN ATTEMPTED TO BE PLACED WITHOUT SUCCESS IN BILAT FEM VEINS DUE TO ANATOMY. ADDITIONALLY, A PERFORATION OCCURRED TO THE LEFT ANTERIOR DESCENDING ARTERY. THE CLINICAL REP STATED THAT THEY DID NOT BELIEVE IT WAS IMPELLA RELATED HOWEVER, THEY WERE NOT SURE WHEN IT OCCURRED. THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: LIMB ISCHEMIA: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿ VENTRICLE PERFORATION: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22592384&pc=MHX | 1220063-2025-00042 | 2025-06-25 04:00 | Death | DRAEGER MEDICAL SYSTEMS, INC | 2025-07-24 04:00 | MHX | INFINITY ACUTE CARE SYSTEM (M540) | No Device Output | Diabetic Ketoacidosis | K113798 |  | 1 | Event Description: A DIABETIC PATIENT WAS ADMITTED WITH HYPERGLYCEMIA AND TAKEN TO THE OPERATING ROOM FOR A RIGHT FOOT WOUND DEBRIDEMENT. THE PROCEDURE WAS COMPLETED, BUT UPON REVERSAL OF ANESTHESIA, THE PATIENT DID NOT RESPOND. RESUSCITATION EFFORTS WERE PERFORMED BUT WERE UNSUCCESSFUL, AND THE PATIENT PASSED AWAY. THE CUSTOMER REPORTED THAT THE HR (HEART RATE) WAVEFORM WAS VISIBLE ON THE M540 PATIENT MONITOR, BUT NO HR VALUES WERE DISPLAYED. ADDITIONALLY, HR VALUES WERE DASHED OUT ON THE PRINTOUTS. FOLLOW-UP INSPECTION (JULY 11, 2025): A SITE VISIT WAS CONDUCTED FOR INSPECTION AND ERROR ANALYSIS. DURING THE VISIT, THE MONITOR FILTER SETTING WAS CORRECTED FOR USE IN AN ELECTRICAL SURGICAL ENVIRONMENT FROM 'MONITOR' TO 'ESU'(ELECTRICAL SURGICAL UNIT), RESOLVING THE ISSUE WITH MISSING HR VALUES. HR VALUES APPEARED ON THE MONITOR AND WERE CORRECTLY PRINTED. THE PATIENT MONITOR WAS TESTED ON-SITE WITH THE CUSTOMER AND SUCCESSFULLY READ NIBP AND SPO2 DATA THROUGHOUT THE TEST. THE DATA WAS PRINTED CORRECTLY, CONFIRMING NORMAL DEVICE FUNCTIONALITY.Event Description: A DIABETIC PATIENT WAS ADMITTED WITH HYPERGLYCEMIA AND TAKEN TO THE OPERATING ROOM FOR A RIGHT FOOT WOUND DEBRIDEMENT. THE PROCEDURE WAS COMPLETED, BUT UPON REVERSAL OF ANESTHESIA, THE PATIENT DID NOT RESPOND. RESUSCITATION EFFORTS WERE PERFORMED BUT WERE UNSUCCESSFUL, AND THE PATIENT PASSED AWAY. THE CUSTOMER REPORTED THAT THE HR (HEART RATE) WAVEFORM WAS VISIBLE ON THE M540 PATIENT MONITOR, BUT NO HR VALUES WERE DISPLAYED. ADDITIONALLY, HR VALUES WERE DASHED OUT ON THE PRINTOUTS. FOLLOW-UP INSPECTION (JULY 11, 2025): A SITE VISIT WAS CONDUCTED FOR INSPECTION AND ERROR ANALYSIS. DURING THE VISIT, THE MONITOR FILTER SETTING WAS CORRECTED FOR USE IN AN ELECTRICAL SURGICAL ENVIRONMENT FROM 'MONITOR' TO 'ESU'(ELECTRICAL SURGICAL UNIT), RESOLVING THE ISSUE WITH MISSING HR VALUES. HR VALUES APPEARED ON THE MONITOR AND WERE CORRECTLY PRINTED. THE PATIENT MONITOR WAS TESTED ON-SITE WITH THE CUSTOMER AND SUCCESSFULLY READ NIBP AND SPO2 DATA THROUGHOUT THE TEST. THE DATA WAS PRINTED CORRECTLY, CONFIRMING NORMAL DEVICE FUNCTIONALITY. Manufacturer Narrative: AN INVESTIGATION HAS STARTED. A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION. Manufacturer Narrative: THIS INVESTIGATION IS ONGOING AS OF JULY 31ST, 2025. A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22592484&pc=NPS | 2135147-2025-04084 | 2018-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-24 04:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P230007 |  | 1 | Event Description: THE ARTICLE "HEART VALVE CENTER APPROACH TO SEVERE TRICUSPID REGURGITATION: REAL-WORLD DATA AND OUTCOMES" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE CLINICAL CHARACTERISTICS AND OUTCOMES OF PATIENTS WITH SEVERE TR MANAGED AT OUR HVC WITH DIFFERENT TREATMENT STRATEGIES, INCLUDING CONSERVATIVE MANAGEMENT, SURGERY, AND TRANSCATHETER REPAIR. DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP AND NON-ABBOTT DEVICES. THE ARTICLE CONCLUDED THAT IN PATIENTS WITH SYMPTOMATIC SEVERE TR MANAGED THROUGH A MODERN HVC APPROACH, INVASIVE TREATMENTS (TV SURGERY OR T-TEER) MAY PROVIDE A PROGNOSTIC BENEFIT OVER CONSERVATIVE MANAGEMENT, PARTICULARLY IN EARLY DISEASE STAGES (TRI-SCORE <6). [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS COSMO GODINO AT SAN RAFFAELE HOSPITAL INSTITUTION WITH CORRESPONDING EMAIL GODINO.COSMO@HSR.IT]. THE TIME FRAME OF THE STUDY WAS JANUARY 2018 TO DECEMBER 2023. A TOTAL OF 476 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 76 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, PREVIOUS STROKE, ATRIAL FIBRILLATION. (B)(4), UNK MITRACLIP, PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION. (B)(4), UNK TRICLIP, PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. ATTACHMENT: ARTICLE TITLED "HEART VALVE CENTER APPROACH TO SEVERE TRICUSPID REGURGITATION: REAL-WORLD DATA AND OUTCOMES".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22592646&pc=NKM | 2135147-2025-04088 | 2019-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-24 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "THE PROGNOSTIC VALUE OF SMALL PSOAS MUSCLE AREA IN PATIENTS UNDERGOING MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A SINGLE CENTER RETROSPECTIVE STUDY, TO EVALUATE THE ASSOCIATION BETWEEN PSOAS MUSCLE AREA (PMA) MEASURED BY COMPUTED TOMOGRAPHY (CT) AS AN INDEX OF SARCOPENIA, AND CLINICAL OUTCOMES FOLLOWING MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR (M-TEER) IN HIGH-RISK PATIENTS. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT DECREASED PMA MEASURED BY PREPROCEDURAL CT IMAGING IN PATIENTS UNDERGOING M-TEER FOR MR WAS ASSOCIATED WITH ADVERSE CLINICAL OUTCOMES. SMALL PMA CAN BE A POTENTIAL POOR PROGNOSTIC INDICATOR IN PATIENTS UNDERGOING M-TEER. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS TAKAO MORIKAWA AT DEPARTMENT OF CARDIOVASCULAR MEDICINE, THE SAKAKIBARA HEART INSTITUTE OF OKAYAMA, JAPAN WITH CORRESPONDING EMAIL TAKAWOMORIKAWA5@GMAIL.COM. THE TIME FRAME OF THE STUDY WAS JANUARY 2019 AND JULY 2023. A TOTAL OF 214 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. AVERAGE AGE WAS 80 AND MAJORITY SEX WAS MALE. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, CHRONIC KIDNEY DISEASE, DIALYSIS, COPD, STROKE, DIABETES, DYSLIPIDEMIA, HYPERTENSION, PERIPHERAL ARTERY DISEASE, CARDIAC RESYNCHRONIZATION THERAPY, PERVIOUS PERCUTANEOUS INTERVENTION, MYOCARDIAL INFARCTION, CARDIOTHORACIC SURGERY, LIVE CIRRHOSIS, HEART FAILURE, CLINICAL FRAILTY 3-8, HEART FAILURE MEDICATION, DEGENERATIVE MITRAL REGURGITATION, TRICUSPID REGURGITATION, MITRAL STENOSIS. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE HOSPITALIZATION, MITRAL STENOSIS. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. ARTICLE TITLED "THE PROGNOSTIC VALUE OF SMALL PSOAS MUSCLE AREA IN PATIENTS UNDERGOING MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22592658&pc=DXY | 2182208-2025-03686 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-07-24 04:00 | DXY | MEDTRONIC LEAD | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING RIGHT VENTRICULAR PACING (RVP) VERSUS LEFT BUNDLE BRANCH AREA PACING (LBBAP). THE AUTHORS DESCRIBED ONE PATIENT DEATH IN THE RVP GROUP DUE TO SUDDEN CARDIAC DEATH SEVEN MONTHS POST IMPLANT. THERE WERE PATIENTS WHO EXPERIENCED PERFORATION, ACUTE MYOCARDIAL INFARCTION, AND LEADS WHICH EXHIBITED DISLODGMENTS, INCREASED THRESHOLDS, AND DECREASES IN R WAVE AMPLITUDE. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/68 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: COMPARISON OF CLINICAL COMPLICATIONS BETWEEN LBBAP AND TRADITIONAL RVP IN LONG-TERM FOLLOW-UP. KARDIOLOGIYA. 2025. 65(1):27¿33. DOI: 10.18087/CARDIO.2025.1.N2809 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22592659&pc=LWS | 2649622-2025-20265 | 2025-06-30 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-24 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Under-Sensing | Arrhythmia | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED AFTER RECEIVING ANTI-TACHYCARDIA PACING (ATP) AND ONE HIGH VOLTAGE THERAPY FOR AN EPISODE OF VENTRICULAR FIBRILLATION (VF). THE PHYSICIAN WANTED TO KNOW WHY THE PATIENT ONLY RECEIVED THE ONE HIGH VOLTAGE THERAPY. POST-THERAPY EPISODES SHOWED THE PATIENT IN VENTRICULAR ARRHYTHMIA WITH RATES IN AND OUT OF DETECTION THAT NEVER MET CRITERIA FOR A SECOND THERAPY. THERE MAY HAVE BEEN SOME FINE RIGHT VENTRICULAR (RV) LEAD UNDERSENSING. THE PHYSICIAN, WHO WAS CURIOUS IF THEY'D PROGRAMMED THE PATIENT CORRECTLY, DID NOT FEEL A DEVICE INTERROGATION WAS NEEDED TO SEE IF ANY THERAPIES WERE EVENTUALLY DELIVERED AFTER THE INITIAL. THE PATIENT WAS BELIEVED TO HAVE EXPIRED DUE TO ARRHYTHMIA. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS; HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED RIGHT VENTRICULAR UNDERSENSING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22592722&pc=NPT | 2015691-2025-06022 | 2025-07-01 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-24 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Degraded | Cardiac Arrest | P140031 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS FIELD CLINICAL SPECIALIST, A PATIENT WITH A FAILED 26MM SAPIEN 3 AORTIC VALVE IMPLANTED FOR 5 YEARS AND 7 MONTHS WAS BEING EVALUATED FOR VALVE-IN-VALVE THERAPY DUE TO SEVERE BIOPROSTHETIC VALVE STENOSIS AND AORTIC INSUFFICIENCY. THE PATIENT EXPIRED PRIOR TO THE SCHEDULED TREATMENT. PER MEDICAL RECORDS, THE PATIENT HAD 3 HOSPITAL ADMISSIONS BETWEEN (B)(6) 2025 AND (B)(6) 2025 FOR ACUTE ON CHRONIC CONGESTIVE HEART FAILURE (CHF). A PERCUTANEOUS CORONARY INTERVENTION WAS PERFORMED ON (B)(6) 2025 AND THE PATIENT WAS SCHEDULED FOR A VALVE-IN-VALVE PROCEDURE WITHIN THE NEXT COUPLE OF WEEKS. BY THE 3RD ADMISSION, THE PATIENT HAD PROGRESSIVE WEAKNESS, SHORTNESS OF BREATH, HYPOTENSION, ELEVATED CREATININE, ELEVATED TROPONIN, AND ELEVATED BNP. THE PATIENT'S PRESENTING FINDINGS WERE PRIMARILY CONCERNING FOR VALVULAR HEART FAILURE DUE TO THE PATIENT'S SEVERE AORTIC VALVE STENOSIS, AORTIC INSUFFICIENCY WITH LOW CARDIAC OUTPUT, HYPOTENSION, AND HEPATORENAL FAILURE WITH INCREASING CREATINE. AFTER CONSULTATION WITH CARE TEAM MEMBERS, PALLIATIVE CARE WAS CONSULTED DUE TO THE PATIENT'S POOR PROGNOSIS. HOWEVER, THE PATIENT'S FAMILY REQUESTED FULL CODE AS WELL AS EXPLORING SURGICAL INTERVENTION. A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE WAS PLANNED ON (B)(6) 2025. OVERNIGHT, THE PATIENT BECAME UNRESPONSIVE, APNEIC AND BRADYCARDIC. CARDIAC ARREST ENSUED. WHILE CPR WAS INITIATED, THE PATIENT SUCCUMBED TO HIS DEATH ON (B)(6) 2025. PER THE HOSPITAL DEATH SUMMARY, THE PATIENT'S FINAL DIAGNOSIS WAS DEATH FROM CARDIOPULMONARY ARREST SECONDARY TO COMBINED SYSTOLIC AND DIASTOLIC CHF IN SETTING OF SEVERE AORTIC VALVE STENOSIS. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), STRUCTURAL VALVE DETERIORATION (WEAR, FRACTURE, CALCIFICATION, LEAFLET TEAR/TEARING FROM THE STENT POSTS, LEAFLET RETRACTION, SUTURE LINE DISRUPTION OF COMPONENTS OF A PROSTHETIC VALVE, THICKENING, STENOSIS), AND DEVICE DEGENERATION ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH THE USE OF BIOPROSTHETIC HEART VALVES. THE IFU CAUTIONS THAT ACCELERATED DETERIORATION OF THE VALVE MAY OCCUR IN PATIENTS WITH AN ALTERED CALCIUM METABOLISM. LONG-TERM DURABILITY HAS NOT BEEN ESTABLISHED FOR THE VALVE. REGULAR MEDICAL FOLLOW-UP IS ADVISED TO EVALUATE VALVE PERFORMANCE. STRUCTURAL VALVE DETERIORATION (SVD) MAY BE MANIFESTED AS STENOSIS WITH THICKENED LEAFLETS. SVD MAY BE MILD AND NOT REQUIRE ANY INTERVENTION OR IT MAY BE MODERATE TO SEVERE. IT CAN CAUSE THE HEART TO WORK HARDER TO EJECT BLOOD FROM THE VENTRICLE. DEPENDING ON THE SEVERITY IT COULD BE AN INDICATION FOR VALVE REPLACEMENT OR MEDICAL INTERVENTION. TISSUE CALCIFICATION IS A COMMON FAILURE MODE. THE MECHANISMS FOR BIOPROSTHETIC HEART VALVE TISSUE CALCIFICATION ARE NOT YET FULLY UNDERSTOOD. MANY FACTORS CAN CONTRIBUTE TO THE ONSET AND PROPAGATION OF CALCIFICATION INCLUDING PATIENT-RELATED (E.G., PATIENT AGE, DISEASE STATE, IMMUNE STATUS, AND OTHER CO-MORBIDITIES), PHARMACOLOGICAL, AND INTRINSIC PROPERTIES OF THE VALVE ITSELF. IT IS WIDELY UNDERSTOOD THAT PATIENTS WITH CHRONIC RENAL DISEASE AND PRIOR HISTORY OF CALCIFIC STENOSIS OF THE NATIVE VALVE MAY BE PREDISPOSED TO BIOPROSTHETIC CALCIFICATION. DURING THE MANUFACTURING PROCESS, ALL SAPIEN THV VALVES ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% FUNCTIONALLY TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A MANUFACTURING NON-CONFORMANCE CONTRIBUTED TO THIS PRODUCT PROBLEM. THE ROOT CAUSE OF THE EVENT REMAINS INDETERMINABLE, BUT PATIENT COMORBIDITIES SUCH AS HYPERLIPIDEMIA MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22592848&pc=DSQ | 2916596-2025-04931 | 2023-05-11 04:00 | Death | THORATEC CORPORATION | 2025-07-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. ADDITIONAL INFORMATION WAS NOT PROVIDED.Event Description: THE PATIENT PASSED AWAY DUE TO MULTIORGAN FAILURE, AND THERE WAS NO PUMP ISSUE NOTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22592901&pc=DSQ | 2916596-2025-04896 | 2025-04-18 04:00 | Death | THORATEC CORPORATION | 2025-07-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22593018&pc=NKM | 2135147-2025-04082 | 2018-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-24 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Off-Label Use | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "HEART VALVE CENTER APPROACH TO SEVERE TRICUSPID REGURGITATION: REAL-WORLD DATA AND OUTCOMES" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE CLINICAL CHARACTERISTICS AND OUTCOMES OF PATIENTS WITH SEVERE TR MANAGED AT OUR HVC WITH DIFFERENT TREATMENT STRATEGIES, INCLUDING CONSERVATIVE MANAGEMENT, SURGERY, AND TRANSCATHETER REPAIR. DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP AND NON-ABBOTT DEVICES. THE ARTICLE CONCLUDED THAT IN PATIENTS WITH SYMPTOMATIC SEVERE TR MANAGED THROUGH A MODERN HVC APPROACH, INVASIVE TREATMENTS (TV SURGERY OR T-TEER) MAY PROVIDE A PROGNOSTIC BENEFIT OVER CONSERVATIVE MANAGEMENT, PARTICULARLY IN EARLY DISEASE STAGES (TRI-SCORE <6). [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS COSMO GODINO AT SAN RAFFAELE HOSPITAL INSTITUTION WITH CORRESPONDING EMAIL GODINO.COSMO@HSR.IT]. THE TIME FRAME OF THE STUDY WAS JANUARY 2018 TO DECEMBER 2023. A TOTAL OF 476 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 76 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, PREVIOUS STROKE, ATRIAL FIBRILLATION. (B)(6), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION. (B)(6), UNK TRICLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. ATTACHMENT: ARTICLE TITLED "HEART VALVE CENTER APPROACH TO SEVERE TRICUSPID REGURGITATION: REAL-WORLD DATA AND OUTCOMES."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22593116&pc=DSQ | 3007042319-2025-01517 | 2016-07-09 04:00 | Death | HEARTWARE, INC. | 2025-07-24 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE A VENTRICULAR ASSIST DEVICE (VAD) PATIENT HAD THEIR VAD DEACTIVATED FOR AN UNKNOWN REASON AND LATER DIED AFTER THE WITHDRAWAL OF CARE. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THE COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS WAS NOT PERFORMED SINCE LOG FILES WERE NOT AVAILABLE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT CARE WAS WITHDRAWN AND THE VAD TURNED OFF FOR UNKNOWN REASONS, RESULTING IN THE PATIENT¿S DEATH. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22593365&pc=DSQ | 2916596-2025-04874 | 2025-01-16 05:00 | Death | THORATEC CORPORATION | 2025-07-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure; Aortic Valve Insufficiency/ Regurgitation; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS FAILURE TO THRIVE AND THE FAMILY DECIDED TO WITHDRAW SUPPORT. IT WAS NOT PROVIDED WHETHER THIS WAS DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED.Event Description: IT WAS REPORTED THAT THE PATIENT HAD AN ECHOCARDIOGRAM ON (B)(6) 2025 NOTED THE RIGHT VENTRICLE BEING MODERATELY REDUCED. THE PATIENT HAD CHRONIC PROGRESSING RIGHT VENTRICULAR FAILURE. THE RIGHT VENTRICLE HAD BEEN MONITORED SINCE IMPLANT. THE RIGHT VENTRICLE WAS MILDLY REDUCED PRIOR TO IMPLANT. THE PATIENT ALSO HAD AORTIC INSUFFICIENCY MAKING IT A CHALLENGE TO INCREASE PUMP SPEED. THE PATIENT REQUIRED FREQUENT DIURETICS AND PARACENTESIS TO ASSIST WITH VOLUME MANAGEMENT. THEY HAD CHRONIC PAIN UNRELATED TO THE DEVICE AND FREQUENT MECHANICAL FALLS. THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS FAILURE TO THRIVE AND THE FAMILY DECIDED TO WITHDRAW SUPPORT. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. THE PATIENT HAD ACUTE CLINICAL SYSTOLIC HEART FAILURE AND END STAGE HEART FAILURE LEADING UP TO DEATH. THIS WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THIS DOCUMENT ALSO WARNS THE USER THAT MODERATE TO SEVERE AORTIC INSUFFICIENCY MUST BE CORRECTED AT THE TIME OF DEVICE IMPLANT. SECTION 6 "PATIENT CARE AND MANAGEMENT" (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6 (UNDER CAUTION!) STATES: "RIGHT HEART FAILURE CAN OCCUR FOLLOWING IMPLANTATION OF THE PUMP. RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH ELEVATED PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP." NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22594821&pc=ITJ | 1417592-2025-00353 | 2025-07-07 04:00 | Death | MEDLINE INDUSTRIES, LP | 2025-07-24 04:00 | ITJ | ROLLATOR,STEEL,KD | Appropriate Term/Code Not Available | Fall |  |  | 1 | Event Description: THE ROLLATOR, WITH THE BRAKES LOCKED, SLID OUT FROM UNDER HER IN THE BATHROOM, CAUSING CUSTOMER TO FALL. Manufacturer Narrative: THERE WAS NO CONTACT INFORMATION PROVIDED IN THE REPORT THAT MEDLINE RECEIVED FROM AMAZON TO FOLLOW UP WITH THE CUSTOMER THEREFORE NO ADDITIONAL INFORMATION IS AVAILABLE TO BE OBTAINED. PER THE REPORT THE ROLLATOR, WITH THE BRAKES LOCKED, SLID OUT FROM UNDER HER IN THE BATHROOM, CAUSING CUSTOMER TO FALL, THIS HAS LED TO AN ADVERSE EVENT SINCE THE CUSTOMER PASSED AWAY FROM THE FALL A WEEK LATER'. IT HAS BEEN DETERMINED THAT THE REPORTED EVENT CAUSED OR CONTRIBUTED TO SERIOUS INJURY REQUIRING MEDICAL INTERVENTION, THEREFORE, THIS MEDWATCH IS BEING FILED. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22596133&pc=DSQ | 2916596-2025-04884 | 2025-02-04 05:00 | Death | THORATEC CORPORATION | 2025-07-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Septic Shock; Multiple Organ Failure; Heart Failure/Congestive Heart Failure; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. ADDITIONAL INFORMATION WAS NOT PROVIDED.Event Description: IT WAS REPORTED THAT THE PATIENT'S CAUSE OF DEATH WAS SEPTIC SHOCK, DECOMPENSATED RIGHT VENTRICULAR FAILURE AND END-ORGAN FAILURE. NO PUMP RELATED ISSUE WAS NOTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS SEPSIS, MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. ALTHOUGH ONLY SEPSIS WAS REPORTED BY THE ACCOUNT, THE IFU ALSO LISTS INFECTION AS AN ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, AS WELL AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK ALSO PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22596660&pc=DSQ | 2916596-2025-04895 | 2025-07-08 04:00 | Death | THORATEC CORPORATION | 2025-07-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIORGAN FAILURE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22596774&pc=DSQ | 2916596-2025-04897 | 2024-04-04 04:00 | Death | THORATEC CORPORATION | 2025-07-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22596971&pc=LWS | 2017865-2025-95902 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-24 04:00 | LWS | FORTIFY ASSURA VR ICD, US | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS CONGESTIVE HEART FAILURE. Manufacturer Narrative: THE DEVICE WAS RETURNED FOR EVALUATION. THE DEVICE VOLTAGE WAS ABOVE ELECTIVE REPLACEMENT INDICATOR (ERI) LEVEL UPON RECEIPT. ANALYSIS OF DEVICE IMAGE INDICATED THE DEVICE WAS WITHIN NORMAL RANGE OF OPERATION. ASSESSMENT OF TOTAL PROJECTED LONGEVITY WAS PERFORMED, AND THE DEVICE WAS FOUND TO HAVE APPROPRIATE REMAINING LONGEVITY. FURTHER ANALYSIS OF THE DEVICE¿S LEAD IMPEDANCE, SENSING, PACING, AND HIGH VOLTAGE CHARGING WERE FOUND TO BE NORMAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22597219&pc=NVY | 2017865-2025-95903 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-24 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS CONGESTIVE HEART FAILURE. Manufacturer Narrative: ANALYSIS WAS NORMAL. NO ANOMALIES WERE NOTED. THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. VISUAL AND X-RAY EXAMINATIONS OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22601166&pc=MKJ | 3030677-2025-001711 | 2025-02-20 05:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-07-24 04:00 | MKJ | HEARTSTART | Failure to Power Up | Cardiac Arrest | P160029 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT THE DEVICE WOULD NOT POWER ON. THE PATIENT DID NOT SURVIVE. Manufacturer Narrative: UPDATED CATALOG NUMBER AND MODEL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22573764&pc=NIQ | 9612164-2025-03567 | 2024-12-10 05:00 | Death | MEDTRONIC IRELAND | 2025-07-23 04:00 | NIQ | ONYX FRONTIER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Ventricular Fibrillation; Thrombosis/Thrombus; Asystole | P160043 |  | 1 | Event Description: DURING THE INDEX PROCEDURE A TOTAL OF SEVEN ONYX FRONTIER DRUG ELUTING STENTS (DES) WERE IMPLANTED. TWO PREVAIL DRUG COATED BALLOONS (DCB) AND THREE LAUNCHER GUIDE CATHETERS WERE ALSO USED DURING THE INDEX PROCEDURE. LESIONS TREATED DURING THE INDEX PROCEDURE INCLUDED THE THE 1ST DIAGONAL ARTERY, THE LEFT MAIN CORONARY ARTERY (LCMA), THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY(LAD), PROXIMAL CIRCUMFLEX (CX), MID AND PROXIMAL RIGHT CORONARY ARTERY (RCA). A BARE METAL STENT WAS IMPLANTED IN THE LAD. THE LCMA WAS TREATED WITH ANGIOPLASTY. APPROXIMATELY 16 MONTHS POST THE INDEX PROCEDURE THE PATIENT SUFFERED FROM A MYOCARDIAL INFARCTION (MI). THE PATIENT COLLAPSED AT HOME WITH CARDIAC ARREST. VENTRICULAR FIBRILLATION WAS SEEN ON THE ELECTROCARDIOGRAM (ECG). INITIALLY THERE WAS ASYSTOLE. THE PATIENT HAD 14 SHOCKS. 10MG OF ADRENALINE AND 450MG OF AMIODARONE WAS ADMINISTERED. THE PATIENT WAS IN ASYSTOLE ON ARRIVAL IN THE EMERGENCY DEPARTMENT. THE PATIENT DIED. THE MI WAS CLASSIFIED AS A TYPE ONE SPONTANEOUS MI. THE LOCATION OF THE MI WAS UNKNOWN. THE DEATH WAS CLASSIFIED AS A SUDDEN CARDIAC DEATH. THE PATIENT WAS NOT TAKING DUAL ANTIPLATELET THERAPY WITHIN 24 HOURS PRIOR TO THE EVENT. THE SITES COMMENTED THAT IT WAS POSSIBLE TO BE A STENT THROMBOSIS RELATED EVENT WITH THE CARDIAC ARREST. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: NO ANGIOGRAM WAS PERFORMED. AS PER PHYSICIAN, THE MI WAS POSSIBLY RELATED TO THE ONYX FRONTIER STENTS IMPLANTED DURING THE INDEX PRO CEDURE. AS PER THE PHYSICIAN, THE CAUSE OF DEATH WAS MYOCARDIAL INFARCTION AND ISCHEMIC HEART DISEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATE RECEIVED: DURING THE INDEX PROCEDURE, A PREVAIL 3.0 × 20 MM BALLOON WAS USED IN THE DISTAL RCA/PLV AT 8 ATM FOR 60 SECONDS, FOLLOWED BY DEPLOYMENT OF AN ONYX FRONTIER 3.5 × 34 MM STENT IN THE DISTAL RCA INTO THE PDA AT 10 ATM FOR 22 SECONDS, AN ONYX FRONTIER 4.0 × 30 MM STENT IN THE OSTIAL RCA AT 12 ATM FOR 24 SECONDS AND AGAIN AT 12 ATM FOR 20 SECONDS, AND A PREVAIL DEB 3.0 × 30 MM IN THE OSTIAL TO MID LEFT CIRCUMFLEX AT 8 ATM FOR 40 SECONDS AND 20 SECONDS; SUBSEQUENTLY, AN ONYX FRONTIER 3.0 × 15 MM STENT WAS DEPLOYED IN THE PROXIMAL DIAGONAL AT 12 ATM FOR 12 SECONDS, AN ONYX FRONTIER 3.0 × 12 MM STENT WAS IMPLANTED IN THE OSTIUM OF THE DIAGONAL AT 12 ATM FOR 10 SECONDS AND 16 ATM FOR 10 SECONDS, AN ONYX FRONTIER 5.0 × 15 MM STENT WAS IMPLANTED FROM THE LMS INTO THE LAD AT 10 ATM FOR 10 SECONDS, AN ONYX FRONTIER 3.5 × 26 MM STENT WAS DEPLOYED IN THE OSTIAL LAD AT 10 ATM FOR 8 SECONDS, AND FINALLY, AN ONYX FRONTIER 4.0 × 15 MM STENT WAS DEPLOYED FROM THE LMS INTO THE OSTIAL LAD AT 16 ATM FOR 12 SECONDS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22573768&pc=NIQ | 9612164-2025-03568 | 2024-12-10 05:00 | Death | MEDTRONIC IRELAND | 2025-07-23 04:00 | NIQ | ONYX FRONTIER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Ventricular Fibrillation; Thrombosis/Thrombus; Asystole | P160043 |  | 1 | Event Description: DURING THE INDEX PROCEDURE A TOTAL OF SEVEN ONYX FRONTIER DRUG ELUTING STENTS (DES) WERE IMPLANTED. TWO PREVAIL DRUG COATED BALLOONS (DCB) AND THREE LAUNCHER GUIDE CATHETERS WERE ALSO USED DURING THE INDEX PROCEDURE. LESIONS TREATED DURING THE INDEX PROCEDURE INCLUDED THE 1ST DIAGONAL ARTERY, THE LEFT MAIN CORONARY ARTERY (LCMA), THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY(LAD), PROXIMAL CIRCUMFLEX (CX), MID AND PROXIMAL RIGHT CORONARY ARTERY (RCA). A BARE METAL STENT WAS IMPLANTED IN THE LAD. THE LCMA WAS TREATED WITH ANGIOPLASTY. APPROXIMATELY 16 MONTHS POST THE INDEX PROCEDURE THE PATIENT SUFFERED FROM A MYOCARDIAL INFARCTION (MI). THE PATIENT COLLAPSED AT HOME WITH CARDIAC ARREST. VENTRICULAR FIBRILLATION WAS SEEN ON THE ELECTROCARDIOGRAM (ECG). INITIALLY THERE WAS ASYSTOLE. THE PATIENT HAD 14 SHOCKS. 10MG OF ADRENALINE AND 450MG OF AMIODARONE WAS ADMINISTERED. THE PATIENT WAS IN ASYSTOLE ON ARRIVAL IN THE EMERGENCY DEPARTMENT. THE PATIENT DIED. THE MI WAS CLASSIFIED AS A TYPE ONE SPONTANEOUS MI. THE LOCATION OF THE MI WAS UNKNOWN. THE DEATH WAS CLASSIFIED AS A SUDDEN CARDIAC DEATH. THE PATIENT WAS NOT TAKING DUAL ANTIPLATELET THERAPY WITHIN 24 HOURS PRIOR TO THE EVENT. THE SITES COMMENTED THAT IT WAS POSSIBLE TO BE A STENT THROMBOSIS RELATED EVENT WITH THE CARDIAC ARREST. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: NO ANGIOGRAM WAS PERFORMED. AS PER PHYSICIAN, THE MI WAS POSSIBLY RELATED TO THE ONYX FRONTIER STENTS IMPLANTED DURING THE INDEX PROCEDURE. AS PER THE PHYSICIAN, THE CAUSE OF DEATH WAS MYOCARDIAL INFARCTION AND ISCHEMIC HEART DISEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATE RECEIVED: DURING THE INDEX PROCEDURE, A PREVAIL 3.0 × 20 MM BALLOON WAS USED IN THE DISTAL RCA/PLV AT 8 ATM FOR 60 SECONDS, FOLLOWED BY DEPLOYMENT OF AN ONYX FRONTIER 3.5 × 34 MM STENT IN THE DISTAL RCA INTO THE PDA AT 10 ATM FOR 22 SECONDS, AN ONYX FRONTIER 4.0 × 30 MM STENT IN THE OSTIAL RCA AT 12 ATM FOR 24 SECONDS AND AGAIN AT 12 ATM FOR 20 SECONDS, AND A PREVAIL DEB 3.0 × 30 MM IN THE OSTIAL TO MID LEFT CIRCUMFLEX AT 8 ATM FOR 40 SECONDS AND 20 SECONDS; SUBSEQUENTLY, AN ONYX FRONTIER 3.0 × 15 MM STENT WAS DEPLOYED IN THE PROXIMAL DIAGONAL AT 12 ATM FOR 12 SECONDS, AN ONYX FRONTIER 3.0 × 12 MM STENT WAS IMPLANTED IN THE OSTIUM OF THE DIAGONAL AT 12 ATM FOR 10 SECONDS AND 16 ATM FOR 10 SECONDS, AN ONYX FRONTIER 5.0 × 15 MM STENT WAS IMPLANTED FROM THE LMS INTO THE LAD AT 10 ATM FOR 10 SECONDS, AN ONYX FRONTIER 3.5 × 26 MM STENT WAS DEPLOYED IN THE OSTIAL LAD AT 10 ATM FOR 8 SECONDS, AND FINALLY, AN ONYX FRONTIER 4.0 × 15 MM STENT WAS DEPLOYED FROM THE LMS INTO THE OSTIAL LAD AT 16 ATM FOR 12 SECONDS. DEVICE LOT NUMBER PROVIDED MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22573773&pc=NIQ | 9612164-2025-03569 | 2024-12-10 05:00 | Death | MEDTRONIC IRELAND | 2025-07-23 04:00 | NIQ | ONYX FRONTIER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Ventricular Fibrillation; Thrombosis/Thrombus; Asystole | P160043 |  | 1 | Event Description: DURING THE INDEX PROCEDURE A TOTAL OF SEVEN ONYX FRONTIER DRUG ELUTING STENTS (DES) WERE IMPLANTED. TWO PREVAIL DRUG COATED BALLOONS (DCB) AND THREE LAUNCHER GUIDE CATHETERS WERE ALSO USED DURING THE INDEX PROCEDURE. LESIONS TREATED DURING THE INDEX PROCEDURE INCLUDED THE 1ST DIAGONAL ARTERY, THE LEFT MAIN CORONARY ARTERY (LCMA), THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY(LAD), PROXIMAL CIRCUMFLEX (CX), MID AND PROXIMAL RIGHT CORONARY ARTERY (RCA). A BARE METAL STENT WAS IMPLANTED IN THE LAD. THE LCMA WAS TREATED WITH ANGIOPLASTY. APPROXIMATELY 16 MONTHS POST THE INDEX PROCEDURE THE PATIENT SUFFERED FROM A MYOCARDIAL INFARCTION (MI). THE PATIENT COLLAPSED AT HOME WITH CARDIAC ARREST. VENTRICULAR FIBRILLATION WAS SEEN ON THE ELECTROCARDIOGRAM (ECG). INITIALLY THERE WAS ASYSTOLE. THE PATIENT HAD 14 SHOCKS. 10MG OF ADRENALINE AND 450MG OF AMIODARONE WAS ADMINISTERED. THE PATIENT WAS IN ASYSTOLE ON ARRIVAL IN THE EMERGENCY DEPARTMENT. THE PATIENT DIED. THE MI WAS CLASSIFIED AS A TYPE ONE SPONTANEOUS MI. THE LOCATION OF THE MI WAS UNKNOWN. THE DEATH WAS CLASSIFIED AS A SUDDEN CARDIAC DEATH. THE PATIENT WAS NOT TAKING DUAL ANTIPLATELET THERAPY WITHIN 24 HOURS PRIOR TO THE EVENT. THE SITES COMMENTED THAT IT WAS POSSIBLE TO BE A STENT THROMBOSIS RELATED EVENT WITH THE CARDIAC ARREST. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: NO ANGIOGRAM WAS PERFORMED. AS PER PHYSICIAN, THE MI WAS POSSIBLY RELATED TO THE ONYX FRONTIER STENTS IMPLANTED DURING THE INDEX PROCEDURE. AS PER THE PHYSICIAN, THE CAUSE OF DEATH WAS MYOCARDIAL INFARCTION AND ISCHEMIC HEART DISEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATE RECEIVED: DURING THE INDEX PROCEDURE, A PREVAIL 3.0 × 20 MM BALLOON WAS USED IN THE DISTAL RCA/PLV AT 8 ATM FOR 60 SECONDS, FOLLOWED BY DEPLOYMENT OF AN ONYX FRONTIER 3.5 × 34 MM STENT IN THE DISTAL RCA INTO THE PDA AT 10 ATM FOR 22 SECONDS, AN ONYX FRONTIER 4.0 × 30 MM STENT IN THE OSTIAL RCA AT 12 ATM FOR 24 SECONDS AND AGAIN AT 12 ATM FOR 20 SECONDS, AND A PREVAIL DEB 3.0 × 30 MM IN THE OSTIAL TO MID LEFT CIRCUMFLEX AT 8 ATM FOR 40 SECONDS AND 20 SECONDS; SUBSEQUENTLY, AN ONYX FRONTIER 3.0 × 15 MM STENT WAS DEPLOYED IN THE PROXIMAL DIAGONAL AT 12 ATM FOR 12 SECONDS, AN ONYX FRONTIER 3.0 × 12 MM STENT WAS IMPLANTED IN THE OSTIUM OF THE DIAGONAL AT 12 ATM FOR 10 SECONDS AND 16 ATM FOR 10 SECONDS, AN ONYX FRONTIER 5.0 × 15 MM STENT WAS IMPLANTED FROM THE LMS INTO THE LAD AT 10 ATM FOR 10 SECONDS, AN ONYX FRONTIER 3.5 × 26 MM STENT WAS DEPLOYED IN THE OSTIAL LAD AT 10 ATM FOR 8 SECONDS, AND FINALLY, AN ONYX FRONTIER 4.0 × 15 MM STENT WAS DEPLOYED FROM THE LMS INTO THE OSTIAL LAD AT 16 ATM FOR 12 SECONDS. DEVICE LOT NUMBER PROVIDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22573776&pc=NIQ | 9612164-2025-03570 | 2024-12-10 05:00 | Death | MEDTRONIC IRELAND | 2025-07-23 04:00 | NIQ | ONYX FRONTIER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Ventricular Fibrillation; Thrombosis/Thrombus; Asystole | P160043 |  | 1 | Event Description: DURING THE INDEX PROCEDURE A TOTAL OF SEVEN ONYX FRONTIER DRUG ELUTING STENTS (DES) WERE IMPLANTED. TWO PREVAIL DRUG COATED BALLOONS (DCB) AND THREE LAUNCHER GUIDE CATHETERS WERE ALSO USED DURING THE INDEX PROCEDURE. LESIONS TREATED DURING THE INDEX PROCEDURE INCLUDED THE 1ST DIAGONAL ARTERY, THE LEFT MAIN CORONARY ARTERY (LCMA), THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY(LAD), PROXIMAL CIRCUMFLEX (CX), MID AND PROXIMAL RIGHT CORONARY ARTERY (RCA). A BARE METAL STENT WAS IMPLANTED IN THE LAD. THE LCMA WAS TREATED WITH ANGIOPLASTY. APPROXIMATELY 16 MONTHS POST THE INDEX PROCEDURE THE PATIENT SUFFERED FROM A MYOCARDIAL INFARCTION (MI). THE PATIENT COLLAPSED AT HOME WITH CARDIAC ARREST. VENTRICULAR FIBRILLATION WAS SEEN ON THE ELECTROCARDIOGRAM (ECG). INITIALLY THERE WAS ASYSTOLE. THE PATIENT HAD 14 SHOCKS. 10MG OF ADRENALINE AND 450MG OF AMIODARONE WAS ADMINISTERED. THE PATIENT WAS IN ASYSTOLE ON ARRIVAL IN THE EMERGENCY DEPARTMENT. THE PATIENT DIED. THE MI WAS CLASSIFIED AS A TYPE ONE SPONTANEOUS MI. THE LOCATION OF THE MI WAS UNKNOWN. THE DEATH WAS CLASSIFIED AS A SUDDEN CARDIAC DEATH. THE PATIENT WAS NOT TAKING DUAL ANTIPLATELET THERAPY WITHIN 24 HOURS PRIOR TO THE EVENT. THE SITES COMMENTED THAT IT WAS POSSIBLE TO BE A STENT THROMBOSIS RELATED EVENT WITH THE CARDIAC ARREST. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: NO ANGIOGRAM WAS PERFORMED. AS PER PHYSICIAN, THE MI WAS POSSIBLY RELATED TO THE ONYX FRONTIER STENTS IMPLANTED DURING THE INDEX PRO CEDURE. AS PER THE PHYSICIAN, THE CAUSE OF DEATH WAS MYOCARDIAL INFARCTION AND ISCHEMIC HEART DISEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATE RECEIVED: DURING THE INDEX PROCEDURE, A PREVAIL 3.0 × 20 MM BALLOON WAS USED IN THE DISTAL RCA/PLV AT 8 ATM FOR 60 SECONDS, FOLLOWED BY DEPLOYMENT OF AN ONYX FRONTIER 3.5 × 34 MM STENT IN THE DISTAL RCA INTO THE PDA AT 10 ATM FOR 22 SECONDS, AN ONYX FRONTIER 4.0 × 30 MM STENT IN THE OSTIAL RCA AT 12 ATM FOR 24 SECONDS AND AGAIN AT 12 ATM FOR 20 SECONDS, AND A PREVAIL DEB 3.0 × 30 MM IN THE OSTIAL TO MID LEFT CIRCUMFLEX AT 8 ATM FOR 40 SECONDS AND 20 SECONDS; SUBSEQUENTLY, AN ONYX FRONTIER 3.0 × 15 MM STENT WAS DEPLOYED IN THE PROXIMAL DIAGONAL AT 12 ATM FOR 12 SECONDS, AN ONYX FRONTIER 3.0 × 12 MM STENT WAS IMPLANTED IN THE OSTIUM OF THE DIAGONAL AT 12 ATM FOR 10 SECONDS AND 16 ATM FOR 10 SECONDS, AN ONYX FRONTIER 5.0 × 15 MM STENT WAS IMPLANTED FROM THE LMS INTO THE LAD AT 10 ATM FOR 10 SECONDS, AN ONYX FRONTIER 3.5 × 26 MM STENT WAS DEPLOYED IN THE OSTIAL LAD AT 10 ATM FOR 8 SECONDS, AND FINALLY, AN ONYX FRONTIER 4.0 × 15 MM STENT WAS DEPLOYED FROM THE LMS INTO THE OSTIAL LAD AT 16 ATM FOR 12 SECONDS. DEVICE LOT NUMBER PROVIDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22573778&pc=NIQ | 9612164-2025-03571 | 2024-12-10 05:00 | Death | MEDTRONIC IRELAND | 2025-07-23 04:00 | NIQ | ONYX FRONTIER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Ventricular Fibrillation; Thrombosis/Thrombus; Asystole | P160043 |  | 1 | Event Description: DURING THE INDEX PROCEDURE A TOTAL OF SEVEN ONYX FRONTIER DRUG ELUTING STENTS (DES) WERE IMPLANTED. TWO PREVAIL DRUG COATED BALLOONS (DCB) AND THREE LAUNCHER GUIDE CATHETERS WERE ALSO USED DURING THE INDEX PROCEDURE. LESIONS TREATED DURING THE INDEX PROCEDURE INCLUDED THE THE 1ST DIAGONAL ARTERY, THE LEFT MAIN CORONARY ARTERY (LCMA), THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY(LAD), PROXIMAL CIRCUMFLEX (CX), MID AND PROXIMAL RIGHT CORONARY ARTERY (RCA). A BARE METAL STENT WAS IMPLANTED IN THE LAD. THE LCMA WAS TREATED WITH ANGIOPLASTY. APPROXIMATELY 16 MONTHS POST THE INDEX PROCEDURE THE PATIENT SUFFERED FROM A MYOCARDIAL INFARCTION (MI). THE PATIENT COLLAPSED AT HOME WITH CARDIAC ARREST. VENTRICULAR FIBRILLATION WAS SEEN ON THE ELECTROCARDIOGRAM (ECG). INITIALLY THERE WAS ASYSTOLE. THE PATIENT HAD 14 SHOCKS. 10MG OF ADRENALINE AND 450MG OF AMIODARONE WAS ADMINISTERED. THE PATIENT WAS IN ASYSTOLE ON ARRIVAL IN THE EMERGENCY DEPARTMENT. THE PATIENT DIED. THE MI WAS CLASSIFIED AS A TYPE ONE SPONTANEOUS MI. THE LOCATION OF THE MI WAS UNKNOWN. THE DEATH WAS CLASSIFIED AS A SUDDEN CARDIAC DEATH. THE PATIENT WAS NOT TAKING DUAL ANTIPLATELET THERAPY WITHIN 24 HOURS PRIOR TO THE EVENT. THE SITES COMMENTED THAT IT WAS POSSIBLE TO BE A STENT THROMBOSIS RELATED EVENT WITH THE CARDIAC ARREST. Manufacturer Narrative: ADDITIONAL INFORMATION: NO ANGIOGRAM WAS PERFORMED. AS PER PHYSICIAN, THE MI WAS POSSIBLY RELATED TO THE ONYX FRONTIER STENTS IMPLANTED DURING THE INDEX PROCEDURE. AS PER THE PHYSICIAN, THE CAUSE OF DEATH WAS MYOCARDIAL INFARCTION AND ISCHEMIC HEART DISEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATE RECEIVED: DURING THE INDEX PROCEDURE, A PREVAIL 3.0 × 20 MM BALLOON WAS USED IN THE DISTAL RCA/PLV AT 8 ATM FOR 60 SECONDS, FOLLOWED BY DEPLOYMENT OF AN ONYX FRONTIER 3.5 × 34 MM STENT IN THE DISTAL RCA INTO THE PDA AT 10 ATM FOR 22 SECONDS, AN ONYX FRONTIER 4.0 × 30 MM STENT IN THE OSTIAL RCA AT 12 ATM FOR 24 SECONDS AND AGAIN AT 12 ATM FOR 20 SECONDS, AND A PREVAIL DEB 3.0 × 30 MM IN THE OSTIAL TO MID LEFT CIRCUMFLEX AT 8 ATM FOR 40 SECONDS AND 20 SECONDS; SUBSEQUENTLY, AN ONYX FRONTIER 3.0 × 15 MM STENT WAS DEPLOYED IN THE PROXIMAL DIAGONAL AT 12 ATM FOR 12 SECONDS, AN ONYX FRONTIER 3.0 × 12 MM STENT WAS IMPLANTED IN THE OSTIUM OF THE DIAGONAL AT 12 ATM FOR 10 SECONDS AND 16 ATM FOR 10 SECONDS, AN ONYX FRONTIER 5.0 × 15 MM STENT WAS IMPLANTED FROM THE LMS INTO THE LAD AT 10 ATM FOR 10 SECONDS, AN ONYX FRONTIER 3.5 × 26 MM STENT WAS DEPLOYED IN THE OSTIAL LAD AT 10 ATM FOR 8 SECONDS, AND FINALLY, AN ONYX FRONTIER 4.0 × 15 MM STENT WAS DEPLOYED FROM THE LMS INTO THE OSTIAL LAD AT 16 ATM FOR 12 SECONDS. DEVICE LOT NUMBER PROVIDED MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22573787&pc=NIQ | 9612164-2025-03572 | 2024-12-10 05:00 | Death | MEDTRONIC IRELAND | 2025-07-23 04:00 | NIQ | ONYX FRONTIER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Ventricular Fibrillation; Thrombosis/Thrombus; Asystole | P160043 |  | 1 | Event Description: DURING THE INDEX PROCEDURE A TOTAL OF SEVEN ONYX FRONTIER DRUG ELUTING STENTS (DES) WERE IMPLANTED. TWO PREVAIL DRUG COATED BALLOONS (DCB) AND THREE LAUNCHER GUIDE CATHETERS WERE ALSO USED DURING THE INDEX PROCEDURE. LESIONS TREATED DURING THE INDEX PROCEDURE INCLUDED THE THE 1ST DIAGONAL ARTERY, THE LEFT MAIN CORONARY ARTERY (LCMA), THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY(LAD), PROXIMAL CIRCUMFLEX (CX), MID AND PROXIMAL RIGHT CORONARY ARTERY (RCA). A BARE METAL STENT WAS IMPLANTED IN THE LAD. THE LCMA WAS TREATED WITH ANGIOPLASTY. APPROXIMATELY 16 MONTHS POST THE INDEX PROCEDURE THE PATIENT SUFFERED FROM A MYOCARDIAL INFARCTION (MI). THE PATIENT COLLAPSED AT HOME WITH CARDIAC ARREST. VENTRICULAR FIBRILLATION WAS SEEN ON THE ELECTROCARDIOGRAM (ECG). INITIALLY THERE WAS ASYSTOLE. THE PATIENT HAD 14 SHOCKS. 10MG OF ADRENALINE AND 450MG OF AMIODARONE WAS ADMINISTERED. THE PATIENT WAS IN ASYSTOLE ON ARRIVAL IN THE EMERGENCY DEPARTMENT. THE PATIENT DIED. THE MI WAS CLASSIFIED AS A TYPE ONE SPONTANEOUS MI. THE LOCATION OF THE MI WAS UNKNOWN. THE DEATH WAS CLASSIFIED AS A SUDDEN CARDIAC DEATH. THE PATIENT WAS NOT TAKING DUAL ANTIPLATELET THERAPY WITHIN 24 HOURS PRIOR TO THE EVENT. THE SITES COMMENTED THAT IT WAS POSSIBLE TO BE A STENT THROMBOSIS RELATED EVENT WITH THE CARDIAC ARREST. Manufacturer Narrative: ADDITIONAL INFORMATION: NO ANGIOGRAM WAS PERFORMED. AS PER PHYSICIAN, THE MI WAS POSSIBLY RELATED TO THE ONYX FRONTIER STENTS IMPLANTED DURING THE INDEX PROCEDURE. AS PER THE PHYSICIAN, THE CAUSE OF DEATH WAS MYOCARDIAL INFARCTION AND ISCHEMIC HEART DISEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATE RECEIVED: DURING THE INDEX PROCEDURE, A PREVAIL 3.0 × 20 MM BALLOON WAS USED IN THE DISTAL RCA/PLV AT 8 ATM FOR 60 SECONDS, FOLLOWED BY DEPLOYMENT OF AN ONYX FRONTIER 3.5 × 34 MM STENT IN THE DISTAL RCA INTO THE PDA AT 10 ATM FOR 22 SECONDS, AN ONYX FRONTIER 4.0 × 30 MM STENT IN THE OSTIAL RCA AT 12 ATM FOR 24 SECONDS AND AGAIN AT 12 ATM FOR 20 SECONDS, AND A PREVAIL DEB 3.0 × 30 MM IN THE OSTIAL TO MID LEFT CIRCUMFLEX AT 8 ATM FOR 40 SECONDS AND 20 SECONDS; SUBSEQUENTLY, AN ONYX FRONTIER 3.0 × 15 MM STENT WAS DEPLOYED IN THE PROXIMAL DIAGONAL AT 12 ATM FOR 12 SECONDS, AN ONYX FRONTIER 3.0 × 12 MM STENT WAS IMPLANTED IN THE OSTIUM OF THE DIAGONAL AT 12 ATM FOR 10 SECONDS AND 16 ATM FOR 10 SECONDS, AN ONYX FRONTIER 5.0 × 15 MM STENT WAS IMPLANTED FROM THE LMS INTO THE LAD AT 10 ATM FOR 10 SECONDS, AN ONYX FRONTIER 3.5 × 26 MM STENT WAS DEPLOYED IN THE OSTIAL LAD AT 10 ATM FOR 8 SECONDS, AND FINALLY, AN ONYX FRONTIER 4.0 × 15 MM STENT WAS DEPLOYED FROM THE LMS INTO THE OSTIAL LAD AT 16 ATM FOR 12 SECONDS. DEVICE LOT NUMBER PROVIDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22573789&pc=NIQ | 9612164-2025-03573 | 2024-12-10 05:00 | Death | MEDTRONIC IRELAND | 2025-07-23 04:00 | NIQ | ONYX FRONTIER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Ventricular Fibrillation; Thrombosis/Thrombus; Asystole | P160043 |  | 1 | Event Description: DURING THE INDEX PROCEDURE A TOTAL OF SEVEN ONYX FRONTIER DRUG ELUTING STENTS (DES) WERE IMPLANTED. TWO PREVAIL DRUG COATED BALLOONS (DCB) AND THREE LAUNCHER GUIDE CATHETERS WERE ALSO USED DURING THE INDEX PROCEDURE. LESIONS TREATED DURING THE INDEX PROCEDURE INCLUDED THE 1ST DIAGONAL ARTERY, THE LEFT MAIN CORONARY ARTERY (LCMA), THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY(LAD), PROXIMAL CIRCUMFLEX (CX), MID AND PROXIMAL RIGHT CORONARY ARTERY (RCA). A BARE METAL STENT WAS IMPLANTED IN THE LAD. THE LCMA WAS TREATED WITH ANGIOPLASTY. APPROXIMATELY 16 MONTHS POST THE INDEX PROCEDURE THE PATIENT SUFFERED FROM A MYOCARDIAL INFARCTION (MI). THE PATIENT COLLAPSED AT HOME WITH CARDIAC ARREST. VENTRICULAR FIBRILLATION WAS SEEN ON THE ELECTROCARDIOGRAM (ECG). INITIALLY THERE WAS ASYSTOLE. THE PATIENT HAD 14 SHOCKS. 10MG OF ADRENALINE AND 450MG OF AMIODARONE WAS ADMINISTERED. THE PATIENT WAS IN ASYSTOLE ON ARRIVAL IN THE EMERGENCY DEPARTMENT. THE PATIENT DIED. THE MI WAS CLASSIFIED AS A TYPE ONE SPONTANEOUS MI. THE LOCATION OF THE MI WAS UNKNOWN. THE DEATH WAS CLASSIFIED AS A SUDDEN CARDIAC DEATH. THE PATIENT WAS NOT TAKING DUAL ANTIPLATELET THERAPY WITHIN 24 HOURS PRIOR TO THE EVENT. THE SITES COMMENTED THAT IT WAS POSSIBLE TO BE A STENT THROMBOSIS RELATED EVENT WITH THE CARDIAC ARREST. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: NO ANGIOGRAM WAS PERFORMED. AS PER PHYSICIAN, THE MI WAS POSSIBLY RELATED TO THE ONYX FRONTIER STENTS IMPLANTED DURING THE INDEX PRO CEDURE. AS PER THE PHYSICIAN, THE CAUSE OF DEATH WAS MYOCARDIAL INFARCTION AND ISCHEMIC HEART DISEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22573834&pc=NVZ | 3008973940-2025-05348 | 2025-06-09 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-07-23 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED LESS THAN ONE MONTH POST IMPLANT OF AN IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: CORRECTED: D9. PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22573838&pc=NVN | 2649622-2025-19974 | 2025-06-09 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-23 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED LESS THAN ONE MONTH POST IMPLANT OF AN IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22573841&pc=DTB | 2649622-2025-19975 | 2025-06-09 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-23 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED LESS THAN ONE MONTH POST IMPLANT OF AN IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22574072&pc=LWR | 9612164-2025-03582 | 2025-07-21 04:00 | Death | MEDTRONIC MEXICO | 2025-07-23 04:00 | LWR | AVALUS AORTIC TISSUE VALVE | Insufficient Information; Patient Device Interaction Problem | Hemorrhage/Bleeding; Unspecified Infection; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P170006 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION DURING IMPLANT OF THIS AVALUS VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE AVALUS VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PHYSICIAN DOES NOT BELIEVE THE MEDTRONIC PRODUCT CAUSED OR CONTRIBUTED TO THIS DEATH. IT WAS REPORTED THAT "POSTOPERATIVE INFECTION BLEEDING" WERE CONTRIBUTING FACTORS TO THE PATIENT DEATH. Manufacturer Narrative: B5 AND H6 UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22574106&pc=PNJ | 9612164-2025-03584 | 2025-06-29 04:00 | Death | MEDTRONIC IRELAND | 2025-07-23 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Dyspnea | P150033 |  | 1 | Event Description: IT WAS REPORTED POSTOPERATIVE THAT THE PATIENT EXPERIENCED DYSPNEA AND THEY WERE OVERHEATING. THE PATIENT PASSED AWAY TWO DAYS POST IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [(B)(6)]). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22574171&pc=MKJ | 3004123209-2025-00199 | 2018-12-03 05:00 | Death | HEARTSINE TECHNOLOGIES LTD | 2025-07-23 04:00 | MKJ | UNKNOWN_MEDICAL - BELFAST_PRODUCT | Failure to Sense | No Clinical Signs, Symptoms or Conditions | P160008 |  | 1 | Event Description: THE CUSTOMER CONTACTED HEARTSINE TO REPORT THAT THEIR DEVICE FAILED TO DETECT A PATIENT IMPEDANCE DURING A PATIENT INVOLVED EVENT. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THE PATIENT INVOLVED IN THE REPORTED EVENT IS DECEASED. Manufacturer Narrative: HEARTSINE'S INVESTIGATION DETERMINED THE ROOT CAUSE OF THE REPORTED FAULT TO BE A POTENTIAL USE ERROR. UPON RECEIPT, A REVIEW OF DEVICE HISTORY LOGS CONFIRMED THE REPORTED FAULT. THE DEVICE THEN UNDERWENT EXTENSIVE TESTING OF ALL CRITICAL FUNCTIONS, PERFORMING TO SPECIFICATION THROUGHOUT. NO FAULT WAS FOUND ON THE RETURNED DEVICE OR PAD-PAK THAT COULD LEAD TO THE REPORTED FAULT. DURING THE INVESTIGATION, THE DEVICE WAS FOUND TO BE CORRECTLY MEASURING IMPEDANCE THROUGHOUT THE RANGE EVEN UNDER THE STRESS OF ELEVATED TEMPERATURE. THEREFORE, THE INVESTIGATION CAN ONLY SUGGEST A POTENTIAL USE ERROR OF THE DEVICE. THE DEVICE HAS BEEN RETAINED BY HEARTSINE AND THE CUSTOMER WAS PROVIDED WITH A REPLACEMENT DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22574232&pc=MVK | 3008642652-2025-07071 | 2025-07-09 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-23 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED 8 APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 09:42:03 ON (B)(6) 2025. AT 22:34:48, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS TACHYCARDIA @ 140 BPM WITH PACS TRANSITIONING TO VT @ 260 BPM. AT 22:35:23, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VT @ 270 BPM. POST SHOCK RHYTHM WAS ASYSTOLE WITH MOTION ARTIFACT FOR 23 SECONDS TRANSITIONING TO SINUS BRADYCARDIA @ 40 BPM. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 22:36:27, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 22:36:58, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 22:38:54, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 22:39:25, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 20 BPM. AT 22:39:58, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BEAT TRANSITIONING TO VT @ 270 BPM. AT 22:40:31, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VT @ 260 BPM. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM TRANSITIONING TO VT @ 270 BPM. AT 22:40:57, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VT @ 270 BPM. POST SHOCK RHYTHM WAS SINUS BEAT TRANSITIONING TO VT @ 270 BPM. AT 22:41:18, THE PATIENT RECEIVED THE SEVENTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VT @ 270 BPM. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA @ 100 BPM TRANSITIONING TO VT @ 270 BPM DEGRADING TO VF / FINE VF. AT 22:53:45, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT/VF. AT 22:54:16, THE PATIENT RECEIVED THE EIGHTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH TACTILE ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE DEVICE WAS SHUT DOWN AT 22:55:43 ON (B)(6) 2025. Manufacturer Narrative: THE MONITOR AND ELECTRODE BELT HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION. THE REVIEW OF THE DATA INDICATED THAT THE DEVICE POWERED ON NORMALLY AND WAS ABLE TO ACQUIRE THE PATIENT¿S ECG SIGNAL ON THE LAST DAY OF USE CAPTURED IN THE DATA DOWNLOAD. NO DEFICIENCIES ALLEGED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22574241&pc=MVK | 3008642652-2025-07072 | 2025-06-13 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-23 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 15:06:22 ON (B)(6) 2025. AT 12:36:49 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY DEGRADING TO VF. AT 12:38:00, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE DEVICE WAS SHUT DOWN AT 12:41:38 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22574249&pc=OZD | 1220648-2025-30603 | 2025-07-05 04:00 | Death | ABIOMED, INC. | 2025-07-23 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT ON IMPELLA CP SUPPORT EXPERIENCED A LARGE STROKE AND CARE WAS WITHDRAWN. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22574255&pc=GEI | 1717344-2025-00942 | 2025-03-30 04:00 | Death | COVIDIEN | 2025-07-23 04:00 | GEI | UNKNOWN LIGASURE INSTRUMENT | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY INCLUDING 178 CASES OF LIVER MALIGNANCIES REQUIRING LAPAROSCOPIC HEPATECTOMY USING EITHER WHITE LIGHT OR INDOCYANINE GREEN FLUORESCENCE GUIDANCE WAS COMPLETED BETWEEN APRIL 2018 AND APRIL 2023. LIVER PARENCHYMAL TRANSECTION WAS PERFORMED USING LIGASURE OR A COMPETITOR DEVICE. DEATH DURING HOSPITALIZATION OCCURRED IN EIGHT CASES WITH REASONS UNSPECIFIED. Manufacturer Narrative: D10 CONCOMITANT PRODUCTS: UNKNOWN LIGASUR, UNKNOWN LIGASURE INSTRUMENT (LOT#: UNKNOWN), UNKNOWN LIGASUR, UNKNOWN LIGASURE INSTRUMENT (LOT#: UNKNOWN) TIANYANG MAO. "INDOCYANINE GREEN FUORESCENCE IMAGING (ICG-FI) IN DIFCULT LAPAROSCOPIC HEPATECTOMY FOR HEPATOCELLULAR CARCINOMA: A RETROSPECTIVE PROPENSITY SCORE¿MATCHED ANALYSIS". 2025. HTTPS://DOI.ORG/10.1007/S00464-025-11707-3 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22574421&pc=MKJ | 0003015876-2025-01479 | 2025-06-27 04:00 | Death | PHYSIO-CONTROL, INC. - 3015876 | 2025-07-23 04:00 | MKJ | LIFEPAK 15 MONITOR/DEFIBRILLATOR | Device Sensing Problem | Insufficient Information | P160026 |  | 1 | Event Description: THE CUSTOMER CONTACTED STRYKER TO REPORT THAT THEIR DEVICE DID NOT DETECT THE PADDLES. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF IT WAS NEEDED. THE PATIENT INVOLVED WITH THE REPORTED EVENT DID NOT SURVIVE. Manufacturer Narrative: STRYKER CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. STRYKER REQUESTS PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. THE CUSTOMER INFORMED STRYKER THAT NO FURTHER PATIENT INFORMATION IS AVAILABLE. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK. STRYKER WILL PERFORM A CLINICAL REVIEW OF THIS EVENT TO DETERMINE DEVICE CONTRIBUTION TO THE PATIENT'S REPORTED OUTCOME. STRYKER CONTINUES TO INVESTIGATE THE REPORTED FAILURE AND WILL SUBMIT A SUPPLEMENTAL REPORT ON THIS EVENT TO THE FDA AS PROVIDED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22574549&pc=OZD | 1220648-2025-30606 | 2025-07-07 04:00 | Death | ABIOMED, INC. | 2025-07-23 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 AND EXHIBITED NEUROLOGICAL CHANGES THAT PROMPTED A COMPUTED TOMOGRAPHY (CT) TO BE PERFORMED. THE CT CONFIRMED AN IRREVERSIBLE BRAIN BLEED. THE PATIENT'S FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)." SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: "TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22574701&pc=OZD | 1220648-2025-30608 | 2025-06-24 04:00 | Death | ABIOMED, INC. | 2025-07-23 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED A GASTROINTESTINAL BLEED. HEPARIN WAS WITHHELD. LATER, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE ROOT CAUSE OF VASCULAR DAMAGE PERIPHERAL VESSEL CANNOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22574856&pc=OZD | 1220648-2025-30610 | 2025-01-06 05:00 | Death | ABIOMED, INC. | 2025-07-23 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Device Sensing Problem | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THE PATIENT WAS ON IMPELLA 5.5 SUPPORT. WHILE ON SUPPORT, PLACEMENT SIGNAL UNKNOWN ALARMS. ECHO PERFORMED SHOWED PROPER PLACEMENT, MID LEFT VENTRICLE POINTED TOWARDS APEX, INLET MEASURING 5.5CM FROM ATRIOVENTRICULAR (AV). PHYSICIAN DID NOT FEEL THE NEED TO REPOSITION OR MAKE ANY CHANGES. THE AUTOMATED IMPELLA CONTROLLER (AIC) WAS RETURNED FOR INVESTIGATION. THIS COMPLAINT WAS CREATED FOR THE AIC. EVENTUALLY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN RECEIVED BUT HAS NOT YET BEEN EVALUATED. WHEN THE DEVICE INVESTIGATION HAS BEEN COMPLETED OR ANY NEW INFORMATION HAS BEEN RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22574894&pc=DTE | 2182208-2025-03653 | 2025-06-11 04:00 | Death | MEDTRONIC, INC. | 2025-07-23 04:00 | DTE | DUAL CHAMBER TEMPORARY PACEMAKER | Use of Device Problem | Arrhythmia; Cardiac Arrest; Ventricular Fibrillation | K150246 |  | 1 | Event Description: IT WAS REPORTED THAT WHILE THE EXTERNAL PULSE GENERATOR (EPG) WAS IN USE, THE PATIENT EXPERIENCED AN R WAVE ON T WAVE EVENT, WHICH PRECIPITATED A VENTRICULAR FIBRILLATION (VF) ARREST. THE PATIENT WAS RESUSCITATED BUT ASPIRATED DURING RE-INTUBATION. IT WAS NOTED THAT HE SUBSEQUENTLY DEVELOPED ADDITIONAL COMPLICATIONS AND, UNFORTUNATELY, PASSED AWAY ON POSTOPERATIVE DAY SIX. IT WAS ALSO NOTED THAT THE PATIENT HAD BEEN ALERT, EXTUBATED, AND AMBULATORY ON POSTOPERATIVE DAY ONE FOLLOWING OPEN-HEART SURGERY, WHICH INCLUDED AN AO RTIC VALVE REPLACEMENT AND ANEURYSM REPAIR. NO FURTHER INFORMATION WAS AVAILABLE. Manufacturer Narrative: MW5171785. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22575108&pc=OZD | 1220648-2025-30612 | 2025-07-06 04:00 | Death | ABIOMED, INC. | 2025-07-23 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT IMPELLA CP WAS PLACED TO SUPPORT PERCUTANEOUS CORONARY INTERVENTION PROCEDURE. DURING PERCUTANEOUS CORONARY INTERVENTION IT WAS NOTED THAT THE PATIENT HAD A HEMATOMA AROUND 14FR SHORT PEEL AWAY. SHEATH WAS REMOVED AND REPO PLACED. AFTER THERE WAS AN ARTERIAL BLEEDING AROUND THE SHEATH. PHYSICIAN PUT LOTS OF FORWARD TENSION ON SHEATH AND HELD MANUAL PRESSURE. THE PHYSICIAN ALSO WENT UP AND OVER WITH A BALLOON TO HELP CONTROL BLEEDING. THE VASCULAR SURGEON WAS CONSULTED AND DID A CUT DOWN. AFTER WORKING FOR AROUND 2-3 HOURS, THE SURGEON WAS ABLE TO GET BLEEDING UNDER CONTROL. DURING SURGERY, THE PATIENT RECEIVED 15 UNITS PACKED RED BLOOD CELLS , 15 UNITS PLASMA, AND 3 UNITS PLATELETS. THE POCKET WAS PARTIALLY CLOSED AND THE BLUE IMPELLA HUB WAS SUTURED DOWN. THROUGHOUT THE PROCEDURE THE PHYSICIAN ALSO HELD PRESSURE ON THE OPPOSITE GROIN BECAUSE THERE WAS A HEMATOMA ON THAT SIDE AS WELL. AT THE END OF THE PROCEDURE, THE PATIENT¿S PRESSOR REQUIREMENTS WERE DECREASING. THE PATIENT WAS TRANSFERRED TO INTENSIVE CARE UNIT. OVER THE NEXT COUPLE OF HOURS, THE PATIENT CODED A FEW TIMES, AND FAMILY MADE THE DECISION TO NOT RESUSCITATE THE PATIENT IF CODED. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA CP WITH SMART ASSIST, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22575589&pc=DQA | MW5173096 | 2024-12-01 05:00 | Death | NELLCOR / COVIDIEN | 2025-07-23 04:00 | DQA | COVIDIEN NELLCOR PM100N BEDSIDE PULSE OXIMETER | Delayed Alarm; Device Alarm System | Bradycardia; Cardiac Arrest; Low Oxygen Saturation; Respiratory Arrest |  |  | 1 | Event Description: I AM WRITING TO FORMALLY REPORT A CONCERN REGARDING THE COVIDIEN NELLCOR PM100N BEDSIDE PULSE OXIMETER, WHICH WE BELIEVE MAY HAVE CONTRIBUTED TO THE DEATH OF OUR DAUGHTER DUE TO A POTENTIALLY DELAYED ALARM NOTIFICATION. OUR DAUGHTER PASSED AWAY ON (B)(6) 2024, AT THE AGE OF FOUR MONTHS. SHE WAS LIVING WITH TRISOMY 13 AND WAS RECEIVING FULL-TIME CARE AT HOME, SUPPORTED BY VARIOUS MEDICAL EQUIPMENT, INCLUDING THE PM100N DEVICE, WHICH WAS MONITORING HER OXYGEN SATURATION AND PULSE RATE. THE DEVICE WAS CONFIGURED TO ALERT US WITH PRIORITY ALARMS IN THE EVENT OF OXYGEN DESATURATION OR BRADYCARDIA. ON JULY 18, 2025, WE RECEIVED A NOTIFICATION REFERENCING A DELAYED PRIORITY ALARM ISSUE WITH THIS DEVICE. THIS NEW INFORMATION RAISES SIGNIFICANT CONCERNS FOR US. IT IS NOW A PAINFUL POSSIBILITY THAT THE ALARM DELAY MAY HAVE AFFECTED OUR ABILITY TO RESPOND QUICKLY DURING A CRITICAL DESATURATION EVENT. HAD THE ALARM SOUNDED PROMPTLY, WE MIGHT HAVE BEEN ABLE TO STIMULATE AND RESUSCITATE HER BEFORE SHE REACHED A STATE FROM WHICH SHE COULD NOT RECOVER. OUR DAUGHTER HAD FREQUENT DESATURATIONS AND BRADYCARDIA EPISODES DURING HER HOSPITAL STAYS, WHICH WERE SUCCESSFULLY ADDRESSED WITH IMMEDIATE RESPONSE AND STIMULATION. WE HAD BEEN TRAINED AND WERE EXPERIENCED IN REACTING TO THESE EVENTS. WE ARE SEEKING ANSWERS TO THE FOLLOWING: WAS THIS A KNOWN ISSUE AT THE TIME OF USE, AND IF SO, WAS THERE A DELAY IN ISSUING THE RECALL OR NOTIFICATION? IS THIS A PERVASIVE ISSUE ACROSS OTHER UNITS OR SETTINGS? COULD THIS DELAY IN ALARM FUNCTIONALITY HAVE REASONABLY CONTRIBUTED TO OR CAUSED THE LACK OF TIMELY INTERVENTION AND ULTIMATELY, HER DEATH? WE RESPECTFULLY REQUEST A THOROUGH INVESTIGATION INTO THIS MATTER, INCLUDING ANY AVAILABLE FINDINGS RELATED TO THIS DEFECT AND ANY STEPS THAT ARE BEING TAKEN TO ADDRESS IT. WE WOULD ALSO APPRECIATE INFORMATION ON HOW OTHER AFFECTED FAMILIES HAVE BEEN NOTIFIED AND SUPPORTED. THANK YOU FOR YOUR ATTENTION TO THIS SERIOUS CONCERN. WE ARE GRIEVING PARENTS TRYING TO UNDERSTAND WHAT MAY HAVE HAPPENED AND HOW TO ENSURE THAT OTHER FAMILIES ARE NOT SIMILARLY AFFECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22576408&pc=CAW | 2515872-2025-00024 | 2025-06-21 04:00 | Death | DEVILBISS HEALTHCARE | 2025-07-23 04:00 | CAW | DRIVE DEVILBISS 10L OXYGEN CONCENTRATOR | Device Emits Odor; Powder Leak | Aspiration Pneumonitis | K071397 |  | 1 | Event Description: DEVILBISS HEALTHCARE WAS NOTIFIED OF AN INCIDENT INVOLVING AN OXYGEN CONCENTRATOR BY AN ATTORNEY, WHO ALLEGED THAT THE DEVICE EXPLODED AND RELEASED A YELLOWISH POWDER AND AN INTENSE NOXIOUS ODOR. THE END USER ALLEGEDLY SUFFERED IMMEDIATE RESPIRATORY DISTRESS AND DIED SHORTLY AFTER. A REVIEW OF OUR RECORDS INDICATED THAT A DIFFERENT END USER REPORTED THE SAME DEVICE AS ALARMING A MONTH PRIOR TO THIS EVENT, AS A RESULT OF WHICH IT WAS RETURNED TO THE DME PROVIDER FOR SERVICE. IT IS UNCLEAR AT THIS TIME WHAT SERVICE WAS PERFORMED ON THE DEVICE AT THAT TIME. DEVILBISS HEALTHCARE IS CURRENTLY INVESTIGATING THE INCIDENT, INCLUDING ARRANGING AN INSPECTION OF THE DEVICE THROUGH THE ATTORNEY. AN UPDATE WILL BE FILED IF ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22576651&pc=OZD | 1220648-2025-30628 | 2025-07-09 04:00 | Death | ABIOMED, INC. | 2025-07-23 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT DEVELOPED HEMOLYSIS. THE DECISION WAS MADE TO WEAN THE PUMP. DURING THE WEANING, THE PLASMA FREE HEMOGLOBIN WAS TRENDING DOWN. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22576793&pc=NKM | 2015691-2025-05978 | 2025-07-14 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-23 04:00 | NKM | EDWARDS PASCAL PRECISION | Difficult to Insert | Cardiac Perforation | P220003 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION OF A PASCAL IN MITRAL PROCEDURE WHERE AN IMPLANT WAS ATTEMPTED AND ABORTED. A PATIENT INJURY (DEATH) OCCURRED RELATED TO THE DEVICE AND PROCEDURAL COMPLICATION. TSP (TRANSSEPTAL PUNCTURE) WAS ATTEMPTED 3 TIMES. THE WIRE VERY EASILY SLIPPED ACROSS THE IAS (INTERATRIAL SEPTUM). TSP WAS REATTEMPTED DUE TO HEIGHT OF THE FIRST TSP. CLINICAL SPECIALIST MENTIONED THAT THE WIRE SLIPPED ACROSS THE IAS LIKE BUTTER YET THEY DID NOT BELIEVE A PERFORATION WAS CREATED DURING THAT PROCEDURAL STEP AS NO EFFUSION WAS SEEN AT THAT TIME. 1 ACE WAS PLACED MEDIAL A2P2. THE SECOND DEVICE WAS INSERTED, AND UPON EXITING THE GUIDE AND CLOSING, IT WAS NOTED ON FLUORO THAT THE IMPLANT APPEARED TO HAVE TENSION. IC WAS EXITING GS (GUIDE SHEATH) AND CLOSING IMPLANT WITHOUT ECHO GUIDANCE AND WAS LOOKING ONLY ON FLUORO. WHILE ON FLUORO, IC CLOSED THE DEVICE AND STARTED TO FLEX DOWN, AND WHEN ECHO WAS VISUALIZED AGAIN, EC (ECHOCARDIOGRAPHER) INFORMED TEAM THAT THE IMPLANT WAS OUTSIDE OF THE PERICARDIUM AND CARDIAC TAMPONADE WAS NOTED. THE TEAM BEGAN CPR, AND IC PROCEEDED TO PERFORM A PERICARDIAL TAP WITH COPIOUS AMOUNTS OF VERY THIN BLOOD. SURGICAL TEAM TOOK OVER AT WHICH POINT THE DEVICE WAS REMOVED. THE PATIENT RECEIVED 8-10 UNITS OF BLOOD OR MORE. IC ADMITTED FAULT AND SHOULD HAVE WAITED FOR IMAGING PRIOR TO MANEUVERING THE IMPLANT. ALSO OF NOTE, PATIENT XARELTO HAD NOT BEEN HELD POTENTIALLY LEADING TO WORSENING BLEEDING UPON PERFORATION OF ATRIUM. THE PATIENT ULTIMATELY DIED AS THE BLEEDING WAS SIGNIFICANT AND UNABLE TO BE CONTROLLED. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22577282&pc=DYB | 1220648-2025-30633 | 2025-07-04 04:00 | Death | ABIOMED, INC. | 2025-07-23 04:00 | DYB | KIT, 23FR INTRODUCER, 11CM, STERILE | Defective Device | Cardiovascular Insufficiency | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE IMPELLA RP FLEX SHEATH SIDE PORT WAS BLOCKED AND COULD NOT FLUSH THE SHEATH. A NEW INTRODUCER KIT WAS SUBSEQUENTLY OPENED FOR INSERTION. IT WAS NOTED THAT THE IMPELLA RP FLEX WAS IMPLANTED AS A BAILOUT WHEN THE PATIENT ARRESTED DURING A THROMBECTOMY. THE PATIENT DID NOT RECOVER AND PASSED AWAY ON SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INTRODUCER WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: H6 HEALTH EFFECT - IMPACT CODE 1802 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30633. Manufacturer Narrative: THE INVESTIGATION OF INTRODUCER DAMAGE ¿ MECHANICAL HAS BEEN COMPLETED. INTRODUCER DAMAGE - MECHANICAL: CLINICAL DETAILS NOTE THAT THE RP FLEX SHEATH COULD NOT BE FLUSHED. A NEW INTRODUCER KIT WAS OPENED. THE PRODUCT WAS RETURNED AND EVALUATED. THE RETURNED SHEATH COULD NOT BE FLUSHED. A GLUE BLOCKAGE WAS FOUND IN THE SHEATH SIDEARM AT THE CONNECTION POINT BETWEEN THE SIDEARM AND WHITE SHEATH HUB. DATA LOGS ARE NOT APPLICABLE. THE ROOT CAUSE OF THE INTRODUCER DAMAGE WAS A BOND FAILURE AS EVIDENCED BY ISSUE REPRODUCTION ON RETURNED PRODUCT AND CONFIRMATION OF GLUE BLOCKAGE AT BONDING POINT BETWEEN HUB AND TUBING. H6 HEALTH EFFECT - IMPACT CODE 1802 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30633.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22577345&pc=BYI | 2518422-2025-108150 | 2025-07-02 04:00 | Death | RESPIRONICS, INC. | 2025-07-23 04:00 | BYI | COUGH ASSIST E70 | No Display/Image | No Clinical Signs, Symptoms or Conditions | K122111 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A CA E70 "DISPLAY WAS NOT WORKING". THE DEVICE WAS SENT TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. THE CUSTOMER COMPLAINT WAS NOT REPRODUCED. THE DEVICE TO BE SENT BACK WITHOUT EVALUATION BASED ON THE CUSTOMER'S DECISION AS THE END CUSTOMER DIED IN THE MEANTIME. A CLINICAL ASSESSMENT DETERMINED BASED ON AVAILABLE INFORMATION AT THIS TIME, THE ALLEGED HARM DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. A GOOD FAITH EFFORT ATTEMPT WILL BE MADE TO DETERMINE THE CAUSE OF DEATH. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22577476&pc=NGV | 2124215-2025-48418 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR UNEXPECTED MEDICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 651 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22577539&pc=MKJ | 1220908-2025-02751 | 2025-07-08 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-07-23 04:00 | MKJ | X SERIES | Incorrect Interpretation of Signal | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE GAVE A "NO SHOCK ADVISED" PROMPT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22577754&pc=NGV | 2124215-2025-48421 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Thromboembolism |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR THROMBOEMBOLISM. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 651 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22578053&pc=NGV | 2124215-2025-48497 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR UNEXPECTED MEDICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 559 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22578055&pc=NGV | 2124215-2025-47578 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR HOSPITALIZATION OR PROLONGED HOSPITALIZATION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 175 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22578079&pc=NGV | 2124215-2025-48500 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Transient Ischemic Attack |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR TRANSIENT ISCHEMIC ATTACK (TIA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 364 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22578101&pc=NGV | 2124215-2025-47774 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22578116&pc=NGV | 2124215-2025-48419 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 707 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22578131&pc=NGV | 2124215-2025-48499 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR A VERY LATE DEVICE THROMBOSIS (MORE THAN 1 YEAR POST IMPLANT). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 559 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22578236&pc=OZD | 1220648-2025-30640 | 2025-07-08 04:00 | Death | ABIOMED, INC. | 2025-07-23 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT HAD A THREE VESSEL CORONARY ARTERY BYPASS GRAFT EARLIER IN THE YEAR. HOWEVER, THEY DID NOT FOLLOW UP WITH CARDIOLOGY OR CONTINUE MEDICATIONS. THEY WERE PRESENTED TO THE HOSPITAL WITH SHORTNESS OF BREATH, PULMONARY CONGESTION, AND HAD A CARDIAC ARREST ON THE FLOOR AND WAS TAKEN TO THE OPERATING ROOM IN CARDIOGENIC SHOCK. AN IMPELLA 5.5 WAS PLACED. THROUGHOUT IMPELLA 5.5 SUPPORT, THE PATIENT EXPERIENCED MULTIPLE EPISODES OF VENTRICULAR TACHYCARDIA AND WAS DEFIBRILLATED SEVERAL TIMES. AMIODARONE WAS STARTED AND THE PUMP WAS REPOSITIONED. HOWEVER, THE PATIENT CONTINUED TO HAVE VENTRICULAR TACHYCARDIA RUNS REQUIRING DEFIBRILLATION. IT WAS NOTED THAT THE PATIENT WAS NOT WAKING UP AND COMPUTED TOMOGRAPHY WAS NEGATIVE. THE PATIENT CONTINUED TO BE JAUNDICE, VENTILATED AND REQUIRING PRESSORS. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22578297&pc=NGV | 2124215-2025-48498 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Thromboembolism |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR THROMBOEMBOLISM. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 559 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22578356&pc=NGV | 2124215-2025-47576 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 201 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22578939&pc=DTB | 3008973940-2025-05376 | 2025-07-17 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-07-23 04:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Asystole | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DURING THE IMPLANT PROCEDURE. THE PATIENT HAD A HISTORY OF DILATED CARDIOMYOPATHY, ATRI OVENTRICULAR BLOCK WITH HYPERKALEMIA, LEFT BUNDLE BRANCH BLOCK, LOW LEFT VENTRICULAR EJECTION FRACTION (35-40%), LEFT BUNDLE BRANCH BLOCK (LBBB), COMPLETE ATRIOVENTRICULAR BLOCK (CAVB) WITH INDICATION FOR CRTD IMPLANT, AND ACUTE KIDNEY INJURY ON CHRONIC KIDNEY DISEASE WITH CORRECTED HYPERKALEMIA. EXPERIENCED MULTIPLE COMPLICATIONS. DURING THE IMPLANT PROCEDURE ISUPREL WAS ADMINISTERED. THE PATIENT HAD A COMPLETE ATRIOVENTRICULAR BLOCK WITH VENTRICULAR ESCAPE AT 59 BPM. VENOUS ACCESS WITH NO ISSUES TO REPORT. THE RV LEAD WAS IMPLANTED. IT WAS NOTED THAT THE FIRST ATTEMPT TO IMPLANT THE RIGHT VENTRICULAR LEAD, HIGH IMPEDANCE (>1400 OHMS) WAS OBSERVED, WITH A THRESHOLD OF 1V AT 0.4 MS AND R WAVES BETWEEN 3 AND 7 MV. THE PHYSICIAN REPOSITIONED THE RIGHT VENTRICULAR LEAD, RESULTING IN IMPEDANCE AROUND 600 OHMS, 0.75V AT 0.4 MS, AND 11 MV R WAVE AMPLITUDE. THE RIGHT ATRIAL LEAD WAS IMPLANTED WITH ALL PARAMETERS WITHIN RANGE. CORONARY SINUS CANNULATION AND VENOGRAM WERE PERFORMED WITHOUT ISSUE. THE LEFT VENTRICULAR LEAD IMPLANT COULD NOT BE PURSUED DUE TO PATIENT AGITATION, AND DESATURATION WAS NOTED, LEADING TO INTERRUPTION OF THE PROCEDURE. INTUBATION WAS PERFORMED, AFTER WHICH A BLOOD PRESSURE DROP OCCURRED, REQUIRING CARDIAC MASSAGE. TAMPONADE FROM PERFORATION WAS IDENTIFIED AND DRAINED, WITH THE PRESENCE OF CLOTS CONFIRMED VIA ECHOGRAPHY. DESPITE RESOLVING THE EFFUSION, ASYSTOLE OCCURRED, AND RESUSCITATION MANEUVERS WERE STOPPED AFTER 45 MINUTES.. Manufacturer Narrative: CONTINUATION OF D10: 6935M62 LEAD IMPLANT DATE: (B)(6) 2025, EXPLANT DATE: (B)(6) 2025 ; 479888 LEAD IMPLANT DATE: (B)(6) 2025, EXPLANT DATE: (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22579043&pc=OZD | 1220648-2025-30643 | 2025-07-08 04:00 | Death | ABIOMED, INC. | 2025-07-23 04:00 | OZD | IMPELLA CP | Positioning Failure | Insufficient Information | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PUMP GOT CAUGHT IN THE TRABECULAR MESHWORK AND WAS THEN UNDER TENSION. THE PUMP WAS EXPLANTED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22579061&pc=LWS | 3008973940-2025-05375 | 2025-07-17 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-07-23 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | High impedance; Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion; Thrombosis/Thrombus; Asystole | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DURING THE IMPLANT PROCEDURE. THE PATIENT HAD A HISTORY OF DILATED CARDIOMYOPATHY, ATRI OVENTRICULAR BLOCK WITH HYPERKALEMIA, LEFT BUNDLE BRANCH BLOCK, LOW LEFT VENTRICULAR EJECTION FRACTION (35-40%), LEFT BUNDLE BRANCH BLOCK (LBBB), COMPLETE ATRIOVENTRICULAR BLOCK (CAVB) WITH INDICATION FOR CRTD IMPLANT, AND ACUTE KIDNEY INJURY ON CHRONIC KIDNEY DISEASE WITH CORRECTED HYPERKALEMIA. EXPERIENCED MULTIPLE COMPLICATIONS. DURING THE IMPLANT PROCEDURE ISUPREL WAS ADMINISTERED. THE PATIENT HAD A COMPLETE ATRIOVENTRICULAR BLOCK WITH VENTRICULAR ESCAPE AT 59 BPM. VENOUS ACCESS WITH NO ISSUES TO REPORT. THE RV LEAD WAS IMPLANTED. IT WAS NOTED THAT THE FIRST ATTEMPT TO IMPLANT THE RIGHT VENTRICULAR LEAD, HIGH IMPEDANCE (>1400 OHMS) WAS OBSERVED, WITH A THRESHOLD OF 1V AT 0.4 MS AND R WAVES BETWEEN 3 AND 7 MV. THE PHYSICIAN REPOSITIONED THE RIGHT VENTRICULAR LEAD, RESULTING IN IMPEDANCE AROUND 600 OHMS, 0.75V AT 0.4 MS, AND 11 MV R WAVE AMPLITUDE. THE RIGHT ATRIAL LEAD WAS IMPLANTED WITH ALL PARAMETERS WITHIN RANGE. CORONARY SINUS CANNULATION AND VENOGRAM WERE PERFORMED WITHOUT ISSUE. THE LEFT VENTRICULAR LEAD IMPLANT COULD NOT BE PURSUED DUE TO PATIENT AGITATION, AND DESATURATION WAS NOTED, LEADING TO INTERRUPTION OF THE PROCEDURE. INTUBATION WAS PERFORMED, AFTER WHICH A BLOOD PRESSURE DROP OCCURRED, REQUIRING CARDIAC MASSAGE. TAMPONADE FROM PERFORATION WAS IDENTIFIED AND DRAINED, WITH THE PRESENCE OF CLOTS CONFIRMED VIA ECHOGRAPHY. DESPITE RESOLVING THE EFFUSION, ASYSTOLE OCCURRED, AND RESUSCITATION MANEUVERS WERE STOPPED AFTER 45 MINUTES.. Manufacturer Narrative: CONTINUATION OF D10: 407652 LEAD IMPLANT DATE: (B)(6) 2025 EXPLANT DATE: (B)(6) 2025; 479888 LEAD IMPLANT DATE: (B)(6) 2025 EXPLANT DATE: (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22579181&pc=OJX | 2649622-2025-20096 | 2025-07-17 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-23 04:00 | OJX | ATTAIN STABILITY QUAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Asystole | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DURING THE IMPLANT PROCEDURE. THE PATIENT HAD A HISTORY OF DILATED CARDIOMYOPATHY, ATRI OVENTRICULAR BLOCK WITH HYPERKALEMIA, LEFT BUNDLE BRANCH BLOCK, LOW LEFT VENTRICULAR EJECTION FRACTION (35-40%), LEFT BUNDLE BRANCH BLOCK (LBBB), COMPLETE ATRIOVENTRICULAR BLOCK (CAVB) WITH INDICATION FOR CRTD IMPLANT, AND ACUTE KIDNEY INJURY ON CHRONIC KIDNEY DISEASE WITH CORRECTED HYPERKALEMIA. EXPERIENCED MULTIPLE COMPLICATIONS. DURING THE IMPLANT PROCEDURE ISUPREL WAS ADMINISTERED. THE PATIENT HAD A COMPLETE ATRIOVENTRICULAR BLOCK WITH VENTRICULAR ESCAPE AT 59 BPM. VENOUS ACCESS WITH NO ISSUES TO REPORT. THE RV LEAD WAS IMPLANTED. IT WAS NOTED THAT THE FIRST ATTEMPT TO IMPLANT THE RIGHT VENTRICULAR LEAD, HIGH IMPEDANCE (>1400 OHMS) WAS OBSERVED, WITH A THRESHOLD OF 1V AT 0.4 MS AND R WAVES BETWEEN 3 AND 7 MV. THE PHYSICIAN REPOSITIONED THE RIGHT VENTRICULAR LEAD, RESULTING IN IMPEDANCE AROUND 600 OHMS, 0.75V AT 0.4 MS, AND 11 MV R WAVE AMPLITUDE. THE RIGHT ATRIAL LEAD WAS IMPLANTED WITH ALL PARAMETERS WITHIN RANGE. CORONARY SINUS CANNULATION AND VENOGRAM WERE PERFORMED WITHOUT ISSUE. THE LEFT VENTRICULAR LEAD IMPLANT COULD NOT BE PURSUED DUE TO PATIENT AGITATION, AND DESATURATION WAS NOTED, LEADING TO INTERRUPTION OF THE PROCEDURE. INTUBATION WAS PERFORMED, AFTER WHICH A BLOOD PRESSURE DROP OCCURRED, REQUIRING CARDIAC MASSAGE. TAMPONADE FROM PERFORATION WAS IDENTIFIED AND DRAINED, WITH THE PRESENCE OF CLOTS CONFIRMED VIA ECHOGRAPHY. DESPITE RESOLVING THE EFFUSION, ASYSTOLE OCCURRED, AND RESUSCITATION MANEUVERS WERE STOPPED AFTER 45 MINUTES.. Manufacturer Narrative: CONTINUATION OF D10: 6935M62 LEAD, IMPLANT DATE: (B)(6) 2025, EXPLANT DATE: (B)(6) 2025, 407652 LEAD IMPLANT DATE: ON (B)(6) 2025, EXPLANT DATE: (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22579505&pc=DSQ | 2916596-2025-04947 | 2025-07-11 04:00 | Death | THORATEC CORPORATION | 2025-07-23 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD END-STAGE RENAL FAILURE (ESRF) AND DID NOT WANT DIALYSIS. THE PATIENT WAS PLACED IN PALLIATIVE CARE AND SUBSEQUENTLY PASSED AWAY.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE ESRF WAS A PROGRESSION OF THE PATIENT'S CHRONIC KIDNEY DISEASE AND WAS NOT DEVICE RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22579588&pc=NGV | 2124215-2025-47814 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 6 | Event Description: THIS REPORT SUMMARIZES 6 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 71 TO 95 YEARS. PATIENT SEX: 2 FEMALES; 4 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 447 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22579619&pc=NGV | 2124215-2025-47850 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Transient Ischemic Attack |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR TRANSIENT ISCHEMIC ATTACK (TIA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 339 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22579676&pc=NGV | 2124215-2025-47584 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 4 | Event Description: THIS REPORT SUMMARIZES 4 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 48 TO 83 YEARS. PATIENT SEX: 2 FEMALES; 2 MALES. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 483 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22579782&pc=NGV | 2124215-2025-47776 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 6 | Event Description: THIS REPORT SUMMARIZES 6 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 76 TO 93 YEARS PATIENT SEX: 2 FEMALE; 4 MALE TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025 AVERAGE TIME TO EVENT: 436 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22579864&pc=MHX | 1218950-2025-000417 | 2025-07-17 04:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-07-23 04:00 | MHX | INTELLIVUE MX40 WLAN | Defective Alarm | No Clinical Signs, Symptoms or Conditions | K113125 |  | 1 | Event Description: IT WAS REPORTED THE DEVICE DID NOT GENERATE AN ALARM FOR A DESATURATION AND THE PATIENT PASSED AWAY. IT WAS INDICATED THE ALARM MAY HAVE BEEN MISSED OR NOT HEARD BY THE STAFF AND THE PATIENT PASSED WAY AT APPROXIMATELY 9:42AM. BIOMED PERFORMED SIMULATED TESTING ON THE MX40 TELEMETRY DEVICE, CONFIRMING THE ALARM FUNCTIONED AND APPEARED ON BOTH THE MX40 AND THE PIC IX CENTRAL STATION. BIOMED ALSO REVIEWED THE CLINICAL AUDIT LOGS, REVEALING A DESATURATION ALARM WAS GENERATED AT 9:20AM FOR AN OXYGEN SATURATION OF 75% AND AT 9:40AM FOR AN OXYGEN SATURATION OF 74%. Manufacturer Narrative: A PHILIPS PRODUCT SUPPORT ENGINEER (PSE WAS PROVIDED SYSTEM LOGS FOR REVIEW OF THIS ISSUE. THE PSE FOUND THE FOLLOWING IN THE 27 MINUTE TIME PERIOD FROM 9:15 UNTIL 9:42: 97 SPO2 INOPERATIVE MESSAGES (INOP TECHNICAL ALARMS) OCCURRED, INCLUDING: SPO2T SENSOR OFF (87), SPO2T NO PULSE (6), SPO2T NO SENSOR (2), SPO2T NOISY SIGNAL (2). THESE INOPERATIVE CONDITIONS PREVENTED MONITORING AND ALARMING FOR A TOTAL DURATION OF 18 MINUTES. THE MONITORING TIME BETWEEN INOPS WAS VERY SHORT (AN AVERAGE OF ABOUT 6 SECONDS BETWEEN INOPS) WHICH PREVENTED THE DEVICE FROM BEING ABLE TO COMPLETE A MEASUREMENT AND ALARM. THE INOP MESSAGES WERE ACKNOWLEDGED FROM THE PIC IX CENTRAL STATION "PICIXVCCOV7" 16 TIMES DURING THE SAME TIME PERIOD. PER THE INSTRUCTIONS FOR USE (IFU) FO THE MX40 DEVICE, "A CONFIGURABLE AVERAGING FILTER IS USED IN THE CALCULATION OF THE RESULT. DISPLAYED RESULTS ARE TYPICALLY UPDATED EVERY SECOND, BUT THE UPDATE PERIOD CAN BE AUTOMATICALLY DELAYED BY UP TO 30 SECONDS IN THE PRESENCE OF NOISE", AND "THE ALARM DELAY TIME IS 0 TO 30 SECONDS (ADJUSTABLE IN ONE SECOND STEPS) WITH A DEFAULT TIME FOR DESAT OF 20 SECONDS AND A DEFAULT TIME FOR HIGH/LOW ALARM OF 10 SECONDS. THE DELAY TIME FOR ALARM AVAILABILITY ON THE NETWORK IS LESS THAN FIVE SECONDS". IT WAS ALSO CONFIRMED THAT THE SPEAKER WAS WORKING AS EXPECTED AND THAT THE VOLUME WAS SET TO A LEVEL DETERMINED BY HOSPITAL POLICY. THE VOLUME WAS SET TO 1. THE PSE PROVIDED THE FOLLOWING INFORMATION FROM THE MX40 IFU: ALARM TEXT: <SPO2 LABEL> SENSOR OFF (NOTE ¿ THE ABILITY OF THE ALGORITHM TO DETECT THIS CONDITION DEPENDS ON THE SENSOR TYPE IN USE.) CONDITION: THE ALGORITHM HAS DETERMINED THAT A SENSOR IS CONNECTED, BUT NOT PROPERLY APPLIED TO THE PATIENT. WHAT TO DO: APPLY THE SENSOR ACCORDING TO THE MANUFACTURER'S INSTRUCTIONS. IF THE CONDITION PERSISTS, RELOCATE THE SENSOR TO A DIFFERENT SITE ON THE PATIENT. ALARM TEXT: <SPO2 LABEL> NO SENSOR (WARNING: SILENCING THIS TECHNICAL ALARM TURNS OFF THE SPO2 MEASUREMENT ON THE MX40 AND AT THE INFORMATION CENTER WHEN OPERATING IN MANUAL OR CONTINUOUS MODE. IT DOES NOT TURN THE MEASUREMENT OFF WHEN OPERATING IN AUTO MODE.) SOURCE - MX40 CONDITION: NO SENSOR ATTACHED TO SPO2 DEVICE. WHAT TO DO: ATTACH SPO2 SENSOR. ALARM TEXT: <SPO2 LABEL> NO PULSE (NOTE ¿ FOR FAST SPO2 ONLY, THIS INOP MAY ALSO BE CONFIGURED TO DISPLAY AS A RED OR YELLOW TECHNICAL ALARM.) SOURCE - MX40 CONDITION: PULSE IS TOO WEAK OR NOT DETECTABLE. WHAT TO DO: CHECK CONNECTION TO PATIENT. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS RELATED TO SPO2 SENSORS NOT BEING ABLE TO READ THE SIGNAL DUE TO EITHER THE SENSOR BEING OFF, OR BEING CAUSED BY INTERFERENCE (NOISY SIGNAL). THE REPORTED PROBLEM OF AN ALARM FAILURE WAS NOT CONFIRMED, AS NUMEROUS INOP ALARMS WERE SEEN, AND THE USERS HAD SILENCED ALARMS MULTIPLE TIMES. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22579882&pc=MKJ | 3030677-2025-001690 | 2025-07-16 04:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-07-23 04:00 | MKJ | FRX (W B) DEFIB, US ENGLISH, EXCHG | Failure to Deliver Shock/Stimulation | No Clinical Signs, Symptoms or Conditions | P180028 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT THE DEVICE DELIVERED ONE SHOCK DURING A CARDIAC EMERGENCY, THEN ANNOUNCED, "LOW BATTERY ¿ REPLACE IMMEDIATELY" AND WAS UNABLE TO DELIVER A SECOND SHOCK; THE PATIENT WAS PRONOUNCED DECEASED AT THE FACILITY BEFORE TRANSFER TO THE HOSPITAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22579892&pc=OZD | 1220648-2025-30654 | 2025-07-11 04:00 | Death | ABIOMED, INC. | 2025-07-23 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS), AND POST IMPLANT, SAME DAY, THE PATIENT EXPERIENCED LIMB ISCHEMIA, ACCESS SITE BLEEDING AND SUBSEQUENTLY EXPIRED. THE PATIENT PRESENTED WITH DIZZINESS AND GENERAL MALAISE. ELECTROCARDIOGRAM REVEALED ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION WHICH PROMPTED EMERGENT INTERVENTION. THE IMPELLA CP DEVICE WAS IMPLANTED GIVEN THE PROFOUND HYPOTENSION AND STARTED IN AUTO. FURTHER WORK UP REVEALED A "LARGE" VENTRICULAR SEPTAL DEFECT. THE TEAM HYPOTHESIZED THAT THE CONDITION WAS A TAKOTSUBO CARDIOMYOPATHY VENTRICULAR SEPTAL DEFECT. THE PATIENT THEN BEGAN REPORTING FOOT PAIN/NUMBNESS. THERE WAS NO DISTAL SIGNAL ON DOPPLER. A PERFUSION CATHETER WAS SUCCESSFULLY PLACED. NOW BLEEDING AT THE ACCESS SITE WAS NOTED, AND CONSEQUENTLY, THE SHEATH WAS REPOSITIONED, ANGLE AUGMENTED, PERCLOSE SUTURES WERE RETIGHTENED, AND MANUAL PRESSURE WAS APPLIED AT THE SITE. THE PATIENT WOULD CONTINUE ON SUPPORT. HOWEVER, TWO DAYS LATER OVERNIGHT, THE PATIENT WENT INTO CARDIAC ARREST AND EXPIRED. CARE WAS WITHDRAWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22580035&pc=NGV | 2124215-2025-47587 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Transient Ischemic Attack |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR TRANSIENT ISCHEMIC ATTACK (TIA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 201 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22580126&pc=NGV | 2124215-2025-47849 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Thromboembolism |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR THROMBOEMBOLISM. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 427 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22580190&pc=NGV | 2124215-2025-47877 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Thromboembolism |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR THROMBOEMBOLISM. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 265 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22580261&pc=NGV | 2124215-2025-48486 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 695 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22580338&pc=NGV | 2124215-2025-47844 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 4 | Event Description: THIS REPORT SUMMARIZES 4 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 56 TO 93 YEARS. PATIENT SEX: 2 FEMALES; 2 MALES. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 289 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22580486&pc=MNS | 2518422-2025-048327 | 2021-05-17 04:00 | Death | RESPIRONICS, INC. | 2025-07-23 04:00 | MNS | RENTAL, DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A RENTAL, DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. DESPITE OF MULTIPLE ATTEMPTS THE DEVICE HAS NOT YET RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22580492&pc=NGV | 2124215-2025-47875 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 5 | Event Description: THIS REPORT SUMMARIZES 5 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 75 TO 85 YEARS, PATIENT SEX: 2 FEMALES; 3 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025 AVERAGE TIME TO EVENT: 290 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22580502&pc=NGV | 2124215-2025-48488 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 4 | Event Description: THIS REPORT SUMMARIZES 4 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 52 TO 81 YEARS. PATIENT SEX: 1 FEMALE; 3 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 588 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22580542&pc=DSQ | 2916596-2025-04869 | 2025-07-13 04:00 | Death | THORATEC CORPORATION | 2025-07-23 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Tachycardia | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS PULSELESS AND UNRESPONSIVE WHEN EMERGENCY MEDICAL SERVICES (EMS) ARRIVED. ADVANCED CARDIAC LIFE SUPPORT (ACLS) WAS STARTED, AND THE PATIENT'S INITIAL PRESENTING RHYTHM WAS VENTRICULAR TACHYCARDIA. THE PATIENT WAS CARDIOVERTED IN THE FIELD AND IT WAS REPORTED THAT THEY SUBSEQUENTLY HAD EPISODES OF ATRIAL FIBRILLATION. THE PATIENT UNDERWENT RESUSCITATION EFFORTS FOR APPROXIMATELY 50 MINUTES PRIOR TO ARRIVAL IN THE EMERGENCY ROOM (ER). THE PATIENT PASSED AWAY IN THE ER ON (B)(6) 2025. LOG FILES WERE SENT FOR REVIEW. THE EVENT LOG CAPTURED INTERMITTENT LOW FLOW EVENTS ON (B)(6) 2025 AT 19:08 THROUGH 20:25 WITH FLOW NOTED AS LOW AS 1.9 LPM. PULSATILITY INDEX (PI) VALUES WERE SLIGHTLY VARIABLE AND NOTED IN THE 2-4 RANGE. THE DRIVELINE WAS DISCONNECTED ON (B)(6) 2025 AT 20:25. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22580666&pc=NIQ | 2024168-2025-07633 | 2023-12-01 05:00 | Death | ABBOTT VASCULAR | 2025-07-23 04:00 | NIQ | XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED THROUGH A RESEARCH ARTICLE REGARDING A PRE-SPECIFIED SUBGROUP ANALYSIS BASED ON COMPLEX PERCUTANEOUS CORONARY INTERVENTION WITH IMPLANTED XIENCE STENTS IN THE 1 YEAR RESULTS OF THE STOPDAPT-3 (SHORT AND OPTIMAL DURATION OF DUAL ANTIPLATELET THERAPY-3) TRIAL, WHICH RANDOMLY COMPARED 1-MONTH DAPT FOLLOWED BY ASPIRIN MONOTHERAPY (ASPIRIN GROUP) WITH 1-MONTH PRASUGREL MONOTHERAPY FOLLOWED BY CLOPIDOGREL MONOTHERAPY (CLOPIDOGREL GROUP). IT WAS NOTED SOME PATIENT DIED BY 30 DAYS OR LOST TO FOLLOW UP AT 30 DAYS BEFORE THE ANALYSIS BEGAN; THEREFORE, THOSE PATIENT'S WERE EXCLUDED FROM THE STUDY. THE MAIN ANALYSIS IN THE PRESENT STUDY WAS THE 30-DAY LANDMARK ANALYSIS. THE CO-PRIMARY ENDPOINTS WERE CARDIOVASCULAR EVENTS (A COMPOSITE OF CARDIOVASCULAR DEATH, MYOCARDIAL INFARCTION, DEFINITE STENT THROMBOSIS, OR ISCHEMIC STROKE AND HEMORRHAGIC STROKE, TARGET LESION REVASCULARIZATION, ) AND MAJOR BLEEDING (BLEEDING ACADEMIC RESEARCH CONSORTIUM 3 OR 5). IN THE 30-DAY LANDMARK ANALYSIS ( N = 5833), THERE WERE 1415 PATIENTS (24.3%) WHO UNDERWENT COMPLEX PCI. THERE WAS A SIGNIFICANT INTERACTION BETWEEN COMPLEX PCI AND THE EFFECT OF ASPIRIN MONOTHERAPY RELATIVE TO CLOPIDOGREL MONOTHERAPY BEYOND 1 MONTH AND UP TO 1 YEAR FOR CARDIOVASCULAR EVENTS DUE TO NUMERICALLY LOWER RISK OF ASPIRIN MONOTHERAPY IN PATIENTS WITH COMPLEX PCI, WHILE THE EFFECT OF ASPIRIN MONOTHERAPY RELATIVE TO CLOPIDOGREL MONOTHERAPY WAS NOT DIFFERENT FOR BLEEDING REGARDLESS OF COMPLEX PCI. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED " ASPIRIN VS. CLOPIDOGREL MONOTHERAPY BEYOND 1 MONTH AFTER COMPLEX PERCUTANEOUS CORONARY INTERVENTION: A PRE-SPECIFIED SUBGROUP ANALYSIS OF THE STOPDAPT-3 TRIAL" NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICES WERE NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBERS AND/OR LOT NUMBERS WERE NOT REPORTED, AND THE DEVICES WERE NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT DEATHS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECTS OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. B2, B3, B6, D6: DATES ESTIMATED D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS REFERENCED IN B5 ARE FILED UNDER A SEPARATE MEDWATCH REPORT. ATTACHMENT: ARTICLE TITLE ¿ASPIRIN VS. CLOPIDOGREL MONOTHERAPY BEYOND 1 MONTH AFTER COMPLEX PERCUTANEOUS CORONARY INTERVENTION: A PRE-SPECIFIED SUBGROUP ANALYSIS OF THE STOPDAPT-3 TRIAL.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22580672&pc=DSQ | 2916596-2025-04852 | 2025-07-13 04:00 | Death | THORATEC CORPORATION | 2025-07-23 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Detachment of Device or Device Component; Infusion or Flow Problem | Hemorrhage/Bleeding; Hemothorax; Hypovolemic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU) ON (B)(6) 2025, EXPERIENCING LOW FLOW ALARMS IN HEMORRHAGIC SHOCK WITH TENSION HEMOTHORAX UNEXPECTEDLY WITHOUT ANTECEDENT TRAUMA OR SURGERY. THE PATIENT WAS TRANSFUSED AND TEMPORARILY STABILIZED WITH CHEST TUBE PLACEMENT, BUT THE BLEEDING CONTINUED. THE PATIENT WAS ON PRESSOR SUPPORT AND FLOWS NORMALIZED. ON COMPUTED TOMOGRAPHY (CT) IMAGING AND ECHOCARDIOGRAM, AN ACTIVE BLEEDING SOURCE COULD NOT BE FOUND. THE PATIENT'S FAMILY CHOSE TO TRANSITION TO COMFORT CARE, AND THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT'S FAMILY AND THE MEDICAL TEAM AGREED TO PERFORM AN AUTOPSY TO HELP IDENTIFY THE CAUSE OF DEATH, AND THE CORONER NOTED ON (B)(6) 2025 THAT IT WAS A PUMP-RELATED ISSUE. THE BEND RELIEF WAS DETACHED AND THERE WAS A VISIBLE HOLE IN THE OUTFLOW GRAFT WHERE BLOOD WOULD HAVE BEEN EXSANGUINATING. COMPUTED TOMOGRAPHY IMAGES WERE ATTACHED WHICH SHOWED THE BEND RELIEF APPEARING ATTACHED ON (B)(6) 2024, HOWEVER SUBSEQUENTLY APPEARING DETACHED ON (B)(6) 2024 AND (B)(6) 2025. LOG FILES WERE SUBMITTED FOR REVIEW, AND THE SITE WOULD LIKE THE PUMP EXAMINED TO SEE IF THERE WAS ANY FLAW WITH THE BEND RELIEFS ABILITY TO LOCK ONTO THE OUTFLOW GRAFT AND PUMP. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22580912&pc=LWS | 1028232-2025-03481 | 2025-07-22 04:00 | Death | BIOTRONIK SE & CO. KG | 2025-07-23 04:00 | LWS | ACTICOR 7 VR-T DX DF4 PROMRI |  | Ventricular Fibrillation | P050023 |  | 1 | Event Description: PATIENT WAS IN HOSPITAL CRITICALLY ILL FROM ISSUES OTHER THAN DEVICE FUNCTION. PATIENT HAD SEVERAL VF EPISODES WITH TERMINATION PROVIDED BY SHOCKS FROM DEVICE STARTING (B)(6) 2025. THE LAST VF EPISODE WAS (B)(6) 2025, AROUND 0200 ESTD (APPROXIMATE TIME.) PATIENT WENT INTO VF AND THE DEVICE SHOCKED WITHOUT CONVERTING TO NORMAL SINUS. UPON INTERROGATION POST EPISODE, INTERROGATION REVEALED DEVICE IN BACK UP MODE. DEVICE FUNCTIONED APPROPRIATELY BUT PATIENT EXPIRED. DEVICE REMAINS IMPLANTED IN PATIENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22582560&pc=BZD | 2518422-2025-108201 | 2025-07-07 04:00 | Death | RESPIRONICS, INC. | 2025-07-23 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THE CAUSE DEATH WAS NOT PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22582865&pc=BZD | 2518422-2025-108199 | 2025-07-07 04:00 | Death | RESPIRONICS, INC. | 2025-07-23 04:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION A USER OF UNKNOWN DEVICE HAD PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED; THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE PATIENT'S DAUGHTER WOULD LIKE TO DONATE THE DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22583613&pc=NIK | 2017865-2025-95807 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-23 04:00 | NIK | GALLANT HF | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE OR LEAD RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.Event Description: NEW INFORMATION RECEIVED INDICATES THAT THERE ARE NO ALLEGATIONS FROM THE HEALTH CARE PROFESSIONAL THAT WOULD INDICATE THAT THE DEATH WAS DEVICE OR LEAD RELATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22583620&pc=NVN | 2017865-2025-95808 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-23 04:00 | NVN | TENDRIL MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE OR LEAD RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.Event Description: NEW INFORMATION RECEIVED INDICATES THAT THERE ARE NO ALLEGATIONS FROM THE HEALTH CARE PROFESSIONAL THAT WOULD INDICATE THAT THE DEATH WAS DEVICE OR LEAD RELATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22583627&pc=NIK | 2017865-2025-95809 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-23 04:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE OR LEAD RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.Event Description: NEW INFORMATION RECEIVED INDICATES THAT THERE ARE NO ALLEGATIONS FROM THE HEALTH CARE PROFESSIONAL THAT WOULD INDICATE THAT THE DEATH WAS DEVICE OR LEAD RELATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22583629&pc=NVY | 2017865-2025-95810 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-23 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE OR LEAD RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED.Event Description: NEW INFORMATION RECEIVED INDICATES THAT THERE ARE NO ALLEGATIONS FROM THE HEALTH CARE PROFESSIONAL THAT WOULD INDICATE THAT THE DEATH WAS DEVICE OR LEAD RELATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22583771&pc=NIQ | 2124215-2025-46721 | 2025-05-22 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-23 04:00 | NIQ | SYNERGY XD | Break | Cardiac Arrest; Obstruction/Occlusion; Foreign Body In Patient |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT UNDERWENT INTRA-ANGIOGRAM FOR CORONARY ARTERY DISEASE. THE TARGET LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD). POST ROTABLATION, INTRAVASCULAR ULTRASOUND WAS PERFORMED FOLLOWED BY MULTIPLE BALLOON INFLATIONS IN THE PROXIMAL AND MID LAD. A 2.50 X 32MM SYNERGY XD DRUG-ELUTING STENT WAS THEN PLACED DOWN THE LAD AND WHILE ATTEMPTING TO FIND THE CORRECT POSITION FOR DEPLOYMENT, A SECTION OF THE SHAFT WITH THE STENT BALLOON AND STENT ATTACHED SHEARED OFF AND LODGED IN THE LAD, OCCLUDING THE VESSEL. THE PATIENT DETERIORATED AND CARDIAC ARREST OCCURRED. THE PATIENT WAS TAKEN FOR URGENT CARDIOTHORACIC SURGERY AND AN URGENT STERNOTOMY AND CORONARY ARTERY BYPASS GRAFTING WAS PERFORMED. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT DEPARTMENT POST-OPERATIVELY BUT AFTER ACTIVE TREATMENT AND HIGH DOSE INOTROPE SUPPORT THE DECISION WAS MADE BY THE FAMILY TO WITHDRAW TREATMENT DUE THE CONTINUED DETERIORATION OF THE PATIENT CONDITION. ON THE NEXT DAY, THE PATIENT DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22584189&pc=MNS | 2518422-2025-108174 | 2025-05-29 04:00 | Death | RESPIRONICS, INC. | 2025-07-23 04:00 | MNS | BIPAP AUTOSV ADV SYSTEM ONE W/HUMID, CAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K090539 |  | 1 | Event Description: THE MANUFACTURER WAS INFORMED THAT A PATIENT USING THE BIPAP AUTOSV ADV SYSTEM ONE W/HUMID, CAN DEVICE HAD PASSED AWAY. THE PATIENT'S SON IS NOW THE SECONDARY USER OF THE DEVICE. THE RESPIRATORY THERAPIST SUSPECTS THAT THE SON HAS SOLD OR PLANS TO SELL THE REPLACEMENT DEVICE. HE WAS ADVISED THAT SELLING THE DEVICE IS ILLEGAL AND UNSAFE. NO CLAIMS SUGGEST THE DEVICE CAUSED THE PATIENT'S DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22584224&pc=NAY | 2955842-2025-31351 | 2025-06-06 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-07-23 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Bowel Perforation | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED CHOLECYSTECTOMY PROCEDURE, THE PATIENT EXPIRED. THE INITIAL PROCEDURE REPORTEDLY WENT WELL WITH NO COMPLICATIONS. THE SITE PERIOPERATIVE DIRECTOR STATED THE PATIENT WAS IN MORE PAIN POST-OPERATIVELY THAN NORMAL AND WAS ADVISED TO STAY OVERNIGHT FOR MONITORING; HOWEVER, THE PATIENT REFUSED AND WENT HOME. AN INTUITIVE SURGICAL CLINICAL REPRESENTATIVE WAS INFORMED THAT THE SURGEON HAD USED A CUTDOWN METHOD FOR INITIAL ENTRY TO THE ABDOMEN [FOR PORT PLACEMENT]. THE PATIENT RETURNED TO THE HOSPITAL AT AN UNSPECIFIED TIME AND WAS RE-ADMITTED FOR A "HURT BOWEL". THE PATIENT WAS THEN TRANSFERRED TO ANOTHER HOSPITAL AND EXPIRED THERE ON AN UNKNOWN DATE. ATTEMPTS WERE MADE TO SPEAK WITH THE SURGEON AND WERE UNSUCCESSFUL. THE PERIOPERATIVE DIRECTOR REPORTED THAT THE PATIENT DEATH HAD NOTHING TO DO WITH THE DA VINCI SYSTEM AND PRODUCTS USED, AND STATED THAT THERE WOULD BE NO FURTHER INFORMATION PROVIDED FOR THIS EVENT. Manufacturer Narrative: THE SPECIFIC PROCEDURE TIME FOR THIS PROCEDURE WAS NOT PROVIDED; THEREFORE, NO DA VINCI SYSTEM OR ACCESSORIES LOG FILES, OR DEVICE HISTORY RECORDS ARE AVAILABLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AFTER A BOWEL INJURY DURING A CHOLECYSTECTOMY. THE INJURY WAS NOT NOTED AT THE TIME OF THE INITIAL SURGERY. HOW THE INJURY OCCURRED IS NOT PROVIDED. THE DETAILS OF HOW THEY DIED AND THE REMAINDER OF THEIR POST OPERATIVE COURSE WAS NOT PROVIDED. NO ALLEGATION HAS BEEN MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SECTION B2: PATIENT DATE OF DEATH IS UNKNOWN AS RETURNED "LATER" TO THE HOSPITAL, THEN TRANSFERRED TO ANOTHER USER FACILITY WHERE THE PATIENT EXPIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22562824&pc=PCU | 3005099803-2025-03278 | 2023-08-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-22 04:00 | PCU | AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Inflammation | K150692 |  | 1 | Event Description: BOSTON SCIENTIFIC CORPORATION BECAME AWARE OF AN EVENT INVOLVING AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEMS AND WALLFLEX BILIARY RX FULLY COVERED STENT SYSTEMS THROUGH THE ARTICLE TITLED, "OPTIMIZING EUS-GUIDED CHOLEDOCHODUODENOSTOMY WITH LUMEN-APPOSING METAL STENTS FOR PRIMARY DRAINAGE OF MALIGNANT DISTAL BILIARY OBSTRUCTION (SCORPION-IIP): A PROSPECTIVE PILOT STUDY", BY JESKA A. FRITZSCHE ET AL. PER THE ARTICLE, BETWEEN AUGUST 2023 TO APRIL 2024, 27 PATIENTS SUFFERING FROM MALIGNANT DISTAL BILE DUCT OBSTRUCTION UNDERWENT EUS-GUIDED CHOLEDOCHODUODENOSTOMY (EUS-CDS) USING AXIOS AS LUMEN-APPOSING METAL STENTS (LAMS) WITH A PLACEMENT OF A WALLFLEX FULLY COVERED SELF-EXPANDING METAL STENTS THROUGH IT. ONE MORTALITY OCCURRED DURING THE TEST PERIOD OF THESE PROCEDURES POTENTIALLY RELATED TO INGROWTH IN THE WALLFLEX DEVICE AND CHOLECYSTITIS. THE PATIENT ALSO HAD CONCOMITANT RENAL FAILURE. Manufacturer Narrative: BLOCK B2: THE EXACT DATE OF THE PATIENTS' DEATH WAS NOT REPORTED. THE RETROSPECTIVE STUDY START DATE OF AUGUST 1, 2023, IS USED FOR THE ESTIMATED DATE OF DEATH. BLOCK B3: THE EXACT DATE OF EVENT WAS NOT REPORTED. THE RETROSPECTIVE STUDY START DATE OF AUGUST 1, 2023, IS USED FOR THE ESTIMATED DATE OF EVENT. BLOCK D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WERE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK G2: LITERATURE SOURCE: JESKA A. FRITZSCHE; PAUL FOCKENS; MARC G. BESSELINK; OLIVER R. BUSCH; FREEK DAAMS; MATTHEUS C. B. WIELENGA; JOHANNA W. WILMINK; ROGIER P. VOERMANS; ROY L. J. VAN WANROOIJ. "OPTIMIZING EUS-GUIDED CHOLEDOCHODUODENOSTOMY WITH LUMEN-APPOSING METAL STENTS FOR PRIMARY DRAINAGE OF MALIGNANT DISTAL BILIARY OBSTRUCTION (SCORPION-IIP): A PROSPECTIVE PILOT STUDY." GASTROINTESTINAL ENDOSCOPY 2025; VOLUME 101, NO. 5: 1009-1016. BLOCK H6: IMDRF PATIENT CODE E2326 CAPTURES THE REPORTABLE EVENT OF CHOLECYSTITIS. IMDRF IMPACT CODE F02 CAPTURES THE PATIENT'S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22563854&pc=OZD | 1220648-2025-30559 | 2025-06-22 04:00 | Death | ABIOMED, INC. | 2025-07-22 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT ON CP SUPPORT FOR HIGH RISK CARDIAC PROCEDURE. THE CP WAS PLACED AND ALLOWED FOR THE CORONARY ARTERY BYPASS GRAFT SURGERY (CABG) AND WAS THEN EXPLANTED. THERE WERE NO ALLEGED ISSUES DURING SUPPORT. PATIENT SURVIVED TO EXPLANT OF THE PUMP PER CASE DOCUMENTATION. HOWEVER, A MONTH POST-IMPLANT, ABIOMED/JNJ LEGAL FORWARDED AN ALLEGATION OF THE PATIENT'S DEATH. THE REPORT/DEATH CERTIFICATE STATES THE DEATH THE DAY AFTER EXPLANT AND CAUSE OF DEATH AS CARDIOGENIC SHOCK AND SEVERE CORONARY ARTERY DISEASE. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22564276&pc=OZD | 1220648-2025-30567 | 2025-06-25 04:00 | Death | ABIOMED, INC. | 2025-07-22 04:00 | OZD | IMPELLA CP | Failure to Advance; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AN IMPELLA CP PUMP WAS UNABLE TO ADVANCE THROUGH THE PATIENT'S VASCULATURE DURING ATTEMPTED DELIVERY. THE IMPLANT WAS SUBSEQUENTLY ABORTED AS A RESULT OF THE NOTED DIFFICULTY. THE PATIENT PASSED LATER THAT NIGHT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION SINCE THE PUMP COULD NOT BE INSERTED DUE TO THE PATIENT ANATOMY AND HIGH AMOUNTS OF INOTROPIC AND VASOACTIVE DROPS PER STAFF. B1 PRODUCT PROBLEM INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30567. Manufacturer Narrative: A4 AND A5 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22564359&pc=OZD | 1220648-2025-30569 | 2025-06-27 04:00 | Death | ABIOMED, INC. | 2025-07-22 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Renal Failure | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED ACUTE RENAL FAILURE. IT IS UNKNOWN IF THIS WAS PREEXISTING CONDITION OR IF IT WAS IMPELLA RELATED. THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22564806&pc=DRY | 3006073153-2025-00174 | 2025-06-15 04:00 | Death | SPECTRUM MEDICAL LTD | 2025-07-22 04:00 | DRY | QUANTUM WORKSTATION 12.1" ELITE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K202557 |  | 1 | Event Description: ADVERSE EVENT OCCURRED THAT RESULTED IN PATIENT DEATH WHILE PATIENT WAS CONNECTED TO THE MEDICAL DEVICE. USER REPORTED THAT BLOOD FLOW WAS UNABLE TO BE RE-ESTABLISHED FOLLOWING CPR AND PATIENT EXPERIENCED HEMORRHAGE BENEATH THE HEART. THE INVESTIGATION DETERMINED THAT THE DEVICE DID NOT MALFUNCTION DURING USE AND WAS OPERATING AS INTENDED. Manufacturer Narrative: THE INVESTIGATION CONFIRMED THE ADVERSE EVENT OCCURRED AS THE USER DESCRIBED; HOWEVER, THERE WAS NO EVIDENCE OF A PROBLEM WITH THE QUANTUM SYSTEM THAT WOULD HAVE RESULTED IN SUCH EVENT. IT WAS CONCLUDED THAT THE QUANTUM SYSTEM WORKED AS INTENDED AND THE ADVERSE EVENT OCCURRED DUE TO EXTERNAL FACTORS FROM THE QUANTUM SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22565005&pc=DSQ | 2916596-2025-04873 | 2025-04-07 04:00 | Death | THORATEC CORPORATION | 2025-07-22 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT LOST POWER DUE TO BAD STORMS AND WAS UNABLE TO HEAR THE ALARMS UNTIL IT WAS TOO LATE. THE PATIENT WAS DEAF WITH HEARING AIDS AND LIVED ALONE. THE PATIENT WAS FOUND DECEASED IN A POSITION THAT LOOKED LIKE THEY WERE TRYING TO GER THE BATTERIES BUT COULD NOT MAKE IT. THE PUMP STOPPED AS A RESULT OF NO EXTERNAL POWER. NO AUTOPSY WAS TO BE PERFORMED. NO ADDITIONAL QUESTIONS WERE TO BE ANSWERED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE ACCOUNT REPORTED THAT HEARTMATE 3 LVAS, SERIAL NUMBER: (B)(6), WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU IS CURRENTLY AVAILABLE. CHAPTER 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22565901&pc=QMG | 3015053858-2025-00088 |  | Death | SHOCKWAVE MEDICAL, INC. | 2025-07-22 04:00 | QMG | SHOCKWAVE C2+ CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | Adverse Event Without Identified Device or Use Problem | Sepsis; Perforation of Vessels | P200039/S008 |  | 1 | Event Description: A SHOCKWAVE CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER WAS USED TO TREAT A LESION IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD). FOLLOWING THE SECOND CYCLE OF IVL PULSES, A PERFORATION OCCURRED. THE PHYSICIAN THEN UTILIZED A PING PONG TECHNIQUE VIA FEMORAL ACCESS, SUCCESSFULLY COMPLETING THE IVL PROCEDURE. HOWEVER, THE PATIENT DEVELOPED SEPSIS FIVE DAYS POST-PROCEDURE AND SUBSEQUENTLY PASSED AWAY. THE DEVICE INFORMATION ASSOCIATED WITH THIS EVENT COULD NOT BE CONFIRMED. THEREFORE, THIS EVENT IS BEING ASSESSED AS A SHOCKWAVE C2+ CORONARY DEVICE BY DEFAULT. THE DATE OF THE PROCEDURE, DATE OF DEATH, AND PATIENT INFORMATION ARE UNKNOWN. Manufacturer Narrative: THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION CANNOT BE PERFORMED. THE CAUSE OF THE PERFORATION, SEPSIS, AND SUBSEQUENT DEATH COULD NOT BE DEFINITIVELY DETERMINED BASED ON THE INFORMATION AVAILABLE. THERE WAS NO IVL DEVICE MALFUNCTION REPORTED. THE LOT NUMBER OF THE DEVICE WAS NOT PROVIDED. THEREFORE, REVIEW OF THE DEVICE MANUFACTURING AND TEST DOCUMENTATION COULD NOT BE COMPLETED. HOWEVER, SHOCKWAVE MEDICAL, INC. HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO DISTRIBUTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22566071&pc=NKM | 2135147-2025-04009 | 2025-06-27 04:00 | Death | ABBOTT MEDICAL | 2025-07-22 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT ON AN UNKNOWN DATE A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT MITRAL REGURGITATION (MR). AN UNKNOWN MITRACLIP WAS DEPLOYED ON THE MITRAL VALVE. ON AN UNKNOWN DATE, THE PATIENT DIED.Event Description: IT WAS REPORTED THAT ON AN UNKNOWN DATE A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT MITRAL REGURGITATION (MR). AN UNKNOWN MITRACLIP WAS DEPLOYED ON THE MITRAL VALVE.
ON AN UNKNOWN DATE, THE PATIENT DIED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS, AND A REVIEW OF THE LOT HISTORY RECORD AND SIMILAR COMPLAINT REVIEW COULD NOT BE PERFORMED AS THE PART AND LOT NUMBER REGARDING THE COMPLAINT DEVICE WAS NOT PROVIDED. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR THE REPORTED DEATH. DEATH IS LISTED IN THE INSTRUCTIONS FOR USE AS A KNOWN POSSIBLE COMPLICATION ASSOCIATED WITH THE MITRACLIP PROCEDURES. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22566567&pc=OZD | 1220648-2025-30592 | 2025-07-10 04:00 | Death | ABIOMED, INC. | 2025-07-22 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED EPISODES OF VENTRICULAR TACHYCARDIA AND VENTRICULAR FIBRILLATION REQUIRING DEFIBRILLATION. THE PATIENT'S FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST SYSTEM LISTS VENTRICULAR ARRYTHMIA AS AN ADVERSE EVENT THAT OCCURRED WHILE IN USE WITH THE IMPELLA DEVICE. IT STATES THAT "ONLY A FRACTION OF THESE RATES WERE ATTRIBUTED TO THE IMPELLA DEVICE AND THE EVENTS RESOLVED WITHOUT RESIDUAL EFFECT IN MOST OF THE CASES, UNLESS THE EVENT OF DEATH OCCURRED. OVERALL, THE RESULTS DID NOT SHOW ANY EVIDENCE OF INCREASED MORBIDITY ASSOCIATED WITH THE IMPELLA SUPPORT IN CARDIOMYOPATHY, MYOCARDITIS, AND PERIPARTUM CARDIOMYOPATHY (PPCM) PATIENTS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22566755&pc=NVY | 1028232-2025-03446 | 2025-07-04 04:00 | Death | BIOTRONIK SE & CO. KG | 2025-07-22 04:00 | NVY | LINOX SD 60/16 | Ambient Noise Problem | Ventricular Fibrillation | P980023 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS NOISE SEEN ON (B)(6) VF RECORDING. PATIENT PASSED AWAY SAME DAY THIS WAS RECORDED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THIS FILE WILL BE UPDATED. Manufacturer Narrative: COMBINATION PRODUCT: YES. Manufacturer Narrative: COMBINATION PRODUCT: YES. THE DEVICE IS CURRENTLY NOT AVAILABLE FOR ANALYSIS. NO CONCLUSION CAN BE DRAWN BASED ON AVAILABLE INFORMATION AT THE MOMENT. THE FILE IS CLOSED. THE INVESTIGATION WILL BE RE-OPENED SHOULD ADDITIONAL DATA BECOME AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22566826&pc=OZD | 1220648-2025-30594 | 2025-07-10 04:00 | Death | ABIOMED, INC. | 2025-07-22 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED TACHYCARDIA. AMIODARONE WAS ADMINISTERED. ULTIMATELY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST SYSTEM LISTS VENTRICULAR ARRYTHMIA AS AN ADVERSE EVENT THAT OCCURRED WHILE IN USE WITH THE IMPELLA DEVICE. IT STATES THAT "ONLY A FRACTION OF THESE RATES WERE ATTRIBUTED TO THE IMPELLA DEVICE AND THE EVENTS RESOLVED WITHOUT RESIDUAL EFFECT IN MOST OF THE CASES, UNLESS THE EVENT OF DEATH OCCURRED. OVERALL, THE RESULTS DID NOT SHOW ANY EVIDENCE OF INCREASED MORBIDITY ASSOCIATED WITH THE IMPELLA SUPPORT IN CARDIOMYOPATHY, MYOCARDITIS, AND PERIPARTUM CARDIOMYOPATHY (PPCM) PATIENTS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22568412&pc=OZD | 1220648-2025-30598 | 2025-07-07 04:00 | Death | ABIOMED, INC. | 2025-07-22 04:00 | OZD | IMPELLA CP | Malposition of Device | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT THE WAVEFORM OF AN IMPLANTED IMPELLA CP BECAME UNCOUPLED DUE TO THE PIGTAIL GETTING CAUGHT IN THE MITRAL VALVE. INITIAL ATTEMPTS TO FREE THE PIGTAIL WERE UNSUCCESSFUL, BUT THE PUMP WAS EVENTUALLY ADVANCED AND THE WAVEFORMS STABILIZED. TWO DAYS LATER, THE LEFT FEMORAL IMPELLA SIDE DEVELOPED SANGUINEOUS DRAINAGE AFTER THE PATIENT MOVED. THE DRESSING WAS CHANGED, HEPARIN WAS WITHHELD, AND ONE UNIT OF PACKED RED BLOOD CELLS WAS TRANSFUSED TO TREAT THE BLEED. THE PATIENT WAS TRANSFERRED TO COMFORT CARE AND TWO DAYS LATER PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22568421&pc=NPT | 2015691-2025-05946 | 2025-07-02 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-22 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED, IT WAS A CASE OF AN IMPLANT OF A 26MM SAPIEN 3 ULTRA RESILIA VALVE, IN AORTIC POSITION BY TRANSFEMORAL APPROACH. DURING THE PROCEDURE, BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED. DIFFICULTIES WERE FOUND CROSSING THE NATIVE VALVE WITH THE DELIVERY SYSTEM AND MULTIPLE ATTEMPTS WERE REQUIRED, BUT THE NATIVE VALVE WAS SUCCESSFULLY CROSSED. THE VALVE WAS SUCCESSFULLY DEPLOYED WITH AN ADDITIONAL +4 ML OF VOLUME. POST-VALVE DEPLOYMENT, A PARAVALVULAR LEAK WAS OBSERVED, PROMPTING A POST-DILATION WITH AN EXTRA 1 ML. SHORTLY THEREAFTER, THE PATIENT EXPERIENCED A SUDDEN DROP IN BLOOD PRESSURE. IMAGING CONFIRMED AN ANNULAR RUPTURE. DESPITE PERICARDIOCENTESIS BEING PERFORMED, THE PATIENT EXPIRED DUE TO RUPTURE OF THE LEFT VENTRICULAR OUTFLOW TRACT (LVOT). THE PERCEIVED ROOT CAUSE OF THE ANNULAR RUPTURE WAS POSSIBLY RELATED TO VALVE POST-DILATION AND LVOT CALCIFICATION. ALTHOUGH THE ANNULUS SIZE WAS WITHIN THE RANGE FOR A 29 MM VALVE, A 26MM VALVE WAS SELECTED DUE TO THE EXTENT OF CALCIFICATION IN THE LVOT. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. AORTIC DISSECTION OR RUPTURE MAY OCCUR WHEN MULTIPLE ATTEMPTS ARE MADE TO CROSS THE STENOTIC NATIVE VALVE, AND/OR WHEN EXCESSIVE FORCE IS USED. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). THE THV TRAINING MANUALS GUIDE TO FACILITATE SAFE CROSSING OF THE NATIVE OR FAILING VALVE, INCLUDING CAMERA PROJECTIONS, HANDLING DURING ADVANCEMENT, AND TROUBLESHOOTING TECHNIQUES IF DIFFICULTY IS ENCOUNTERED. AS STATED, EXCESSIVE FORCE SHOULD NOT BE USED WHEN THE DEVICE HAS DIFFICULTY CROSSING THE STENOTIC VALVE. ADDING TENSION TO THE WIRE, PULLING BACK THE SYSTEM TO RE-ORIENT THE VALVE, AS NEEDED, AND TORQUING THE FLEX CATHETER MAY HELP SOLVE THE PROBLEM. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION THAT A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THE PERCEIVED ROOT CAUSE OF THE ANNULAR RUPTURE WAS POSSIBLY RELATED TO VALVE POST-DILATION AND LVOT CALCIFICATION. HOWEVER, THE CAUSE OF THE EVENT IS UNKNOWN DUE TO THE LIMITED INFORMATION AVAILABLE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22568589&pc=NRY | MW5173055 | 2025-07-05 04:00 | Death | WALLABY MEDICAL | 2025-07-22 04:00 | NRY | ESPERANCE CATHETER | Detachment of Device or Device Component | Hemorrhage/Bleeding; Foreign Body Embolism |  |  | 1 | Event Description: EMERGENT THROMBECTOMY PERFORMED. DURING THE THIRD PASS FOR THROMBECTOMY, PART OF THE ESPERANCE ASPIRATION CATHETER BROKE OFF INTO PATIENT'S RIGHT MIDDLE CEREBRAL ARTERY. THE PIECE WAS RETRIEVED USING A STENT RETRIEVER; HOWEVER, SIGNIFICANT BLEEDING WAS NOTED. ESPERANCE CATHETER: MANUFACTURER: WALLABY MEDICAL LOT#: ASP00572, REF#: ASP3F120KIT, EXP: 2025.12.06 THE WIRE USED WITH THE CATHETER: DRIVEWIRE24. MANUFACTURER: RAPID MEDICAL LOT#: 250301DW01, REF#: DRPP2411, EXP: 01.31.2026.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22568618&pc=QBJ | MW5173056 | 2024-06-23 04:00 | Death | DEXCOM INC | 2025-07-22 04:00 | QBJ | DEXCOM RECEIVER | Defective Device | Hyperglycemia; Vomiting; Diabetic Ketoacidosis |  |  | 1 | Event Description: MY SON WAS ALONE AT HIS APARTMENT HE GOT VERY SICK AND STARTED VOMITING HE PASSED AWAY I FOUND HIS DEXCOM NEXT TO HIM. I HAVE AN AUTOPSY REPORT AND POLICE REPORTS VERIFYING THE VOMIT ALL OVER THE HOUSE HE WAS 22 YEARS OLD HE DIED OF DIABETIC KETOACIDOSIS. PRIOR TO THIS WE WERE AT EMERGENCY ROOM WHEN THE DOCTOR DID LAB WORK AND INFORMED US HIS DEVICE WAS NOT WORKING AT ALL IT SAID HE WAS FINE, BUT THE LAB WORK SAID DIFFERENTLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22568632&pc=KRH | MW5173057 | 2025-06-04 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-07-22 04:00 | KRH | DURAN ANCORE ANNULOPLASTY BAND | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Unspecified Infection |  |  | 1 | Event Description: I AM SUBMITTING THIS ON BEHALF OF MY FATHER, (B)(6). HE PASSED AWAY ON (B)(6) 2025 AT 8:24AM. HE WAS BEING TREATED AT (B)(6) FOR VEGETATION ON HIS MITRAL VALVE REPLACEMENT. HE WAS TO HAVE SURGERY ON (B)(6) TO REPLACE THE INFECTED VALVE AND REDO HIS BYPASS SURGERY DUE TO THE SEVERITY OF THE INFECTION. IN SPEAKING WITH HIS DOCTOR, THIS IS A KNOWN SIDE EFFECT WITH ~9% OF THESE PORCINE MITRAL VALVES DEVELOPING VEGETATION. SADLY, MY DAD IS PART OF THE ~9% WHICH LEAD TO HIS DEATH. IT IS UNDERSTOOD THAT SOMETIME IN THE EARLY MORNING HOURS OF (B)(6) THE VEGETATION BROKE OFF, CAUSING A FATAL HEART ATTACK. THE VEGETATION COULD NOT BE FOUND ON AN ECHO AFTER FAILED RESUSCITATION AND PUTTING HIM ON ECMO. THUS, LEADING TO THE BELIEF THAT THE VEGETATION CAUSED THE FATAL HEART ATTACK. CONVERSATIONS WITH THE DOCTOR REGARDING MY DAD'S QUALITY OF LIFE, WE AGREED THAT IT WAS IN HIS BEST INTEREST TO TAKE HIM OFF LIFE SUPPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22568633&pc=FKX | 0008030665-2025-01651 | 2025-07-02 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-07-22 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K181108 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S SPOUSE THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 WHILE UNDERGOING CCPD THERAPY. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME WHILE UNDERGOING CCPD THERAPY ON (B)(6) 2025. THE PATIENT WAS DISCOVERED BY HER SPOUSE WHEN HE AWOKE IN THE MORNING. THE PDRN REPORTED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA FROM (B)(6) 2025 INDICATED THE PATIENT COMPLETED TREATMENT, AND NO ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING IN THE HOME AND TRANSPORTED THE PATIENT TO THE MORGUE. THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST SECONDARY TO HER EXTENSIVE CARDIAC COMORBID CONDITIONS (E.G., ATRIAL FIBRILLATION, CORONARY ARTERY DISEASE, UNDERWENT MULTIPLE CARDIAC CATHETERIZATIONS WITH STENTING. THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. Manufacturer Narrative: CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING CCPD THERAPY WHEN HE EXPIRED. THE PDRN ATTRIBUTED CAUSALITY TO THE PATIENTS¿ SIGNIFICANT CARDIAC COMORBID CONDITIONS. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (INCLUDING SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. PER THE PDRN, THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22568769&pc=MVK | 3008642652-2025-07051 | 2025-06-09 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-22 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT HAD PASSED AWAY ON (B)(6) 2025. THE PATIENT'S FAMILY MEMBER REPORTED THAT THE PATIENT WAS NOT WEARING THE LIFEVEST AT THE TIME OF PASSING DUE TO REPORTED BLISTERS AND SORES. THE PATIENT'S FAMILY MEMBER REPORTED THAT THE PATIENT WAS VERY SICK AND THE DEATH WAS NOT CARDIAC RELATED. REPORTING OUT OF AN ABUNDANCE OF CAUTION AS THE CAUSE OF THE BLISTERS AND SORES ARE UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22568788&pc=MVK | 3008642652-2025-07052 | 2025-07-11 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-22 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 04:57:14 ON (B)(6) 2025. AT 07:10:09, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. AT 07:10:47, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM DEGRADING TO ASYSTOLE WITH UNCONDUCTED P WAVES AND MOTION ARTIFACT. AT 07:12:59, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 07:13:33, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS AF FROM 30-75 BPM WITH CPR/MOTION ARTIFACT DEGRADING TO ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 07:30:05 ON (B)(6) 2025. Manufacturer Narrative: THE MONITOR AND ELECTRODE BELT HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION. THE REVIEW OF THE DATA INDICATED THAT THE DEVICE POWERED ON NORMALLY AND WAS ABLE TO ACQUIRE THE PATIENT¿S ECG SIGNAL ON THE LAST DAY OF USE CAPTURED IN THE DATA DOWNLOAD. NO DEFICIENCIES ALLEGED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22568870&pc=NVN | 2017865-2025-94663 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-22 04:00 | NVN | TENDRIL SDX LEAD | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED FOR ANALYSIS. ELECTRICAL TESTING THAT COULD BE MEASURED DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22568878&pc=LWP | 2017865-2025-94662 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-22 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22569113&pc=BZD | 2518422-2025-108088 | 2025-06-26 04:00 | Death | RESPIRONICS, INC. | 2025-07-22 04:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER BECAME AWARE OF A DS2ADV AUTO CPAP STATING THAT THE PATIENT'S SIGNIFICANT OTHER ADVISED THAT THE PATIENT SADLY PASSED AWAY AND THAT THERE IS NO ALLEGATION THAT THE DEVICE HAS CONTRIBUTED TO THE DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22569463&pc=DSQ | 3004582654-2025-00040 | 2025-06-20 04:00 | Death | BERLIN HEART GMBH | 2025-07-22 04:00 | DSQ | EXCOR® | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P160035 |  | 1 | Event Description: BERLIN HEART GMBH CLINICAL AFFAIRS WAS INFORMED BY THE POLAND DISTRIBUTOR ON 2025-06-26 THAT A PATIENT SUFFERED A MINOR HEMORRHAGIC STROKE ON (B)(6) 2025, FOLLOWED BY A MASSIVE HEMORRHAGIC STROKE ON (B)(6) 2025, WHICH UNFORTUNATELY RESULTED IN DEATH. THE PATIENT WAS BEING SUPPORTED WITH AN EXCOR BLOOD PUMP PU VALVES; 15 ML IN/OUT; Ø 9 MM. ACCORDING TO THE SITE, THE EXCOR BLOOD PUMP MAINTAINED COMPLETE FILLING AND EJECTION AND NO DEPOSITS IN THE EXCOR BLOOD PUMP OR ANTICOAGULATION PROBLEM WERE REPORTED. Manufacturer Narrative: THE LVAD EXCOR BLOOD PUMP PU VALVES; 15 ML IN/OUT; Ø 9 MM; SN (B)(6) WAS IN USE ON THE PATIENT FROM (B)(6) 2025 UNTIL IT WAS REMOVED FROM THE PATIENT ON (B)(6) 2025. THE EVENT OCCURRED ON (B)(6) 2025, WHICH IS (31 DAYS) AFTER THE PUMP WAS PLACED ON THE PATIENT. WE HAVE REVIEWED THE PRODUCTION RECORDS OF THE EXCOR BLOOD PUMP AND THE PUMP WAS PRODUCED ACCORDING TO OUR SPECIFICATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22569673&pc=DSQ | 3004582654-2025-00041 | 2025-05-08 04:00 | Death | BERLIN HEART GMBH | 2025-07-22 04:00 | DSQ | EXCOR® | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P160035 |  | 1 | Event Description: BERLIN HEART GMBH CLINICAL AFFAIRS WAS INFORMED BY THE CLINIC ON 2025-07-03 THAT A PATIENT WAS PRESENTED WITH A HEADACHE, AND A CT SCAN WAS PERFORMED. A SMALL HEMORRHAGIC STROKE WAS IDENTIFIED ON (B)(6) 2025. ANTICOAGULATION WAS STOPPED BUT OVER THE NEXT DAYS THE BLEEDING INCREASED, AND THE PATIENT WAS DECLARED BRAIN DEAD AND DIED ON (B)(6) 2025 DUE TO MASSIVE CEREBRAL HEMORRHAGE. THE PATIENT WAS BEING SUPPORTED WITH AN EXCOR BLOOD PUMP PU VALVES; 50 ML IN/OUT; Ø 12 MM. NO DEPOSITS IN THE EXCOR BLOOD PUMP WERE RECORDED. Manufacturer Narrative: THE LVAD EXCOR BLOOD PUMP PU VALVES; 50 ML IN/OUT; Ø 12 MM; SN (B)(6) WAS IN USE ON THE PATIENT FROM (B)(6) 2025 UNTIL IT WAS REMOVED FROM THE PATIENT ON (B)(6) 2025. THE EVENT OCCURRED ON 2025-05-08, WHICH IS (63 DAYS) AFTER THE PUMP WAS PLACED ON THE PATIENT. WE HAVE REVIEWED THE PRODUCTION RECORDS OF THE EXCOR BLOOD PUMP AND THE PUMP WAS PRODUCED ACCORDING TO OUR SPECIFICATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22571229&pc=NPU | 2015691-2025-05960 | 2025-06-30 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-22 04:00 | NPU | SAPIEN 3 TRANSCATHETER HEART VALVE | Migration or Expulsion of Device; Perivalvular Leak | Cardiogenic Shock; Foreign Body In Patient | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS AFFILIATES IN ITALY, DURING A PROCEDURE INVOLVING THE IMPLANTATION OF A 29MM SAPIEN 3 VALVE INTO A NON-EDWARDS RING, A MILD TO MODERATE PARAVALVULAR LEAK WAS OBSERVED FOLLOWING VALVE DEPLOYMENT. NO CORRECTIVE ACTIONS WERE TAKEN AT THAT TIME. ONE WEEK LATER, THE PATIENT DEVELOPED A SEVERE PARAVALVULAR LEAK, NECESSITATING RE-INTERVENTION. DURING THE VALVULOPLASTY, THE VALVE EMBOLIZED INTO THE LEFT ATRIUM, AND THE PATIENT UNFORTUNATELY PASSED AWAY DUE TO CARDIOGENIC SHOCK. THE IMPLANT POSITION IS CURRENTLY UNKNOWN.Event Description: AS REPORTED BY EDWARDS AFFILIATES IN ITALY, A TRANSCATHETER MITRAL VALVE REPLACEMENT (TMVR) PROCEDURE WAS PERFORMED USING A 29 MM SAPIEN 3 VALVE IMPLANTED IN A NON-EDWARDS ANNULOPLASTY RING IN THE MITRAL POSITION. FOLLOWING VALVE DEPLOYMENT, A MILD TO MODERATE PARAVALVULAR LEAK (PVL) WAS OBSERVED. NO IMMEDIATE CORRECTIVE ACTIONS WERE TAKEN AT THAT TIME. APPROXIMATELY ONE WEEK POST-PROCEDURE, THE PATIENT DEVELOPED A SEVERE PARAVALVULAR LEAK AND REQUIRED EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SUPPORT. A RE-INTERVENTION WAS PLANNED, CONSISTING OF A VALVULOPLASTY FOLLOWED BY THE PLACEMENT OF AN AMPLATZER OCCLUSION DEVICE. DURING THE VALVULOPLASTY, THE IMPLANTED VALVE BECAME UNSTABLE AND EMBOLIZED INTO THE LEFT ATRIUM. SURGICAL INTERVENTION WAS CONSIDERED; HOWEVER, THE PATIENT EXPERIENCED CARDIOGENIC SHOCK AND PASSED AWAY BEFORE ANY FURTHER ACTION COULD BE TAKEN. Manufacturer Narrative: THE VALVE SERIAL NUMBER WAS PROVIDED. THEREFORE THE D4 AND H4 FIELDS HAVE BEEN UPDATED. Manufacturer Narrative: UPDATE TO B5. CORRECTION TO D2, D3, AND H6; IMPACT CODE. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PARAVALVULAR LEAK (PVL) IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE/PRE-STENT AND THE LANDING ZONE, AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE. SOME PVL IS NOT UNCOMMON POST-DEPLOYMENT. MANY CASES ARE MILD TO MODERATE, AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS. OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION. THE PATIENT SCREENING MANUAL AND THE PROCEDURE DIDACTIC IDENTIFY SEVERAL PROCEDURAL AND ANATOMICAL FACTORS WHICH COULD CONTRIBUTE TO PVL, INCLUDING DEVICE MALPOSITION, INACCURATE MEASUREMENT OF THE LANDING ZONE, UNEVEN DISTRIBUTION OF CALCIUM ON THE LANDING ZONE, BULKY OR SEVERE CALCIFICATION, AN ELLIPTICAL ANNULUS SHAPE, AND VALVE UNDER-SIZING. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). THE DEVICE TRAINING MANUALS INSTRUCTS THE OPERATOR ON PROPER IMAGING SCREENING REQUIREMENTS, INCLUDING THE USE OF GOOD QUALITY ECHOCARDIOGRAPHY AND/OR COMPUTED TOMOGRAPHY (CT) TO APPROPRIATELY MEASURE THE LANDING ZONE, ASSESS THE CONTENT AND DISTRIBUTION OF CALCIUM, AND OTHER PATIENT ANATOMICAL FACTORS. MULTIPLE IMAGING MODALITIES SHOULD BE CONSIDERED DURING VALVE SIZE SELECTION. CONTRAINDICATIONS, IMPORTANT CONSIDERATIONS WHEN ASSESSING THE VALVE, AND CHOOSING THE PROPER TRANSCATHETER HEART VALVE ARE ALSO DISCUSSED. THE DEVICE TRAINING MANUALS ALSO INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING PATIENTS SCREENING AND PROCEDURAL CONSIDERATIONS. DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES ARE ALSO INCLUDED. PER THE INSTRUCTIONS FOR USE (IFU), VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION OF THE DEVICE, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, LOSS OF PACING CAPTURE, A NARROW SINOTUBULAR JUNCTION (IN AORTIC POSITION PROCEDURES), RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE LANDING ZONE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THAT UNKNOWN PROCEDURAL FACTORS MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22571577&pc=LGW | 3006630150-2025-05753 | 2025-07-01 04:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-07-22 04:00 | LGW | INFINION? PRO | Adverse Event Without Identified Device or Use Problem | Abscess; Bacterial Infection | P030017 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED AN INFECTION AT THE INCISION SITE. THE PATIENT WAS ADMINISTERED WITH ANTIBIOTICS AND UNDERWENT A SPINAL CORD STIMULATOR (SCS) SYSTEM EXPLANT PROCEDURE. THE EXPLANTED DEVICE COMPONENTS WERE NOT RETURNED.Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED AN INFECTION AT THE SPINAL CORD STIMULATION (SCS) LEAD INCISION SITE. THE PATIENT WAS HOSPITALIZED AND UNDERWENT A PROCEDURE IN WHICH THE ENTIRE SCS SYSTEM WAS EXPLANTED. IN ADDITION, THE PATIENT WAS ADMINISTERED ANTIBIOTICS. THE EXPLANTED DEVICES WILL NOT BE RETURNED AS THEY WERE DISCARDED BY THE MEDICAL FACILITY. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS NOT DISCHARGED FOLLOWING THE EXPLANT PROCEDURE. THE PATIENT WAS TRANSFERRED TO ANOTHER MEDICAL FACILITY, AND ANOTHER PHYSICIAN TOOK OVER HER CARE AND TREATED HER EPIDURAL ABSCESS. THE PATIENT PASSED AWAY. THE PATIENT WAS NO LONGER IMPLANTED WITH ANY SCS DEVICES AT THE TIME OF PASSING. THE SPECIFIC REASON FOR PASSING IS UNKNOWN, HOWEVER THE PHYSICIAN ASSESSED THE PASSING WAS RELATED TO AN EPIDURAL ABSCESS THAT WAS THE RESULT OF THE SPINAL CORD STIMULATION SYSTEM. Manufacturer Narrative: BLOCK D2B; QRB ADDITIONAL SUSPECT MEDICAL DEVICE COMPONENT INVOLVED IN THE EVENT: PRODUCT FAMILY: SCS-LINEAR LEADS. UPN: M365SC2318700. MODEL: SC-2318-70. SERIAL: (B)(6). BATCH: (B)(6). UDI: (B)(4). PRODUCT FAMILY: SCS-IPG-R-MRI. UPN: M365SC12320. MODEL: SC-1232. SERIAL: (B)(6). BATCH: (B)(6). UDI: (B)(4). Manufacturer Narrative: ADDITIONAL SUSPECT MEDICAL DEVICE COMPONENT INVOLVED IN THE EVENT: PRODUCT FAMILY: SCS-LINEAR LEADS. UPN: M365SC2318700. MODEL: SC-2318-70. SERIAL: (B)(6). BATCH: 5004773. UDI: (B)(4). PRODUCT FAMILY: SCS-IPG-R-MRI. UPN: M365SC12320. MODEL: SC-1232. SERIAL: (B)(6). BATCH: 786946. UDI: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22571900&pc=NVN | 2017865-2025-94711 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-22 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, A COMPLETE LEAD WAS RETURNED FOR ANALYSIS. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22571902&pc=LWP | 2017865-2025-94709 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-22 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22571906&pc=NVN | 2017865-2025-94716 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-22 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A COMPLETE LEAD WAS RETURNED IN ONE PIECE. VISUAL AND X-RAY EXAMINATIONS OF THE LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22571924&pc=DSQ | 2916596-2025-04830 | 2025-07-14 04:00 | Death | THORATEC CORPORATION | 2025-07-22 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Airway Obstruction; Respiratory Failure; Bronchial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEVERE RESPIRATORY FAILURE AFTER BRONCHIAL BLEEDING WITH BLOCKING OF AIRWAYS. THE DEATH WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22553086&pc=MIH | 9612164-2025-03531 | 2025-07-04 04:00 | Death | MEDTRONIC IRELAND | 2025-07-21 04:00 | MIH | VALIANT CAPTIVIA - FF | Positioning Problem | Pain; Perforation of Vessels | P100040 |  | 1 | Event Description: A VALIANT CAPTIVIA STENT GRAFT WAS INTENDED TO BE IMPLANTED DURING THE ENDOVASCULAR TREATMENT OF A 200MM DISSECTION. DURING THE INDEX PROCEDURE, IT WAS REPORTED THAT THE THORACIC AORTIC STENT COULD NOT BE ADVANCED TO THE DESIGNATED POSITION IN THE ARCH AND COULD NOT BE IMPLANTED SUCCESSFULLY. THE LOWER DELIVERY SYSTEM WAS POSITIONED PROXIMALLY, AND A VASCULAR CAPTURE SYSTEM WAS INSERTED VIA THE AXILLARY ARTERY APPROACH TO CAPTURE THE TIP AND PULL IT PROXIMALLY INTO THE AORTIC ARCH. THE PHYSICIAN ALSO ATTEMPTED TO ACCESS THE DESIRED AREA USING A DOUBLE-HARDENED GUIDEWIRE TO ADDRESS THE TWISTING OF THE BLOOD VESSEL; HOWEVER, THE STENT STILL COULD NOT REACH THE INTENDED POSITION. AFTER REPEATED ATTEMPTS, THE PATIENT REPORTED CHEST PAIN THAT COULD NOT BE ALLEVIATED, LEADING TO THE DECISION TO HALT THE PROCEDURE. THE STENT WAS REMOVED, AND THE PROCEDURE WAS ABORTED. IT WAS NOTED THAT THE PATIENT'S CHEST PAIN WAS BEARABLE AND SLIGHTLY RELIEVED AFTERWARD. IT WAS REPORTED THAT ONE DAY LATER THE PATIENT EXPIRED. PER THE PHYSICIAN, THE CAUSE OF THE POSITIONING DIFFICULTIES REMAINS UNDETERMINED; HOWEVER, VASCULAR TORTUOSITY IS SUSPECTED. IT IS BELIEVED THAT THE CAUSE OF DEATH MAY HAVE BEEN DUE TO RUPTURE. THE PHYSICIAN ALSO NOTED THAT THE VASCULAR RESPONSE DURING THE ADVANCEMENT OF THE DELIVERY SYSTEM, COMBINED WITH THE INABILITY TO EFFECTIVELY TREAT THE PATIENT, WERE CONTRIBUTING FACTORS TO THE PATIENT'S DEATH. Manufacturer Narrative: DEVICE PHOTO EVALUATION SUMMARY: 6 STILL IMAGES WERE RECEIVED FOR ANALYSIS. 2 IMAGES DEPICTED A VALIANT SHELF CARTON AND LABEL. LABELLING INFORMATION MATCHES THAT REPORTED IN THE COMPLAINT FILE. 4 IMAGES WERE RECEIVED DEPICTING THE DEVICE ALONGSIDE 3 GUIDEWIRES. A BEND IS EVIDENT TO THE GRAFT COVER OVER THE STENT GRAFT. DEFORMATION IS EVIDENT TO THE GRAFT COVER OVER THE STENT STOP. NO DEFORMATION EVIDENT TO HANDLE OR SCREW GEAR. THE REPORTED POSITIONING DIFFICULTIES COULD BOT BE CONFIRMED FROM THE IMAGES PROVIDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ANALYSIS CONCLUSION THE REPORTED DIFFICULTY WITH DEVICE POSITIONING COULD NOT BE FULLY CONFIRMED BASED ON THE LIMITED IMAGING AVAILABLE. WHILE PROCEDURAL ANGIOGRAM VIDEOS SHOWING ATTEMPTS TO ADVANCE THE DEVICE WERE NOT AVAILABLE FOR REVIEW, THE SEVERE TORTUOSITY WITH THREE DISTINCT AREAS OF SIGNIFICANT ANGULATIONS (AT THE DISTAL ARCH/PROXIMAL DESCENDING AORTA, PROXIMAL/MIDDLE, AND MIDDLE/DISTAL DESCENDING AORTA) LIKELY CONTRIBUTED TO THE REPORTED POSITIONING DIFFICULTY. THE REPORTED PAIN, VESSEL RUPTURE, AND FATAL OUTCOME COULD NOT BE ASSESSED FROM THE AVAILABLE IMAGES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22553776&pc=OZD | 1220648-2025-30525 | 2025-07-03 04:00 | Death | ABIOMED, INC. | 2025-07-21 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT WITH ANTERIOR ST-ELEVATION MYOCARDIAL INFARCTION WAS FOUND TO BE IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK. THE DECISION WAS MADE TO PLACE AN IMPELLA 5.5. DURING PLACEMENT, BEFORE THE GRAFT COULD BE ANASTOMOSED, THE PATIENT DECOMPENSATED. THE PATIENT WAS CANNULATED FOR PERIPHERAL EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WHEN THE RIGHT VENTRICLE FREE WALL AND INTERVENTRICULAR SEPTUM RUPTURED SIMULTANEOUSLY. THE PATIENT WAS CRASHED ONTO A PERIPHERAL VENOARTERIAL ECMO AND AN EXTENSIVE REPAIR SURGERY WAS COMPLETED. ECMO WAS RECONFIGURED TO CENTRAL ECMO WITH LEFT VENTRICLE (LV) VENT EXTENDING FROM THE RIGHT UPPER PULMONARY VEIN INTO THE LV THROUGH THE MITRAL VALVE. TWO DAYS POST PROCEDURE, THE DECISION WAS MADE TO BRING THE PATIENT BACK TO HYBRID OPERATING ROOM FOR IMPELLA 5.5 PLACEMENT. THE IMPELLA 5.5 WAS SUCCESSFULLY PLACED. DURING SUPPORT, IT WAS REPORTED THAT THE PATIENT HAD COMPUTED TOMOGRAPHY IMAGES TAKEN AND THERE WERE THREE AREAS OF INFARCT AND A BRAIN BLEED NOTED. ULTIMATELY CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22554012&pc=OZD | 1220648-2025-30530 | 2025-07-03 04:00 | Death | ABIOMED, INC. | 2025-07-21 04:00 | OZD | IMPELLA CP | Positioning Failure | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AN IMPELLA CP WAS PLACED WITH BEST PRACTICES. THE PHYSICIAN FELT LIKE THE IMPELLA CP WAS UNDERNEATH THE PAPILLARY MUSCLE AND WANTED TO REPOSITION. WHILE ATTEMPTING TO REPOSITION, THE IMPELLA CP WAS PULLED BACK INTO THE AORTA AND WAS UNABLE TO RECROSS. THE PHYSICIAN WAS ADVICED TO USE A NEW IMPELLA BUT DECIDED TO USE THE SAME IMPELLA CP FOR PLACEMENT AGAIN. THE IMPELLA CP WAS RINSED IN HEPARINIZED SALINE. THE PUMP WAS EXAMINED FOR A CLOT OR DEFECT AND WAS REINSERTED WITH NO ISSUES. IMPELLA CP TURNED ON AND FLOWS OF 3.5 WERE ACHIEVED. ON DAY FOUR OF IMPELLA CP SUPPORT, AN ECHOCARDIOGRAM SHOWED A THROMBUS IN THE LEFT VENTRICLE. THE PATIENT WAS NOT A CANDIDATE FOR ADVANCED THERAPY AND CARE WAS WITHDRAWN. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION, SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." Manufacturer Narrative: THE INVESTIGATION FOR THE THROMBUS AND POSITIONING ISSUE HAS BEEN COMPLETED. PRODUCT WAS NOT RETURNED FOR REVIEW. BASED ON CLINICAL DESCRIPTION, THE ROOT CAUSE OF POSITIONING ISSUE WAS MOST LIKELY USE RELATED SINCE THE PUMP PULLED OUT OF THE LV WHILE REPOSITIONING OF THE PUMP. WITHOUT SUFFICIENT CLINICAL DETAILS, THE ROOT CAUSE OF THE THROMBUS WAS UNABLE TO BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22554133&pc=BTO | 3012307300-2025-08601 | 2025-06-01 04:00 | Death | SMITHS MEDICAL CZECH REPUBLIC A. S | 2025-07-21 04:00 | BTO | PORTEX BLUSELECT SUCTIONAID CUFFED TRACHEOSTOMY | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K173384 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED DUE TO PROBLEMS WITH THE DEVICE. THE BALLOON WAS DEFLATING, AND THE PATIENT WAS MICROASPIRATING, WHICH CONTRIBUTED TO THE DEATH. ACCORDING TO THE CAREGIVER, THE PATIENT ASPIRATED MULTIPLE TIMES, ULTIMATELY RESULTING IN DEATH BY BRONCHOASPIRATION CAUSED BY A CANNULA OBSTRUCTING THE AIRWAY. Manufacturer Narrative: NO PRODUCT WAS RETURNED. THE REPORTED ISSUE CANNOT BE CONFIRMED AS NO PRODUCT WAS RETURNED FOR INVESTIGATION. IF THE PRODUCT IS RETURNED, THIS COMPLAINT WILL BE REOPENED FOR FURTHER INVESTIGATION. A DEVICE HISTORY RECORD (DHR) REVIEW WAS CONDUCTED WHICH INDICATED ALL INSPECTIONS WERE COMPLETED AND NO ISSUES WERE NOTED DURING MANUFACTURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22554471&pc=GDW | 1219930-2025-03103 | 2025-04-26 04:00 | Death | COVIDIEN SURGICAL | 2025-07-21 04:00 | GDW | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY INCLUDING 192 CASES OF LAPAROSCOPIC RIGHT COLECTOMY FOR NEOPLASIA USING EITHER EX TRACORPOREAL ANASTOMOSIS OR INTRACORPOREAL ANASTOMOSIS WERE COMPLETED BETWEEN APRIL 2013 AND JANUARY 2024. TRANSECTION OF THE COLON WAS ACHIEVED USING 60MM STAPLE LOADS ATTACHED TO IDRIVE, SIGNIA, OR A COMPETITOR STAPLER PRODUCT. CONVERSIONS TO OPEN (17 CASES) AND 90 DAYS POSTOPERATIVE MORTALITIES (3 CASES) WERE MENTIONED WITH UNDISCLOSED CAUSES. Manufacturer Narrative: D10 CONCOMITANT PRODUCT: UNKNOWN EGIA SU, UNKNOWN ENDO GIA SULU (LOT#UNKNOWN) MINIMIZING INCISIONAL HERNIA: INTRACORPOREAL ANASTOMOSIS MAKES THE DIFFERENCE AFTER LAPAROSCOPIC RIGHT COLECTOMY: ERNESTO DE GIULIO, INTERNATIONAL JOURNAL OF COLORECTAL DISEASE (2025) 40:112, HTTPS://DOI.ORG/10.1007/S00384-025-04903-Z, PUBLISHED ONLINE: 08 MAY 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22554479&pc=OZD | MW5172995 | 2023-10-15 04:00 | Death | ABIOMED, INC. | 2025-07-21 04:00 | OZD | P9 IMPELLA SWAN CL | Adverse Event Without Identified Device or Use Problem | Tachycardia; Cardiogenic Shock; Obstruction/Occlusion; Pallor; Respiratory Failure |  |  | 1 | Event Description: ON (B)(6) 2023, 76-YEAR-OLD (B)(6) WALKED INTO (B)(6) HOSPITAL, LOCATED AT (B)(6) COMPLAINING OF SHORTNESS OF BREATH. WITHIN MINUTES HE WAS DIAGNOSED WITH A CONGESTIVE HEART FAILURE EXACERBATION AND WAS SCHEDULED FOR A CATHETERIZATION THE NEXT DAY DUE TO FLUID SURROUNDING HEART AND LUNGS. HE WAS GIVEN THE DRUG LASIX TO HELP. IT WAS DELAYED ANOTHER DAY DUE TO WAITING FOR FLUID TO DRAIN. DR. (B)(6) PERFORMED THE CATHETERIZATION AROUND 11AM AND CAME TO THE CONCLUSION THERE WERE NO BLOCKAGES, AND IT WAS NOT A "HEART PROBLEM." HIS HOSPITALIZATION THEN BECAME COMPLICATED BY CARDIOGENIC SHOCK AND TRANSFERRED TO INTENSIVE CARE UNIT. ON (B)(6), NEPHROLOGY WAS CONSULTED, DIURETICS WERE STOPPED, AND THEY ARE STARTED CRRT AND PRESSOR THERAPY. THIS DIALYSIS WAS SUPPOSED TO BE TEMPORARY. HE HAD SUSTAINED V TACH ON (B)(6) REQUIRING BRIEF CPR, SUBSEQUENT DEFIBRILLATION, THEN INTUBATION; ALSO TEMPORARY. HE WENT IN FOR ANOTHER CATHETERIZATION WITH DR. (B)(6), WHO FOUND FIVE BLOCKAGES AND PLACED FIVE STENTS. HE THEN PLACED THE IMPELLA SUPPORTED PCI. ON (B)(6) HIS COLORING WAS GOOD AND ACCORDING TO THE NURSE ON DUTY, HE WAS TRENDING POSITIVE. WHEN I ARRIVED ON (B)(6) HIS COLORING WAS GRAY, HIS BLOOD WAS PURPLE/BROWN, AND I HAD NOTICED THAT HIS BANK OF MONITORS HAD BEEN CHANGED AND WERE PLUGGED IN ON THE OTHER SIDE OF THE ROOM. SOON AFTER, THE DOCTORS WERE LEADING ME BY THE ARM TO A QUIET PLACE TO TELL ME I NEEDED TO MAKE A DECISION OF WHETHER OR NOT TO TAKE HIM OFF OF LIFE SUPPORT, AS HE WAS GETTING WORSE, AND THERE WAS NO HOPE FOR HIM. AFTER CONSULTING WITH FAMILY, THE DECISION WAS MADE TO TAKE HIM OFF OF LIFE SUPPORT. HE DIED WITHIN MINUTES ON (B)(6) 2023.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22554672&pc=GEI | 9610614-2025-00043 | 2025-06-25 04:00 | Death | ERBE ELEKTROMEDIZIN GMBH | 2025-07-21 04:00 | GEI | ERBE VIO 3 | Appropriate Term/Code Not Available | Hemorrhage/Bleeding | K190823 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT INCIDENT OCCURRED WITH THE ELECTROSURGICAL UNIT (ESU/GENERATOR) AND AN ERBE TRISECT RAPIDE INSTRUMENT (PART NUMBER 21150-402, LOT NUMBER WO433553.) DURING A TOTAL PELVIC EXENTERATION (TPE) PROCEDURE (NOTE: THE ERBE TRISECT RAPIDE INSTRUMENT IS NOT CLEARED FOR DISTRIBUTION IN THE U.S. AND IS NOT DISTRIBUTED IN THE U.S.). NO INFORMATION WAS PROVIDED REGARDING ANY OTHER ACCESSORIES USED IN THE OPERATION. ALSO, THE ESU SETTINGS EMPLOYED WERE NOT CONVEYED TO ERBE. DURING THE POSTOPERATIVE PERIOD, A DELAYED PERFORATION OF THE AORTA WAS DISCOVERED WHICH RESULTED IN MASSIVE BLEEDING AND THE PATIENT DIED (NOTE: IT APPEARS THAT A THERMAL LESION ON THE AORTA OCCURRED DURING THE TPE PROCEDURE WHILE USING THE ERBE TRISECT RAPIDE INSTRUMENT.). Manufacturer Narrative: THE ESU WAS THOROUGHLY INSPECTED/TESTED (NOTE: THE ERBE TRISECT RAPIDE INSTRUMENT WAS DISPOSED OF AFTER THE PROCEDURE AND WAS NOT AVAILABLE FOR AN EVALUATION.). THE GENERATOR WAS FOUND TO BE FUNCTIONING AS INTENDED. THE EVALUATION INCLUDED AN ELECTRICAL SAFETY CHECK, A FUNCTIONAL CHECK OF EACH OF THE EQUIPMENT'S FEATURES, AND A POWER OUTPUT CHECK. THE UNIT WAS/IS WITHIN SPECIFICATIONS AND ALL FEATURES WERE/ARE WORKING PROPERLY. IN ADDITION, NO ANOMALIES WERE FOUND IN THE DEVICE HISTORY RECORD (DHR) OF THE ESU OR THE TRISECT RAPIDE INSTRUMENT.). IN CONCLUSION, NO ERBE EQUIPMENT PROBLEM WAS FOUND THAT WOULD HAVE CAUSED OR CONTRIBUTED TO THE INCIDENT. BASED UPON THE REPORTED INFORMATION, MOST LIKELY THE THERMAL INSULT TO THE AORTA WAS CAUSED BY INADVERTENT CONTACT BETWEEN THE BRANCHES OF THE TRISECT RAPIDE INSTRUMENT AND THE AORTA. THE USER MANUAL OF THE ESU WARNS THAT BURNS OR FIRES MAY OCCUR DUE TO ACCIDENTALLY ACTIVATING AN INSTRUMENT. ALSO, THE MANUAL WARNS THAT UNWANTED TISSUE TRAUMA CAN OCCUR IF TISSUE COMES INTO CONTRACT WITH A INSTRUMENT THAT IS STILL HOT. THE TRISECT RAPIDE INSTRUMENT NOTES ON USE WARNS THAT TISSUE IN THE JOINT AREA MAY BE THERMALLY DAMAGED WHEN THE INSTRUMENT IS ACTIVATED WITH THE BRANCHES WIDE OPEN AND THAT TISSUE CONTACT WITH THE JOINT AREA OF THE INSTRUMENT SHOULD BE AVOIDED IN THIS CASE. ERBE USA, INC. IS NOW CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22554746&pc=BTO | 3012307300-2025-08620 | 2025-02-12 05:00 | Death | ICU MEDICAL, INC. | 2025-07-21 04:00 | BTO | PORTEX TUBES PDT UNIPERC | Improper or Incorrect Procedure or Method; Installation-Related Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE BECAME DISPLACED FOLLOWING SURGICAL INSERTION IN ONE OF THE ICU PATIENTS. THE MECHANISM/'CLIP' USED TO SECURE THE TRACHEOSTOMY TUBE ITSELF TO THE FLANGE CAUSED A POTENTIAL SAFETY CONCERN. THE TRACHEOSTOMY WAS MISTAKENLY INSERTED UPSIDE DOWN, WITH THE "CLIP" FACING UPWARDS. THIS RESULTED IN IT BECOMING KNOCKED OPEN WHEN THE PATIENT WAS MOVED. THE TRACHEOSTOMY BECAME DISPLACED AND THE PATIENT RAPIDLY DESATURATED DIFFICULTIES REPLACING IT THROUGH A FRESH STOMA. THE TRACHEOSTOMY BECAME DISPLACED THE MORNING AFTER IT WAS SURGICALLY PLACED ON (B)(6) 2025. IT WAS ALSO DIFFICULT TO RE-INTUBATE HER, BUT SURGICAL ASSISTANCE WAS ON HAND TO PLACE A SMALL (SIZE 6) ENDOTRACHEAL TUBE THROUGH THE TRACHEOSTOMY STOMA PRIOR TO IT BEING FORMALLY REPLACED IN THEATRE. THE PATIENT SUBSEQUENTLY DIED BUT HER DEATH WAS NOT DIRECTLY RELATED TO THE DISPLACED TRACHEOSTOMY. THE STATUS OF THE EVENT IS RESOLVED. THERE WAS PATIENT INVOLVEMENT AND ADVERSE HARM REPORTED. Manufacturer Narrative: B2: DATE OF DEATH UNKNOWN. H3: NO DEVICE OR DEVICE PHOTO WAS RETURNED FOR INVESTIGATION. ALTHOUGH THE COMPLAINT MENTIONS A PHOTO, ONLY A CATALOGUE IMAGE WAS PROVIDED, NOT A PHOTO OF THE ACTUAL USED PRODUCT. DUE TO THIS, NO ROOT CAUSE WAS DETERMINED. THE DEVICE HISTORY REVIEW OF THE REPORTED LOT NUMBER WAS NOT POSSIBLE AS NO LOT NUMBER WAS PROVIDED. IF THE PRODUCT IS RETURNED, THIS COMPLAINT WILL BE REOPENED FOR FURTHER INVESTIGATION. THE CUSTOMER ALSO HAD REPORTED A POTENTIAL ISSUE RELATED TO THE FLANGE LOCKING MECHANISM. HOWEVER, IT WAS ALSO STATED THAT THE PRODUCT WAS INSERTED INCORRECTLY, WHICH IS THE MOST LIKELY ROOT CAUSE OF THE REPORTED INCIDENT. IT WAS FURTHER NOTED THAT THE TRACHEOSTOMY DISPLACEMENT WAS NOT CONSIDERED TO BE THE DIRECT CAUSE OF THE PATIENT'S DEATH. Manufacturer Narrative: B2: DEATH DATE IS UNKNOWN. DEVICE EVALUATION: INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22554870&pc=DSQ | 2916596-2025-04785 | 2020-11-30 05:00 | Death | THORATEC CORPORATION | 2025-07-21 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2020. THE CAUSE OF DEATH WAS UNKNOWN, NO AUTOPSY WOULD BE PERFORMED, AND THE PUMP WAS OPERATING AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22554962&pc=OZD | 1220648-2025-30537 | 2025-07-05 04:00 | Death | ABIOMED, INC. | 2025-07-21 04:00 | OZD | IMPELLA CP | Positioning Failure | Insufficient Information | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WITH ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT. AFTER REPOSITIONING, THE AUTOMATED IMPELLA CONTROLLER (AIC) DISPLAYED AN ¿IMPELLA IN AORTA¿ ALARM. ATTEMPTS TO REPOSITION THE DEVICE WERE UNSUCCESSFUL, AND THE IMPELLA PUMP WAS SUBSEQUENTLY EXPLANTED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22555058&pc=GKT | 1722028-2025-00163 | 2025-02-18 05:00 | Death | TERUMO BCT | 2025-07-21 04:00 | GKT | RIKA | Insufficient Information | No Clinical Signs, Symptoms or Conditions | BK241097 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT THE GRANDMOTHER OF A DONOR NOTIFIED ON 02/21/2025 THAT A DONOR HAS PASSED AWAY ON (B)(6) 2025 (SUSPECTED) AFTER DONATING ON A RIKA DEVICE ON (B)(6) 2025. DONOR ID: (B)(6) THE COLLECTION SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. THIS REPORT IS BEING FILED DUE TO PATIENT DEATH, ALTHOUGH PER CURRENT INFORMATION THERE IS NO DETECTABLE MALFUNCTION WITH THE TERUMO BCT DEVICE OR ALLEGATION OF A MALFUNCTION.Event Description: THE CUSTOMER REPORTED THAT THE GRANDMOTHER OF A DONOR NOTIFIED ON 02/21/2025 THAT A DONOR HAS PASSED AWAY ON (B)(6) 2025 (SUSPECTED) AFTER DONATING ON A RIKA DEVICE ON (B)(6) 2025. DONOR ID: (B)(6) THE COLLECTION SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. THIS REPORT IS BEING FILED DUE TO PATIENT DEATH, ALTHOUGH PER CURRENT INFORMATION THERE IS NO DETECTABLE MALFUNCTION WITH THE TERUMO BCT DEVICE OR ALLEGATION OF A MALFUNCTION. Manufacturer Narrative: INVESTIGATION: PER THE CUSTOMER, NO SERVICE WAS INITIALLY REQUESTED AS THERE WERE NO ISSUES WITH THE DEVICE. INVESTIGATION IS IN PROCESS, A FOLLOW-UP REPORT WILL BE PROVIDED. Manufacturer Narrative: THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION IN H.6 AND H.11. INVESTIGATION: PER THE CUSTOMER, NO SERVICE WAS INITIALLY REQUESTED AS THERE WERE NO ISSUES WITH THE DEVICE. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) FOR THIS UNIT SHOWED NO IRREGULARITIES DURING MANUFACTURING THAT WERE RELEVANT TO THIS ISSUE. A SERVICE TECHNICIAN CHECKED OUT THE DEVICE AT THE CUSTOMER SITE, THE TECH COMPLETED A FUNCTIONAL CHECKOUT AND THE INVESTIGATION CONCLUDED THAT THE DEVICE WAS WORKING AS INTENDED AND NOTHING ANOMALOUS WAS FOUND. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE LOG FILE SHOWS THAT NO UNPLANNED OPERATOR INTERACTIONS OCCURRED DURING THE 30-MINUTE PROCEDURE. THE AC FLUID DETECTOR, DONOR LINE FLUID DETECTOR, LINE SENSOR, APS, RETURN CPS, AND DRAW CPS SIGNALS WERE ANALYZED, AND NOTHING ANOMALOUS WAS FOUND. PER COMMUNICATION WITH THE CUSTOMER, THE FUNERAL HOME AND CORONER OFFICE WOULD NOT RELEASE ANY INFORMATION. INVESTIGATION IS IN PROCESS, A FOLLOW-UP REPORT WILL BE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22555506&pc=NGV | 2124215-2025-47586 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-21 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR UNSPECIFIED INFECTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 175 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22555720&pc=NGV | 2124215-2025-47781 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-21 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Transient Ischemic Attack |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR TRANSIENT ISCHEMIC ATTACK (TIA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 160 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22555860&pc=NGV | 2124215-2025-47818 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-21 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Transient Ischemic Attack |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR TRANSIENT ISCHEMIC ATTACK (TIA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX DAPT CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 380 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE WATCHMAN FLX DAPT POST APPROVAL PLAN (WATCHMAN FLX-DAPT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22555906&pc=MFA | 3007284006-2025-00140 | 2025-07-02 04:00 | Death | SPECTRANETICS CORPORATION | 2025-07-21 04:00 | MFA | GLIDELIGHT LASER SHEATH | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation | P960042 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV), RIGHT ATRIAL (RA), AND LEFT VENTRICULAR (LV) LEADS DUE TO CIED SYSTEM/POCKET INFECTION. DUE TO SIGNIFICANT COMORBIDITIES, THE PATIENT WAS DEEMED UNFIT FOR SURGICAL INTERVENTION OR SURGICAL BACKUP. PRE-PROCEDURAL VENOGRAPHY SHOWED THAT THE LEADS WERE BORDERING THE LATERAL WALL OF THE SUPERIOR VENA CAVA (SVC), CAUSING CONCERN FOR VESSEL EXTERNALIZATION. SPECTRANETICS LLD EZ LEAD LOCKING DEVICES (LLD EZS), ALONG WITH SUTURE TIES, WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 16F GLIDELIGHT LASER SHEATH ON THE RV LEAD, THE LEAD WAS SUCCESSFULLY REMOVED. NEXT, SWITCHING TO THE RA LEAD, USING THE SAME TECHNIQUE, ADVANCEMENT WAS MADE TO THE INNOMINATE, WHEN THE LEAD DETACHED AND WAS REMOVED. THEN, TARGETING THE LV LEAD, THE LLD EZ BROKE AT THE POCKET (MDR #3007284006-2025-00141), AND THE SUTURE TIES WERE USED TO MAINTAIN TRACTION. WHEN PROGRESS WAS MADE TO THE SVC/RA JUNCTION, THE PATIENT¿S BLOOD PRESSURE DROPPED, AND RELEASING THE TRACTION, THE PATIENT STABILIZED, BUT THEN DECLINED AGAIN. THERE WAS NO PERICARDIAL EFFUSION DETECTED VIA TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE); HOWEVER, THERE WAS FREE FLUID WITHIN THE CHEST, AND AN SVC PERFORATION WAS CONFIRMED ON FURTHER IMAGING. DUE TO THE PATIENT¿S FRAGILITY AND INELIGIBILITY FOR SURGICAL INTERVENTION, THE PROCEDURE WAS TERMINATED, AND THE PATIENT DID NOT SURVIVE. THIS REPORT CAPTURES THE 16F GLIDELIGHT DEVICE IN USE WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF THE GLIDELIGHT DEVICE IN USE DURING THE PROCEDURE. Manufacturer Narrative: A2) PATIENT DATE OF BIRTH - NOT AVAILABLE FROM FACILITY. A4) PATIENT WEIGHT - NOT AVAILABLE FROM FACILITY. A5/A6) PATIENT ETHNICITY AND RACE - NOT AVAILABLE FROM FACILITY. B6) RELEVANT TESTS/LABORATORY DATA - NOT AVAILABLE FROM FACILITY. D4) DEVICE SERIAL NUMBER - NOT AVAILABLE FROM FACILITY. H3) THE GLIDELIGHT WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) GREAT VESSEL PERFORATION AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE GLIDELIGHT DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22555987&pc=NPT | 2015691-2025-05905 | 2025-07-01 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-21 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES AFFILIATE IN ITALY, REGARDING A 20 MM SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE WAS IMPLANTED IN THE AORTIC POSITION VIA TRANSFEMORAL APPROACH, FOLLOWING SUCCESSFUL DEPLOYMENT OF THE VALVE, THE PATIENT EXPERIENCED A SUDDEN ONSET OF RESPIRATORY ARREST, WHICH RAPIDLY PROGRESSED TO CARDIAC ARREST AND ULTIMATELY RESULTED IN THE PATIENT'S DEATH. ACCORDING TO THE IMPLANTING PHYSICIANS, THE CLINICAL DETERIORATION MAY HAVE BEEN CAUSED BY AN ACUTE WORSENING OF MITRAL VALVE FUNCTION, POTENTIALLY DUE TO RUPTURE OF A PAPILLARY MUSCLE. THIS LIKELY LED TO SEVERE MITRAL REGURGITATION AND PULMONARY CONGESTION, CULMINATING IN ACUTE RESPIRATORY COMPROMISE. DESPITE IMMEDIATE RESUSCITATIVE EFFORTS, INCLUDING CPAP AND INTUBATION ATTEMPTS, THE PATIENT COULD NOT BE REVIVED. THE PAPILLARY MUSCLE RUPTURE WAS THE HYPOTHESIS FORMULATED BY THE OPERATORS AND IT WAS CAUSED BY THE PROCEDURAL WIRE. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ADDITIONAL INFORMATION PROVIDE. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: B.4, G.3, G.6, H.2 AND CORRECTED H.6 CLINICAL CODE:4855, INVESTIGATION FINDINGS CORRECTED TO 3233, AND INVESTIGATION CONCLUSIONS CORRECTED TO 11. H.10 HAS BEEN CORRECTED TO THE INVESTIGATION IS ONGOING. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIAC ARREST IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BALLOON VALVULOPLASTY, THE USE OF LOCAL AND/OR GENERAL ANESTHESIA, BIOPROSTHETIC HEART VALVES, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. CARDIAC ARREST OCCURS WHEN A MAJOR CARDIOVASCULAR, RESPIRATORY, OR METABOLIC EVENT RESULTS IN THE INABILITY OF CARDIAC MUSCLE TO GENERATE SUFFICIENT FORCE IN RESPONSE TO ELECTRICAL DEPOLARIZATION. IT MAY BE CAUSED BY A PROFOUND CARDIOVASCULAR INSULT. EXAMPLES INCLUDE SEVERE PROLONGED HYPOXIA OR ACIDOSIS, EXTREME HYPOVOLEMIA, FLOW-RESTRICTING PULMONARY EMBOLUS, CARDIAC TAMPONADE, THROMBOSIS (CORONARY OR PULMONARY). THESE EVENTS MAY BE RELATED TO THE EDWARDS DEVICE IF IT IS ASSOCIATED WITH CARDIAC TAMPONADE, ANNULAR RUPTURE, OR OTHER EDWARDS'S DEVICE-RELATED BLEED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT AND PROCEDURAL FACTORS (ACUTE WORSENING OF MITRAL VALVE FUNCTION, POTENTIALLY DUE TO RUPTURE OF A PAPILLARY MUSCLE) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22556080&pc=DSQ | 3007042319-2025-01490 | 2025-07-16 04:00 | Death | HEARTWARE, INC. | 2025-07-21 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Diarrhea; Low Blood Pressure/ Hypotension; Unspecified Infection; Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT PRESENTED TO A PERIPHERAL HOSPITAL OUTSIDE THE IMPLANTING CENTER. UPON ARRIVAL AT THE HOSPITAL, THE PATIENT WAS PLACED IN THE SHOCK ROOM AND THE VAD WAS PROVIDING CIRCULATION BUT THERE WAS NO CONTRACTILITY OF THE LEFT VENTRICLE (LV) ON ECHOCARDIOGRAM. THE PATIENT WAS HYPOVOLEMIC FROM DIARRHEA DUE TO A GASTROINTESTINAL INFECTION. THE MEAN ARTERIAL PRESSURE (MAP) WAS 50 MMHG AND THE DECISION WAS MADE NOT TO RESUSCITATE THE PATIENT. THE PATIENT DIED AND THE VADWAS POWERED DOWN AFTER DEATH. THE OFFICIAL CAUSE OF DEATH WAS MULTI ORGAN FAILURE. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION SUMMARY. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO MULTI ORGAN FAILURE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, INFECTION, MULTI ORGAN FAILURE, AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22556281&pc=NGV | 2124215-2025-48539 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-21 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 3 | Event Description: THIS REPORT SUMMARIZES 3 DEATH EVENTS FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 71 TO 87 YEARS. PATIENT SEX: 2 FEMALES; 1 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 621 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22556300&pc=OZD | 1220648-2025-30546 | 2025-07-08 04:00 | Death | ABIOMED, INC. | 2025-07-21 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT ENDURED VENTRICULAR TACHYCARDIA DURING IMPELLA CP SUPPORT. AN AMINO DRIP WAS STARTED AND TWO SHOCKS WERE GIVEN TO STABILIZE THE PATIENT. THE PUMP WAS ALSO REPOSITIONED VIA ECHO TO HELP PREVENT FUTURE RECURRENCE. AFTER THREE DAYS OF SUPPORT, THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22556365&pc=MNS | 2518422-2025-107999 | 2025-07-17 04:00 | Death | RESPIRONICS, INC. | 2025-07-21 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE RENTAL DREAMSTATION ST30 DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22556443&pc=DSQ | 2916596-2025-04781 | 2025-06-30 04:00 | Death | THORATEC CORPORATION | 2025-07-21 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH PATIENT OUTCOME THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS A SPONTANEOUS INTRACRANIAL BLEED. NO DEVICE RELATED ISSUE LED TO THE BLEED. NO ANTICOAGUATION ISSUES WERE REPORTED. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) FUNCTIONED AS EXPECTED, AND THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING, AND DEATH THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE IFU, UNDER ¿ANTICOAGULATION¿, PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22556572&pc=KDJ | 3030306055-2025-00114 |  | Death | VANTIVE US HEALTHCARE LLC | 2025-07-21 04:00 | KDJ | NI | Adverse Event Without Identified Device or Use Problem | Sepsis; Peritonitis | NI |  | 1 | Event Description: A PERITONEAL DIALYSIS (PD) PATIENT EXPERIENCED "3RD ROUND OF PERITONITIS". THE CAUSE OF THE PERITONITIS WAS UNKNOWN. THE PATIENT WAS TREATED WITH UNSPECIFIED ANTIBIOTICS INTRAVENOUSLY. SUBSEQUENTLY, IT WAS REPORTED THAT THE PATIENT EXPERIENCED SEPTICEMIA AND HAD BEEN IN THE HOSPITAL ABOUT 7 MONTHS AND PASSED AWAY. IT WAS NOT REPORTED IF THE PERITONITIS WAS RESOLVED. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. IT WAS NOT REPORTED IF PD THERAPY WAS ONGOING PRIOR TO DEATH OR AT THE TIME OF DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: B2 DATE OF DEATH: IT WAS REPORTED THAT THE PATIENT "PASSED FIRST WEEK ON (B)(6) 2025". B3 DATE OF EVENT: IT WAS REPORTED THAT THE PATIENT HAD BEEN IN THE HOSPITAL FROM ON (B)(6) 2024 DUE TO SEPTICEMIA. IT WAS FURTHER REPORTED THAT THE PATIENT EXPERIENCED THE 3RD ROUND OF PERITONITIS AND PASSED AWAY THE FIRST WEEK ON (B)(6) 2025. H11: THE DEVICE WAS NOT RETURNED, AND THE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22556598&pc=DSQ | 2916596-2025-04831 | 2025-07-11 04:00 | Death | THORATEC CORPORATION | 2025-07-21 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT NEVER RECOVERED AFTER IMPLANTATION AND PASSED AWAY. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WOULD NOT BE PERFORMED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: SECTION E3 - OCCUPATION: CORRECTED. SECTION G2 - REPORT SOURCE: CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22556599&pc=DSQ | 2916596-2025-04860 | 2025-07-15 04:00 | Death | THORATEC CORPORATION | 2025-07-21 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE PATIENT OUTCOME, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT COMMUNICATED. THE PATIENT WAS VERY FRAIL AND INDICATION OF IMPLANTATION WAS A LAST RESORT ATTEMPT. THAT UNFORTUNATELY FAILED. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WILL NOT BE EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: SECTION A: ADDITIONAL PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER: (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22556741&pc=DQY | 2124215-2025-47789 | 2025-07-14 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-21 04:00 | DQY | WATCHMAN FXD CURVE? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Hematoma; Low Blood Pressure/ Hypotension; Pain; Vascular Dissection | K212228 |  | 1 | Event Description: IT WAS REPORTED THAT A CARDIAC DISSECTION OCCURRED. THE PROCEDURE WAS CANCELLED. DURING A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE, A VERSACROSS CONNECT (VCC) LAAC WAS SELECTED FOR USE ALONGSIDE A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS). THE PHYSICIAN WAS HAVING DIFFICULTY TO PERFORM TRANSEPTAL AND STRUGGLED TO ENGAGE SEPTUM WITHIN THE FOSSA. IT WAS ATTEMPTED MULTIPLE TIMES BY READVANCING UP THE SUPERIOR VENA CAVA (SVC) WITH THE VCC WIRE LEADING AND COMING DOWN TO ENGAGE THE FOSSA LIKE NORMAL. THE PHYSICIAN PULLED OUT THE WAS/VCC SYSTEM AND RESHAPED IT TO GIVE IT GREATER BEND. AFTER ANOTHER FEW ATTEMPTS LIKE BEFORE, THE WAS/VCC SYSTEM WAS PULLED OUT ONE MORE TIME TO ADD MORE BEND PROXIMALLY. THIS TIME THE PHYSICIAN WAS ABLE TO ENGAGE THE SEPTUM AND GET ACROSS NORMALLY. HOWEVER ONCE ACROSS AND WERE ATTEMPTING TO MOVE THE VCC WIRE, THE ANESTHESIOLOGIST NOTICED SOMETHING IN THE RIGHT ATRIUM (RA). A LARGE FLUID SAC IN THE RA WAS NOTED AND THE DECISION WAS MADE TO ABORT THE CASE. THE WAS/VCC WAS WITHDRAWN, FLUID AND PRESSORS WERE GIVEN BY ANESTHESIA TO HELP INCREASE BLOOD PRESSURE, AN ARTERIAL LINE WAS STARTED, AND THEN THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) FOR MONITORING. CARDIOTHORACIC SURGERY WAS CONSULTED HOWEVER OPERATION WAS NOT RECOMMENDED DUE TO THE PATIENT AGE. IN THE PHYSICIAN'S OPINION, WHILE TRYING TO GO TRANSEPTAL, THE PHYSICIAN INADVERTENTLY CAUSED SOME SORT OF INJURY TO THE RA WALL WHICH CAUSED BLEEDING AND A DISSECTION OF THE RA WALL, CREATING A SUB-LUMINAL HEMATOMA.Event Description: IT WAS REPORTED THAT A CARDIAC DISSECTION OCCURRED. THE PROCEDURE WAS CANCELLED. DURING A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE, A VERSACROSS CONNECT (VCC) LAAC WAS SELECTED FOR USE ALONGSIDE A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS). THE PHYSICIAN WAS HAVING DIFFICULTY TO PERFORM TRANSEPTAL AND STRUGGLED TO ENGAGE SEPTUM WITHIN THE FOSSA. IT WAS ATTEMPTED MULTIPLE TIMES BY READVANCING UP THE SUPERIOR VENA CAVA (SVC) WITH THE VCC WIRE LEADING AND COMING DOWN TO ENGAGE THE FOSSA LIKE NORMAL. THE PHYSICIAN PULLED OUT THE WAS/VCC SYSTEM AND RESHAPED IT TO GIVE IT GREATER BEND. AFTER ANOTHER FEW ATTEMPTS LIKE BEFORE, THE WAS/VCC SYSTEM WAS PULLED OUT ONE MORE TIME TO ADD MORE BEND PROXIMALLY. THIS TIME THE PHYSICIAN WAS ABLE TO ENGAGE THE SEPTUM AND GET ACROSS NORMALLY. HOWEVER ONCE ACROSS AND WERE ATTEMPTING TO MOVE THE VCC WIRE, THE ANESTHESIOLOGIST NOTICED SOMETHING IN THE RIGHT ATRIUM (RA). A LARGE FLUID SAC IN THE RA WAS NOTED AND THE DECISION WAS MADE TO ABORT THE CASE. THE WAS/VCC WAS WITHDRAWN, FLUID AND PRESSORS WERE GIVEN BY ANESTHESIA TO HELP INCREASE BLOOD PRESSURE, AN ARTERIAL LINE WAS STARTED, AND THEN THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) FOR MONITORING. CARDIOTHORACIC SURGERY WAS CONSULTED HOWEVER OPERATION WAS NOT RECOMMENDED DUE TO THE PATIENT AGE. IN THE PHYSICIAN'S OPINION, WHILE TRYING TO GO TRANSEPTAL, THE PHYSICIAN INADVERTENTLY CAUSED SOME SORT OF INJURY TO THE RA WALL WHICH CAUSED BLEEDING AND A DISSECTION OF THE RA WALL, CREATING A SUB-LUMINAL HEMATOMA. IT WAS FURTHER REPORTED THAT THE PATIENT DIED EIGHT (8) DAYS POST INDEX PROCEDURE. THE PHYSICIAN STATED THAT THE HEMATOMA FROM THE PROCEDURE WAS IMPROVING IN SIZE ON SERIAL IMAGING AND EXPECTED THE PATIENT TO FULLY RECOVER FROM THE COMPLICATION. THE PHYSICIAN STATED THE PATIENT WAS EXPERIENCING A LOT OF BACK PAIN AND PALLIATIVE/COMFORT CARE WAS BROUGHT ON BOARD AND EVENTUALLY THE PATIENT EXPIRED SHORTLY AFTER. IT IS UNCLEAR WHAT CAUSED THE PATIENT TO ULTIMATELY EXPIRE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22559399&pc=MVK | 3008642652-2025-07013 | 2025-06-20 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-21 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal; Structural Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE STARTED UP AT 06:02:35 ON (B)(6) 2025. AT 18:44:32 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS AF/AF WITH RVR FROM 70-170 BPM WITH NSVT/PVCS. AT 18:45:06, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. AF WITH RVR AND NSVT CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS AF WITH RVR @ 160 BPM WITH PVCS. POST SHOCK RHYTHM WAS ASYSTOLE WITH UNCONDUCTED P WAVES FOR 8 SECONDS TRANSITIONING TO SINUS BRADYCARDIA FROM 15-40 BPM WITH PVCS AND HB/UNCONDUCTED P WAVES. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE DEVICE SHUT DOWN AT 18:53:31 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. DEVICE EVALUATION OF ELECTRODE BELT HAS BEEN COMPLETED. UPON INVESTIGATION THE ELECTRODE BELT DID NOT PASS A THERAPY ELECTRODE RECOGNITION TEST. THE CAUSE FOR NOT PASSING WAS ISOLATED TO BROKEN PULSE WIRE PINS. THE ROOT CAUSE FOR THE DAMAGED PULSE WIRE PINS COULD NOT BE POSITIVELY IDENTIFIED. THE DAMAGED ELECTRODE BELT DID NOT CAUSE OR CONTRIBUTE TO THE ADVERSE EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22559586&pc=LPB | 2029046-2025-50002 | 2024-10-25 04:00 | Death | STOCKERT GMBH | 2025-07-21 04:00 | LPB | SMARTABLATE¿ SYSTEM RF GENERATOR (EU) | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Fistula |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION ABLATION WITH A THERMOCOOL® SMART TOUCH® SF BI-DIRECTIONAL NAVIGATION CATHETER AND SMARTABLATE¿ SYSTEM RF GENERATOR (EU) AND THE PATIENT EXPERIENCED ESOPHAGEAL FISTULA AND STROKE. THE PATIENT DIED. A SUSPECTED ESOPHAGEAL FISTULA WAS NOTICED. THE PATIENT HAD AN AFIB CASE ON(B)(6) 2024, THEN THE PATIENT ARRIVED AT THE EMERGENCY ROOM AFTER WITH SYMPTOMS OF A STROKE. THE PATIENT WAS UNDERGOING NEUROLOGICAL TESTS. THE PATIENTS LAST KNOWN STATUS WAS THAT THEY WERE UNRESPONSIVE. THE ADVERSE EVENT OCCURRED (B)(6) 2024 (24 DAYS POST PROCEDURE). PATIENT WAS AWAITING SURGICAL INTERVENTION BUT UNFORTUNATELY THE PATIENT PASSED AWAY ON THE SAME DAY HE WAS ADMITTED BEFORE SURGERY. PATIENT PASSED AWAY DUE TO A STOKE FROM SUSPECTED AIR EMBOLI ENTERING THE LEFT ATRIUM THROUGH A SUSPECTED ESOPHAGEAL FISTULA. ESOPHAGEAL INJURY WAS NOT CONFIRMED AS THE PATIENT PASSED AWAY PRIOR TO SURGERY. THE PHYSICIAN STATED THAT GIVEN THAT THE PATIENT¿S STROKE OCCURRED WHILE ON FULL ANTICOAGULATION 24 DAYS AFTER A PULMONARY VEIN ISOLATION, THE PRESENCE OF AIR BUBBLES IN THE LEFT ATRIUM WAS SUFFICIENT TO CONFIRM SUSPICION OF AN ESOPHAGEAL FISTULA. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT WAS PROCEDURE AND PATIENT CONDITION. AFTER REVIEWING THE CASE, PHYSICIAN STATED THAT HE WOULD NOT HAVE CHANGED ANY ASPECT OF THE PROCEDURE. HE TOOK EVERY POSSIBLE PRECAUTION TO AVOID ABLATING NEAR THE ESOPHAGUS; HOWEVER, ESOPHAGEAL FISTULAS REMAIN A KNOWN RISK ASSOCIATED WITH PULMONARY VEIN ISOLATION. HE SUSPECTS THAT AN ESOPHAGEAL FISTULA CAUSED THE STOKE LEADING TO DEATH OF THE PATIENT AND THAT THIS OUTCOME MAY BE ATTRIBUTED TO THE PATIENT¿S UNIQUE TISSUE CHARACTERISTICS AND OVERALL CONDITION. TEMPERATURE MONITORING WAS DONE BY THE ANESTHESIOLOGIST DURING THE AF (ATRIAL FIBRILLATION) PROCEDURE. TEMPERATURE PROBE WAS VERIFIED TO BE IN THE CORRECT POSITION BY VISUALIZING USING INTRACARDIAC ULTRASOUND AND THIS STRUCTURE WAS CONTOURED USING CARTOSOUND TO MARK THE LOCATION OF THE ESOPHAGUS RELATIVE TO THE LEFT ATRIUM. AS PER NORMAL PRACTICE, ABLATION OVER THE TEMPERATURE PROBE/ ESOPHAGUS REGION WAS AVOIDED/ MINIMIZED. IF TEMPERATURE INCREASED, PHYSICIAN WOULD PAUSE ABLATION AND MOVE TO A DIFFERENT AREA UNTIL THE TEMPERATURE CAME DOWN. SMARTABLATE GENERATOR SERVICING IS NOT NEEDED. THERE WAS NO GENERATOR MALFUNCTION. PHYSICIAN NOTED THAT THE PROCEDURE OCCURRED WITHOUT ANY ACUTE COMPLICATIONS AND ALL PRECAUTIONS WERE TAKEN TO AVOID ABLATING NEAR THE ESOPHAGUS. THIS PATIENT OUTCOME WAS BELIEVED TO NOT BE DUE TO ANY SPECIFIC ISSUE DURING THE PROCEDURE. GENERATOR PARAMETERS INCLUDED: STSF CATHETER WITH SMARTABLATE GENERATOR USING STSF PRESET IN POWER CONTROL MODE - ABLATION PARAMETERS WERE 50W LESIONS FOR 6-12 SECONDS. HIGH POWER SHORT DURATION. FORCE VISUALIZATION FEATURES USED: VECTOR, VISITAG. PARAMETERS FOR STABILITY USED FOR VISITAG MODULE: 3MM MAXIMUM DISTANCE CHANGE AND 4S MINIMUM TIME WITH RESPIRATORY ADJUSTMENT ON. COLOR OPTIONS: IMPEDANCE DROP WITH A RANGE OF 5-10 OHMS. THERE ARE TWO REPORTS FOR THIS EVENT. SINCE A REPORT HAS ALREADY BEEN SUBMITTED FOR THE THERMOCOOL SMARTTOUCH, THIS REPORT IS FOR THE SMARTABLATE GENERATOR. Manufacturer Narrative: ON 21-JUL-2025, THE PRODUCT INVESTIGATION WAS COMPLETED AS SERVICE FOR THE DEVICE WAS DECLINED. DEVICE INVESTIGATION DETAILS: NO WORK ORDER OR ANY OTHER EVIDENCE OF ANALYSIS FOR THIS DEVICE WAS SENT TO JOHNSON & JOHNSON MEDTECH FOR EVALUATION. THE SERVICE WAS DECLINED. NO FURTHER INVESTIGATION ON THIS DEVICE WAS PERFORMED. ACCORDING TO CUSTOMER INFORMATION, THERE WAS NO GENERATOR MALFUNCTION. PHYSICIAN NOTED THAT THE PROCEDURE OCCURRED WITHOUT ANY ACUTE COMPLICATIONS AND ALL PRECAUTIONS WERE TAKEN TO AVOID ABLATING NEAR THE ESOPHAGUS. HOWEVER, HE BELIEVES THAT ESOPHAGEAL FISTULAS ARE A KNOWN RISK ASSOCIATED WITH PULMONARY VEIN ISOLATION AND THIS PATIENT OUTCOME WAS NOT DUE TO ANY SPECIFIC ISSUE DURING THE PROCEDURE. A DEVICE HISTORY RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER (B)(6), AND NO INTERNAL ACTIONS RELATED TO THE REPORTED CONDITION WERE IDENTIFIED. AS PART OF THE QUALITY PROCESS, THE DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: D4: UDI: THE MANUFACTURING DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC. OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22559801&pc=OZD | 1220648-2025-30554 | 2025-07-09 04:00 | Death | ABIOMED, INC. | 2025-07-21 04:00 | OZD | IMPELLA 5.5 | Failure to Advance; Patient-Device Incompatibility | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT NOTED WITH POSTCARDIOTOMY CARDIOGENIC SHOCK (PCCS)/LOW CARDIAC OUTPUT SYNDROME (LCOS) WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PUMP WAS BACKLOADED OVER ABIOMED 0.018 WIRE. THE PUMP WAS UNABLE TO PASS PATIENT'S ANATOMY. DECISION WAS THEN MADE TO TRANSITION TO THE LEFT. INSERTION WAS DIFFICULT BUT PUMP WAS EVENTUALLY PLACED INTO LEFT VENTRICLE AND WIRED REMOVED. IT WAS FURTHER REPORTED THAT THE PATIENT¿S CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. HOWEVER, CLINICAL DETAILS PROVIDED WERE SUFFICIENT TO DETERMINE THE ROOT CAUSE. THE ROOT CAUSE OF THE DELIVERY ISSUE WAS DUE TO PATIENT CONDITION. UPDATES TO THE INITIAL MFR REPORT: G1 MDR REPORTING CONTACT EMAIL. H6 MED DEV PROB CODE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22559872&pc=NPT | 2025587-2025-05178 | 2025-07-17 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-21 04:00 | NPT | EVOLUT PRO PLUS VALVE | Fluid/Blood Leak; Perivalvular Leak; Insufficient Information; Activation Failure; Patient Device Interaction Problem | Cardiogenic Shock; Valvular Insufficiency/ Regurgitation; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE, THE DEVICE APPEARED CONSTRICTED DURING DEPLOYMENT, INITIALLY RESEMBLING AN INFOLD; HOWEVER, IT WAS DETERMINED TO BE CAUSED BY SEVERE NATIVE CALCIFICATION, WHICH DEFLECTED THE VALVE TO THE SIDE. THE DEVICE WAS THEN REPOSITIONED AND RECAPTURED TWICE BEFORE FINAL DEPLOYMENT. FOLLOWING DEPLOYMENT, SEVERE AORTIC INSUFFICIENCY (AI) DEVELOPED, "ATTRIBUTED TO TRAUMA CAUSED BY THE GUIDEWIRE." THE GUIDEWIRE WAS THEN REMOVED AND MECHANICAL AND MEDICATION RESUSCITATION WAS INITIATED DUE TO AN UNKNOWN REASON. HEMODYNAMICS WERE THEN STABILIZED, WITH NORMALIZATION OF BLOOD PRESSURE, PULSE, AND SKIN COLOR. ADDITIONALLY, AN ECHOCARDIOGRAM WAS PERFORMED WHICH CONFIRMED THE AI AND REVEALED A PARAVALVULAR LEAK IN THE LEFT SINUS, CORRESPONDING TO A KNOWN AREA OF SIGNIFICANT CALCIFICATION. WHILE THE ACCESS SITE WAS BEING CLOSED, THE PATIENT WENT INTO SHOCK AND ULTIMATELY DIED. PER THE PHYSICIAN, THE PROBABLE CAUSE OF DEATH WAS CARDIOGENIC SHOCK.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE VALVE IMPLANT, FOLLOWING DEPLOYMENT, THE SEVERE AORTIC INSUFFICIENCY WAS CLARIFIED TO BE PRE-EXISTING MILD AI. THREE IMPLANT ATTEMPTS WERE MADE AND THE VALVE WAS RECAPTURED TREE TIMES. DURING THE IMPLANT ATTEMPTS, THE VALVE REMAINED UNDER EXPANDED DUE TO LEAFLET CALCIFICATION AND MILD PARAVALVULAR LEAK (PVL) WAS NOTED. THE GUIDEWIRE USED WAS NOT A MEDTRONIC GUIDEWIRE. THE PATIENT WENT INTO SHOCK WITH NO APPARENT CAUSE AND MULTIPLE RESUSCITATION MANEUVERS, CARDIOVERSION AND PHARMACOLOGICAL RESUSCITATION WERE INITIATED. Manufacturer Narrative: CONTINUATION OF D10: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID {{GWBC30}}; PRODUCT LOT/SERIAL NUMBER {{UNKNOWN}}; PRODUCT TYPE: {{0195 HEART VALVES}}; IMPLANT DATE {{NA}}; EXPLANT DATE {{NA}} PRODUCT ID {{D-EVOLUTFX-2329}}; PRODUCT LOT/SERIAL NUMBER {{UNKOWN}}; PRODUCT TYPE: {{0195 HEART VALVES}}; IMPLANT DATE {{NA}}; EXPLANT DATE {{NA}} SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. H6. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5 CONTINUATION OF D10: THE OTHER MEDICAL PRODUCT PREVIOUSLY REPORTED (PRODUCT ID {{GWBC30}}) NO LONGER MEETS REPORTING CRITERIA. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22560431&pc=FDT | 9610595-2025-14717 |  | Death | AIZU OLYMPUS CO., LTD. | 2025-07-21 04:00 | FDT | EVIS EXERA III DUODENOVIDEOSCOPE | Adverse Event Without Identified Device or Use Problem | Septic Shock | K250701 |  | 1 | Event Description: OLYMPUS MEDICAL SYSTEMS CORP. (OMSC) RECEIVED A LITERATURE TITLED ""DIRECT COMPARISON OF SIMULTANEOUS AND SEQUENTIAL ENDOSCOPIC METALLIC BILATERAL STENTING IN MALIGNANT HILAR BILIARY OBSTRUCTION." THE AIM OF THE STUDY IS TO EVALUATE A POSSIBLE SUPERIORITY OF SIMULTANEOUS ¿SIDE BY SIDE¿ (SBS) BILIARY DRAINAGE IN UNRESECTABLE MHBO. BACKGROUND: ENDOSCOPIC BILATERAL BILIARY DRAINAGE IS A FIRST LINE PALLIATIVE TREATMENT FOR UNRESECTABLE MALIGNANT HILAR BILIARY OBSTRUCTION (MHBO) BUT REMAINS TECHNICALLY CHALLENGING. THE EMERGENCE OF SELF-EXPANDABLE METALLIC STENTS CARRIED BY AN ULTRATHIN (6 FR OR SMALLER) DELIVERY SYSTEM NOW PERMITS SIMULTANEOUS BILATERAL STENT PLACEMENT. TO DATE, ONLY A FEW STUDIES HAVE COMPARED THIS NEW METHOD WITH CONVENTIONAL SEQUENTIAL BILATERAL STENTING. METHODS THIS STUDY IDENTIFIED 135 PATIENTS WHO BENEFITED FROM BILATERAL DRAINAGE USING UNCOVERED SELF-EXPANDABLE METALLIC STENTS BETWEEN 2010 AND 2023. AMONG THEM, 62 BENEFITED FROM SIMULTANEOUS SBS BILATERAL DRAINAGE BETWEEN 2017 AND 2023, AND 73 BENEFITED FROM SEQUENTIAL BILATERAL DRAINAGE [38 USING ¿STENT IN STENT¿ (SIS) TECHNIQUE AND 35 USING SBS TECHNIQUE BETWEEN 2010 AND 2017]. RESULTS: TECHNICAL SUCCESS WAS SIGNIFICANTLY INCREASED IN SIMULTANEOUS DRAINAGE COMPARED WITH SEQUENTIAL DRAINAGE (94% VS 75%, P = 0.008). HOWEVER, SIMULTANEOUS SBS DRAINAGE AND SEQUENTIAL SIS DRAINAGE HAD A SIMILAR TECHNICAL SUCCESS (94% VS 95%). WE OBSERVED NO DIFFERENCES REGARDING CLINICAL SUCCESS, PROCEDURE DURATION AND RECURRENT BILIARY OBSTRUCTION RATE. STENT PATENCY WAS SHORTER IN THE SIS GROUP COMPARED WITH THE SIMULTANEOUS GROUP (103 DAYS VS 144 DAYS). EARLY ADVERSE EVENTS WERE MORE FREQUENT IN THE SEQUENTIAL GROUP (31% VS 21%, P = 0.205), WITH NO DIFFERENCES REGARDING SIS OR SBS TECHNIQUE. TECHNICAL FAILURE WAS ASSOCIATED WITH A HIGHER RATE OF INFECTIOUS FATAL ADVERSE EVENTS (9.5% VS 1.7%, P = 0.02). REINTERVENTION AFTER RECURRENT BILIARY OBSTRUCTION SEEMS TO BE MORE SUCCESSFUL AFTER USING SBS RATHER THAN SIS TECHNIQUES (83% VS 75%, P = 0.53). CONCLUSION: SIMULTANEOUS SBS METALLIC STENT PLACEMENT USING AN ULTRA-THIN DELIVERY SYSTEM WAS TECHNICALLY EASIER AND AS EFFICIENT AS SEQUENTIAL BILATERAL STENTING IN UNRESECTABLE MHBO TO ACHIEVE BILATERAL DRAINAGE. THE SIS PROCEDURE REMAINS A GOOD OPTION IN UNRESECTABLE MHBO."" TYPE OF ADVERSE EVENTS/NUMBER OF PATIENTS FOUR PATIENTS (3%) DIED DUE TO INFECTIOUS COMPLICATIONS (CHOLANGITIS) LEADING TO MULTIVISCERAL FAILURE AND SEPTIC SHOCK. THESE DEATHS OCCURRED AFTER TECHNICAL FAILURE OF THE STENTING PROCEDURE, WHICH WAS SIGNIFICANTLY ASSOCIATED WITH A HIGHER RISK OF FATAL OUTCOMES. THE SEQUENTIAL GROUP HAD THREE FATAL COMPLICATIONS, WHILE THE SIMULTANEOUS GROUP HAD ONE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION IDENTIFIED IN THE ARTICLE. Manufacturer Narrative: THE DEVICE HAS NOT BEEN RETURNED TO OLYMPUS FOR EVALUATION. BASED ON THE RESULTS OF THE INVESTIGATION, THE RELATIONSHIP BETWEEN THE DEVICE AND THE ADVERSE EVENT CANNOT BE CONFIRMED. THERE IS NO EVIDENCE OF AN OLYMPUS DEVICE MALFUNCTION IN THE LITERATURE. THEREFORE, THE ROOT CAUSE CANNOT BE DETERMINED. THE COMPLAINT IS RELATED TO THE FOLLOWING PATIENT IDENTIFIERS: (B)(6). THE LITERATURE DESCRIBED "TJF-190" AND OLYMPUS SELECTED "TJF-Q190V" AS THE REPRESENTATIVE MODEL. HECC- MULTIVISCERAL FAILURE. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22560454&pc=FDT | 9610595-2025-14719 |  | Death | AIZU OLYMPUS CO., LTD. | 2025-07-21 04:00 | FDT | EVIS EXERA DUODENOVIDEOSCOPE | Adverse Event Without Identified Device or Use Problem | Septic Shock | K024033 |  | 1 | Event Description: OLYMPUS MEDICAL SYSTEMS CORP. (OMSC) RECEIVED A LITERATURE TITLED "DIRECT COMPARISON OF SIMULTANEOUS AND SEQUENTIAL ENDOSCOPIC METALLIC BILATERAL STENTING IN MALIGNANT HILAR BILIARY OBSTRUCTION." THE AIM OF THE STUDY IS TO EVALUATE A POSSIBLE SUPERIORITY OF SIMULTANEOUS ¿SIDE BY SIDE¿ (SBS) BILIARY DRAINAGE IN UNRESECTABLE MHBO. BACKGROUND: ENDOSCOPIC BILATERAL BILIARY DRAINAGE IS A FIRST LINE PALLIATIVE TREATMENT FOR UNRESECTABLE MALIGNANT HILAR BILIARY OBSTRUCTION (MHBO) BUT REMAINS TECHNICALLY CHALLENGING. THE EMERGENCE OF SELF-EXPANDABLE METALLIC STENTS CARRIED BY AN ULTRATHIN (6 FR OR SMALLER) DELIVERY SYSTEM NOW PERMITS SIMULTANEOUS BILATERAL STENT PLACEMENT. TO DATE, ONLY A FEW STUDIES HAVE COMPARED THIS NEW METHOD WITH CONVENTIONAL SEQUENTIAL BILATERAL STENTING. METHODS THIS STUDY IDENTIFIED 135 PATIENTS WHO BENEFITED FROM BILATERAL DRAINAGE USING UNCOVERED SELF-EXPANDABLE METALLIC STENTS BETWEEN 2010 AND 2023. AMONG THEM, 62 BENEFITED FROM SIMULTANEOUS SBS BILATERAL DRAINAGE BETWEEN 2017 AND 2023, AND 73 BENEFITED FROM SEQUENTIAL BILATERAL DRAINAGE [38 USING ¿STENT IN STENT¿ (SIS) TECHNIQUE AND 35 USING SBS TECHNIQUE BETWEEN 2010 AND 2017]. RESULTS: TECHNICAL SUCCESS WAS SIGNIFICANTLY INCREASED IN SIMULTANEOUS DRAINAGE COMPARED WITH SEQUENTIAL DRAINAGE (94% VS 75%, P = 0.008). HOWEVER, SIMULTANEOUS SBS DRAINAGE AND SEQUENTIAL SIS DRAINAGE HAD A SIMILAR TECHNICAL SUCCESS (94% VS 95%). WE OBSERVED NO DIFFERENCES REGARDING CLINICAL SUCCESS, PROCEDURE DURATION AND RECURRENT BILIARY OBSTRUCTION RATE. STENT PATENCY WAS SHORTER IN THE SIS GROUP COMPARED WITH THE SIMULTANEOUS GROUP (103 DAYS VS 144 DAYS). EARLY ADVERSE EVENTS WERE MORE FREQUENT IN THE SEQUENTIAL GROUP (31% VS 21%, P = 0.205), WITH NO DIFFERENCES REGARDING SIS OR SBS TECHNIQUE. TECHNICAL FAILURE WAS ASSOCIATED WITH A HIGHER RATE OF INFECTIOUS FATAL ADVERSE EVENTS (9.5% VS 1.7%, P = 0.02). REINTERVENTION AFTER RECURRENT BILIARY OBSTRUCTION SEEMS TO BE MORE SUCCESSFUL AFTER USING SBS RATHER THAN SIS TECHNIQUES (83% VS 75%, P = 0.53). CONCLUSION: SIMULTANEOUS SBS METALLIC STENT PLACEMENT USING AN ULTRA-THIN DELIVERY SYSTEM WAS TECHNICALLY EASIER AND AS EFFICIENT AS SEQUENTIAL BILATERAL STENTING IN UNRESECTABLE MHBO TO ACHIEVE BILATERAL DRAINAGE. THE SIS PROCEDURE REMAINS A GOOD OPTION IN UNRESECTABLE MHBO." TYPE OF ADVERSE EVENTS/NUMBER OF PATIENTS. FOUR PATIENTS (3%) DIED DUE TO INFECTIOUS COMPLICATIONS (CHOLANGITIS) LEADING TO MULTIVISCERAL FAILURE AND SEPTIC SHOCK. THESE DEATHS OCCURRED AFTER TECHNICAL FAILURE OF THE STENTING PROCEDURE, WHICH WAS SIGNIFICANTLY ASSOCIATED WITH A HIGHER RISK OF FATAL OUTCOMES. THE SEQUENTIAL GROUP HAD THREE FATAL COMPLICATIONS, WHILE THE SIMULTANEOUS GROUP HAD ONE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION IDENTIFIED IN THE ARTICLE. Manufacturer Narrative: THE DEVICE HAS NOT BEEN RETURNED TO OLYMPUS FOR EVALUATION. BASED ON THE RESULTS OF THE INVESTIGATION, THE RELATIONSHIP BETWEEN THE DEVICE AND THE ADVERSE EVENT CANNOT BE CONFIRMED. THERE IS NO EVIDENCE OF AN OLYMPUS DEVICE MALFUNCTION IN THE LITERATURE. THEREFORE, THE ROOT CAUSE CANNOT BE DETERMINED. THE COMPLAINT IS RELATED TO THE FOLLOWING PATIENT IDENTIFIERS: (B)(6). THE LITERATURE DESCRIBED "TJF-160" AND SO OLYMPUS SELECTED ¿TJF-160VR " AS THE REPRESENTATIVE MODEL. HECC- MULTIVISCERAL FAILURE. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22560661&pc=FRN | 3004548776-2025-00306 | 2024-07-03 04:00 | Death | FRESENIUS VIAL S.A.S. | 2025-07-21 04:00 | FRN | AGILIA SP MC WIFI GB | Excess Flow or Over-Infusion | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: THE FOLLOWING HAS BEEN REPORTED BY CUSTOMER: EMAIL (NOT FORMAL) CARDIFF COMPLAINT - NNU LIPIDS - ATTENDED A CALL WITH [NURSE] FROM CARDIFF TODAY TO ASSIST WITH CONFIGURATION QUERY. SHE RAISED THAT THIS WAS REQUIRED DUE TO AN INCIDENT WITH A NEONATAL LIPID INFUSION, WHERE THE LIPID WAS GIVEN AT AN INCORRECT FLOW RATE. [NURSE] CONFIRMED THAT THIS WAS A CLINICAL ERROR, AS THE LIPID WAS INFUSED AT THE RATE OF THE SECOND INFUSION (AQUEOUS SOLUTION). THE TRUST HAVE CARRIED OUT THEIR REVIEW OF THE DEVICE, DATIX INFO AND PUMP DATA AND HAVE STATED THEY DO NOT REQUIRE SUPPORT WITH THIS AS IT WAS DEEMED A CLINICAL MISTAKE AND DUE TO IT BEING A FEW MONTHS BACK, HAS BEEN RESOLVED. UNFORTUNATELY, THE PATIENT DIED AS A RESULT OF THIS INCIDENT. REPORTING AS A CONSERVATIVE MEASURE. ADVERSE EFFECTS WERE REPORTED. MORE INFORMATION IS NEEDED TO COMPLETE THE INVESTIGATION. AS DEVICE SERIAL NUMBER WAS NOT PROVIDED, DEVICE HISTORY RECORD COULD NOT BE REVIEWED. DEVICE LOG WAS NOT PROVIDED; THEREFORE, NO REVIEW COULD BE PERFORMED. THE DEVICE WAS NOT RECEIVED IN OUR LABORATORY; HOWEVER, THE INVESTIGATION WAS CONDUCTED BASED ON THE DESCRIPTION PROVIDED. BASED ON THIS DESCRIPTION, THERE HAS BEEN AN INCIDENT WITH A NEONATAL LIPID INFUSION, WHERE THE LIPID WAS GIVEN AT AN INCORRECT FLOW RATE. [NURSE] CONFIRMED THAT THIS WAS A CLINICAL ERROR, AS THE LIPID WAS INFUSED AT THE RATE OF THE SECOND INFUSION (AQUEOUS SOLUTION). FOR THIS COMPLAINT, THE ROOT CAUSE HAS BEEN IDENTIFIED AS AN INCORRECT FLOW RATE SETTING BY THE USER, WHICH CONSTITUTES A USE ERROR. THIS COMPLAINT HAS BEEN INCLUDED IN A TREND ANALYSIS. THIS COMPLAINT IS CONSIDERED AS: USE ERROR. THE TREND IS: NORMAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22560683&pc=FRN | 3000240707-2025-00297 | 2024-07-03 04:00 | Death | FRESENIUS VIAL S.A.S. | 2025-07-21 04:00 | FRN | AGILIA SP MC WIFI GB | Excess Flow or Over-Infusion | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K121613 |  | 1 | Event Description: THE FOLLOWING HAS BEEN REPORTED BY CUSTOMER: EMAIL (NOT FORMAL) CARDIFF COMPLAINT - NNU LIPIDS - ATTENDED A CALL WITH [NURSE] FROM CARDIFF TODAY TO ASSIST WITH CONFIGURATION QUERY. SHE RAISED THAT THIS WAS REQUIRED DUE TO AN INCIDENT WITH A NEONATAL LIPID INFUSION, WHERE THE LIPID WAS GIVEN AT AN INCORRECT FLOW RATE. [NURSE] CONFIRMED THAT THIS WAS A CLINICAL ERROR, AS THE LIPID WAS INFUSED AT THE RATE OF THE SECOND INFUSION (AQUEOUS SOLUTION). THE TRUST HAVE CARRIED OUT THEIR REVIEW OF THE DEVICE, DATIX INFO AND PUMP DATA AND HAVE STATED THEY DO NOT REQUIRE SUPPORT WITH THIS AS IT WAS DEEMED A CLINICAL MISTAKE AND DUE TO IT BEING A FEW MONTHS BACK, HAS BEEN RESOLVED. UNFORTUNATELY, THE PATIENT DIED AS A RESULT OF THIS INCIDENT. REPORTING AS A CONSERVATIVE MEASURE. ADVERSE EFFECTS WERE REPORTED. MORE INFORMATION IS NEEDED TO COMPLETE THE INVESTIGATION. AS DEVICE SERIAL NUMBER WAS NOT PROVIDED, DEVICE HISTORY RECORD COULD NOT BE REVIEWED. DEVICE LOG WAS NOT PROVIDED; THEREFORE, NO REVIEW COULD BE PERFORMED. THE DEVICE WAS NOT RECEIVED IN OUR LABORATORY; HOWEVER, THE INVESTIGATION WAS CONDUCTED BASED ON THE DESCRIPTION PROVIDED. BASED ON THIS DESCRIPTION, THERE HAS BEEN AN INCIDENT WITH A NEONATAL LIPID INFUSION, WHERE THE LIPID WAS GIVEN AT AN INCORRECT FLOW RATE. [NURSE] CONFIRMED THAT THIS WAS A CLINICAL ERROR, AS THE LIPID WAS INFUSED AT THE RATE OF THE SECOND INFUSION (AQUEOUS SOLUTION). FOR THIS COMPLAINT, THE ROOT CAUSE HAS BEEN IDENTIFIED AS AN INCORRECT FLOW RATE SETTING BY THE USER, WHICH CONSTITUTES A USE ERROR. THIS COMPLAINT HAS BEEN INCLUDED IN A TREND ANALYSIS. THIS COMPLAINT IS CONSIDERED AS: USE ERROR. THE TREND IS: NORMAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22561655&pc=BZD | 2518422-2025-108031 | 2023-04-26 04:00 | Death | RESPIRONICS, INC. | 2025-07-21 04:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: IT WAS REPORTED THAT THE DREAMSTATION AUTO CPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22561966&pc=BZD | 2518422-2025-108025 | 2025-06-22 04:00 | Death | RESPIRONICS, INC. | 2025-07-21 04:00 | BZD | DREAMSTATION AUTO | Material Integrity Problem | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A DEATH HAS OCCURRED. WIFE CALLED IN STATING HER HUSBAND GOT LUNG CANCER FROM THE DEVICE & PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22562013&pc=HRS | 0001032347-2025-00284 | 2024-10-10 04:00 | Death | BIOMET MICROFIXATION | 2025-07-21 04:00 | HRS | UNK WIRE | Adverse Event Without Identified Device or Use Problem | Sepsis | NI |  | 1 | Event Description: IT WAS REPORTED IN A JOURNAL ARTICLE STUDYING RIGID PLATE FIXATION SYSTEMS THAT 1 PATIENT DIED DUE TO UNKNOWN SEPSIS, NO FURTHER INFORMATION PROVIDED. ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE AT THIS TIME. Manufacturer Narrative: (B)(4). D10: UNK STERNALOCK PLATE CAT#UNK LOT#UNK G2: JAPAN THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MEDWATCH 3500A WILL BE SUBMITTED. NAKAMURA, H., MIURA, Y., YOSHIDA, K., EDO, N., SAITO, R., & ORIHASHI, K. (2024). EFFECTIVENESS OF RIGID PLATE FIXATION FOR STERNAL CLOSURE IN PATIENTS WITH A HIGH RISK OF DEEP STERNAL WOUND INFECTION. JOURNAL OF INTERNATIONAL MEDICAL RESEARCH, 52(10). HTTPS://DOI.ORG/10.1177/03000605241281915.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22562018&pc=CAW | 2518422-2025-108032 | 2025-07-11 04:00 | Death | RESPIRONICS, INC. | 2025-07-21 04:00 | CAW | EVERFLO 230V, OPI, AUS, ULTRAFILL | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K061261 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED ABOUT A DECEASED PATIENT. LAW ENFORCEMENT INFORMED THEM THAT THE DEVICE IS INVOLVED IN A HOMICIDE INVESTIGATION. THE DEVICE IS NOW IN LAW ENFORCEMENT CUSTODY. CURRENTLY, THERE IS NO EVIDENCE SUGGESTING THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS STILL IN PROGRESS, AND A FOLLOW-UP REPORT WILL BE PROVIDED ONCE IT IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22562022&pc=HRS | 0001032347-2025-00285 | 2024-10-10 04:00 | Death | BIOMET MICROFIXATION | 2025-07-21 04:00 | HRS | UNK STERNALOCK PLATE | Adverse Event Without Identified Device or Use Problem | Sepsis | NI |  | 1 | Event Description: IT WAS REPORTED IN A JOURNAL ARTICLE STUDYING RIGID PLATE FIXATION SYSTEMS THAT 1 PATIENT DIED DUE TO UNKNOWN SEPSIS, NO FURTHER INFORMATION PROVIDED. ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE AT THIS TIME. Manufacturer Narrative: (B)(4). D10: UNK WIRE CAT# UNK; LOT# UNK. G2: JAPAN. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MEDWATCH 3500A WILL BE SUBMITTED. NAKAMURA, H., MIURA, Y., YOSHIDA, K., EDO, N., SAITO, R., & ORIHASHI, K. (2024). EFFECTIVENESS OF RIGID PLATE FIXATION FOR STERNAL CLOSURE IN PATIENTS WITH A HIGH RISK OF DEEP STERNAL WOUND INFECTION. JOURNAL OF INTERNATIONAL MEDICAL RESEARCH, 52(10). HTTPS://DOI.ORG/10.1177/03000605241281915.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22562176&pc=QZI | 2124215-2025-48769 | 2025-01-29 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-21 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Fall; Brain Injury |  |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER THE PATIENT PASSED AWAY. SIX DAYS AFTER THE SUCCESSFUL COMPLETION OF A PFA ABLATION PROCEDURE THE PATIENT PASSED AWAY DUE TO A TRAUMATIC BRAIN INJURY SUFFERED AS THE RESULT OF A FALL. THE CAUSE OF THE FALL HAS NOT BEEN REPORTED AND THUS A COMPLICATION OF THE PROCEDURE CANNOT BE RULED OUT. NO FURTHER INFORMATION HAS YET BEEN PROVIDED REGARDING THE CAUSE. THE CATHETER IS NOT EXPECTED TO BE RETURNED AS IT WOULD HAVE BEEN DISPOSED AFTER THE PROCEDURE WAS COMPLETED. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. GOOD FAITH EFFORTS TO OBTAIN THE INFORMATION ARE IN PROGRESS. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE FILED IF THE INFORMATION IS RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22562378&pc=KPO | 3005162618-2025-00001 | 2022-06-30 04:00 | Death | OREGON MANUFACTURING PLANT | 2025-07-21 04:00 | KPO | COLLECTIVE CONCENTRATE- NATURALYTE | Adverse Event Without Identified Device or Use Problem | Non specific EKG/ECG Changes | K192017 |  | 1 | Event Description: ON 24/JUN/2025 FRESENIUS WAS MADE AWARE OF AN ARTICLE TITLED ¿UNFAMILIAR PRODUCT AND BARCODE SCANNING ERROR LEADS TO PATIENT DEATH¿ WHICH WAS PUBLISHED IN THE AMERICAN JOURNAL OF NURSING AUGUST 2024, VOLUME 124, NUMBER 8. THE ARTICLE REVIEWED A RECENT ALERT FROM THE INSTITUTE FOR SAFE MEDICATION PRACTICES (ISMP) WHICH HIGHLIGHTED A CASE IN WHICH AN INTENSIVE CARE NURSE (ICU) ACCIDENTALLY ADMINISTERED NATURALYTE® (A HEMODIALYSIS PRODUCT) INSTEAD OF THE PRESCRIBED BOWEL PREPARATION PRODUCT (SUPREP BOWEL PREP) TO A PATIENT. THE PATIENT LATER DIED. ADDITIONAL DETAILS WERE OBTAINED FROM A NEWS ARTICLE WRITTEN BY LEIGH SEARCY AND MORGAN EADS ON 25/JUL/2023 FOR WLEX18, LEXINGTON, KENTUCKY. THE ARTICLE PRESENTED A TIMELINE OF EVENTS WITH DATES AND PATIENT INFORMATION. AN 81-YEAR-OLD MALE PATIENT WAS TRANSFERRED TO CHI ST. JOSEPH¿S HEATH MAIN IN LEXINGTON, KY FOR HIGHER LEVEL CARE DUE TO GASTROINTESTINAL (GI) BLEED (DATE UNKNOWN). THIS PATIENT WAS ADMITTED TO THE ICU. ON 30/JUN/2022 A BOWEL PREP WAS ORDERED FOR THIS PATIENT. THE DIALYSIS TEAM HAD PREVIOUSLY LEFT BEHIND A LARGE PLASTIC CONTAINER OF NATURALYTE®, WHICH HAD BEEN PRESCRIBED TO A DIFFERENT PATIENT. THE CONTAINER HAD BEEN LEFT IN THE ICU PHARMACY DELIVERY AREA. THE CONTAINER REPORTEDLY COULD HAVE BEEN THERE FOR UP TO THREE DAYS PER DOCUMENTS CITED IN THE NEWS ARTICLE. THIS PATIENT¿S NURSE SAW THE NATURALYTE® CONTAINER AND ASSUMED THAT IT WAS THE SAME AS GOLYTELY. GOLYTELY IS AN ELECTROLYTE-BASED BOWEL PREP WITH SIMILAR PACKAGING AS NATURALYTE®. ADDITIONALLY, THEY CONTAIN SIMILAR ELECTROLYTES SUCH AS MAGNESIUM, POTASSIUM, AND SODIUM. FURTHERMORE, THIS NURSE ASSUMED THAT GOLYTELY WAS THE SAME GENERIC PRODUCT AS SUPREP BOWEL PREP. THE NURSE DID NOT READ THE LABELING CAREFULLY AS THE NATURALYTE® LABEL DOES STATE THAT IT IS INTENDED FOR HEMODIALYSIS. THE NURSE DID NOT CONFIRM THAT THE PRODUCTS WERE IDENTICAL WITH THE PHARMACY OR BY USING AN ONLINE INDEX. THE NURSE ATTEMPTED TO SCAN THE PHARMACY BARCODE ON THE NATURALYTE® CONTAINER (THINKING IT WAS THE GOLYTELY). THE BARCODE DID NOT SCAN. THE PHARMACY BARCODE DID NOT CONTAIN A PATIENT NAME OR IDENTIFYING FACTOR. THE NURSE CALLED THE PHARMACY AT 5:35PM AND STATED THAT THE CONTAINER WOULD NOT SCAN. THE PHARMACY DID NOT SEND THE CORRECT PRODUCT (SUPREP BOWEL PREP) WITH A LABEL FOR THAT SPECIFIC PATIENT NOR COME UP TO THE ICU TO INSPECT THE MEDICATION. INSTEAD, THE PHARMACY SENT A NEW BARCODE LABEL TO THE ICU THROUGH THE TUBE DELIVERY SYSTEM. THE NEW BARCODE LABEL ARRIVED APPROXIMATELY FIVE MINUTES AFTER THE CALL TO THE PHARMACY. THE NURSE PROCEEDED TO ADMINISTER APPROXIMATELY 8OZ OF THE NATURALYTE® TO THE PATIENT, ORALLY, BEFORE THE END OF HER SHIFT. THE PATIENT WAS UNABLE TO TOLERATE THE LIQUID AS THE TASTE WAS UNPLEASANT AND IT PRODUCED NAUSEA. DOCUMENTED NOTES ANNOTATED THAT THE PHYSICIAN WHO INITIALLY ORDERED THE COLONOSCOPY PREP SAID THE PATIENT MUST TAKE THE FULL AMOUNT OF THE CONTAINER. IT WAS ORDERED THAT THE PREP BE ADMINISTERED VIA FEEDING TUBE. AFTER THE INITIAL NURSE ENDED HER SHIFT, ANOTHER NURSE ATTENDED TO THE PATIENT. THAT NURSE DID NOT CAREFULLY READ THE PREPRINTED LABEL ON THE CONTAINER OF NATURALYTE® AND ADMINISTERED THE REMAINDER OF THE LIQUID THROUGH A FEEDING BAG. THE PATIENT BEGAN TO DETERIORATE. ANOTHER PHYSICIAN ASSESSED THE PATIENT AND MISTOOK THE NATURALYTE® FOR GOLYTELY. SIGNIFICANT ELECTROCARDIOGRAM CHANGES DEVELOPED. THE MEDICATION ERROR WAS CAUGHT AT ABOUT MIDNIGHT. THE PATIENT PASSED AWAY AT 7:35AM ON (B)(6) 2022. THE CORONER¿S REPORT DOCUMENTED THE CAUSE OF DEATH AS COMPLICATIONS OF INADVERTENT ADMINISTRATION OF NATURALYTE® IN THE SETTING OF GI HEMORRHAGE. THE INITIAL NURSE CARING FOR THE PATIENT WAS ASSIGNED TO ADDITIONAL PATIENTS IN THE ICU THAT NIGHT WHICH IS SAID TO HAVE BEEN A FACTOR IN THE MEDICATION ERROR. Manufacturer Narrative: CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN THE ERRONEOUS ADMINISTRATION OF NATURALYTE® TO THIS ICU PATIENT PRESCRIBED A COLONOSCOPY BOWEL PREP WITH SUPREP BOWEL PREP AND ADVERSE OUTCOME OF DEATH. THERE IS NO EVIDENCE OF NATURALYTE® PRODUCT DEFICIENCY OR MANUFACTURING DEFICIENCY AND WAS THE RESULT OF MEDICATION ERROR. NURSES DID NOT PERFORM RIGHTS OF MEDICATION ADMINISTRATION PRIOR TO ADMINISTERING NATURALYTE®. EACH NURSE IS RESPONSIBLE TO CHECK FOR RIGHT PATIENT, DRUG, ROUTE, TIME, AND DOSE PRIOR TO ADMINISTERING MEDICATION. THERE WAS NO LABEL ON THE CONTAINER WHICH IDENTIFIED NATURALYTE® AS BEING PRESCRIBED FOR THAT PATIENT; THERE WAS ONLY A PHARMACY BARCODE. NURSES DID NOT READ MANUFACTURER'S LABEL ON NATURALYTE® CONTAINER WHICH INDICATES WHAT PRODUCT IS IN CONTAINER: ¿NATURALYTE® LIQUID ACID CONCENTRATE FOR BICARBONATE HEMODIALYSIS.¿ INDICATIONS FOR USE STATES: ¿NATURALYTE LIQUID ACID CONCENTRATE IS INDICATED FOR USE IN PATIENTS UNDERGOING EXTRACORPOREAL BICARBONATE HEMODIALYSIS FOR ACUTE AND CHRONIC RENAL FAILURE. NATURALYTE LIQUID ACID CONCENTRATE IS INTENDED TO BE USED AS ONE COMPONENT IN THE PREPARATION OF DIALYSATE IN A THREE-STREAM PROPORTIONING HEMODIALYSIS MACHINE ACCORDING TO A PHYSICIAN¿S PRESCRIPTION.¿ THERE IS A WARNING ON LABEL: ¿WARNING FAILURE TO FOLLOW THESE INSTRUCTIONS FOR USE MAY RESULT IN PATIENT INJURY OR DEATH.¿ GOLYTELY PREP AND SUPREP BOWEL PREPS REQUIRE RECONSTITUTING POWDER OR LIQUID WITH ANOTHER LIQUID WHICH NURSES WOULD HAVE HAD TO DO THEMSELVES OR CHECK IT WAS COMPLETED. HAD NURSES STOPPED TO CHECK THE CORRECTLY LABELED MANUFACTURER LABEL ON NATURALYTE®, COMPLETED FIVE RIGHTS OF MEDICATION ADMINISTRATION AND NOT RELIED ON BARCODE LABEL FROM PHARMACY, THE MEDICATION ERROR WOULD NOT HAVE BEEN MADE. BASED ON AVAILABLE INFORMATION, IT CAN BE CONCLUDED THE PATIENT DEATH RESULTED FROM ADMINISTRATION ERROR BY MEDICAL STAFF AND NOT BY ANY NATURALYTE® PRODUCT OR MANUFACTURING DEFICIENCY. PLANT INVESTIGATION IN PROCESS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22552171&pc=NAY | 2955842-2025-30856 | 2025-06-18 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-07-20 04:00 | NAY | DAVINCI XI | Insufficient Information | Insufficient Information | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI-ASSISTED INGUINAL HERNIA REPAIR, THE PATIENT WAS DISCHARGED HOME. AN INTUITIVE CLINICAL TERRITORY ASSOCIATE WAS LATER INFORMED BY A HOSPITAL NURSE THAT THE PATIENT HAD PRESENTED TO THE EMERGENCY DEPARTMENT EIGHT HOURS LATER WITH UNSPECIFIED SYMPTOMS AND EXPIRED ON POST-OPERATIVE DAY 1. NO INFORMATION REGARDING THE CAUSE OF DEATH OR TREATMENT RENDERED WAS PROVIDED. THE ROBOTIC PROCEDURE WAS REPORTEDLY COMPLETED WITHOUT ANY INTRAOPERATIVE COMPLICATIONS OR DA VINCI SYSTEM ERRORS. ADDITIONAL INFORMATION WAS REQUESTED; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. Manufacturer Narrative: REVIEW OF THE SYSTEM LOGS FOUND NO RELATED ERRORS WERE LOGGED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO RELATED ERRORS OCCURRED. THE SYSTEM FUNCTIONED NORMALLY, NO INTRAOPERATIVE EVENTS OR ISSUES FOUND. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, MONOPOLAR CURVED SCISSORS, FORCE BIPOLAR, FENESTRATED BIPOLAR FORCEPS, MEGA SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THESE PRODUCTS. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE INSTRUMENTS CONFIRMED THAT THERE WERE NO RELATED NON-CONFORMANCES DOCUMENTED. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING AN INGUINAL HERNIA REPAIR. ALTHOUGH NO INTRAOPERATIVE ISSUES OR COMPLICATIONS WERE REPORTED, THE REASON FOR THE RETURN TO THE EMERGENCY ROOM JUST 8 HOURS AFTER SURGERY DISCHARGE IS NOT PROVIDED. THE CAUSE OF DEATH IS ALSO NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION A2: PATIENT AGE IS ESTIMATED; REPORTED TO BE "EARLY TO MID 70'S."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22552174&pc=NVY | 2649622-2025-19656 | 2025-06-21 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-20 04:00 | NVY | EPSILA EV¿ MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P220012 |  | 1 | Event Description: IT WAS REPORTED THAT ONE DAY POST IMPLANT OF AN EXTRAVASCULAR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (EV-ICD) SYSTEM THE PATIENT PAS SED AWAY. THE PHYSICIAN DOES NOT BELIEVE THE PATIENT'S DEATH WAS DIRECTLY RELATED TO ANYTHING RELATED TO THE DEVICE OR PROCEDURE AT THIS TIME. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID DVEA3E4 ((B)(6)); PRODUCT TYPE: 0235-ICD; IMPLANT DATE (B)(6) 2025-. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22552175&pc=LWS | 9614453-2025-02835 | 2025-06-21 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-07-20 04:00 | LWS | AURORA EV-ICD¿ MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P220012 |  | 1 | Event Description: IT WAS REPORTED THAT ONE DAY POST IMPLANT OF AN EXTRAVASCULAR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (EV-ICD) SYSTEM THE PATIENT PASSED AWAY. THE PHYSICIAN DOES NOT BELIEVE THE PATIENT'S DEATH WAS DIRECTLY RELATED TO ANYTHING RELATED TO THE DEVICE OR PROCEDURE AT THIS TIME. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: EV240163 (EVL011163V); PRODUCT TYPE: 0264-LEAD-HV; IMPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22539754&pc=MFA | 3007284006-2025-00144 | 2025-05-01 04:00 | Death | SPECTRANETICS CORPORATION | 2025-07-18 04:00 | MFA | GLIDELIGHT LASER SHEATH | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism | P960042 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) AND A RIGHT ATRIAL (RA) LEAD DUE TO BACTEREMIA. PRIOR TO THE PROCEDURE, A LARGE PERICARDIAL EFFUSION FROM END STAGE RENAL DISEASE WAS NOTED. UTILIZING A SPECTRANETICS 14F GLIDELIGHT LASER SHEATH AND A COOK MEDICAL NEEDLE¿S EYE SNARE, THE RV AND RA LEAD WERE REMOVED. DURING THE EXTRACTION, VEGETATION WAS ENCOUNTERED ON ONE OF THE LEADS, DISLODGING IT AND LATER CAUSING IT TO EMBOLIZE IN THE PATIENT¿S LUNG; THEREFORE, THE PATIENT WENT INTO PULMONARY ARREST AFTER THE PROCEDURE. RESCUE EFFORTS BEGAN, INCLUDING CPR, MEDICATIONS, AND PERICARDIOCENTESIS. HOWEVER, DESPITE RESCUE EFFORTS, THE PATIENT NEVER RECOVERED, AND THE DECISION WAS MADE TO REMOVE FROM LIFE SUPPORT. THIS REPORT CAPTURES THE 14F GLIDELIGHT IN USE WHEN THE VEGETATION BROKE FREE FROM THE LEAD, REQUIRING INTERVENTION, BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF THE SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: A2) PATIENT DATE OF BIRTH AND AGE - NOT AVAILABLE FROM FACILITY. A3B) PATIENT GENDER - NOT AVAILABLE FROM FACILITY. A4) PATIENT WEIGHT - NOT AVAILABLE FROM FACILITY. A5/A6) PATIENT ETHNICITY AND RACE - NOT AVAILABLE FROM FACILITY. B6) RELEVANT TESTS/LABORATORY DATE - NOT AVAILABLE FROM FACILITY. D4/H4) THE LOT NUMBER WAS NOT PROVIDED, THUS THE FOLLOWING INFORMATION IS UNKNOWN: DEVICE SERIAL NUMBER, UNIQUE ID, EXPIRATION DATE, AND MANUFACTURE DATE. H3) THE GLIDELIGHT WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) MIGRATION OF VEGETATION AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE GLIDELIGHT DEVICE. THE IFU WARNS "WHEN MARKED CALCIFICATION THAT MOVES WITH THE LEAD TO BE EXTRACTED IS SEEN ON FLUOROSCOPY, PARTICULARLY IN THE ATRIUM, THE AVAILABILITY OF IMMEDIATE SURGICAL ASSISTANCE IS PARAMOUNT IF A PROBLEM PRESENTS ITSELF AS A RESULT OF THE EXTRACTION PROCEDURE." SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22541066&pc=MIH | 2017233-2025-06427 | 2025-06-20 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-07-18 04:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Patient-Device Incompatibility | Vascular Dissection | P040043 |  | 1 | Event Description: THE FIELD SALES ASSOCIATE (FSA) REPORTED THE FOLLOWING TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE TO TREAT AN ANEURYSM UTILIZING THE GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) AND THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (CTAG) IN ZONE 6. IT WAS REPORTED THAT ALL TAMBE AND CTAG DEVICES WERE IMPLANTED PRIOR TO THE PATIENT HAVING CARDIAC ARREST. THE PHYSICIAN THEN ATTEMPTED TO ADVANCE A GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX) THROUGH THE LEFT SUBCLAVIAN ARTERY (LSA) WHEN THE PATIENT CODED. THE PHYSICIAN AND FSA DETERMINED THAT THE PATIENT MAY HAVE HAD A DISSECTION IN THE MID DESCENDING AORTA. THE PHYSICIAN STATED THAT THE DISSECTION WAS PROXIMAL TO THE CTAG BUT COULD NOT VISUALLY SEE FAR ENOUGH DOWN THE PROXIMAL DESCENDING AORTA TO MAKE A DETERMINATION AS TO EXACT LOCATION OF RETROGRADE TEAR. AT THIS TIME, THE CARDIAC TEAM WAS ACTIVATED, CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED AND THE PATIENT WAS PLACED ON BYPASS. THE PHYSICIAN PERFORMED AN OPEN THORACOTOMY WITH TWO ARCH ANASTOMOSES. ONCE THE PATIENT CAME OFF BYPASS, THEY EXPIRED. AT THE TIME OF THE CARDIAC ARREST AND SUSPECTED DISSECTION, THE VBX DEVICE WAS BEING ROUTED TO ITS INTENDED LOCATION AND HAD NOT YET BEEN DEPLOYED. THE FSA STATED THE VBX, TAMBE, AND CTAG DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT DEATH. THE PATIENT ULTIMATELY EXPIRED AFTER MULTIPLE UNSUCCESSFUL ATTEMPTS TO WEAN FROM BYPASS SUPPORT. Manufacturer Narrative: W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: ACCORDING TO THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE DISSECTION. AS IT IS UNKNOWN WHICH DEVICE(S) IS DIRECTLY IMPLICATED IN THIS ENDOLEAK EVENT, THE FOLLOWING DEVICE(S) WILL ALSO BE INCLUDED IN THIS REPORT: CATALOG #TGMR404010 / SERIAL #(B)(6) / UDI #(B)(4). CATALOG #TGM342815 / SERIAL #(B)(6) / UDI #(B)(4). ADDED G3/G4, H1/H2, H6: INVESTIGATION FINDING AND INVESTIGATION CONCLUSION CODE. CORRECTED H6: HEATH EFFECT CLINICAL AND MEDICAL DEVICE PROBLEM CODE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22541071&pc=OZD | 1220648-2025-30477 | 2025-07-02 04:00 | Death | ABIOMED, INC. | 2025-07-18 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED SUPRAVENTRICULAR TACHYCARDIA WITH A WIDE QRS COMPLEX FOLLOWED BY ASYSTOLE. THE PATIENT WOULD DECLINE AND SUBSEQUENTLY EXPIRE. THE PATIENT WAS FOUND TO HAVE LEFT ANTERIOR DESCENDING ARTERY DISEASE AND CARDIOTHORACIC SURGERY WAS CONSULTED. CORONARY ARTERY BYPASS GRAFT SURGERY WAS SCHEDULED TWO DAYS PRIOR TO THE IMPELLA CP IMPLANT BUT WAS POSTPONED DUE TO CREATININE AND WHITE BLOOD CELLS INCREASE. OVER THE NIGHT, THE PATIENT'S RESPIRATORY STATUS DECLINED AND WAS INTUBATED DUE TO POOR O2 SATURATION. WHEN THE CATHETERIZATION LAB STAFF ARRIVED AT THE UNIT TODAY, THE PT HAD A MEAN ARTERIAL PRESSURE OF 30 WITH NEOSYNEPHRINE AT 50MCG. THE DECISION WAS MADE RUSH THE PATIENT TO THE CATHETERIZATION LAB FOR IMPELLA MCS. UPON ARRIVAL TO THE CATHETERIZATION LAB THE PATIENT HAD A MEAN ARTERIAL PRESSURE OF 20. THE IMPELLA WAS PLACED VIA THE RIGHT FEMORAL ARTERY AND STARTED. THE PATIENT CONVERTED OUT OF ATRIAL FIBRILLATION AND MEAN ARTERIAL PRESSURE IMPROVED TO 80. DRIPS WERE THEN WEANED AND DECISION MADE TO STENT THE LEFT ANTERIOR DESCENDING ARTERY USING SINGLE ACCESS. ONE DRUG-ELUTING STENT WAS PLACED, AND THE PATIENT HEMODYNAMICALLY LOOKED MUCH IMPROVED. THE PATIENT WAS SENT TO THE UNIT TO REST AND RECOVER AND PLANS TO START CONTINUOUS RENAL REPLACEMENT THERAPY AS TOLERATED. APPROXIMATELY TWO DAYS AFTER START OF THE IMPELLA MCS, THE PATIENT ARRESTED FOR APPROXIMATELY THREE MINUTES BEFORE RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED; SUPRAVENTRICULAR TACHYCARDIA CONVERTED TO A WIDE QRS COMPLEX THEN ASYSTOLE. THE PATIENT'S OXYGENATION HAD ALSO DECLINED. THE NEXT MORNING AN ECHOCARDIOGRAM DEMONSTRATED NEW RIGHT VENTRICULAR FAILURE, WHICH MADE IT DIFFICULT TO FLOW ABOVE IMPELLA SUPPORT LEVEL P6-P7. THE PLAN WAS TO RESTART CRRT, REPLACE THE SWAN, AND LIGHTEN SEDATION TO BETTER ASSESS NEURO STATUS POST ARREST. THE PATIENT, HOWEVER, CONTINUED TO DECLINE THROUGHOUT THE DAY, BECAME INCREASINGLY HYPOTENSIVE AND EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22541120&pc=MFE | 2029214-2025-01628 | 2024-12-17 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-07-18 04:00 | MFE | ONYX | Adverse Event Without Identified Device or Use Problem | Hematoma |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ANGIOARCHITECTURE AND ENDOVASCULAR THERAPY OF INFANTILE DURAL ARTERIOVENOUS FISTULAS THE FOLLOWING MEDTRONIC DEVICES WERE USED: ONYX DEATHS OCCURRED IN THE STUDY POPULATION; HOWEVER, THERE AS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. THE CAUSE OF DEATH WAS ACUTE EPIDURAL HEMATOMA AT TORCULA. AMONG ALL PATIENT ADVERSE EVENTS / DEVICE PRODUCT PERFORMANCE ISSUES INCLUDED: ONE PATIENT DIED OF ACUTE EPIDURAL HEMATOMA AT TORCULA. THIS PATIENT DIED OF A SUDDEN FORAMEN MAGNUM HERNIATION 30 MIN AFTER THE TREATMENT. HEAD CT SHOWED AN ACUTE EPIDURAL HEMATOMA AT TORCULA. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: LIANG, S., LV, X.. ANGIOARCHITECTURE AND ENDOVASCULAR THERAPY OF INFANTILE DURAL ARTERIOVENOUS FISTULAS. NEUROLOGY INDIA 2024. DOI: 10.4103/NEUROL-INDIA.NEUROLINDIA-D-23-00527 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22541245&pc=PNJ | 9612164-2025-03523 | 2025-07-01 04:00 | Death | MEDTRONIC IRELAND | 2025-07-18 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE FOR THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE PATIENT EXPERIENCED A CARDIAC PERFORATION THAT LED TO A PERICARDIAL EFFUSION. THE PATIENT'S BLOOD PRESSURE DROPPED, AND AN ELECTROCARDIOGRAM (ECG) WAS PERFORMED. PERICARDIOCENTESIS AND RESUSCITATION EFFORTS WERE INITIATED. DESPITE THESE INTERVENTIONS, THE PATIENT PASSED AWAY. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6)). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22541254&pc=NPT | 2015691-2025-05845 | 2025-06-27 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-18 04:00 | NPT | COMMANDER DELIVERY SYSTEM | Difficult to Remove; Failure to Advance | Foreign Body In Patient | P140031 |  | 1 | Event Description: PER THE FIELD CLINICAL SPECIALIST (FCS), DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE UTILIZING A 23MM SAPIEN 3 ULTRA RESILIA VALVE VIA TRANSFEMORAL APPROACH, THE VALVE BECAME LODGED INFRARENALLY DUE TO CALCIFICATION AND COULD NOT BE ADVANCED OR RETRACTED. PERIPHERAL BALLOONS (6MM AND 8MM) WERE DEPLOYED VIA RADIAL ACCESS TO DILATE AROUND THE SAPIEN 3 VALVE, FOLLOWED BY THE USE OF AN 8MM LITHOTRIPSY BALLOON. A 2CC VOLUME WAS INTRODUCED INTO THE DISTAL BALLOON OF THE COMMANDER DELIVERY SYSTEM IN AN ATTEMPT TO ACT AS A BUMPER; HOWEVER, THE VALVE REMAINED STUCK INFRARENALLY. ATTEMPTS TO RETRACT THE VALVE INTO THE ESHEATH WERE UNSUCCESSFUL DUE TO VESSEL TORTUOSITY, WHICH WOULD NOT ALLOW THE VALVE TO LINE UP WITH THE SHEATH TO CAPTURE. A LUNDERQUIST WIRE WAS DEPLOYED VIA RADIAL ACCESS TO STRAIGHTEN THE ANATOMY AND FACILITATE VALVE ADVANCEMENT, BUT THIS MANEUVER ALSO FAILED. ULTIMATELY, THE VALVE WAS DEPLOYED IN THE DISTAL AORTA BELOW THE RENAL ARTERIES. AN APPARENT AORTIC RUPTURE OCCURRED. A TAMPONADE BALLOON WAS INFLATED SUPERIOR TO THE SAPIEN 3 VALVE AND THE RUPTURE SITE. A VBX STENT WAS DEPLOYED ACROSS THE SAPIEN 3 VALVE IN THE DISTAL AORTA AND EXTENDED INTO THE BILATERAL ILIAC ARTERIES. DESPITE THESE EFFORTS, BLEEDING PERSISTED. VASCULAR SURGERY WAS CONSULTED FOR ABDOMINAL EVACUATION AND AORTIC REPAIR. FOLLOW UP RECEIVED FROM THE FCS INDICATED THAT THE INJURY WAS AT THE DISTAL AORTA AND THE PHYSICIAN CHOSE TO STENT BILATERAL ILIACS ARTERIES AS WELL. ONCE THE VALVE WAS DEPLOYED IN THE AORTA, THE COMMANDER AND SHEATH CAME OUT WITHOUT ISSUE. THE PERCEIVED ROOT CAUSE WAS DUE TO CALCIUM AND NARROW DIAMETER OF THE DISTAL AORTA WHICH WAS ABOUT 10MM AND SEVERELY CALCIFIED. THE PATIENT WAS CONVERTED TO OPEN VASCULAR SURGERY AND EVERYTHING FROM THE CATH TABLE WAS DISCARDED. THE PATIENT'S FAMILY WITHDREW CARE THE NEXT DAY. Manufacturer Narrative: CORRECTION TO H6; TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSION. THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, DIFFICULTY OR INABILITY IN TRACKING THE DELIVERY SYSTEM TO THE LANDING ZONE, INCLUDING TRAVERSING THE AORTIC ARCH AND CROSSING THE NATIVE ANNULUS/BIOPROSTHESIS HAS NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE OF THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PATIENT FACTORS, POOR GUIDEWIRE MANAGEMENT, DAMAGED GUIDEWIRE, AND/OR EXCESSIVE MANIPULATION OF THE FLEX WHEEL. THE REPORTED EVENT OF TRACKING DIFFICULTY HAS BEEN CONFIRMED BASED ON THE INFORMATION AVAILABLE TO DATE. ANY UPDATES OR ADDITIONAL FINDINGS WILL BE SUBMITTED IN ACCORDANCE WITH FDA REPORTING REQUIREMENTS. AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (CALCIFICATION) LIKELY CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT THE VALVE BECAME LODGED INTRARENAL DUE TO CALCIFICATION AND COULD NOT BE ADVANCED OR RETRACTED. PATIENT FACTORS (I.E. CALCIFICATION, TORTUOSITY, SMALL VESSEL SIZE, PRE-EXISTING VASCULAR STENT) MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERE WITH PASSAGE OF THE DELIVERY SYSTEM AND CAUSES DIFFICULTY DURING TRACKING/CROSSING. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING DIFFICULTY/INABILITY WITHDRAWING CATHETER WITH CRIMPED VALVE THROUGH PATIENT VASCULATURE OR THROUGH THE SHEATH HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PATIENT FACTORS, NON-COAXIAL WITHDRAWAL, ALTERED CRIMP PROFILE, DAMAGED SHEATH, NOT CENTERING THE THV ON THE FLEX TIP, AND/OR WITHDRAWING THE THV THROUGH THE SHEATH HUB. THE REPORTED EVENT OF WITHDRAWAL DIFFICULTIES HAS BEEN CONFIRMED BASED ON THE INFORMATION AVAILABLE TO DATE. ANY UPDATES OR ADDITIONAL FINDINGS WILL BE SUBMITTED IN ACCORDANCE WITH FDA REPORTING REQUIREMENTS. AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (TORTUOSITY) AND/OR PROCEDURAL FACTORS (NON-COAXIAL WITHDRAWAL) LIKELY CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT "ATTEMPTS TO RETRACT THE VALVE INTO THE ESHEATH WERE UNSUCCESSFUL DUE TO VESSEL TORTUOSITY, WHICH WOULD NOT ALLOW THE VALVE TO LINE UP WITH THE SHEATH TO CAPTURE. A LUNDERQUIST WIRE WAS DEPLOYED VIA RADIAL ACCESS TO STRAIGHTEN THE ANATOMY AND FACILITATE VALVE ADVANCEMENT, BUT THIS MANEUVER ALSO FAILED. ULTIMATELY, THE VALVE WAS DEPLOYED IN THE DISTAL AORTA BELOW THE RENAL ARTERIES." IN THIS CASE, PATIENT FACTORS (I.E. CALCIFICATION, TORTUOSITY, OR SMALL VESSEL SIZE) MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERE WITH PASSAGE OF THE DELIVERY SYSTEM, POTENTIALLY LEADING TO NON-COAXIAL WITHDRAWAL, CAUSING THE THV TO INTERACT WITH THE SHEATH. THIS CAN RESULT IN WITHDRAWAL DIFFICULTY AS REPORTED. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22541618&pc=OZD | 1220648-2025-30486 | 2025-07-04 04:00 | Death | ABIOMED, INC. | 2025-07-18 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT EXPERIENCED MULTIPLE COMPLICATIONS DURING IMPELLA CP SUPPORT. AFTER ONE DAY OF SUPPORT, IT WAS NOTED THAT THE PATIENT EXPERIENCED BLEEDING ISSUES FROM THEIR ACCESS SITE, A GASTROINTESTINAL BLEED, AND BLOODY OUTPUT FROM THEIR NASOGASTRIC TUBING. BLOOD PRODUCTS WERE GIVEN TO ADDRESS THE NG OUTPUT AND GI BLEED. THE DRESSING WAS CHANGED AND A SANDBAG WAS PLACED TO MANAGE THE OOZING FROM THE ACCESS SITE. IN ADDITION, IT WAS NOTED THAT THE PATIENT'S IMPELLA LIMB WAS COOLER WITH NO PULSE IN THE RIGHT DORSALIS PEDIS OR POSTERIOR TIBIAL. A PLAN WAS MADE TO WEAN PRESSORS TO STIMULATE FLOW TO THE LEG. SUPPORT CONTINUED FOLLOWING THE INTERVENTIONS, HOWEVER THE PATIENT EXPIRED TWO DAYS LATER. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA CP WITH SMART ASSIST SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22541636&pc=BZD | 2518422-2025-107948 | 2025-07-17 04:00 | Death | RESPIRONICS, INC. | 2025-07-18 04:00 | BZD | RENTAL-DREAMSTATION AUTO CPAP, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION AUTO CPAP DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE EXAMINATION HAS NOT BEGUN BECAUSE THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22541736&pc=LWS | 2649622-2025-19565 | 2025-03-27 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-18 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Capturing Problem; Impedance Problem | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT A TRANSMISSION ALERT FROM THE PATIENT¿S DEVICE NOTED SUDDEN ELEVATED RIGHT VENTRICULAR (RV) LEAD IMPEDANCE AND RV THRESHOLDS. A NEIGHBOR WENT TO CHECK ON THE PATIENT, CALLED EMERGENCY MEDICAL SERVICES (EMS), AND THE PATIENT WAS PRONOUNCED DEAD. IT WAS FURTHER REPORTED THAT THE PATIENT MAY HAVE BEEN DECEASED AT THE TIME OF THE ALERT, BUT CONCLUSIVE EVIDENCE WAS UNKNOWN. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22542017&pc=MLV | 2135147-2025-03923 | 1999-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-18 04:00 | MLV | AMPLATZER PFO OCCLUDER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P120021 |  | 1 | Event Description: THE ARTICLE, "15-YEAR OUTCOMES OF PFO CLOSURE IN PATIENTS WITH CRYPTOGENIC EMBOLISM", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, MULTICENTER STUDY TO REPORT THE EXTENDED CLINICAL OUTCOMES OF PATIENTS WHO UNDERWENT TRANSCATHETER PATENT FORAMEN OVALE (PFO) CLOSURE FOR CRYPTOGENIC EMBOLISM. DEVICES INCLUDED IN THE STUDY WERE AMPLATZER PFO OCCLUDER, OCCLUTECH FIGULLA FLEX II, NMT CARDIOSEAL, CARDIA INTRASEPT, NMT BIOSTAR, ST. JUDE PREMERE, W.L. GORE CARDIOFORM SEPTAL OCCLUDER, COHEREX FLATSTENT, PFM NIT-OCCLUD, AND OTHER UNKNOWN. THE ARTICLE CONCLUDED THE STUDY CONFIRMED THE LONG-TERM EFFICACY AND SAFETY OF TRANSCATHETER PFO CLOSURE FOR PATIENTS WITH CRYPTOGENIC EMBOLISM AND PFO IN A REAL-WORLD SETTING. [THE PRIMARY AND CORRESPONDING AUTHOR WAS COSMO GODINO, CARDIOLOGY UNIT, SAN RAFFAELE HOSPITAL, IRCCS, VIA OLGETTINA 60, 20132 MILAN, ITALY, WITH CORRESPONDING E-MAIL: GODINO.COSMO@HSR.IT]. THE TIME FRAME OF THE STUDY WAS FROM 1999 TO 2013. THE TOTAL NUMBER OF PATIENTS INVOLVED IN THE STUDY WAS 1245, OF WHICH ONLY 1114 HAD AVAILABLE DEVICE DATA. OF THE 1114 PATIENTS WITH AVAILABLE DEVICE DATA, 908 (81.5%) RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 47 YEARS, AND THE MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES MELLITUS, SMOKING, HYPERLIPIDEMIA, DEEP VEIN THROMBOSIS, PULMONARY EMBOLISM, ORAL CONTRACEPTIVES, MIGRAINE, ATRIAL SEPTAL ANEURYSM. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF 15-YEAR OUTCOMES OF PFO CLOSURE IN PATIENTS WITH CRYPTOGENIC EMBOLISM WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES MELLITUS, SMOKING, HYPERLIPIDEMIA, DEEP VEIN THROMBOSIS, PULMONARY EMBOLISM, ORAL CONTRACEPTIVES, MIGRAINE, ATRIAL SEPTAL ANEURYSM. SOME OF THE COMPLICATIONS REPORTED WERE HEMORRHAGE, EMBOLISM, THROMBUS, STROKE AND DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: 15-YEAR OUTCOMES OF PFO CLOSURE IN PATIENTS WITH CRYPTOGENIC EMBOLISM: INSIGHTS FROM THE PROLONG REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22542036&pc=LWR | 2135147-2025-03925 | 2025-06-20 04:00 | Death | ABBOTT MEDICAL | 2025-07-18 04:00 | LWR | EPIC¿ VALVE (AORTIC) | Incomplete Coaptation; Patient Device Interaction Problem | Angina; Bradycardia; Dyspnea; Low Blood Pressure/ Hypotension; Unspecified Infection; Pneumonia; Pulmonary Edema; Obstruction/Occlusion; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P040021 |  | 1 | Event Description: THE ARTICLE, "RAPIDLY PROGRESSIVE OBSTRUCTIVE BIOPROSTHETIC VALVE THROMBOSIS AFTER SURGICAL AORTIC VALVE REPLACEMENT: A CASE REPORT", WAS REVIEWED. THE ARTICLE PRESENTED A CASE STUDY OF A 65-YEAR-OLD MALE PATIENT WITH ALCOHOLIC LIVER DISEASE, GOUT, CHRONIC KIDNEY DISEASE, AND PROSTATE CANCER. IT WAS REPORTED THAT ON AN UNKNOWN DATE, A 25MM EPIC AORTIC VALVE WAS CHOSEN FOR IMPLANT FOR AORTIC VALVE REPLACEMENT (AVR). ON POST-OPERATIVE DAY 113 (20 DAYS FOLLOWING WARFARIN DISCONTINUATION), THE PATIENT RETURNED WITH CHEST PAIN AND DYSPNEA AT REST. THE PATIENT ALSO EXHIBITED HYPOTENSION, WITH A BLOOD PRESSURE OF 80/59MMHG, A HEART RATE OF 115 BPM, AND A RESPIRATORY RATE OF 30/MIN. PHYSICAL EXAMINATION REVEALED PERIPHERAL COLDNESS AND CYANOSIS. PULMONARY AUSCULTATION DEMONSTRATED BILATERAL CRACKLES. THE PATIENT EXPERIENCED IN-HOSPITAL CARDIOPULMONARY ARREST (CPA) AND UNDERWENT EXTRACORPOREAL CARDIOPULMONARY RESUSCITATION (ECPR). PATIENT WAS DIAGNOSED WITH HEART FAILURE SECONDARY TO OBSTRUCTIVE BIOPROSTHETIC VALVE THROMBOSIS (BPVT). SUBSEQUENT COMPUTED TOMOGRAPHY IMAGING REVEALED SEVERE PULMONARY CONGESTION. A DECISION WAS MADE TO PERFORM REDO-AVR. THROMBUS FORMATION WAS OBSERVED ON THE EXPLANTED EPIC VALVE AT ALL THREE LEAFLETS, ESPECIALLY SEVERE ON THE LEFT CORONARY CUSP AND RIGHT CORONARY CUSP. A 25MM EDWARDS INSPIRIS RESILIA VALVE WAS IMPLANTED. ON POST-OPERATIVE DAY 7, HYPOTENSION DEVELOPED DUE TO INFECTION AND BLEEDING FROM THE WOUND AND ALVEOLI, NECESSITATING ADDITIONAL MECHANICAL CIRCULATORY SUPPORT. DUE TO THE PATIENT¿S BLEEDING TENDENCY, IMPELLA WAS AVOIDED, AND AN INTRA-AORTIC BALLOON PUMP (IABP) WAS SELECTED FOR HEMODYNAMIC SUPPORT. EXTRACORPOREAL MEMBRANE OXYGENATION WAS GRADUALLY WEANED BY POST-OPERATIVE DAY 10, AND THE IABP WAS REMOVED ON POST-OPERATIVE DAY 15. HOWEVER, THE PATIENT DEVELOPED MULTI-ORGAN FAILURE SECONDARY TO PNEUMONIA AND DIED ON POST-OPERATIVE DAY 25. [THE PRIMARY AUTHOR WAS SHIN HASEGAWA, DEPARTMENT OF CARDIOLOGY, NAGOYA TOKUSHUKAI GENERAL HOSPITAL, 2-52 KOZOJICHO KITA, KASUGAI, AICHI 487-0016, JAPAN. THE CORRESPONDING AUTHOR WAS SOH HOSOBA, DEPARTMENT OF CARDIOVASCULAR SURGERY, NAGOYA TOKUSHUKAI GENERAL HOSPITAL, 2-52 KOZOJICHO KITA, KASUGAI, AICHI 487-0016, JAPAN, WITH CORRESPONDING EMAIL: SOH.HOSOBA@GMAIL.COM]. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: RAPIDLY PROGRESSIVE OBSTRUCTIVE BIOPROSTHETIC VALVE THROMBOSIS AFTER SURGICAL AORTIC VALVE REPLACEMENT: A CASE REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22542078&pc=PYX | 1220648-2025-30491 | 2025-06-30 04:00 | Death | ABIOMED, INC. | 2025-07-18 04:00 | PYX | IMPELLA RP | Failure to Advance; Patient-Device Incompatibility | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P170011 |  | 1 | Event Description: IT WAS REPORTED THAT IMPLANTATION OF AN IMPELLA RP WAS COMPLICATED AS THE PATIENT WAS A HANCOCK IMPLANT WITH FOUR STENTS. THE TEAM HAD MANY ATTEMPTS AND TROUBLESHOOTING WITH MULTIPLE WIRE EXCHANGES. IMPLANTATION WAS ABORTED AND THE PATIENT EXPIRED. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF DELIVERY ISSUE IS DETERMINED TO BE PATIENT CONDITION BECAUSE PATIENT HAD UNFAVORABLE ANATOMY AND AFTER MULTIPLE ATTEMPTS AND AFTER CHANGE IN GUIDE WIRE AS PER CLINICAL DETAILS AND INSERTION WAS ABORTED. B1 PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30491. D6A AND D6B WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30491. G1 REPORTING CONTACT ADDRESS AND COUNTRY WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30491. G1 MANUFACTURING SITE NAME, ADDRESS, COUNTRY, AND FAX NUMBER WERE ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30491. H6 HEALTH EFFECT - CLINICAL CODE 4582 WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30491.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22542155&pc=DSQ | 2916596-2025-04809 | 2025-04-17 04:00 | Death | THORATEC CORPORATION | 2025-07-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY ON (B)(6) 2025. CAUSE OF DEATH REMAINED UNKNOWN, AND NO AUTOPSY WAS PERFORMED. THE PUMP OPERATED AS INTENDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. WHETHER THE PATIENT¿S OUTCOME WAS CONSIDERED TO BE DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. NO ALARMS WERE REPORTED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22542243&pc=DSQ | 2916596-2025-04800 | 2023-08-12 04:00 | Death | THORATEC CORPORATION | 2025-07-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2023 FROM RIGHT VENTRICLE FAILURE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION H6: HEALTH EFFECT - CLINICAL CODE CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED RIGHT HEART FAILURE AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22542574&pc=DSQ | 2916596-2025-04814 | 2024-09-30 04:00 | Death | THORATEC CORPORATION | 2025-07-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. NO AUTOPSY WAS PERFORMED. THE PUMP WAS OPERATING AS INTENDED. IT WAS NOT PROVIDED WHETHER OR NOT THIS WAS DEVICE RELATED. Manufacturer Narrative: SECTION A: DATE OF BIRTH CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, TO DATE, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22542578&pc=NPT | 2025587-2025-05134 | 2025-06-24 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-18 04:00 | NPT | EVOLUT FX PLUS VALVE | Difficult or Delayed Positioning; Insufficient Information; Patient Device Interaction Problem | Non specific EKG/ECG Changes; Hematoma; Low Blood Pressure/ Hypotension; Shock; Cardiac Tamponade; Cardiac Perforation; Multiple Organ Failure; Heart Block; No Clinical Signs, Symptoms or Conditions | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE ATTEMPTED IMPLANT OF TRANSCATHETER AORTIC VALVE, THERE WAS DIFFICULTY POSITIONING THE VALVE AND MULTIPLE RECAPTURES WERE PERFORMED. HOWEVER, THE VALVE WAS THEN IMPLANTED AND FOLLOWING THE IMPLANT, A LEFT BUNDLE BRANCH BLOCK (LBBB) OCCURRED. SUBSEQUENTLY, TEMPORARY PACING WAS INITIATED. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012659427); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NA ; EXPLANT DATE: NA PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012529123); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NA ; EXPLANT DATE: NA SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22542618&pc=NPT | 2015691-2025-05850 | 2025-06-27 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-18 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation | P140031 |  | 1 | Event Description: PER THE FIELD CLINICAL SPECIALIST (FCS), DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE UTILIZING A 23MM SAPIEN 3 ULTRA RESILIA VALVE VIA TRANSFEMORAL APPROACH, THE VALVE BECAME LODGED INFRARENALLY DUE TO CALCIFICATION AND COULD NOT BE ADVANCED OR RETRACTED. PERIPHERAL BALLOONS (6MM AND 8MM) WERE DEPLOYED VIA RADIAL ACCESS TO DILATE AROUND THE SAPIEN 3 VALVE, FOLLOWED BY THE USE OF AN 8MM LITHOTRIPSY BALLOON. A 2CC VOLUME WAS INTRODUCED INTO THE DISTAL BALLOON OF THE COMMANDER DELIVERY SYSTEM IN AN ATTEMPT TO ACT AS A BUMPER; HOWEVER, THE VALVE REMAINED STUCK INFRARENALLY. ATTEMPTS TO RETRACT THE VALVE INTO THE ESHEATH WERE UNSUCCESSFUL DUE TO VESSEL TORTUOSITY, WHICH WOULD NOT ALLOW THE VALVE TO LINE UP WITH THE SHEATH TO CAPTURE. A LUNDERQUIST WIRE WAS DEPLOYED VIA RADIAL ACCESS TO STRAIGHTEN THE ANATOMY AND FACILITATE VALVE ADVANCEMENT, BUT THIS MANEUVER ALSO FAILED. ULTIMATELY, THE VALVE WAS DEPLOYED IN THE DISTAL AORTA BELOW THE RENAL ARTERIES. AN APPARENT AORTIC RUPTURE OCCURRED. A TAMPONADE BALLOON WAS INFLATED SUPERIOR TO THE SAPIEN 3 VALVE AND THE RUPTURE SITE. A VBX STENT WAS DEPLOYED ACROSS THE SAPIEN 3 VALVE IN THE DISTAL AORTA AND EXTENDED INTO THE BILATERAL ILIAC ARTERIES. DESPITE THESE EFFORTS, BLEEDING PERSISTED. VASCULAR SURGERY WAS CONSULTED FOR ABDOMINAL EVACUATION AND AORTIC REPAIR. FOLLOW UP RECEIVED FROM THE FCS INDICATED THAT THE INJURY WAS LOCATED AT DISTAL AORTA AND THE PHYSICIAN CHOSE TO STENT BILATERAL ILIACS ARTERIES AS WELL. ONCE THE VALVE WAS DEPLOYED IN THE AORTA, THE COMMANDER AND SHEATH CAME OUT WITHOUT ISSUE. THE PERCEIVED ROOT CAUSE WAS DUE TO CALCIUM AND NARROW DIAMETER OF THE DISTAL AORTA WHICH WAS ABOUT 10MM AND SEVERELY CALCIFIED. PATIENT WAS CONVERTED TO OPEN VASCULAR SURGERY AND EVERYTHING FROM THE CATH LAB TABLE WAS DISCARDED. THE PATIENT'S FAMILY WITHDREW CARE THE NEXT DAY. Manufacturer Narrative: CORRECTION TO A1 AND H6: IMPACT CODE. DURING AN ADMINISTRATIVE REVIEW, IT WAS NOTICED THAT THE A1 AND H6 FIELDS WERE INADVERTENTLY LEFT OFF OF THE PREVIOUS SUBMISSION. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THAT PATIENT FACTORS (CALCIUM, TORTUOUS ANATOMY, AND NARROW DIAMETER OF THE DISTAL AORTA) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22543104&pc=BZD | 2518422-2025-048236 | 2023-03-16 04:00 | Death | RESPIRONICS, INC. | 2025-07-18 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22543154&pc=DSQ | 2916596-2025-04822 | 2025-07-09 04:00 | Death | THORATEC CORPORATION | 2025-07-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Dyspnea; Muscle Weakness; Asystole | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE HOSPICE TEAM SAW THE PATIENT THE NIGHT BEFORE DUE TO WORSENING SHORTNESS OF BREATH (SOB). THE PATIENT HAD NOT BEEN COMPLIANT WITH THEIR NEBULIZER TREATMENTS. THEY WERE VERY WEAK. IT TOOK 2 PEOPLE TO GET THEM BACK TO BED FROM THE BATHROOM. NO LEFT VENTRICULAR ASSIST DEVICE (LVAD) ALARMS AT THAT TIME. THE HOSPICE NURSE LEFT AT 22:30 AND THE PATIENT WAS STABLE AND IN BED ON THEIR BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP). THE PATIENT'S SPOUSE SAW THEM AT 01:30 TO HELP GIVE THEM MEDICATIONS AND REPOSITION FOR COMFORT. NO LVAD ALARMS AT THAT TIME. THE SPOUSE NOTED AROUND 04:04 THAT THE LVAD HAD ALARMED AND WENT TO THEIR ROOM. THEY FOUND THE PATIENT FACE DOWN ON THE FLOOR UNRESPONSIVE. THE SPOUSE CALLED 911 AND THE HOSPICE TEAM. THEY DID NOT CALL THE LVAD TEAM. EMERGENCY MEDICAL SERVICES (EMS) CALLED LVAD AND WANTED GUIDANCE ON COMPRESSIONS. NO ORDERS FOR COMPRESSIONS OR RESUSCITATION EFFORTS AS THE PATIENT WAS ALREADY ON HOSPICE AND IT HAD BEEN 20+ MINUTES OF STRAIGHT ASYSTOLE. THE PATIENT REMAINED UNRESPONSIVE WITH LVAD ALARMING LOW FLOW. BY THE TIME EMS ARRIVED, THE PATIENT HAD NO SPONTANEOUS BREATH SOUNDS, NO HEART TONES AND PUPILS WERE FIXED AND DILATED. ENDCO2 WAS DOWN TO 18 AND FLOW FROM LVAD WAS AT 0.8 LITERS PER MINUTE (LPM). THE LVAD PUMP WAS DEACTIVATED, AND TIME OF DEATH WAS NOTED TO BE 04:41. PUMP HAD APPEARED TO HAVE BEEN WORKING UP UNTIL THE PATIENT WAS FOUND DOWN. THE REASON FOR DEATH WAS CARDIOPULMONARY FAILURE. NO EQUIPMENT WOULD BE RETURNED AS THE PATIENT HAD BEEN CREMATED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22543330&pc=DSQ | 2916596-2025-04791 | 2024-03-23 04:00 | Death | THORATEC CORPORATION | 2025-07-18 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Head Injury; Intracranial Hemorrhage | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED ON (B)(6) 2024 FROM SEVERE HEAD INJURY/HEMORRHAGE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22543341&pc=DSQ | 2916596-2025-04803 | 2025-07-07 04:00 | Death | THORATEC CORPORATION | 2025-07-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Pulmonary Dysfunction; Renal Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH PATIENT OUTCOME THAT THE PATIENT PASSED AWAY DUE TO PULMONARY FAILURE AND RENAL FAILURE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE, RENAL FAILURE, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: THERE WERE NO OTHER UNUSUAL EVENTS RECORDED IN THE LOG FILE EVENT HISTORY. THE MECHANICAL CIRCULATORY SUPPORT EQUIPMENT APPEARED TO BE OPERATING AS INTENDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22543497&pc=NVN | 2124215-2025-48118 | 2025-06-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-18 04:00 | NVN | INGEVITY+ | Failure to Capture | Perforation of Vessels; Heart Block | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) SYSTEM UNDERWENT A REVISION PROCEDURE. IT HAD BEEN REPORTED IN JANUARY THAT THE DEVICE WAS NEARING ELECTIVE REPLACEMENT INDICATOR (ERI) BATTERY STATUS AND THAT THE RV LEAD AND LEFT VENTRICULAR (LV) LEADS HAD HIGH OUTPUTS PROGRAMMED. ADDITIONALLY, THE RV LEAD EXHIBITED LOSS OF CAPTURE AT HIGH OUTPUTS IN BOTH UNIPOLAR AND BIPOLAR CONFIGURATIONS. THE RV LEAD WAS EXTRACTED WITH NO ISSUES, BUT THEN 22.5 MINUTES POST-LASER THE PATIENT'S CONDITION DETERIORATED. A TEAR TO THE SUPERIOR VENA CAVA (SVC) WAS DETERMINED AND THE PATIENT DIED BEFORE THE CHEST COULD BE OPENED TO ATTEMPT A REPAIR OF THE SVC. THE EXPLANTED PRODUCTS WERE EXPECTED TO BE RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22543532&pc=DQY | 2124215-2025-47688 | 2025-04-28 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-18 04:00 | DQY | CROSSBOSS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED VIA LITERATURE ARTICLE THAT PATIENT DEATH OCCURRED. A COMPARATIVE ANALYSIS OF ANTEGRADE DISSECTION AND RE-ENTRY CTO PCIS FROM THE PROGRESS-CTO REGISTRY FROM 50 CENTERS (2012-2024) WAS PERFORMED. THREE DISSECTION STRATEGIES WERE COMPARED: (A) KNUCKLE WIRING: KNUCKLE WIRES WITHOUT CARLINO TECHNIQUE OR CROSSBOSS CATHETER; (B) THE CROSSBOSS TECHNIQUE: USE OF THE CROSSBOSS CATHETER; AND (C) THE CARLINO TECHNIQUE: USE OF THE CARLINO TECHNIQUE. IN-STENT CTO PCIS WERE EXCLUDED FROM THE ANALYSIS. IN TOTAL, 1575 (74.6%) CASES USED KNUCKLE WIRING, 427 (20.2%) THE CROSSBOSS, AND 110 (5.2%) THE CARLINO TECHNIQUE. KNUCKLE WIRING WAS THE MOST FREQUENT STRATEGY WITH INCREASING UTILIZATION OVER TIME. THE CROSSBOSS CATHETER WAS MORE COMMON IN LESIONS WITH LOWER J-CTO SCORES AND WAS ASSOCIATED WITH HIGHER TECHNICAL SUCCESS AND SIMILAR MAJOR ADVERSE CARDIAC EVENTS RATES, BUT LOWER PERFORATION RATES. TIME TO CROSSING WAS LONGER WHEN CARLINO WAS USED. USE OF DIFFERENT TYPES OF KNUCKLE WIRES WAS ASSOCIATED WITH SIMILAR SUCCESS RATES. KNUCKLE WIRING IS THE MOST COMMONLY USED ANTEGRADE DISSECTION STRATEGY. THE CROSSBOSS CATHETER WAS USED IN LESS COMPLEX CASES AND WAS ASSOCIATED WITH HIGHER SUCCESS, WHEREAS THE OPPOSITE WAS TRUE FOR THE CARLINO TECHNIQUE. THE IN-HOSPITAL OUTCOMES INDICATED THAT OF THE 426 CROSSBOSS PATIENTS, 3 DEATHS OCCURRED. Manufacturer Narrative: B3 DATE OF EVENT: USED THE DATE THAT THE JOURNAL ARTICLE WAS PUBLISHED. ALEXANDROU M, STREPKOS D, CARVALHO PEP, MUTLU D, SER OS, POOMMIPANIT P, GORGULU S, KHELIMSKII D, KRESTYANINOV O, AHMAD Y, JAMIL Y, ALASWAD K, BASIR MB, AZZALINI L, KEARNEY KE, KHATRI JJ, YOUNG L, OZDEMIR R, ULUGANYAN M, RAJ LM, KUMAR S, MASTRODEMOS OC, RANGAN BV, JALLI S, BURKE MN, SANDOVAL Y, BRILAKIS ES. DISSECTION TECHNIQUES IN CHRONIC TOTAL OCCLUSION PERCUTANEOUS CORONARY INTERVENTION. CATHETER CARDIOVASC INTERV. 2025 JUL;106(1):540-549. DOI: 10.1002/CCD.31573. EPUB 2025 MAY 11. PMID: 40350793. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22543545&pc=QZI | 2124215-2025-48031 | 2025-05-21 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-18 04:00 | QZI | FARAWAVE | Off-Label Use | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A PULSED FIELD ABLATION (PFA) PROCEDURE WAS COMPLETED TO TREAT AN OFF-LABEL PERSISTENT ATRIAL FIBRILLATION. PULMONARY VEINS, POSTERIOR WALL AND 2 CARDIOVERSIONS WERE PERFORMED SUCCESSFULLY AND THE PROCEDURE WAS COMPLETED WITHOUT ISSUES. THE PATIENT WAS DISCHARGED THE SAME DAY. AROUND TWO WEEKS AFTER THE PROCEDURE, THE PATIENT WAS FOUND UNRESPONSIVE AT HOME, RESUSCITATION WAS ATTEMPTED BUT WITHOUT SUCCESS.Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A PULSED FIELD ABLATION (PFA) PROCEDURE WAS COMPLETED TO TREAT AN OFF-LABEL PERSISTENT ATRIAL FIBRILLATION. PULMONARY VEINS, POSTERIOR WALL AND 2 CARDIOVERSIONS WERE PERFORMED SUCCESSFULLY AND THE PROCEDURE WAS COMPLETED WITHOUT ISSUES. THE PATIENT WAS DISCHARGED THE SAME DAY. AROUND TWO WEEKS AFTER THE PROCEDURE, THE PATIENT WAS FOUND UNRESPONSIVE AT HOME, RESUSCITATION WAS ATTEMPTED BUT WITHOUT SUCCESS. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. UPDATED FIELD H6 PATIENT CODE: APPROPRIATE TERM / CODE NOT AVAILABLE E2402 REPLACED WITH NO CLINICAL SIGNS, SYMPTOMS OR CONDITIONS E2403. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22543772&pc=NGV | 2124215-2025-46511 | 2025-07-07 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-18 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT DEATH OCCURRED. A CONCOMITANT PULMONARY VEIN ISOLATION (PVI) / POSTERIOR WALL (PW) ABLATION WITH LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. FARAPULSE WAS USED FOR THE ABLATION. TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) IMAGING NOTED THERE A SMALL BASELINE PERICARDIAL EFFUSION (PE) THAT REMAINED UNCHANGED THROUGHOUT THE ENTIRE PROCEDURE. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) WAS INSERTED AND USED TO DELIVER A 27MM WATCHMAN FLX PRO CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) TO THE LAA. THE WDS WAS DEPLOYED ONCE, WITH NO RECAPTURES, AND THE CLOSURE DEVICE WAS IMPLANTED. COMPRESSION WAS AN AVERAGE OF 30% WITH ZERO (0) SHOULDER AT ALL ANGLES. THERE WERE NO COMPLICATIONS INTRAPROCEDURE AND THE PROCEDURE WAS CONCLUDED. A LIMITED ECHOCARDIOGRAM WAS PERFORMED PRIOR THE SAME-DAY PATIENT DISCHARGE WHICH WAS NORMAL. THE FOLLOWING MORNING POST INDEX PROCEDURE, THE PATIENT WAS FOUND DEAD BY HIS FAMILY IN HIS HOME. CAUSE OF DEATH IS UNKNOWN. IN THE PHYSICIANS OPINION, THE DEATH IS LIKELY RELATED TO THE ABLATION PORTION OF THE PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22543801&pc=QZI | 2124215-2025-48398 | 2025-06-25 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-18 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A PULSED FIELD ABLATION (PFA) PROCEDURE WAS PERFORMED WITHOUT ISSUES TO TREAT AN ATRIAL FIBRILLATION. THE ISSUE OCCURRED THE MORNING AFTER THE PROCEDURE, BEFORE DISCHARGING THE PATIENT. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22543811&pc=DSQ | 2916596-2025-04810 | 2025-07-11 04:00 | Death | THORATEC CORPORATION | 2025-07-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Hemolysis; Hemoptysis; Hemorrhage/Bleeding; Shock; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS READMITTED REQUIRING INTUBATION WITH ONGOING INPATIENT CARE.Event Description: IT WAS LATER REPORTED THAT LOG FILES WERE SUBMITTED SINCE THE PATIENT WAS COMING WITH HEMOPTYSIS, WITH CONTINUOUS LOW FLOW ALARM DESPITE FLUID RESUSCITATION. THEY HAD PRIOR EXTRINSIC OUTFLOW GRAFT OBSTRUCTION (EOGO) STENT PLACED IN 2024 (MANUFACTURER REPORT NUMBER 2916596-2024-02593). THERE WAS CONCERN ABOUT PUMP THROMBOSIS. THE LOG CAPTURED MULTIPLE LOW FLOW ALARMS AS WELL AS BRIEF LOW FLOW ESTIMATES ON (B)(6) 2025. THE ESTIMATED FLOWS WERE AVERAGING FROM 1.5 TO 2.4 LITER PER MINUTE (LPM) AND PULSATILITY INDEX (PI) WAS AVERAGING FROM 6.2 TO 10 DURING THESE EVENTS. EVENT LOG ALSO CAPTURED LOW SPEED ADVISORY THE SET SPEED WAS CHANGE TO 4000 REVOLUTION PER MINUTE (RPM) WITH LOW-SPEED SETTING OF 4300-4900 RPM AT 14:41. ULTIMATELY, THE PATIENT TRANSITIONED TO COMFORT CARE AND PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22543944&pc=KNT | 9611594-2025-00162 | 2025-03-06 05:00 | Death | AVANOS MEDICAL INC. | 2025-07-18 04:00 | KNT | MIC KEY LOW PROFILE TUBE | Malposition of Device | Pain; Sepsis; Peritonitis; Pallor | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED VIA FDA MEDWATCH / FDA USER FACILITY REPORT MW REPORT MW5168243 THE FOLLOWING INFORMATION: THE PATIENT'S PEG (PERCUTANEOUS ENDOSCOPIC GASTROSTOMY) WAS REMOVED AND REPLACED WITH A MICKEY BUTTON PER PROTOCOL WITHOUT COMPLICATION (MORE FORCE THAN USUAL WAS NEEDED, HOWEVER). PATIENT TOLERATED THIS WELL. THE PATIENT'S CAREGIVER CALLED LATER, STATING THE PATIENT WAS HAVING PAIN AND LOOKED PALE. PER REVIEW OF THE NOTES, THE PATIENT WAS EVALUATED AT HIS LOCAL ER AND THEN ADMITTED TO HIS LOCAL HOSPITAL FOR CONTINUED CARE AND TREATMENT. A CT SCAN REVEALED FINDINGS CONSISTENT WITH TUBE FEEDINGS WITHIN THE ABDOMINAL CAVITY. THE PATIENT DEVELOPED SEPSIS AND PERITONITIS AND PASSED AWAY. Manufacturer Narrative: THE ACTUAL COMPLAINT PRODUCT WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD IS NOT POSSIBLE AS NO LOT NUMBER WAS PROVIDED; THEREFORE, THE UDI-PI IS UNAVAILABLE. THE LOT NUMBER WAS NOT PROVIDED BY THE CUSTOMER. ROOT CAUSE COULD NOT BE DETERMINED. ALL INFORMATION REASONABLY KNOWN AS OF 18-JUL-2025 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AVANOS MEDICAL, INC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AVANOS MEDICAL, INC. AVANOS MEDICAL, INC. HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AVANOS MEDICAL, INC. COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN AVANOS MEDICAL, INC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22543993&pc=KNT | 3006646024-2025-00018 | 2025-03-06 05:00 | Death | AVANOS MEDICAL INC. | 2025-07-18 04:00 | KNT | KIT INTRODUCER PEG 20FR PULL | Malposition of Device | Pain; Sepsis; Peritonitis; Pallor | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED VIA FDA MEDWATCH / FDA USER FACILITY REPORT MW REPORT MW5168243 RECEIVED ON 19JUN2025, THE FOLLOWING INFORMATION: THE PATIENT'S PEG (PERCUTANEOUS ENDOSCOPIC GASTROSTOMY) WAS REMOVED AND REPLACED WITH A MICKEY BUTTON PER PROTOCOL WITHOUT COMPLICATION (MORE FORCE THAN USUAL WAS NEEDED, HOWEVER). PATIENT TOLERATED THIS WELL. THE PATIENT'S CAREGIVER CALLED LATER, STATING THE PATIENT WAS HAVING PAIN AND LOOKED PALE. PER REVIEW OF THE NOTES, THE PATIENT WAS EVALUATED AT HIS LOCAL ER AND THEN ADMITTED TO HIS LOCAL HOSPITAL FOR CONTINUED CARE AND TREATMENT. A CT SCAN REVEALED FINDINGS CONSISTENT WITH TUBE FEEDINGS WITHIN THE ABDOMINAL CAVITY. THE PATIENT DEVELOPED SEPSIS AND PERITONITIS AND PASSED AWAY. Manufacturer Narrative: THE ACTUAL COMPLAINT PRODUCT WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD IS NOT POSSIBLE AS NO LOT NUMBER WAS PROVIDED; THEREFORE, THE UDI-PI IS UNAVAILABLE. THE LOT NUMBER WAS NOT PROVIDED BY THE CUSTOMER. ROOT CAUSE COULD NOT BE DETERMINED. ALL INFORMATION REASONABLY KNOWN AS OF 18-JUL-2025 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AVANOS MEDICAL, INC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AVANOS MEDICAL, INC. AVANOS MEDICAL, INC. HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AVANOS MEDICAL, INC. COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN AVANOS MEDICAL, INC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22544064&pc=DSQ | 2916596-2025-04843 | 2025-07-12 04:00 | Death | THORATEC CORPORATION | 2025-07-18 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICULAR FAILURE. IT WAS NOTED THAT THE PUMP WAS WORKING AS INTENDED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO AUTOPSY WAS PERFORMED, AND THE PUMP WAS NOT EXPLANTED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER "RIGHT HEART FAILURE"), DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22544188&pc=CHL | 3002637618-2025-00055 | 2025-07-06 04:00 | Death | EPOCAL INC. | 2025-07-18 04:00 | CHL | EPOC BLOOD ANALYSIS SYSTEM | High Test Results; Erratic Results | Insufficient Information | K200107 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT THEY RECEIVED A DISCREPANT HIGH HCO3 RESULT OF 12.8 ON A SINCE DECEASED PATIENT THAT DID NOT MATCH CLINICAL PICTURE OF THE PATIENT. THE CUSTOMER DOES NOT KNOW IF THE PATIENT WAS TREATED ON THE DISCREPANT RESULT. REPEATED TESTING OF A NEW SAMPLE GAVE A RESULT OF -17.3. THE CUSTOMER HAS NOT PROVIDED WHAT THEIR EXPECTED RESULT WAS FOR THIS PATIENT. Manufacturer Narrative: SIEMENS HAS REQUESTED MORE INFORMATION AND INSTRUMENT FILES FOR FURTHER INVESTIGATION. INVESTIGATION WILL COMMENCE ONCE INFORMATION HAS BEEN RECEIVED. THE CUSTOMER HAS NOT ALLEGED THAT THE DISCREPANT RESULT CAUSED OR CONTRIBUTED TO THE PATIENT DEATH. HOWEVER, SIEMENS IS FILING THIS REPORT IN AN ABUNDANCE OF CAUTION DUE TO THE PASSING OF THE PATIENT. THE CAUSE OF THIS EVENT IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22544191&pc=DSQ | 2916596-2025-04586 | 2025-06-30 04:00 | Death | THORATEC CORPORATION | 2025-07-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AFTER THEY EXPERIENCED LOW FLOW ALARMS DUE TO LOW BLOOD PRESSURE. ADDITIONAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22545396&pc=DSQ | 2916596-2025-04804 | 2020-08-07 04:00 | Death | THORATEC CORPORATION | 2025-07-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Intracranial Hemorrhage; Tachycardia; Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2020 DUE TO VENTRICULAR TACHYCARDIA ARREST AND AN INTRACRANIAL HEMORRHAGE. THE DEVICE WAS OPERATING AS EXPECTED AND WILL NOT BE RETURNED.Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2020 DUE TO RESPIRATORY ARREST, VENTRICULAR TACHYCARDIA ARREST, AND AN INTRACRANIAL HEMORRHAGE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: B5 - DESCRIBE EVENT OR PROBLEM - CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA, STROKE, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. SECTION 6 OF THE IFU ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22545864&pc=DSQ | 2916596-2025-04867 | 2025-07-12 04:00 | Death | THORATEC CORPORATION | 2025-07-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Bradycardia; Cardiac Arrest; Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS RECENTLY IMPLANTED AND WAS RECOVERING WELL ON THE STEPDOWN UNIT. THE PATIENT WAS EXTUBATED ON POD (POST-OPERATIVE DAY) 0. ON (B)(6) 2025, AT ABOUT 15:30 PATIENT HAD NEW ALTERED MENTAL STATUS (AMS). NARCAN WAS ADMINISTERED SINCE THE PATIENT HAD RECENTLY BEEN GIVEN OXYCODONE. SHORTLY THEREAFTER, THE PATIENT BEGAN HAVING AGONAL BREATHING AND LOW FLOW ALARMS WITH HEART RATE IN THE 40S AND WAS STAT CALLED. THE PATIENT CODED ON (B)(6) 2025 AT 15:51. THE PERMANENT PACEMAKER (PPM) LOST CAPTURE WITH RATES 0-40. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. A HEAD COMPUTED TOMOGRAPHY (CT) SHOWED SEVERAL ACUTE INFARCTIONS IN A WATERSHED TERRITORY. BEDSIDE VENTRICULAR ASSIST DEVICE (VAD) INTERROGATION WAS UNREVEALING. LOG FILES WERE SENT FOR REVIEW. THE EVENT LOG CAPTURED ROUTINE EVENTS. THE VAD AND PERIPHERAL EQUIPMENT APPEARED TO BE FUNCTIONING AS INTENDED. NO LOW FLOW ALARMS NOTED IN THE EVENT LOG. THE PERIODIC LOG CAPTURED A DIP IN FLOW VALUES AND A LOW FLOW EVENT BACK ON (B)(6) 2025 AT 21:14 WITH FLOW NOTED AT 2.3 LITERS PER MINUTE (LPM). CARE WAS WITHDRAWN DUE TO DEVASTATING CEREBROVASCULAR ACCIDENT (CVA) NOTED ON (B)(6) 2025 AND THE PATIENT PASSED PEACEFULLY THAT EVENING. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22545885&pc=CHL | 3002637618-2025-00056 | 2025-07-06 04:00 | Death | EPOCAL INC. | 2025-07-18 04:00 | CHL | EPOC BLOOD ANALYSIS SYSTEM | Low Test Results | Insufficient Information | K200107 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT THEY RECEIVED A DISCREPANT LOW HCO3 RESULT OF -17.3 ON A SINCE DECEASED PATIENT THAT DID NOT MATCH CLINICAL PICTURE OF THE PATIENT. THE CUSTOMER DOES NOT KNOW IF THE PATIENT WAS TREATED ON THE DISCREPANT RESULT. REPEATED TESTING OF A NEW SAMPLE GAVE A RESULT OF 12.8. THE CUSTOMER HAS NOT PROVIDED WHAT THEIR EXPECTED RESULT WAS FOR THIS PATIENT. NOTE THIS IS THE SAME PATIENT AS MDR 3002637618-2025-00055. THIS IS FOR THE OTHER INSTRUMENT. Manufacturer Narrative: SIEMENS HAS REQUESTED MORE INFORMATION AND INSTRUMENT FILES FOR FURTHER INVESTIGATION. INVESTIGATION WILL COMMENCE ONCE INFORMATION HAS BEEN RECEIVED. THE CUSTOMER HAS NOT ALLEGED THAT THE DISCREPANT RESULT CAUSED OR CONTRIBUTED TO THE PATIENT DEATH. HOWEVER, SIEMENS IS FILING THIS REPORT IN AN ABUNDANCE OF CAUTION DUE TO THE PASSING OF THE PATIENT. THE CAUSE OF THIS EVENT IS UNKNOWN. NOTE THIS IS THE SAME PATIENT AS MDR 3002637618-2025-00055.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22546059&pc=MAE | 2135147-2025-03972 | 2021-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-18 04:00 | MAE | AMPLATZER PICCOLO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P020024 |  | 1 | Event Description: THE ARTICLE, "PICCOLO DEVICE FOR TREATING PATENT DUCTUS ARTERIOSUS BEYOND PREMATURE NEWBORN PATIENTS, A NOVEL EXPERIENCE IN SOUTH AMERICA: A RETROSPECTIVE STUDY", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO PRESENT THE EXPERIENCE WITH THE AMPLATZER PICCOLO DEVICE IN A PEDIATRIC COHORT AT A REFERRAL CENTER IN COLOMBIA AND TO COMPARE THESE RESULTS WITH THE INFORMATION AVAILABLE IN SIMILAR STUDIES. DEVICES USED IN THE STUDY WERE SOLELY AMPLATZER PICCOLO. THE ARTICLE CONCLUDED THAT FINDINGS SUGGEST A POTENTIAL EXTENSION OF THE UTILITY OF THE PICCOLO DEVICE BEYOND THE NEONATAL PHASE, DEMONSTRATING ITS SAFETY AND EFFICACY IN OLDER PEDIATRIC PATIENTS WEIGHING OVER 2 KG, WITHOUT A SIGNIFICANT INCREASE IN COMPLICATIONS OR MORTALITY RATES. [THE PRIMARY AUTHOR WAS ANA ARISTIZABAL, FUNDACIÓN VALLE DEL LILI, DEPARTAMENTO MATERNO INFANTIL, SERVICIO DE CARDIOLOGÍA PEDIÁTRICA, CALI, COLOMBIA. THE CORRESPONDING AUTHOR WAS CARLOS GUZMÁN-SERRANO, FUNDACIÓN VALLE DEL LILI, CENTRO DE INVESTIGACIONES CLÍNICAS, CALI, COLOMBIA] THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2021 TO DECEMBER 2023. A TOTAL OF 29 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 24 MONTHS, THE AVERAGE WEIGHT WAS 10.4KG, AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED PRE-TERM BIRTH, PATENT DUCTUS ARTERIOSUS, DOWN SYNDROME, OTHER CONCOMITANT CARDIAC DEFECT, BICUSPID AORTIC VALVE, ATRIAL SEPTAL DEFECT, VENTRICULAR SEPTAL DEFECT, RESPIRATORY DISTRESS. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF PICCOLO DEVICE FOR TREATING PATENT DUCTUS ARTERIOSUS BEYOND PREMATURE NEWBORN PATIENTS, A NOVEL EXPERIENCE IN SOUTH AMERICA WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PRE-TERM BIRTH, PATENT DUCTUS ARTERIOSUS, DOWN SYNDROME, OTHER CONCOMITANT CARDIAC DEFECT, BICUSPID AORTIC VALVE, ATRIAL SEPTAL DEFECT, VENTRICULAR SEPTAL DEFECT, RESPIRATORY DISTRESS. SOME OF THE COMPLICATIONS REPORTED WERE INFLAMMATION, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: PICCOLO DEVICE FOR TREATING PATENT DUCTUS ARTERIOSUS BEYOND PREMATURE NEWBORN PATIENTS, A NOVEL EXPERIENCE IN SOUTH AMERICA: A RETROSPECTIVE STUDY. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED B3: DATE OF EVENT IS ESTIMATED D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: PICCOLO DEVICE FOR TREATING PATENT DUCTUS ARTERIOSUS BEYOND PREMATURE NEWBORN PATIENTS, A NOVEL EXPERIENCE IN SOUTH AMERICA: A RETROSPECTIVE STUDY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22546068&pc=OZD | 1220648-2025-30514 | 2025-07-07 04:00 | Death | ABIOMED, INC. | 2025-07-18 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WAS ADMITTED FOR INFERIOR ST-ELEVATION MYOCARDIAL INFARCTION. THE IMPELLA WAS INSERTED INTO LEFT VENTRICLE AND POSITION WAS CONFIRMED. DURING SUPPORT, A HEMATOMA WAS OBSERVED. MANUAL PRESSURE WAS HELD IN ADDITION TO RE-SUTURING AND HEPARIN WAS STOPPED. THE GROIN WAS BRUISED. ULTIMATELY THE PATIENT EXPIRED ON SUPPORT Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ ¿POTENTIAL ADVERSE EVENTS (UNITED STATES) ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND VASCULAR INJURY¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22546077&pc=DSQ | 2916596-2025-04845 | 2025-06-01 04:00 | Death | THORATEC CORPORATION | 2025-07-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Electrical Power Problem; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cellulitis; Muscle Weakness; Vomiting; Dizziness; Malaise | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT IT WAS FOUND THAT THE PATIENT'S AORTIC VALVE WAS NOT OPENING EVERY BEAT. THE PATIENT NEEDED AND EXTENSIVE HOSPITAL STAY POST LEFT VENTRICULAR ASSIST DEVICE (LVAD) REQUIRING TRACHEOSTOMY AND REHABILITATION. THE PATIENT WAS DISCHARGED FOR A SHORT PERIOD OF TIME BEFORE BEING READMITTED TO THE EMERGENCY DEPARTMENT WITH FLUID OVERLOAD/ CELLULITIS BILATERAL LOWER EXTREMITIES (BLE). THAT REQUIRED EXTENSIVE DIURESIS. TH LVAD WAS SET TO 5800 AND THE PATIENT WAS STARTED ON ANTIBIOTICS AND BUMEX DRIP. THE PATIENT WAS THE TRANSITIONED TO INPATIENT REHAB WHERE THEY SPENT 3 WEEKS AND DID VERY WELL. NO CHANGES TO LVAD PARAMETERS OR LOW FLOWS WERE NOTED DURING THEIR TIME IN REHAB. IT WAS NOTED THAT PATIENT LOST SIGNIFICANT AMOUNT OF WEIGHT AND AMBULATING MORE THAN PRIOR TO LVAD SURGERY. THE PATIENT WAS INITIAL SET TO BE DISCHARGED ON (B)(6) 2025 BUT ON (B)(6) 2025 THE PATIENT FELT SICK AND HAD MEAN ARTERIAL PRESSURES (MAPS) IN THE 50S, LOW FLOWS, DIZZINESS, WEAKNESS, GENERAL MALAISE, AND EMESIS. THEY WERE GIVEN 500ML FLUID WITH NO CHANGE. PATIENT HAD EMESIS AND FEELINGS OF IMPENDING DOOM. THE PATIENT WAS TRANSFERRED TO THE EMERGENCY DEPARTMENT AND ANOTHER 500 ML NORMAL SALINE GIVEN, AND THEIR DOPPLER MAP 72. CARDIOLOGY ASSESSED THE PATIENT, AND IT WAS FOUND THAT THE PATIENT WAS VERY SOMNOLENT AND INTERMITTENTLY AROUSABLE AND DENIED CONSTRICTIVE PERICARDITIS (CP) WITHIN 10 MIN OF ARRIVAL. BEDSIDE ECHO WAS ATTEMPTED BUT LARGE PROJECTILE EMESIS OCCURRED, AND THE PATIENT IMMEDIATELY BECAME UNRESPONSIVE AND ULTIMATELY INTUBATED. DOBUTAMINE STARTED. CARDIAC STANDSTILL ON ECHO WITH VERY DILATED RIGHT VENTRICLE WITH NO PERICARDIAL EFFUSION. LOG FILES WERE SUBMITTED FOR REVIEW. THE EVENT LOG FILE CAPTURED PERSISTENT LOW FLOW EVENTS WITH THE FLOW NOTED TO BE AS LOW AS 0 LPM AT TIMES. MOST OF THESE LOWER FLOWS WERE ASSOCIATED WITH ELEVATED PULSATILITY INDEX (PI) VALUES. ALSO NOTED THAT WHEN THE ESTIMATED FLOWS WERE 0 THAT THE PI VALUES DROPPED AND NOTED A DROP IN PUMP POWER BY ABOUT A HALF WATT. IT WAS ALSO NOTED THAT THE DRIVELINE WAS DISCONNECTED THIS MORNING AT 0143. IT WAS THEN COMMUNICATED ON 01JUL2025 THAT THE REASON FOR THE DRIVELINE DISCONNECT WAS BECAUSE THE SYSTEM CONTROLLER WAS TURNED OFF DUE TO THE PATIENT PASSING AWAY. IT WAS SAID THERE WAS NO IMAGING COMPLETED PRIOR TO THE PATIENT PASSING. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22546240&pc=NPT | 2025587-2025-05154 | 2025-07-11 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-18 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Vascular Dissection; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT A TRANSCATHETER AORTIC VALVE WAS IMPLANTED IN A PATIENT. AFTER WITHDRAWAL OF THE DELIVERY SYSTEM AND HEMODYNAMIC ASSESSMENT, A DROP IN AORTIC PRESSURE WAS NOTED. ULTRASOUND EXAMINATION REVEALED A PERICARDIAL EFFUSION, WHICH WAS PUNCTURED AND DRAINED. SUBSEQUENT ANGIOGRAPHY AND TRANSESOPHAGEAL ECHOCARDIOGRAPHY CONFIRMED THE PRESENCE OF AN AORTIC DISSECTION. THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE AND REQUIRED RESUSCITATION. DURING CARDIOPULMONARY RESUSCITATION, THE PATIENT WAS EMERGENTLY TRANSFERRED TO CARDIAC SURGERY AND RECEIVED OPERATIVE CARE. THE PATIENT DIED AS A RESULT OF THE CONSEQUENCES OF THE IMPLANTATION.Event Description: IT WAS REPORTED THAT A TRANSCATHETER AORTIC VALVE WAS IMPLANTED IN A PATIENT. AFTER WITHDRAWAL OF THE DELIVERY SYSTEM AND HEMODYNAMIC ASSESSMENT, A DROP IN AORTIC PRESSURE WAS NOTED. ULTRASOUND EXAMINATION REVEALED A PERICARDIAL EFFUSION, WHICH WAS PUNCTURED AND DRAINED. SUBSEQUENT ANGIOGRAPHY AND TRANSESOPHAGEAL ECHOCARDIOGRAPHY CONFIRMED THE PRESENCE OF AN AORTIC DISSECTION. THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE AND REQUIRED RESUSCITATION. DURING CARDIOPULMONARY RESUSCITATION, THE PATIENT WAS EMERGENTLY TRANSFERRED TO CARDIAC SURGERY AND RECEIVED OPERATIVE CARE. THE PATIENT DIED AS A RESULT OF THE CONSEQUENCES OF THE IMPLANTATION. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT IT WAS UNKNOWN WHAT CAUSED THE AORTIC DISSECTION, SUBSEQUENT PERICARDIAL EFFUSION AND DEATH. PER THE PHYSICIAN, NEITHER THE MEDTRONIC VALVE OR DELIVERY CATHETER SYSTEM COULD BE EXCLUDED AS POTENTIAL CONTRIBUTING FACTORS. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT FX DCS; PRODUCT ID D-EVOLUTFX-2329 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0195-HEART VALVES; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. H6. UPDATED THE SYSTEM REPORTED DCS EVAL CODE METHOD - THE DEVICE WAS DISCARDED BY THE FACILITY. PRODUCT ANALYSIS IMAGE REVIEW: PROCEDURAL IMAGES INCLUDING A ONE 3-SECOND MOVIE CLIP, TRANSTHORACIC ECHOCARDIOGRAM (TTE) IMAGE OF DISSECTION SEEN IN AORTA, AND TWELVE ANGIOGRAPHIC MEDIA FILES WERE SUBMITTED TO MEDTRONIC FOR REVIEW. NO DICOM IMAGING WAS PROVIDED NOR WAS THE PATIENT¿S EXECUTIVE SUMMARY PROVIDED FOR ANATOMICAL REVIEW. THE SUBMITTED IMAGES SHOWED THE VALVE BEING IMPLANTED AND A POSSIBLE DISSECTION ORIGINATING ABOVE THE LEFT CORONARY CUSP, AND FINAL ANGIOGRAM REVEALED AN AORTIC DISSECTION IN THE ASCENDING AORTA. A PRE-PROCEDURE AORTOGRAM WAS NOT PROVIDED AND THE INFORMATION PROVIDED IS LIMITED; THEREFORE, IT WAS NOT POSSIBLE TO VALIDATE THE CAUSE OF THE DISSECTION. AS STATED IN THE DEVICE INSTRUCTIONS FOR USE, POTENTIAL RISKS ASSOCIATED WITH THE IMPLANTATION OF THE EVOLUT FX+ BIOPROSTHESIS MAY INCLUDE, BUT ARE NOT LIMITED TO, THE FOLLOWING: EMERGENT SURGICAL OR TRANSCATHETER INTERVENTION (FOR EXAMPLE, CORONARY ARTERY BYPASS, HEART VALVE REPLACEMENT, VALVE EXPLANT, PERCUTANEOUS CORONARY INTERVENTION [PCI], BALLOON VALVULOPLASTY); DEATH; CARDIAC TAMPONADE, CARDIOVASCULAR INJURY (INCLUDING RUPTURE, PERFORATION, TISSUE EROSION, OR DISSECTION OF VESSELS, ASCENDING AORTA TRAUMA, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION). CONCLUSION: THE INVESTIGATION REMAINS IN PROGRESS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22546393&pc=DSQ | 2916596-2025-04847 | 2025-07-10 04:00 | Death | THORATEC CORPORATION | 2025-07-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT WAS DISCHARGED TO HOSPICE CARE AND PASSED AWAY ON (B)(6) 2025. ADDITIONAL INFORMATION RECEIVED ON INDICATED THAT THE PATIENT SUFFERED FROM PSEUDOMONAS BACTEREMIA PRIOR TO THEIR DEATH.Event Description: THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: SECTION B5 CORRECTION. SECTION H6 HEALTH EFFECT - CLINICAL CODE CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22546487&pc=MVK | 3008642652-2025-06976 | 2025-06-17 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-18 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE STARTED UP AT 08:58:20 ON (B)(6) 2025. AT 09:04:56 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 09:05:33, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS ASYSTOLE. THE ELECTRODE BELT WAS DISCONNECTED AT 09:17:40 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22546521&pc=BZD | 2518422-2025-107961 | 2024-09-05 04:00 | Death | RESPIRONICS, INC. | 2025-07-18 04:00 | BZD | BIPAP PRO BIFLEX, W/SD CARD, US-RENTAL | Degraded | Insufficient Information | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY 4 YEARS AGO. MEDICAL INTERVENTION WAS NOT SPECIFIED. THERE IS NO REPORT OF ANY DEVICE MALFUNCTION RELATED TO THE PATIENT¿S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22546607&pc=MFE | 2029214-2025-01639 | 2023-08-28 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-07-18 04:00 | MFE | ONYX AVM | Adverse Event Without Identified Device or Use Problem | Aneurysm; Hematoma; Convulsion/Seizure; Movement Disorder; Respiratory Insufficiency; Swelling/ Edema | P030004 |  | 1 | Event Description: YANG, Y., MA, H., HU, H., LI, L.,TANG, J.; PEDIATRIC DISCOVERY; 2024; 2:E80; NEWLY FORMED AND RUPTURED INFECTIOUS INTRACRANIAL ANEURYSM IN FEW DAYS FOLLOWING FIRST INTRACRANIAL ANEURYSM EMBOLIZATION CONCURRENT WITH MIDDLE CEREBRAL ARTERY OCCLUSION: A REGRETTABLE CASE IN A BABY; DOI.ORG/10.1002/PDI3.80 LITERATURE WAS REVIEWED REGARDING: INFECTIOUS INTRACRANIAL ANEURYSM (IIA) AND ITS COMPLICATIONS IN CHILDREN SUFFERING FROM INFECTIVE ENDOCARDITIS. THE STUDY DISCUSSES A CASE REPORT OVER APPROXIMATELY ONE WEEK (FROM ADMISSION TO THE SEVENTH DAY POST-EMBOLIZATION). SPECIFIC BROADER TIME FRAMES FOR THE STUDY ARE NOT MENTIONED. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: ONYX 18. THE PATIENT IN THE STUDY DIED. THE BABY DIED THE DAY AFTER BEING TAKEN HOME, BUT SPECIFIC CAUSES OF DEATH ARE NOT DETAILED IN THE PROVIDED TEXT. AMONG PATIENT ADVERSE EVENTS INCLUDED: MAJOR SEIZURE WITH OPISTHOTONOS (SEIZURE ACTIVITY). LOW OXYGEN SATURATION. MASSIVE HEMATOMA IN THE RIGHT BASAL GANGLIA WITH SERIOUS EDEMA. BIFURCATE ANEURYSM OF THE RIGHT MIDDLE CEREBRAL ARTERY (MCA). DEATH NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22546831&pc=NGV | 2124215-2025-48493 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-18 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 4 | Event Description: THIS REPORT SUMMARIZES 4 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 64 TO 87. PATIENT SEX: 1 FEMALE; 3 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 495 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22546837&pc=NGV | 2124215-2025-48495 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-18 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR SURGICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025 AVERAGE TIME TO EVENT: 670 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22546954&pc=QMG | 3015053858-2025-00086 | 2025-06-19 04:00 | Death | SHOCKWAVE MEDICAL, INC. | 2025-07-18 04:00 | QMG | SHOCKWAVE C2+ CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200039/S008 |  | 1 | Event Description: A SHOCKWAVE C2+ CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER WAS USED TO TREAT A PATIENT WITH A KNOWN POOR EJECTION FRACTION (EF) AND CARDIAC CONDITION. APPROXIMATELY, THIRTY (30) CYCLES WERE DELIVERED. DURING THE PROCEDURE, THE PATIENT EXPERIENCED AN UNSPECIFIED EVENT THAT REQUIRED CHEST COMPRESSION. ONCE THE PATIENT WAS STABILIZED, THE PHYSICIAN WAS ABLE TO COMPLETE THE PROCEDURE BY STENT PLACEMENT. THE FOLLOWING DAY, THE PHYSICIAN REPORTED THAT THE PATIENT HAD PASSED AWAY POST-PROCEDURE. THE PHYSICIAN INDICATED MULTIPLE TIMES THAT THE USE OF THE IVL DEVICE DID NOT CONTRIBUTE TO THE PATIENT'S DEATH AND ATTRIBUTED THE OUTCOME TO PRE-EXISTING MEDICAL CONDITIONS. Manufacturer Narrative: THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION CANNOT BE PERFORMED. BASED ON THE REPORTED INFORMATION, NO INFORMATION WAS PROVIDED SUGGESTING THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. THERE WAS NO REPORTED IVL CATHETER ISSUE. SHOCKWAVE MEDICAL HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO BEING DISTRIBUTED. A REVIEW OF THE MANUFACTURING AND TEST DOCUMENTATION FOR SUBJECT LOT DOES NOT REVEAL ANY ISSUES WITH THE MANUFACTURING OF THE DEVICE. THE DEVICE PASSED ALL OF SHOCKWAVE MEDICAL, INC. ACCEPTANCE CRITERIA PRIOR TO SHIPPING. COMPLAINT TRENDS WILL CONTINUE TO BE MONITORED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22548038&pc=NIK | 1028232-2025-03412 | 2025-07-03 04:00 | Death | BIOTRONIK SE & CO. KG | 2025-07-18 04:00 | NIK | INTICA 7 HF-T DF-1 IS-1 PROMRI | Over-Sensing; Inappropriate/Inadequate Shock/Stimulation | Ventricular Fibrillation; Shock from Patient Lead(s) | P050023 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT RECEIVED SEVERAL SHOCKS. OVERSENSING WAS INITIALLY SUSPECTED, FOLLOWED BY ADDITIONAL INFORMATION THAT THE PATIENT EXPERIENCED ELECTROLYTE IMBALANCE, LEADING TO SEVERAL EPISODES OF VENTRICULAR FIBRILLATION. THE MEDICATION WAS ADAPTED. ON (B)(6), IT WAS REPORTED THAT THE PATIENT HAD SINCE PASSED AWAY. Manufacturer Narrative: THE ICD WAS NOT RETURNED FOR ANALYSIS. THE ANALYSIS IS THEREFORE BASED ON THE INSPECTION OF THE QUALITY DOCUMENTS ASSOCIATED WITH THE MANUFACTURE OF THIS PARTICULAR DEVICE AS WELL AS ON THE RETURNED DEVICE DATA. THE MANUFACTURING PROCESS FOR THIS DEVICE WAS REVIEWED. ALL PRODUCTION STEPS HAD BEEN PERFORMED ACCORDINGLY. THERE WAS NO SIGN OF ANY INCONSISTENCY DURING THE MANUFACTURING PROCESS. THE FINAL ACCEPTANCE TEST PROVED THE ICD FUNCTIONS TO BE AS SPECIFIED. THE RETURNED DEVICE DATA HAVE BEEN ANALYZED. THE ANALYSIS DID NOT SHOW ANY ANOMALIES. IN PARTICULAR, ALL DOCUMENTED EPISODES WERE TRIGGERED EXCLUSIVELY BY INTRINSIC HEART SIGNALS. NO NOISE WAS PRESENT IN THE RECORDED IEGMS. THE LAST DOCUMENTED EPISODE (NUMBER: (B)(4), FROM ON (B)(6) 2025 AT 02:38 H, WAS A VF WITH ONE ATP AND ONE 40 J SHOCK DELIVERY. THE IEGM OF THIS EPISODE DOCUMENTS LACK OF SENSED SIGNAL IN THE RA AND LV CHANNELS AND A VERY SMALL SENSED SIGNAL IN THE RV CHANNEL. ANALYSIS INDICATES A NORMAL AND EXPECTED FUNCTIONALITY OF THE ICD. IN CONCLUSION, THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE REVIEW OF THE QUALITY DOCUMENTS CONFIRMED A REGULAR DEVICE MANUFACTURING. THE ANALYSIS OF THE RETURNED DEVICE DATA REVEALED NO INDICATION OF AN ICD MALFUNCTION. Manufacturer Narrative: THE DEVICE IS CURRENTLY NOT AVAILABLE FOR ANALYSIS. NO CONCLUSION CAN BE DRAWN BASED ON AVAILABLE INFORMATION AT THE MOMENT. THE FILE IS CLOSED. THE INVESTIGATION WILL BE RE-OPENED SHOULD ADDITIONAL DATA BECOME AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22549016&pc=NIQ | 2124215-2025-45055 | 2025-06-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-18 04:00 | NIQ | PROMUS ELITE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE 90% STENOSED AND DEFUSE LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING ARTERY AND WAS DILATED WITH 2.5 X 15MM BALLOON FOLLOWED BY INTRAVASCULAR ULTRASOUND (IVUS) IN WHICH DISSECTION TO THE DISTAL LEFT MAIN WAS FOUND. A 38 X 3.00MM PROMUS ELITE MR DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT. THE PATIENT DEVELOPED HYPOTENSION AND WENT INTO CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS IMMEDIATELY STARTED. HOWEVER, ATTEMPTS WERE UNSUCCESSFUL AND THE PATIENT EXPIRED.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE 90% STENOSED AND DEFUSE LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING ARTERY AND WAS DILATED WITH 2.5 X 15MM BALLOON FOLLOWED BY INTRAVASCULAR ULTRASOUND (IVUS) IN WHICH DISSECTION TO THE DISTAL LEFT MAIN WAS FOUND. A 38 X 3.00MM PROMUS ELITE MR DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT. THE PATIENT DEVELOPED HYPOTENSION AND WENT INTO CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS IMMEDIATELY STARTED. HOWEVER, ATTEMPTS WERE UNSUCCESSFUL AND THE PATIENT EXPIRED. IT WAS FURTHER REPORTED THAT PHYSICIAN BELIEVE THAT THE CAUSE OF DEATH WAS CARDIAC ARREST AND NOT DUE TO THE DEVICE. Manufacturer Narrative: E1: INITIAL REPORTER ADDRESS: (B)(6). E1: INITIAL REPORTER PHONE: (B)(6). GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED. Manufacturer Narrative: E1: INITIAL REPORTER ADDRESS: (B)(6). E1: INITIAL REPORTER PHONE: (B)(6). GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22549783&pc=NAY | 2955842-2025-30846 | 2025-06-25 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-07-18 04:00 | NAY | SUREFORM | Insufficient Information | Cardiac Arrest; Myocardial Infarction | K173721 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED TRANSTHORACIC ESOPHAGECTOMY WITH CHEST ANASTOMOSIS PROCEDURE, THE SUREFORM 60 STAPLER INSTRUMENT WAS IN THE PROCESS OF FIRING WHEN COINCIDENTALLY, THE PATIENT REQUIRED IMMEDIATE RESUSCITATION DUE TO SYMPTOMS OF A HEART ATTACK. AT THAT TIME, THE STAPLING CYCLE HAD ALREADY BEEN INITIATED AND COULD NOT BE INTERRUPTED. ALTHOUGH THE ROBOTIC SYSTEM WAS UNDOCKED, BOTH THE STAPLER INSTRUMENT AND THE ENDOSCOPE REMAINED INSIDE THE PATIENT THROUGHOUT RESUSCITATION EFFORTS. AN INTUITIVE TECHNICAL SUPPORT ENGINEER (TSE) WAS CONTACTED DURING THE EVENT AND ADVISED THE SURGICAL TEAM TO OBTAIN VISUAL ACCESS TO THE STAPLER INSTRUMENT INSIDE THE ABDOMEN TO ASSESS WHETHER THE EMERGENCY RELEASE FUNCTION COULD BE SAFELY ACTIVATED. MULTIPLE UNSUCCESSFUL ATTEMPTS WERE MADE TO UNCLAMP AND RELEASE THE STAPLER INSTRUMENT FROM THE SMALL INTESTINE. THE TSE ADVISED THE CUSTOMER TO CONTINUE MOVING THE RELEASE KNOB TO ALLOW THE INSTRUMENT TO BE UNCLAMPED FROM THE TISSUE. RELEASE OF THE STAPLER INSTRUMENT WAS THEN SUCCESSFUL. IT WAS REPORTED THAT THE PATIENT ULTIMATELY EXPIRED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOUND THAT A FORCE FIRE FAILURE OCCURRED, RESULTING IN THE SUREFORM 60 STAPLER INSTRUMENT BEING FORCE EXPIRED. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, TIP-UP FENESTRATED GRASPER, FENESTRATED BIPOLAR FORCEPS, VESSEL SEALER EXTEND. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THESE INSTRUMENTS USED DURING THIS PROCEDURE. THE SUREFORM 60 STAPLER INSTRUMENT WAS RECEIVED FOR EVALUATION; THE SUREFORM 60 RELOAD WAS NOT RETURNED. THE INSTRUMENT WAS FOUND TO HAVE A BLOCKED RADIO FREQUENCY IDENTIFIER (RFID) CHIP DUE TO THE USE OF THE INSTRUMENT¿S MANUAL RELEASE KNOB. AFTER RESETTING THE RFID CHIP, THE INSTRUMENT WAS PLACED AND RECOGNIZED BY THE IN-HOUSE ROBOT, THE INSTRUMENT WAS DRIVEN AND PASSED THE INITIALIZATION, RECOGNITION, AND ENGAGEMENT TESTS MULTIPLE TIMES. THE INSTRUMENT MOVED INTUITIVELY WITH FULL RANGE OF MOTION IN ALL DIRECTIONS, AND OPERATED AS EXPECTED: THE JAWS OPENED AND CLOSED PROPERLY, AND THE INSTRUMENT CLAMPED, FIRED A NEW RELOAD, AND UNCLAMPED WITHOUT ISSUE. A VISUAL INSPECTION REVEALED NO SIGNS OF DAMAGE. THE INSTRUMENT WAS FULLY FUNCTIONAL, AND NO PRODUCT ISSUE WAS IDENTIFIED. REVIEW OF THE SUREFORM 60 STAPLER LOG SHOWS THAT THE INSTRUMENT WAS INSTALLED 8 TIMES AND FIRED 8 RELOADS. EACH INSTALL CLAMPED AND FIRED SUCCESSFULLY, WITH THE EXCEPTION OF THE LAST INSTALL, WHERE FIRING FAILED, STOPPING AT APPROXIMATELY 94% COMPLETION AFTER APPROXIMATELY 51 SECONDS. THE LOGS SHOW AN ERROR REPRESENTING THE FAILURE, AND NO UNCLAMPING WAS ATTEMPTED. APPROXIMATELY 10 SECONDS LATER, THE GRIP RELEASE TOOL WAS DETECTED, THOUGH THE E-STOP HAD NOT YET BEEN PRESSED, AND THE STAPLER WAS THEN FORCE EXPIRED BY THE SYSTEM AS EXPECTED. APPROXIMATELY 18 MINUTES LATER, THE INSTRUMENT WAS RE-INSTALLED, AND LESS THAN A MINUTE LATER, THE E-STOP BUTTON WAS PRESSED ON THE SURGEON SIDE CONSOLE. THERE WERE NO FURTHER ERRORS OR EVENTS RELATED TO THE STAPLER FOUND IN THE LOGS. PER THE CUSTOMER AND TSE REPORT, THE INSTRUMENT WAS THEN SUCCESSFULLY RELEASED WITH THE GRIP RELEASE TOOL. THE DA VINCI SYSTEM AND INSTRUMENT MANUALS PROVIDE THE FOLLOWING WARNINGS: **WARNING: THE SYSTEM NEEDS TO BE IN A FAULTED STATE IN ORDER TO PERFORM MANUAL STAPLER RELEASE TO PREVENT PATIENT INJURY. **WARNING: DO NOT PERFORM GRIP RELEASE ON A NON-FAULTED SYSTEM WITHOUT FIRST PRESSING THE EMERGENCY STOP BUTTON. FAILURE TO OBSERVE THIS WARNING MAY RESULT IN UNINTENDED INSTRUMENT MOTION OR DAMAGE TO THE GRIP RELEASE MECHANISM. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING AN INTRAOPERATIVE EVENT. THE STAPLER WAS IN THE PROCESS OF BEING USED ON THE ESOPHAGUS WHEN THE PATIENT REQUIRED EMERGENT RESUSCITATION. THE DETAILS OF WHAT PRECIPITATED THAT EVENT ARE NOT PROVIDED. IT IS UNKNOWN IF THE USE OF THE STAPLER AT THAT MOMENT CONTRIBUTED TO THE EVENT. THE SURGEON COULD NOT RELEASE THE STAPLER AT THAT TIME. HOW THAT MAY HAVE AFFECTED EITHER THE SITUATION OR RESUSCITATIVE EFFORTS IS NOT PROVIDED. THE PATIENT ULTIMATELY EXPIRED. HOW THEY DIED AND THE REMAINDER OF THE EVENTS THAT LED TO THE DEATH ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SECTION B2: DATE OF DEATH IS ESTIMATED TO BE THE SAME DATE AS THE PROCEDURE, BUT WAS NOT CONFIRMED. . Manufacturer Narrative: SECTION B5: ADDITIONAL INFORMATION RECEIVED FROM THE SURGEON CLARIFIED THAT THE EVENT OCCURRED DURING A DA VINCI-ASSISTED SLEEVE GASTRECTOMY. THE PATIENT SUFFERED AN INTRAOPERATIVE MYOCARDIAL INFARCTION DURING THE STAPLING CYCLE AND REQUIRED MECHANICAL AND ELECTRICAL RESUSCITATION. AS A RESULT, THE DA VINCI UNIVERSAL SURGICAL MANIPULATORS COULD NOT BE DISCONNECTED, AND THE STAPLER REMAINED IN THE PATIENT, THOUGH IT DID NOT INTERFERE WITH RESUSCITATION EFFORTS. THE SURGEON REPORTED THERE WAS NO CONNECTION BETWEEN THE EVENT AND THE DA VINCI SYSTEM. THE AUTOPSY DETERMINED THE CAUSE OF DEATH TO BE AN INTRAOPERATIVE MYOCARDIAL INFARCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22549796&pc=LGW | 3006630150-2025-05681 | 2025-06-23 04:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-07-18 04:00 | LGW | INFINION? 16 | Adverse Event Without Identified Device or Use Problem | Muscle Weakness; Nausea; Vomiting; Diaphoresis | P030017 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS DOING WELL POST OPERATIVELY DURING THE TRIAL PHASE OF THE SPINAL CORD STIMULATOR (SCS) SYSTEM. ON DAY FOUR THE PATIENT WAS SCHEDULED FOR A MID TRIAL REPROGRAMMING. IT WAS LATER REPORTED THAT THE PATIENT EXPERIENCED NAUSEA AND VOMITING, STIMULATION WAS TURNED OFF AND THE PATIENT WAS TRANSPORTED TO THE HOSPITAL. IT WAS THEN STATED THAT THE PATIENT PASSED AWAY. IT IS UNKNOWN IF THE PATIENT DEATH IS DEVICE OR PROCEDURE, AND THERE IS NO INFORMATION AS TO WHAT MAY HAVE CAUSED IT DESPITE GOOD FAITH EFFORT REQUESTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22528571&pc=BZD | 2518422-2025-048165 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-07-17 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE; CANCER; DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN CONFIRMED NO EVIDENCE OF FOAM PARTICLES IN THE AIR PATH DURING THE DEVICE EVALUATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22528709&pc=MFA | 3007284006-2025-00142 | 2025-06-18 04:00 | Death | SPECTRANETICS CORPORATION | 2025-07-17 04:00 | MFA | GLIDELIGHT LASER SHEATH | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation; Cardiac Perforation; Pericardial Effusion | P960042 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) AND RIGHT ATRIAL (RA) LEAD DUE TO TRICUSPID VALVE REPAIR. SPECTRANETICS LLD LEAD LOCKING DEVICES (LLDS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. WHILE UTILIZING MULTIPLE SPECTRANETICS DEVICES (14F GLIDELIGHT LASER SHEATH, MEDIUM VISISHEATH DILATOR SHEATH, AND TIGHTRAIL ROTATING DILATOR SHEATH), AND MOVING BACK AND FORTH ON THE RV AND RA LEAD, THE RA LEAD WAS REMOVED. NEXT, USING THE GLIDELIGHT, PROGRESS WAS MADE TO THE RV APEX, AND WITH GENTLE TRACTION, THE LEAD CAME FREE. HOWEVER, A PERICARDIAL EFFUSION WAS DETECTED VIA TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE), AND RESCUE EFFORTS BEGAN, INCLUDING RESCUE BALLOON, CPR, AND STERNOTOMY. A PERFORATION IN THE RV AND SMALL PERFORATION OF THE SUPERIOR VENA CAVA (SVC) WERE DISCOVERED AND REPAIRED. THE PATIENT STABILIZED, THUS, THE DECISION WAS MADE TO CONTINUE WITH THE TRICUSPID VALVE REPAIR. UNFORTUNATELY, AFTER THE REPAIR, THE PATIENT DECLINED AND WAS PUT ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO), BUT THE PATIENT DEVELOPED PULSELESS LEGS, SEPSIS, ISCHEMIC LEGS, AND RENAL FUNCTION FAILURE. THE PATIENT DID NOT SURVIVE AND PASSED LATER THAT DAY. THIS REPORT CAPTURES THE 14F GLIDELIGHT IN USE WHEN THE PERFORATIONS OCCURRED, REQUIRING INTERVENTION, BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: A2) PATIENT DATE OF BIRTH - NOT AVAILABLE FROM FACILITY. A4) PATIENT WEIGHT - NOT AVAILABLE FROM FACILITY. A5/A6) PATIENT ETHNICITY AND RACE - NOT AVAILABLE FROM FACILITY. B6/B7) PATIENT INFORMATION REGARDING RELEVANT TESTS/LABORATORY DATA OR MEDICAL HISTORY - NOT AVAILABLE FROM FACILITY. D4/H4) THE LOT NUMBER WAS NOT PROVIDED, THUS THE FOLLOWING INFORMATION IS UNKNOWN: DEVICE SERIAL NUMBER, UNIQUE ID, EXPIRATION DATE, AND MANUFACTURE DATE. H3) THE DEVICE WAS DISCARDED, THUS NO INVESTIGATION COULD BE COMPLETED. H6) GREAT VESSEL, CARDIAC PERFORATION, AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE GLIDELIGHT DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22528788&pc=DWF | 9617601-2025-00645 | 2024-09-08 04:00 | Death | MEDTRONIC MEXICO | 2025-07-17 04:00 | DWF | BIO-MEDICUS NG ARTERIAL CANNULA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K142673 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING "COMPARISON OF ECMO, IABP AND ECMO + IABP IN THE POSTOPERATIVE PERIOD IN PATIENTS WITH POSTCARDIOTOMY SHOCK." THE OBJECTIVE WAS TO ASSESS OUTCOMES AND COMPLICATIONS OF PATIENTS WHO RECEIVED VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA-ECMO) AND INTRA-AORTIC BALLOON PUMP (IABP) SUPPORT AFTER CARDIAC SURGERY. THE TIME FRAME OF THIS STUDY WAS BETWEEN 2000 AND 2023. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: ECMO CANNULAE (BIO-MEDICUS FEMORAL). DEATHS OCCURRED IN THE STUDY POPULATION. BASED ON THE AVAILABLE INFORMATION, THESE DEATHS MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. SPECIFIC CAUSES OF DEATH ARE NOT PROVIDED. AMONG PATIENT ADVERSE EVENTS INCLUDED: STROKE SEPSIS REOPERATION DIALYSIS LIMB ISCHEMIA LIVER FAILURE NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: 96570-115 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0183-CANNULA; IMPLANT DATE N/A; EXPLANT DATE N/A G2: CITATION: AUTHORS: CAGDAS BARAN, EVREN OZCINAR, AHMET KAYAN. TITLE OF THE ARTICLE: COMPARISON OF ECMO, IABP AND ECMO + IABP IN THE POSTOPERATIVE PERIOD IN PATIENTS WITH POSTCARDIOTOMY SHOCK. JOURNAL NAME: JOURNAL OF CARDIOVASCULAR DEVELOPMENT AND DISEASE ISSUE: 11, 283 YEAR: 2024 LITERATURE REFERENCE: DOI.ORG/10.3390/JCDD11090283. THE EARLIEST DATE OF PUBLICATION HAS BEEN USED FOR B3 (DATE OF EVENT). NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22529218&pc=DTB | 2649622-2025-19359 | 2025-06-27 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-17 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Chest Pain; Dyspnea; Pleural Effusion; Tachycardia; Cardiac Tamponade; Discomfort; Cardiac Perforation; Convulsion/Seizure | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC PERFORATION, CARDIAC TAMPONADE, DISCOMFORT, CHEST PAIN, AND BREATHING DIFFICULTY APPROXIMATELY TEN DAYS POST IMPLANT OF A RIGHT VENTRICULAR (RV) LEAD. A PERICARDIOCENTESIS DRAIN WAS INSERTED INTO THE HEART. ECHOCARDIOGRAM SHOWED A SMALL PLEURAL EFFUSION. A SMALL AMOUNT OF BLOOD WAS DRAINED AND THE DECISION WAS MADE TO REMOVE THE RV LEAD. AFTER THE LEAD WAS REMOVED AN ECHO SHOWED THE EFFUSION GETTING BIGGER. A PERICARDIAL WINDOW WAS PERFORMED TO ACCESS THE HEART TO REPAIR THE HOLE IN THE HEART. THE PATIENT CONDITION BECAME WORSE AND THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA (VT). AFTER THE PATIENT WAS SHOCKED, THE PATIENT CONVERTED FROM ATRIAL FIBRILLATION (AF) INTO A SINUS RHYTHM. THE DECISION WAS MADE TO PERFORM A THORACOTOMY FOR BETTER ACCESS TO THE HEART AND POSSIBLE CARDIAC MASSAGE. WHILE THE HEART WAS EXPOSED AN EPICARDIAL LEAD WAS ATTACHED, THE PATIENT'S CONDITION STABILIZED. THE CHEST WAS CLOSED AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). A POST OP CHECK, SHOWED THE DEVICE FUNCTION WAS APPROPRIATE AT THAT TIME. THE PATIENT WAS EXHIBITING SIGNS OF SEIZURES, AN ELECTROENCEPHALOGRAM (EEG) STUDY WAS ORDERED. THE FAMILY CHOSE COMFORT CARE AFTER THE EEG WAS PERFORMED. IT WAS REPORTED THAT THE PATIENT DIED. CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: CONTINUATION OF D10: 5071-35 LEAD, IMPLANTED (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22529347&pc=DSQ | 2916596-2025-04750 | 2025-06-26 04:00 | Death | THORATEC CORPORATION | 2025-07-17 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED DUE TO DATA PRIVACY LAW. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22529699&pc=OZD | 1220648-2025-30415 | 2025-06-30 04:00 | Death | ABIOMED, INC. | 2025-07-17 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Cardiac Perforation | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND EXPERIENCED LOW FLOWS AND A PERFORATED LEFT VENTRICLE. IMPLANTATION WAS COMPLICATED AND REQUIRED REWIRING THE CATHETER TO AVOID THE PAPILLARY MUSCLE. THE PERFORATION WAS REPAIRED SURGICALLY, BUT THE PATIENT EXPIRED OVERNIGHT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22529786&pc=OZD | 1220648-2025-30416 | 2025-06-30 04:00 | Death | ABIOMED, INC. | 2025-07-17 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Malposition of Device | Hemorrhage/Bleeding; Tachycardia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND EXPERIENCED A GASTROINTESTINAL BLEED. THE PATIENT WAS TRANSFUSED A BOLUS AND TWO UNITS OF PACKED RED BLOOD CELLS. ANTICOAGULATION WAS WITHHELD. THE PATIENT ALSO EXPERIENCED VENTRICULAR TACHYCARDIA SO THE P LEVEL WAS REDUCED. THERE WERE PERSISTENT SUCTION ALARMS AND THE PUMP WAS FOUND TO BE POSITIONED UNDER THE PAPILLARY MUSCLE. IT IS UNKNOWN IF THE POSITIONING ISSUE OR SUCTION ISSUE WAS RESOLVED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL, POSITIONING ISSUES, VF/VT/AF/AT, AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. VASCULAR DAMAGE - PERIPHERAL: THE CAUSE OF THE GI BLEED WAS NOT CONFIRMED, HOWEVER. PRODUCT WAS NOT RETURNED FOR INVESTIGATION. SINCE INSUFFICIENT CLINICAL DETAILS WERE PROVIDED, THE ROOT CAUSE OF THE VASCULAR DAMAGE COULD NOT BE DETERMINED. LOW PUMP FLOW: DATA LOGS WERE REVIEWED AND THE SUDDEN SUCTION ALARMS WERE CONFIRMED. AT 01:30 ON (B)(6) 2025, THERE WAS A SUDDEN DROP IN PLACEMENT SIGNAL (PS) VALUES BUT INCREASED PULSATILITY THAT ALIGNED WITH A VERY SUDDEN DROP IN MEAN MOTOR CURRENT (MC) BUT INCREASED MODULATION. THOUGH THE PS VALUES DROPPED, THE PRESSURE TRACE STILL HAD AN AORTIC CHARACTERISTIC, AND MC MODULATION INCREASED, SO IT DOESN'T SEEM LIKELY THAT THE PUMP DROPPED INTO THE VENTRICLE. THE PULSATILE SUCTION ALGORITHM, OR CONTINUOUS SUCTION TRIGGERED SHORTLY AFTER. THERE WAS ALSO A CONCENTRATION OF POSITION IN AORTA ALARMS LATER THAT DAY, HOWEVER, THEY APPEAR TO BE UNRELATED TO THE INITIAL ONSET OF THE SUCTION ALARMS. PUMP FLOWS REMAINED LOW FOR THE REMAINDER OF THE CASE AND THE PS CONTINUED TO TREND DOWN. BASED ON THE CLINICAL DETAILS PROVIDED, PUMP POSITION SEEMS TO BE THE MOST LIKELY CAUSE OF THE SUCTION ALARMS. HOWEVER, IT IS UNCLEAR WHETHER OR NOT THE VERY SUDDEN DROP IN PS AND MC AT THE TIME SUCTION BEGAN TO TRIGGER CORRELATES WITH ANY PUMP POSITIONING ISSUES. ADDITIONALLY, THE PS ON DATA LOGS TRENDED DOWN FROM THAT POINT FORWARD, ALIGNING WITH THE PATIENT HAVING A GI BLEED AND THEN EXPIRING ON SUPPORT THAT DAY. SINCE THERE ARE INSUFFICIENT CLINICAL DETAILS, PRODUCT WASN'T RETURNED FOR INVESTIGATION, AND DATA LOGS WERE INCONCLUSIVE, THE ROOT CAUSE OF THE LOW PUMP FLOWS COULD NOT BE DETERMINED. POSITIONING ISSUES: DATA LOGS WERE REVIEWED AND THERE WERE 3 POSITION IN AORTA ALARMS ON (B)(6) 2025. BASED ON THE SUDDEN CHANGE IN MC PULSATILITY, IT APPEARS THIS SHIFT OF THE PUMP WAS INTENTIONAL (REPOSITIONING) BUT THERE WERE NO REPOSITIONING EVENTS REPORTED. THERE WERE ALSO 2 POSITION IN VENTRICLE ALARMS AT THE END OF THE CASE, SO THEY WERE LIKELY DURING PUMP EXPLANT. PRODUCT WASN'T RETURNED FOR INVESTIGATION. SINCE INSUFFICIENT CLINICAL DETAILS WERE PROVIDED REGARDING HOW THE PUMP CAME OUT OF POSITION AND PRODUCT WASN'T RETURNED, THE ROOT CAUSE OF THE POSITIONING ISSUE COULD NOT BE DETERMINED. VF/VT: THE ROOT CAUSE OF THE VT/VF/AT/AF WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22529861&pc=NIQ | 2024168-2025-07342 | 2025-01-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-07-17 04:00 | NIQ | XIENCE SKYPOINT¿ | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P110019 |  | 1 | Event Description: THE RESET TRIAL, A RANDOMIZED CONTROLLED STUDY COMPARING THE SECOND-GENERATION XIENCE¿ STENT WITH THE FIRST-GENERATION CYPHER¿ STENT NOTED 10 YEAR FOLLOW UP RESULTS SHOWING TARGET LESION FAILURE (TLF)¿A COMPOSITE OF CARDIAC DEATH, TARGET VESSEL-RELATED MYOCARDIAL INFARCTION (MI), AND CLINICALLY-DRIVEN TARGET LESION REVASCULARIZATION (TLR)¿WAS SIGNIFICANTLY LOWER IN THE XIENCE¿ GROUP. ADDITIONAL ADVERSE EVENTS NOTED WERE THROMBOTIC/EMBOLIC EVENTS, AND BLEEDING EVENTS. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE SKYPOINT EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. LITERATURE ATTACHMENT: ARTICLE TITLE: "UNKNOWN ARTICLE TITLE" THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER SEPARATE MEDWATCH REPORT. B3: DATE OF DEATH WILL BE ESTIMATED AS 1/01/2025, B3: PROCEDURE WILL BE ESTIMATED AS 1/01/2025 , D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A: DATE OF IMPLANT WILL BE ESTIMATED AS (B)(6) 2015.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530114&pc=OZD | 1220648-2025-30420 | 2025-07-01 04:00 | Death | ABIOMED, INC. | 2025-07-17 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT, THE IMPELLA DEVICE WAS UNABLE TO GENERATE FLOWS ABOVE P4 WITHOUT TRIGGERING SUCTION ALARMS. THE PATIENT EXPIRED WHILE ON MECHANICAL CIRCULATORY SUPPORT. Manufacturer Narrative: B.2 UPDATED OUTCOMES ATTRIBUTED TO ADVERSE EVENT TO NOT REFLECT OTHER SERIOUS (IMPORTANT MEDICAL EVENTS) AS DEATH IS THE ADVERSE EVENT SINCE MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30420 WAS SUBMITTED. H.6 UPDATED HEALTH EFFECT - CLINICAL CODE TO 4580 SINCE MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30420 WAS SUBMITTED. H.10 ADDITIONAL MANUFACTURER NARRATIVE INSTRUCTIONS FOR USE (IFUS) WERE REPORTED ON THE PREVIOUSLY SUBMITTED MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30420 INCORRECTLY. NO IFUS ARE NEEDED TO BE REPORTED FOR THIS REPORT IN ACCORDANCE WITH UPDATED PROCEDURES. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530333&pc=NVZ | 2017865-2025-93315 | 2025-04-27 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-17 04:00 | NVZ | ASSURITY RF DR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P880086 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530334&pc=NVY | 2017865-2025-93316 | 2025-04-27 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-17 04:00 | NVY | ISOFLEX OPTIM LEAD | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960030 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. VISUAL EXAMINATION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530342&pc=NVN | 2017865-2025-93317 | 2025-04-27 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-17 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL ANALYSIS DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530348&pc=PHX | 0001825034-2025-02155 | 2025-06-17 04:00 | Death | ZIMMER BIOMET, INC. | 2025-07-17 04:00 | PHX | COMP AUG MINI BSPLT W TPR SM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K172502 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN RECOVERY FOLLOWING A REVERSE TOTAL SHOULDER REPLACEMENT WITH THE COMPREHENSIVE SYSTEM. AS FAR AS THE SURGEON WAS CONCERNED, THE OPERATION WAS COMPLETED AS EXPECTED. IT WAS REPORTED THAT NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). G2: FOREIGN - THE EVENT OCCURRED IN AUSTRALIA. D10: CONCOMITANT MEDICAL PRODUCTS, PART (LOT): 115310(J7836529), 115395(66792976), 180552(66890490), 180553(67271033), 180550(67338919), 180550(67338919), 113613(66840669). ASSOCIATED PRODUCT INFORMATION: 110031424(67233121), 110031399(67351423). MULTIPLE MDR REPORTS WERE FILED FOR THIS EVENT. ASSOCIATED REPORTS ARE AVAILABLE FOR THE APPLICABLE CONCOMITANT PRODUCTS BELOW IN THE D10 NARRATIVE. THE CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION DUE TO HOSPITAL POLICY. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. IF ANY FURTHER INFORMATION IS FOUND THAT WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530354&pc=PHX | 0001825034-2025-02156 | 2025-06-17 04:00 | Death | ZIMMER BIOMET, INC. | 2025-07-17 04:00 | PHX | COMP RVRS SHLDR GLNSP STD 36MM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K193373 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN RECOVERY FOLLOWING A REVERSE TOTAL SHOULDER REPLACEMENT WITH THE COMPREHENSIVE SYSTEM. AS FAR AS THE SURGEON WAS CONCERNED, THE OPERATION WAS COMPLETED AS EXPECTED. IT WAS REPORTED THAT NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). G2: FOREIGN - THE EVENT OCCURRED IN AUSTRALIA. D10: CONCOMITANT MEDICAL PRODUCTS, PART (LOT): 110032410(67185082), 115395(66792976), 180552(66890490), 180553(67271033), 180550(67338919), 180550(67338919), 113613(66840669). ASSOCIATED PRODUCT INFORMATION: 110031424(67233121), 110031399(67351423). PROPOSED ANNEX G: MECHANICAL (G04) - HEAD. MULTIPLE MDR REPORTS WERE FILED FOR THIS EVENT. ASSOCIATED REPORTS ARE AVAILABLE FOR THE APPLICABLE CONCOMITANT PRODUCTS BELOW IN THE D10 NARRATIVE. THE CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION DUE TO HOSPITAL POLICY. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. IF ANY FURTHER INFORMATION IS FOUND THAT WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530359&pc=PHX | 0001825034-2025-02157 | 2025-06-17 04:00 | Death | ZIMMER BIOMET, INC. | 2025-07-17 04:00 | PHX | COMP RVS CNTRL 6.5X25MM ST/RST | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K193373 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN RECOVERY FOLLOWING A REVERSE TOTAL SHOULDER REPLACEMENT WITH THE COMPREHENSIVE SYSTEM. AS FAR AS THE SURGEON WAS CONCERNED, THE OPERATION WAS COMPLETED AS EXPECTED. IT WAS REPORTED THAT NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). G2: FOREIGN - THE EVENT OCCURRED IN AUSTRALIA. D10: CONCOMITANT MEDICAL PRODUCTS, PART (LOT): 115310(J7836529), 110032410(67185082), 180552(66890490), 180553(67271033), 180550(67338919), 180550(67338919), 113613(66840669). ASSOCIATED PRODUCT INFORMATION: 110031424(67233121), 110031399(67351423). MULTIPLE MDR REPORTS WERE FILED FOR THIS EVENT. ASSOCIATED REPORTS ARE AVAILABLE FOR THE APPLICABLE CONCOMITANT PRODUCTS BELOW IN THE D10 NARRATIVE. THE CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION DUE TO HOSPITAL POLICY. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. IF ANY FURTHER INFORMATION IS FOUND THAT WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530363&pc=PHX | 0001825034-2025-02158 | 2025-06-17 04:00 | Death | ZIMMER BIOMET, INC. | 2025-07-17 04:00 | PHX | COMP LK SCR 3.5HEX 4.75X25 ST | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K132239 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN RECOVERY FOLLOWING A REVERSE TOTAL SHOULDER REPLACEMENT WITH THE COMPREHENSIVE SYSTEM. AS FAR AS THE SURGEON WAS CONCERNED, THE OPERATION WAS COMPLETED AS EXPECTED. IT WAS REPORTED THAT NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). G2: FOREIGN - THE EVENT OCCURRED IN AUSTRALIA. D10: CONCOMITANT MEDICAL PRODUCTS, PART (LOT): 115310(J7836529), 110032410(67185082), 115395(66792976), 180553(67271033), 180550(67338919), 180550(67338919), 113613(66840669). ASSOCIATED PRODUCT INFORMATION: 110031424(67233121), 110031399(67351423). MULTIPLE MDR REPORTS WERE FILED FOR THIS EVENT. ASSOCIATED REPORTS ARE AVAILABLE FOR THE APPLICABLE CONCOMITANT PRODUCTS BELOW IN THE D10 NARRATIVE. THE CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION DUE TO HOSPITAL POLICY. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. IF ANY FURTHER INFORMATION IS FOUND THAT WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530370&pc=PHX | 0001825034-2025-02159 | 2025-06-17 04:00 | Death | ZIMMER BIOMET, INC. | 2025-07-17 04:00 | PHX | COMP LK SCR 3.5HEX 4.75X30 ST | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K132239 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN RECOVERY FOLLOWING A REVERSE TOTAL SHOULDER REPLACEMENT WITH THE COMPREHENSIVE SYSTEM. AS FAR AS THE SURGEON WAS CONCERNED, THE OPERATION WAS COMPLETED AS EXPECTED. IT WAS REPORTED THAT NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). G2: FOREIGN - THE EVENT OCCURRED IN AUSTRALIA. D10: CONCOMITANT MEDICAL PRODUCTS, PART (LOT): 115310(J7836529), 110032410(67185082), 115395(66792976), 180552(66890490), 180550(67338919), 180550(67338919), 113613(66840669). ASSOCIATED PRODUCT INFORMATION: 110031424(67233121), 110031399(67351423). MULTIPLE MDR REPORTS WERE FILED FOR THIS EVENT. ASSOCIATED REPORTS ARE AVAILABLE FOR THE APPLICABLE CONCOMITANT PRODUCTS BELOW IN THE D10 NARRATIVE. THE CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION DUE TO HOSPITAL POLICY. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. IF ANY FURTHER INFORMATION IS FOUND THAT WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530377&pc=PHX | 0001825034-2025-02160 | 2025-06-17 04:00 | Death | ZIMMER BIOMET, INC. | 2025-07-17 04:00 | PHX | COMP LK SCR 3.5HEX 4.75X15 ST | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K132239 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN RECOVERY FOLLOWING A REVERSE TOTAL SHOULDER REPLACEMENT WITH THE COMPREHENSIVE SYSTEM. AS FAR AS THE SURGEON WAS CONCERNED, THE OPERATION WAS COMPLETED AS EXPECTED. IT WAS REPORTED THAT NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). G2: FOREIGN - THE EVENT OCCURRED IN AUSTRALIA. D10: CONCOMITANT MEDICAL PRODUCTS, PART (LOT): 115310(J7836529), 110032410(67185082), 115395(66792976), 180552(66890490), 180553(67271033), 180550(67338919), 113613(66840669). ASSOCIATED PRODUCT INFORMATION: 110031424(67233121), 110031399(67351423). MULTIPLE MDR REPORTS WERE FILED FOR THIS EVENT. ASSOCIATED REPORTS ARE AVAILABLE FOR THE APPLICABLE CONCOMITANT PRODUCTS BELOW IN THE D10 NARRATIVE. THE CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION DUE TO HOSPITAL POLICY. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. IF ANY FURTHER INFORMATION IS FOUND THAT WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530379&pc=PHX | 0001825034-2025-02161 | 2025-06-17 04:00 | Death | ZIMMER BIOMET, INC. | 2025-07-17 04:00 | PHX | COMP LK SCR 3.5HEX 4.75X15 ST | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K132239 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN RECOVERY FOLLOWING A REVERSE TOTAL SHOULDER REPLACEMENT WITH THE COMPREHENSIVE SYSTEM. AS FAR AS THE SURGEON WAS CONCERNED, THE OPERATION WAS COMPLETED AS EXPECTED. IT WAS REPORTED THAT NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). G2: FOREIGN - THE EVENT OCCURRED IN AUSTRALIA. D10: CONCOMITANT MEDICAL PRODUCTS, PART (LOT): 115310(J7836529), 110032410(67185082), 115395(66792976), 180552(66890490), 180553(67271033), 180550(67338919), 113613(66840669). ASSOCIATED PRODUCT INFORMATION: 110031424(67233121), 110031399(67351423). MULTIPLE MDR REPORTS WERE FILED FOR THIS EVENT. ASSOCIATED REPORTS ARE AVAILABLE FOR THE APPLICABLE CONCOMITANT PRODUCTS BELOW IN THE D10 NARRATIVE. THE CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION DUE TO HOSPITAL POLICY. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. IF ANY FURTHER INFORMATION IS FOUND THAT WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530391&pc=MBF | 0001825034-2025-02162 | 2025-06-17 04:00 | Death | ZIMMER BIOMET, INC. | 2025-07-17 04:00 | MBF | COMP PRIMARY STEM 13MM MICRO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K193038 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN RECOVERY FOLLOWING A REVERSE TOTAL SHOULDER REPLACEMENT WITH THE COMPREHENSIVE SYSTEM. AS FAR AS THE SURGEON WAS CONCERNED, THE OPERATION WAS COMPLETED AS EXPECTED. IT WAS REPORTED THAT NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). G2: FOREIGN - THE EVENT OCCURRED IN AUSTRALIA. D10: CONCOMITANT MEDICAL PRODUCTS, PART (LOT): 115310(J7836529), 110032410(67185082), 115395(66792976), 180552(66890490), 180553(67271033), 180550(67338919), 180550(67338919). ASSOCIATED PRODUCT INFORMATION: 110031424(67233121), 110031399(67351423). MULTIPLE MDR REPORTS WERE FILED FOR THIS EVENT. ASSOCIATED REPORTS ARE AVAILABLE FOR THE APPLICABLE CONCOMITANT PRODUCTS BELOW IN THE D10 NARRATIVE. THE CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION DUE TO HOSPITAL POLICY. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. IF ANY FURTHER INFORMATION IS FOUND THAT WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530447&pc=DSQ | 22530447 | 2025-06-20 04:00 | Death | THORATEC CORPORATION | 2025-07-17 04:00 | DSQ | HEARTMATE 3 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Intracranial Hemorrhage; Respiratory Arrest |  |  | 1 | Event Description: FDA [THE FOOD AND DRUG ADMINISTRATION] REPORTABLE EVENT DUE TO SAH [SUBARACHNOID HEMORRHAGE] IN LVAD [LEFT VENTRICULAR ASSIST DEVICE] PATIENT. PATIENT IS WELL KNOWN TO OUR TEAM S/P [STATUS POST] VAD [VENTRICULAR ASSIST DEVICE] FOR ADVANCED NICM C/B RV [NON-ISCHEMIC CARDIOMYOPATHY CAUSED BY RIGHT VENTRICULAR] FAILURE (REQUIRING PROLONGED INOTROPE SUPPORT), RESPIRATORY FAILURE REQUIRING TRACH (NOW DECANNULATED), AKI [ACUTE KIDNEY INJURY] REQUIRING CRRT [CONTINUOUS RENAL REPLACEMENT THERAPY] THEN OFF OF HD [HEMODIALYSIS] FOR SEVERAL MONTHS. PATIENT PRESENTED WITH DYSPHAGIA AND ASPIRATION PNA [PNEUMONIA]. PATIENT HAD A CHANGE IN MENTAL STATUS AND WAS HYPERCARBIC REQUIRING INTUBATION. WORSENING RENAL FAILURE FELT [REDACTED] RV [RIGHT VENTRICLE] DYSFUNCTION/CARDIORENAL REQUIRING IV DIURETICS (HIGH DOSE) AND DOBUTAMINE. CT HEAD REVEALED A SAH [SUBARACHNOID HEMORRHAGE]. NEUROSURGERY WAS INVOLVED, ANTICOAGULATION WAS HELD, AND PATIENT RECEIVED VITAMIN K. SERIAL HEAD CTS WERE STABLE, AND INR [INTERNATIONAL NORMALIZED RATIO] TRENDED DOWN TO 1.5. UNFORTUNATELY, PATIENT HAD AN ACUTE CHANGE IN NEURO STATUS EARLY [REDACTED] AM AND WAS FOUND TO HAVE A SIGNIFICANT INTRAPARENCHYMAL HEMORRHAGE THAT WAS NEUROLOGICALLY DEVASTATING. THE FAMILY ELECTED TO NOT ESCALATE CARE, AND PATIENT PASSED AWAY PEACEFULLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530772&pc=DQY | 22530772 | 2025-06-11 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-17 04:00 | DQY | WATCHMAN FXD CURVE¿ ACCESS SYSTEM | Gas/Air Leak | Air Embolism; Bradycardia; Cardiac Arrest; Respiratory Arrest |  |  | 1 | Event Description: DURING THE LAAC [LEFT ATRIAL APPENDAGE CLOSURE] PROCEDURE, AFTER PLACEMENT OF THE WATCHMAN DEVICE AND REMOVAL OF THE DILATOR FROM THE SHEATH, AUDIBLE SUCTION OF AIR INTO THE SHEATH WAS NOTED. SIGNIFICANT AIR WAS NOTED IN THE LA [LEFT ATRIUM]. FLUID WAS PULLED BACK INTO THE SHEATH AND THE SHEATH WAS RETRACTED INTO THE RA [RIGHT ATRIUM]. PATIENT DEVELOPED BRADYCARDIA. A PACING CATH [CATHETER] WAS PLACED INTO THE RV [RIGHT VENTRICLE]. PATIENT DETERIORATED AND NO PULSE WAS FOUND. CPR [CARDIOPULMONARY RESUSCITATION] AND ACLS [ADVANCED CARDIAC LIFE SUPPORT] WERE INITIATED. INTERVENTIONAL CARD RESPONDED AND PERFORMED A LHC [LEFT HEART CATHETERIZATION] WHICH REVEALED SIGNIFICANT AIR IN THE RCA [RIGHT CORONARY ARTERY] AND L [LEFT] CORONARY SYSTEM. AIR WAS ASPIRATED AS MUCH AS POSSIBLE. AFTER 45 MIN AND SEVERAL ROUNDS OF ACLS, PATIENT EXPIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530785&pc=MVK | 3008642652-2025-06932 | 2025-06-06 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-17 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. PER CLINICAL REVIEW OF THE CONTINUOUS ECG RECORDING, THE DEVICE WAS STARTED UP AT 16:16:11 ON (B)(6) /2025. AN ARRHYTHMIA WAS DETECTED AT 08:46:01 ON (B)(6) 2025. ECG SHOWS SINUS RHYTHM @ 60 BPM DEGRADING TO ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. THE DETECTION STOPPED AT 08:46:14. ASYSTOLE WAS DETECTED TWO TIMES FROM 10:07:46 TO 10:08:46. ECG SHOWS SINUS BRADYCARDIA @ 20 BPM TRANSITIONING TO AN IDIOVENTRICULAR RHYTHM @ 60 BPM. THE RHYTHM THEN DEGRADES TO FINE VF WITH VARYING AMPLITUDES. AN ARRHYTHMIA WAS DETECTED AT 10:09:11. ECG SHOWS FINE VF WITH VARYING AMPLITUDES. THE DETECTION STOPPED AT 10:09:15. ASYSTOLE WAS DETECTED AT 11:05:10. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT AND INTERMITTENT CARDIAC ACTIVITY. THE RHYTHM THEN DEGRADES TO ASYSTOLE. PER ECG DATA, THE PATIENT WAS IN FINE VF WITH VARYING AMPLITUDES. THE RHYTHM THEN SLOWS TO AN IDIOVENTRICULAR RHYTHM FROM 60-40 BPM. THE RHYTHM THEN DEGRADES TO ASYSTOLE WITH CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF FROM APPROX. 10:09:11 UNTIL THE ELECTRODE BELT DISCONNECTED AT 11:12:46 ON (B)(6) 2025. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE DEVICE PROPERLY DETECTED VF. VARYING AMPLITUDE PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE DEVICE PROPERLY DETECTED ASYSTOLE. WHICH IS A NON-SHOCKABLE RHYTHM. FINE VF WAS SEEN DURING ASYSTOLE. THE VF FREQUENCY WAS LESS THAN THE RATE THRESHOLD PREVENTING THE LIFEVEST FROM TREATING THE PATIENT. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530820&pc=PNJ | 9612164-2025-03503 | 2025-06-25 04:00 | Death | MEDTRONIC IRELAND | 2025-07-17 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING IMPLANT THE MICRA INTRODUCER WAS INSERTED AND THE DELIVERY SYSTEM WAS PLACED IN THE RIGHT VENTRICLE. THERE HAD BEEN TWO UNSUCCESSFUL DEPLOYMENTS OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) AND THE THIRD WAS BEING ATTEMPTED. DURING THIS PLACEMENT THERE WAS DIFFICULTY READING THE PATIENTS BLOOD PRESSURE. A PACED RHYTHM WAS OBSERVED ON THE SURFACE ELECTROCARDIOGRAM (ECG). FEMORAL PULSE WAS ABSENT. INITIALLY CHEST COMPRESSIONS WHERE STARTED MANUALLY, THEN A CHEST COMPRESSION DEVICE WAS USED. THE PATIENT CONTINUED TO SHOW PULSELESS ELECTRICAL ACTIVITY, ALTHOUGH FEMORAL PULSE WAS NOTED DURING CHEST COMPRESSIONS. ALL ATTEMPTS AT RESUSCITATION WERE INEFFECTIVE AND THE PATIENT WAS RULED DECEASED. THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), DELIVERY SYSTEM AND INTRODUCER REMAINED IN THE PATIENT AFTER THE PATIENT EXPIRED Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME MICRA PRODUCT ID MC2AVR1-DELSYS SERIAL: (B)(6). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530821&pc=DYB | 2953200-2025-00008 | 2025-06-25 04:00 | Death | MEDTRONIC VASCULAR | 2025-07-17 04:00 | DYB | MICRA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT DURING IMPLANT THE MICRA INTRODUCER WAS INSERTED AND THE DELIVERY SYSTEM WAS PLACED IN THE RIGHT VENTRICLE. THERE HAD BEEN TWO UNSUCCESSFUL DEPLOYMENTS OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) AND THE THIRD WAS BEING ATTEMPTED. DURING THIS PLACEMENT THERE WAS DIFFICULTY READING THE PATIENTS BLOOD PRESSURE. A PACED RHYTHM WAS OBSERVED ON THE SURFACE ELECTROCARDIOGRAM (ECG). FEMORAL PULSE WAS ABSENT. INITIALLY CHEST COMPRESSIONS WERE STARTED MANUALLY, THEN A CHEST COMPRESSION DEVICE WAS USED. THE PATIENT CONTINUED TO SHOW PULSELESS ELECTRICAL ACTIVITY, ALTHOUGH FEMORAL PULSE WAS NOTED DURING CHEST COMPRESSIONS. ALL ATTEMPTS AT RESUSCITATION WERE INEFFECTIVE AND THE PATIENT WAS RULED DECEASED. THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), DELIVERY SYSTEM AND INTRODUCER REMAINED IN THE PATIENT AFTER THE PATIENT EXPIRED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22530863&pc=PHX | 0001822565-2025-02578 | 2025-06-17 04:00 | Death | ZIMMER BIOMET, INC. | 2025-07-17 04:00 | PHX | MINI TRAY STD COCR +0 OFFSET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181611 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN RECOVERY FOLLOWING A REVERSE TOTAL SHOULDER REPLACEMENT WITH THE COMPREHENSIVE SYSTEM. AS FAR AS THE SURGEON WAS CONCERNED, THE OPERATION WAS COMPLETED AS EXPECTED. IT WAS REPORTED THAT NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). G2: FOREIGN - THE EVENT OCCURRED IN AUSTRALIA. D10: CONCOMITANT MEDICAL PRODUCTS, PART (LOT): 110031424(67233121). ASSOCIATED PRODUCT INFORMATION: 115310(J7836529), 110032410(67185082), 115395(66792976), 180552(66890490), 180553(67271033), 180550(67338919), 180550(67338919), 113613(66840669). MULTIPLE MDR REPORTS WERE FILED FOR THIS EVENT. ASSOCIATED REPORTS ARE AVAILABLE FOR THE APPLICABLE CONCOMITANT PRODUCTS BELOW IN THE D10 NARRATIVE. THE CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION DUE TO HOSPITAL POLICY. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. IF ANY FURTHER INFORMATION IS FOUND THAT WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22531163&pc=PHX | 0001822565-2025-02579 | 2025-06-17 04:00 | Death | ZIMMER BIOMET, INC. | 2025-07-17 04:00 | PHX | CR VIVACIT-E 36MM BRNG STD | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181611 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN RECOVERY FOLLOWING A REVERSE TOTAL SHOULDER REPLACEMENT WITH THE COMPREHENSIVE SYSTEM. AS FAR AS THE SURGEON WAS CONCERNED, THE OPERATION WAS COMPLETED AS EXPECTED. IT WAS REPORTED THAT NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). G2: FOREIGN - THE EVENT OCCURRED IN AUSTRALIA. D10: CONCOMITANT MEDICAL PRODUCTS, PART (LOT): 110031399(67351423). ASSOCIATED PRODUCT INFORMATION: 115310(J7836529), 110032410(67185082), 115395(66792976), 180552(66890490), 180553(67271033), 180550(67338919), 180550(67338919), 113613(66840669). MULTIPLE MDR REPORTS WERE FILED FOR THIS EVENT. ASSOCIATED REPORTS ARE AVAILABLE FOR THE APPLICABLE CONCOMITANT PRODUCTS BELOW IN THE D10 NARRATIVE. THE CUSTOMER HAS INDICATED THAT THE PRODUCT WILL NOT BE RETURNED TO ZIMMER BIOMET FOR INVESTIGATION DUE TO HOSPITAL POLICY. THE INVESTIGATION IS IN PROCESS. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. IF ANY FURTHER INFORMATION IS FOUND THAT WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22531383&pc=KNT | 3010757606-2025-00317 | 2025-06-01 04:00 | Death | ABBVIE - MEDICAL DEVICE CENTER | 2025-07-17 04:00 | KNT | DUODOPA_DUOPA | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K142793 |  | 1 | Event Description: THE REPORTED EVENT INVOLVES A PATIENT IN GERMANY WHO UNDERWENT A PROCEDURE FOR THE REPLACEMENT OF A PERCUTANEOUS ENDOSCOPIC GASTROSTOMY (PEG) TUBE WITH JEJUNAL (PEG-J) TUBE ON UNKNOWN DATE. ON AN UNKNOWN DATE, THE PATIENT WAS HOSPITALIZED BECAUSE THE PEG HAD BECOME DETACHED AND A TUBE REPLACEMENT WAS CARRIED OUT. ON (B)(6) 2025, THE PATIENT DIED. THE CONSUMER REPORTER DID NOT KNOW WHETHER THERE WAS A CONNECTION. THE TUBING REPLACEMENT DATE WAS NOT PROVIDED. AN AUTOPSY WAS NOT PERFORMED. DESPITE MULTIPLE QUERIES, THERE IS NO CLEAR EVIDENCE LINKING THE DEVICE TO THE REPORTED EVENT, AND NO ADDITIONAL INFORMATION REGARDING DEVICE-RELATED ALLEGATIONS, MALFUNCTIONS, COMPLICATIONS, DIAGNOSTIC TESTING, OR TREATMENT IS AVAILABLE. THEREFORE, OUT OF ABUNDANCE OF CAUTION, ABBVIE IS CONSERVATIVELY REPORTING THE EVENT WITHOUT ATTRIBUTING CAUSALITY TO THE DEVICE. Manufacturer Narrative: CATALOG NUMBER IN D4 IS THE INTERNATIONAL LIST NUMBER WHICH IS SIMILAR TO US LIST NUMBER OF 062912. THE DEVICE INVOLVED IN THE EVENT WAS NOT RETURNED; THEREFORE, A RETURN SAMPLE EVALUATION IS UNABLE TO BE PERFORMED. POTENTIAL POST PROCEDURAL COMPLICATION IS A KNOWN COMPLICATION OF A PEG- J TUBE PLACEMENT. H6 CODE OF E2402 WAS CHOSEN TO CAPTURE THE EVENT OF POSSIBLE POST PROCEDURAL COMPLICATION. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22531420&pc=MNS | 2518422-2025-107870 | 2025-07-15 04:00 | Death | RESPIRONICS, INC. | 2025-07-17 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE EXAMINATION HAS NOT BEGUN BECAUSE THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22531437&pc=DSQ | 2916596-2025-04782 | 2024-12-06 05:00 | Death | THORATEC CORPORATION | 2025-07-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22531569&pc=DSQ | 2916596-2025-04724 | 2025-06-30 04:00 | Death | THORATEC CORPORATION | 2025-07-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Cardiogenic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO SEPSIS SUPERIMPOSED ON CARDIOGENIC SHOCK. THE PATIENT'S FAMILY DECLINED AN AUTOPSY. THE TEAM REPORTED THAT THE EVENT WAS PATIENT RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE PATIENT'S SEPSIS COULD NOT CONCLUSIVELY BE DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING SEPSIS, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22531657&pc=DSQ | 2916596-2025-04671 | 2025-07-06 04:00 | Death | THORATEC CORPORATION | 2025-07-17 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Tachycardia | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT DURING POST INTENSIVE CARE UNIT (ICU) STAY, THE PATIENT PERFORMED MULTIPLE VENTRICULAR TACHYCARDIA (VT) EPISODES NOT WELL TOLERATED. THE PATIENT WAS SHOCKED SEVERAL TIMES BEFORE GOING BACK TO A SINUS RHYTHM. THE PATIENT CONTINUED TO HAVE THE VT ALONG WITH LOW FLOW ALARMS STORM DESPITE ANTI ARRHYTHMIC DRUG INTRODUCTION. THE PATIENT WAS READMITTED TO THE ICU, INTUBATED AND VENTILATED. VT ABLATION WAS PERFORMED BUT PATIENT HAD OTHER VT THAT WAS NOT WELL TOLERATED. ON (B)(6) 2025, THE PATIENT HAD A CARDIAC ARREST AND DID NOT RECOVER ULTIMATELY PASSING AWAY. FLOW THE WHOLE TIME WAS STABLE AROUND 4LPM. THE DEATH WAS NOT DEVICE OR THERAPY RELATED AND THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22531666&pc=DSQ | 2916596-2025-04769 | 2025-04-22 04:00 | Death | THORATEC CORPORATION | 2025-07-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS SAID TO BE UNKNOWN. Manufacturer Narrative: A4: PATIENT WEIGHT NOT PROVIDED. PENDING FOLLOW-UP WITH SITE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22532355&pc=NPT | 2025587-2025-05085 | 2025-06-21 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-17 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Non specific EKG/ECG Changes; Low Blood Pressure/ Hypotension | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE OF A TRANSCATHETER VALVE IN A PATIENT WITH A HORIZONTAL AORTA AND ENLARGED ASCENDING AORTA, THE VALVE WAS IMPLANTED IN A SINGLE ATTEMPT AT AN OPTIMAL POSITION AT AN IMPLANT DEPTH OF 3 MILLIMETERS (MM) ON THE NON-CORONARY CUSP (NCC) AND 4 MM ON THE LEFT CORONARY CUSP (LCC) WITHOUT RECAPTURE. FOLLOWING THE IMPLANT OF THE VALVE, A NARROW QRS WAS OBSERVED AND THERE WERE NO ISSUES WITH ATRIO-VENTRICULAR (AV) CONDUCTION. A POST-OPERATIVE ECHOCARDIOGRAPHY WAS PERFORMED THAT DID NOT REVEAL ANY ISSUES. THE FOLLOWING DAY, AN ECHOCARDIOGRAPHIC EXAMINATION WAS CONDUCTED AND NO ABNORMALITIES WERE IDENTIFIED. TWO DAYS FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT'S CONDITION RAPIDLY DETERIORATED. PERCUTANEOUS CARDIOPULMONARY SUPPORT (PCPS) WAS INSERTED; HOWEVER, THE PATIENT'S BLOOD PRESSURE COULD NOT BE STABILIZED. UPON FLUOROSCOPY, AN ASCENDING AORTIC DISSECTION WAS OBSERVED AND THE PATIENT LATER DIED. IT WAS SUSPECTED THAT FOLLOWING THE IMPLANT PROCEDURE, THE UPPER EDGE OF THE VALVE CAME INTO CONTACT WITH THE ENLARGED ASCENDING AORTIC WALL DUE TO PULSATION CAUSED BY RESPIRATION THAT MAY HAVE CAUSED THE DISSECTION. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT FX DCS; PRODUCT ID D-EVOLUTFX-2329 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22532413&pc=NPT | 2025587-2025-05088 | 2025-06-10 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-17 04:00 | NPT | EVOLUT FX DCS | Insufficient Information; Patient Device Interaction Problem | Hemorrhage/Bleeding | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT ON THE SAME DAY AS A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE, MASSIVE HEMORRHAGING AND HEMORRHAGING AT THE ACCESS SITE OCCURRED.Event Description: IT WAS REPORTED THAT ON THE SAME DAY AS A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE, MASSIVE HEMORRHAGING AND HEMORRHAGING AT THE ACCESS SITE OCCURRED. IT WAS REPORTED THAT APPROXIMATELY 1 MONTH AND 2 WEEKS LATER, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS SEPSIS DUE TO BEDSORES OR THE "CUTDOWN" PROCEDURE. PER THE PHYSICIAN, NEITHER THE DEVICE NOR THE PROCEDURE CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID {{EVFXPLUS}}; PRODUCT LOT/SERIAL NUMBER {{UNKNOWN}}; PRODUCT TYPE: {{0195 HEART VALVES}}; IMPLANT DATE {{B)(6) 2025}}; EXPLANT DATE {{NA}} H6. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2. B5. ADDED SECOND PARAGRAPH. H1. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22532415&pc=OZD | 1220648-2025-30444 | 2025-07-01 04:00 | Death | ABIOMED, INC. | 2025-07-17 04:00 | OZD | IMPELLA CP | Positioning Failure; Restricted Flow rate | Hemolysis; Tachycardia |  |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT WAS IMPLANTED WITH IMPELLA CP WITH ACUTE MYOCARDIAL INFARCTION AND CORONARY ARTERY DISEASE FOR SUPPORT DURING A HIGH RISK CARDIAC PROCEDURE. DURING THE PROCEDURE, THE PATIENT WAS NOTED TO HAVE MULTIPLE BLOCKAGES AND WITH THE PREVIOUS STENTS PLACED. THE PATIENT WAS ON MULTIPLE PRESSORS, INTUBATED, AND CONTINUED TO DECLINE. IT WAS THEN DECIDED TO PLACE THE IMPELLA CP WITH THE GOAL TO LEAVE IN PLACE TO SUPPORT THE PATIENT PRIOR TO PROCEEDING WITH ANY FURTHER INTERVENTIONS. SHORTLY AFTER THE REPOSITIONING SHEATH WAS ADVANCED, THE PATIENT WAS NOTED TO BE IN VENTRICULAR TACHYCARDIA AND REQUIRED TWO SHOCKS PER THE DEFIBRILLATOR. THE PATIENT WAS GIVEN BOLUS AND STARTED ON A DRIP AT THIS TIME. THE PHYSICIAN NOTED THE COLOR AND WARMTH TO THE RIGHT FOOT AND WAS UNABLE TO OBTAIN DOPPLER PULSE. HOWEVER, IT WAS STATED WITH CURRENT PRESSOR SUPPORT AND DECREASED PULSATILITY DUE TO AMIO BOLUS, THERE WERE SATISFIED WITH THE RIGHT FOOT BEING WARM. THE PATIENT HAD SUCTION EVENTS AND GROSS HEMOLYSIS OF LAB. ECHOCARDIOGRAM MEASURED THE IMPELLA AT 2.0 CENTIMETERS FROM ATRIOVENTRICULAR ANNULUS. THE PUMP WAS REPOSITIONED DUE TO THE LABS HEMOLYZING. AFTER REPOSITIONING, THE IMPELLA MEASURED 3.6 CENTIMETERS. UNFORTUNATELY, THE PATIENT CONTINUED TO DECOMPENSATE. CONTINUOUS RENAL REPLACEMENT THERAPY WAS ATTEMPTED. HOWEVER, IT ONLY RAN FOR 30 MINUTES DUE TO VOLUME ISSUES. FLUIDS WERE GIVEN AND BICARBONATE DRIP WAS STARTED. SUCTION CONTINUED AND P-LEVEL WAS DROPPED AS NEEDED. ULTIMATELY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION. ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22532438&pc=OZD | 1220648-2025-30445 | 2025-07-03 04:00 | Death | ABIOMED, INC. | 2025-07-17 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia; Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT (PT) PRESENTED AS AN OUT OF HOSPITAL CARDIAC ARREST. DOWNTIME 15-20 MINUTES. RETURN OF SPONTANEOUS CIRCULATION ACHIEVED BY EMERGENCY MEDICAL SERVICES. PT INTUBATED DUE TO BEING UNRESPONSIVE. ELECTROCARDIOGRAM SHOWED ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION. PT WAS TRANSFERRED FOR LEFT HEART CATHETERIZATION EVALUATION. IN THE CATH LAB, PT HAD HEPARIN DRIP AND DOPAMINE DRIP INFUSING. THE PT CONTINUED TO BE HYPOTENSIVE AFTER INTERVENTIONS. DECISION MADE TO PLACE IMPELLA CP. THE IMPELLA WAS PREPPED AND INSERTED PER IFU. PEEL-AWAY SHEATH REMOVED, REPO ADVANCED, AND SITE SECURED. DRESSING APPLIED. DURING SUPPORT IT WAS NOTED THAT THE PT HAD ABSENT DOPPLER PULSE TO DISTAL EXTREMITY AND WAS RAPIDLY DETERIORATING HAVING MULTIPLE RUNS OF VENTRICULAR TACHYCARDIA (VT). THE PHYSICIAN STATED LIKELY RELATED TO INFERIOR MYOCARDIAL INFARCTION. PT WENT PULSELESS VT AND ADVANCED CARDIAC LIFE SUPPORT MEASURES WERE INITIATED AND GIVEN FOR SEVERAL MINUTES. CODE ENDED AND MEDICAL DOCTOR PRONOUNCED TIME OF DEATH. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA CP WITH SMART ASSIST. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22532567&pc=BZD | 2518422-2025-107575 | 2025-05-19 04:00 | Death | RESPIRONICS, INC. | 2025-07-17 04:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE RELATIVE OF A PATIENT INQUIRED ABOUT RETURNING A PATIENT'S DEVICE REPORTING THE PATIENT PASSED AWAY. THE PATIENT'S CAUSE OF DEATH WAS NOT PROVIDED. ADDITIONAL INFORMATION IS BEING REQUESTED FROM THE REPORTER SURROUNDING THE DETAILS OF THE PATIENT'S DEATH, AND IF THERE WAS ANY KNOWN DEVICE ISSUE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22532605&pc=DSQ | 2916596-2025-04756 | 2025-06-25 04:00 | Death | THORATEC CORPORATION | 2025-07-17 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT COMMUNICATED. THE DECISION WAS MADE TO END THERAPY DUE TO PATIENT CONDITION BEING TOO SEVERE. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT BE EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22533491&pc=FKX | 0008030665-2025-01620 | 2025-06-28 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-07-17 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K222318 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER ON (B)(6) 2025. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION WERE UNSUCCESSFUL.Event Description: ADDITIONAL INFORMATION WAS PROVIDED BY THE CLINIC MANAGER. THE PATIENT WAS IN TREATMENT AT HOME ON (B)(6) 2025 UTILIZING THE LIBERTY SELECT CYCLER. THE PATIENT WAS IN THE LAST CYCLE DRAIN PHASE WHEN SHE WAS FOUND WITH NO VITAL SIGNS BY HER GRANDDAUGHTER. 911 WAS CALLED. WHEN FIRST RESPONDERS ARRIVED, IT WAS DETERMINED THE PATIENT WAS DECEASED. NO RESUSCITATIVE EFFORTS WERE ADMINISTERED. THE PATIENT WAS TRANSPORTED TO THE FUNERAL HOME. THE CAUSE OF DEATH WAS DETERMINED TO BE CARDIAC ARREST RELATED TO EXISTING COMORBID CONDITIONS. NO PROBLEMS WERE REPORTED WITH THE LIBERTY SELECT CYCLER. THE DEATH WAS NOT RELATED TO USE OF THE LIBERTY SELECT CYCLER OR OTHER FRESENIUS PRODUCT(S). Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION OF A CAUSAL RELATIONSHIP BETWEEN THE PATIENT DEATH AND USE OF THE LIBERTY SELECT CYCLER. IT IS UNKNOWN IF THE PATIENT WAS FOUND DECEASED OR IF RESUSCITATIVE EFFORTS WERE ADMINISTERED. THE PATIENT¿S CAUSE OF DEATH WAS NOT PROVIDED. DIALYSIS PATIENTS ARE AT EXTRAORDINARILY HIGH RISK FOR DEATH WITH CARDIAC DISEASE BEING THE MAJOR CAUSE OF DEATH AND THE SINGLE, LARGEST, SPECIFIC CAUSE OF DEATH BEING ATTRIBUTED TO ARRHYTHMIC MECHANISMS OR SUDDEN CARDIAC ARREST (SCA). BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION A2, A3, A4, B3, B5, AND B7. CORRECTION PROVIDED IN D10 AND G4.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22533668&pc=NGV | 2135147-2025-03892 | 2025-06-24 04:00 | Death | ABBOTT MEDICAL | 2025-07-17 04:00 | NGV | AMPLATZER AMULET | Patient-Device Incompatibility | Cardiac Tamponade; Pericardial Effusion | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, AN 18MM AMPLATZER AMULET WAS SELECTED FOR IMPLANT USING A 12F AMPLATZER AMULET DELIVERY SHEATH. NO PERICARDIAL EFFUSION WAS OBSERVED PRIOR TO PROCEDURE. THE TRANSSEPTAL PUNCTURE WAS MID ANTERO POSTERIOR AND INFERIOR. THE AMULET WAS IMPLANTED UPON THE FIRST ATTEMPT AND WITH "MINIMUM OF MANIPULATIONS INSIDE THE LEFT APPENDAGE." CLOSE CRITERIA WAS MET AND STABILITY TEST WAS PERFORMED AND PASSED. POST-IMPLANT TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED AND NEGATIVE FOR PERICARDIAL EFFUSION. APPROXIMATELY ONE HOUR POST-PROCEDURE, "THE PATIENT COLLAPSED" WHILE IN THE INTENSIVE CARE UNIT (ICU). TRANSTHORACIC ECHOCARDIOGRAM (TTE) WAS PERFORMED, WHICH SHOWED A "MASSIVE" LOCALIZED POSTERIOR LARGE PERICARDIAL EFFUSION; CARDIAC TAMPONADE WAS CONFIRMED PRESENT. THE CAUSE OF THE EFFUSION WAS THOUGHT TO BE EITHER DUE TO THE TRANSSEPTAL PUNCTURE OR TO AMULET MANIPULATION; "NO CLEAR ANSWER" WAS PROVIDED. THE PERICARDIAL SPACE WAS DRAINED. DESPITE THIS AND RESUSCITATION, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED AS "MASSIVE TAMPONADE."Event Description: N/A. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: IT WAS REPORTED THAT PERICARDIAL EFFUSION, CARDIAC TAMPONADE AND PATIENT DEATH WERE REPORTED. THERE ARE MULTIPLE RISK FACTORS FOR PERICARDIAL EFFUSION AFTER A LEFT ATRIAL APPENDAGE CLOSURE PROCEDURE, INCLUDING ANTICOAGULATION STATUS OF THE PATIENT, MANEUVERING OF THE GUIDEWIRE OR SHEATHS WITHIN THE HEART, THE TRANS-SEPTAL PUNCTURE, OR THE STABILIZING WIRES ON THE AMULET PUNCTURING THE LAA. NO INFORMATION WAS RECEIVED TO INDICATE THAT THE PATIENT¿S MONITORED ACTIVATED CLOTTING TIME WAS ABNORMAL. INFORMATION FROM FIELD INDICATED THAT THE AMULET WAS IMPLANTED WITH MINIMUM MANIPULATIONS. ABBOTT REQUESTED FOR PROCEDURE IMAGING TO REVIEW, THIS WAS NOT PROVIDED BY THE SITE. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE REMAINS IMPLANTED AND WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE CAUSE OF THE REPORTED PERICARDIAL EFFUSION DUE TO THE TRANSSEPTAL PUNCTURE OR TO AMULET MANIPULATION. HOWEVER, THIS CANNOT BE CONFIRMED. THE CAUSE OF CARDIAC TAMPONADE AND DEATH ARE MAYBE CASCADE EVENTS OF EFFUSION. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES AS PERICARDIOCENTESIS WAS PERFORMED. THERE WERE NO COMPLAINTS ASSOCIATED WITH ANY OTHER DEVICES FROM THE LOT. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. FROM MEDICAL REVIEW: WITHOUT ANALYSIS OF PROCEDURAL IMAGING IT IS NOT POSSIBLE TO ASSESS THE LOCATION OF THE TRANSSEPTAL PUNCTURE OR THE MANIPULATION OF THE AMULET DEVICE WITHIN THE LEFT ATRIAL APPENDAGE. AS A RESULT, THE MECHANISM OF CAUSATION OF THE PERICARDIAL EFFUSION CANNOT BE ESTABLISHED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22533779&pc=NPT | 2015691-2025-05829 | 2025-06-25 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Migration or Expulsion of Device | Foreign Body In Patient; Respiratory Arrest | P140031 |  | 1 | Event Description: AS REPORTED BY OUR AFFILIATES IN THE UNITED KINGDOM, THIS WAS A CASE WITH A 23MM SAPIEN 3 ULTRA RESILIA VALVE, IN AORTIC POSITION BY TRANSFEMORAL APPROACH. THE VALVE WAS PREPPED WITHOUT ISSUES. HOWEVER, DURING DEPLOYMENT, THE VALVE EMBOLIZED INTO AORTA, AND WAS LEFT IMPLANTED IN THE ASCENDING AORTA. ALTHOUGH A SECOND VALVE WAS PREPARED, IT COULD NOT BE IMPLANTED AS THE PATIENT EXPERIENCED RESPIRATORY ARREST. THE CASE WAS ABANDONED. CONSEQUENTLY, THE DECISION WAS MADE TO PLACE THE PATIENT ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) OVERNIGHT, WITH THE INTENTION OF PROCEEDING TO SURGERY THE FOLLOWING DAY. THE PATIENT WAS IN THE INTENSIVE THERAPY UNIT. FINALLY, THE PATIENT DID NOT UNDERGO SURGERY AND DECEASED. THE CAUSE OF DEATH WAS AN ENCEPHALOPATHY AND CHRONIC CORTICAL DYSFUNCTION OR ATROPHY. ACCORDING TO MEDICAL OPINION, THE PERCEIVED ROOT CAUSE OF THIS EVENT WAS TOO LITTLE CALCIFICATION TO ANCHOR THE VALVE IN THE NATIVE VALVE. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION OF THE DEVICE, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, LOSS OF PACING CAPTURE, A NARROW SINOTUBULAR JUNCTION (IN AORTIC POSITION PROCEDURES), RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE LANDING ZONE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE TOO LITTLE CALCIFICATION TO ANCHOR THE VALVE IN THE NATIVE VALVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22534572&pc=MLV | 2135147-2025-03895 | 2025-06-25 04:00 | Death | ABBOTT MEDICAL | 2025-07-17 04:00 | MLV | AMPLATZER SEPTAL OCCLUDER | Patient-Device Incompatibility | Erosion; Cardiac Arrest; Cardiac Tamponade; Pericardial Effusion | P000039 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2024, A 20MM AMPLATZER SEPTAL OCCLUDER WAS SELECTED FOR IMPLANT USING A 10F AMPLATZER TREVISIO INTRAVASCULAR DELIVERY SYSTEM. PRIOR TO PROCEDURE, THE PATIENT'S CONDITION WAS GOOD. NO PERICARDIAL EFFUSION WAS OBSERVED PRIOR TO PROCEDURE. "VARIOUS STABILITY TESTS WERE PERFORMED," AND THE OCCLUDER WAS RELEASED FOR IMPLANT. POST-IMPLANT, NO PERICARDIAL EFFUSION WAS OBSERVED. THE PATIENT WAS REPORTED TO HAVE BEEN DOING WELL IN THE AFTERNOON FOLLOWING IMPLANT. ON (B)(6) 2025, IN THE MORNING, CARDIAC TAMPONADE WAS OBSERVED. ULTRASOUND WAS PERFORMED, WHICH SHOWED CIRCUMFERENTIAL PERICARDIAL EFFUSION THAT WAS SUBSEQUENTLY DRAINED SUCCESSFULLY AND THE PATIENT STABILIZED. HOWEVER, APPROXIMATELY ONE HOUR LATER, A SECOND TAMPONADE WAS OBSERVED AND THE PATIENT EXPERIENCED CARDIAC ARREST. THORACIC SURGERY ATTEMPTED TO OPEN THE CHEST TO VIEW THE WOUND BUT NOTED PULSATILE RED BLOOD EJECTION. THE PATIENT WAS REPORTED TO HAVE PASSED. THE CAUSE OF DEATH WAS SUSPECTED AS "PROBABLY AORTA EROSION."Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 20MM AMPLATZER SEPTAL OCCLUDER WAS SELECTED FOR IMPLANT USING A 10F AMPLATZER TREVISIO INTRAVASCULAR DELIVERY SYSTEM. PRIOR TO PROCEDURE, THE PATIENT'S CONDITION WAS GOOD. NO PERICARDIAL EFFUSION WAS OBSERVED PRIOR TO PROCEDURE. "VARIOUS STABILITY TESTS WERE PERFORMED," AND THE OCCLUDER WAS RELEASED FOR IMPLANT. POST-IMPLANT, NO PERICARDIAL EFFUSION WAS OBSERVED. THE PATIENT WAS REPORTED TO HAVE BEEN DOING WELL IN THE AFTERNOON FOLLOWING IMPLANT. ON (B)(6) 2025, IN THE MORNING, CARDIAC TAMPONADE WAS OBSERVED. ULTRASOUND WAS PERFORMED, WHICH SHOWED CIRCUMFERENTIAL PERICARDIAL EFFUSION THAT WAS SUBSEQUENTLY DRAINED SUCCESSFULLY AND THE PATIENT STABILIZED. HOWEVER, APPROXIMATELY ONE HOUR LATER, A SECOND TAMPONADE WAS OBSERVED AND THE PATIENT EXPERIENCED CARDIAC ARREST. THORACIC SURGERY ATTEMPTED TO OPEN THE CHEST TO VIEW THE WOUND BUT NOTED PULSATILE RED BLOOD EJECTION. THE PATIENT WAS REPORTED TO HAVE PASSED. THE CAUSE OF DEATH WAS SUSPECTED AS "PROBABLY AORTA EROSION." Manufacturer Narrative: AN EVENT OF CIRCUMFERENTIAL PERICARDIAL EFFUSION, CARDIAC TAMPONADE, PATIENT DEATH ONE DAY POST SEPTAL OCCLUDER IMPLANT WAS REPORTED. THE PRESUMED CAUSE OF DEATH WAS REPORTED AS AORTA EROSION. INFORMATION FROM FIELD INDICATED THAT PFO CLOSURE WAS ATTEMPTED INITIALLY USING AN 30-25 MM AMPLATZER TALISMAN PFO OCCLUDER, HOWEVER, THE DEVICE WAS NOT STABLE, AND A DECISION WAS MADE TO IMPLANT AN ASO. THE DEVICE REMAINED IMPLANTED AND WILL NOT RETURN TO ABBOTT FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. THE CAUSE OF REPORTED INCIDENT AND PATIENT EFFECTS COULD NOT BE DETERMINED. BASED ON THE INFORMATION RECEIVED AND MEDICAL REVIEW, 20 MM ASO TO CLOSE THE PFO IN THIS PATIENT WAS TOO LARGE FOR THE ENCOUNTERED ANATOMY AND RESULTED IN IMPINGEMENT INTO THE AORTA WITH SUBSEQUENT EROSION. HOWEVER, THIS COULD NOT BE CONFIRMED. THE REPORTED PATIENT EFFECTS OF PERICARDIAL EFFUSION, CARDIAC TAMPONADE, EROSION, CARDIAC ARREST AND PATIENT DEATH APPEAR TO BE A CASCADING EFFECT. THE REPORTED UNEXPECTED MEDICAL INTERVENTION IS A RESULT OF CASE-SPECIFIC CIRCUMSTANCE AS PERICARDIOCENTESIS WAS PERFORMED TO DRAIN THE EFFUSION. THERE WERE NO RELATED COMPLAINTS ASSOCIATED WITH ANY OTHER DEVICES FROM THE LOT. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22534609&pc=NPT | 2025587-2025-05096 | 2025-07-09 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-17 04:00 | NPT | EVOLUT FX DCS | Failure to Advance; Insufficient Information; Patient Device Interaction Problem | Exsanguination; Hemorrhage/Bleeding; Hemothorax; Low Blood Pressure/ Hypotension | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE, THE VALVE WAS INTRODUCED INTO THE PATIENT BY A NON MEDTRONIC GUIDEWIRE (INNOWI). AN ADDITIONAL BACKUP WIRE (MAYER STIFF) WAS PUT IN PLACE AS A BUDDY WIRE TO STRAIGHTEN THE KINK. WHILE PASSING THE THORACIC AORTA, NO CHALLENGES WERE NOTED. ONCE TRYING TO CROSS THE AORTIC ARCH, THE DELIVERY CATHETER SYSTEM (DCS) BECAME STUCK. AFTER TWO ATTEMPTS CROSSING THE AORTIC ARCH FAILED, THE GUIDEWIRE WAS EXCHANGED FOR A NEW NON-MEDTRONIC GUIDEWIRE (LUNDERQUIST). DURING THE NEW ATTEMPT TO CROSS THE ARCH WITH THE NEW GUIDEWIRE, THE DCS BECAME STUCK AGAIN. WHILE ADVANCING WITH MORE PRESSURE TO THE DCS, THE DCS LOST "SUDDEN BACKUP IN THE KINK." A DROP IN BLOOD PRESSURE WAS OBSERVED (TO 50/30 MILLIMETERS OF MERCURY (MMHG)). CARDIOPULMONARY RESUSCITATION AND INTUBATION WERE PERFORMED. A TRANSESOPHAGEAL ECHOCARDIOGRAM SHOWED NO PERICARDIAL EFFUSION, ANGIOGRAPHY COULD NOT CONFIRM THE PERFORATION, AND A TRANSTHORACIC ECHOCARDIOGRAM CONFIRMED A SEVERE HEMATOTHORAX ON THE RIGHT LUNG SIDE. A PERFORATION OF THE DESCENDING AORTA OCCURRED DUE TO SEVERE KINKING OF THE VESSEL ASSOCIATED WITH SCOLIOSIS. SEVERE BLOOD LOSS WAS OBSERVED, RESULTING IN A HEMATOTHORAX ON THE RIGHT SIDE. A THORAX DRAINAGE WAS APPLIED, AND APPROXIMATELY 2 LITERS OF BLOOD WERE DRAINED. THE PATIENT WAS DECLARED DEAD DUE TO SEVERE BLOOD LOSS AND HEMATOTHORAX RESULTING FROM THE AORTIC PERFORATION.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE, THE VALVE WAS INTRODUCED INTO THE PATIENT BY A NON MEDTRONIC GUIDEWIRE (INNOWI). AN ADDITIONAL BACKUP WIRE (MAYER STIFF) WAS PUT IN PLACE AS A BUDDY WIRE TO STRAIGHTEN THE KINK. WHILE PASSING THE THORACIC AORTA, NO CHALLENGES WERE NOTED. ONCE TRYING TO CROSS THE AORTIC ARCH, THE DELIVERY CATHETER SYSTEM (DCS) BECAME STUCK. AFTER TWO ATTEMPTS CROSSING THE AORTIC ARCH FAILED, THE GUIDEWIRE WAS EXCHANGED FOR A NEW NON-MEDTRONIC GUIDEWIRE (LUNDERQUIST). DURING THE NEW ATTEMPT TO CROSS THE ARCH WITH THE NEW GUIDEWIRE, THE DCS BECAME STUCK AGAIN. WHILE ADVANCING WITH MORE PRESSURE TO THE DCS, THE DCS LOST "SUDDEN BACKUP IN THE KINK." A DROP IN BLOOD PRESSURE WAS OBSERVED (TO 50/30 MILLIMETERS OF MERCURY (MMHG)). CARDIOPULMONARY RESUSCITATION AND INTUBATION WERE PERFORMED. A TRANSESOPHAGEAL ECHOCARDIOGRAM SHOWED NO PERICARDIAL EFFUSION, ANGIOGRAPHY COULD NOT CONFIRM THE PERFORATION, AND A TRANSTHORACIC ECHOCARDIOGRAM CONFIRMED A SEVERE HEMATOTHORAX ON THE RIGHT LUNG SIDE. A PERFORATION OF THE DESCENDING AORTA OCCURRED DUE TO SEVERE KINKING OF THE VESSEL ASSOCIATED WITH SCOLIOSIS. SEVERE BLOOD LOSS WAS OBSERVED, RESULTING IN A HEMATOTHORAX ON THE RIGHT SIDE. A THORAX DRAINAGE WAS APPLIED, AND APPROXIMATELY 2 LITERS OF BLOOD WERE DRAINED. THE PATIENT WAS DECLARED DEAD DUE TO SEVERE BLOOD LOSS AND HEMATOTHORAX RESULTING FROM THE AORTIC PERFORATION. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE SEVERE CALCIFICATION AND TORTUOUS PATIENT ANATOMY CONTRIBUTED TO THE PERFORATION. IT WAS ALSO NOTED THE CAUSE OF THE HEMOTHORAX WAS DUE TO THE PERFORATION OF THE DESCENDING AORTA WITH THE DCS. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVFXPLUS-34 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID L-EVOLUTFX-34 (LOT: 0012768946); PRODUCT TYPE: 0195-HEART VALVES; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW: THERE WERE FIVE MEDIA IMAGES OF THE EVENT REPORTED. THERE WAS NO COMPUTED TOMOGRAPHY (CT) PROVIDED FOR ANATOMICAL CONSIDERATIONS EXECUTIVE SUMMARY WAS PROVIDED AND EVIDENCE SHOWS TORTUOSITY IN THE AORTA. A FLUROSCOPY VALVE LOAD INSPECTION WAS PROVIDED AND EVIDENCE SHOWS THERE IS NO RESEMBLANCE OF A MISLOAD. IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE, THE VALVE WAS INTRODUCED INTO THE PATIENT BY A NON MEDTRONIC GUIDEWIRE. AN ADDITIONAL BACKUP WIRE WAS PUT IN PLACE AS A BUDDY WIRE TO STRAIGHTEN THE KINK . WHILE PASSING THE THORACIC AORTA, NO CHALLENGES WERE NOTED. ONCE TRYING TO CROSS THE AORTIC ARCH, THE DELIVERY CATHETER SYSTEM (DCS) BECAME STUCK. AFTER TWO ATTEMPTS CROSSING THE AORTIC ARCH FAILED, THE GUIDEWIRE WAS EXCHANGED FOR A NEW NON-MEDTRONIC GUIDEWIRE. DURING THE NEW ATTEMPT TO CROSS THE ARCH WITH THE NEW GUIDEWIRE, THE DCS BECAME STUCK AGAIN. WHILE ADVANCING WITH MORE PRESSURE TO THE DCS, THE DCS LOST "SUDDEN BACKUP IN THE KINK." A DROP IN BLOOD PRESSURE WAS OBSERVED (TO 50/30 MILLIMETERS OF MERCURY (MMHG)). CARDIOPULMONARY RESUSCITATION AND INTUBATION WERE PERFORMED. AS STATED IN THE MEDTRONIC INSTRUCTIONS FOR USE (IFU); POTENTIAL RISKS ASSOCIATED WITH THE IMPLANTATION OF THE EVOLUT FX BIOPROSTHESIS MAY INCLUDE, BUT ARE NOT LIMITED TO, THE FOLLOWING: - DEATH-CARDIOVASCULAR INJURY (INCLUDING RUPTURE, PERFORATION, TISSUE EROSION, OR DISSECTION OF VESSELS, ASCENDING AORTA TRAUMA, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION) -MAJOR OR MINOR BLEEDING THAT MAY REQUIRE TRANSFUSION OR INTERVENTION (INCLUDING LIFE-THREATENING OR DISABLING BLEEDING). IT WAS REPORTED THAT, A PERFORATION OF THE DESCENDING AORTA OCCURRED DUE TO SEVERE KINKING OF THE VESSEL ASSOCIATED WITH SCOLIOSIS. WITH IMAGES PROVIDED IT CANNOT BE CONFIRMED OR DENIED PERFORATION. EVIDENCE DOES SHOW SEVER TORTUOSITY IN THE AORTA. SEVERE BLOOD LOSS WAS OBSERVED, RESULTING IN A HEMATOTHORAX ON THE RIGHT SIDE. A THORAX DRAINAGE WAS APPLIED, AND APPROXIMATELY 2 LITERS OF BLOOD WERE DRAINED. THE PATIENT WAS DECLARED DEAD DUE TO SEVERE BLOOD LOSS AND HEMATOTHORAX RESULTING FROM THE AORTIC PERFORATION. UPDATED H.6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. CORRECTED DATA: H6. THE ANNEX E CODE REPLACED THE PREVIOUSLY REPORTED ANNEX E E0506. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22534655&pc=BZD | 2518422-2025-048200 | 2022-07-12 04:00 | Death | RESPIRONICS, INC. | 2025-07-17 04:00 | BZD | REP DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A REP DREAMSTATION AUTO BIPAP. THE PATIENT HAS DIED OF HEART ATTACK. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. AN EXTERNAL AND INTERNAL VISUAL INSPECTION OF THE DEVICE WAS COMPLETED BY THE MANUFACTURER AND FOUND DUST-LIKE CONTAMINATION, HAIRLIKE MATERIALS, AN UNKNOWN CONTAMINANT, YELLOW COLORED CONTAMINATION TO THE DEVICE. THE MANUFACTURER FOUND LIQUID INGRESS TO THE BLOWER AND BLOWER BOX. THE MANUFACTURER ALSO FOUND BROWN COLORED CONTAMINATION, WHITE COLORED CONTAMINATION, FIBER LIKE CONTAMINATION, DUST CONTAMINATION AND LIQUID INGRESS IN THE HUMIDIFIER. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THE MANUFACTURER ALSO CONFIRMED LIQUID INGRESS TO THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22534787&pc=OZD | 1220648-2025-30458 | 2025-07-03 04:00 | Death | ABIOMED, INC. | 2025-07-17 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA 5.5 SUFFERED BLOOD LOSS DURING GRAFT ANASTOMOSIS AND WAS TRANSFUSED TWO UNITS OF PACKED RED BLOOD CELLS. ADDITIONALLY, THE PATIENT HAD SUBSTANTIAL BLEEDING FROM A FASCIOTOMY SITE ON THE CALF REQUIRING VOLUME REPLACEMENT AND PRESSORS. THE PATIENT REMAINS ON IMPELLA SUPPORT.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER. ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: ACCESS SITE BLEEDING: THE ROOT CAUSE OF ACCESS SITE BLEEDING WAS MOST LIKELY USE ISSUE SINCE NON-ABIOMED PRODUCT WAS USED WHICH IS NOT RECOMMENDED PER THE INSTRUCTIONS FOR USE. VASCULAR DAMAGE PERIPHERAL VESSEL: THE ROOT CAUSE OF VASCULAR DAMAGE PERIPHERAL VESSEL CANNOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. ADDITIONAL INFORMATION RECEIVED REPORTING THE PATIENT OUTCOME OF DEATH. THE FOLLOWING FIELDS WERE UPDATED: B2 DEATH AND DATE OF DEATH ADDED. B5 PATIENT OUTCOME ADDED. D6B EXPLANT DATE. H1 TYPE OF EVENT UPDATED. H6 HEALTH EFFECT - IMPACT CODE 1802 ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22535237&pc=OZD | 1220648-2025-30460 | 2025-07-03 04:00 | Death | ABIOMED, INC. | 2025-07-17 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Arrhythmia; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ON IMPELLA CP SUPPORT DEVELOPED INCREASING RUNS OF VENTRICULAR TACHYCARDIA THROUGH THE NIGHT. THE PATIENT WAS CANNULATED FOR ECMO TO RESOLVE THE ISSUE. THE PATIENT DEVELOPED A LARGE CLOT IN THE AORTIC ROOT AND ON TOP OF IMPELLA CANNULA AND THE IMPELLA HAD LOW PUMP FLOW. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST SYSTEM LISTS VENTRICULAR ARRYTHMIA AS AN ADVERSE EVENT THAT OCCURRED WHILE IN USE WITH THE IMPELLA DEVICE. IT STATES THAT "ONLY A FRACTION OF THESE RATES WERE ATTRIBUTED TO THE IMPELLA DEVICE AND THE EVENTS RESOLVED WITHOUT RESIDUAL EFFECT IN MOST OF THE CASES, UNLESS THE EVENT OF DEATH OCCURRED. OVERALL, THE RESULTS DID NOT SHOW ANY EVIDENCE OF INCREASED MORBIDITY ASSOCIATED WITH THE IMPELLA SUPPORT IN CARDIOMYOPATHY, MYOCARDITIS, AND PERIPARTUM CARDIOMYOPATHY (PPCM) PATIENTS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22535499&pc=LPB | 2029046-2025-02374 | 2025-06-24 04:00 | Death | BIOSENSE WEBSTER INC | 2025-07-17 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade | P030031 |  | 1 | Event Description: IT WAS REPORTED AT THE END OF A PERSISTENT ATRIAL FIBRILLATION (AFIB) PROCEDURE WITH A THERMOCOOL SMARTTOUCH SF CATHETER, THE PATIENT EXPERIENCED A CARDIAC TAMPONADE TREATED WITH A PERICARDIAL DRAINAGE. IN ADDITION, THE PATIENT HAD A CARDIAC ARREST AND SUBSEQUENTLY THE PATIENT EXPIRED. AT THE END OF THE PROCEDURE, TAMPONADE WAS DETECTED. THE PATIENT WAS NOT WAKING UP, WITH A DROP IN BLOOD PRESSURE. TRANSTHORACIC ECHOCARDIOGRAM TO CONFIRM TAMPONADE AND PERICARDIAL DRAINAGE. CARDIAC ARREST ON THE TABLE AND PATIENT DEATH REPORTED. POSSIBLE CAUSES WERE ABLATION IN THE CORONARY SINUS, LOSS OF TRANSSEPTAL ACCESS FOLLOWED BY MULTIPLE ATTEMPTS TO RE-ENTER THE LEFT ATRIUM. THE BIOSENSE EQUIPMENT WAS NOT IMPLICATED. MULTIPLE ATTEMPTS HAVE BEEN MADE TO OBTAIN CLARIFICATION OF THIS COMPLAINT. HOWEVER, NO FURTHER INFORMATION HAS BEEN MADE AVAILABLE. Manufacturer Narrative: DURING AN INTERNAL REVIEW ON 17-JUL-2025, NOTED A CORRECTION TO THE 3500A INITIAL, UNDER THE D4. PRIMARY UDI NUMBER AS IT SHOULD HAVE BEEN PROCESSED AS (B)(4). THEREFORE, CORRECTED. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER LOT 31620964L AND NO INTERNAL ACTION RELATED TO THE COMPLAINT WAS FOUND DURING THE REVIEW. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22535716&pc=DSQ | 2916596-2025-04813 | 2025-05-29 04:00 | Death | THORATEC CORPORATION | 2025-07-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Asystole | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 AFTER BEING FOUND IN ASYSTOLE. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22535932&pc=FRN | 1416980-2025-03809 | 2025-06-18 04:00 | Death | BAXTER HEALTHCARE CORPORATION | 2025-07-17 04:00 | FRN | NOVUM IQ | Insufficient Flow or Under Infusion | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K211122 |  | 1 | Event Description: IT WAS REPORTED A NOVUM IQ LARGE VOLUME PUMP UNDER-INFUSED. WHILE IN THE INTENSIVE CARE UNIT (ICU) AT 06:35, THE PATIENT RECEIVED A NEW BAG OF NOREPINEPHRINE 8MG/250ML. THE NOREPINEPHRINE INFUSED AT 10 ML/HOUR FROM 06:35 UNTIL 16:50 THAT SAME DAY. THE PATIENT¿S BLOOD PRESSURE (BP) ¿DROPPED.¿ OVER A FOUR-HOUR PERIOD, THE PATIENT¿S BP WAS ¿UNSTABLE¿ AND ¿NUMEROUS ADJUSTMENTS¿ WERE MADE TO THE NOREPINEPHRINE INFUSION RATE. IN RESPONSE TO THE BP DROPPING, THE NURSE TITRATED THE NOREPINEPHRINE UP TO 50ML/HOUR BETWEEN 16:50 AND 20:45. VASOPRESSIN WAS ADDED WHEN THE NOREPINEPHRINE REACHED 50ML/HOUR. AT 20:45 THE PUMP ALARMED THAT THERE IS A ¿LOW VOLUME REMAINING¿ (25 ML OF THE TOTAL PROGRAMMED 250 ML). THE PUMP DISPLAYED 224 ML WAS INFUSED; HOWEVER, THE NOREPINEPHRINE BAG APPEARS TO CONTAIN MORE THAN 25ML. THE MEDICAL TEAM MEASURED THE REMAINING VOLUME IN THE BAG AT 100ML. THEREFORE, THE PATIENT RECEIVED 150ML AND NOT 224ML AS THE PUMP INDICATED. NO FURTHER INFORMATION WAS AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: B5: UPON FOLLOW UP, THE PATIENT REPORTEDLY PASSED AWAY (NO FURTHER INFORMATION PROVIDED ON DATE OF DEATH, CAUSE, INTERVENTION OR IF AN AUTOPSY WAS PERFORMED). H11: A REVIEW OF THE EVENT HISTORY LOG (EHL) DURING THE EVENT DATE RANGE. THE SET WAS LOADED FOR 112 HOURS. PER THE COMPLAINT REPORT THE ALLEGED INCIDENT OCCURRED ON (B)(6) 2025 BETWEEN 06:35 AM UNTIL 16:50 THE SAME DAY AND WAS INFUSING AT 10ML/HOUR. THE DEVICE HISTORY INDICATES THE VOLUME TO BE INFUSED WAS SET TO 240ML AND THE RATE WAS CHANGED THREE TIMES DURING THIS PERIOD AND WAS PROGRAMED TO DELIVER 101ML. THE INFUSION CONTINUED FOR AN ADDITIONAL 4HR 46MIN WHERE THE RATE WAS CHANGED NINE TIMES AND WAS PROGRAMMED TO DELIVER 113.5ML. REVIEW OF THE EHL INDICATES THAT THE REPORTED INCIDENT OCCURRED AFTER THE DEVICE HAD BEEN RUNNING CONTINUOUSLY OR OVER 96 HOURS WHICH EXCEEDS THE PUMP SPECIFICATIONS FOR VOLUMETRIC ACCURACY STATED IN THE PRODUCT LABELING. IN ADDITION, THE DEVICE WAS PROGRAMED TO DELIVER 214ML WHICH IS BELOW THE CLAIMED INFUSION RATE. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22536081&pc=NPT | 2015691-2025-05834 | 2025-06-26 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-17 04:00 | NPT | SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Bradycardia | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES AFFILIATE IN FRANCE, A PATIENT UNDERWENT A TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WITH A 29 MM SAPIEN 3 VALVE. PRE-PROCEDURAL PLANNING HAD INCLUDED CONSIDERATION OF A 26 MM SAPIEN 3 ULTRA VALVE; HOWEVER, A 29 MM SAPIEN 3 VALVE WAS SELECTED FOR IMPLANTATION BASED ON INTRA-PROCEDURAL ASSESSMENT. DURING THE TAVR PROCEDURE, FOLLOWING VALVE DEPLOYMENT, HEMODYNAMIC INSTABILITY OCCURRED, AND A LEFT MAIN CORONARY ARTERY OCCLUSION WAS NOTED DESPITE THE CORONARY HEIGHT BEING VERIFIED PRIOR TO THE PROCEDURE TO BE BETWEEN 13 MM AND 14 MM. THE PATIENT DEVELOPED BRADYCARDIA AND WAS IMMEDIATELY PACED OVER THE WIRE. AN EXTRA BACKUP (EBU) CATHETER WITH RAPID CROSSING OF THE LEFT MAIN CORONARY ARTERY USING A J-SHAPED WIRE WAS ADVANCED INTO THE DISTAL LEFT ANTERIOR DESCENDING (LAD) ARTERY AND A STENT WAS PLACED WITH GOOD RESULTS AND SUCCESSFUL REVASCULARIZATION OF THE LEFT MAIN CORONARY ARTERY. THERE WAS HEMODYNAMIC IMPROVEMENT FOLLOWING INTUBATION, VENTILATION, AND ADMINISTRATION OF AN ADRENALINE BOLUS. SUBSEQUENTLY, THE PATIENT EXPERIENCED HEMODYNAMIC DETERIORATION WITH ONSET OF ELECTROMECHANICAL DISSOCIATION AND DID NOT RESPOND TO RESUSCITATION MANEUVERS. A TRANSTHORACIC ECHOCARDIOGRAM (TTE) CONFIRMED THE VALVE WAS IN THE CORRECT POSITION, NO LEAK, DIFFUSE HYPOKINESIA, AND A LEFT VENTRICULAR EJECTION FRACTION (LVEF) OF 40% UNDER ADRENALINE. A SMALL PERICARDIAL EFFUSION WAS OBSERVED ADJACENT TO THE RIGHT CHAMBERS, BUT WITHOUT COMPRESSIVE SIGNS ON THE RIGHT-SIDED CAVITIES. DESPITE MULTIPLE RESUSCITATION ATTEMPTS, THE PATIENT PASSED AWAY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OR MORE OF THE CORONARY ARTERIES, AND/OR AN AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT TO A CORONARY OBSTRUCTION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, SIGNIFICANT VALVE OVER SIZING, AND VALVE MALPOSITION COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, LANDING ZONE TO CORONARY OSTIA DISTANCE, LENGTH OF THE NATIVE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BALLOON VALVULOPLASTY (BV), PATENCY OF CORONARIES DURING BV, AND EXPANDED HEIGHT OF THE INTENDED VALVE. ASSESSMENT FOR CORONARY COMPRESSION RISK PRIOR TO IMPLANTATION IS RECOMMENDED FOR VALVES IMPLANTED IN PULMONARY POSITION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PROCEDURAL FACTORS (VALVE OVER SIZING) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22537240&pc=MNS | 2518422-2025-107924 | 2025-07-07 04:00 | Death | RESPIRONICS, INC. | 2025-07-17 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION ST30 DEVICE. THE PATIENT¿S STEPDAUGHTER CALLED TO INFORM US THAT THE POLICE HAVE TAKEN THE DEVICE AS PART OF AN INVESTIGATION INTO THE PATIENT¿S SUDDEN DEATH. THEY ARE CHECKING WHETHER THE PATIENT WAS WEARING THE MACHINE WHEN HE DIED. SHE CONFIRMED THAT HE WAS NOT WEARING IT AT THE TIME, ALTHOUGH HE WAS SUPPOSED TO BE USING IT. THE POLICE ARE INVESTIGATING HOW LONG HE MAY HAVE BEEN WITHOUT THE DEVICE. THE STEPDAUGHTER STATED THE PATIENT MOVED FROM THE HOSPITAL TO A CARE HOME AND DIED SHORTLY AFTER ARRIVING. THE PATIENT HAD BEEN SENT A NEW MACHINE BECAUSE HE FELT THE FIRST ONE WASN¿T WORKING. RECORDS SHOW THAT THE PATIENT WAS CONTACTED REGARDING THE ISSUE, AT WHICH TIME THE MACHINE WORKED PROPERLY. THERE IS NO REPORT OF ANY DEVICE MALFUNCTION RELATED TO THE PATIENT¿S DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW- UP REPORT WILL BE FILED
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22537800&pc=BZD | 2518422-2025-107927 | 2024-07-21 04:00 | Death | RESPIRONICS, INC. | 2025-07-17 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Cancer | K113068 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO AN ALLEGATION OF A DEATH OF A PATIENT ON (B)(6) 2024 DUE TO ESOPHAGEAL CANCER THAT SPREAD THROUGH HIS BODY. THE PATIENT'S WIFE BELIEVES THAT THE CPAP MAY HAVE CONTRIBUTED TO THE CANCER. IT IS NOT YET CLEAR IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH OF THE PATIENT. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22537839&pc=QZK | 3007284313-2025-04122 | 2025-06-28 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-07-17 04:00 | QZK | GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: ON (B)(6) 2025, THIS PATIENT UNDERWENT ENDOVASCULAR TREATMENT AND FOR A ZONE 4 DESCENDING THORACIC AORTIC ANEURYSM AND WAS IMPLANTED WITH GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESES (TAMBE). THE COMPLETE TAMBE SYSTEM IS COMPRISED OF THE TAMBE AORTIC COMPONENT, THE BRANCH COMPONENTS (GORE® VIABAHN® (VBX) BALLOON EXPANDABLE ENDOPROSTHESIS), A DISTAL BIFURCATED COMPONENT (GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS COMPONENT) AND CONTRALATERAL LEGS (GORE® EXCLUDER® CONTRALATERAL LEGS). THE PATIENT TOLERATED THE PROCEDURE WITH GOOD RESULTS. ON JUNE 28, 2025, THE ATTENDING PHYSICIAN REPORTED THE PATIENT'S UNEXPECTED PASSING. THE STATED CAUSE OF DEATH WAS CEREBRAL HERNIATION. NOTABLY, THE PATIENT HAD A SPINAL DRAIN PLACED PRIOR TO THE PROCEDURE AND REMAINED STABLE THROUGH THE NIGHT OF (B)(6). ON THE MORNING OF (B)(6), ADDITIONAL CEREBROSPINAL FLUID WAS DRAINED, FOLLOWING WHICH THE PATIENT BECAME ACUTELY UNSTABLE AND SUBSEQUENTLY EXPIRED. HERNIATION OF THE SPINE WAS ALSO REPORTED. A GORE REPRESENTATIVE INVOLVED IN PROCEDURAL SUPPORT NOTED INTRAOPERATIVE LOSS OF PERFUSION TO THE RIGHT HYPOGASTRIC ARTERY, LIKELY CAUSED BY PLAQUE DISPLACEMENT ("SNOW PLOWING") AT THE VESSEL¿S ORIGIN THOUGHT TO BE DUE TO SHEATH USE. DESPITE THIS EVENT, FINAL ASSESSMENT CONFIRMED ADEQUATE PERFUSION TO ALL VISCERAL VESSELS AND A SUCCESSFUL DEVICE SEAL. Manufacturer Narrative: W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22537966&pc=DYB | 3007284313-2025-04123 | 2025-06-28 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-07-17 04:00 | DYB | GORE® DRYSEAL FLEX INTRODUCER SHEATH | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K160254 |  | 1 | Event Description: ON (B)(6) 2025, THIS PATIENT UNDERWENT ENDOVASCULAR TREATMENT AND FOR A ZONE 4 DESCENDING THORACIC AORTIC ANEURYSM AND WAS IMPLANTED WITH GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESES (TAMBE). THE COMPLETE TAMBE SYSTEM IS COMPRISED OF THE TAMBE AORTIC COMPONENT, THE BRANCH COMPONENTS (GORE® VIABAHN® (VBX) BALLOON EXPANDABLE ENDOPROSTHESIS), A DISTAL BIFURCATED COMPONENT (GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS COMPONENT) AND CONTRALATERAL LEGS (GORE® EXCLUDER® CONTRALATERAL LEGS). THE PATIENT TOLERATED THE PROCEDURE WITH GOOD RESULTS. ON (B)(6) 2025, THE ATTENDING PHYSICIAN REPORTED THE PATIENT'S UNEXPECTED PASSING. THE STATED CAUSE OF DEATH WAS CEREBRAL HERNIATION. NOTABLY, THE PATIENT HAD A SPINAL DRAIN PLACED PRIOR TO THE PROCEDURE AND REMAINED STABLE THROUGH THE NIGHT OF (B)(6). ON THE MORNING OF (B)(6), ADDITIONAL CEREBROSPINAL FLUID WAS DRAINED, FOLLOWING WHICH THE PATIENT BECAME ACUTELY UNSTABLE AND SUBSEQUENTLY EXPIRED. HERNIATION OF THE SPINE WAS ALSO REPORTED. A GORE REPRESENTATIVE INVOLVED IN PROCEDURAL SUPPORT NOTED INTRAOPERATIVE LOSS OF PERFUSION TO THE RIGHT HYPOGASTRIC ARTERY, LIKELY CAUSED BY PLAQUE DISPLACEMENT ("SNOW PLOWING") AT THE VESSEL¿S ORIGIN THOUGHT TO BE DUE TO SHEATH USE. DESPITE THIS EVENT, FINAL ASSESSMENT CONFIRMED ADEQUATE PERFUSION TO ALL VISCERAL VESSELS AND A SUCCESSFUL DEVICE SEAL. CODED 1500-E: -OTHER/UNKNOWN TO CAPTURE SUSPECTED "PLAQUE DISPLACEMENT = LOSS OF PERFUSION OF RIGHT HYPOGASTRIC ARTERY. Manufacturer Narrative: W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22538322&pc=BZD | 2518422-2025-107912 | 2025-07-09 04:00 | Death | RESPIRONICS, INC. | 2025-07-17 04:00 | BZD | DREAMSTATION | Degraded | Eye Injury; Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT PASSED AWAY FROM SQUAMOUS CELL CARCINOMA OF THE NOSE, REMOVAL OF THE RIGHT EYE, AND SINONASAL CARCINOMA. MEDICAL INTERVENTION WAS NOT SPECIFIED THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22538473&pc=NGV | 2135147-2025-03905 | 2021-08-01 04:00 | Death | ABBOTT MEDICAL | 2025-07-17 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200049 |  | 1 | Event Description: THE ARTICLE, "CONCOMITANT CATHETER ABLATION AND LEFT ATRIAL APPENDAGE OCCLUSION WITH THE AMULET OCCLUDER FROM EMERGE LAA POST-APPROVAL STUDY", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, MULTICENTER STUDY TO ASSESS OUTCOMES IN PATIENTS UNDERGOING CONCOMITANT CATHETER ABLATION (CA) WITH AN AMULET OCCLUDER IN THE EMERGE LEFT ATRIAL APPENDAGE (EMERGE LAA) POST-APPROVAL STUDY. DEVICES INCLUDED IN THE STUDY WERE SOLELY AMPLATZER AMULET. THE ARTICLE CONCLUDED THAT COMBINED ABLATION WITH LAAO APPROACH IS FEASIBLE AND SAFE WITH THE DUAL-SEAL AMULET OCCLUDER, PROVIDING AN EFFICIENT MEANS TO MANAGE ATRIAL FIBRILLATION AND STROKE RISK. [THE PRIMARY AND CORRESPONDING AUTHOR WAS CHRISTOPHER ELLIS, VANDERBILT HEART AND VASCULAR INSTITUTE, VANDERBILT UNIVERSITY MEDICAL CENTER, 1211 21ST AVENUE SOUTH, MCE 5414, NASHVILLE, TENNESSEE 37232-8802, USA. E-MAIL: CHRISTOPHER.ELLIS@VUMC.ORG]. THE TIME FRAME OF THE STUDY WAS FROM 14 AUGUST 2021 TO 15 DECEMBER 2023. A TOTAL OF 12,151 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. IN THE LAAO+CA GROUP, THE AVERAGE AGE WAS 74.7 YEARS AND THE MAJORITY GENDER WAS MALE. IN THE LAAO-CA GROUP, THE AVERAGE AGE WAS 77.1 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, PRIOR ABLATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF CONCOMITANT CATHETER ABLATION AND LEFT ATRIAL APPENDAGE OCCLUSION WITH THE AMULET OCCLUDER FROM EMERGE LAA POST-APPROVAL STUDY WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING ATRIAL FIBRILLATION, PRIOR ABLATION. SOME OF THE COMPLICATIONS REPORTED WERE STROKE, PERICARDIAL EFFUSION, THROMBOSIS, EMBOLISM, DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: CONCOMITANT CATHETER ABLATION AND LEFT ATRIAL APPENDAGE OCCLUSION WITH THE AMULET OCCLUDER FROM EMERGE LAA POST-APPROVAL STUDY
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22538616&pc=QFG | 3019004087-2025-00656 | 2025-04-15 04:00 | Death | BETA BIONICS, INC | 2025-07-17 04:00 | QFG | ILET BIONIC PANCREAS | Improper or Incorrect Procedure or Method; Obstruction of Flow; Output Problem; Patient Device Interaction Problem | Diabetic Ketoacidosis | K231485 |  | 1 | Event Description: ON 15-APR-2025, BETA BIONICS WAS NOTIFIED BY A HEALTHCARE PROVIDER (HCP) THAT THE USER HAD PASSED AWAY. ACCORDING TO THE HCP, THE ILET WAS NOT CONSIDERED A CONTRIBUTING FACTOR. THE USER HAD REPORTEDLY BEEN HOSPITALIZED DAYS PRIOR FOR DIABETIC KETOACIDOSIS (DKA) AND HAD A DOCUMENTED HISTORY OF MULTIPLE DKA HOSPITALIZATIONS. THE HCP NOTED ONGOING CONCERNS REGARDING INADEQUATE DIABETES MANAGEMENT BY THE USER'S CAREGIVER AND INDICATED THAT THESE CONCERNS HAD BEEN REPORTED TO CHILD PROTECTIVE SERVICES IN THE PAST. THE ILET WAS NOT RETURNED, AND DEVICE LOGS WERE LAST SYNCED ON (B)(6) 2025. Manufacturer Narrative: THE DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THAT THE ILET PASSED ALL MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. DEVICE LOGS WERE NOT AVAILABLE FOR THE DATE OF THE REPORTED DEATH, BUT AVAILABLE DATA SHOWED NO MALFUNCTION ALERTS OR DOSING ANOMALIES. THE ILET RESPONDED TO CGM GLUCOSE VALUES AS INTENDED, AND INSULIN DELIVERY REQUESTS WERE OBSERVED AT REGULAR INTERVALS. NO EVIDENCE OF DEVICE MALFUNCTION WAS IDENTIFIED. BASED ON INFORMATION PROVIDED BY THE HCP AND ENGINEERING LOG REVIEW, THE CAUSE OF DEATH COULD NOT BE LINKED TO A DEVICE-RELATED FAILURE. DEVICE DATA INDICATE THAT ON (B)(6) 2025, THE USER EXPERIENCED HYPOGLYCEMIA BEGINNING AT 6:41 PM, LASTING APPROXIMATELY 40 MINUTES, FOLLOWED BY A SUSTAINED PERIOD OF HYPERGLYCEMIA. THIS MAY HAVE BEEN DRIVEN BY BOTH OVERTREATMENT OF THE LOW- EVIDENCED BY A CGM RISE FROM 69 MG/DL AT 7:16 PM TO 126 MG/DL FIVE MINUTES LATER, WITH CONTINUED ELEVATION THEREAFTER REACHING >400 MG/DL BY 10:40 PM - AND INADEQUATE INSULIN DELIVERY FOLLOWING A 5:54 PM OCCLUSION ALARM. NO CARTRIDGE OR INFUSION SET CHANGE WAS PERFORMED DURING THIS PERIOD, AND MULTIPLE ALARMS WERE TRIGGERED, INCLUDING HIGH GLUCOSE, CGM SIGNAL LOSS, AND LOW BATTERY. THE DEVICE POWERED DOWN DUE TO BATTERY DEPLETION ON (B)(6) 2025 AT 10:47[?]AM, WITH CGM GLUCOSE REMAINING SUSTAINED AT >400MG/DL IN THE HOURS PRECEDING SHUTDOWN. IT REMAINS UNCLEAR WHETHER THE USER RESUMED THERAPY WITH THE ILET THEREAFTER DUE TO LACK OF DEVICE LOG DATA. WHILE THIS SEQUENCE OF EVENTS MAY HAVE CONTRIBUTED TO THE RISK FOR DKA, THE EVENTS THAT OCCURRED BETWEEN DEVICE SHUTDOWN AND THE USER'S DEATH ON (B)(6) 2025 ARE UNKNOWN. OTHER CONTRIBUTING FACTORS ARE LIKELY, INCLUDING A HISTORY OF RECURRENT DKA AND DOCUMENTED SOCIAL CONCERNS FROM THE HEALTHCARE PROVIDER. BASED ON AVAILABLE DEVICE LOG DATA, NO DEVICE MALFUNCTION WAS IDENTIFIED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22516689&pc=OZP | 2032227-2025-213190 | 2025-06-10 04:00 | Death | MEDTRONIC MINIMED | 2025-07-16 04:00 | OZP | UNKNOWN NGP | Adverse Event Without Identified Device or Use Problem | Hypoglycemia; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2025, AND THE CAUSE OF DEATH WAS CARDIOGENIC SHOCK. THE CUSTOMER REPORTED A BLOOD GLUCOSE VALUE OF 21 MG/DL. THE CUSTOMER REPORTED THE HYPOGLYCEMIA WAS TREATED WITH HOSPITALIZATION. THE EVENT INVOLVED PRODUCTS UNK_NGP AND UNOMEDICAL. TROUBLESHOOTING WAS PERFORMED AND FOUND THAT THE DOCUMENT OF ANY HEALTH ISSUES OR ILLNESS THAT MAY HAVE CONTRIBUTED OR LED UP TO PASSING WAS THE LOW BLOOD GLUCOSE. THE CUSTOMER WAS STATED THAT THE PUMP WAS WORN AT THE TIME OF PASSING. IT WAS UNKNOWN WHETHER THE CUSTOMER HAS BEEN USING THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED LOW BLOOD GLUCOSE EVENT OR NOT. IT WAS UNKNOWN WHETHER THE CUSTOMER WAS USED THE AUTO MODE FEATURE AT THE TIME OF THE EVENT OR NOT. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. NO PRODUCT RETURN IS REQUIRED FOR UNK_NGP AND UNOMEDICAL. Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22516717&pc=DSQ | 2916596-2025-04732 | 2025-07-07 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS TRANSITIONED TO COMFORT CARE AND PASSED AWAY. THE PUMP WAS ON AND WORKING AS INTENDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22516720&pc=DSQ | 2916596-2025-04733 | 2025-06-28 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO RIGHT VENTRICULAR FAILURE AND RENAL FAILURE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, RENAL DYSFUNCTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22517536&pc=OZO | 2032227-2025-213495 | 2025-05-24 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-16 04:00 | OZO | 630G INSULIN PUMP MMT-1715KL 630G | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Renal Failure; Urinary Tract Infection; Heart Failure/Congestive Heart Failure | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2025, AND THE CAUSE OF DEATH WAS NATURAL CAUSES. THE EVENT INVOLVED PRODUCTS MMT-1715KL, UNK_RESERVOIR, AND UNOMEDICAL. THE CUSTOMER REPORTED THE HEART FAILURE, RENAL FAILURE, HYPERGLYCEMIA/HYPOGLYCEMIA, HIGH BLOOD PRESSURE, AND URINARY TRACT INFECTION WAS TREATED WITH HOSPITALIZATION. TROUBLESHOOTING WAS PERFORMED, AND IT WAS FOUND THAT THE DOCUMENT OF ANY HEALTH ISSUES OR ILLNESSES THAT MAY HAVE CONTRIBUTED TO OR LED UP TO PASSING WAS THE HEART FAILURE, RENAL FAILURE, DIABETES, HIGH BLOOD PRESSURE, AND URINARY TRACT INFECTION. IT WAS REPORTED THAT THE PUMP WAS NOT WORN AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-1715KL, UNK_RESERVOIR, AND UNOMEDICAL.Event Description: UPDATED SUMMARY:- AS PER CASE NOTES, PUMP WAS NOT WORN AT TIME OF PASSING, THE PUMP WAS LAST WORN ON DATED (B)(6) 2024, CUSTOMER PASSING ON DATED (B)(6) 2025 AND CAUSE OF CUSTOMER PASSING WAS NATURAL CAUSES, SO HERE EVENT SUBMITTED UNDER REGULATORY REPORT 2032227-2025-213495 WAS IN ERROR. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL REVIEW OF THE EVENT AND/OR ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DID NOT AND DOES NOT MEET THE REPORTING REQUIREMENTS STIPULATED IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22518805&pc=MHX | 1218950-2025-000395 | 2025-06-12 04:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-07-16 04:00 | MHX | INTELLIVUE MX40 WLAN | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Syncope/Fainting; Insufficient Information | K113125 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT PASSED AWAY WHILE BEING MONITORED ON THE MX40 TELEMETRY DEVICE AND THE CUSTOMER REQUESTED ASSISTANCE WITH AUDIT LOG INTERPRETATION. THE PATIENT WAS ADMITTED FOR OBSERVATION POST-SYNCOPAL EPISODE WITH A FALL. THE PATIENT THEN PASSED AWAY WHILE ON TELEMETRY AND THE CUSTOMER REQUESTED INTERPRETATION OF THE CLINICAL AUDIT LOGS. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. REPORTING INSTITUTION PHONE # (B)(6). Manufacturer Narrative: ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT THE CUSTOMER HAS ONLY PROVIDED LIMITED INFORMATION AT THIS TIME. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: PATIENT INVOLVEMENT WAS REPORTED. THE PATIENT PASSED AWAY WHILE BEING MONITORED ON THE MX40 TELEMETRY DEVICE AND THE CUSTOMER REQUESTED ASSISTANCE WITH AUDIT LOG INTERPRETATION. THE PATIENT WAS ADMITTED FOR OBSERVATION POST-SYNCOPAL EPISODE WITH A FALL. THE PATIENT THEN PASSED AWAY WHILE ON TELEMETRY AND THE CUSTOMER REQUESTED INTERPRETATION OF THE CLINICAL AUDIT LOGS. IT IS NOT ALLEGED THAT THERE WAS ANY MALFUNCTION WITH ANY PHILIPS DEVICE. A PHILIPS CLINICAL PRODUCT SPECIALIST (CPS) REVIEWED THE LOGS PROVIDED. ACCORDING TO THE LOGS AND THE PIC IX CENTRAL STATION CONFIGURATION, THE RF AUTOSHUTOFF WAS "ON" MEANING THAT 10 MINUTES AFTER LEADS ARE OFF, THE TRANSMITTER TURNS OFF AND SENDS "TRANSMITTER OFF INOP" TO THE PIC IX. THE LEADS REMAINED OFF THE PATIENT FOR THE AFTERNOON INTO THE EVENING AND IT KEPT GENERATING "ECG LEADS OFF" AND THEN TRANSMITTER OFF REPEATEDLY. RF AUTOSHUTOFF ENABLED/ON WAS THE DEFAULT IN PIC IX B, FROM WHAT CPS COULD SEE IT WAS WORKING AS DESIGNED AND CONFIGURED. IT WAS ALSO CONFIRMED IN THE LOGS BY PHILIPS RESEARCH & DEVELOPMENT (R&D), WHEN THE PATIENT ARRIVED ON THE FLOOR, THE DEVICE WAS NOT TAKEN OUT OF STANDBY, THE DEVICE CAME OUT OF STANDBY 10 MIN INTO MONITORING DUE TO THE DETECTION OF THE VALID ECG SOURCE. THE FIRST LEADS OFF INSTANCE WAS A VALID ECG BECAUSE THE ECG LEADS OFF WENT FROM CYAN TO YELLOW, THE YELLOW ECG LEADS OFF WAS ACKNOWLEDGED AND THE DEVICE TURNED OFF 10 MINUTES LATER. THE DEVICE CAME BACK ON EVERY HOUR OR SO DUE TO THE DEVICE DETECTING SOME SIGNAL TO WAKE THE DEVICE UP BUT THE DEVICE NEVER SAW A VALID ECG SIGNAL, THE ECG LEADS OFF REMAINED CYAN EACH TIME IT WOKE UP. IT ALARMED CYAN INOP ONLY. THERE WAS NO FURTHER ECG AVAILABLE FOR THIS PATIENT. IT IS REITERATED THAT THE STAFF DID NOT TAKE THE DEVICE OUT OF STANDBY INITIALLY AND DID NOT RESPOND TO ECG RELATED INOP ALARMS. THE DEVICE WAS FUNCTIONING AS CONFIGURED AND DESIGNED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22518942&pc=OZD | 1220648-2025-30377 | 2025-06-28 04:00 | Death | ABIOMED, INC. | 2025-07-16 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Arrhythmia; Hemorrhage/Bleeding; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WITH ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK WAS INITIALLY SUPPORTED WITH AN IMPELLA CP DEVICE AND SUBSEQUENTLY ESCALATED TO AN IMPELLA 5.5 FOR MECHANICAL CIRCULATORY SUPPORT. DURING 5.5 SUPPORT, THE DEVICE BEGAN TRIGGERING SUDDEN "PURGE BLOCKED" ALARMS. THE TEAM INITIATED TISSUE PLASMINOGEN ACTIVATOR IN THE PURGE LINE. DESPITE THIS INTERVENTION, THERE HAS BEEN NO SIGNIFICANT IMPROVEMENT IN PURGE PRESSURES. THE PATIENT REMAINED ON ESCALATING VASOPRESSOR SUPPORT. THE PATIENT'S CONDITION APPEARS TO BE WORSENING, WITH MULTIPLE EPISODES OF VENTRICULAR TACHYCARDIA AND VENTRICULAR FIBRILLATION REQUIRING DEFIBRILLATION. THE PATIENT WAS RECEIVING AMIODARONE INFUSION BUT CANNOT TOLERATE A BOLUS DUE TO BRADYCARDIA AFTER CONVERSION. ULTIMATELY, CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE, VF/VT/AF/AT, AND VASCULAR DAMAGE (PERIPHERAL VESSEL) HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. HIGH PURGE PRESSURE (HPP): CLINICAL DETAILS STATE THAT THERE WERE PURGE SYSTEM BLOCKED/PURGE PRESSURE HIGH ALARMS ON (B)(6) 2025. NO KINKS/BLOCKAGES WERE OBSERVED. PURGE SOLUTION WAS D5W WITH BICARB. LYSIS SOLUTION WAS ADDED TO PURGE SOLUTION AND INFUSION CONTINUED FOR 30 HOURS BEFORE PURGE FLOW WAS RESTORED AND PRESSURE NORMALIZED. DATA LOG REVIEW SHOWED SMALL TEMPORARY INCREASE IN MOTOR CURRENT LATER FOLLOWED BY A 13 MINUTE PURGE PRESSURE (PP) INCREASE. HPP WAS SUSTAINED FOR ABOUT 33 HOURS BEFORE RESOLVING OVER A 6 HOUR PERIOD. THERE WERE NO MOTOR CURRENT SPIKES PRIOR TO THE PP RISE. THERE WAS A LONG CASSETTE CHANGE BUT IT OCCURRED AFTER THE RISE IN PP. NO PRODUCT WAS RETURNED. THE CAUSE OF THE HIGH PURGE PRESSURE WAS MOST LIKELY BIOMATERIAL BUILDUP AS THERE WAS A GRADUAL RISE IN PP OVER 13 MINUTES AND NO MOTOR CURRENT SPIKES BASED ON DATA LOGS. VF/VT/AF/AT: THE ROOT CAUSE OF THE VT/VF/AT/AF WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. VASCULAR DAMAGE (PERIPHERAL VESSEL): CLINICAL DETAILS STATE THAT THE PATIENT EXPERIENCED INCREASED BLEEDING FROM ORAL AND NASAL CAVITIES AND A DROP IN HEMOGLOBIN REQUIRING 4 UNITS OF PACKED RED BLOOD CELLS (PRBC). TEAM SUSPECTED RETROPERITONEAL BLEED FROM EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) CANNULAS. IT WAS NOTED THAT THE HEPARIN DRIP WAS CONTRIBUTING TO BLEEDING RISK. DATA LOGS ARE NOT APPLICABLE AND PRODUCT WAS NOT RETURNED. THE CAUSE OF THE VASCULAR DAMAGE WAS NOT DETERMINED SINCE INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. H6 HEALTH EFFECT - CLINICAL CODE 1888 AND HEALTH EFFECT - IMPACT CODE 4643 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER. ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22519149&pc=OZD | 1220648-2025-30378 | 2025-06-29 04:00 | Death | ABIOMED, INC. | 2025-07-16 04:00 | OZD | IMPELLA 5.5 | Device Sensing Problem; Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED HAD A 5.5 PUMP SUPPORT PLACED FOR THE PATIENT WITH CGS AND CARDIOMYOPATHY. THE PUMP WAS SUPPORTING WHEN AFTER 18 DAYS THE PUMP WAS FOUND TO BE HAVING PS ISSUES AND THE ALARM WAS DISABLED. ADDITIONALLY THERE WAS VENTRICULAR TACHYCARDIA THAT WAS NEW TO THE PATIENT AND POSSIBLY DUE TO PUMP POSITIONING. THE CARE WAS WITHDRAWN AND PATIENT EXPIRED. Manufacturer Narrative: OPTICAL SIGNAL ISSUE: CLINICAL DETAILS NOTE THAT THE PATIENT WAS HAVING CONSTANT POSITION IN VENTRICLE ALARMS ON 03-JULY-2025. ECHO VERIFIED VERY GOOD PLACEMENT. NO OTHER DETAILS ARE KNOWN AROUND THIS TIME. DATA LOG REVIEW SHOWS A DECREASE IN MINIMUM SNR AROUND THE TIME OF THE FIRST IMPELLA IN VENTRICLE ALARMS WITH PLACEMENT SIGNAL WAVEFORMS LOOKING VENTRICULAR AND MOTOR CURRENT LOSING PULSATILITY. THE RAW OPTICAL SIGNAL DID NOT SHOW ANY CLEAR ISSUES WITH THE OPTICAL SENSOR. PRODUCT WAS NOT RETURNED. THE CAUSE OF THE OPTICAL SIGNAL ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. VF/VT/AF/AT: DATA LOGS ARE NOT APPLICABLE AND PRODUCT WAS NOT RETURNED. I THE ROOT CAUSE OF THE VT/VF/AT/AF WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. H6 MEDICAL DEVICE PROBLEM CODE 2917 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-30378. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22519342&pc=OZD | 1220648-2025-30380 | 2025-06-30 04:00 | Death | ABIOMED, INC. | 2025-07-16 04:00 | OZD | IMPELLA CP | Failure to Advance; Patient-Device Incompatibility | Ischemia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED DIFFICULT PLACEMENT AS THE PUMP COULD NOT BE ADVANCED BEYOND A LARGE ABDOMINAL ANEURYSM. A COOK 18FR SHEATH WAS USED TO DELIVER THE PUMP. ADDITIONALLY, THE PATIENT'S LIMB WAS PERFUSED AFTER LOW PULSATILITY WAS OBSERVED. LATER, THE FAMILY WITHDREW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22519787&pc=OZD | 1220648-2025-30383 | 2025-06-26 04:00 | Death | ABIOMED, INC. | 2025-07-16 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia; Hemorrhagic Stroke | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED CLOTTED DISTAL PERFUSION CATHETERS AND THE DECISION WAS MADE TO EXPLANT THE PUMP. SYSTEMIC HEPARIN HAD BEEN WITHHELD DUE TO BLEEDING FROM THE MOUTH. FOUR DAYS AFTER EXPLANTATION THE PATIENT EXHIBITED NEUROLOGICAL CHANGES AND A HEAD CT CONFIRMED AN INTRAPARENCHYMAL HEMORRHAGE WITH INTERVENTRICULAR AND SUBARACHNOID EXTENSION RESULTING IN HERNIATION. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION FOR STROKE/CVA , MOUTH BLEEDING, AND LIMB ISCHEMIA HAS BEEN COMPLETED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS AND THE DEVICE EVALUATION, THE ROOT CAUSE OF MOUTH BLEEDING, STROKE, NOR THE LIMB ISCHEMIA CAN BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22519812&pc=DSQ | 2916596-2025-04639 | 2025-07-01 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS SEPSIS. THE ONSET/IDENTIFICATION OF THE SEPSIS WAS (B)(6) 2025, WHICH WAS A DAY AFTER LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION. THE SOURCE OF THE SEPSIS WAS NOT KNOWN SO AN AUTOPSY WAS PLANNED. THE DEVICE WORKED UP TO CUSTOMER EXPECTATIONS, THEREFORE IT IS NOT DEVICE RELATED. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WILL BE PERFORMED. IT IS UNKNOWN WHETHER THE DEVICE WILL BE EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THE HOSPITAL WAS NOT WILLING TO SHARE FURTHER PATIENT RELATED INFORMATION DUE TO DATA PRIVACY POLICY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22519982&pc=LWS | 2017865-2025-93132 | 2025-07-02 04:00 | Death | ABBOTT | 2025-07-16 04:00 | LWS | FORTIFY ASSURA DR ICD, US | Inappropriate or Unexpected Reset | Arrhythmia | P910023 |  | 1 | Event Description: THE PATIENT WAS SEEN IN THE HOSPITAL DUE TO CLINICAL CONDITIONS, DURING THE VISIT IT WAS OBSERVED THAT THE PATIENT EXPERIENCED ARRYTHMIA THAT WAS TREATED WITH EXTERNAL DEFIBRILLATOR. THE PHYSICIAN TRIED TO INTERROGATE THE DEVICE BUT THE DEVICE WAS IN BACK-UP MODE WITH HIGH VOLTAGE THERAPY DISABLED. TECHNICAL SUPPORT WAS CONTACTED AND SUSPECTED THAT THE BACK-UP MODE WAS DUE TO THE EXTERNAL SHOCK. NO FURTHER INTERVENTION HAS BEEN PERFORMED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT HAD PASSED AWAY AFTER EXPERIENCING ARRYTHMIA EPISODES. THE CAUSE OF DEATH WAS SUSPECTED TO BE DUE TO LOW CARDIAC OUTPUT AND AN ARRHYTHMIC STORM. THE DEVICE IS NOT SUSPECTED BY THE PHYSICIAN TO HAVE CAUSE OR CONTRIBUTED TO THE PATIENTS PASSING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22519989&pc=NPT | 2025587-2025-05062 | 2025-07-08 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-16 04:00 | NPT | EVOLUT FX PLUS VALVE | Fluid/Blood Leak; Device Dislodged or Dislocated; Activation Failure; Patient Device Interaction Problem; Central Regurgitation | Cardiac Arrest; High Blood Pressure/ Hypertension; Valvular Insufficiency/ Regurgitation; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE USING THE TRANSCATHETER AORTIC VALVE EVFXPLUS-26 FOR A PATIENT WITH BICUSPID ANATOMY AND SIGNIFICANT CALCIFICATION NEAR THE LEFT CORONARY CUSP, SEVERAL COMPLICATIONS OCCURRED. THE PATIENT REMAINED HYPERTENSIVE THROUGHOUT THE PROCEDURE BUT WAS STABLE WITH ONGOING INTRAVENOUS FLUIDS. THERE WAS DIFFICULTY CROSSING THE VALVE, REQUIRING MULTIPLE CATHETERS AND WIRES BEFORE SUCCESSFUL CROSSING. THE NATIVE VALVE WAS PRE-DILATED SUCCESSFULLY, AND THE VALVE WAS IMPLANTED WITH ONE RECAPTURE TO ADJUST DEPTH. THE VALVE WAS NOTED TO BE CONSTRAINED FROM THE SKIRT, LEADING THE CONSULTANT TO POST-DILATE WITH A 23MM SEMI-COMPLIANT BALLOON. DURING BALLOON INFLATION, THERE WAS A LOSS OF CAPTURE ON THE PACING, AND THE BALLOON POPPED UPWARDS, DRAGGING THE VALVE ABOVE THE NON-CORONARY CUSP. THIS RESULTED IN AORTIC REGURGITATION. WHILE PREPARING TO IMPLANT A NON-MEDTRONIC VALVE TO ADDRESS THE REGURGITATION, THE PATIENT EXPERIENCED CARDIAC ARREST AND DID NOT RECOVER.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE USING THE TRANSCATHETER AORTIC VALVE EVFXPLUS-26 FOR A PATIENT WITH BICUSPID ANATOMY AND SIGNIFICANT CALCIFICATION NEAR THE LEFT CORONARY CUSP, SEVERAL COMPLICATIONS OCCURRED. THE PATIENT REMAINED HYPERTENSIVE THROUGHOUT THE PROCEDURE BUT WAS STABLE WITH ONGOING INTRAVENOUS FLUIDS. THERE WAS DIFFICULTY CROSSING THE VALVE, REQUIRING MULTIPLE CATHETERS AND WIRES BEFORE SUCCESSFUL CROSSING. THE NATIVE VALVE WAS PRE-DILATED SUCCESSFULLY, AND THE VALVE WAS IMPLANTED WITH ONE RECAPTURE TO ADJUST DEPTH. THE VALVE WAS NOTED TO BE CONSTRAINED FROM THE SKIRT, LEADING THE CONSULTANT TO POST-DILATE WITH A 23MM SEMI-COMPLIANT BALLOON. DURING BALLOON INFLATION, THERE WAS A LOSS OF CAPTURE ON THE PACING, AND THE BALLOON POPPED UPWARDS, DRAGGING THE VALVE ABOVE THE NON-CORONARY CUSP. THIS RESULTED IN AORTIC REGURGITATION (AR). WHILE PREPARING TO IMPLANT A NON-MEDTRONIC VALVE TO ADDRESS THE REGURGITATION, THE PATIENT EXPERIENCED CARDIAC ARREST AND DID NOT RECOVER. ADDITIONAL INFORMATION WAS RECEIVED THAT THE VALVE WAS CONSTRAINED BUT NOT INFOLDED, DUE TO THE SEVERE CALCIFICATION EVEN AFTER THE PRE-IMPLANT BALLOON DILATION. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS SEVERE CENTRAL AR FROM VALVE EMBOLIZATION DUE TO THE LOSS OF CAPTURE ON PACING DURING POST-IMPLANT BALLOON DILATION. THE PHYSICIAN ATTRIBUTED THE DEATH TO A PROCEDURAL COMPLICATION THAT CAUSED THE SEVERE AR AND DID NOT ATTRIBUTE THE DEATH DIRECTLY TO THE VALVE OR DCS. BICUSPID ANATOMY AND SEVERE CALCIFICATION OF THE PATIENT'S VALVE WERE REPORTED TO BE CONTRIBUTING FACTORS TO THE EVENT.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE USING THE TRANSCATHETER AORTIC VALVE EVFXPLUS-26 FOR A PATIENT WITH BICUSPID ANATOMY AND SIGNIFICANT CALCIFICATION NEAR THE LEFT CORONARY CUSP, SEVERAL COMPLICATIONS OCCURRED. THE PATIENT REMAINED HYPERTENSIVE THROUGHOUT THE PROCEDURE BUT WAS STABLE WITH ONGOING INTRAVENOUS FLUIDS. THERE WAS DIFFICULTY CROSSING THE VALVE, REQUIRING MULTIPLE CATHETERS AND WIRES BEFORE SUCCESSFUL CROSSING. THE NATIVE VALVE WAS PRE-DILATED SUCCESSFULLY, AND THE VALVE WAS IMPLANTED WITH ONE RECAPTURE TO ADJUST DEPTH. THE VALVE WAS NOTED TO BE CONSTRAINED FROM THE SKIRT, LEADING THE CONSULTANT TO POST-DILATE WITH A 23MM SEMI-COMPLIANT BALLOON. DURING BALLOON INFLATION, THERE WAS A LOSS OF CAPTURE ON THE PACING, AND THE BALLOON POPPED UPWARDS, DRAGGING THE VALVE ABOVE THE NON-CORONARY CUSP. THIS RESULTED IN AORTIC REGURGITATION (AR). WHILE PREPARING TO IMPLANT A NON-MEDTRONIC VALVE TO ADDRESS THE REGURGITATION, THE PATIENT EXPERIENCED CARDIAC ARREST AND DID NOT RECOVER. ADDITIONAL INFORMATION WAS RECEIVED THAT THE VALVE WAS CONSTRAINED BUT NOT INFOLDED, DUE TO THE SEVERE CALCIFICATION EVEN AFTER THE PRE-IMPLANT BALLOON DILATION. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS SEVERE CENTRAL AR FROM VALVE EMBOLIZATION DUE TO THE LOSS OF CAPTURE ON PACING DURING POST-IMPLANT BALLOON DILATION. THE PHYSICIAN ATTRIBUTED THE DEATH TO A PROCEDURAL COMPLICATION THAT CAUSED THE SEVERE AR AND DID NOT ATTRIBUTE THE DEATH DIRECTLY TO THE VALVE OR DCS. BICUSPID ANATOMY AND SEVERE CALCIFICATION OF THE PATIENT'S VALVE WERE REPORTED TO BE CONTRIBUTING FACTORS TO THE EVENT. ADDITIONAL INFORMATION WAS RECEIVED THAT CONFIRMED THE CONNECTOR USED TO PACE OVER THE GUIDEWIRE WAS NOT MANUFACTURED BY MEDTRONIC. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 (THIRD PARAGRAPH). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012768939); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012768953); PRODUCT TYPE: 0195-HEART VALVES; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22520052&pc=OZP | 2032227-2025-214015 | 2025-06-08 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-16 04:00 | OZP | PUMP MMT-1880L MM770G BLE MG | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia | P160017 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2025 IN HOME, AND THE CAUSE OF DEATH WAS DUE TO DIABETES. THE BLOOD GLUCOSE VALUE WAS UNKNOWN. THE EVENT INVOLVED PRODUCTS MMT-1880L, MMT-332A, MMT-7911NA AND MMT-378A. TROUBLESHOOTING WAS PERFORMED. THE CUSTOMER PASSING WAS DUE TO COMPLICATIONS ON DIABETES. THE CUSTOMER WAS WORN THE PUMP AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-7911NA . NO PRODUCT RETURN IS REQUIRED FOR MMT-378A. PRODUCT RETURN FOR MMT-1880L IS EXPECTED AND THE CUSTOMER RESPONSE WAS THE DEVICE WILL BE RETURNED. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22520178&pc=DQY | 3007284313-2025-04112 | 2025-06-19 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-07-16 04:00 | DQY | GORE® MOLDING & OCCLUSION BALLOON CATHETER | Use of Device Problem | Rupture | K172567 |  | 1 | Event Description: THE FIELD SALES ASSOCIATE (FSA) REPORTED THE FOLLOWING TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE TO TREAT AN ANEURYSM - RUPTURE UTILIZING THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM. IT WAS REPORTED THAT THE ANEURYSM WAS EMERGENT, IT HAD NOT YET RUPTURED HOWEVER, IT WAS EXTREMELY ENLARGED. IT WAS OVER 10CM. AFTER IMPLANTATION OF THE TRUNK IPSILATERAL COMPONENT, AN AORTIC OCCLUSION BALLOON WAS USED TO DILATE THE MAIN BODY PORTION OF THE DEVICE. AFTER THE BALLOON WAS DEFLATED, THE PATIENTS PRESSURE MARKEDLY DROPPED. AN ANGIOGRAM WAS PERFORMED WHICH SHOWED A LARGE AMOUNT OF EXTRALUMINAL CONTRAST ORIGINATING FROM THE DISTAL AORTA NECK. AORTIC EXTENDERS WERE DEPLOYED PROXIMALLY TO TRY AND SEAL THE LARGE TYPE 1A ENDOLEAK WHICH WAS PRESENT DUE TO CALCIUM IN THE PROXIMAL NECK. THE PATIENT WAS NOT A CANDIDATE FOR CONVERSION TO OPEN REPAIR SO COMFORT CARE WAS CHOSEN, AND THE PATIENT EXPIRED A COUPLE HOURS AFTER THE PROCEDURE. THE PHYSICIAN BELIEVES THAT CAUSE OF THE REPORTED EVENT WAS DUE TO THE DISTAL AORTIC NECK RUPTURE DURING BALLOON DILATION OF THE PROXIMAL INFRARENAL AORTIC NECK, CALCIUM PRESENT IN THE PROXIMAL SEALING ZONE, AND A LARGE 1A ENDOLEAK THAT COULDN'T BE RESOLVED THAT LED TO EXSANGUINATION. Manufacturer Narrative: W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: ADDED G3/G4, H1/H2, H6. ACCORDING TO THE GORE® MOLDING & OCCLUSION BALLOON CATHETER INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE TRAUMA TO THE VESSEL WALL, INCLUDING RUPTURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22520179&pc=MIH | 3007284313-2025-04113 | 2025-06-19 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-07-16 04:00 | MIH | GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS |  | Rupture | P200030 |  | 1 | Event Description: THE FIELD SALES ASSOCIATE (FSA) REPORTED THE FOLLOWING TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE TO TREAT AN ANEURYSM - RUPTURE UTILIZING THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM. IT WAS REPORTED THAT THE ANEURYSM WAS EMERGENT, IT HAD NOT YET RUPTURED HOWEVER, IT WAS EXTREMELY ENLARGED. IT WAS OVER 10CM. AFTER IMPLANTATION OF THE TRUNK IPSILATERAL COMPONENT, AN AORTIC OCCLUSION BALLOON WAS USED TO DILATE THE MAIN BODY PORTION OF THE DEVICE. AFTER THE BALLOON WAS DEFLATED, THE PATIENTS PRESSURE MARKEDLY DROPPED. AN ANGIOGRAM WAS PERFORMED WHICH SHOWED A LARGE AMOUNT OF EXTRALUMINAL CONTRAST ORIGINATING FROM THE DISTAL AORTA NECK. AORTIC EXTENDERS WERE DEPLOYED PROXIMALLY TO TRY AND SEAL THE LARGE TYPE 1A ENDOLEAK WHICH WAS PRESENT DUE TO CALCIUM IN THE PROXIMAL NECK. THE PATIENT WAS NOT A CANDIDATE FOR CONVERSION TO OPEN REPAIR SO COMFORT CARE WAS CHOSEN, AND THE PATIENT EXPIRED A COUPLE HOURS AFTER THE PROCEDURE. THE PHYSICIAN BELIEVES THAT CAUSE OF THE REPORTED EVENT WAS DUE TO THE DISTAL AORTIC NECK RUPTURE DURING BALLOON DILATION OF THE PROXIMAL INFRARENAL AORTIC NECK, CALCIUM PRESENT IN THE PROXIMAL SEALING ZONE, AND A LARGE 1A ENDOLEAK THAT COULDN'T BE RESOLVED THAT LED TO EXSANGUINATION. Manufacturer Narrative: W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY HAVE NOT BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: ADDED H1/H2, H4, H6, G3/G4: DATE RECEIVED BY MANUFACTURER. CORRECTED D1, D4, G3/G4: PMA/510(K)NUMBER. ACCORDING TO THE GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE ENDOLEAK AND RUPTURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22520760&pc=DXF | 2124215-2025-47521 | 2025-06-30 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-07-16 04:00 | DXF | NRG RF TRANSSEPTAL KIT | Display or Visual Feedback Problem; Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Septic Shock; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | K073326 |  | 1 | Event Description: IT WAS REPORTED A CARDIAC PERFORATION, PERICARDIAL EFFUSION AND CARDIAC TAMPONADE OCCURRED. THE PROCEDURE WAS CANCELLED. DURING A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A NRG TRANSSEPTAL NEEDLE WAS SELECTED FOR USE. PRIOR TO THE PROCEDURE, A TRACE PERICARDIAL EFFUSION WAS NOTED ON BASELINE TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). THE IMPLANTING PHYSICIAN EXPERIENCED DIFFICULTY OBTAINING A CLEAR ANTERIOR-POSTERIOR (AP) VIEW ON ICE BUT OBSERVED A GOOD TENTING OF THE SEPTUM IN THE INFERIOR/SUPERIOR VIEW. THE NRG NEEDLE CROSSED THE SEPTUM EVEN PRIOR TO USING RADIO FREQUENCY ENERGY. THE NEEDLE WAS THEN RETRACTED TO THE RIGHT ATRIUM. AFTER REVIEWING TEE IMAGING, THE PHYSICIAN PROCEEDED WITH SEPTAL CROSSING USING RF ENERGY. AS THE SHEATH WAS EXCHANGED FOR A WAS DOUBLE CURVE SHEATH AND THE NON-BOSTON SCIENTIFIC PIGTAIL CATHETER WAS ADVANCED, THE TEE IMAGING PHYSICIAN NOTED THAT THE PREVIOUSLY OBSERVED TRACE EFFUSION HAD INCREASED IN SIZE. A PERICARDIOCENTESIS TRAY WAS PREPARED, AND CONTRAST IMAGING OF THE LEFT ATRIAL APPENDAGE (LAA) WAS PERFORMED. NO CONTRAST LEAK WAS OBSERVED FROM THE APPENDAGE OR INTO THE PERICARDIAL SPACE. PERICARDIOCENTESIS WAS COMPLETED AND THE PATIENT WAS TREATED WITH PROTAMINE, IV FLUIDS, AND HEMODYNAMIC SUPPORT PROVIDED BY THE CRNA (CERTIFIED REGISTERED NURSE ANESTHETISTS). A CARDIOTHORACIC SURGEON WAS CALLED FOR CONSULTATION. THE PATIENT RECEIVED TWO UNITS OF PACKED RED BLOOD CELLS AND WAS TRANSFERRED TO THE OPERATING ROOM FOR SURGICAL MANAGEMENT. INTRAOPERATIVELY, A PERFORATION WAS IDENTIFIED IN THE LEFT ATRIAL WALL NEAR THE RIGHT INFERIOR PULMONARY VEIN, WHICH WAS DETERMINED TO BE THE SOURCE OF THE EFFUSION. THE PERFORATION WAS SURGICALLY REPAIRED, AND THE LAA WAS LIGATED. THE PATIENT TOLERATED THE PROCEDURE WELL, WITH NO FURTHER COMPLICATIONS, AND WAS TRANSFERRED TO RECOVERY IN STABLE CONDITION. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS.Event Description: IT WAS REPORTED A CARDIAC PERFORATION, PERICARDIAL EFFUSION AND CARDIAC TAMPONADE OCCURRED. THE PROCEDURE WAS CANCELLED. DURING A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A NRG TRANSSEPTAL NEEDLE WAS SELECTED FOR USE. PRIOR TO THE PROCEDURE, A TRACE PERICARDIAL EFFUSION WAS NOTED ON BASELINE TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). THE IMPLANTING PHYSICIAN EXPERIENCED DIFFICULTY OBTAINING A CLEAR ANTERIOR-POSTERIOR (AP) VIEW ON ICE BUT OBSERVED A GOOD TENTING OF THE SEPTUM IN THE INFERIOR/SUPERIOR VIEW. THE NRG NEEDLE CROSSED THE SEPTUM EVEN PRIOR TO USING RADIO FREQUENCY ENERGY. THE NEEDLE WAS THEN RETRACTED TO THE RIGHT ATRIUM. AFTER REVIEWING TEE IMAGING, THE PHYSICIAN PROCEEDED WITH SEPTAL CROSSING USING RF ENERGY. AS THE SHEATH WAS EXCHANGED FOR A WAS DOUBLE CURVE SHEATH AND THE NON-BOSTON SCIENTIFIC PIGTAIL CATHETER WAS ADVANCED, THE TEE IMAGING PHYSICIAN NOTED THAT THE PREVIOUSLY OBSERVED TRACE EFFUSION HAD INCREASED IN SIZE. A PERICARDIOCENTESIS TRAY WAS PREPARED, AND CONTRAST IMAGING OF THE LEFT ATRIAL APPENDAGE (LAA) WAS PERFORMED. NO CONTRAST LEAK WAS OBSERVED FROM THE APPENDAGE OR INTO THE PERICARDIAL SPACE. PERICARDIOCENTESIS WAS COMPLETED AND THE PATIENT WAS TREATED WITH PROTAMINE, IV FLUIDS, AND HEMODYNAMIC SUPPORT PROVIDED BY THE CRNA (CERTIFIED REGISTERED NURSE ANESTHETISTS). A CARDIOTHORACIC SURGEON WAS CALLED FOR CONSULTATION. THE PATIENT RECEIVED TWO UNITS OF PACKED RED BLOOD CELLS AND WAS TRANSFERRED TO THE OPERATING ROOM FOR SURGICAL MANAGEMENT. INTRAOPERATIVELY, A PERFORATION WAS IDENTIFIED IN THE LEFT ATRIAL WALL NEAR THE RIGHT INFERIOR PULMONARY VEIN, WHICH WAS DETERMINED TO BE THE SOURCE OF THE EFFUSION. THE PERFORATION WAS SURGICALLY REPAIRED, AND THE LAA WAS LIGATED. THE PATIENT TOLERATED THE PROCEDURE WELL, WITH NO FURTHER COMPLICATIONS, AND WAS TRANSFERRED TO RECOVERY IN STABLE CONDITION. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS. IT WAS FURTHER CONFIRMED THAT THE PHYSICIAN DID NOT INTEND TO CROSS THE FIRST TIME AND COULD NOT CLEARLY VISUALIZE WHERE THE NEEDLE CROSSED SO HE PULLED IT BACK TO GET THE IDEAL CROSSING. THE PHYSICIAN MADE MENTION THAT THE ANATOMY WAS TOUGH, HEART ROTATED. ACT WAS THERAPEUTIC WHEN CHECKED HOWEVER PROTAMINE WAS GIVEN ONCE THE EFFUSION WAS BEING TAPPED. IT WAS REPORTED THAT THE PATIENT MADE IT THROUGH SURGERY, WAS EXTUBATED ON POST-OP DAY ONE HOWEVER BEGAN TO DECLINE AND PASSED AWAY ON POST-OP DAY THREE BELIEVED TO BE DUE TO SEPTIC SHOCK. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). Manufacturer Narrative: DUE TO ADDITIONAL INFORMATION RECEIVED, THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR THE UPDATED FIELDS DESCRIBE EVENT OR PROBLEM (B5) AND OUTCOMES ATTRIB TO ADV EVENT (B2), IN SECTION B - PATIENT CODES, IMPACT CODES AND DEVICE CODES (F10 AND H6), IN SECTIONS F - USER FACILITY / IMPORTER REPORT INFORMATION AND H - DEVICE MANUFACTURERS ONLY. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. GOOD FAITH EFFORTS TO OBTAIN THE INFORMATION ARE IN PROGRESS. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE FILED IF THE INFORMATION IS RECEIVED. GOOD FAITH EFFORT ATTEMPTS TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT ARE IN PROGRESS, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED YET. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22520864&pc=DSQ | 2916596-2025-04736 | 2025-06-03 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF THE DEATH WAS UNKNOWN. ACCORDING TO THE VENTRICULAR ASSIST DEVICE (VAD) COORDINATOR THE DEATH WAS NOT DEVICE RELATED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS REPORTED THAT NO DETAILS LEADING UP TO THE PATIENT DEATH COULD BE PROVIDED. Manufacturer Narrative: CORRECTED DATA: SECTION G2: REPORT SOURCE CORRECTED TO INCLUDE FOREIGN. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22520967&pc=DSQ | 2916596-2025-04721 | 2025-07-01 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-ORGAN FAILURE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. WHETHER THE PATIENT¿S OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: H6-3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521017&pc=DSQ | 2916596-2025-04739 | 2025-07-09 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Septic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PASSED AWAY DUE TO SEPTIC SHOCK. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: SECTION B2: OUTCOMES ATTRIBUTED TO ADVERSE CORRECTED. SECTION E: INITIAL REPORTER INFORMATION CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521091&pc=OZD | 1220648-2025-30390 | 2025-06-22 04:00 | Death | ABIOMED, INC. | 2025-07-16 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Tachycardia; Hematuria; Ischemia Stroke; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED AN ARRHYTHMIC EVENT, HEMATURIA, THROMBUS AND A STROKE. THE PATIENT WOULD EVENTUALLY PASS AWAY. THE PATIENT PRESENTED TO THE EMERGENCY ROOM WITH CHEST PAIN AND SHORTNESS OF BREATH. THE PATIENT UNDERWENT A CARDIAC CATHETERIZATION WHICH REVEALED MULTIVESSEL DISEASE. THE LEFT ANTERIOR DESCENDING ARTERY WAS THE CULPRIT AND STENTED. THE DECISION WAS MADE TO PLACE AN IMPELLA CP AND SWAN WHICH WAS SUCCESSFULLY PLACE. DUE TO THE PATIENT'S EJECTION FRACTION PERCENTAGE, THE PLAN WAS TO TRANSFER THE PATIENT TO A DIFFERENT HOSPITAL FOR ADVANCE HEART FAILURE WORK UP. APPROXIMATELY NINE HOURS LATER, THERE WERE SUCTION ALARMS. THE PATIENT WAS MOVING AND WAS TACHYCARDIC. THE IMPELLA CP WAS REPOSITIONED INITIALLY MEASURING 5.2 CENTIMETERS WITH SUCTION ALARMS AND AFTER REPOSITIONING MEASURING 3.3 CENTIMETERS WITHIN MID LEFT VENTRICLE AND FREE OF STRUCTURES. THE PERFORMANCE LEVEL WAS AT P-7. IT WAS NOTED THAT URINE WAS PREVIOUSLY PINK TINGED BUT CLEARED TO YELLOW. THE PATIENT WAS ACCEPTED AT THE HIGHER LEVEL OF CARE HOSPITAL AND PLANNED FOR TRANSFER, WHICH WAS COMPLETED. THE FOLLOWING DAY IT WAS NOTED THAT THE PATIENT HAD UNDERGONE A COMPUTED TOMOGRAPHY ANGIOGRAM WHICH WAS POSITIVE FOR A CEREBRAL VASCULAR ACCIDENT. THE PATIENT WAS UNABLE TO MOVE THE LEFT SIDE. THE CARE TEAM IS UNCLEAR AS TO WHERE THE CLOT ORIGINATED AND WAS NOTED THAT THE PATIENT DID NOT RECEIVE PRE-IMPELLA CP PLACEMENT ANGIOGRAM. EVENTUALLY, THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." Manufacturer Narrative: STROKE: THE ROOT CAUSE OF THE STROKE/CVA WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. THROMBUS: THE ROOT CAUSE OF THE THROMBUS WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. VF/VT/AF/AT: THE ROOT CAUSE OF THE VT/VF/AT/AF WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. HEMOLYSIS: THE PUMP WAS NOT RETURNED. DATA LOG REVIEW SHOWED SUCTION ALARMS MOSTLY RESOLVING AFTER A PLACEMENT SIGNAL SPIKE, AROUND THE TIME PUMP NOTED TO BE REPOSITIONED. THE CAUSE OF THE HEMOLYSIS WAS MOST LIKELY IMPROPER POSITIONING OF THE DEVICE AS REPOSITIONING THE PUMP FROM 5.2 TO 3.3CM RESOLVED THE ISSUE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521285&pc=LWS | 2649622-2025-19257 | 2025-06-20 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-16 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Intermittent Capture | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF ATRIAL TACHYCARDIA (AT)/ATRIAL FIBRILLATION (AF) RESULTING IN INAPPROPRIATE ATRIAL PACING. ADDITIONALLY, THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED INTERMITTENT LOSS OF CAPTURE. THE LEADS REMAIN IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THE PATIENT PASSED AWAY TWO WEEKS LATER. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521329&pc=NPT | 2015691-2025-05787 | 2025-04-11 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-16 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemorrhage/Bleeding | P140031 |  | 1 | Event Description: AS REPORTED BY THE CLINICAL SPECIALIST, APPROXIMATELY TWO TO THREE DAYS POST SUCCESSFUL TRANSFEMORAL TAVR THE PATIENT PRESENTED TO THE HOSPITAL WITH CHANGES IN THEIR MENTAL STATUS. IMAGING REVEALED THE PATIENT HAD A STROKE. THE PATIENT'S HOSPITAL STAY WAS COMPLICATED BY BLEEDING AFTER BEING ANTICOAGULATED. THERE WAS NO ALLEGATION THAT THE BLEEDING OCCURRED AS A RESULT OF AN INJURY ASSOCIATED WITH AN EDWARDS DEVICE AND NO INFORMATION WAS PROVIDED REGARDING ANY ACTIONS TO TREAT THE BLEEDING, SUCH AS A BLOOD TRANSFUSION. THEIR CONDITION CONTINUED TO DECLINE AND THE FAMILY DECIDED TO WITHDRAW CARE. THE CAUSE OF DEATH WAS ATTRIBUTED TO BOTH THE STROKE AND BLEEDING. OF NOTE, THE PATIENT HAD A HISTORY OF AFIB AND ANTICOAGULATION WAS HELD PRIOR TO THE TAVR PROCEDURE BUT WAS GIVEN IN RESPONSE TO THE STROKE DIAGNOSIS, WHICH THEN LED TO THE BLEEDING. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE EXACT CAUSE OF THE STROKE AND SUBSEQUENT BLEEDING IS UNKNOWN BUT PATIENT FACTORS, SUCH AS THEIR ADVANCED AGE AND ATRIAL FIBRILLATION, MAY HAVE CONTRIBUTED TO THE EVENT, ALONG WITH THE MECHANISMS ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521465&pc=NRY | 2029214-2025-01600 | 2025-06-08 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-07-16 04:00 | NRY | SOLITAIRE FR4 | Break; Adverse Event Without Identified Device or Use Problem; Physical Resistance/Sticking | Ischemia Stroke; No Clinical Signs, Symptoms or Conditions | K183022 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION REGARDING A SOLITAIRE X STENT RETRIEVER (SFR4-6-40-10, LOT: B791192) THAT BROKE DURING A PROCEDURE. IT WAS REPORTED THAT PATIENT WAS UNDERGOING A MECHANICAL THROMBECTOMY PROCEDURE TO TREAT A NON-HEMORRHAGIC STROKE. THE SOLITAIRE AND ALL ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THE SOLITAIRE WAS DEPLOYED FOR CLOT RETRIEVAL. HOWEVER, DURING RETRIEVAL, THERE WAS RESISTANCE IN THE DISTAL END OF THE CATHETER (REBAR, 105-5081-153, LOT: B813957) AND THE SOLITAIRE BROKE AND REMAINED IN THE CLOT. THE SURGEON USED ANOTHER SOLITAIRE X (SFR4-4-40-10, LOT: B777662) TO RETRIEVE THE BROKEN SOLITAIRE ALONG WITH THE CLOT. IT APPEARS THE BROKEN DEVICE WAS SUCCESSFULLY RETRIEVED. HOWEVER, IT WAS NOTED THAT THE PATIENT ULTIMATELY HAD AN OUTCOME OF DEATH.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED AN AUTOPSY WAS NOT PERFORMED. BASELINE CONDITION TICI 0; NO BLOOD FLOW PERFUSION. POST-THROMBECTOMY CONDITION: TICI 1: PARTIAL REPERFUSION, WITH LIMITED BLOOD FLOW RESTORATION TO THE AFFECTED AREA. BASED ON THE INFORMATION FROM PHYSICIAN, THERE ARE MANY FACTOR THAT CAN BE THE CAUSE OF THE DEATH, NOT ONLY BECAUSE THE DEFECT OF THE ITEM. THE CAUSE OF THE SOLITAIRE RESISTANCE WAS NOT DEFINITELY KNOWN; POSSIBLY DUE TO EXTREMELY HARD CLOT. CONTINUOUS SALINE FLUSH WAS ADMINISTERED AND MAINTAINED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THERE WAS NO DAMAGE OBSERVED TO THE REBAR CATHETER. THE SOLITAIRE BROKE IN THE FIRST PASS. THE SOLITAIRE DEVICE WAS NOT TORQUED OR REPOSITIONED DURING DELIVERY OR RETRIEVAL. THE PATIENT DID NOT HAVE VESSEL STENOSIS PROXIMAL TO THE THROMBUS SITE. THE MICROCATHETER TIP COVERED THE SOLITAIRE DEVICE PROXIMAL MARKER DURING THE ATTEMPTED RETRIEVAL. DURING THE PROCEDURE, THE PATIENT REMAINED STABLE, WITH VITAL SIGNS INCLUDING BLOOD PRESSURE WITHIN NORMAL LIMITS. UPON COMPLETION OF THE PROCEDURE AT APPROXIMATELY 2:00 AM, THE PATIENT CONTINUED TO EXHIBIT STABLE VITAL SIGNS. IN ACCORDANCE WITH STANDARD HOSPITAL PROTOCOL, THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) FOR POST-PROCEDURAL RECOVERY. AT APPROXIMATELY 10:00 AM, THE PATIENT WAS REPORTED TO HAVE PASSED AWAY IN THE ICU DUE TO SHOCK. THE SOLITAIREX DEVICE WAS UTILIZED DURING THE PROCEDURE. THE ATTENDING PHYSICIAN HAS CONFIRMED THAT THE PATIENT'S DEATH WAS NOT ATTRIBUTABLE TO ANY PRODUCT DEFECT OR DAMAGE RELATED TO SOLITAIREX. THE DAMAGED COMPONENT OF THE SOLITAIREX DEVICE WAS SUCCESSFULLY RETRIEVED DURING THE PROCEDURE, AND NO REMNANTS WERE LEFT IN THE PATIENT'S BODY. BASED ON THE ATTENDING PHYSICIAN¿S ASSESSMENT AND PROCEDURAL DOCUMENTATION, THE PATIENT¿S DEATH WAS NOT CAUSED BY DEVICE MALFUNCTION OR PROCEDURAL ERROR. ALL STANDARD PROTOCOLS WERE FOLLOWED, AND THE PATIENT WAS MONITORED APPROPRIATELY THROUGHOUT THE PROCESS. IT WAS CLARIFIED THE PATIENT WAS ALREADY UNCONSCIOUS PRIOR TO THE THROMBECTOMY PROCEDURE, ALTHOUGH VITAL SIGNS REMAINED STABLE. THE FAMILY REQUESTED THAT THE THROMBECTOMY STILL BE PERFORMED AND AGREED TO ALL ASSOCIATED RISKS, AS THE PATIENT'S CONDITION WOULD HAVE FURTHER DETERIORATED WITHOUT THE PROCEDURE. AFTER THE THROMBECTOMY PROCEDURE WAS COMPLETED, THE PATIENT'S CONDITION REMAINED THE SAME AS IT WAS PRIOR TO THE OPERATION. THE PATIENT DID NOT PASS AWAY DUE TO SHOCK, BUT DUE TO INSUFFICIENT RESTORATION OF BLOOD FLOW IN THE DISTAL MIDDLE CEREBRAL ARTERY. THE ATTENDING PHYSICIAN HAS CLINICALLY CONFIRMED THAT THERE WAS NO BLEEDING OR VESSEL RUPTURE RESULTING FROM THE THROMBECTOMY PROCEDURE. WHAT WAS MEANT BY "SHOCK" WAS ACTUALLY "DETERIORATION".Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE INDEX PROCEDURE WAS MORE THAN 24 HOURS AFTER THE STROKE BEGAN. THE PATIENT WAS UNCONSCIOUS WHEN BROUGHT TO THE HOSPITAL. THE CLOT WAS HARD. PART OF THE CLOT WAS RETRIEVED WITH THE FIRST SOLITAIRE (SFR4-6-40-10) WHICH IS WHEN THE PUSHWIRE BROKE. THERE WAS NO DAMAGE TO THE MICROCATHETER. THE SOLITAIRE SFR4-6-40-10 WHICH HAS BROKEN WAS SUCCESSFULLY RETRIEVED WITH THE SOLITAIRE X (SFR4-4-40-10). AFTER THAT, THE PHYSICIAN ATTEMPTED ASPIRATION OF THE REMAINING CLOT USING A NON-MEDTRONIC SYSTEM. DESPITE THESE INTERVENTIONS, COMPLETE REVASCULARIZATION WAS NOT ACHIEVED. ANCILLARY DEVICES: PENUMBRA ASPIRATION SYSTEM.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE INDEX PROCEDURE WAS MORE THAN 24 HOURS AFTER THE STROKE BEGAN. THE PATIENT WAS UNCONSCIOUS WHEN BROUGHT TO THE HOSPITAL. THE CLOT WAS HARD. PART OF THE CLOT WAS RETRIEVED WITH THE FIRST SOLITAIRE (SFR4-6-40-10) WHICH IS WHEN THE PUSHWIRE BROKE. THERE WAS NO DAMAGE TO THE MICROCATHETER. THE SOLITAIRE SFR4-6-40-10 WHICH HAS BROKEN WAS SUCCESSFULLY RETRIEVED WITH THE SOLITAIRE X (SFR4-4-40-10). AFTER THAT, THE PHYSICIAN ATTEMPTED ASPIRATION OF THE REMAINING CLOT USING A NON-MEDTRONIC SYSTEM. DESPITE THESE INTERVENTIONS, COMPLETE REVASCULARIZATION WAS NOT ACHIEVED. CORRECTION NOTED THAT TICI WAS 1 PRE-PROCEDURE. IT WAS ALSO NOTED THAT TICI REMAINED 1 AFTER THE PROCEDURE. ANCILLARY DEVICES: PENUMBRA ASPIRATION SYSTEM. Manufacturer Narrative: B5 INFORMATION CORRECTED/UPDATED H3. PRODUCT ANALYSIS #(B)(4): EQUIPMENT USED: VIDEO INSPECTION SYSTEM (M-81805), RULER (M-83360), CAMERA (PANASONIC LUMIX DMC-ZS5) DRAWING(S) REFERENCED: N/A AS FOUND CONDITION: THE SOLITAIRE FR4-6-40-10 STENT DEVICE WAS RETURNED FOR ANALYSIS INSIDE THE OPENED INNER POUCH, SEALED BIOHAZARD BAG, AND A SHIPPING BOX. THE REBAR 18 MICROCATHETER AND GUIDE CATHETER WERE NOT RETURNED. DAMAGED LOCATION DETAILS: THE SOLITAIRE FR4-6-40-10 STENT¿S WORKING LENGTH STRUTS WERE OBSERVED TO BE TANGLED, WHILE THE NON-WORKING LENGTH STRUTS WERE IN GOOD CONDITION. NO IRREGULARITIES WERE FOUND OUTSIDE THE PROXIMAL MARKER, AND NO DEFECTS WERE FOUND ON THE MARKER COIL. THE PUSHER WIRE WAS OBSERVED TO BE BROKEN AT ~16.0CM FROM THE DISTAL TIP WITH THE PTFE OUTER JACKET STILL INTACT. BOTH BROKEN ENDS OF THE PUSHER WIRE WERE SENT OUT FOR SCANNING ELECTRON MICROSCOPY (SEM) FAILURE ANALYSIS TESTING. NO OTHER ANOMALIES WERE FOUND. TESTING/ANALYSIS: DUE TO ITS DAMAGED CONDITION, THE RETURNED SOLITAIRE FR4-6-40-10 STENT DEVICE CANNOT BE USED FOR RESISTANCE TESTING. THE BROKEN ENDS OF THE PUSHER WIRE WERE SENT OUT FOR SCANNING ELECTRON MICROSCOPY (SEM) FAILURE ANALYSIS. THE SEM RESULTS INDICATED THAT, BASED ON THE DIMPLE RUPTURE FEATURES ON THE FRACTURE SURFACE OF BOTH ENDS, IT IS BELIEVED TO HAVE FAILED VIA OVERLOAD. CONCLUSION: BASED ON THE REPORTED INFORMATION, THE CUSTOMER¿S COMPLAINT WAS CONFIRMED. IT WAS OBSERVED THAT THE WORKING LENGTH STRUTS WERE TANGLED, AND THE PUSHER WIRE WAS BROKEN. THE STRUT ENTANGLEMENT LIKELY OCCURRED WHEN ANOTHER SOLITAIRE DEVICE RETRIEVED THE SOLITAIRE FR4-6-40-10 STENT DEVICE. HOWEVER, THE CAUSE COULD NOT BE DETERMINED. IN ADDITION, THE PUSHER WIRE WAS ALSO FOUND TO BE BROKEN AT ~16.0 CM FROM THE DISTAL TIP; IT IS BELIEVED TO HAVE FAILED VIA OVERLOAD. LIKELY CAUSES WERE RETRIEVAL FRICTION WITH THE MICROCATHETER OR HARD CLOTS. HOWEVER, THE CAUSE COULD NOT BE DETERMINED. SINCE THE MICROCATHETER OR GUIDE CATHETER WAS NOT RETURNED, ANY CONTRIBUTION OF THE CATHETERS TO THE COMPLAINT COULD NOT BE DETERMINED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H3. PRODUCT ANALYSIS: EQUIPMENT USED: VIDEO INSPECTION SYSTEM (M-81805), RULER (M-83360), CAMERA (PANASONIC LUMIX DMC-ZS5) AS FOUND CONDITION: THE SOLITAIRE FR4-6-40-10 STENT DEVICE WAS RETURNED FOR ANALYSIS INSIDE THE OPENED INNER POUCH, SEALED BIOHAZARD BAG, AND A SHIPPING BOX. THE REBAR 18 MICROCATHETER AND GUIDE CATHETER WERE NOT RETURNED. DAMAGED LOCATION DETAILS: THE SOLITAIRE FR4-6-40-10 STENT¿S WORKING LENGTH STRUTS WERE OBSERVED TO BE TANGLED, WHILE THE NON-WORKING LENGTH STRUTS WERE IN GOOD CONDITION. NO IRREGULARITIES WERE FOUND OUTSIDE THE PROXIMAL MARKER, AND NO DEFECTS WERE FOUND ON THE MARKER COIL. THE PUSHER WIRE WAS OBSERVED TO BE BROKEN AT ~16.0CM FROM THE DISTAL TIP WITH THE PTFE OUTER JACKET STILL INTACT. BOTH BROKEN ENDS OF THE PUSHER WIRE WERE SENT OUT FOR SCANNING ELECTRON MICROSCOPY (SEM) FAILURE ANALYSIS TESTING. NO OTHER ANOMALIES WERE FOUND. TESTING/ANALYSIS: DUE TO ITS DAMAGED CONDITION, THE RETURNED SOLITAIRE FR4-6-40-10 STENT DEVICE CANNOT BE USED FOR RESISTANCE TESTING. THE BROKEN ENDS OF THE PUSHER WIRE WERE SENT OUT FOR SCANNING ELECTRON MICROSCOPY (SEM) FAILURE ANALYSIS. THE SEM RESULTS INDICATED THAT, BASED ON THE DIMPLE RUPTURE FEATURES ON THE FRACTURE SURFACE OF BOTH ENDS, IT IS BELIEVED TO HAVE FAILED VIA OVERLOAD. CONCLUSION: BASED ON THE REPORTED INFORMATION, THE CUSTOMER¿S COMPLAINT WAS CONFIRMED. IT WAS OBSERVED THAT THE WORKING LENGTH STRUTS WERE TANGLED, AND THE PUSHER WIRE WAS BROKEN. THE STRUT ENTANGLEMENT LIKELY OCCURRED WHEN ANOTHER SOLITAIRE DEVICE RETRIEVED THE SOLITAIRE FR4-6-40-10 STENT DEVICE. HOWEVER, THE CAUSE COULD NOT BE DETERMINED. IN ADDITION, THE PUSHER WIRE WAS ALSO FOUND TO BE BROKEN AT ~16.0 CM FROM THE DISTAL TIP; IT IS BELIEVED TO HAVE FAILED VIA OVERLOAD. LIKELY CAUSES WERE RETRIEVAL FRICTION WITH THE MICROCATHETER OR HARD CLOTS. HOWEVER, THE CAUSE COULD NOT BE DETERMINED. SINCE THE MICROCATHETER OR GUIDE CATHETER WAS NOT RETURNED, ANY CONTRIBUTION OF THE CATHETERS TO THE COMPLAINT COULD NOT BE DETERMINED. H6. CODING UPDATED BASED ON ADDITIONAL INFORMATION RECEIVED AND ANALYSIS FINDINGS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: 105-5081-153, (B813957); IMPLANT DATE N/A; EXPLANT DATE N/A; PRODUCT ID: 103-0606-200, (B810426); IMPLANT DATE: N/A; EXPLANT DATE: N/A; PRODUCT ID: SFR4-4-40-10, (B777662); IMPLANT DATE: N/A; EXPLANT DATE: N/A, MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521504&pc=KRA | 2029214-2025-01601 | 2025-06-08 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-07-16 04:00 | KRA | REBAR | Adverse Event Without Identified Device or Use Problem; Physical Resistance/Sticking | Ischemia Stroke; No Clinical Signs, Symptoms or Conditions | K093750 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION REGARDING A SOLITAIRE X STENT RETRIEVER (SFR4-6-40-10, LOT: B791192) THAT BROKE DURING A PROCEDURE. IT WAS REPORTED THAT PATIENT WAS UNDERGOING A MECHANICAL THROMBECTOMY PROCEDURE TO TREAT A NON-HEMORRHAGIC STROKE. THE SOLITAIRE AND ALL ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THE SOLITAIRE WAS DEPLOYED FOR CLOT RETRIEVAL. HOWEVER, DURING RETRIEVAL, THERE WAS RESISTANCE IN THE DISTAL END OF THE CATHETER (REBAR, 105-5081-153, LOT: B813957) AND THE SOLITAIRE BROKE AND REMAINED IN THE CLOT. THE SURGEON USED ANOTHER SOLITAIRE X (SFR4-4-40-10, LOT: B777662) TO RETRIEVE THE BROKEN SOLITAIRE ALONG WITH THE CLOT. IT APPEARS THE BROKEN DEVICE WAS SUCCESSFULLY RETRIEVED. HOWEVER, IT WAS NOTED THAT THE PATIENT ULTIMATELY HAD AN OUTCOME OF DEATH.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED AN AUTOPSY WAS NOT PERFORMED. BASELINE CONDITION TICI 0; NO BLOOD FLOW PERFUSION. POST-THROMBECTOMY CONDITION: TICI 1: PARTIAL REPERFUSION, WITH LIMITED BLOOD FLOW RESTORATION TO THE AFFECTED AREA. BASED ON THE INFORMATION FROM PHYSICIAN, THERE ARE MANY FACTOR THAT CAN BE THE CAUSE OF THE DEATH, NOT ONLY BECAUSE THE DEFECT OF THE ITEM. THE CAUSE OF THE SOLITAIRE RESISTANCE WAS NOT DEFINITELY KNOWN; POSSIBLY DUE TO EXTREMELY HARD CLOT. CONTINUOUS SALINE FLUSH WAS ADMINISTERED AND MAINTAINED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THERE WAS NO DAMAGE OBSERVED TO THE REBAR CATHETER. THE SOLITAIRE BROKE IN THE FIRST PASS. THE SOLITAIRE DEVICE WAS NOT TORQUED OR REPOSITIONED DURING DELIVERY OR RETRIEVAL. THE PATIENT DID NOT HAVE VESSEL STENOSIS PROXIMAL TO THE THROMBUS SITE. THE MICROCATHETER TIP COVERED THE SOLITAIRE DEVICE PROXIMAL MARKER DURING THE ATTEMPTED RETRIEVAL. DURING THE PROCEDURE, THE PATIENT REMAINED STABLE, WITH VITAL SIGNS INCLUDING BLOOD PRESSURE WITHIN NORMAL LIMITS. UPON COMPLETION OF THE PROCEDURE AT APPROXIMATELY 2:00 AM, THE PATIENT CONTINUED TO EXHIBIT STABLE VITAL SIGNS. IN ACCORDANCE WITH STANDARD HOSPITAL PROTOCOL, THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) FOR POST-PROCEDURAL RECOVERY. AT APPROXIMATELY 10:00 AM, THE PATIENT WAS REPORTED TO HAVE PASSED AWAY IN THE ICU DUE TO SHOCK. THE SOLITAIREX DEVICE WAS UTILIZED DURING THE PROCEDURE. THE ATTENDING PHYSICIAN HAS CONFIRMED THAT THE PATIENT'S DEATH WAS NOT ATTRIBUTABLE TO ANY PRODUCT DEFECT OR DAMAGE RELATED TO SOLITAIREX. THE DAMAGED COMPONENT OF THE SOLITAIREX DEVICE WAS SUCCESSFULLY RETRIEVED DURING THE PROCEDURE, AND NO REMNANTS WERE LEFT IN THE PATIENT'S BODY. BASED ON THE ATTENDING PHYSICIAN¿S ASSESSMENT AND PROCEDURAL DOCUMENTATION, THE PATIENT¿S DEATH WAS NOT CAUSED BY DEVICE MALFUNCTION OR PROCEDURAL ERROR. ALL STANDARD PROTOCOLS WERE FOLLOWED, AND THE PATIENT WAS MONITORED APPROPRIATELY THROUGHOUT THE PROCESS. IT WAS CLARIFIED THE PATIENT WAS ALREADY UNCONSCIOUS PRIOR TO THE THROMBECTOMY PROCEDURE, ALTHOUGH VITAL SIGNS REMAINED STABLE. THE FAMILY REQUESTED THAT THE THROMBECTOMY STILL BE PERFORMED AND AGREED TO ALL ASSOCIATED RISKS, AS THE PATIENT'S CONDITION WOULD HAVE FURTHER DETERIORATED WITHOUT THE PROCEDURE. AFTER THE THROMBECTOMY PROCEDURE WAS COMPLETED, THE PATIENT'S CONDITION REMAINED THE SAME AS IT WAS PRIOR TO THE OPERATION. THE PATIENT DID NOT PASS AWAY DUE TO SHOCK, BUT DUE TO INSUFFICIENT RESTORATION OF BLOOD FLOW IN THE DISTAL MIDDLE CEREBRAL ARTERY. THE ATTENDING PHYSICIAN HAS CLINICALLY CONFIRMED THAT THERE WAS NO BLEEDING OR VESSEL RUPTURE RESULTING FROM THE THROMBECTOMY PROCEDURE. WHAT WAS MEANT BY "SHOCK" WAS ACTUALLY "DETERIORATION".Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE INDEX PROCEDURE WAS MORE THAN 24 HOURS AFTER THE STROKE BEGAN. THE PATIENT WAS UNCONSCIOUS WHEN BROUGHT TO THE HOSPITAL. THE CLOT WAS HARD. PART OF THE CLOT WAS RETRIEVED WITH THE FIRST SOLITAIRE (SFR4-6-40-10) WHICH IS WHEN THE PUSHWIRE BROKE. THERE WAS NO DAMAGE TO THE MICROCATHETER. THE SOLITAIRE SFR4-6-40-10 WHICH HAS BROKEN WAS SUCCESSFULLY RETRIEVED WITH THE SOLITAIRE X (SFR4-4-40-10). AFTER THAT, THE PHYSICIAN ATTEMPTED ASPIRATION OF THE REMAINING CLOT USING A NON-MEDTRONIC SYSTEM. DESPITE THESE INTERVENTIONS, COMPLETE REVASCULARIZATION WAS NOT ACHIEVED. ANCILLARY DEVICES: PENUMBRA ASPIRATION SYSTEM.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE INDEX PROCEDURE WAS MORE THAN 24 HOURS AFTER THE STROKE BEGAN. THE PATIENT WAS UNCONSCIOUS WHEN BROUGHT TO THE HOSPITAL. THE CLOT WAS HARD. PART OF THE CLOT WAS RETRIEVED WITH THE FIRST SOLITAIRE (SFR4-6-40-10) WHICH IS WHEN THE PUSHWIRE BROKE. THERE WAS NO DAMAGE TO THE MICROCATHETER. THE SOLITAIRE SFR4-6-40-10 WHICH HAS BROKEN WAS SUCCESSFULLY RETRIEVED WITH THE SOLITAIRE X (SFR4-4-40-10). AFTER THAT, THE PHYSICIAN ATTEMPTED ASPIRATION OF THE REMAINING CLOT USING A NON-MEDTRONIC SYSTEM. DESPITE THESE INTERVENTIONS, COMPLETE REVASCULARIZATION WAS NOT ACHIEVED. CORRECTION NOTED THAT TICI WAS 1 PRE-PROCEDURE. IT WAS ALSO NOTED THAT TICI REMAINED 1 AFTER THE PROCEDURE. ANCILLARY DEVICES: PENUMBRA ASPIRATION SYSTEM Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID SFR4-6-40-10 (B791192); PRODUCT TYPE: DATE N/A PRODUCT ID 103-0606-200 (B810426); PRODUCT TYPE: PRODUCT ID SFR4-4-40-10 (B777662); PRODUCT TYPE: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5 INFORMATION CORRECTED/UPDATED MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521522&pc=DQX | 2029214-2025-01602 | 2025-06-08 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-07-16 04:00 | DQX | 0.14 GUIDEWIRES | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Ischemia Stroke; No Clinical Signs, Symptoms or Conditions | K113454 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION REGARDING A SOLITAIRE X STENT RETRIEVER (SFR4-6-40-10, LOT: B791192) THAT BROKE DURING A PROCEDURE. IT WAS REPORTED THAT PATIENT WAS UNDERGOING A MECHANICAL THROMBECTOMY PROCEDURE TO TREAT A NON-HEMORRHAGIC STROKE. THE SOLITAIRE AND ALL ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THE SOLITAIRE WAS DEPLOYED FOR CLOT RETRIEVAL. HOWEVER, DURING RETRIEVAL, THERE WAS RESISTANCE IN THE DISTAL END OF THE CATHETER (REBAR, 105-5081-153, LOT: B813957) AND THE SOLITAIRE BROKE AND REMAINED IN THE CLOT. THE SURGEON USED ANOTHER SOLITAIRE X (SFR4-4-40-10, LOT: B777662) TO RETRIEVE THE BROKEN SOLITAIRE ALONG WITH THE CLOT. IT APPEARS THE BROKEN DEVICE WAS SUCCESSFULLY RETRIEVED. HOWEVER, IT WAS NOTED THAT THE PATIENT ULTIMATELY HAD AN OUTCOME OF DEATH. Manufacturer Narrative: INFORMATION RECEIVED INDICATED PLI30 AND PLI40 ARE NO LONGER REPORTABLE. SINCE A REGULATORY REPORT HAS BEEN SENT GCH FUNCTIONALITY DOES NOT ALLOW THE COMPLAINT FLAG TO BE CHANGED TO NO. SHOULD FURTHER INFORMATION MAKE THE EVENT REPORTABLE AGAIN THE EVENT WILL BE REOPENED AND A REGULATORY REPORT WILL BE SENT. SEE SECTION H.6. FOR CORRECTED CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID SFR4-6-40-10 (B791192); IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID 105-5081-153 (B813957); IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID SFR4-4-40-10 (B777662); IMPLANT DATE N/A; EXPLANT DATE N/A MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521543&pc=NRY | 2029214-2025-01603 | 2025-06-08 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-07-16 04:00 | NRY | SOLITAIRE FR4 | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Ischemia Stroke; No Clinical Signs, Symptoms or Conditions | K183022 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION REGARDING A SOLITAIRE X STENT RETRIEVER (SFR4-6-40-10, LOT: B791192) THAT BROKE DURING A PROCEDURE. IT WAS REPORTED THAT PATIENT WAS UNDERGOING A MECHANICAL THROMBECTOMY PROCEDURE TO TREAT A NON-HEMORRHAGIC STROKE. THE SOLITAIRE AND ALL ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THE SOLITAIRE WAS DEPLOYED FOR CLOT RETRIEVAL. HOWEVER, DURING RETRIEVAL, THERE WAS RESISTANCE IN THE DISTAL END OF THE CATHETER (REBAR, 105-5081-153, LOT: B813957) AND THE SOLITAIRE BROKE AND REMAINED IN THE CLOT. THE SURGEON USED ANOTHER SOLITAIRE X (SFR4-4-40-10, LOT: B777662) TO RETRIEVE THE BROKEN SOLITAIRE ALONG WITH THE CLOT. IT APPEARS THE BROKEN DEVICE WAS SUCCESSFULLY RETRIEVED. HOWEVER, IT WAS NOTED THAT THE PATIENT ULTIMATELY HAD AN OUTCOME OF DEATH. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID SFR4-6-40-10 (B791192); PRODUCT TYPE: PRODUCT ID 105-5081-153 (B813957); PRODUCT TYPE: PRODUCT ID 103-0606-200 (B810426); PRODUCT TYPE: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INFORMATION RECEIVED INDICATED PLI30 AND PLI40 ARE NO LONGER REPORTABLE. SINCE A REGULATORY REPORT HAS BEEN SENT GCH FUNCTIONALITY DOES NOT ALLOW THE COMPLAINT FLAG TO BE CHANGED TO NO. SHOULD FURTHER INFORMATION MAKE THE EVENT REPORTABLE AGAIN THE EVENT WILL BE REOPENED AND A REGULATORY REPORT WILL BE SENT. SEE SECTION H.6. FOR CORRECTED CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521562&pc=NVZ | 2182208-2025-03477 | 2025-06-01 04:00 | Death | MEDTRONIC, INC. | 2025-07-16 04:00 | NVZ | MEDTRONIC IPG | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CONDUCTION SYSTEM PACING COMPARED WITH BIVENTRICULAR PACING FOR CARDIAC RESYNCHRONIZATION THERAPY (CRT). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DESCRIBED AS ALL-CAUSE MORTALITY. THERE WERE PATIENTS WHO EXPERIENCED HEART FAILURE HOSPITALIZATIONS AND PROCEDURAL COMPLICATIONS SUCH AS PERICARDIAL EFFUSION, PNEUMOTHORAX, INFECTION, PERFORATION, AND HEMATOMA WHICH REQUIRED INTERVENTION. THERE WERE LEADS WHICH EXHIBITED INCREASES IN THRESHOLDS AND DISLODGEMENTS AND REQUIRED LEAD REVISIONS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/73 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CONDUCTION SYSTEM PACING COMPARED WITH BIVENTRICULAR PACING FOR CARDIAC RESYNCHRONIZATION THERAPY IN PATIENTS WITH HEART FAILURE AND MILDLY REDUCED LEFT VENTRICULAR EJECTION FRACTION: RESULTS FROM INTERNATIONAL COLLABORATIVE LBBAP STUDY (I-CLAS) GROUP. HEART RHYTHM. 2025; 22:1512¿1522. DOI: 10.1016/J.HRTHM.2024.09.030 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521571&pc=OZD | 1220648-2025-30393 | 2025-06-27 04:00 | Death | ABIOMED, INC. | 2025-07-16 04:00 | OZD | IMPELLA CP | Positioning Failure; Restricted Flow rate; Fluid/Blood Leak | Low Blood Pressure/ Hypotension | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WITH ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK WAS IMPLANTED WITH A IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT. DURING TRANSPORT, A REGISTERED NURSE WAS INSTRUCTED TO APPLY MANUAL PRESSURE TO THE ACCESS SITE DUE TO A DEVELOPING HEMATOMA AND WAS KNEELING ON THE PATIENT'S BED. SUBSEQUENTLY, THE PATIENT DEVELOPED HYPOTENSION, ACCOMPANIED BY BOTH A SUCTION ALARM AND A PURGE SYSTEM OPEN ALARM ON THE IMPELLA CONSOLE. TROUBLESHOOTING REVEALED THAT THE PURGE SYSTEM SIDE ARM HAD BECOME LODGED BETWEEN THE BED RAILING, AS REPORTED BY THE NURSE, AND HAD FRACTURED INTO TWO PIECES. A 5 FRENCH MICROPUNCTURE SHEATH WAS INSERTED INTO THE DAMAGED PURGE TUBING TO RESTORE FLOW TO THE CATHETER. DURING THIS MANEUVER, THE IMPELLA CATHETER MIGRATED PROXIMALLY AND RETRACTED INTO THE AORTA. DESPITE REPOSITIONING ATTEMPTS, THE PATIENT'S CONDITION DETERIORATED AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION, SECTION: CLEANING, ¿DO NOT CLEAN WITH OR EXPOSE ANY PART OF THE CLEAR SIDEARM OF THE IMPELLA CATHETER (E.G. INFUSION FILTER, PRESSURE RESERVOIR) TO ALCOHOL. ALCOHOL HAS BEEN SHOWN TO CAUSE CRACKS AND LEAKS IN THESE COMPONENTS. CAREFULLY READ LABELS ON COMMON SKIN PREPS AND LOTIONS TO AVOID USING ANY ALCOHOL-CONTAINING PRODUCTS IN THE AREA OF THE INFUSION FILTER OR PRESSURE RESERVOIR." ¿CLEAN THE CONNECTOR CABLE WITH 70% ISOPROPYL ALCOHOL.¿ Manufacturer Narrative: THE INVESTIGATION OF PURGE LEAK ¿ PUMP, POSITIONING ISSUES, AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE CAUSE OF THE ISSUE WAS A LEAK FROM THE SIDEARM BUT THE EXACT LOCATION OF THE LEAK WAS UNKNOWN. LOGS SHOW PUMP WAS SET TO P9 BUT COULD NOT REACH DESIRED FLOWS AND WAS AT P2 TO P6 RANGE. THERE WERE NO MOTOR CURRENT SPIKES. THERE WERE NO POSITIONING ISSUES OR ALARMS BEFORE LOW PUMP FLOW ONSET BUT AN 'IMPELLA POSITION IN VENTRICLE' THAT SOON RESOLVED OCCURRED. NO OTHER DEVICES OR PATIENT CONDITION ISSUES WERE NOTED. THE ROOT CAUSE OF THE POSITIONING ISSUE AND LOW PUMP FLOW WAS NOT DETERMINED AS LIMITED CLINICAL INFORMATION WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521573&pc=DXY | 2182208-2025-03478 | 2025-06-01 04:00 | Death | MEDTRONIC, INC. | 2025-07-16 04:00 | DXY | MEDTRONIC LEAD | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CONDUCTION SYSTEM PACING COMPARED WITH BIVENTRICULAR PACING FOR CARDIAC RESYNCHRONIZATION THERAPY (CRT). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DESCRIBED AS ALL-CAUSE MORTALITY. THERE WERE PATIENTS WHO EXPERIENCED HEART FAILURE HOSPITALIZATIONS AND PROCEDURAL COMPLICATIONS SUCH AS PERICARDIAL EFFUSION, PNEUMOTHORAX, INFECTION, PERFORATION, AND HEMATOMA WHICH REQUIRED INTERVENTION. THERE WERE LEADS WHICH EXHIBITED INCREASES IN THRESHOLDS AND DISLODGEMENTS AND REQUIRED LEAD REVISIONS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/73 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CONDUCTION SYSTEM PACING COMPARED WITH BIVENTRICULAR PACING FOR CARDIAC RESYNCHRONIZATION THERAPY IN PATIENTS WITH HEART FAILURE AND MILDLY REDUCED LEFT VENTRICULAR EJECTION FRACTION: RESULTS FROM INTERNATIONAL COLLABORATIVE LBBAP STUDY (I-CLAS) GROUP. HEART RHYTHM. 2025; 22:1512¿1522. DOI: 10.1016/J.HRTHM.2024.09.030 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521702&pc=NGV | 2124215-2025-46190 | 2025-01-25 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-16 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: THE PATIENT WAS A PART OF A CLINICAL STUDY. IT WAS REPORTED THAT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 27MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED WITH COMPLETE DEVICE SEAL NOTED. THE PATIENT WAS DISCHARGED ON RIVAROXABAN. 218 DAYS POST INDEX PROCEDURE, THE PATIENT DIED. AT THE TIME OF THE EVENT, THE PATIENT WAS ON RIVAROXABAN. THE CAUSE OF DEATH REMAINS UNKNOWN, NO DIAGNOSTICS AND NO ACTION WAS TAKEN TO TREAT THE EVENT. PER ADDITIONAL INFORMATION RECEIVED, IT WAS STATED "I HAVE ATTEMPTED TO CONTACT THE FAMILY FOR A CAUSE. THERE IS NO CHARTED DOCUMENTATION OF CAUSE ONLY TO STATE THE PATIENT PASSED AWAY" AND NO FURTHER DETAILS WERE MADE AVAILABLE AT THE TIME OF THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521773&pc=BZD | 2518422-2025-107694 | 2025-07-14 04:00 | Death | RESPIRONICS, INC. | 2025-07-16 04:00 | BZD | RENTAL-DREAMSTATION AUTO CPAP, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION AUTO CPAP DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE EXAMINATION HAS NOT BEGUN BECAUSE THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521778&pc=MNS | 2518422-2025-107691 | 2025-07-14 04:00 | Death | RESPIRONICS, INC. | 2025-07-16 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521792&pc=DSQ | 2916596-2025-04651 | 2025-07-06 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Decreased Pump Speed; Infusion or Flow Problem | Cardiac Arrest; Cardiogenic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS REQUIRING INCREASED SUPPORT CHEMICALLY AND MECHANICALLY. THE PATIENT REPORTED A LOW FLOW ALARM AROUND 2 OR 3AM ON (B)(6) 2025. THE PATIENT WAS OFF VASOPRESSORS ON (B)(6) 2025 AND HAD TO BE STARTED AGAIN ON (B)(6) 2025. THE RP IMPELLA WAS REMOVED ON (B)(6) 2025 AND HAD TO BE ADDED BACK ON (B)(6) 2025. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS DONE TO SEE THE FUNCTION OF THE RIGHT VENTRICLE (RV). IT APPEARED THE RV WAS DILATED, AND THE SEPTUM WAS BOWING TOWARD THE LV. A DECISION WAS MADE TO REINSERT THE SWAN/PA CATHETER AND THE RP IMPELLA. LOG FILES WERE SENT FOR REVIEW. THE LOG FILES SHOWED FREQUENT PULSATILITY INDEX (PI) EVENTS THAT WERE OCCURRING FROM 06JUL2025 4:15:33 UNTIL 13:24:30. THE PI EVENTS CAUSED THE VENTRICULAR ASSIST DEVICE (VAD) SPEED TO DROP BELOW ITS LOW-SPEED LIMIT OF 5200 RPMS. LOW SPEED ADVISORY EVENTS WERE ALSO RECORDED WHILE SPEED ADJUSTMENTS WERE BEING MADE. THE PATIENT EXPERIENCED PULSELESS ELECTRICAL ACTIVITY (PEA) ARREST APPROXIMATELY 1245/1250 WITH LOW FLOW ALARMS, MEDICINAL CODE AND NO COMPRESSIONS. THE PATIENT'S PH WAS 7.37, LACTIC 4.7, SODIUM 160, POTASSIUM 4.1, CA++ 1.11, CL- 106, HEMOGLOBIN 7.6, HEMATOCRIT (HCT) 23. LOG FILES WERE SENT FOR REVIEW. LOG FILES FROM 08JUL2025 CAPTURED LOW FLOW ALARM EVENTS ON 08JUL2025 FROM 12:51:28 UNTIL 12:55:34. THERE WERE ALSO SOME MOMENTARY LOW FLOW EVENTS RECORDED.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT PASSED AWAY FROM REFRACTORY CARDIOGENIC SHOCK ON (B)(6) 2025. THE DEVICE WILL NOT BE RETURNING FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B2, B5, H1, REPORT TYPE, AND REPORT SUBTYPE WAS UPDATED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS OF RIGHT VENTRICULAR (RV) FAILURE, PULSELESS ELECTRICAL ACTIVITY (PEA) ARREST, AND REFRACTORY CARDIOGENIC SHOCK COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW ALARMS OCCURRING ON (B)(6) 2025; HOWEVER, THE REPORTED ALARMS FROM (B)(6) 2025 COULD NOT BE CONFIRMED. ACCORDING TO THE ACCOUNT, THE REPORTED LOW FLOW ALARMS OCCURRED IN THE SETTING OF RIGHT VENTRICULAR FAILURE AND PEA. THE SYSTEM CONTROLLER EVENT LOG FILES CONTAINED EVENTS FROM 04:15:33 ON (B)(6) 2025 THROUGH 13:09:48 ON (B)(6) 2025, PER THE TIMESTAMPS. EVENTS CAPTURED PRIOR TO THIS TIMEFRAME WERE OVERWRITTEN BY NEWER INCOMING EVENTS, PER DESIGN. MULTIPLE, TRANSIENT LOW FLOW HAZARD ALARMS WERE CAPTURED BETWEEN 12:52 AND 12:55 ON (B)(6) 2025. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE SET SPEED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. CHAPTER 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING RIGHT HEART FAILURE. THIS CHAPTER ALSO ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. CHAPTER 4 OF THE IFU DESCRIBES PUMP FLOW AND HAZARD ALARMS. THE IFU STATES THAT PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. THIS SECTION ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. CHAPTER 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, STATES THAT SUDDEN RIGHT VENTRICULAR FAILURE MAY DEVELOP DURING OR SHORTLY AFTER PUMP IMPLANTATION AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. CHAPTER 7 OF THE IFU, "ALARMS AND TROUBLESHOOTING", AND SECTION 5 OF THE PATIENT HANDBOOK, ""ALARMS AND TROUBLESHOOTING"", ADDRESS ALL SYSTEM ALARM CONDITIONS, INCLUDING THE LOW FLOW HAZARD, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22521796&pc=DTB | 2649622-2025-19264 | 2025-06-20 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-16 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Pacing Problem; Under-Sensing | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF ATRIAL TACHYCARDIA (AT)/ATRIAL FIBRILLATION (AF)RESULTING IN INAPPROPRIATE ATRIAL PACING. ADDITIONALLY, THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED INTERMITTENT LOSS OF CAPTURE. THE LEADS REMAIN IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THE PATIENT PASSED AWAY TWO WEEKS LATER. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22522024&pc=DSQ | 2916596-2025-04705 | 2025-07-07 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THROUGH PATIENT OUTCOME THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU) A IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION D4: EXPIRATION DATE AND MANUFACTURE DATE (H4) WILL BE PROVIDED UPON INVESTIGATION. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22522156&pc=MNS | 2518422-2025-107699 | 2025-07-14 04:00 | Death | RESPIRONICS, INC. | 2025-07-16 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE EXAMINATION HAS NOT BEGUN BECAUSE THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22522542&pc=DSQ | 2916596-2025-04655 | 2025-07-06 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN THE CARDIOVASCULAR SURGICAL INTENSIVE CARE UNIT (CVSICU) DUE TO RIGHT VENTRICULAR FAILURE (RVF). THE VENTRICULAR ASSIST DEVICE (VAD) HAD FUNCTIONED AS INTENDED AND WAS DECOMMISSIONED.Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED AN INTRACARDIAC CLOT WHICH CAUSED RIGHT VENTRICULAR FAILURE (RVF). THE RIGHT HEART FAILURE DID NOT EXIST PRE-LVAD IMPLANT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, THROMBOEMBOLISM (INCLUDING VENOUS AND ARTERIAL NON-CENTRAL NERVOUS SYSTEM), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THIS SECTION ALSO PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INTERNATIONAL NORMALIZED RATIO (INR) VALUES. ADDITIONALLY, SECTION 6 OF THE IFU (UNDER ¿RIGHT HEART FAILURE¿) STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THIS SECTION (UNDER ¿CAUTION!¿) ALSO EXPLAINS THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22523151&pc=DSQ | 2916596-2025-04657 | 2025-07-07 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Dyspnea; Respiratory Failure; Heart Failure/Congestive Heart Failure; Skin Infection; Swelling/ Edema | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A HISTORY OF IMPLANTABLE CARDIAC MONITOR (ICM) (DILATED CARDIOMYOPATHY, FAMILIAL), CHRONIC HEART FAILURE WITH REDUCED EJECTION FRACTION (HFREF) STATUS POST LEFT VENTRICULAR ASSIST DEVICE (LVAD) ON (B)(6) 2019 (HEARTWARE) WITH LVAD COMPLICATION (RECURRENT CEREBROVASCULAR ACCIDENT (CVA)) AND DEVICE EXCHANGE TO HEARTMATE 3 ON (B)(6) 2022.THE PATIENT HAD A HISTORY OF ARRHYTHMIA PRE-LEFT VENTRICULAR ASSIST DEVICE. THE PATIENT ALSO HAD LEFT PARIETAL/OCCIPITAL CVA IN (B)(6) 2021, RIGHT PARIETAL CVA (B)(6) 2022, SYNCOPAL EPISODES WITH NONSUSTAINED VENTRICULAR TACHYCARDIA (NSVT) STATUS POST MEDTRONIC ICD PLACEMENT, COMPLEX MIGRAINES, TYPE 2 DIABETES MELLITUS (T2DM), HYPERTENSION (HTN), PULMONARY EMBOLISM (PE), PRIOR SUICIDE ATTEMPTS (TURNING OFF LVAD), STAPH EPI BACTEREMIA IN 2019, NOCTURNAL HYPOXEMIA, HEPATIC CYSTS, DIVERTICULOSIS, STAGE 4 CHRONIC KIDNEY DISEASE (CKD), SEGMENTAL/SUBSEGMENTAL ACUTE PE IN 2019, RIGHT INTERNAL JUGULAR VEIN DEEP VEIN THROMBOSIS (RIJ DVT) IN 2019, ANEMIA, RECURRENT CHEST PAIN WHO PRESENTS WITH 50LB WEIGHT GAIN, SHORTNESS OF BREATH AND WORSENING ABDOMINAL AND BILATERAL LOWER EXTREMITY (BLE) EDEMA. THE ONSET DATE OF THE RENAL DYSFUNCTION WAS (B)(6) 2019. THE RENAL DYSFUNCTION WAS NOT DEVICE OR THERAPY RELATED. THE PATIENT WAS ON INTRAVENOUS (IV) DIURESIS AS TRANSITIONED TO BY MOUTH (PO) DIURETICS ON (B)(6) 2025. THE PATIENT NOTED TO HAVE ALTERED MENTAL STATUS ON (B)(6) 2025 AND WAS TRANSFERRED TO CARDIAC INTENSIVE CARE UNI T(CICU). THE PATIENT HAD HEAD CT(CTH) WITHOUT ACUTE ABNORMALITY. THE PATIENT STARTED ON LINEOLID/ZOSYN AND CONSULTED NEUROLOGY AND INTENSIVE CARE UNIT (ICU) PSYCH, AND NEPHROLOGY FOLLOWING FOR ACUTE KIDNEY INJURY (AKI). THE PATIENT'S MENTAL STATUS IMPROVED, AND THE PATIENT WAS TRANSFERRED TO THE FLOOR WITH CONTINUED DIURESIS AND RESUMPTION OF GUIDELINE-DIRECTED MEDICAL THERAPY (GDMT). THE PATIENT HAD A REPEAT EPISODE OF ALTERED MENTAL STATUS (AMS)/TRANSIENT UNRESPONSIVENESS ON (B)(6) 2025. THE PATIENT HAD CTH AGAIN WITHOUT ACUTE ABNORMALITY. THE PATIENT HAD A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WITH DIRECT CURRENT CARDIOVERSION (DCCV) FOR TYPICAL ATRIAL FLUTTER ON (B)(6) 2025 WITH CONVERSION TO NORMAL SINUS RHYTHM (NSR). THE PATIENT WENT BACK INTO AN ATRIAL TACHYCARDIA EARLY MORNING OF (B)(6) 2025. RHC WITH RAMP ON (B)(6) 2025 WITH ELEVATED RIGHT SIDED FILLING PRESSURES, AND THE LVAD SPEED DECREASED TO 6000. THE PATIENT HAD A REPEAT EPISODE OF AMS/TRANSIENT UNRESPONSIVENESS ON (B)(6) 2025, NEUROLOGY AND PSYCHIATRY RE-ENGAGED. NEPHROLOGY RE-ENGAGED AS WELL FOR WORSENING RENAL FUNCTION AND OLIGURIA. THE PATIENT WAS GIVEN DOBUTAMINE 5MCG/KG/MIN ON (B)(6) 2025 DUE TO CONCERN FOR RIGHT HEART FAILURE (RHF) IN SUPPORT OF RISING CREATINE, LOW URINE OUTPUT (UOP), RISING INTERNATIONAL NORMALIZED RATIO (INR). THE PATIENT WAS HYPOTENSIVE WITH MEAN ARTERIAL PRESSURE (MAP) 56, RECEIVED 500CC BOLUS AND DIURETICS HELD OVERNIGHT ON (B)(6) 2025 UNTIL (B)(6) 2025. ON (B)(6) 2025, THE PATIENT HAD ANOTHER EPISODE OF TRANSIENT UNRESPONSIVENESS/AMS. THE PATIENT WITHDREW TO PAINFUL STIMULUS, OPENED EYES SLIGHTLY AND SAID, "I JUST FEEL SLEEPY." THE PATIENT'S INR WAS 5.2. THE PATIENT HAD A COMPUTED TOMOGRAPHY (CT) AND WAS TRANSFERRED TO CICU. POSITIVE BLOOD CULTURES REVEALED A DRIVELINE INFECTION. THE PATIENT STARTED BUMEX 2 GLUCOSE TOLERANCE TEST (GTT) WITH AUGMENTATION WITH METOLAZONE. A BEDSIDE PULMONARY ARTERY (PA) CATHETER WAS PLACED ON (B)(6) 2025. THE PATIENT HAD CONTINUING DIURESIS WITH AUGMENTATION AND WEANING DOBUTAMINE. ON (B)(6) 2025, THE PATIENT HAD A PALLIATIVE CARE CONSULT FOR GOALS OF CARE AND NEW DIARRHEA WITH C. DIFF PANEL ORDERED WHICH WAS NEGATIVE. THE PATIENT REQUIRED DOBUTAMINE 2.5 OVERNIGHT AND WAS PUT ON DIURIL 500 BID AND BUMEX 5 GTT. ON (B)(6) 2025, THE PATIENT HAD DIURIL 1G BID, BY MOUTH ACETAZOLAMIDE, 5 TO 4 BUMEX GTT. THE PATIENT CONTINUED DOBUTAMINE 2.5. THE PATIENT HAD A GOAL OF CARE (GOC) CONVERSATION FOR DIALYSIS VERSUS HOSPICE WITH SISTER INVOLVED. ON (B)(6) 2025 THE PATIENT DISCONTINUED BUMEX AND DIEURL AND DOBUTAMINE WAS 5. THE CRRT WAS UNABLE TO HAVE VOLUME PULLED OFF; LEVOPHED AND VASOPRESSIN NOW ON. GOC CONVERSATION HAD WITH SISTER IN PERSON. THE PATIENT STARTED CRRT ON (B)(6) 2025. THE PATIENT HAD REENGAGED PALLIATIVE CARE AND PLAN FOR GOC ON (B)(6) 2025. ON (B)(6) 2025 THE PATIENT STARTED OCTREOTIDE FOR NOSEBLEEDS AND MELENA, SDS, EPINEPHRINE DRIP, BLOOD TRANSFUSION. THE PATIENT SIGNED THE DNR/DNI. THE PATIENT WAS TRANSITIONED TO COMFORT CARE AFTER DISCUSSIONS WITH FAMILY ON (B)(6) 2025. THE PATIENT PASSED AWAY AT 19:51. THE PRELIMINARY CAUSE OF DEATH WAS ADVANCED HEART FAILURE COMPLICATED BY CARDIORENAL SYNDROME. RIGHT HEART FAILURE DID NOT EXIST PRE-LVAD IMPLANT. A DEVICE RELATED ISSUE DID NOT CAUSE OR CONTRIBUTE TO RIGHT HEART FAILURE.Event Description: IT WAS REPORTED THAT THE PATIENT HAD A HISTORY OF ISCHEMIC CARDIOMYOPATHY (ICM). Manufacturer Narrative: H6 ADDITIONAL CLINICAL CODES: 1721 - ARRHYTHMIA, 4868 - HEMODYNAMIC INSTABILITY, 1888 - BLEEDING. SECTION E1: REPORTER EMAIL ADDRESS:(B)(6). NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: H6 ADDITIONAL CLINICAL CODES: 1721 - ARRHYTHMIA, 4868 - HEMODYNAMIC INSTABILITY, 1888 - BLEEDING. SECTION B5 CORRECTION. SECTION E3 CORRECTION. SECTION H6 HEALTH EFFECT - IMPACT CODE 4641 - UNEXPECTED MEDICAL INTERVENTION ADDED. SECTION H6 HEALTH EFFECT - IMPACT CODE 4643 - BLOOD TRANSFUSION ADDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. ADDITIONALLY, A SPECIFIC CAUSE FOR THE REPORTED INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. IT WAS REPORTED THAT HEARTMATE 3 LVAS, SERIAL NUMBER: (B)(6) WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH, OTHER NEUROLOGICAL EVENTS (NOT STROKE-RELATED), DRIVELINE INFECTION, LOCALIZED INFECTION, RENAL DYSFUNCTION, CARDIAC ARRHYTHMIA, RIGHT HEART FAILURE, AND BLEEDING. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ INCLUDES INFORMATION FOR CARING FOR THE DRIVELINE EXIT SITE AS WELL AS INFORMATION REGARDING HOW TO PREVENT AND CONTROL INFECTION. THIS SECTION (UNDER ¿RIGHT HEART FAILURE¿) ALSO OUTLINES INDICATIONS OF RIGHT HEART FAILURE AS WELL AS POSSIBLE TREATMENTS. SECTION 6 PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. A, CONTAINS SEVERAL SECTIONS WITH INFORMATION ABOUT HOW TO PREVENT AND CONTROL INFECTION AS WELL AS INFORMATION FOR CARING FOR THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22523254&pc=NPT | 2015691-2025-05795 | 2025-06-11 04:00 | Death | EDWARDS LIFESCIENCES | 2025-07-16 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Degraded | Chest Pain; Pseudoaneurysm | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST, A PATIENT WITH A 20MM S3UR AORTIC VALVE IMPLANTED FOR 1 YEAR AND 9 MONTHS EXPIRED DURING A SURGICAL PROCEDURE TO DO A HEMI ARCH AND POSSIBLE EXPLANT OF THE 20MM S3UR VALVE. PER MEDICAL RECORDS, THE PATIENT SOUGHT TREATMENT AT THE ED FOR CHEST PAIN AND WAS ADMITTED FOR ANEURYSM AORTIC WITHOUT RUPTURE, NUMEROUS PSEUDOANEURYSM, AND PROSTHETIC AORTIC VALVE STENOSIS. A CARDIAC CT DURING THIS ADMISSION REVEALED MULTILOBULATED ANEURYSM INVOLVING THE AORTIC ARCH AND DESCENDING AORTA. COMPARED TO A PRIOR CT, THERE WAS NO SIGNIFICANT CHANGE IN SIZE OR APPEARANCE OF THE PSEUDOANEURYSM AT THE AORTIC ROOT OR CHANGE IN THE SIZE OR APPEARANCE OF MULTIPLE ADDITIONAL PSEUDOANEURYSMS OF THE DISTAL ASCENDING AORTA AND AORTIC ARCH. THERE WAS A FOCAL OUTPOUCHING (14 X 16 X 9 MM), LOCATED BELOW THE RIGHT SINUS OF VALSALVA COMMUNICATING WITH THE LVOT (MYCOTIC PSEUDOANEURYSM VERSUS POSTSURGICAL). THERE WAS SURGICAL MATERIAL IN THE ADJACENT MEMBRANOUS INTERVENTRICULAR SEPTUM WITH MILDLY THICKENED LEAFLETS. THE PATIENT UNDERWENT A PROCEDURE FOR BENTALL ROOT REPLACEMENT, AORTIC ANNULUS ENLARGEMENT, RECONSTRUCTION OF THE LVOT, INTERPOSITION SVG TO LM AND RCA, CABG X 2, REPAIR OF MULTIPLE ANEURYSMS ASCENDING AORTA, HEMIARCH REPLACEMENT, AND PLACEMENT OF IABP. THERE WERE COMPLICATIONS INTRAOPERATIVELY WITH POST BYPASS VF REQUIRING CPR, IABP PLACEMENT AND REQUIRING VASOPRESSIN DRUGS. THERE WAS NO LV FUNCTION AND MINIMAL RV FUNCTION. THE HEART DID NOT RECOVER. ATTEMPTS TO WEAN OFF BYPASS WERE UNSUCCESSFUL AND THE PATIENT PASSED AWAY. Manufacturer Narrative: UPDATED SECTION H6- TYPE OF INVESTIGATION, FINDINGS, AND CONCLUSIONS. THE DEVICE WAS NOT RETURNED FOR EVALUATION AS IT REMAINED IMPLANTED. THE COMPLAINT FOR LEAFLET THICKENED/HALT WAS CONFIRMED BASED ON THE PROVIDED MEDICAL RECORDS. A REVIEW OF THE DHR, LOT HISTORY AND COMPLAINT HISTORY DID NOT PROVIDE ANY INDICATION THAT A MANUFACTURING NON-CONFORMANCE CONTRIBUTED TO THE COMPLAINT EVENT. A REVIEW OF THE IFU REVEALED NO DEFICIENCIES. DURING THE MANUFACTURING PROCESS, ALL SAPIEN 3 ULTRA RESILIA VALVES ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. HYPOATTENUATED LEAFLET THICKENING (HALT) MAY BE CAUSED BY SEVERAL PATIENT-RELATED FACTORS INCLUDING EARLY VALVE DETERIORATION (SVD) (E.G. STENOSIS/CALCIFICATION), NON-STRUCTURAL DYSFUNCTION (E.G. PANNUS), OR THROMBOSIS (FORMATION OF BLOOD CLOTS ON THE VALVE). IN THIS CASE, LIMITED INFORMATION WAS PROVIDED; THEREFORE, A DEFINITIVE ROOT CAUSE IS UNABLE TO BE DETERMINED AT THIS TIME. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. A PSEUDOANEURYSM IS A LEAKAGE OF ARTERIAL BLOOD FROM AN ARTERY INTO THE SURROUNDING TISSUE WITH A PERSISTENT COMMUNICATION BETWEEN THE ORIGINATING ARTERY AND THE RESULTANT ADJACENT CAVITY. THIS MAY OCCUR AFTER ARTERIAL PUNCTURE FOR A DIAGNOSTIC CARDIAC CATHETERIZATION OR AN ARTERIOGRAM BUT IS MORE COMMON AFTER AN ARTERIAL INTERVENTION. CATHETER-DIRECTED INTERVENTIONS MORE COMMONLY REQUIRE LARGER ARTERIAL SHEATHS TO BE USED, AND THE ANTICOAGULATION OR ANTIPLATELET AGENTS THAT ARE ADMINISTERED CAN INTERFERE WITH NORMAL SEALING OF THE PUNCTURE SITE. SOME PSEUDOANEURYSMS RESOLVE THEMSELVES, THOUGH OTHERS REQUIRE TREATMENT TO PREVENT HEMORRHAGE, AN UNCONTROLLED LEAK, OR OTHER COMPLICATIONS. SURGERY IS SOMETIMES REQUIRED. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT FACTORS SUCH AS MARFAN SYNDROME AND ENDOCARDITIS MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22523261&pc=DSQ | 2916596-2025-04677 | 2025-07-04 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND DEAD AT HOME OVER THE WEEKEND. LOG FILES WERE OBTAINED FROM THEIR SYSTEM CONTROLLER. IT APPEARED THAT THE PATIENT HAD LOW POWER FOR THEIR EXTERNAL BATTERIES AND ENDED UP WITH AN EMERGENCY BACKUP BATTERY (EBB) FAULT ALARM FOR A PERIOD OF TIME BEFORE THE DRIVELINE WAS DISCONNECTED. IT APPEARED THAT THE PATIENT ATTEMPTED A SYSTEM CONTROLLER EXCHANGE. FOLLOWING REVIEW OF THE LOG FILES BY TECH SERVICES, IT APPEARED ON (B)(6) 2025 THAT THERE WERE EBB FAULTS BEGINNING AT ABOUT 9:32 AM. THE EBB DID NOT PASS THE LOAD TEST, WHICH CAUSED THE ALARMS. THE PUMP APPEARED TO HAVE BEEN OPERATING WITHIN NORMAL PARAMETERS DURING THAT TIME FRAME. THE EBB FAULT REMAINED ACTIVE, ALONG WITH A FEW UNUSUAL POWER CABLE DISCONNECT ALARMS WHILE THE PATIENT WAS ON BATTERY POWER. THE ALARMS CONTINUED UNTIL THE DRIVELINE WAS DISCONNECTED FROM THE CONTROLLER LATER THAT SAME MORNING AT ABOUT 10:01 AM. THE CONTROLLER APPEARED TO HAVE BEEN SHUT DOWN AT ABOUT 11:30 AM ON (B)(6) 2025. PRIOR TO THE DRIVELINE DISCONNECT EVENT, THE LOG CONTAINED MULTIPLE LOW VOLTAGE AND POWER CABLE DISCONNECT ALARMS WHILE ON BATTERY POWER. THE ALARMS APPEARED TO HAVE BEEN THE RESULT OF RELATIVE STATE OF CHARGE (RSOC) BATTERY DATA INVALID FLAGS. ADDITIONAL INFORMATION RECEIVED ON 08JUL2025 REPORTED NO PLAN TO SEND EQUIPMENT BACK Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22523265&pc=NIK | 2182208-2025-03487 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-07-16 04:00 | NIK | MEDTRONIC CRT-D | Appropriate Term/Code Not Available | Cardiac Arrest |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING EFFECTIVE CARDIAC RESYNCHRONIZATION THERAPY (CRT) RESPONSE. THE AUTHORS DESCRIBED PATIENT DEATHS IN BOTH GROUPS OF PACING PERCENTAGE CRT RESPONSE; THERE WERE TWO PATIENTS IN EACH GROUP THAT DIED DUE TO SUDDEN CARDIAC DEATH AND OTHER UNKNOWN CAUSES OF DEATH. THERE WERE PATIENTS WHO EXPERIENCED HEART FAILURE HOSPITALIZATIONS. THERE WERE LEFT VENTRICULAR (LV) LEADS WHICH EXHIBITED LOSS OF CAPTURE WITH ONE DISLODGEMENT WHICH WAS REIMPLANTED, AND THE OTHER LEADS WERE REPROGRAMMED. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: PRESENCE OF INEFFECTIVE CARDIAC RESYNCHRONIZATION THERAPY PACING PROVIDES INSIGHTS INTO HIDDEN CAUSES AND THERAPEUTIC TARGETS OF NONRESPONDER. JOURNAL OF CARDIOVASCULAR ELECTROPHYSIOLOGY. 2025; 36:925¿934. DOI: 10.1111/JCE.16605 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22523278&pc=DXY | 2182208-2025-03488 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-07-16 04:00 | DXY | MEDTRONIC LEAD | Appropriate Term/Code Not Available | Cardiac Arrest |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING EFFECTIVE CARDIAC RESYNCHRONIZATION THERAPY (CRT) RESPONSE. THE AUTHORS DESCRIBED PATIENT DEATHS IN BOTH GROUPS OF PACING PERCENTAGE CRT RESPONSE; THERE WERE TWO PATIENTS IN EACH GROUP THAT DIED DUE TO SUDDEN CARDIAC DEATH AND OTHER UNKNOWN CAUSES OF DEATH. THERE WERE PATIENTS WHO EXPERIENCED HEART FAILURE HOSPITALIZATIONS. THERE WERE LEFT VENTRICULAR (LV) LEADS WHICH EXHIBITED LOSS OF CAPTURE WITH ONE DISLODGEMENT WHICH WAS REIMPLANTED, AND THE OTHER LEADS WERE REPROGRAMMED. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: PRESENCE OF INEFFECTIVE CARDIAC RESYNCHRONIZATION THERAPY PACING PROVIDES INSIGHTS INTO HIDDEN CAUSES AND THERAPEUTIC TARGETS OF NONRESPONDER. JOURNAL OF CARDIOVASCULAR ELECTROPHYSIOLOGY. 2025; 36:925¿934. DOI: 10.1111/JCE.16605 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22523289&pc=DSQ | 3007042319-2025-01465 | 2025-07-06 04:00 | Death | HEARTWARE, INC. | 2025-07-16 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A VENTRICULAR ASSIST DEVICE (VAD) PATIENT WAS FOUND DOWN AT HOME AND LATER PRONOUNCED DEAD. THE PATIENT WAS LAST SEEN 20 MINUTES BEFORE BEING FOUND DOWN. UPON ARRIVAL OF EMERGENCY MEDICAL SERVICES (EMS), THE PATIENT'S PUPILS WERE FIXED AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED BY EMS. A LOW FLOW ALARM WAS NOTED ON THE VENTRICULAR ASSIST DEVICE, AND EMS PERFORMED A CONTROLLER EXCHANGE, WHICH WAS NOT INDICATED AS THE CORRECT REASON FOR EXCHANGE. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM CONTROLLER 2.0 D4: MODEL #: 1420 / CATALOG #: 1420 / EXPIRATION DATE: UNKNOWN / SERIAL OR LOT#: UNKNOWN D9: NO H3: NO H4: MFG DATE: UNKNOWN H5: NO H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22523349&pc=DSQ | 2916596-2025-04678 | 2025-07-04 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN INTRACRANIAL HEMORRHAGE WITHDRAWAL FROM SUPPORT. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: SECTION E1: REPORTER PHONE NUMBER (B)(6). NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22523527&pc=OZD | 1220648-2025-30406 | 2025-07-02 04:00 | Death | ABIOMED, INC. | 2025-07-16 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED VASCULAR COMPLICATIONS THAT RESULTED IN HEMORRHAGE, CARDIAC ARREST, AND DEATH. Manufacturer Narrative: A4 AND A5 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22523560&pc=DTB | 2649622-2025-19288 | 2025-05-03 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-07-16 04:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Over-Sensing | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) TRIGGERED A FALSE OBSERVATION FOR AN INCREASE IN THE LEFT VENTRICULAR (LV) LEAD THRESHOLD INDICATING THAT THE INCREASE WAS GREATER THAN THE AMPLITUDE SAFETY MARGIN AND MAY HAVE COMPROMISED CAPTURE. THE OBSERVATION DISPLAYED QUESTION MARKS WHERE A NUMERICAL VALUE WOULD TYPICALLY BE DISPLAYED. ADDITIONALLY, THE RIGHT ATRIAL (RA) LEAD EXHIBITED INTERMITTENT FAR FIELD R-WAVE (FFRW) OVERSENSING THAT MOST LIKELY CAUSED FALSE ATRIAL TACHYCARDIA/ATRIAL FIBRILLATION (AT/AF) DETECTION. THE PATIENT LATER PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: DTMA1QQ CRT-D IMPLANTED: (B)(6) 2021. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22523561&pc=DSQ | 2916596-2025-04695 | 2025-07-09 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Disconnection; Infusion or Flow Problem | Loss of consciousness | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS TRANSITIONING TO HOSPICE WHEN THEY WERE FOUND UNRESPONSIVE BY FAMILY AND PRESENTED TO THE EMERGENCY DEPARTMENT. THE PATIENT PASSED AWAY IN THE EMERGENCY DEPARTMENT. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. IT IS UNKNOWN WHETHER THE DEVICE WILL BE EXPLANTED OR RETURNED FOR EVALUATION. LOG FILES WERE REVIEWED, AND THEY CAPTURED SOME PERIODS OF LOW FLOW EVENTS LATE EVENING OF (B)(6) 2025 @21:29 AND CONTINUED INTERMITTENTLY UNTIL THE SYSTEM CONTROLLER WAS DISCONNECTED ON (B)(6) 2025 @22:54. FLOW WAS NOTED AS LOW AS 0.4 LITERS PER MINUTE (LPM) AT TIMES, AND PULSATILITY INDEX (PI) VALUES WERE VARIABLE AND ON THE HIGH SIDE DURING THESE LOW FLOW PERIODS. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW HAZARD ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THE ALARMS COULD NOT BE CONCLUSIVELY DETERMINED. THE CONTROLLER EVENT LOG FILES CONTAINED EVENTS FROM (B)(6) 2025 THROUGH (B)(6) 2025. LOW FLOW HAZARD ALARMS WERE CAPTURED THROUGHOUT THE FILE, WHICH WERE ASSOCIATED WITH ELEVATED PUISATILITY INDEX VALUES. THE POWER CABLES WERE BOTH DISCONNECTED ON (B)(6) 2025 BY 22:54:46, AND THE DRIVELINE WAS DISCONNECTED AT 22:54:56, CONSISTENT WITH THE REPORTED TRANSITION TO HOSPICE AND WITHDRAWAL OF LEFT VENTRICULAR ASSIST DEVICE SUPPORT. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENTS AND PRODUCT RETURN; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE. THIS SECTION ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿HEARTMATE TOUCH COMMUNICATION SYSTEM¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LPM, A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. THIS SECTION ALSO OUTLINES SITUATIONS THAT CAN RESULT IN A LOW FLOW ALARM AND FURTHER EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING," AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿ OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22523686&pc=MHX | 9610816-2025-000621 | 2025-06-05 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-07-16 04:00 | MHX | PATIENT INFORMATION CENTER IX | Use of Device Problem | Insufficient Information | K211900 |  | 1 | Event Description: IT WAS REPORTED THE USER DID NOT HEAR OR RESPOND TO THE ALARMS AND THE PATIENT PASSED WAY; HOWEVER, THERE WAS NO ALLEGATION OF ALARM FAILURE. IT WAS INDICATED THE USER DID NOT HEAR THE ECG LEADS OFF ALARMS AND DID NOT RESPOND TO THE ECG LEADS OFF YELLOW TECHNICAL INOP ALARMS. THE CUSTOMER REQUESTED ASSISTANCE OBTAINING THE UNFILTERED AUDIT LOGS, WHICH REVEALED THERE WERE TWO HOURS OF ECG LEADS OFF ALARMS WITH NO ALARM ACKNOWLEDGEMENT AND THE ALARMS WERE PAGED APPROPRIATELY. RESUSCITATION WAS ATTEMPTED BUT UNSUCCESSFUL AND THE PATIENT PASSED AWAY. THE NURSE MANAGER INDICATED THE ALARM VOLUME FOR YELLOW ALARMS IS SET LOWER THAN RED ALARMS AND THE STAFF IS NOT RESPONDING TO THE YELLOW ALARMS. THE CUSTOMER WILL BE REQUESTING THE ECG LEADS OFF TECHNICAL INOP BE CHANGED TO A RED ALARM VIA THE INTERNAL ALARM COUNCIL. Manufacturer Narrative: THE CLINICAL AUDIT LOGS WERE PULLED FOR INVESTIGATION. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22523698&pc=NPT | 2025587-2025-05076 | 2025-06-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-16 04:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Insufficient Information; Patient Device Interaction Problem; Thickening of Material; Device Stenosis; Central Regurgitation | Heart Failure/Congestive Heart Failure; Aortic Valve Insufficiency/ Regurgitation; Mitral Valve Insufficiency/ Regurgitation; Unspecified Gastrointestinal Problem; No Clinical Signs, Symptoms or Conditions | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 YEARS AND 8 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED THAT REVEALED THE VALVE FAILED DUE TO AN UNKNOWN REASON. ADDITIONALLY, THICKENING OF THE VALVE LEAFLETS WAS OBSERVED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 YEARS AND 8 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED THAT REVEALED THE VALVE FAILED DUE TO AN UNKNOWN REASON. ADDITIONALLY, THICKENING OF THE VALVE LEAFLETS WAS OBSERVED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE PATIENT HAD BEEN ADMITTED DUE TO GASTROINTESTINAL BLEEDING, HEART FAILURE, MITRAL INSUFFICIENCY, AND MITRAL STENOSIS. THE FAILURE OF THE VALVE WAS NOTED TO BE MODERATE CENTRAL AORTIC REGURGITATION AND AORTIC STENOSIS. THE PATIENT PASSED AWAY PRIOR TO INTERVENTION. Manufacturer Narrative: UPDATED DATA: B1. B2. B5. ADDED SECOND PARAGRAPH. H1. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22523787&pc=DQY | 2124215-2025-47036 | 2025-07-09 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-16 04:00 | DQY | WATCHMAN TRUSTEER? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | K240018 |  | 1 | Event Description: IT WAS REPORTED THAT PERICARDIAL EFFUSION (PE), CARDIAC TAMPONADE, AND DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING NOTED THERE WAS NO BASELINE PE. GROIN ACCESS OBTAINED AND NON-BOSTON SCIENTIFIC VASCULAR PRE-CLOSURE DEVICE WAS USED. A VERSACROSS RADIOFREQUENCY (RF) WIRE WAS INSERTED UP TO THE SUPERIOR VENA CAVA (SVC), A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND VERSACROSS CONNECT DILATOR TO FOLLOW. NO TRANSSEPTAL PUNCTURE (TSP) WAS ACHIEVED ON THE FIRST DROP DOWN. ON THE SECOND DROPDOWN TSP WAS PERFORMED WITHOUT RF DELIVERY REQUIRED. TSP LOCATION APPEARED TO BE MID/MID-POSTERIOR BASED ON ICE IMAGING. NO PE WAS NOTED AT THIS POINT IN THE PROCEDURE AND PATIENT VITALS WERE STABLE. TSP LOCATION WAS DILATED WITH THE WAS AND THEN WAS AND VCC DILATOR RETRACTED FOR LEFT ATRIUM (LA) ACCESS OF THE ICE CATHETER. TSP ACCESS OF ICE CATHETER WAS STILL UNSUCCESSFUL. A PE WAS THEN DISCOVERED FOUR (4) MINUTES LATER. WAS REMOVED FROM THE GROIN AND ICE VISUALIZED THE PE LOCATION ON THE RIGHT SIDE OF THE HEART. THE PROCEDURE WAS ABORTED. CARDIAC TAMPONADE OCCURRED, WHICH CAUSED THE NEED FOR CARDIOPULMONARY RESUSCITATION (CPR), BLOOD TRANSFUSION, PERICARDIOCENTESIS REMOVING BRIGHT RED BLOOD OF OVER THREE (3) LITERS, FRESH FROZEN PLASMA (FFP), AND CRYOPRECIPITATE. THE FAMILY WAS CONTACTED A LITTLE OVER TWO (2) HOURS LATER AND AS THE PATIENT WAS A DO NOT RESUSCITATE (DNR) STATUS AND THEY REQUESTED NO ADDITIONAL LIFE-SAVING MEASURES, SO TIME OF DEATH WAS CALLED BY THE PHYSICIAN. THE OFFICIAL CAUSE OF DEATH IS PE THAT RESULTED IN DEATH. IT WAS NOTED THE PATIENT WAS ON ELIQUIS AND STOPPED TAKING ELIQUIS TWO (2) DAYS PRIOR TO THE PROCEDURE. LAST RECORDED INR LAB RESULT WAS 1.4.Event Description: IT WAS REPORTED THAT PERICARDIAL EFFUSION (PE), CARDIAC TAMPONADE, AND DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING NOTED THERE WAS NO BASELINE PE. GROIN ACCESS OBTAINED AND NON-BOSTON SCIENTIFIC VASCULAR PRE-CLOSURE DEVICE WAS USED. A VERSACROSS RADIOFREQUENCY (RF) WIRE WAS INSERTED UP TO THE SUPERIOR VENA CAVA (SVC), A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND VERSACROSS CONNECT DILATOR TO FOLLOW. NO TRANSSEPTAL PUNCTURE (TSP) WAS ACHIEVED ON THE FIRST DROP DOWN. ON THE SECOND DROPDOWN TSP WAS PERFORMED WITHOUT RF DELIVERY REQUIRED. TSP LOCATION APPEARED TO BE MID/MID-POSTERIOR BASED ON ICE IMAGING. NO PE WAS NOTED AT THIS POINT IN THE PROCEDURE AND PATIENT VITALS WERE STABLE. TSP LOCATION WAS DILATED WITH THE WAS AND THEN WAS AND VCC DILATOR RETRACTED FOR LEFT ATRIUM (LA) ACCESS OF THE ICE CATHETER. TSP ACCESS OF ICE CATHETER WAS STILL UNSUCCESSFUL. A PE WAS THEN DISCOVERED FOUR (4) MINUTES LATER. WAS REMOVED FROM THE GROIN AND ICE VISUALIZED THE PE LOCATION ON THE RIGHT SIDE OF THE HEART. THE PROCEDURE WAS ABORTED. CARDIAC TAMPONADE OCCURRED, WHICH CAUSED THE NEED FOR CARDIOPULMONARY RESUSCITATION (CPR), BLOOD TRANSFUSION, PERICARDIOCENTESIS REMOVING BRIGHT RED BLOOD OF OVER THREE (3) LITERS, FRESH FROZEN PLASMA (FFP), AND CRYOPRECIPITATE. THE FAMILY WAS CONTACTED A LITTLE OVER TWO (2) HOURS LATER AND AS THE PATIENT WAS A DO NOT RESUSCITATE (DNR) STATUS AND THEY REQUESTED NO ADDITIONAL LIFE-SAVING MEASURES, SO TIME OF DEATH WAS CALLED BY THE PHYSICIAN. THE OFFICIAL CAUSE OF DEATH IS PE THAT RESULTED IN DEATH. IT WAS NOTED THE PATIENT WAS ON ELIQUIS AND STOPPED TAKING ELIQUIS TWO (2) DAYS PRIOR TO THE PROCEDURE. LAST RECORDED INR LAB RESULT WAS 1.4. IT WAS FURTHER REPORTED A CARDIAC PERFORATION WAS NOT VISUALIZED, BUT A RIGHT SIDED PERFORATION WAS SUSPECTED BY THE PHYSICIANS. Manufacturer Narrative: B6 INFORMATION ADDED. H6 PATIENT CODE ADDED: CARDIAC PERFORATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22523791&pc=MCX | 3004742232-2025-00204 | 2025-06-27 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-07-16 04:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Perforation of Vessels | P130005 |  | 1 | Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A DISTAL LEFT MAIN ARTERY (LM) AND LEFT ANTERIOR DESCENDING ARTERY (LAD) WITH MULTIPLE CALCIUM NODULES. MULTIPLE LOW SPEED TREATMENTS WERE PERFORMED. ANGIOGRAPHIC IMAGING SHOWED GOOD RESULTS. DURING THE NEXT TREATMENT IN THE PROXIMAL LAD, A PERFORATION WAS OBSERVED AND THE PATIENT'S BLOOD PRESSURE DROPPED. A GRAFTMASTER WAS USED IN AN ATTEMPT TO MANAGE THE PERFORATION, HOWEVER, THE PATIENT EXPIRED. IN THE PHYSICIAN'S OPINION, ORBITAL ATHERECTOMY CAUSED THE PERFORATION. IN THE PHYSICIAN'S OPINION, THE CAUSE OF THE PATIENT DEATH WAS DUE TO SUDDEN DROP IN BLOOD PRESSURE, PERFORATION, AND UNABLE TO MANAGE THE PERFORATION IN TIME WITH THE GRAFTMASTER. IT WAS NOTED THE GRAFTMASTER STENT DID SEAL THE AREA IT WAS IMPLANTED; HOWEVER, THE PERFORATION WAS LARGER THAN EXPECTED.Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A DISTAL LEFT MAIN ARTERY (LM) AND LEFT ANTERIOR DESCENDING ARTERY (LAD) WITH MULTIPLE CALCIUM NODULES. MULTIPLE LOW SPEED TREATMENTS WERE PERFORMED. ANGIOGRAPHIC IMAGING SHOWED GOOD RESULTS. DURING THE NEXT TREATMENT IN THE PROXIMAL LAD, A PERFORATION WAS OBSERVED AND THE PATIENT'S BLOOD PRESSURE DROPPED. A GRAFTMASTER WAS USED IN AN ATTEMPT TO MANAGE THE PERFORATION, HOWEVER, THE PATIENT EXPIRED. NO ADDITIONAL INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING IN THE PHYSICIAN'S OPINION, ORBITAL ATHERECTOMY CAUSED THE PERFORATION. IN THE PHYSICIAN'S OPINION, THE CAUSE OF THE PATIENT DEATH WAS DUE TO SUDDEN DROP IN BLOOD PRESSURE, PERFORATION, AND UNABLE TO MANAGE THE PERFORATION IN TIME WITH THE GRAFTMASTER. IT WAS NOTED THE GRAFTMASTER STENT DID SEAL THE AREA IT WAS IMPLANTED; HOWEVER, THE PERFORATION AS LARGER THAN EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED PATIENT EFFECTS OF PERFORATION OF VESSELS, LOW BLOOD PRESSURE, DELAY TO TREATMENT, UNEXPECTED MEDICAL INTERVENTION AND DEATH APPEAR TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE IT IS POSSIBLE THAT THE REPORTED PATIENT EFFECTS OF PERFORATION OF VESSELS, LOW BLOOD PRESSURE, DELAY TO TREATMENT, UNEXPECTED MEDICAL INTERVENTION AND DEATH ARE A RESULT OF THE PATIENT¿S DISEASE SEVERITY AND/OR USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. UPDATED: B5, H6 (CODES 4118, 3233, AND 11 NO LONGER APPLY).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22524204&pc=BZD | 2518422-2025-107776 | 2025-07-02 04:00 | Death | RESPIRONICS, INC. | 2025-07-16 04:00 | BZD | DREAMSTATION 2 ADVANCED AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: A RELATIVE OF THE PATIENT INQUIRED ABOUT THE OWNER OF THE DEVICE, A DREAMSTATION 2 ADVANCED AUTO CPAP. THE RELATIVE OF THE PATIENT REPORTED THE PATIENT PASSED AWAY. THE REPORTER STATED THE PATIENT WAS NOT ON THE DEVICE AT THE TIME OF DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22524390&pc=DSQ | 2916596-2025-04755 | 2025-06-25 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT WAS LOST TO FOLLOW UP. BASED ON A REVIEW OF THE AVAILABLE PATIENT RECORDS AND THE REPORTED PATIENT OUTCOME THERE WERE NO DEVICE ISSUES AT THE TIME OF PATIENT DEATH.Event Description: THE CAUSE OF DEATH WAS EXHAUSTED PHYSIOLOGICAL RESERVES OF THE PATIENT. THE PATIENT IN PALLIATIVE CARE. THE DEVICE OPERATED AS EXPECTED. NO PRODUCTS WOULD BE RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THERE WERE NO DEVICE RELATED ISSUES WITH HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL (B)(6), REPORTED OR IDENTIFIED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. B, IS ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22524469&pc=DSQ | 2916596-2025-04744 | 2025-04-21 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS CARDIAC ARREST. SUPPORT WAS WITHDRAWN BECAUSE THE PATIENT HAD "MSIF" AND DEVELOPED AN ACUTE HEPATIC ARTERY BLEED, WHICH WAS ASSOCIATED WITH THE CARDIAC ARREST. THE DEVICE FUNCTIONED AS INTENDED. THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED AND AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS REPORTED THAT SUPPORT WAS WITHDRAWN BECAUSE THE PATIENT HAD MULTI-SYSTEM ORGAN FAILURE (MSOF) AND DEVELOPED AN ACUTE HEPATIC ARTERY BLEED, WHICH WAS ASSOCIATED WITH THE CARDIAC ARREST. THE ACUTE HEPATIC ARTERY BLEED WAS IDENTIFIED USING A COMPUTED TOMOGRAPHY (CT) ANGIOGRAPHY OF THE ABDOMEN/PELVIS. BLEEDING WAS CONFIRMED AND SUCCESSFULLY TREATED WITH INTERVENTIONAL RADIOLOGY (IR) EMBOLIZATION OF THE LEFT HEPATIC ARTERY ON (B)(6) 2025.Event Description: IT WAS LATER REPORTED THAT THE MSOF BEGAN ON (B)(6) 2025 AND WAS RELATED TO THE DEVICE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS DEATH AND BLEEDING AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RIGHT HEART FAILURE, RESPIRATORY FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22524937&pc=DSQ | 2916596-2025-04647 | 2025-03-18 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENTS; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CUSTOMER. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU, REV. A IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22525568&pc=DSQ | 2916596-2025-04703 | 2025-07-07 04:00 | Death | THORATEC CORPORATION | 2025-07-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. DUE TO HOSPITAL POLICY, NO FURTHER INFORMATION WAS PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. IT WAS UNKNOWN WHETHER THE DEVICE WOULD BE EXPLANTED, AN AUTOPSY WOULD BE PERFORMED, OR IF IT WOULD BE RETURNED FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER: (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22526837&pc=BYJ | 3004972304-2025-00004 | 2025-06-29 04:00 | Death | CAIRE INC. | 2025-07-16 04:00 | BYJ | LIBERATOR | Insufficient Information | Insufficient Information | K800742 |  | 1 | Event Description: AS REPORTED: AN ADDITIONAL OXYGEN TANK WAS DELIVERED ON (B)(6) 2025 TO A PATIENT'S HOME DUE TO INCREASED OXYGEN NEEDS. THE TANK WAS FIRST CONNECTED AND USED ON FRIDAY, ON (B)(6) 2025. IT FUNCTIONED NORMALLY: THE LED INDICATORS WERE FULLY GREEN, AND OXYGEN WAS BEING DELIVERED PROPERLY. ON SATURDAY, ON (B)(6) 2025, AROUND 11:00 P.M., A ROUTINE CHECK WAS PERFORMED. ALL CONNECTIONS WERE IN ORDER, THE LED INDICATORS REMAINED GREEN, AND OXYGEN APPEARED TO BE AVAILABLE IN SUFFICIENT QUANTITY. THE PATIENT WAS SLEEPING IN A HOSPITAL BED IN THE LIVING ROOM, WEARING A NASAL CANNULA. OTHER HOUSEHOLD MEMBERS WERE ASLEEP IN THEIR RESPECTIVE ROOMS. ON THE MORNING OF SUNDAY, ON (B)(6) 2025, AT APPROXIMATELY 6:40 A.M., THE PATIENT WAS FOUND DECEASED IN BED. THE NASAL CANNULA WAS STILL IN PLACE AND CORRECTLY CONNECTED TO THE TANK, BUT NO OXYGEN WAS FLOWING, DESPITE THE LED INDICATORS STILL SHOWING GREEN. ANOTHER TANK WAS IMMEDIATELY CONNECTED, AND OXYGEN FLOWED NORMALLY. CPR WAS ATTEMPTED BUT WAS UNSUCCESSFUL. THE PATIENT'S MOBILE PHONE, USUALLY KEPT WITHIN REACH TO CALL FOR HELP, WAS FOUND ON THE FLOOR, SUGGESTING AN ATTEMPT TO CALL DURING THE NIGHT. ON (B)(6) 2025, A CALL WAS MADE TO THE SUPPLIER. AS IT WAS A SUNDAY, A TECHNICIAN ANSWERED AND GAVE A CONFUSED EXPLANATION, MENTIONING THAT AN ALARM SHOULD HAVE SOUNDED IN CASE OF OXYGEN FAILURE. HOWEVER, THE TECHNICIAN WAS REFERRING TO A DIFFERENT DEVICE. THE CALLER WAS ADVISED TO CONTACT THE MAIN OFFICE THE NEXT MORNING. ON 30/06/2025, THE MAIN OFFICE WAS CONTACTED. IT WAS CONFIRMED THAT THE MODEL IN QUESTION, THE LIBERATOR, DOES NOT HAVE A BUILT-IN ALARM SYSTEM, WHICH THE FAMILY WAS ALREADY AWARE OF AFTER SEVERAL YEARS OF USE. THAT SAME DAY, THE PATIENT'S GENERAL PRACTITIONER INITIATED A JUDICIAL INVESTIGATION. THE PUBLIC PROSECUTOR'S OFFICE WAS NOTIFIED. THE OXYGEN TANKS WERE SEALED ON-SITE BY THE POLICE AND PLACED IN A SEPARATE ROOM IN THE HOME. THE OFFICERS ATTEMPTED TO WEIGH THE TANKS USING A HOUSEHOLD SCALE, WHICH WAS NOT SUITABLE FOR SUCH HEAVY EQUIPMENT. ONE OF THE OFFICERS EXPRESSED DOUBT ABOUT THE RELIABILITY OF THIS METHOD, BUT THE WEIGHING WAS CARRIED OUT FOR REPORTING PURPOSES. THE TANKS REMAIN SEALED AT THE RESIDENCE, WITH A POLICE NOTICE ATTACHED INDICATING A "SUSPICIOUS DEATH." AN AUTOPSY WAS PERFORMED ON (B)(6) 2025 AT (B)(6) HOSPITAL. THE CASE IS REGISTERED UNDER A PV NUMBER WITH THE PUBLIC PROSECUTOR'S OFFICE, UNDER THE SUPERVISION OF A DEPUTY PUBLIC PROSECUTOR. CONTACT DETAILS ARE AVAILABLE UPON REQUEST FOR OFFICIAL FOLLOW-UP. THE INCIDENT WAS ALSO REPORTED TO THE FEDERAL PUBLIC SERVICE (FPS) HEALTH VIA EMAIL ON 08/07/2025. A RESPONSE WAS RECEIVED FROM THE DEPUTY PROSECUTOR, BUT IT IS RECOMMENDED TO CONTACT HER DIRECTLY FOR FURTHER INFORMATION. IT IS UNDERSTOOD THAT THE PUBLIC PROSECUTOR'S OFFICE WILL NOT CONDUCT A TECHNICAL INVESTIGATION. THEREFORE, AN INDEPENDENT TECHNICAL ANALYSIS IS NECESSARY TO DETERMINE WHETHER A MALFUNCTION OCCURRED THAT CAUSED AN INTERRUPTION IN OXYGEN DELIVERY, DESPITE THE LED INDICATORS SHOWING NORMAL STATUS. SUCH AN INVESTIGATION IS CRUCIAL TO PREVENT SIMILAR INCIDENTS IN THE FUTURE. THE SEALED TANKS REMAIN IN THE HOME, WHICH IS A DAILY EMOTIONAL BURDEN FOR THE FAMILY. A THOROUGH TECHNICAL EXAMINATION OF THE TANK IS STRONGLY REQUESTED. Manufacturer Narrative: PURSUANT TO TITLE 21 - FOOD AND DRUGS, CHAPTER I - FOOD AND DRUG ADMINISTRATION DEPARTMENT OF HEALTH AND HUMAN SERVICES, SUBCHAPTER H -0 MEDICAL DEVICE, PART 803 - MEDICAL DEVICE REPORTING, SUBPART A - GENERAL PROVISIONS, SECTION 803.16, NEITHER THIS REPORT NOR ANY INFORMATION SUBMITTED HEREIN CONSTITUTES AN ADMISSION BY CAIRE INC. THAT THE DEVICE STATED IN THIS REPORT, CAIRE INC., OR CAIRE INC.'S EMPLOYEES, CAUSED OR CONTRIBUTED TO THE REPORTABLE EVENT STATED HEREIN. CAIRE CUSTOMER SERVICE CONTACTED THE DISTRIBUTOR, OXYSPHAIR, TO GATHER ADDITIONAL INFORMATION REGARDING ALL CAIRE DEVICES IN THE PATIENT'S HOME. OXYSPHAIR AND CAIRE WILL REQUEST THAT THE DEVICE BE RELEASED FROM POLICE CUSTODY AND SENT TO CAIRE'S LANGENFELD, GERMANY FACILITY FOR EVALUATION. CAIRE WILL CONDUCT AN EVALUATION OF THE DEVICE, IF POSSIBLE. A FOLLOW UP AND FINAL REPORT WILL BE SUBMITTED. NOTE: THE INTENDED USE OF LIBERATOR LIQUID OXYGEN BASE IS SUPPLEMENTAL OXYGEN TO INCREASE BLOOD OXYGEN SATURATION. IT IS NOT INDICATED FOR LIFE SUSTAINING OR LIFE SUPPORT ACTIVITIES. LIBERATOR DOES NOT HAVE PATIENT SURVEILLANCE CAPABILITIES. PER THE INITIAL REPORT, "THE PATIENT'S MOBILE PHONE, USUALLY KEPT WITHIN REACH TO CALL FOR HELP, WAS FOUND ON THE FLOOR, SUGGESTING AN ATTEMPT TO CALL DURING THE NIGHT." A MOBILE PHONE IS NOT AN ADEQUATE PATIENT SURVEILLANCE SYSTEM. SECTION D4 PRIMARY UNIQUE DEVICE IDENTIFIER (UDIS) # IS INTENTIONALLY BLANK PER RULE 21 CFR 801.30(8). THE DEVICE WAS INTENDED FOR EXPORT ONLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22673850&pc=OZD | 1220648-2025-30876 | 2025-07-17 04:00 | Death | ABIOMED, INC. | 2025-07-31 04:00 | OZD | IMPELLA 5.5 | Positioning Failure; Pumping Stopped; Adverse Event Without Identified Device or Use Problem | Stroke/CVA; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THROUGH COMMUNICATION WITH THE ABIOMED FIELD REPRESENTATIVE, IT WAS DISCOVERED THAT A DIFFERENT COMPLAINT EVENT OCCURRED WITH THIS DEVICE (PUMP SN (B)(6)). ACCORDING TO THE FIELD REPRESENTATIVE, "THE PATIENT HAD A STROKE AFTER THE SECOND PUMP WAS PUT IN." THE PUMP STOP EVENT IS ASSOCIATED WITH THE FIRST PUMP THAT WAS IMPLANTED. PLEASE SEE MANUFACTURER REPORT NUMBER FOR IMPELLA 5.5 PUMP ONE 1220648-2025-45710. UPON IMPLANTING THE SECOND PUMP, THE OLD GRAFT (WHICH WAS ON THE INNOMINATE ARTERY), WAS REMOVED AND A NEW ONE WAS PLACED IN THE SAME SPOT. WITHIN 12 HOURS, THE PATIENT HAD A STROKE, AND THUS, WAS REMOVED FROM THE TRANSPLANT LIST. THE PATIENT WAS WITHDRAWN FROM CARE AND FOUR DAYS LATER EXPIRED.Event Description: THE USER FACILITY REPORTED AN IMPELLA 5.5 HAD POSITIONAL ISSUES. WHILE REPOSITIONING, THE PUMP'S CABLE WAS UNPLUGGED TO RELIEVE THE TORQUE. WHEN PLUGGED BACK IN THE PUMP FAILED TO RE-START. THE TEAM MADE THE DECISION TO REMOVE AND REPLACE THE PUMP. THE PATIENTS OUTCOME IS CRITICAL. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: DIRECT AORTIC INSERTION ¿MAINTAINING ACT AT OR ABOVE 250 SECONDS WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: WARNINGS, CAUTIONS, AND PRECAUTIONS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ THE FOLLOWING FIELDS WERE CORRECTED BASED ON THE ADDITIONAL INFORMATION RECEIVED BY THE FIELD REPRESENTATIVE, THAT THE INITIAL MANUFACTURER DEVICE REPORT 1220648-2025-30876 REPORTED A MALFUNCTION THAT SHOULD HAVE BEEN AGAINST THE FIRST PUMP IMPLANTED, WHICH HAS BEEN REPORTED IN MANUFACTURER DEVICE REPORT 1220648-2025-45710. THIS REPORT IS FOR THE SECOND PUMP THAT WAS IMPLANTED AND THE PATIENT EXPERIENCED A STROKE AND EXPIRED. B.1 PRODUCT PROBLEM WAS ERRONEOUSLY SELECTED. B.2 DEATH AND DATE OF DEATH WAS INADVERTENTLY OMITTED. B.5 CORRECTIVE STATEMENT AND PATIENT NARRATIVE WERE UPDATED TO REFLECT THE CORRECT INFORMATION. H.1 TYPE OF REPORT WAS ERRONEOUSLY SELECTED. H.6 HEALTH EFFECT - CLINICAL CODE, HEALTH EFFECT - IMPACT CODE, AND MEDICAL DEVICE PROBLEM CODE WERE ALL ERRONEOUSLY ENTERED. H.10 INSTRUCTIONS FOR USE WERE UPDATED TO REFLECT THE NEW INFORMATION PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22673900&pc=MVK | 3008642652-2025-07347 | 2025-07-03 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-07-31 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON 7/3/2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED EIGHT APPROPRIATE TREATMENTS AND FOUR INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 10:40:17 ON (B)(6) 2025. AT 03:45:42 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS IDIOVENTRICULAR RHYTHM @ 80 BPM DEGRADING TO VT @ 210 BPM DEGRADING TO VF. AT 03:46:17, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 290 BPM. POST SHOCK RHYTHM WAS ONE SINUS BEAT DEGRADING TO VT @ 250 BPM. AT 03:46:44, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 270 BPM. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 70 BPM DEGRADING TO VT @ 250 BPM. AT 03:47:18, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 60 BPM. AT 03:48:07, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 200 BPM. AT 03:48:37, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 180 BPM. POST SHOCK RHYTHM WAS SINUS RHYTHM/IDIOVENTRICULAR RHYTHM @ 70 BPM DEGRADING TO VT @ 200 BPM. AT 03:49:03, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 70 BPM DEGRADING TO VT @ 200 BPM. AT 03:49:30, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS VT @ 280 BPM. AT 03:50:00, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 260 BPM. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 70 BPM DEGRADING TO VT @ 170 BPM. AT 03:50:27, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS VT @ 170 BPM. AT 03:53:04, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 03:53:34, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH PVC¿S TRANSITIONING TO AN IDIOVENTRICULAR RHYTHM @ 80 BPM DEGRADING TO VT @ 190 BPM. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 03:54:06, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 03:54:51, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY/CPR/MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 03:59:02, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 03:59:32, THE PATIENT RECEIVED THE SEVENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH AN IDIOVENTRICULAR RHYTHM BEAT DEGRADING TO VF. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 04:00:03, THE PATIENT RECEIVED THE EIGHTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY/CPR/MOTION ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE ELECTRODE BELT WAS DISCONNECTED AT 04:19:10 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22674038&pc=DSQ | 2916596-2025-05078 | 2025-07-19 04:00 | Death | THORATEC CORPORATION | 2025-07-31 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Stroke/CVA; Hemorrhage/Bleeding; Muscle Weakness; Thrombosis/Thrombus | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A MIDDLE CEREBRAL ARTERY STROKE. THE PATIENT WAS ANTICOAGULATED WITH INTERNATIONAL NORMALIZED RATIO OF 3.0 UPON EMERGENCY DEPARTMENT ARRIVAL WITH NEW ONSET STROKE SYMPTOMS ON (B)(6) 2024. THE PATIENT HAD SYMPTOMS OF LEFT SIDED FACIAL DROOP, LEFT ARM FLACCID, AND THE INABILITY TO MOVE LEFT LEG. LOG FILES WERE SUBMITTED FOR REVIEW AND CAPTURED CLUSTERS OF PULSATILITY INDEX (PI) EVENTS AS WELL AS ROUTINE POWER SOURCE CHANGES. NO UNUSUAL ROTOR FAULTS, PUMP TEMPERATURE, OR ANY OTHER ABNORMAL PUMP FAULTS WERE SEEN IN THE LOG. THE PATIENT WAS BROUGHT TO INTERVENTIONAL RADIOLOGY FOR MECHANICAL THROMBECTOMY WITH ¿PARTIAL RESTORATION OF FLOW¿. ADDITIONAL INFORMATION RECEIVED STATED THAT LEFT LEG DEFICIT OCCURRED DURING EVENT AND IMPROVED POST MECHANICAL THROMBECTOMY. THE LEFT ARM DEFICIT UNFORTUNATELY DID NOT SIGNIFICANTLY IMPROVE POST THROMBECTOMY. THE PATIENT WAS THEN ADMITTED TO THE INTENSIVE CARE UNIT WHERE THEY UNDERWENT MULTIPLE HEAD COMPUTED TOMOGRAPHY SCANS. ANTICOAGULATION WAS SAID TO HAVE BEEN HELP DUE TO CONCERN OF SMALL SUBARACHNOID HEMORRHAGE POST PROCEDURE. ON 24JUL2025, IT WAS REPORTED THAT DUE TO THE PATIENT'S RECENT STROKE, ANTICOAGULATION WAS HELD AND PATIENT HAD ATRIAL FIBRILLATION ON AND OFF. THE PATIENT WAS SITTING UP IN THE CHAIR WHEN THEIR FLOWS SUDDENLY DROPPED TO 0 LITERS PER MINUTE. DESPITE A SYSTEM CONTROLLER EXCHANGE, FLOW CONTINUED AT 0. THE VENTRICULAR ASSIST DEVICE COORDINATORS (VAD) REPORTED THE ABILITY TO HEAR THE REVVING NOISE, AND THE PATIENT BECAME PULSATILE. THE DECISION WAS MADE TO TURN VAD OFF, AND THE PATIENT PASSED AWAY FEW MINUTES AFTER THE LEFT VENTRICULAR ASSIST DEVICE STOPPAGE. LOG FILES WERE SUBMITTED FOR REVIEW. THE EVENT LOG FILE CAPTURED FREQUENT PI EVENTS ON (B)(6) 2025 FROM 0:55 TO 13:20. THE EVENT LOG FILE THEN CAPTURED LOW FLOW ALARMS FROM STARTING ON 13:20 TO 13:33 THEN THE DRIVELINE WAS DISCONNECTED. THE FLOW WAS 0 AND THE PI VALUES WERE NON-VARIABLE AND ROUTINELY IN THE RANGE 2.4 TO 4.3. THE BACKUP SYSTEM CONTROLLER CONTAINED DATA FROM (B)(6) 2025 AT 12:36 TO 12:59. THE EVENT LOG FILE CAPTURED ON SET OF LOW FLOW ALARMS FROM STARTING ON 12:36 TO 12:58 WHEN THE INTENTIONAL PUMP STOP WAS ACTIVATED ON HEARTMATE TOUCH. THE FLOW WAS 0 AND THE PI VALUES WERE NON-VARIABLE AND ROUTINELY IN THE 1.8 TO 3.5. THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO PRESUMED PUMP THROMBOSIS IN THE "EYE OF THE PUMP" DUE TO ATRIAL FIBRILLATION AND BEING OFF OF ANTICOAGULATION.Event Description: IT WAS LATER REPORTED THROUGH A MEDWATCH REPORT THAT THE PATIENT WAS FOUND TO HAVE A DISTAL RIGHT M1 OCCLUSION THAT REQUIRED MECHANICAL THROMBECTOMY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: THE REPORTED PUMP THROMBOSIS WAS UNABLE TO BE CONFIRMED THROUGH THIS EVALUATION. REVIEW OF THE SUBMITTED LOG FILES CONFIRMED THE REPORTED LOW FLOW ALARMS, WHICH THE ACCOUNT CONTRIBUTED TO THE SUSPECTED PUMP THROMBOSIS IN THE EYE OF THE PUMP. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILES COLLECTIVELY CONTAINED EVENTS FROM (B)(6) 2025. LOW FLOW HAZARD ALARMS WERE CAPTURED ON (B)(6) 2025 WHEN THE ESTIMATED FLOW SUDDENLY DECREASED TO 0 LITERS PER MINUTE (LPM). NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE DEVICE APPEARED TO OPERATE AS INTENDED AT THE FIXED SPEED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. D AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS, INCLUDING PUMP THROMBOSIS, STROKE, CARDIAC ARRHYTHMIA, BLEEDING, AND DEATH, WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THIS SECTION ALSO ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW AND PULSATILITY INDEX (PI). IN REFERENCE TO PI, SECTIONS 1 AND 4 OF THE IFU STATE THAT THE PI CALCULATION REPRESENTS CARDIAC PULSATILITY AND VALUES TYPICALLY RANGE FROM 1 TO 10. IN GENERAL, THE MAGNITUDE OF THE PI VALUE IS RELATED TO THE AMOUNT OF ASSISTANCE PROVIDED BY THE PUMP. HIGHER VALUES INDICATE MORE VENTRICULAR FILLING AND HIGHER PULSATILITY (I.E., THE PUMP IS PROVIDING LESS SUPPORT TO THE LEFT VENTRICLE). SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 5 OF THE IFU, "SURGICAL PROCEDURES", INSTRUCTS TO INSPECT THE VENTRICULAR CHAMBER FOR MURAL THROMBI AND CROSSING TRABECULAE FOLLOWING REMOVAL OF THE CORE AND TO ADDRESS ONE OR BOTH, AS NEEDED. THIS SECTION ALSO INSTRUCTS TO INSPECT THE VENTRICLE AND REMOVE ANY PREVIOUSLY FORMED CLOTS THAT MAY CAUSE EMBOLISM OR ANY TRABECULAE THAT MAY IMPEDE FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS THROMBOEMBOLISM AND ARRHYTHMIA AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. THIS SECTION, UNDER ¿ANTICOAGULATION¿, ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INR RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. FOR PATIENTS WITH LOW FLOW SOFTWARE SYSTEM CONTROLLERS, THE HEARTMATE 3 LVAS POCKET GUIDES TO ALARMS FOR PATIENTS (REV. B) AND CLINICIANS (REV. B) EXPLAIN THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.0 LPM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22674232&pc=OZD | 1220648-2025-30878 | 2025-07-17 04:00 | Death | ABIOMED, INC. | 2025-07-31 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT BEGAN BLEEDING FROM THEIR FEMORAL ACCESS SITE DURING IMPELLA CP SUPPORT. THE HEPARIN DRIP WAS HELD TO MANAGE THE CONDITION. LATER THAT SAME DAY, THE FAMILY MADE THE DECISION TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)." SECTION: GENERAL PATIENT CARE CONSIDERATIONS: "ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA." SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: "REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22674689&pc=IKY | 3007420694-2025-00188 | 2025-07-02 04:00 | Death | ARJOHUNTLEIGH POLSKA SP. Z O.O. | 2025-07-31 04:00 | IKY | ARJO FOAM MATTRESS | Fire | Burn(s) |  |  | 1 | Event Description: THE ALLEGATION IS THAT THE PATIENT WAS ON THE ARJO MATTRESS WHEN THE FIRE HAPPENED. THE EXACT CAUSE OF THE FIRE WAS NOT CURRENTLY ESTABLISHED; HOWEVER, IT IS SUSPECTED THAT A PATIENT SMOKED A CIGARETTE, WHICH WAS THE CAUSE. THE PATIENT PASSED AWAY AS A RESULT OF A FIRE. ACCORDING TO THE INFORMATION RECEIVED BY AN ARJO EMPLOYEE, THE MATTRESS THAT WAS ALLEGEDLY IN USE FOR THE PATIENT TREATMENT WAS THE ARJO FOAM MATTRESS. Manufacturer Narrative: THE INVESTIGATION IS STILL ONGOING. FURTHER INFORMATION WILL BE AVAILABLE IN THE FINAL REPORT. H3 OTHER TEXT : DEVICE IS NOT AVAILABLE FOR THE INSPECTION. Manufacturer Narrative: THE INVESTIGATION IS ONGOING. FURTHER INFORMATION WILL BE AVAILABLE IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22674690&pc=IKY | 1419652-2025-00042 | 2025-07-02 04:00 | Death | ARJOHUNTLEIGH POLSKA SP. Z O.O. | 2025-07-31 04:00 | IKY | ARJO FOAM MATTRESS | Fire | Burn(s) |  |  | 1 | Event Description: THE ALLEGATION IS THAT THE PATIENT WAS ON THE ARJO MATTRESS WHEN THE FIRE HAPPENED. THE EXACT CAUSE OF THE FIRE WAS NOT CURRENTLY ESTABLISHED; HOWEVER, IT IS SUSPECTED THAT A PATIENT SMOKED A CIGARETTE, WHICH WAS THE CAUSE. THE PATIENT PASSED AWAY AS A RESULT OF A FIRE. ACCORDING TO THE INFORMATION RECEIVED BY AN ARJO EMPLOYEE, THE MATTRESS THAT WAS ALLEGEDLY IN USE FOR THE PATIENT TREATMENT WAS THE ARJO FOAM MATTRESS. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22675386&pc=MCW | 3004742232-2025-00217 | 2025-05-30 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-07-31 04:00 | MCW | STEALTH | Unintended System Motion; Material Separation; Activation Problem | Vascular Dissection | K220568 |  | 1 | Event Description: IT WAS REPORTED THE STEALTH 360 PERIPHERAL ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A CALCIFIED, TORTUOUS COMMON ILIAC ARTERY. DURING DISTAL TO PROXIMAL TREATMENT, THE OAD CROWN SEEMED TO BE JUMPING SLIGHTLY AND NOT SPINNING AS PER USUAL. UPON REMOVAL, IT WAS OBSERVED THE OAD CROWN HAD DETACHED AND REMAINED ON THE VIPERWIRE GUIDE WIRE. A DISSECTION WAS OBSERVED. A STENT WAS PLACED AT THE DISSECTION SITE. TWO DAYS POST PROCEDURE, THE PATIENT EXPIRED FROM COMPLICATIONS. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22675665&pc=OZD | 1220648-2025-30890 | 2025-07-18 04:00 | Death | ABIOMED, INC. | 2025-07-31 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Deformation Due to Compressive Stress | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD AN IMPELLA CP PLACED TO SUPPORT A PATIENT ADMITTED IN AFTER BEING TRANSFERRED FROM COMMUNITY TO THE TERTIARY CENTER. AN IMPELLA 14 X 25 CENTIMETER SHEATH WAS INSERTED. THE IMPELLA CP WAS INSERTED AND ADVANCED OVER THE WIRE WITH NO ISSUES. THE DECISION WAS MADE TO LEAVE THE DISLODGED PORTION OF PREVIOUS PUMP UNTIL THE PATIENT WAS STABILIZED AND TAKE TO SURGERY. WHILE EXCHANGING THE PEEL AWAY FROM THE REPOSITIONING SHEATH, THE IMPELLA CP WAS WATCHED CLOSELY UNDER FLUOROSCOPY. WHILE ON FLUOROSCOPY, DURING REPOSITIONING, AN ODD SHAPE OF THE CATHETER WAS NOTICED AND A KINK WAS LOCATED BETWEEN THE MOTOR AND BEND OF THE CANNULA. FLOWS WERE DIMINISHED AND THE DECISION WAS MADE TO LEAVE ACROSS THE VALVE AS IT WAS THE LAST IMPELLA ON-SITE. THE PATIENT WAS SENT TO THE INTENSIVE CARE UNIT (ICU) WITH A POSSIBLE PLAN OF EXTRACTION OF THE DISLODGED IMPELLA IF THE PATIENT STABILIZED. THE PATIENT DECLINED IN THE ICU AND PASSED AWAY ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AS IT WAS LEFT INSIDE THE PATIENT WHEN THEY PASSED AWAY. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22675987&pc=OZD | 1220648-2025-30891 | 2025-07-17 04:00 | Death | ABIOMED, INC. | 2025-07-31 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT HAD AN IMPELLA 5.5 PLACED IN A PATIENT WITH NON-ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION AFTER THE INTRA-AORTIC BALLOON PUMP WAS EXPLANTED AND A 4-VESSEL CORONARY ARTERY BYPASS GRAFT WAS PLANNED AT THE SECOND, TERTIARY HOSPITAL. THE PATIENT SUFFERED FROM VENTRICULAR TACHYCARDIA THAT WAS TREATED BY DEFIBRILLATION ON THE SECOND DAY OF IMPELLA 5.5 SUPPORT. TWO UNITS OF RED BLOOD CELLS WERE GIVEN AND ALBUMIN. TWO AMIODARONE BOLUSES WERE DONE. THE PATIENT WAS CARDIOVERTED TWICE FOR RAPID VENTRICULAR RESPONSE. THE PUMP REMAINED ON DESPITE THIS UNTIL CARE WAS WITHDRAWN AND THE PATIENT EXPIRED AFTER JUST UNDER FOUR DAYS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22676561&pc=OZD | 1220648-2025-30893 | 2025-07-17 04:00 | Death | ABIOMED, INC. | 2025-07-31 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, VENTRICULAR TACHYCARDIA OCCURRED. THE PATIENT WAS CODED AND SHOCKED MULTIPLE TIMES. THE DECISION WAS MADE TO PLACE VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION WITH IMPELLA 5.5 SUPPORT. HOWEVER, THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22676624&pc=DQY | 2124215-2025-52155 | 2025-07-23 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-31 04:00 | DQY | WATCHMAN FXD CURVE? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Pericardial Effusion | K212228 |  | 1 | Event Description: IT WAS REPORTED A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE WAS BEING PERFORMED ON A PATIENT WITH MARFAN SYNDROME. A WATCHMAN FXD DOUBLE CURVE ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS SELECTED FOR USE. A SMALL EFFUSION WAS NOTED ON THE RIGHT VENTRICLE PRIOR TO THE START OF THE PROCEDURE. AFTER THE TRANSEPTAL PUNCTURE (TSP) USING THE VERSACROSS CONNECT SYSTEM AND THE WAS DOUBLE CURVE SHEATH, THE PATIENT DEVELOPED LARGER EFFUSION AROUND THE WHOLE PERICARDIUM ALONG WITH CARDIAC TAMPONADE. CARDIOPULMONARY RESUSCITATION (CPR) WAS STARTED. A PERICARDIAL TAP WAS ATTEMPTED TO TREAT THE EFFUSION BUT WAS UNSUCCESSFUL. RESUSCITATION WAS DONE FOR THIRTY (30) MINUTES, HOWEVER THE PATIENT DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22676834&pc=DSQ | 2916596-2025-04827 | 2025-06-05 04:00 | Death | THORATEC CORPORATION | 2025-07-31 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT COMMUNICATED AND WAS STATED TO BE END OF LIFE-PALLIATIVE CARE. THE DEATH WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22677231&pc=OZD | 1220648-2025-30897 | 2025-07-18 04:00 | Death | ABIOMED, INC. | 2025-07-31 04:00 | OZD | IMPELLA CP | Malposition of Device | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED A PLACEMENT SIGNAL ALARM THAT THE PUMP WAS POSITIONED DEEP. THE DOCTOR DID NOT WANT TO ADJUST THE DEPTH AND WAS LEFT THE PUMP AS IS. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22677338&pc=DSQ | 2916596-2025-05058 | 2025-07-23 04:00 | Death | THORATEC CORPORATION | 2025-07-31 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE WAS NOT PROVIDED. THE PATIENT WAS VERY FRAIL AT THE TIME OF IMPLANTATION. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV C, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV B, ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22677570&pc=DSQ | 2916596-2025-05069 | 2023-11-08 05:00 | Death | THORATEC CORPORATION | 2025-07-31 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE PATIENT OUTCOME, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS MULTI ORGAN FAILURE. THE PUMP WAS WORKING AS EXPECTED. NO PUMP RELATED ISSUES WERE REPORTED. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT BE EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) REVISION, REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22677767&pc=DRE | 3007284006-2025-00152 | 2025-07-09 04:00 | Death | SPECTRANETICS CORPORATION | 2025-07-31 04:00 | DRE | TIGHTRAIL ROTATING DILATOR SHEATH | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation; Cardiac Perforation; Pericardial Effusion | K150360 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) LEAD DUE TO FRACTURE. A RIGHT ATRIAL (RA), LEFT VENTRICULAR (LV), AND A CAPPED RV LEAD WERE PRESENT IN THE PATIENT, BUT NOT TARGETED FOR EXTRACTION. A SPECTRANETICS LLD LEAD LOCKING DEVICE (LLD) WAS INSERTED INTO THE LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 11F TIGHTRAIL ROTATING DILATOR SHEATH, SOME RESISTANCE WAS ENCOUNTERED; THEREFORE, A COOK MEDICAL INTRODUCER SHEATH WAS USED. THE LEAD CAME FREE FROM THE WALL OF THE RV AND A MERIT MEDICAL EN SNARE DEVICE WAS USED TO REPOSITION THE LEAD TO IMPROVE ALIGNMENT AND TRACTION. FURTHER RESISTANCE WAS MET AT THE COIL OF THE LEAD, AND UPSIZED TO A SPECTRANETICS 13F TIGHTRAIL ROTATING DILATOR SHEATH. AFTER NO FURTHER ADVANCEMENT COULD BE MADE, THE COOK MEDICAL INTRODUCER SHEATH WAS LOADED ONTO THE TIGHTRAIL. WITH SOME DIFFICULTY GOING AROUND THE INNOMINATE/SUPERIOR VENA CAVA (SVC) JUNCTION, THE TIGHTRAIL MOVED FURTHER ALONG THE LEAD PAST THE SVC, WHEN PROGRESS STALLED. UPON REMOVAL OF THE TIGHTRAIL, A PERICARDIAL EFFUSION WAS DETECTED VIA TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE), AND RESCUE EFFORTS BEGAN, INCLUDING BYPASS AND STERNOTOMY. A SVC PERFORATION THAT EXTENDED INTO THE RA WAS DISCOVERED AND ATTEMPTS WERE MADE TO REPAIR; HOWEVER, BLEEDING CONTINUED. DESPITE LIFESAVING EFFORTS, THE PATIENT DID NOT SURVIVE THE PROCEDURE. THIS REPORT CAPTURES THE 13F TIGHTRAIL IN USE IN THE AREA WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION, BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: A4) PATIENT WEIGHT - NOT AVAILABLE FROM FACILITY. A5) PATIENT ETHNICITY - NOT AVAILABLE FROM FACILITY. B7) OTHER RELEVANT HISTORY - NOT AVAILABLE FROM FACILITY. H3) THE TIGHTRAIL WAS DISCARDED, THUS NO PRODUCT INVESTIGATION WAS PERFORMED. H6) GREAT VESSEL PERFORATION, CARDIAC PERFORATION, AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE TIGHTRAIL DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22677777&pc=NGV | 2135147-2025-04197 | 2022-05-05 04:00 | Death | ABBOTT MEDICAL | 2025-07-31 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction | P200049 | GEN2201026 | 3 | Event Description: IT WAS REPORTED THROUGH AMULET LAAO REGISTRY DATA THAT AMPLATZER AMULET DEVICES MAY BE RELATED TO 3 ACUTE MYOCARDIAL INFARCTION WHICH ARE CONSIDERED DEATH EVENTS. THE RELATIONSHIP OF THE DEATH TO THE AMPLATZER AMULET DEVICES COULD NOT BE DETERMINED BASED ON THE LIMITED DATA RECEIVED FROM THE REGISTRY. AMULET LAAO REGISTRY DATA IS REPORTED AS A SUMMARY PER SUMMARY REPORTING EXEMPTION APPROVAL NUMBER - (B)(4). THE DATA RECEIVED INDICATED THAT THE PROCEDURE DATE RANGE WAS BETWEEN (B)(6) 2022 - (B)(6) 2023. PATIENTS¿ MEAN AGE IS 69 YEARS, RANGING FROM 55 - 83 YEARS. 67% OF THE PATIENTS WERE MALE, 33% OF THE PATIENTS WERE FEMALE. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN (B)(6) 2025. AS OF (B)(6) 2025, (B)(4) PATIENTS WERE TREATED WITH THE AMULET DEVICE. Manufacturer Narrative: 3 EVENTS OF DEATH DUE TO ACUTE MYOCARDIAL INFARCTION WERE REPORTED IN THE CURRENT REPORTING PERIOD. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS LIMITED INFORMATION WAS PROVIDED INCLUDING THAT THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE LIMITED INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. HOWEVER, RISK ASSESSMENT WAS COMPLETED FOR THE REPORTED INCIDENTS. (B)(4) SUBJECTS EXPERIENCED DEATH DUE TO ACUTE MYOCARDIAL INFARCTION SINCE THE BEGINNING OF THE REGISTRY. THE DEATH EVENTS AND RATES REPORTED IN THIS STUDY WERE FOUND TO BE ANTICIPATED IN THE STUDY POPULATION AND ARE IN ALIGNMENT WITH RATES OBSERVED IN OTHER CLINICAL STUDIES. NO NEW RISKS HAVE BEEN IDENTIFIED FOR THIS POPULATION. NO NEW HAZARDS OR HARMS WERE IDENTIFIED THAT WOULD IMPACT THE BENEFIT/RISK PROFILE. BASED ON A REVIEW OF THE AVAILABLE INFORMATION, THERE WAS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. THEREFORE, NO REMEDIAL ACTION IS REQUIRED AT THIS TIME. B2 - DATE OF DEATH: THE EARLIEST DATE OF DEATH WAS USED D4. PRIMARY UDI NUMBER: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A - IMPLANT DATE: THE EARLIEST IMPLANT DATE WAS USED. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22677833&pc=NPT | 9617601-2025-00763 | 2025-07-09 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-07-31 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Hemorrhagic Stroke | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE, ACCESS WAS INITIALLY PLANNED WITH A NON-MEDTRONIC SHEATH; HOWEVER, DUE TO THE "POOR VESSEL CONDITION," A NON-MEDTRONIC SHEATH WAS USED INSTEAD. BLEEDING THEN OCCURRED AT THE ACCESS SITE, AND TWO NON-MEDTRONIC DEVICES WERE DEPLOYED, BUT ONE FAILED TO SECURE THE VESSEL. COMPRESSION HEMOSTASIS WAS ATTEMPTED FOR NEARLY AN HOUR BUT WAS UNSUCCESSFUL. AS A RESULT, SURGICAL REPAIR WAS PERFORMED WITH AN APPROXIMATELY 3 CENTIMETER (CM) INCISION TO SUTURE THE VESSEL, AND THE ACCESS SITE BLEEDING WAS "WELL-CONTROLLED" FOLLOWING THE PROCEDURE. IT WAS FURTHER REPORTED THAT 5 DAYS FOLLOWING THE PROCEDURE, THE PATIENT'S CONDITION DETERIORATED AND A CEREBRAL HEMORRHAGE OCCURRED. SUBSEQUENTLY, 6 DAYS FOLLOWING THE PROCEDURE, DESPITE EMERGENCY INTERVENTIONS, THE PATIENT DIED.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE, ACCESS WAS INITIALLY PLANNED WITH A NON-MEDTRONIC SHEATH; HOWEVER, DUE TO THE "POOR VESSEL CONDITION," A NON-MEDTRONIC SHEATH WAS USED INSTEAD. BLEEDING THEN OCCURRED AT THE ACCESS SITE, AND TWO NON-MEDTRONIC DEVICES WERE DEPLOYED, BUT ONE FAILED TO SECURE THE VESSEL. COMPRESSION HEMOSTASIS WAS ATTEMPTED FOR NEARLY AN HOUR BUT WAS UNSUCCESSFUL. AS A RESULT, SURGICAL REPAIR WAS PERFORMED WITH AN APPROXIMATELY 3 CENTIMETER (CM) INCISION TO SUTURE THE VESSEL, AND THE ACCESS SITE BLEEDING WAS "WELL-CONTROLLED" FOLLOWING THE PROCEDURE. IT WAS FURTHER REPORTED THAT 5 DAYS FOLLOWING THE PROCEDURE, THE PATIENT'S CONDITION DETERIORATED AND A CEREBRAL HEMORRHAGE OCCURRED. SUBSEQUENTLY, 6 DAYS FOLLOWING THE PROCEDURE, DESPITE EMERGENCY INTERVENTIONS, THE PATIENT DIED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT BLEEDING FROM THE ACCESS VESSEL WAS WELL-CONTROLLED POST-OPERATIVELY. THE PATIENT AWAKENED FOLLOWING THE PROCEDURE WITHOUT ISSUE. 5 DAYS FOLLOWING THE PROCEDURE THE PATIENT BEGAN REHABILITATION IN THE MORNING. THE CEREBRAL HEMORRHAGE OCCURRED THAT NIGHT. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012529124); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID D-EVOLUTFX-2329(B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NA; EXPLANT DATE: NA SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. ADDED SECOND PARAGRAPH. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22677847&pc=NPT | 9612164-2025-03754 | 2025-07-09 04:00 | Death | MEDTRONIC IRELAND | 2025-07-31 04:00 | NPT | EVOLUT FX DCS | Patient Device Interaction Problem | Hemorrhage/Bleeding; Hemorrhagic Stroke | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE, ACCESS WAS INITIALLY PLANNED WITH A NON-MEDTRONIC SHEATH; HOWEVER, DUE TO THE "POOR VESSEL CONDITION," A NON-MEDTRONIC SHEATH WAS USED INSTEAD. BLEEDING THEN OCCURRED AT THE ACCESS SITE, AND TWO NON-MEDTRONIC DEVICES WERE DEPLOYED, BUT ONE FAILED TO SECURE THE VESSEL. COMPRESSION HEMOSTASIS WAS ATTEMPTED FOR NEARLY AN HOUR BUT WAS UNSUCCESSFUL. AS A RESULT, SURGICAL REPAIR WAS PERFORMED WITH AN APPROXIMATELY 3 CENTIMETER (CM) INCISION TO SUTURE THE VESSEL, AND THE ACCESS SITE BLEEDING WAS "WELL-CONTROLLED" FOLLOWING THE PROCEDURE. IT WAS FURTHER REPORTED THAT 5 DAYS FOLLOWING THE PROCEDURE, THE PATIENT'S CONDITION DETERIORATED AND A CEREBRAL HEMORRHAGE OCCURRED. SUBSEQUENTLY, 6 DAYS FOLLOWING THE PROCEDURE, DESPITE EMERGENCY INTERVENTIONS, THE PATIENT DIED.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE, ACCESS WAS INITIALLY PLANNED WITH A NON-MEDTRONIC SHEATH; HOWEVER, DUE TO THE "POOR VESSEL CONDITION," A NON-MEDTRONIC SHEATH WAS USED INSTEAD. BLEEDING THEN OCCURRED AT THE ACCESS SITE, AND TWO NON-MEDTRONIC DEVICES WERE DEPLOYED, BUT ONE FAILED TO SECURE THE VESSEL. COMPRESSION HEMOSTASIS WAS ATTEMPTED FOR NEARLY AN HOUR BUT WAS UNSUCCESSFUL. AS A RESULT, SURGICAL REPAIR WAS PERFORMED WITH AN APPROXIMATELY 3 CENTIMETER (CM) INCISION TO SUTURE THE VESSEL, AND THE ACCESS SITE BLEEDING WAS "WELL-CONTROLLED" FOLLOWING THE PROCEDURE. IT WAS FURTHER REPORTED THAT 5 DAYS FOLLOWING THE PROCEDURE, THE PATIENT'S CONDITION DETERIORATED AND A CEREBRAL HEMORRHAGE OCCURRED. SUBSEQUENTLY, 6 DAYS FOLLOWING THE PROCEDURE, DESPITE EMERGENCY INTERVENTIONS, THE PATIENT DIED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT BLEEDING FROM THE ACCESS VESSEL WAS WELL-CONTROLLED POST-OPERATIVELY. THE PATIENT AWAKENED FOLLOWING THE PROCEDURE WITHOUT ISSUE. 5 DAYS FOLLOWING THE PROCEDURE THE PATIENT BEGAN REHABILITATION IN THE MORNING. THE CEREBRAL HEMORRHAGE OCCURRED THAT NIGHT. Manufacturer Narrative: UPDATED DATA: B5. ADDED SECOND PARAGRAPH. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012529124); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVFXPLUS-29 (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025: NA SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22677975&pc=OZD | 1220648-2025-30906 | 2025-07-19 04:00 | Death | ABIOMED, INC. | 2025-07-31 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hematoma; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA 5.5 HAD A LARGE HEMATOMA ON THE LATERAL THORAX/AXILLA. THE HEMATOMA WAS EVACUATED IN THE OPERATING ROOM AND THE PATIENT WAS TRANSFUSED FIVE UNITS OF PACKED RED BLOOD CELLS AND TWO UNITS OF FRESH FROZEN PLASMA. THE PURGE WAS SWITCHED FROM HEPARIN TO SODIUM BICARB. THE FOLLOWING DAY, THE PATIENT CONTINUED TO BLEED AND A GASTROINTESTINAL BLEED WAS IDENTIFIED. SEVEN UNITS OF PACKED RED BLOOD CELLS WERE TRANSFUSED. THE PATIENT WENT INTO PULSELESS ELECTRICAL ACTIVITY, AND THE FAMILY DECIDED TO WITHDRAW CARE. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22679337&pc=OZD | 1220648-2025-30909 | 2025-07-18 04:00 | Death | ABIOMED, INC. | 2025-07-31 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND HAD BLEEDING FROM THE ACCESS SITE DUE TO EXCESSIVE MOVEMENT. ALBUMIN WAS GIVEN AND A FEMSTOP WAS APPLIED WITH A SANDBAG. LATER, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22679438&pc=DSQ | 2916596-2025-04875 | 2025-05-26 04:00 | Death | THORATEC CORPORATION | 2025-07-31 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. IT WAS REPORTED THAT THERE WAS NO PUMP ISSUES AND THAT THIS WAS A NATURAL DEATH. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22679447&pc=OZD | 1220648-2025-30910 | 2025-07-19 04:00 | Death | ABIOMED, INC. | 2025-07-31 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Application Program Problem | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PLACEMENT SIGNAL ERROR ALARM APPEARED ON THE AUTOMATED IMPELLA CONTROLLER (AIC) DURING IMPELLA 5.5 SUPPORT. THE ALARM SELF RESOLVED. ULTIMATELY, CARE WAS WITHDRAWN AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: A.2, A.4, AND A.5 WERE LEFT BLANK AS PATIENT INFORMATION IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22679580&pc=NGV | 2135147-2025-04216 | 2022-03-03 05:00 | Death | ABBOTT MEDICAL | 2025-07-31 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200049 | GEN2201026 | 31 | Event Description: IT WAS REPORTED THROUGH AMULET LAAO REGISTRY DATA THAT AMPLATZER AMULET DEVICES MAY BE RELATED TO 31 UNKNOWN DEATHS WHICH ARE CONSIDERED DEATH EVENTS. THE RELATIONSHIP OF THE DEATH TO THE AMPLATZER AMULET DEVICES COULD NOT BE DETERMINED BASED ON THE LIMITED DATA RECEIVED FROM THE REGISTRY. AMULET LAAO REGISTRY DATA IS REPORTED AS A SUMMARY PER SUMMARY REPORTING EXEMPTION APPROVAL NUMBER - GEN2201026. THE DATA RECEIVED INDICATED THAT THE PROCEDURE DATE RANGE WAS BETWEEN 03MARCH2022 - 14DECEMBER2023. PATIENTS¿ MEAN AGE IS 81 YEARS, RANGING FROM 63 - 93 YEARS. 71% OF THE PATIENTS WERE MALE, 29% OF THE PATIENTS WERE FEMALE. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025. AS OF 04 APRIL 2025, 12,180 PATIENTS WERE TREATED WITH THE AMULET DEVICE. Manufacturer Narrative: 31 EVENTS OF DEATH DUE TO UNKNOWN CAUSE WERE REPORTED IN THE CURRENT REPORTING PERIOD. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS LIMITED INFORMATION WAS PROVIDED INCLUDING THAT THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE LIMITED INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. HOWEVER, RISK ASSESSMENT WAS COMPLETED FOR THE REPORTED INCIDENTS. 95 OUT OF 12,180 (0.78%) SUBJECTS EXPERIENCED DEATH DUE TO UNKNOWN CAUSE SINCE THE BEGINNING OF THE REGISTRY. THE DEATH EVENTS AND RATES REPORTED IN THIS STUDY WERE FOUND TO BE ANTICIPATED IN THE STUDY POPULATION AND ARE IN ALIGNMENT WITH RATES OBSERVED IN OTHER CLINICAL STUDIES. NO NEW RISKS HAVE BEEN IDENTIFIED FOR THIS POPULATION. NO NEW HAZARDS OR HARMS WERE IDENTIFIED THAT WOULD IMPACT THE BENEFIT/RISK PROFILE. BASED ON A REVIEW OF THE AVAILABLE INFORMATION, THERE WAS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. THEREFORE, NO REMEDIAL ACTION IS REQUIRED AT THIS TIME. B2 - DATE OF DEATH: THE EARLIEST DATE OF DEATH WAS USED. D4. PRIMARY UDI NUMBER: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A - IMPLANT DATE: THE EARLIEST IMPLANT DATE WAS USED. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22679633&pc=DSQ | 2916596-2025-04930 | 2023-08-15 04:00 | Death | THORATEC CORPORATION | 2025-07-31 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FROM VENTRICULAR FIBRILLATION AFTER RECEIVING MULTIPLE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR SHOCKS. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22679638&pc=NGV | 2135147-2025-04218 | 2022-05-18 04:00 | Death | ABBOTT MEDICAL | 2025-07-31 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P200049 | GEN2201026 | 12 | Event Description: IT WAS REPORTED THROUGH AMULET LAAO REGISTRY DATA THAT AMPLATZER AMULET DEVICES MAY BE RELATED TO 12 CAUSE OF DEATH: HEART FAILURE WHICH ARE CONSIDERED DEATH EVENTS. THE RELATIONSHIP OF THE DEATH TO THE AMPLATZER AMULET DEVICES COULD NOT BE DETERMINED BASED ON THE LIMITED DATA RECEIVED FROM THE REGISTRY. AMULET LAAO REGISTRY DATA IS REPORTED AS A SUMMARY PER SUMMARY REPORTING EXEMPTION APPROVAL NUMBER - GEN2201026. THE DATA RECEIVED INDICATED THAT THE PROCEDURE DATE RANGE WAS BETWEEN 18MAY2022 - 07DECEMBER2023. PATIENTS¿ MEAN AGE IS 81 YEARS, RANGING FROM 61 - (B)(6) YEARS. 58% OF THE PATIENTS WERE MALE, 42% OF THE PATIENTS WERE FEMALE. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025. AS OF 04 APRIL 2025, 12,180 PATIENTS WERE TREATED WITH THE AMULET DEVICE. Manufacturer Narrative: 12 EVENTS OF DEATH DUE TO HEART FAILURE WERE REPORTED IN THE CURRENT REPORTING PERIOD. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS LIMITED INFORMATION WAS PROVIDED INCLUDING THAT THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE LIMITED INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. HOWEVER, RISK ASSESSMENT WAS COMPLETED FOR THE REPORTED INCIDENTS. 64 OUT OF 12,180 (0.52%) SUBJECTS EXPERIENCED DEATH DUE TO HEART FAILURE SINCE THE BEGINNING OF THE REGISTRY. THE DEATH EVENTS AND RATES REPORTED IN THIS STUDY WERE FOUND TO BE ANTICIPATED IN THE STUDY POPULATION AND ARE IN ALIGNMENT WITH RATES OBSERVED IN OTHER CLINICAL STUDIES. NO NEW RISKS HAVE BEEN IDENTIFIED FOR THIS POPULATION. NO NEW HAZARDS OR HARMS WERE IDENTIFIED THAT WOULD IMPACT THE BENEFIT/RISK PROFILE. BASED ON A REVIEW OF THE AVAILABLE INFORMATION, THERE WAS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. THEREFORE, NO REMEDIAL ACTION IS REQUIRED AT THIS TIME. B2 - DATE OF DEATH: THE EARLIEST DATE OF DEATH WAS USED. D4. PRIMARY UDI NUMBER: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A - IMPLANT DATE: THE EARLIEST IMPLANT DATE WAS USED. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22679664&pc=NGV | 2135147-2025-04219 | 2022-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-07-31 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P200049 | GEN2201026 | 12 | Event Description: IT WAS REPORTED THROUGH AMULET LAAO REGISTRY DATA THAT AMPLATZER AMULET DEVICES MAY BE RELATED TO 12 INFECTION WHICH ARE CONSIDERED DEATH EVENTS. THE RELATIONSHIP OF THE DEATH TO THE AMPLATZER AMULET DEVICES COULD NOT BE DETERMINED BASED ON THE LIMITED DATA RECEIVED FROM THE REGISTRY. AMULET LAAO REGISTRY DATA IS REPORTED AS A SUMMARY PER SUMMARY REPORTING EXEMPTION APPROVAL NUMBER - GEN2201026. THE DATA RECEIVED INDICATED THAT THE PROCEDURE DATE RANGE WAS BETWEEN (B)(6) 2022 - (B)(6) 2023. PATIENTS¿ MEAN AGE IS 76 YEARS, RANGING FROM 59 - 87 YEARS. 75% OF THE PATIENTS WERE MALE, 25% OF THE PATIENTS WERE FEMALE. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN (B)(6) 2025. AS OF (B)(6) 2025, 12,180 PATIENTS WERE TREATED WITH THE AMULET DEVICE. Manufacturer Narrative: 12 EVENTS OF DEATH DUE TO INFECTION WERE REPORTED IN THE CURRENT REPORTING PERIOD. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS LIMITED INFORMATION WAS PROVIDED INCLUDING THAT THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE LIMITED INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. HOWEVER, RISK ASSESSMENT WAS COMPLETED FOR THE REPORTED INCIDENTS. A RISK ASSESSMENT WAS COMPLETED FOR THE REPORTED INCIDENTS. 53 OUT OF 12,180 (0.44%) SUBJECTS EXPERIENCED DEATH DUE TO INFECTION SINCE THE BEGINNING OF THE REGISTRY. THE DEATH EVENTS AND RATES REPORTED IN THIS STUDY WERE FOUND TO BE ANTICIPATED IN THE STUDY POPULATION AND ARE IN ALIGNMENT WITH RATES OBSERVED IN OTHER CLINICAL STUDIES. NO NEW RISKS HAVE BEEN IDENTIFIED FOR THIS POPULATION. NO NEW HAZARDS OR HARMS WERE IDENTIFIED THAT WOULD IMPACT THE BENEFIT/RISK PROFILE. BASED ON A REVIEW OF THE AVAILABLE INFORMATION, THERE WAS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. THEREFORE, NO REMEDIAL ACTION IS REQUIRED AT THIS TIME. B2 - DATE OF DEATH: THE EARLIEST DATE OF DEATH WAS USED D4. PRIMARY UDI NUMBER: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A - IMPLANT DATE: THE EARLIEST IMPLANT DATE WAS USED THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22679819&pc=NGV | 2135147-2025-04223 | 2022-04-29 04:00 | Death | ABBOTT MEDICAL | 2025-07-31 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Brain Injury | P200049 | GEN2201026 | 12 | Event Description: IT WAS REPORTED THROUGH AMULET LAAO REGISTRY DATA THAT AMPLATZER AMULET DEVICES MAY BE RELATED TO 12 NEUROLOGICAL WHICH ARE CONSIDERED DEATH EVENTS. THE RELATIONSHIP OF THE DEATH TO THE AMPLATZER AMULET DEVICES COULD NOT BE DETERMINED BASED ON THE LIMITED DATA RECEIVED FROM THE REGISTRY. AMULET LAAO REGISTRY DATA IS REPORTED AS A SUMMARY PER SUMMARY REPORTING EXEMPTION APPROVAL NUMBER - GEN2201026. THE DATA RECEIVED INDICATED THAT THE PROCEDURE DATE RANGE WAS BETWEEN 29APRIL2022 - 16NOVEMBER2023. PATIENTS¿ MEAN AGE IS 79 YEARS, RANGING FROM 61 - 87 YEARS. 58% OF THE PATIENTS WERE MALE, 42% OF THE PATIENTS WERE FEMALE. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025. AS OF 04 APRIL 2025, 12,180 PATIENTS WERE TREATED WITH THE AMULET DEVICE. Manufacturer Narrative: 12 EVENTS OF DEATH DUE TO NEUROLOGICAL WERE REPORTED IN THE CURRENT REPORTING PERIOD. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS LIMITED INFORMATION WAS PROVIDED INCLUDING THAT THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE LIMITED INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. HOWEVER, RISK ASSESSMENT WAS COMPLETED FOR THE REPORTED INCIDENTS. 28 OUT OF 12,180 (0.23%) SUBJECTS EXPERIENCED DEATH DUE TO NEUROLOGICAL SINCE THE BEGINNING OF THE REGISTRY. THE DEATH EVENTS AND RATES REPORTED IN THIS STUDY WERE FOUND TO BE ANTICIPATED IN THE STUDY POPULATION AND ARE IN ALIGNMENT WITH RATES OBSERVED IN OTHER CLINICAL STUDIES. NO NEW RISKS HAVE BEEN IDENTIFIED FOR THIS POPULATION. NO NEW HAZARDS OR HARMS WERE IDENTIFIED THAT WOULD IMPACT THE BENEFIT/RISK PROFILE. BASED ON A REVIEW OF THE AVAILABLE INFORMATION, THERE WAS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. THEREFORE, NO REMEDIAL ACTION IS REQUIRED AT THIS TIME. B2 - DATE OF DEATH: THE EARLIEST DATE OF DEATH WAS USED. D4. PRIMARY UDI NUMBER: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A - IMPLANT DATE: THE EARLIEST IMPLANT DATE WAS USED. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22679836&pc=NGV | 2135147-2025-04224 | 2022-05-17 04:00 | Death | ABBOTT MEDICAL | 2025-07-31 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P200049 | GEN2201026 | 13 | Event Description: IT WAS REPORTED THROUGH AMULET LAAO REGISTRY DATA THAT AMPLATZER AMULET DEVICES MAY BE RELATED TO 13 OTHER CARDIOVASCULAR REASON WHICH ARE CONSIDERED DEATH EVENTS. THE RELATIONSHIP OF THE DEATH TO THE AMPLATZER AMULET DEVICES COULD NOT BE DETERMINED BASED ON THE LIMITED DATA RECEIVED FROM THE REGISTRY. AMULET LAAO REGISTRY DATA IS REPORTED AS A SUMMARY PER SUMMARY REPORTING EXEMPTION APPROVAL NUMBER - GEN2201026. THE DATA RECEIVED INDICATED THAT THE PROCEDURE DATE RANGE WAS BETWEEN 17MAY2022 - 06NOVEMBER2023. PATIENTS¿ MEAN AGE IS 75 YEARS, RANGING FROM 60 - 86 YEARS. 62% OF THE PATIENTS WERE MALE, 38% OF THE PATIENTS WERE FEMALE. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025. AS OF 04 APRIL 2025, 12,180 PATIENTS WERE TREATED WITH THE AMULET DEVICE. Manufacturer Narrative: 13 EVENTS OF DEATH DUE TO OTHER CARDIOVASCULAR REASON WERE REPORTED IN THE CURRENT REPORTING PERIOD. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS LIMITED INFORMATION WAS PROVIDED INCLUDING THAT THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE LIMITED INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. HOWEVER, RISK ASSESSMENT WAS COMPLETED FOR THE REPORTED INCIDENTS. 68 OUT OF 12,180 (0.56%) SUBJECTS EXPERIENCED DEATH DUE TO OTHER CARDIOVASCULAR REASON SINCE THE BEGINNING OF THE REGISTRY. THE DEATH EVENTS AND RATES REPORTED IN THIS STUDY WERE FOUND TO BE ANTICIPATED IN THE STUDY POPULATION AND ARE IN ALIGNMENT WITH RATES OBSERVED IN OTHER CLINICAL STUDIES. NO NEW RISKS HAVE BEEN IDENTIFIED FOR THIS POPULATION. NO NEW HAZARDS OR HARMS WERE IDENTIFIED THAT WOULD IMPACT THE BENEFIT/RISK PROFILE. BASED ON A REVIEW OF THE AVAILABLE INFORMATION, THERE WAS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. THEREFORE, NO REMEDIAL ACTION IS REQUIRED AT THIS TIME. B2 - DATE OF DEATH: THE EARLIEST DATE OF DEATH WAS USED. D4. PRIMARY UDI NUMBER: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A - IMPLANT DATE: THE EARLIEST IMPLANT DATE WAS USED. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22679839&pc=NGV | 2135147-2025-04225 | 2022-02-01 05:00 | Death | ABBOTT MEDICAL | 2025-07-31 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Renal Failure | P200049 | GEN2201026 | 5 | Event Description: IT WAS REPORTED THROUGH AMULET LAAO REGISTRY DATA THAT AMPLATZER AMULET DEVICES MAY BE RELATED TO 5 RENAL WHICH ARE CONSIDERED DEATH EVENTS. THE RELATIONSHIP OF THE DEATH TO THE AMPLATZER AMULET DEVICES COULD NOT BE DETERMINED BASED ON THE LIMITED DATA RECEIVED FROM THE REGISTRY. AMULET LAAO REGISTRY DATA IS REPORTED AS A SUMMARY PER SUMMARY REPORTING EXEMPTION APPROVAL NUMBER - GEN2201026. THE DATA RECEIVED INDICATED THAT THE PROCEDURE DATE RANGE WAS BETWEEN 01FEBRUARY2022 - 19OCTOBER2023. PATIENTS¿ MEAN AGE IS 83 YEARS, RANGING FROM 75 - 89 YEARS. 60% OF THE PATIENTS WERE MALE, 40% OF THE PATIENTS WERE FEMALE. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025. AS OF 04 APRIL 2025, 12,180 PATIENTS WERE TREATED WITH THE AMULET DEVICE. Manufacturer Narrative: 5 EVENTS OF DEATH DUE TO RENAL WERE REPORTED IN THE CURRENT REPORTING PERIOD. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS LIMITED INFORMATION WAS PROVIDED INCLUDING THAT THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE LIMITED INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. HOWEVER, RISK ASSESSMENT WAS COMPLETED FOR THE REPORTED INCIDENTS. 25 OUT OF 12,180 (0.21%) SUBJECTS EXPERIENCED DEATH DUE TO RENAL SINCE THE BEGINNING OF THE REGISTRY. THE DEATH EVENTS AND RATES REPORTED IN THIS STUDY WERE FOUND TO BE ANTICIPATED IN THE STUDY POPULATION AND ARE IN ALIGNMENT WITH RATES OBSERVED IN OTHER CLINICAL STUDIES. NO NEW RISKS HAVE BEEN IDENTIFIED FOR THIS POPULATION. NO NEW HAZARDS OR HARMS WERE IDENTIFIED THAT WOULD IMPACT THE BENEFIT/RISK PROFILE. BASED ON A REVIEW OF THE AVAILABLE INFORMATION, THERE WAS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. THEREFORE, NO REMEDIAL ACTION IS REQUIRED AT THIS TIME. B2 - DATE OF DEATH: THE EARLIEST DATE OF DEATH WAS USED. D4. PRIMARY UDI NUMBER: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A - IMPLANT DATE: THE EARLIEST IMPLANT DATE WAS USED. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22679858&pc=NGV | 2135147-2025-04227 | 2022-11-07 05:00 | Death | ABBOTT MEDICAL | 2025-07-31 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P200049 | GEN2201026 | 1 | Event Description: IT WAS REPORTED THROUGH AMULET LAAO REGISTRY DATA THAT AMPLATZER AMULET DEVICES MAY BE RELATED TO 1 TRAUMA WHICH IS CONSIDERED A DEATH EVENT. THE RELATIONSHIP OF THE DEATH TO THE AMPLATZER AMULET DEVICES COULD NOT BE DETERMINED BASED ON THE LIMITED DATA RECEIVED FROM THE REGISTRY. AMULET LAAO REGISTRY DATA IS REPORTED AS A SUMMARY PER SUMMARY REPORTING EXEMPTION APPROVAL NUMBER - GEN2201026. THE DATA RECEIVED INDICATED THAT THE PROCEDURE DATE 07NOVEMBER2022. PATIENT WAS 83 YEARS OLD. PATIENT WAS A MALE. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025. AS OF 04 APRIL 2025, 12,180 PATIENTS WERE TREATED WITH THE AMULET DEVICE. Manufacturer Narrative: ONE EVENT OF DEATH DUE TO TRAUMA WAS REPORTED IN THE CURRENT REPORTING PERIOD. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS LIMITED INFORMATION WAS PROVIDED INCLUDING THAT THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE LIMITED INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. HOWEVER, RISK ASSESSMENT WAS COMPLETED FOR THE REPORTED INCIDENTS. 15 OUT OF 12,180 (0.12%) SUBJECTS EXPERIENCED DEATH DUE TO TRAUMA SINCE THE BEGINNING OF THE REGISTRY. THE DEATH EVENTS AND RATES REPORTED IN THIS STUDY WERE FOUND TO BE ANTICIPATED IN THE STUDY POPULATION AND ARE IN ALIGNMENT WITH RATES OBSERVED IN OTHER CLINICAL STUDIES. NO NEW RISKS HAVE BEEN IDENTIFIED FOR THIS POPULATION. NO NEW HAZARDS OR HARMS WERE IDENTIFIED THAT WOULD IMPACT THE BENEFIT/RISK PROFILE. BASED ON A REVIEW OF THE AVAILABLE INFORMATION, THERE WAS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. THEREFORE, NO REMEDIAL ACTION IS REQUIRED AT THIS TIME. D4. PRIMARY UDI NUMBER: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22679862&pc=NGV | 2135147-2025-04226 | 2023-01-30 05:00 | Death | ABBOTT MEDICAL | 2025-07-31 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P200049 | GEN2201026 | 4 | Event Description: IT WAS REPORTED THROUGH AMULET LAAO REGISTRY DATA THAT AMPLATZER AMULET DEVICES MAY BE RELATED TO 4 STROKE WHICH ARE CONSIDERED DEATH EVENTS. THE RELATIONSHIP OF THE DEATH TO THE AMPLATZER AMULET DEVICES COULD NOT BE DETERMINED BASED ON THE LIMITED DATA RECEIVED FROM THE REGISTRY. AMULET LAAO REGISTRY DATA IS REPORTED AS A SUMMARY PER SUMMARY REPORTING EXEMPTION APPROVAL NUMBER - GEN2201026. THE DATA RECEIVED INDICATED THAT THE PROCEDURE DATE RANGE WAS BETWEEN 30JANUARY2023 - 19OCTOBER2023. PATIENTS¿ MEAN AGE IS 75 YEARS, RANGING FROM 71 - 78 YEARS. 75% OF THE PATIENTS WERE MALE, 25% OF THE PATIENTS WERE FEMALE. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025. AS OF ON (B)(6) 2025, 12,180 PATIENTS WERE TREATED WITH THE AMULET DEVICE. Manufacturer Narrative: 4 EVENTS OF DEATH DUE TO STROKE WERE REPORTED IN THE CURRENT REPORTING PERIOD. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS LIMITED INFORMATION WAS PROVIDED INCLUDING THAT THE DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE LIMITED INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. HOWEVER, RISK ASSESSMENT WAS COMPLETED FOR THE REPORTED INCIDENTS. 17 OUT OF 12,180 (0.14%) SUBJECTS EXPERIENCED DEATH DUE TO STROKE SINCE THE BEGINNING OF THE REGISTRY. THE DEATH EVENTS AND RATES REPORTED IN THIS STUDY WERE FOUND TO BE ANTICIPATED IN THE STUDY POPULATION AND ARE IN ALIGNMENT WITH RATES OBSERVED IN OTHER CLINICAL STUDIES. NO NEW RISKS HAVE BEEN IDENTIFIED FOR THIS POPULATION. NO NEW HAZARDS OR HARMS WERE IDENTIFIED THAT WOULD IMPACT THE BENEFIT/RISK PROFILE. BASED ON A REVIEW OF THE AVAILABLE INFORMATION, THERE WAS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. THEREFORE, NO REMEDIAL ACTION IS REQUIRED AT THIS TIME. B2: DATE OF DEATH: THE EARLIEST DATE OF DEATH WAS USED. D4: PRIMARY UDI NUMBER: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A: IMPLANT DATE: THE EARLIEST IMPLANT DATE WAS USED. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22679867&pc=OZD | 1220648-2025-45705 | 2025-06-28 04:00 | Death | ABIOMED, INC. | 2025-07-31 04:00 | OZD | IMPELLA 5.5 | Malposition of Device; Device Sensing Problem | Bradycardia; Hemorrhage/Bleeding; Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT WAS ON AGGRASTAT FOR PERCUTANEOUS CORONARY INTERVENTION AND THEREFORE, HEMOSTASIS TOOK A LITTLE LONGER TO ACHIEVE. A JACKSON-PRATT BULB WAS PLACED TO ASSESS FOR BLEEDING AND A UNIT OF PACKED RED BLOOD CELLS (PRBC) WERE ADMINISTERED. THE PATIENT WAS TRANSFERRED TO THE UNIT ON SUPPORT AND A TIME THEREAFTER, THE PATIENT WAS ADMINISTERED TWO ADDITIONAL UNITS OF PRBCS WITH PLATELETS, CRYOPRECIPITATE, AND DESMOPRESSIN WHICH ACHIEVED HEMOSTASIS WITH ONLY 120ML OF DRAIN OUTPUT. THE PATIENT, PREVIOUSLY PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AND CONTINUOUS RENAL REPLACEMENT THERAPY, CONTINUED ON IMPELLA MCS. APPROXIMATELY FIVE DAYS LATER, THE PATIENT WAS SLOWING IMPROVING AND WAS TAKEN OFF ECMO. THE PATIENT WAS STILL ON DOBUTAMINE BUT ALL OTHER INOPRESSORS WERE OFF. TWO DAYS LATER, IT WAS NOTED THE PATIENT REMAINED INTUBATED AND OFF SEDATION WITH INCONSISTENT RESPONSES. FOLEY CATHETER WAS REMOVED; PATIENT REMAINED ON CRRT. IT WAS FURTHER NOTED A NASAL GASTRIC TUBE WAS PLACED AND CAUSED SIGNIFICANT NOSEBLEED REQUIRING TWO UNITS OF PRBCS FOR A HEMOGLOBIN OF 6.7. DOBUTAMINE WAS ATTEMPTED TO BE WEANED; HOWEVER, THE PATIENT WAS HYPOTENSIVE. AGGRASTAT WAS CHANGED OUT FOR BRILINTA AND HEPARIN WAS ON HOLD DUE TO THE NOSEBLEED. THE FOLLOWING DAY, THE PATIENT STARTED COUGHING AND BRADY DOWN AND THEN VENTRICULAR TACHYCARDIA ARRESTED. CARDIOPULMONARY RESUSCITATION WAS ADMINISTERED FOR TWO MINUTES WITH 100MG OF LIDOCAINE AND DEFIBRILLATOR X200 JOULES. THE PATIENT WAS CONVERTED BACK TO NORMAL SINUS RHYTHM. LATER THIS DAY, THE PHYSICIAN ENCOUNTERED DIFFICULTY WITH THE LEFT VENTRICULAR (LV) PLACEMENT SIGNAL AND THE DEVICE¿S POSITION. THE ADJUST LV SIGNAL WAS REPERFORMED MULTIPLE TIMES BY THE PHYSICIAN MAKING HIS OWN ADJUSTMENTS. THE CARE TEAM INFORMED OF POSSIBLE ISSUES RELATED TO THE ADJUSTMENTS AND NEED TO MONITOR PATIENT HEMODYNAMICS AND IMAGING FOR POSITION. TWO DAYS LATER, A TEMPORARY PACER CATHETER WAS PLACED DUE TO EPISODES OF BRADYCARDIA AND EXTENDED PAUSES IN RHYTHM. THE PATIENT EXPERIENCED A SHORT RUN OF VENTRICULAR TACHYCARDIA WITH NO INTERVENTION REQUIRED. HOWEVER, IT WAS NOTED THE PATIENT WOULD DEVELOP ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE. DAYS FOLLOWING, THE PATIENT WAS NOTED TO BE DOING WELL. THE PATIENT WAS EXTUBATED, SITTING UP AND WORKED ON STRENGTHENING ACTIVITIES. HOWEVER LATER THAT NIGHT, THE PATIENT ASPIRATED AFTER VOMITING AND REQUIRED EMERGENT REINTUBATION. THE PATIENT WOULD CONTINUE IMPELLA 5.5 MCS AND NOW A COUPLE OF DAYS LATER, THE PATIENT DEVELOPED A ¿MASSIVE¿ GASTROINTESTINAL BLEED (MALLORY-WEISS TEAR). HEPARIN WAS STOPPED. THE GASTROINTESTINAL BLEED CAUSED THE PATIENT TO GO INTO CARDIAC ARREST, AND THE DECISION WAS MADE BY THE FAMILY TO WITHDRAW CARE. THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22680039&pc=LWR | 2135147-2025-04230 | 2025-07-06 04:00 | Death | ABBOTT MEDICAL | 2025-07-31 04:00 | LWR | EPIC¿ SUPRA VALVE (AORTIC) | Material Puncture/Hole | Insufficient Information | P040021 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 23MM EPIC VALVE WAS SELECTED FOR IMPLANT. AFTER THE VALVE WAS IMPLANTED AND THE CHEST WAS CLOSED, IT WAS NOTED THAT THERE WAS A SMALL HOLE ON THE LEAFLET. THE DEVICE WAS REMOVED AND REPLACED WITH A 21MM EPIC VALVE TO RESOLVE THE EVENT. THE PATIENT REMAINED HEMODYNAMICALLY STABLE THROUGHOUT THE PROCEDURE AND THERE WAS NO CLINICALLY SIGNIFICANT DELAY. THE PATIENT WAS NOTED TO HAVE SINCE PASSED AWAY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22680322&pc=NPT | 2135147-2025-04232 | 2025-07-08 04:00 | Death | ABBOTT MEDICAL | 2025-07-31 04:00 | NPT | NAVITOR VISION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM NAVITOR VALVE WAS SELECTED FOR IMPLANT. THE PATIENT HAD A RELEVANT PAST MEDICAL HISTORY OF SEVERE AORTIC STENOSIS, SEVERE MITRAL REGURGITATION, SEVERE TRICUSPID REGURGITATION, EF20%, AND SEVERE PULMONARY HYPERTENSION. PRIOR TO PROCEDURE, THE PATIENT WAS IN STABLE CONDITION. AFTER BALLOON VALVULOPLASTY (BVP) WAS PERFORM WITH AN 18MM OSYPKA BALLOON, THE PATIENT COLLAPSED DUE TO HYPOTENSION. FAST IMPLANTATION OF THE NAVITOR VALVE WAS PERFORMED WHILE THE PATIENT WAS UNDER CARDIOPULMONARY RESUSCITATION (CPR). THE VALVE WAS IMPLANT WAS AT A GOOD HEIGHT WITH NO GRADIENT AND NO PARAVALVULAR LEAK; IT WAS CONSIDERED SUCCESSFULLY IMPLANTED. HOWEVER, THE PATIENT PASSED AWAY DUE TO "PUMP FAILURE." Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22680408&pc=QBJ | 3004753838-2025-205494 | 2024-06-23 04:00 | Death | DEXCOM, INC. | 2025-07-31 04:00 | QBJ | DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | Appropriate Term/Code Not Available | Vomiting; Diabetic Ketoacidosis | K213919 |  | 1 | Event Description: IT WAS REPORTED VIA MAUDE (MW5173056) AND FOLLOW UP CONTACT WITH THE PATIENT'S PARENT THAT THE PATIENT PASSED AWAY. THE PATIENT'S PARENT NOTED THAT SHE REPORTED THIS VIA MAUDE AFTER RECEIVING THE RECEIVER FIELD ACTION LETTER IN THE MAIL AND NOTING THAT THE SERIAL NUMBER ON THE LETTER MATCHED THE PATIENT'S RECEIVER. THE PATIENT'S PARENT REPORTED THAT THE PATIENT USED AN INSULIN PUMP BUT DID NOT KNOW WHICH ONE AND BELIEVED IT TO HAVE BEEN IN MODIFIED CLOSED LOOP. USE OF AN INSULIN PUMP DURING THE EVENT IS NOT CONFIRMED AND NO FURTHER DETAILS WERE AVAILABLE. THE PATIENT HAD REPORTEDLY BEEN SICK AND VOMITING. HE WAS NOTED TO HAVE BEEN ALONE IN HIS APARTMENT WHEN HE WAS FOUND AFTER BEING DECEASED 3 DAYS. THE PARENT REPORTED FINDING THE DEXCOM RECEIVER NEXT TO HIM, BUT COULD NOT RECALL ANY CGM READINGS, ALERTS, OR ALARMS FROM THE DEVICE AT THAT TIME. AN AUTOPSY NOTED THAT HE HAD DIED FROM DIABETIC KETOACIDOSIS (DKA). THE PATIENT HAD BEEN EVALUATED AND TREATED IN THE EMERGENCY DEPARTMENT FOR SYMPTOMATIC HYPERGLYCEMIA DUE TO INACCURACIES ON THE CGM APPROXIMATELY 8 MONTHS PRIOR TO HIS DEATH. THIS PREVIOUS ISSUE COUPLED WITH THE LETTER SHE RECEIVED MADE THE PARENT SUSPICIOUS THAT THE PATIENT'S DEATH WAS THE RESULT OF A DEXCOM MALFUNCTION, ALTHOUGH SHE COULD NOT SPECIFY WHAT SHE BELIEVED THE PROBLEM TO BE. THE PATIENT'S PARENT NOTED THAT SHE HAD THE PATIENT'S RECEIVER IN HER POSSESSION AND WOULD CONSIDER RETURNING IT FOR INVESTIGATION. DEXCOM HAS SENT THE PATIENT'S MOTHER A RETURN GOODS AUTHORIZATION TO SEND PRODUCT IN FOR INVESTIGATION. IF PRODUCT IS RETURNED, A FOLLOW UP REPORT WILL BE SUBMITTED UPON INVESTIGATION COMPLETION. NO PRODUCT OR DATA WAS PROVIDED FOR INVESTIGATION. PROBLEM COULD NOT BE CONFIRMED AND PROBABLE CAUSE CANNOT BE DETERMINED. NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22680502&pc=OZD | 1220648-2025-45709 | 2025-07-27 04:00 | Death | ABIOMED, INC. | 2025-07-31 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia; Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. IT WAS REPORTED WHILE ON IMPELLA CP SUPPORT, THE PATIENT EXPERIENCED SEVERE LIMB ISCHEMIA IN THE LEFT LEG WHERE THE IMPELLA CP WAS INSERTED. A VASCULAR SURGEON PLACED A 6 FRENCH DISTAL REPERFUSION SHEATH AND FASCIOTOMY AS PERFORMED. THE PATIENT WAS THEN SENT TO THE ICU IN CRITICAL CONDITION. IT WAS ADDITIONALLY REPORTED THAT THE IMPELLA WAS FOUND TO BE SLIGHTLY DEEP AT 5.7 CM, HOWEVER, THE PHYSICIAN DID NOT WANT TO REPOSITION AS THE PATIENT EXPERIENCED LONG RUNS OF VENTRICULAR TACHYCARDIA WHEN THE PUMP HAD BEEN REPOSITIONED IN THE CARDIAC CATHETERIZATION LABORATORY. ON (B)(6) 2025 THE DECISION WAS MADE TO ESCALATE THE PATIENT TO AN IMPELLA 5.5 FOR INCREASED SUPPORT. THE PATIENT WAS SUCCESSFULLY ESCALATED TO THE IMPELLA 5.5 AND WAS NOTED TO BE IN STABLE CONDITION FOLLOWING IMPLANT OF THE IMPELLA 5.5, HOWEVER THE FOLLOWING DAY THE DECISION WAS MADE TO WITHDRAW CARE. THE PATIENT¿S CARE WAS WITHDRAWN, AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22680668&pc=MFA | 3007284006-2025-00155 | 2025-07-22 04:00 | Death | SPECTRANETICS CORPORATION | 2025-07-31 04:00 | MFA | GLIDELIGHT LASER SHEATH | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation | P960042 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) AND A RIGHT ATRIAL (RA) LEAD DUE TO NON-FUNCTION. SPECTRANETICS LLD EZ LEAD LOCKING DEVICES (LLD EZS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. USING MULTIPLE SPECTRANETICS DEVICES (14F GLIDELIGHT LASER SHEATH, 11F TIGHTRAIL SUB-C ROTATING DILATOR SHEATH, AND 16F GLIDELIGHT LASER SHEATH), THE RA LEAD WAS REMOVED. SWITCHING TO THE RV LEAD WITH THE 16F GLIDELIGHT, THE LEAD WAS REMOVED; HOWEVER, THE PATIENT'S BLOOD PRESSURE DROPPED. RESCUE EFFORTS BEGAN, INCLUDING RESCUE BALLOON AND STERNOTOMY. A PERFORATION IN THE INNOMINATE VEIN WAS DISCOVERED AND REPAIRED, BUT THE PATIENT WOULD NOT STABILIZE, AND DID NOT SURVIVE THE PROCEDURE. THIS REPORT CAPTURES THE 16F GLIDELIGHT IN USE WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION, RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICE IN USE DURING THE PROCEDURE. Manufacturer Narrative: A2) PATIENT DATE OF BIRTH - NOT AVAILABLE FROM FACILITY. A4) PATIENT WEIGHT - NOT AVAILABLE FROM FACILITY. A5/A6) PATIENT ETHNICITY AND RACE - NOT AVAILABLE FROM FACILITY. B6/B7) PATIENT INFORMATION REGARDING RELEVANT TESTS/LABORATORY DATA OR MEDICAL HISTORY - NOT AVAILABLE FROM FACILITY. D4/H4) THE LOT NUMBER WAS NOT PROVIDED; THUS THE FOLLOWING INFORMATION IS UNKNOWN: DEVICE SERIAL NUMBER, UNIQUE ID, EXPIRATION DATE, AND MANUFACTURE DATE. H3) THE GLIDELIGHT WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) GREAT VESSEL PERFORATION AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE GLIDELIGHT DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22680819&pc=NGV | 2124215-2025-50311 | 2025-07-23 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-31 04:00 | NGV | SYSTEM, APPENDAGE CLOSURE, LEFT ATRIAL | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED VIA JOURNAL ARTICLE THAT DEATHS OCCURRED. THE FOLLOWING EVENT WAS OBTAINED VIA A RETROSPECTIVE ARTICLE OF THREE (3) YEAR OUTCOMES FROM FOLLOWING 220 PATIENTS WHO UNDERWENT A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WHERE EITHER A NON-BOSTON SCIENTIFIC LAA CLOSURE DEVICE, A WATCHMAN CLOSURE DEVICE, OR A WATCHMAN FLX CLOSURE DEVICE WAS IMPLANTED ACROSS EIGHT (8) CENTERS. THIS WAS A RANDOMIZED CONTROL TRIAL. IT WAS REPORTED THE FOLLOWING OCCURRED AND RESULTS WERE COMPILED AT THE THREE (3) YEAR CLINICAL FOLLOW UP: PATIENT CARDIOVASCULAR DEATH. LITERATURE CITATION: GALEA, R. ET AL 2025) 3-YEAR CLINICAL OUTCOMES COMPARING THE AMULET VS WATCHMAN FLX FOR LEFT ATRIAL APPENDAGE CLOSURE IN PATIENTS WITH ATRIAL FIBRILLATION. 86 (1). JOURNAL OF THE AMERICAN COLLEGE OF CARDIOLOGY. Manufacturer Narrative: B3: USED BSC AWARE DATE AS EVENT DATE AS IT IS UNKNOWN. D1: BRAND NAME IS BLANK BECAUSE THE DEVICE IS KNOWN TO BE A WATCHMAN LAA CLOSURE DEVICE, BUT THE EXACT BRAND NAME IS UNKNOWN. D4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22681327&pc=DSQ | 2916596-2025-04823 | 2025-03-19 04:00 | Death | THORATEC CORPORATION | 2025-07-31 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FROM CEREBROVASCULAR ACCIDENT (CVA). Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22681362&pc=DXF | 2124215-2025-53097 | 2025-07-09 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-07-31 04:00 | DXF | VERSACROSS CONNECT LAAC ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Pericardial Effusion | K150709 |  | 1 | Event Description: IT WAS REPORTED THAT PERICARDIAL EFFUSION (PE), CARDIAC TAMPONADE, AND DEATH OCCURRED. THE PROCEDURE WAS CANCELLED. DURING A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A VERSACROSS CONNECT LAAC WAS SELECTED FOR USE. INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING NOTED THERE WAS NO BASELINE PE. GROIN ACCESS OBTAINED AND NON-BOSTON SCIENTIFIC VASCULAR PRE-CLOSURE DEVICE WAS USED. A VERSACROSS RF WIRE WAS INSERTED UP TO THE SUPERIOR VENA CAVA (SVC), A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND VERSACROSS DILATOR TO FOLLOW. THE TRANSSEPTAL PUNCTURE (TSP) WAS PERFORMED WITHOUT RF DELIVERY REQUIRED. TSP LOCATION APPEARED TO BE MID/MID-POSTERIOR BASED ON ICE IMAGING. NO PE WAS NOTED AT THIS POINT IN THE PROCEDURE AND PATIENT VITALS WERE STABLE. TSP LOCATION WAS DILATED WITH THE WAS AND THEN THE SHEATH AND DILATOR WERE RETRACTED FOR LEFT ATRIUM (LA) ACCESS OF THE ICE CATHETER. TSP ACCESS OF ICE CATHETER WAS STILL UNSUCCESSFUL. A PE WAS THEN DISCOVERED FOUR (4) MINUTES LATER. WAS REMOVED FROM THE GROIN AND ICE VISUALIZED THE PE LOCATION ON THE RIGHT SIDE OF THE HEART. THE PROCEDURE WAS ABORTED. CARDIAC TAMPONADE OCCURRED, WHICH CAUSED THE NEED FOR CARDIOPULMONARY RESUSCITATION (CPR), BLOOD TRANSFUSION, PERICARDIOCENTESIS REMOVING BRIGHT RED BLOOD OF OVER THREE (3) LITERS, FRESH FROZEN PLASMA (FFP), AND CRYOPRECIPITATE. THE FAMILY WAS CONTACTED A LITTLE OVER TWO (2) HOURS LATER AND AS THE PATIENT WAS A DO NOT RESUSCITATE (DNR) STATUS AND THEY REQUESTED NO ADDITIONAL LIFE-SAVING MEASURES, SO TIME OF DEATH WAS CALLED BY THE PHYSICIAN. IT WAS NOTED THE PATIENT WAS ON ELIQUIS AND STOPPED TAKING ELIQUIS TWO (2) DAYS PRIOR TO THE PROCEDURE. LAST RECORDED INR LAB RESULT WAS 1.4. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22681373&pc=DXF | 2124215-2025-53130 | 2025-06-11 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-07-31 04:00 | DXF | VERSACROSS ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Loss of consciousness; Ischemia Stroke | K150709 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED STROKE WHILE IN RECOVERY. THE PATIENT PASSED AWAY. DURING A PULSE FIELD ABLATION PROCEDURE, A VERSACROSS ACCESS SOLUTION WAS SELECTED FOR USE. A CLOT WAS RULED OUT PRIOR TO THE PROCEDURE, BY TEE IMAGE PERFORMED. TRANSEPTAL PUNCTURE (TSP) WAS DONE USING VERSACROSS RF WIRE TO GAIN ACCESS TO THE LEFT ATRIUM, WITH NO ISSUES REPORTED. SUCCESSFUL PULSE FIELD ABLATION DONE TO ALL FOUR PULMONARY VEINS. THE PHYSICIAN WAS COMPLETED WITHOUT ANY ISSUE FROM THE PHYSICIAN. DURING POST-PROCEDURE, WHILE IN RECOVERY, THE PATIENT HAD A STROKE, CONFIRMED BY A CT SCAN. ON THE DAY OF THE PROCEDURE, ACTIVATED CLOTTING TIME (ACT) WAS MAINTAINED ABOVE 300 SECONDS WITH HEPARIN ADMINISTERED PRIOR TO THE TRANSSEPTAL PUNCTURE. LEFT ATRIAL DWELL TIME WAS KEPT UNDER 15 MINUTES. PULMONARY VEIN ISOLATION (PVI) WAS PERFORMED EXCLUSIVELY, REQUIRING A TOTAL OF 48 APPLICATIONS. FOLLOWING PVI, A CAVOTRICUSPID ISTHMUS (CTI) LINE WAS CREATED USING A RADIOFREQUENCY (RF) CATHETER. THE ONLY ADDITIONAL CATHETER PLACED WAS IN THE CORONARY SINUS. POST-PROCEDURE, THE PATIENT EXPERIENCED DIFFICULTY AWAKENING FROM ANESTHESIA AND WAS CHALLENGING TO EXTUBATE. A CT SCAN FOLLOWED BY AN MRI REVEALED MULTIPLE CEREBRAL INFARCTS CONSISTENT WITH AN EMBOLIC SHOWER. THE PATIENT DID NOT REGAIN CONSCIOUSNESS FOLLOWING GENERAL ANESTHESIA. NO TREATMENT WAS ADMINISTERED BEYOND INTRACRANIAL PRESSURE (ICP) RELIEF, AS THE STROKE INVOLVED MULTIPLE AREAS WITH ASSOCIATED MASS EFFECTS. THE PATIENT HAS SINCE BEEN TRANSITIONED TO COMFORT CARE, WITH NO EXPECTATION OF RECOVERY. THE PHYSICIAN WAS UNABLE TO DETERMINE THE CAUSE OF THE STROKE. THE TEE WAS REVIEWED AND THOROUGHLY EXAMINED, SHOWING NO EVIDENCE OF A LEFT ATRIAL APPENDAGE (LAA) THROMBUS. IT IS HIGHLY UNLIKELY THAT THE CEREBROVASCULAR ACCIDENT (CVA) WAS DUE TO A MISSED THROMBUS, AS THE INFARCTS WERE DISTRIBUTED THROUGHOUT THE BRAIN. THE PATIENT PASSED AWAY. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22682167&pc=NPT | 2135147-2025-04234 | 2023-06-08 04:00 | Death | ABBOTT MEDICAL | 2025-07-31 04:00 | NPT | PORTICO TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemorrhage/Bleeding; Unspecified Infection; Myocardial Infarction; Renal Failure; Brain Injury; Heart Failure/Congestive Heart Failure; Respiratory Insufficiency; Insufficient Information | P190023 | RWD2300134 | 99 | Event Description: IT WAS REPORTED THROUGH STS/ACC TVT REGISTRY DATA THAT PORTICO/NAVITOR DEVICES MAY BE RELATED TO 99 ADVERSE EVENTS WHICH ARE CONSIDERED DEATH INCLUDING: HEART FAILURE, SUDDEN CARDIAC DEATH, OTHER NON-CARDIOVASCULAR REASON, UNKNOWN, INFECTION, OTHER CARDIOVASCULAR REASON, CARDIOVASCULAR PROCEDURE, PULMONARY, HEMORRHAGE, NEUROLOGICAL, STROKE, ACUTE MYOCARDIAL INFARCTION, TRAUMA, HEPATOBILIARY, RENAL. THE RELATIONSHIP OF THE DEATH TO THE PORTICO/NAVITOR DEVICES COULD NOT BE DETERMINED BASED ON THE LIMITED DATA RECEIVED FROM THE REGISTRY. STS/ACC TVT REGISTRY DATA IS REPORTED AS A SUMMARY PER SUMMARY REPORTING EXEMPTION APPROVAL NUMBER - RWD2300134. THE DATA RECEIVED INDICATED THAT THE PROCEDURE DATE RANGE WAS BETWEEN 08JUNE2023 ¿ 13DECEMBER2024. PATIENTS¿ MEAN AGE IS 83 YEARS, RANGING FROM 46 TO (B)(6) YEARS. 57% OF THE PATIENTS WERE MALE, 43% OF THE PATIENTS WERE FEMALE. THIS REPORT IS FOR QUARTER 2 TIMEFRAME BETWEEN 01 APRIL 2025 - 30 JUNE 2025. AS OF 16 MAY 2025, 3396 PATIENTS WERE TREATED WITH THE NAVITOR DEVICE AND 584 WITH THE PORTICO DEVICE IN THE STS/ACC TVT REGISTRY.Event Description: Unique Complaint ID Number,Event Date,Date Entered,Manufacturer Aware Date,Brand Name,Generic Name,Model Number,Lot Number,Catalog Number,Serial Number,UDI Number,PMA / 510K Number,Date Returned to Manufacturer,Type of Reportable Event,Event Description,Manufacturer Narrative,Medical History,Patient Age,Patient Gender,Patient Weight,Date Implanted,Date Explanted,Health Effect Clinical Code,Health Effect Impact Code,Device Problem Code,Device Component Code,Investigation Type Code,Investigation Findings Code,Investigation Conclusion Code,Remedial Action Type,Latest Line Item Version
<br>42110917,3/11/2024,7/31/2025,5/2/2025,Navitor Titan,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor Titan was implanted on 02/07/2024 in a 63 year old Male. On 03/11/2024, patient death was reported due to Acute myocardial infarction. The relationship of the Death to the Navitor Titan could not be determined based on the limited data received from the registry.","An event of death due to Acute myocardial infarction was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. Five out of 3396 (0.1%) subjects experienced death due to Acute myocardial infarction since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,63,Male,65.8,2/7/2024,NI,1969,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>43091966,11/12/2024,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 11/11/2024 in a 77 year old Female. On 11/12/2024, patient death was reported due to Acute myocardial infarction. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Acute myocardial infarction was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. Five out of 3396 (0.1%) subjects experienced death due to Acute myocardial infarction since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,77,Female,118,11/11/2024,NI,1969,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>42369915,7/3/2024,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 07/03/2024 in a 88 year old Female. On 07/03/2024, patient death was reported due to Cardiovascular procedure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Cardiovascular procedure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 10 out of 3396 (0.3%) subjects experienced death due to Cardiovascular procedure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,88,Female,47.6,7/3/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>42535020,8/16/2024,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 08/16/2024 in a 79 year old Female. On 08/16/2024, patient death was reported due to Cardiovascular procedure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Cardiovascular procedure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 10 out of 3396 (0.3%) subjects experienced death due to Cardiovascular procedure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,79,Female,96.1,8/16/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>43032776,11/18/2024,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 11/18/2024 in a 94 year old Male. On 11/18/2024, patient death was reported due to Cardiovascular procedure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Cardiovascular procedure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 10 out of 3396 (0.3%) subjects experienced death due to Cardiovascular procedure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,94,Male,83.1,11/18/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>42887473,11/22/2024,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 11/20/2024 in a 87 year old Male. On 11/22/2024, patient death was reported due to Cardiovascular procedure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Cardiovascular procedure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 10 out of 3396 (0.3%) subjects experienced death due to Cardiovascular procedure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,87,Male,66.8,11/20/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>43350155,7/17/2023,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 07/07/2023 in a 83 year old Female. On 07/17/2023, patient death was reported due to Heart failure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 19 out of 3396 (0.6%) subjects experienced death due to Heart failure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,83,Female,96.8,7/7/2023,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41195315,2/28/2024,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 08/07/2023 in a 90 year old Male. On 02/28/2024, patient death was reported due to Heart failure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 19 out of 3396 (0.6%) subjects experienced death due to Heart failure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,90,Male,87.2,8/7/2023,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>42016018,12/21/2024,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 12/11/2023 in a 75 year old Female. On 12/21/2024, patient death was reported due to Heart failure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 19 out of 3396 (0.6%) subjects experienced death due to Heart failure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,75,Female,112,12/11/2023,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>42403339,3/7/2025,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 05/07/2024 in a 77 year old Male. On 03/07/2025, patient death was reported due to Heart failure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 19 out of 3396 (0.6%) subjects experienced death due to Heart failure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,77,Male,83.7,5/7/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>42998300,10/7/2024,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 08/22/2024 in a 87 year old Female. On 10/07/2024, patient death was reported due to Heart failure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 19 out of 3396 (0.6%) subjects experienced death due to Heart failure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,87,Female,60.9,8/22/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>43071813,10/16/2024,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 09/17/2024 in a 92 year old Female. On 10/16/2024, patient death was reported due to Heart failure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 19 out of 3396 (0.6%) subjects experienced death due to Heart failure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,92,Female,56.7,9/17/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>42776705,10/5/2024,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 10/02/2024 in a 91 year old Female. On 10/05/2024, patient death was reported due to Heart failure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 19 out of 3396 (0.6%) subjects experienced death due to Heart failure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,91,Female,61.9,10/2/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>43271944,12/16/2024,7/31/2025,5/2/2025,Navitor Titan,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor Titan was implanted on 11/20/2024 in a 81 year old Male. On 12/16/2024, patient death was reported due to Heart failure. The relationship of the Death to the Navitor Titan could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 19 out of 3396 (0.6%) subjects experienced death due to Heart failure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,81,Male,68.7,11/20/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>42975517,12/28/2024,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 12/04/2024 in a 78 year old Male. On 12/28/2024, patient death was reported due to Heart failure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 19 out of 3396 (0.6%) subjects experienced death due to Heart failure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,78,Male,74.8,12/4/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>43371506,1/18/2025,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 12/05/2024 in a 81 year old Female. On 01/18/2025, patient death was reported due to Heart failure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 19 out of 3396 (0.6%) subjects experienced death due to Heart failure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,81,Female,79.7,12/5/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>43005109,12/13/2024,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 12/09/2024 in a 91 year old Male. On 12/13/2024, patient death was reported due to Heart failure. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Heart failure was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. 19 out of 3396 (0.6%) subjects experienced death due to Heart failure since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,91,Male,91,12/9/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41378369,3/24/2024,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor was implanted on 06/09/2023 in a 95 year old Female. On 03/24/2024, patient death was reported due to Hemorrhage. The relationship of the Death to the Navitor could not be determined based on the limited data received from the registry.","An event of death due to Hemorrhage was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. Two out of 3396 (<0.1%) subjects experienced death due to Hemorrhage since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,95,Female,71.4,6/9/2023,NI,1888,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>43034863,12/3/2024,7/31/2025,5/2/2025,Navitor Titan,"Aortic Valve, Prosthesis, Percutaneously Delivered",NA,Unknown,UNK NAVITOR,NI,NI,P190023,NA,D,"It was reported through Portico TVT Registry data that a Navitor Titan was implanted on 11/15/2024 in a 75 year old Male. On 12/03/2024, patient death was reported due to Hepatobiliary. The relationship of the Death to the Navitor Titan could not be determined based on the limited data received from the registry.","An event of death due to Hepatobiliary was reported. A more comprehensive assessment could not be performed as limited information was provided including that the device was not returned for analysis. Based on the limited information received, the cause of the reported incident could not be conclusively determined. A risk assessment was completed for the reported incident. One out of 3396 (<0.1%) subjects experienced death due to Hepatobiliary since the beginning of the registry. The death events and rates reported in this study were found to be anticipated in the study population and are in alignment with rates observed in other clinical studies.<br><br>No new risks have been identified for this population. No new hazards or harms were identified that would impact the benefit/risk profile. Based on a review of the available information, there was no indication of a product quality issue with regards to manufacture, design, or labeling. Therefore, no remedial action is required at this time. Therefore, no remedial action is required at this time.",No Information,75,Male,77.9,11/15/2024,NI,4446,1802,2993,4755,4114;4119,3221,4315,NA,0
<br>41568886,1/30/2024,7/31/2025,5/2/2025,Navitor,"Aortic Valve, Prosthesis, Percutaneously ... Edited for length. See MAUDE Adverse Event Report for complete event narratives.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22682695&pc=DSQ | 2916596-2025-05065 | 2023-03-10 05:00 | Death | THORATEC CORPORATION | 2025-07-31 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS MULTI SYSTEM ORGAN FAILURE. THE PUMP WAS WORKING AS EXPECTED. THERE WERE NO PUMP RELATED ISSUES THAT WERE PROVIDED. THIS WAS NOT DEVICE OR THERAPY RELATED. THE PUMP WOULD NOT BE EXPLANTED OR RETURNED.Event Description: NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH AND MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION. NO FURTHER INFORMATION WAS PROIVDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22682972&pc=NAY | 2955842-2025-32534 | 2023-12-27 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-07-31 04:00 | NAY | DAVINCI XI | Insufficient Information | Failure to Anastomose; Sepsis; Bowel Perforation | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT A DA VINCI-ASSISTED RADICAL CYSTECTOMY WITH INTRACORPOREAL ILEAL CONDUIT CREATION, BILATERAL URETERAL STENT PLACEMENT, AND BILATERAL EXTENDED PELVIC LYMPH NODE DISSECTION WAS PERFORMED. THERE WERE NO INTRAOPERATIVE COMPLICATIONS. THE INITIAL POST-OPERATIVE COURSE WAS UNCOMPLICATED AND THE PATIENT WAS DISCHARGED ON POST-OPERATIVE DAY (POD) 5. THE PATIENT ULTIMATELY EXPIRED ON POD 50 DUE TO SUBSEQUENT COMPLICATIONS. THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT ON POD 8 WITH INTRA-ABDOMINAL FLUID COLLECTION AND SUBSEQUENTLY DEVELOPED SIGNS OF SEPTIC SHOCK AND ACIDOSIS AND WAS STABILIZED WITH VASOPRESSORS. DURING EMERGENT EXPLORATORY SURGERY 2 LITERS OF FRANK STOOL WAS FOUND IN THE ABDOMEN, AND A LARGE ABSCESS IN THE PELVIS WAS IRRIGATED AND SUCTIONED OUT. A 1 CM MID-ILEAL ANTI-MESENTERIC SMALL BOWEL PERFORATION WAS FOUND AND RESECTED. ADDITIONALLY, THE COMMON ENTEROTOMY AND DISTAL ILEAL-ILEAL ANASTOMOSIS STAPLE LINE WAS COMPLETELY DISRUPTED WITH FRANK LEAKAGE OF STOOL; THE AREA WAS RESECTED (20CM TOTAL) AND REPAIRED. THE PATIENT REMAINED CRITICALLY ILL AFTER THIS PROCEDURE. A SUBSEQUENT EXPLORATORY LAPAROTOMY ON POD 10 OCCURRED AND PER THE OPERATIVE NOTES, A RIGHT COLECTOMY, ABDOMINAL WASHOUT, ILEOCOLIC ANASTOMOSIS, SMALL-BOWEL TO SMALL-BOWEL ANASTOMOSIS, PLACEMENT OF A WOUND VAC AND FASCIAL CLOSURE WERE COMPLETED. EXAMINATION OF THE ILEAL CONDUIT FROM THE ORIGINAL ROBOTIC PROCEDURE BY UROLOGY FOUND NO SIGNS OF LEAKAGE; THE ANASTOMOSES AT THE CONDUIT ITSELF AS WELL AS THE URETERS APPEARED VIABLE WITHOUT SIGNS OF NECROSIS OR ISCHEMIA. NOTES AT POD 33 DOCUMENT A POST-PROCEDURE COURSE THAT INCLUDED ACUTE RENAL FAILURE REQUIRING DIALYSIS, ACUTE HYPOXIC RESPIRATORY FAILURE REQUIRING VENTILATION, URINE LEAK REQUIRING BILATERAL NEPHROSTOMY TUBE/DRAIN PLACEMENT, SEPTIC SHOCK REQUIRING ONGOING VASOPRESSORS, MULTIPLE ANTIBIOTICS RELATED TO ANASTOMOTIC LEAK, FUNGEMIA, AND LOWER ABDOMINAL MIDLINE WOUND DEHISCENCE WITH PURULENT DRAINAGE. NO ADDITIONAL INFORMATION WAS PROVIDED UP TO THE PROVIDED DEATH CERTIFICATE AT POD 50 WHICH LISTED CARDIAC ARREST, SEPSIS AND ABDOMINAL INFECTION AS CAUSE OF DEATH. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE FIVE SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE EVENT. REVIEW FOUND THAT THE SYSTEM FUNCTIONED NORMALLY, NO INTRAOPERATIVE EVENTS OR ISSUES WERE OBSERVED. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, VESSEL SEALER EXTEND, MONOPOLAR CURVED SCISSORS, CADIERE FORCEPS, MARYLAND BIPOLAR FORCEPS, SUREFORM 60 STAPLER, LARGE NEEDLE DRIVER, LARGE SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO RELATED COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE FOUND NO NON-CONFORMANCE'S WERE IDENTIFIED TO BE RELATED TO THIS EVENT. THE STAPLER INSTRUMENT LOGS FOR THIS PROCEDURE FOUND A SUREFORM 60 STAPLER INSTRUMENT WAS USED TO FIRE 12 RELOADS WITHOUT ANY ISSUES OBSERVED IN THE LOGS. THE VESSEL SEALER INSTRUMENT LOGS FOUND ONE CUT FAILURE ON THE INSTRUMENT'S 15TH CUT ATTEMPT. THE LOGS INDICATE THERE WERE TWO TIMES WHERE THE VESSEL SEALER EXTEND JAWS WERE OPEN TOO MUCH TO INITIATE A CUT OF THE TISSUE WITHIN THE JAWS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A PRIOR HISTORY OF PROSTATE CANCER AND PRIOR PROSTATECTOMY WHO WAS UNDERGOING A CYSTECTOMY FOR BLADDER CANCER WITH CREATION OF AN ILEAL CONDUIT. THE PROCEDURE REPORTEDLY WENT WELL AND NO INTRAOPERATIVE ISSUES WERE NOTED. POST-OPERATIVELY, THE PATIENT DEVELOPED A SMALL BOWEL LEAK SEVERAL DAYS AFTER THE ROBOTIC PROCEDURE FROM THE ANASTOMOSIS WHERE THE ILEAL CONDUIT WAS HARVESTED. THE REASON FOR THE LEAK AND POTENTIAL FACTORS WHICH MAY HAVE LED TO THE LEAK IS NOT PROVIDED. THE PATIENT UNDERWENT MULTIPLE ADDITIONAL PROCEDURES AS NOTED ABOVE FOLLOWING THE INITIAL COMPLICATION AND EVENTUALLY DIED SEVERAL WEEKS LATER. NO ALLEGATION WAS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT FROM THE INITIAL PROCEDURE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22683648&pc=MFE | 2029214-2025-01735 | 2024-11-21 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-07-31 04:00 | MFE | ONYX | Adverse Event Without Identified Device or Use Problem | Hematoma |  |  | 1 | Event Description: LIU, J. , NI, W. , ZUO, Q. , YANG, H. , PENG, Y. , LIN, Z. , LI, Z. , WANG, J. , ZHEN, Y. , LUO, J. , LIN, Y. , CHEN, J. , HUA, X. , LU, H. , ZHONG, M. , LIU, M. , ZHANG, J. , WANG, Y. , WAN, J. , LI,; THE NEW ENGLAND JOURNAL OF MEDICINE; 2024; 391 (20), 1901-1912; MIDDLE MENINGEAL ARTERY EMBOLIZATION FOR NONACUTE SUBDURAL HEMATOMA; DOI: 10.1056/NEJMOA2401201 LITERATURE WAS REVIEWED REGARDING "MIDDLE MENINGEAL ARTERY EMBOLIZATION FOR NONACUTE SUBDURAL HEMATOMA." THE TIME FRAME OF THIS STUDY WAS WITHIN 90 DAYS AFTER RANDOMIZATION, WITH THE FINAL PROTOCOL COMPLETED ON SEPTEMBER 21, 2022, AND STATISTICAL ANALYSIS ON DECEMBER 2, 2023. THE MEDIAN FOLLOW-UP TIME WAS 91 DAYS. MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: ONYX LIQUID EMBOLIC MATERIAL, WHICH WAS PROVIDED FOR THE EMBOLIZATION PROCEDURE. 2 DEATHS OCCURRED IN THE STUDY POPULATION. THE CAUSES OF DEATH WERE: 1 HEMATOMA-RELATED, ANOTHER FOR ANOTHER REASON. AMONG PATIENT ADVERSE EVENTS INCLUDED: SYMPTOMATIC RECURRENCE OR PROGRESSION OF SUBDURAL HEMATOMA WAS REPORTED IN 24 OF 360 PATIENTS. 17 PATIENTS (4.7%) HAD SYMPTOMATIC HEMATOMA RECURRENCE. 7 (1.9%) HAD SYMPTOMATIC HEMATOMA PROGRESSION. REPEAT OR RESCUE SURGERY WAS PERFORMED WITHIN 90 DAYS IN 11 PATIENTS. 1 PATIENT HAD FACIAL-NERVE PARALYSIS. 2 PATIENTS HAD AN ALLERGIC REACTION TO THE CONTRAST AGENT. 29 ADVERSE EVENTS OF SPECIAL INTEREST WITHIN 90 DAYS- ASYMPTOMATIC INTRACRANIAL HEMORRHAGE, NON¿CENTRAL NERVOUS SYSTEM INFE CTION, PERIOPERATIVE INFECTION (UNDETERMINED CAUSE), ADVERSE DRUG REACTION (LIVER INJURY), ADVERSE DRUG REACTION (ALLERGY), EPILEPSY, POOR INCISION HEALING, DEEP VENOUS THROMBOSIS, NEW-ONSET ISCHEMIC STROKE, AND OTHER. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. A2 AGE IS AVERAGE AGE OF COHORT. A3B GENDER IS MAJORITY GENDER OF COHORT. SEE ATTACHMENTS FOR LITERATURE ARTICLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22661729&pc=DTZ | 8010762-2025-0000325 |  | Death | MAQUET CARDIOPULMONARY GMBH | 2025-07-30 04:00 | DTZ | TUBING SET | Adverse Event Without Identified Device or Use Problem; Air/Gas in Device | Insufficient Information | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN CHINA. THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT FOR TREATMENT DUE TO A CAR ACCIDENT. IT WAS REPORTED THAT THE HLS SET COULD NOT BE PRIMED DUE TO AIR BUBBLES IN THE OXYGENATOR. THIS HLS SET WAS NEVER CONNECTED TO THE PATIENT. AS THE FIRST SET COULD NOT BE USED IMMEDIATELY, THERE WAS A DELAY IN TREATMENT. A SECOND HLS SET WAS USED AND HAS BEEN CONNECTED THE PATIENT TO PROCESS FURTHER. THE PATIENT PASSED AWAY WHILE BEING CONNECTED TO THE SECOND HLS SET AFTER SEVERAL HOURS. AS THE PATIENT PASSED AWAY A REPORT IS REQUIRED. COMPLAINT ID# (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. THIS EVENT OCCURRED ON THE CHINESE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069076.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22662383&pc=HGM | 9610816-2025-000677 | 2021-12-27 05:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-07-30 04:00 | HGM | AVALON FETAL MONITOR FM30 | Incorrect Measurement | Hypoxia in Utero | K140535 |  | 1 | Event Description: IT WAS REPORTED THE MONITOR MEASURED THE MATERNAL HEART RATE INSTEAD OF THE FETAL HEART RATE ON TWO SEPARATE CTG READINGS AND THE FETUS WAS STILLBORN AT TERM. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. Manufacturer Narrative: THE MANUFACTURING DATE FOR THE REPORTED DEVICE IS PRIOR TO 24SEPT2016 SO NO UDI REQUIRED. THE REPORTER'S NAME AND EMAIL WERE PROVIDED; HOWEVER, DUE TO PRIVACY LAWS WERE NOT INCLUDED IN THE REPORT. E1 REPORTER PHONE#: (B)(6). PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22662798&pc=NIQ | 9612164-2025-03707 | 2025-07-06 04:00 | Death | MEDTRONIC IRELAND | 2025-07-30 04:00 | NIQ | ONYX FRONTIER | Device Damaged by Another Device | Gastrointestinal Hemorrhage | P160043 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING ONE ONYX FRONTIER CORONARY DRUG-ELUTING STENT, THE STENT WAS DAMAGED BY ANOTHER DEVICE. THE LESION BEING TREATED WAS MODERATELY TORTUOUS AND MODERATELY CALCIFIED LOCATED IN THE OSTIUM OF THE LEFT MAIN (LM) CORONARY ARTERY. THE STENT WAS PLACED AND SUCCESSFULLY DEPLOYED IN THE LM. ANOTHER STENT WAS THEN PLACED IN THE LEFT ANTERIOR DESCENDING A RTERY(LAD). ANOTHER STENT WAS PLACED IN THE RAMUS USING A NON-MEDTRONIC GUIDEWIRE. WHEN IT WAS ATTEMPTED TO RETRIEVE THE GUIDEWIRE, IT LOCKED UP WITH THE DEPLOYED STENT IN THE LM. THE GUIDEWIRE WAS PULLED AND THE STENT WAS EXPANDED AND REMOVED FROM THE PATIENT STUCK TO THE GUIDEWIRE. THE STENT DID NOT PASS THROUGH A PREVIOUSLY DEPLOYED STENT. NO RESISTANCE WAS ENCOUNTERED WHEN ADVANCING THE DEVICE AND NO EXCESSIVE FORCE WAS USED DURING DELIVERY. THE STENT WAS INSPECTED PRIOR TO USE AND NEGATIVE PREP WAS PERFORMED WITH NO ISSUES. THE PATIENT DIED THE FOLLOWING DAY. THE PHYSICIAN ASSESSED THAT THE DEATH WAS RELATED TO GASTROINTESTINAL (GI) BLEED AND WAS NOT DUE TO THE ISSUES ENCOUNTERED WITH THE STENT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22663097&pc=OZP | 2032227-2025-221944 | 2025-07-01 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-30 04:00 | OZP | MINIMED 770G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON 8TH JUN 2025 AT HOME. THE CAUSE OF DEATH WAS DIABETIC COMPLICATIONS. THE INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, UNOMEDICAL, MMT-1880. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING. IT WAS UNKNOWN WHETHER THE CUSTOMER HAD BEEN USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF REPORTED EVENT. THERE IS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A, UNOMEDICAL, MMT-1880. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22663329&pc=BZD | 2518422-2025-048635 | 2024-01-22 05:00 | Death | RESPIRONICS, INC. | 2025-07-30 04:00 | BZD | DREAMSTATION CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY CANCER DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22663433&pc=DSQ | 3007042319-2025-01621 |  | Death | HEARTWARE, INC. | 2025-07-30 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22663535&pc=DTZ | 3013876692-2025-00062 |  | Death | MAQUET CARDIOPULMONARY GMBH | 2025-07-30 04:00 | DTZ | TUBING SET | Adverse Event Without Identified Device or Use Problem; Air/Gas in Device | Insufficient Information | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN CHINA. THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT FOR TREATMENT DUE TO A CAR ACCIDENT. IT WAS REPORTED THAT THE HLS SET COULD NOT BE PRIMED DUE TO AIR BUBBLES IN THE OXYGENATOR. THIS HLS SET WAS NEVER CONNECTED TO THE PATIENT. AS THE FIRST SET COULD NOT BE USED IMMEDIATELY, THERE WAS A DELAY IN TREATMENT. A SECOND HLS SET WAS USED AND HAS BEEN CONNECTED THE PATIENT TO PROCESS FURTHER. THE PATIENT PASSED AWAY WHILE BEING CONNECTED TO THE SECOND HLS SET AFTER SEVERAL HOURS. AS THE PATIENT PASSED AWAY A REPORT IS REQUIRED. COMPLAINT ID # (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. THIS EVENT OCCURRED ON THE CHINESE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069076.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22663906&pc=OZD | 1220648-2025-30848 | 2025-07-15 04:00 | Death | ABIOMED, INC. | 2025-07-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT HAD A CP PUMP PLACED TO SUPPORT THE PATIENT ADMITTED IN WITH CARDIOGENIC SHOCK AND CARDIOMYOPATHY. DURING THE SUPPORT THERE WERE SIGNS OF HEMOLYSIS THE TEAM TRIED TO TROUBLESHOOT AWAY AND TREAT. THE DECISION WAS MADE TO WITHDRAW ALL CARE AND PATIENT EXPIRED AFTER 63 HOURS OF SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS: INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES: MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22664168&pc=OZD | 1220648-2025-30852 | 2025-07-15 04:00 | Death | ABIOMED, INC. | 2025-07-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT ON EXTRACORPOREAL MEMBRANE OXYGENATION WITH EJECTION FRACTION OF 10%, WAS PLACED ON IMPELLA CP. THE IMPELLA WAS INSERTED FROM THE LEFT SUBCLAVIAN AND EXTRACORPOREAL MEMBRANE OXYGENATION WITH IMPELLA (ECPELLA) WAS ADMINISTERED. DURING SUPPORT IT WAS REPORTED THAT THE PATIENT HAD A CEREBRAL HEMORRHAGE AND THAT DUE TO A POOR NEUROLOGICAL PROGNOSIS OF THE CEREBRAL HEMORRHAGE, IT WAS DECIDED NOT TO UNDERGO ADDITIONAL TREATMENT, AND WAS TREATED FOR MULTIPLE ORGAN FAILURE. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA CP WITH SMART ASSIST SYSTEM & IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22664384&pc=OZD | 1220648-2025-30855 | 2025-07-13 04:00 | Death | ABIOMED, INC. | 2025-07-30 04:00 | OZD | IMPELLA 5.5 | Positioning Failure | Hemolysis; Thrombosis/Thrombus; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT DEVELOPED HEMOLYSIS. THE PUMP MIGRATED TOO DEEP AND WAS REPOSITIONED. STAFF MONITORED PATIENT'S LABS. HEPARIN WAS EVENTUALLY SWITCHED TO BIVAL WHICH HELPED THE PATIENT'S HEMOLYSIS. ADDITIONALLY, THE PATIENT HAD AN EXTENSIVE THROMBUS IN TRANSVERSE ARCH INTO DESCENDING. HEPARIN WAS SWITCHED TO BIVAL AND THE STAFF MONITORED. FURTHERMORE, THE PATIENT HAD A GASTROINTESTINAL BLEED. LATER CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." . IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22664532&pc=MHX | 3009156722-2025-00016 | 2025-07-07 04:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-07-30 04:00 | MHX | BENEVISION N1 PATIENT MONITOR | Device Alarm System; No Audible Alarm; Alarm Not Visible | Asystole | K213799 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, AT 10:00 P.M., A PATIENT WAS CONNECTED TO A BENEVISION N12 PATIENT MONITOR (SERIAL NUMBER: (B)(6) IN USE WITH A BENEVISION N1 (SERIAL NUMBER: (B)(6). THE PATIENT EXPERIENCED ASYSTOLE; HOWEVER, NO ALARM WAS GENERATED. Manufacturer Narrative: NARRATIVE MINDRAY REPRESENTATIVES TESTED THE OPERATION OF ALL BENEVISION N12 AND N1 MONITORS IN THE DEPARTMENT. THE DEVICES ARE FUNCTIONING NORMALLY, AND NO ALARM FAILURE PHENOMENON WAS FOUND. MORE MONITOR LOGS ARE BEING COLLECTED FOR INVESTIGATION, AND THIS ISSUE IS UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665226&pc=NGV | 2124215-2025-50979 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 6 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665241&pc=NGV | 2124215-2025-50580 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 6 | Event Description: THIS REPORT SUMMARIZES 6 DEATH EVENTS FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 56 TO 91YEARS. PATIENT SEX: 2 FEMALES; 4 MALES. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 39 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665242&pc=NGV | 2124215-2025-50977 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR UNSPECIFIED INFECTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 98 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665248&pc=NGV | 2124215-2025-50978 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 4 | Event Description: THIS REPORT SUMMARIZES 4 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 68 TO 83 YEARS. PATIENT SEX: 1 FEMALE; 3 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 57 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665490&pc=NGV | 2124215-2025-50966 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR LEFT ATRIAL THROMBOSIS. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 11 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665673&pc=NGV | 2124215-2025-50961 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR UNEXPECTED MEDICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 56 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665674&pc=NGV | 2124215-2025-50964 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR CARDIOVASCULAR HEMORRHAGE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 11 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665675&pc=NGV | 2124215-2025-50962 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR CARDIAC TAMPONADE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665676&pc=NGV | 2124215-2025-50958 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR DEVICE EXPLANTATION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 56 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665677&pc=NGV | 2124215-2025-50782 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Unspecified Tissue Injury |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR HEART VALVE DAMAGE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 11 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665678&pc=NGV | 2124215-2025-50965 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Pericardial Effusion |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR PERICARDIAL EFFUSION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. THE PERICARDIAL EFFUSION DEATH EVENT REQUIRED ADDITIONAL DEVICE (DRAINAGE) AND INVOLVED A CARDIAC TAMPONADE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665679&pc=NGV | 2124215-2025-50955 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR IMPLANT EMBOLIZATION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 56 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665680&pc=NGV | 2124215-2025-50957 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR DEATH CARDIOVASCULAR PROCEDURE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 56 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665682&pc=NGV | 2124215-2025-50959 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR HOSPITALIZATION OR PROLONGED HOSPITALIZATION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 56 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665683&pc=NGV | 2124215-2025-50956 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR ADDITIONAL DEVICE REQUIRED. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22665693&pc=NGV | 2124215-2025-50960 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR SURGICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 56 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22666041&pc=NIK | 9614453-2025-03005 | 2025-06-14 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-07-30 04:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Appropriate Term/Code Not Available | Heart Failure/Congestive Heart Failure | P010031 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S FAMILY MEMBER THAT TWO DAYS POST ROUTINE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) RE PLACEMENT PROCEDURE, THE PATIENT BEGAN "DETERIORATING" AND WAS EVENTUALLY HOSPITALIZED. THE CRT-D WAS PROGRAMMED OFF AND THE PATIENT DIED THREE WEEKS LATER. CAUSE OF DEATH WAS NOTED TO BE HEART FAILURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22666157&pc=FKX | 0008030665-2025-01708 | 2025-07-18 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-07-30 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Vomiting; Insufficient Information | K181108 |  | 1 | Event Description: A NURSE REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PD CLINIC. THE PATIENT¿S SPOUSE CONTACTED THE PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) ON (B)(6) 2025 AROUND 10:30AM TO REPORT THAT THE PATIENT HAD PASSED AWAY. THE PATIENT HAD BEEN IN TREATMENT CONNECTED TO THE LIBERTY SELECT CYCLER WHEN SHE BEGAN VOMITING. THE VOMIT WAS BLACK IN COLOR. THE SPOUSE DID AN EMERGENCY DISCONNECT FROM THE TREATMENT AND CALLED 911. WHEN EMERGENCY MEDICAL SERVICES (EMS) ARRIVED, THE PATIENT WAS NOT RESPONSIVE. EMS ATTEMPTED TO RESUSCITATE THE PATIENT BUT WERE UNSUCCESSFUL. THE CAUSE OF DEATH IS UNKNOWN. THERE WERE NO REPORTED ISSUES WITH THE LIBERTY SELECT CYCLER OR THE PATIENT¿S TREATMENT PRIOR TO THE EVENT. NO ADDITIONAL INFORMATION WAS AVAILABLE PERTAINING TO THE DEATH. Manufacturer Narrative: CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE PATIENT REPORTEDLY BEGAN VOMITING BLACK MATTER PRIOR TO BECOMING UNRESPONSIVE. ALTHOUGH THE CAUSE OF DEATH WAS NOT CONFIRMED, THIS IS AN INDICATION OF A GASTROINTESTINAL (GI) BLEED. IF THE PATIENT DID HAVE A SIGNIFICANT BLEED, SHE COULD HAVE GONE INTO HYPOVOLEMIC SHOCK CAUSING CARDIAC ARREST AND DEATH WHICH WOULD HAVE BEEN WHY THE PATIENT WAS UNRESPONSIVE WHEN PARAMEDICS ARRIVED. HOWEVER, WITHOUT ADDITIONAL INFORMATION, THAT CANNOT BE CONFIRMED. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22666553&pc=NIK | 9614453-2025-03008 | 2025-04-05 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-07-30 04:00 | NIK | COBALT¿ HF CRT-D MRI SURESCAN¿ | Failure to Deliver Shock/Stimulation | Cardiac Arrest | P010031 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S REPRESENTATIVE THAT THE PATIENT'S CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) DID NOT SHOCK THE PATIENT WHEN THE PATIENT WENT INTO SUDDEN CARDIAC ARREST (SCA). THE DEVICE REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: FOLLOW UP WITH THE CLINIC INDICATED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: 200SH18 PULMONARY VALVE IMPLANTED: (B)(6) 2009 4592-45 LEAD IMPLANTED: (B)(6) 2003 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22666646&pc=NGV | 2124215-2025-50925 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Pericardial Effusion |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR PERICARDIAL EFFUSION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. OF THE 2 PERICARDIAL EFFUSION DEATH EVENTS: BOTH REQUIRED AN ADDITIONAL DEVICE (DRAINAGE). BOTH INVOLVED A CARDIAC TAMPONADE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 11 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22666652&pc=NGV | 2124215-2025-50924 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  | E2017028 | 3 | Event Description: THIS REPORT SUMMARIZES 3 DEATH EVENTS FOR UNSPECIFIED INFECTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 71 TO 81 YEARS. PATIENT SEX: 3 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 38 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22666764&pc=NGV | 2124215-2025-50923 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Endocarditis |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR ENDOCARDITIS. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 79 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22666787&pc=NGV | 2124215-2025-50920 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR CARDIAC ARREST. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22666788&pc=NGV | 2124215-2025-50922 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 6 | Event Description: THIS REPORT SUMMARIZES 6 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 71 TO 81 YEARS. PATIENT SEX: 1 FEMALE; 5 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 132 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22666790&pc=NGV | 2124215-2025-50921 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR CARDIAC TAMPONADE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 11 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22666793&pc=NGV | 2124215-2025-50913 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR ADDITIONAL DEVICE REQUIRED. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 11 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22666796&pc=NGV | 2124215-2025-50917 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR DEATH CARDIOVASCULAR PROCEDURE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22666803&pc=NGV | 2124215-2025-50915 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR SUDDEN CARDIAC DEATH. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 1 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22666810&pc=NGV | 2124215-2025-50919 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR SURGICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22666957&pc=NGV | 2124215-2025-48854 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR HOSPITALIZATION OR PROLONGED HOSPITALIZATION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 577 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX DEVICE SURVEILLANCE POST APPROVAL ANALYSIS PLAN (WATCHMAN FLX-SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667012&pc=NGV | 2124215-2025-51340 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FXD CURVE? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Hematoma |  | E2017028 | 3 | Event Description: THIS REPORT SUMMARIZES 3 DEATH EVENTS FOR HEMATOMA. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 73 TO 90 YEARS. PATIENT SEX: 2 FEMALES; 1 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 22 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667016&pc=NGV | 2124215-2025-51350 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FXD CURVE? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Pseudoaneurysm |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR VESSEL PSEUDOANEURYSM. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667051&pc=OZD | 1220648-2025-30867 | 2025-07-18 04:00 | Death | ABIOMED, INC. | 2025-07-30 04:00 | OZD | IMPELLA CP | Restricted Flow rate | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT ADMITTED IN WITH THIRD DEGREE HEART BLOCK, SEIZURES, NO PULSE, AND ONGOING CARDIOPULMONARY RESUSCITATION HAD AN IMPELLA CP IMPLANTED. HOWEVER, THE PUMP FLOW WAS NOT APPROPRIATE AND THE MEDICAL TEAM MADE THE DECISION TO WITHDRAW CARE. THE PATIENT EXPIRED. Manufacturer Narrative: A4 AND A5 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667136&pc=BZD | 2518422-2025-108420 | 2025-07-22 04:00 | Death | RESPIRONICS, INC. | 2025-07-30 04:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: A REPORT WAS RECEIVED FROM THE FATHER OF A PATIENT STATING THE PATIENT PASSED AWAY SEVERAL YEARS AGO. A REPLACEMENT DEVICE SENT TO THE PATIENT FROM PHILIPS PROMPTED THE CALL FROM THE REPORTER (THE PATIENT'S FATHER) TO ADVISE THERE IS NO NEED FOR THE DEVICE AND INQUIRED ABOUT RETURNING THE DEVICE. DETAILS SURROUNDING THE PATIENT'S DEATH WERE NOT PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667145&pc=DSQ | 2916596-2025-04986 | 2025-07-11 04:00 | Death | THORATEC CORPORATION | 2025-07-30 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ischemia; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE ON (B)(6) 2025. THE PATIENT REQUIRED MULTIPLE PRESSORS POST OP WHICH LED TO ISCHEMIC BOWEL. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. CAUSE OF DEATH-ON-DEATH CERTIFICATE MULTIPLE SYSTEM ORGAN FAILURE. VAD WITHOUT ANY ISSUES. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667167&pc=FNL | 9616091-2025-00012 | 2025-06-09 04:00 | Death | INVAMEX | 2025-07-30 04:00 | FNL | IVC FULL ELECTRIC LOW BED | Adverse Event Without Identified Device or Use Problem | Physical Entrapment |  |  | 1 | Event Description: PER FDA MEDWATCH FORM, AN 80 YEAR OLD MALE WEIGHING 196LBS WAS FOUND DECEASED AND LYING ON THE FLOOR, WITH HIS ARM CAUGHT IN THE BED RAIL. THE PATIENT'S HEAD WAS ON THE FLOOR WITH HIS LOWER BODY SUSPENDED IN THE AIR.Event Description: PER FDA MEDWATCH FORM, AN 80 YEAR OLD MALE WEIGHING 196LBS WAS FOUND DECEASED AND LYING ON THE FLOOR, WITH HIS ARM CAUGHT IN THE BED RAIL. THE PATIENT'S HEAD WAS ON THE FLOOR WITH HIS LOWER BODY SUSPENDED IN THE AIR. Manufacturer Narrative: THIS FOLLOW UP MEDWATCH IS BEING SUBMITTED PER A REQUEST BY THE FDA TO INCLUDE THE MISSING RELATED REPORT NUMBERS FROM H10 OF THE INITIAL MEDWATCH SUBMITTED ON 7/30/25. THE FACILITY FILLED OUT TWO REPORTS, ONE FOR EACH COMPONENT OF THE BED INVOLVED, FOR THIS EVENT. THERE WAS ONLY ONE ADVERSE EVENT REGARDING THE 5410LOW BED. Manufacturer Narrative: THE FACILITY (B)(6) REPORTED AN ENTRAPMENT INCIDENT THAT LED TO A PATIENT'S DEATH. THE EVENT INVOLVES THE IVC 5410LOW ELECTRIC BED, MANUFACTURED BY INVACARE SANFORD (FLORIDA) ON MAY 1, 2022. THE HALF RAILS AND MATTRESS USED AT THE TIME OF THE INCIDENT WERE UNKNOWN. THE FACILITY WAS UNABLE TO PROVIDE FURTHER INFORMATION REGARDING THE MANUFACTURER AND OR SERIAL NUMBER OR DATE CODE FOR THESE ACCESSORIES. THERE WAS NO MALFUNCTION OR ISSUE ALLEGED WITH THE BED OR HALF RAILS. THE FACILITY HAS REMOVED FROM BED FROM USE. THE PATIENT HAD MULTIPLE PRE-EXISTING MEDICAL CONDITIONS DOCUMENTED IN THIS REPORT THAT MAY HAVE CONTRIBUTED TO THIS INCIDENT. BED RAIL ENTRAPMENT IS A KNOWN RISK IN THE USE OF BED¿S EQUIPPED WITH BED RAILS. THE BED RAIL ENTRAPMENT RISK NOTIFICATION GUIDE PROVIDED WITH THE BED ADVISES, ¿PROPER PATIENT ASSESSMENT, EQUIPMENT SELECTION, FREQUENT PATIENT MONITORING, AND COMPLIANCE WITH INSTRUCTIONS, WARNINGS AND THIS BED RAIL ENTRAPMENT RISK NOTIFICATION GUIDE IS ESSENTIAL TO REDUCE THE RISK OF ENTRAPMENT. ONLY THE PATIENT¿S MEDICAL CARE PROVIDER AND/OR THE DEALER FROM WHOM YOU OBTAINED THIS EQUIPMENT, UPON PROPER ASSESSMENT OF THE PATIENT¿S MEDICAL CONDITION AND NEEDS, CAN EVALUATE WHETHER THIS EQUIPMENT IS APPROPRIATE FOR USE BY ANY PARTICULAR PATIENT AND ASSIST THE PATIENT, THE PATIENT¿S FAMILY AND/OR THE PATIENT¿S PRIMARY DAY-TO-DAY CAREGIVER IN ASSESSING THE RISK OF ENTRAPMENT.¿ SHOULD ANY ADDITIONAL INFORMATION BECOME AVAILABLE, A FOLLOW UP REPORT WILL BE FILED. NOTE: THE MANUFACTURER OF THE SUBJECT BED IS INVACARE SANFORD (FLORIDA) WHICH IS NO LONGER IN OPERATION. THIS DEVICE IS NOW MANUFACTURED AT THE INVACARE INVAMEX LOCATION IN MEXICO; THEREFORE, THE INVAMEX ADDRESS AND CFN NUMBER IS BEING UTILIZED FOR THE MDR FILING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667178&pc=NGV | 2124215-2025-50598 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR HOSPITALIZATION OR PROLONGED HOSPITALIZATION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED.
TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE
BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 117 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667404&pc=NGV | 2124215-2025-50603 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR SURGICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667630&pc=OZD | 1220648-2025-30872 | 2025-05-25 04:00 | Death | ABIOMED, INC. | 2025-07-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ENDURED VENTRICULAR TACHYCARDIA DURING PUMP SUPPORT. LATER, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: TO DETERMINE THE TRUE ROOT CAUSE OF THE VENTRICULAR TACHYCARDIA THE CLINICAL INFORMATION WOULD NEED TO BE ASSESSED IN CONJUNCTION WITH EVIDENCE FROM THE RETURNED DEVICE. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VENTRICULAR TACHYCARDIA COULD NOT DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667644&pc=NGV | 2124215-2025-50759 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  | E2017028 | 14 | Event Description: THIS REPORT SUMMARIZES 14 DEATH EVENTS FOR CEREBRAL VASCULAR ACCIDENT (CVA). NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 65 TO (B)(6) YEARS. PATIENT SEX: 5 FEMALE; 9 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 68 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667645&pc=NGV | 2124215-2025-50764 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Endocarditis |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR ENDOCARDITIS. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 139 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667649&pc=NGV | 2124215-2025-50606 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR UNEXPECTED MEDICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 227 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667650&pc=NGV | 2124215-2025-50708 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest |  | E2017028 | 4 | Event Description: THIS REPORT SUMMARIZES 4 DEATH EVENTS FOR CARDIAC ARREST. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 68 TO 91 YEARS. PATIENT SEX: 3 FEMALE; 1 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667699&pc=OZD | 1220648-2025-30873 | 2025-07-18 04:00 | Death | ABIOMED, INC. | 2025-07-30 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT UNDERWENT CORONARY ARTERY BYPASS GRAFT, AORTIC VALVE REPLACEMENT, MAZE. THE PATIENT REQUIRED CARDIAC MASSAGE DUE TO PRE-BYPASS VENTRICULAR FIBRILLATION EVENT. INTRA-AORTIC BALLOON PUMP WAS PLACED TO RIGHT FEMORAL. DUE TO BI-VENTRICULAR FAILURE AND UNABLE TO SEPARATE FROM CARDIO-PULMONARY BYPASS, THE PATIENT UNDERWENT RIGHT AXILLARY IMPELLA 5.5 PLACEMENT WITH AL1 AND STRAIGHT WIRE UNDER TRANSESOPHAGEAL ECHOCARDIOGRAM GUIDANCE. THERE WAS A LEFT VENTRICLE PERFORATION NOTICED BY SURGEON AND PRIMARILY REPAIRED WITH SINGLE STITCH. THE PATIENT SUBSEQUENTLY UNABLE TO SEPARATE FROM CARDIO-PULMONARY BYPASS WITH IMPELLA AND EXPIRED ON THE TABLE. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667861&pc=NGV | 2124215-2025-50795 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Thromboembolism |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR THROMBOEMBOLISM. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025 AVERAGE TIME TO EVENT: 227 DAYS DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667862&pc=NGV | 2124215-2025-50770 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  | E2017028 | 2 | Event Description: THIS REPORT SUMMARIZES 2 DEATH EVENTS FOR UNSPECIFIED INFECTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 85 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667864&pc=NGV | 2124215-2025-50775 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Unspecified Tissue Injury |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR HEART VALVE DAMAGE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22667969&pc=LPB | 2029046-2025-02489 |  | Death | BIOSENSE WEBSTER INC | 2025-07-30 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030031/S078 |  | 1 | Event Description: THIS COMPLAINT IS FROM A DATABASE RELATED RESEARCH ACTIVITY (DRRA), "INDICATION EXTENSION FOR MEDICAL DEVICES USING RWE FROM NESTCC NETWORK COLLABORATORS: SAFETY AND EFFECTIVENESS OF CARDIAC ABLATION OF ISCHEMIC VENTRICULAR TACHYCARDIA USING THE THERMOCOOL STSF CATHETER CLINICAL STUDY REPORT AMENDMENT: UPDATED ANALYSES". ADVERSE EVENTS REPORTED FOR THERMOCOOL SMARTTOUCH® SF CATHETER: QTY 3: DEATH WITHIN 30 DAYS OF THE INDEX ABLATION PROCEDURE (DEATH)(NO INFORMATION AVAILABLE). Manufacturer Narrative: D4. CATALOG: UNK SMART TOUCH UNIDIRECTIONAL SF. A DATE OF DEATH WAS NOT PROVIDED, AS SUCH, BEST ESTIMATED DATE OF (B)(6) 2025 HAS BEEN ADDED TO FIELD B2. DATE OF DEATH. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. THIS COMPLAINT IS FROM A DATABASE RELATED RESEARCH ACTIVITY (DRRA), "INDICATION EXTENSION FOR MEDICAL DEVICES USING RWE FROM NESTCC NETWORK COLLABORATORS: SAFETY AND EFFECTIVENESS OF CARDIAC ABLATION OF ISCHEMIC VENTRICULAR TACHYCARDIA USING THE THERMOCOOL STSF CATHETER CLINICAL STUDY REPORT AMENDMENT: UPDATED ANALYSES". NO DEVICE WAS RECEIVED FOR ANALYSIS AT THE TIME OF SUBMISSION OF THE INITIAL 3500A. SINCE THE PRODUCT WAS NOT RETURNED FOR ANALYSIS, NO PRODUCT FAILURE ANALYSIS CAN BE CONDUCTED, AND NO DETERMINATION OF POSSIBLE CONTRIBUTING FACTORS COULD BE MADE. DEVICE HISTORY RECORD (DHR) REVIEW CANNOT BE CONDUCTED BECAUSE NO LOT NUMBER WAS PROVIDED BY THE CUSTOMER. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REF #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22668027&pc=LPB | 2029046-2025-02486 |  | Death | BIOSENSE WEBSTER INC | 2025-07-30 04:00 | LPB | THERMOCOOL SMARTTOUCH | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030031/S053 |  | 1 | Event Description: THIS COMPLAINT IS FROM A DATABASE RELATED RESEARCH ACTIVITY (DRRA), "INDICATION EXTENSION FOR MEDICAL DEVICES USING RWE FROM NESTCC NETWORK COLLABORATORS: SAFETY AND EFFECTIVENESS OF CARDIAC ABLATION OF ISCHEMIC VENTRICULAR TACHYCARDIA USING THE THERMOCOOL STSF CATHETER CLINICAL STUDY REPORT AMENDMENT: UPDATED ANALYSES." ADVERSE EVENTS REPORTED FOR THERMOCOOL SMARTTOUCH® CATHETER: QTY 2: DEATH WITHIN 30 DAYS OF THE INDEX ABLATION PROCEDURE (DEATH)(NO INFORMATION AVAILABLE). Manufacturer Narrative: A DATE OF DEATH WAS NOT PROVIDED, AS SUCH, BEST ESTIMATED DATE OF (B)(6) 2025 HAS BEEN ADDED TO FIELD B2. DATE OF DEATH. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. THIS COMPLAINT IS FROM A DATABASE RELATED RESEARCH ACTIVITY (DRRA), "INDICATION EXTENSION FOR MEDICAL DEVICES USING RWE FROM NESTCC NETWORK COLLABORATORS: SAFETY AND EFFECTIVENESS OF CARDIAC ABLATION OF ISCHEMIC VENTRICULAR TACHYCARDIA USING THE THERMOCOOL STSF CATHETER CLINICAL STUDY REPORT AMENDMENT: UPDATED ANALYSES". NO DEVICE WAS RECEIVED FOR ANALYSIS AT THE TIME OF SUBMISSION OF THE INITIAL 3500A. SINCE THE PRODUCT WAS NOT RETURNED FOR ANALYSIS, NO PRODUCT FAILURE ANALYSIS CAN BE CONDUCTED, AND NO DETERMINATION OF POSSIBLE CONTRIBUTING FACTORS COULD BE MADE. DEVICE HISTORY RECORD (DHR) REVIEW CANNOT BE CONDUCTED BECAUSE NO LOT NUMBER WAS PROVIDED BY THE CUSTOMER. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REF #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22668158&pc=OZD | 1220648-2025-30875 | 2025-07-18 04:00 | Death | ABIOMED, INC. | 2025-07-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP HAD LITTLE TO NO FLOW DOWN THE LEG. THE PLAN WAS TO REMAIN ON IMPELLA OVER THE WEEKEND, AND THE STAFF WAS MONITORING THE DISTAL PULSES, BUT THE FAMILY DECIDED TO WITHDRAW CARE. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22668260&pc=MIH | 3007284313-2025-04155 | 2025-07-04 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-07-30 04:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem |  | P020004 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON AN UNKNOWN DATE IN 2014, PATIENT UNDERWENT AN UNKNOWN ENDOVASCULAR PROCEDURE UTILIZING A GORE® EXCLUDER® AAA ENDOPROSTHESIS (TRUNK IPSILATERAL LEG AND CONTRALATERAL LEG). ON (B)(6) 2019, PATIENT UNDERWENT AN UNKNOWN REINTERVENTION UTILIZING A GORE® EXCLUDER® AAA ENDOPROSTHESIS (CONTRALATERAL LEG AND AN AORTIC EXTENDER), AND GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS (INTERNAL ILIAC COMPONENT AND ILIAC BRANCH COMPONENT). ON (B)(6) 2025, PATIENT UNDERWENT A REINTERVENTION ENDOVASCULAR PROCEDURE ON A PRIOR ENDOVASCULAR ANEURYSM REPAIR (EVAR) TO TREAT A TYPE 1A AND 1B ENDOLEAKS UTILIZING A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM, GORE® EXCLUDER® AAA ENDOPROSTHESIS (THREE CONTRALATERAL LEGS AND AORTIC EXTENDER), AND GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS (AORTIC EXTENDER). REPORTEDLY, PHYSICIAN WAS NOT ABLE TO GET CONTRAST SCANS DUE TO PATIENT'S POOR RENAL FUNCTION, WHICH WAS DUE TO PATIENT CONDITION. THE TYPE 1B ENDOLEAK WAS ON THE TRUNK IPSILATERAL LEG FROM 2014 AND THE TYPE 1B ENDOLEAKS WAS ON THE CONTRALATERAL LEG IMPLANTED IN 2019. DESPITE USING MULTIPLE CONTRALATERAL LEGS AND CONFIRMABLE AORTIC EXTENDERS DURING THE REINTERVENTION, THERE WAS STILL A TYPE 3 ENDOLEAK BETWEEN THEM WHICH WAS FIXED BY USING A THORACIC CONFORMABLE STENT GRAFT IN AN OFF-LABEL USE. REINTERVENTION PROCEDURE WENT WELL AND PATIENT WAS DOING OKAY. ON (B)(6) 2025, FIELD SALES ASSOCIATE WAS MADE AWARE THAT PATIENT PASSED AWAY ON (B)(6) 2025. CAUSE OF DEATH WAS NOT PROVIDED AND FURTHER INFORMATION IS NOT AVAILABLE. Manufacturer Narrative: H6: CODE B20 -THE DEVICE REMAINS IMPLANTED AND WAS THEREFORE NOT AVAILABLE FOR ENGINEERING EVALUATION BY GORE. H6: C20- A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE CONDUCTED BECAUSE THE SERIAL/LOT NUMBER REMAINS UNKNOWN. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22668266&pc=MIH | 3013164176-2025-02625 | 2025-07-04 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-07-30 04:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem |  | P020004 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON AN UNKNOWN DATE IN 2014, PATIENT UNDERWENT AN UNKNOWN ENDOVASCULAR PROCEDURE UTILIZING A GORE® EXCLUDER® AAA ENDOPROSTHESIS (TRUNK IPSILATERAL LEG AND CONTRALATERAL LEG). ON (B)(6) 2019, PATIENT UNDERWENT AN UNKNOWN REINTERVENTION UTILIZING A GORE® EXCLUDER® AAA ENDOPROSTHESIS (CONTRALATERAL LEG AND AN AORTIC EXTENDER), AND GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS (INTERNAL ILIAC COMPONENT AND ILIAC BRANCH COMPONENT). ON (B)(6) 2025, PATIENT UNDERWENT A REINTERVENTION ENDOVASCULAR PROCEDURE ON A PRIOR ENDOVASCULAR ANEURYSM REPAIR (EVAR) TO TREAT A TYPE 1A AND 1B ENDOLEAKS UTILIZING A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM, GORE® EXCLUDER® AAA ENDOPROSTHESIS (THREE CONTRALATERAL LEGS AND AORTIC EXTENDER), AND GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS (AORTIC EXTENDER). REPORTEDLY, PHYSICIAN WAS NOT ABLE TO GET CONTRAST SCANS DUE TO PATIENT'S POOR RENAL FUNCTION, WHICH WAS DUE TO PATIENT CONDITION. THE TYPE 1B ENDOLEAK WAS ON THE TRUNK IPSILATERAL LEG FROM 2014 AND THE TYPE 1B ENDOLEAKS WAS ON THE CONTRALATERAL LEG IMPLANTED IN 2019. DESPITE USING MULTIPLE CONTRALATERAL LEGS AND CONFIRMABLE AORTIC EXTENDERS DURING THE REINTERVENTION, THERE WAS STILL A TYPE 3 ENDOLEAK BETWEEN THEM WHICH WAS FIXED BY USING A THORACIC CONFORMABLE STENT GRAFT IN AN OFF-LABEL USE. REINTERVENTION PROCEDURE WENT WELL AND PATIENT WAS DOING OKAY. ON JULY 07, 2025, FIELD SALES ASSOCIATE WAS MADE AWARE THAT PATIENT PASSED AWAY ON (B)(6) 2025. CAUSE OF DEATH WAS NOT PROVIDED AND FURTHER INFORMATION IS NOT AVAILABLE. Manufacturer Narrative: H6: CODE C19 - A REVIEW OF THE MANUFACTURING RECORDS INDICATED THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE B20 -THE DEVICE REMAINS IMPLANTED AND WAS THEREFORE NOT AVAILABLE FOR ENGINEERING EVALUATION BY GORE. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22668475&pc=DSQ | 2916596-2025-05023 | 2025-07-19 04:00 | Death | THORATEC CORPORATION | 2025-07-30 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE WAS NOT PROVIDED; PATIENT WAS FOUND DECEASED AT HOME. THE PATIENT WAS NOTED TO HAVE BEEN SLEEPING ON BATTERY(S). WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. IT IS UNKNOWN IF THE DEVICE WILL BE RETURNED FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22668619&pc=NGV | 2124215-2025-51840 | 2025-02-18 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-30 04:00 | NGV | WATCHMAN LAA CLOSURE DEVICE & DELIVERY SYSTEM | Failure to Seal | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: REPORTED VIA LITERATURE ARTICLE. IT WAS REPORTED THAT DEATH OCCURRED. THE FOLLOWING EVENT WAS OBTAINED VIA A RETROSPECTIVE ARTICLE FOLLOWING THE ANALYSIS OF 236 PATIENTS TO THE TEST OR CONTROL DEVICE FOR LEFT ATRIAL APPENDAGE OCCLUSION (LAAO). THE EFFECTIVENESS ENDPOINTS INCLUDED 12-MONTH SUCCESSFUL LEFT ATRIAL APPENDAGE (LAA) SEALING WITH RESIDUAL FLOW LESS OR EQUAL TO 5MM AND CLINICAL SUCCESS COMPOSITE OF ISCHEMIC STROKE, TRANSIENT ISCHEMIC ATTACK, AND SYSTEMIC EMBOLISM. SAFETY ENDPOINTS WERE ALSO STUDIED. A NONINFERIORITY MARGIN OF MINUS 7 PERCENT WAS SET BETWEEN THE DEVICES. LAAO WAS PERFORMED UNDER ANGIOGRAPHIC AND TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) GUIDANCE. AFTER LAAO, THE PATIENTS WERE ADMINISTERED A POST IMPLANTATION DRUG REGIMEN CONSISTING OF WARFARIN FOR THE FIRST 3 MONTHS, ASPIRIN PLUS CLOPIDOGREL FOR THE SECOND 3 MONTHS, FOLLOWED BY ASPIRIN OR CLOPIDOGREL ALONE FOR 6 MONTHS. THE CONTROL GROUP HAD TWO CASES OF UNSUCCESSFUL IMPLANTATION DUE TO THE DISCOVERY OF A SHORT LAA LENGTH ON DIGITAL SUBTRACTION ANGIOGRAPHY (DSA) DURING THE PROCEDURE BEFORE IMPLANTATION. ADVERSE EVENTS (AES) WERE ADJUDICATED BY AN INDEPENDENT CLINICAL EVENTS COMMITTEE BLINDED TO TREATMENT ASSIGNMENT AND DEVICE SELECTION. THE SAFETY END POINTS INCLUDED THE12-MONTH INCIDENCES OF ALL-CAUSE DEATH AND STROKE/TIA. THE OTHER SAFETY ENDPOINTS WERE ASSESSED BEFORE DISCHARGE OR WITHIN 7 DAYS AND AT 1, 3, 6, AND 12MONTHS POST-LAAO. ALL-CAUSE AES WERE DEFINED AS ADVERSE MEDICAL EVENTS THAT OCCURRED DURING THE TRIAL, REGARDLESS OF WHETHER THEY WERE RELATED TO THE EXPERIMENTAL MEDICAL DEVICE. ALL-CAUSE SERIOUS ADVERSE EVENTS (SAES) WERE DEFINED AS ANY COMPLICATION THAT LED TO DEATH; CAUSED SERIOUS HEALTH DETERIORATION OR THAT DID NOT CAUSE DEATH BUT THREATENED LIFE OR REQUIRED HOSPITALIZATION. IN THE CURRENT STUDY, THE RATE OF 12 MONTH ALL CAUSE DEATH WITH THE TEST DEVICE WAS 0.9 PERCENT, AND NO CASES OF DRT WERE OBSERVED. Manufacturer Narrative: B3: ESTIMATED AS 2/18/2025 AS THE PUBLISHED DATE FOR THE ARTICLE IS NOTED AS FEBRUARY 18, 2025. D4 UNIQUE IDENTIFIER (UDI) #: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. LITERATURE CITATION: FAN, Y., MA, D., WANG, C., LUO, J., LING, Z., LI, S., PENG, X., ZHANG, Z., CHI, H., & WANG, J. (2025). A MEMBRANE MODIFICATION TECHNIQUE FOR LEFT ATRIAL APPENDAGE OCCLUSION. JACC, VOLUME 5, NO. 3. HTTPS://DOI.ORG/10.1016/J.JACASI.2024.12.008.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22668627&pc=DSQ | 2916596-2025-05119 | 2025-07-22 04:00 | Death | THORATEC CORPORATION | 2025-07-30 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Ischemia Stroke; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS TRANSFERRED TO ANCILLARY FACILITY AND PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) POST-LEFT VENTRICULAR ASSIST DEVICE (LVAD) INSERTION. CARE WAS WITHDRAWN, AND THE PATIENT PASSED AWAY DUE TO A STROKE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. ADDITIONAL INFORMATION WAS UNKNOWN.Event Description: ADDITIONAL INFORMATION WAS PROVIDED ON 04AUG2025 THAT THEY WERE NOT RESPONDING TO TRANSITION TO THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) FROM PUMP WITH CLOSING OF THEIR CHEST. THEY WERE PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) FOR SUPPORT AND LVAD FLOWS COMING OUT OF THE OPERATING ROOM WERE 0.0 LITERS PER MINUTE (LPM), PULSATILITY INDEX (PI) WAS 8.6, POWER WAS 2.8 WATTS, AND SPEED WAS 4800 ROTATIONS PER MINUTE (RPM). THEY WERE EVENTUALLY TRANSITIONED TO IMPELLA RP AND OFF ECMO WITH ONLY RP AND LVAD RUNNING. THE PATIENT WAS ANTICOAGULATED THE ENTIRE TIME WITH HEPARIN. THEY HAD MULTIPLE ISCHEMIC INFARCTS NOTED ON COMPUTED TOPOGRAPHY (CT) OF HEAD. THE DEATH WAS POSSIBLY DEVICE RELATED WITH POTENTIAL CLOTS, BUT IT WAS UNDETERMINED. THE DEVICE OPERATED APPROPRIATELY, FOR 2 DAYS WAS AT 0.0 LPM FLOW WITH FULL ECMO SUPPORT. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE REPORT OF POTENTIAL CLOTS IN THE DEVICE WAS UNABLE TO BE CONFIRMED THROUGH THIS EVALUATION. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING PUMP THROMBOSIS, STROKE, RIGHT HEART FAILURE, AND DEATH, THAT MAY BE ASSOCIATED WITH THE SUE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 5 OF THE IFU, "SURGICAL PROCEDURES", INSTRUCTS TO INSPECT THE VENTRICULAR CHAMBER FOR MURAL THROMBI AND CROSSING TRABECULAE FOLLOWING REMOVAL OF THE CORE AND TO ADDRESS ONE OR BOTH, AS NEEDED. THIS SECTION ALSO INSTRUCTS TO INSPECT THE VENTRICLE AND REMOVE ANY PREVIOUSLY FORMED CLOTS THAT MAY CAUSE EMBOLISM OR ANY TRABECULAE THAT MAY IMPEDE FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THIS SECTION, UNDER ¿ANTICOAGULATION¿, ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INR RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 6 OF THE IFU, "PATIENT CARE AND MANAGEMENT" (UNDER "RIGHT HEART FAILURE"), DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22669434&pc=NGV | 2135147-2025-04176 | 2025-07-07 04:00 | Death | ABBOTT MEDICAL | 2025-07-30 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 22MM AMPLATZER AMULET WAS SELECTED FOR IMPLANT. THIS DEVICE WAS MIS-SIZED AND UPSIZED TO A 25MM AMPLATZER AMULET, WHICH WAS IMPLANTED USING A 14F AMPLATZER STEERABLE DELIVERY SHEATH. PRIOR TO PROCEDURE, THE PATIENT WAS REPORTED TO BE IN STABLE CONDITION. POST-IMPLANT TRANSTHORACIC ECHOCARDIOGRAM (TTE) WAS PERFORMED PRIOR TO DISCHARGE WITH NO FINDINGS. THE PATIENT WAS DISCHARGED HOME, HAD DINNER, WATCHED A MOVIE, THEN WENT TO BED. AT APPROXIMATELY 2AM, THE PATIENT AWOKE WITH NECK PAIN, BY THE TIME TEA WAS MADE, THE PATIENT REPORTED FEELING BETTER AND WENT BACK TO SLEEP. IN THE MORNING, THE PATIENT WAS FOUND TO HAVE PASSED. NO AUTOPSY IS PLANNED. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22670717&pc=CCW | 2020550-2025-01307 | 2025-06-15 04:00 | Death | KARL STORZ SE & CO. KG | 2025-07-30 04:00 | CCW | C-MAC® POCKET MONITOR | Optical Problem | Insufficient Information |  |  | 1 | Event Description: ACCORDING TO THE INFORMATION RECEIVED, IMAGE DROPOUTS. IT WAS CONFIRMED BY DOCTOR THAT THE PATIENT HAD ALREADY DIED BY THE TIME PARAMEDIC TEAM ARRIVED, BUT ATTEMPTS WERE STILL MADE TO SECURE THE AIRWAY. DATE OF DEATH IS UNKNOWN. DOCTOR/PARAMEDIC CONFIRMED THAT THE DEVICE FAILURE NEITHER CAUSED NOR CONTRIBUTED TO THE PATIENT'S DEATH. APPROPRIATE REPLACEMENT EQUIPMENT WAS AVAILABLE AT THE TIME OF USE. Manufacturer Narrative: THE AFFECTED DEVICE HAS BEEN REQUESTED FOR INVESTIGATION BY THE MANUFACTURER. THE DEVICE WAS NOT YET RETURNED FOR INVESTIGATION. IT WAS CONFIRMED THAT THE DEVICE FAILURE NEITHER CAUSED NOR CONTRIBUTED TO THE PATIENT'S DEATH. THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22670796&pc=CCW | 9610617-2025-01307 | 2025-06-15 04:00 | Death | KARL STORZ SE & CO. KG | 2025-07-30 04:00 | CCW | C-MAC® POCKET MONITOR | Optical Problem | Insufficient Information | EXEMPT |  | 1 | Event Description: ACCORDING TO THE INFORMATION RECEIVED, IMAGE DROPOUTS. IT WAS CONFIRMED BY DOCTOR THAT THE PATIENT HAD ALREADY DIED BY THE TIME PARAMEDIC TEAM ARRIVED, BUT ATTEMPTS WERE STILL MADE TO SECURE THE AIRWAY. DATE OF DEATH IS UNKNOWN. DOCTOR/PARAMEDIC CONFIRMED THAT THE DEVICE FAILURE NEITHER CAUSED NOR CONTRIBUTED TO THE PATIENT'S DEATH. APPROPRIATE REPLACEMENT EQUIPMENT WAS AVAILABLE AT THE TIME OF USE. Manufacturer Narrative: THE AFFECTED DEVICE HAS BEEN REQUESTED FOR INVESTIGATION BY THE MANUFACTURER. THE DEVICE WAS NOT YET RETURNED FOR INVESTIGATION. IT WAS CONFIRMED THAT THE DEVICE FAILURE NEITHER CAUSED NOR CONTRIBUTED TO THE PATIENT'S DEATH. THE EVENT IS FILED UNDER INTERNAL KARL STORZ COMPLAINT ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22672018&pc=BZD | 2518422-2025-108448 | 2025-07-23 04:00 | Death | RESPIRONICS, INC. | 2025-07-30 04:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION A USER OF UNKNOWN DEVICE HAD PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED; THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22672597&pc=DSQ | 2916596-2025-05089 | 2025-07-21 04:00 | Death | THORATEC CORPORATION | 2025-07-30 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest; Pneumonia; Ventricular Fibrillation; Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS PNEUMONIA AND BACTEREMIA. THE PATIENT WAS ON VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO). THE PATIENT'S FAMILY WITHDREW CARE. ADDITIONAL INFORMATION WAS NOT PROVIDED.Event Description: PNEUMONIA WAS NOT ATTRIBUTED TO LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY. THE SOURCE OF INFECTION WAS VENTILATOR ASSOCIATED PNEUMONIA. THE PATIENT HAD CARDIAC ARREST THAT WAS BOTH RESPIRATORY AND ASSOCIATED WITH VENTRICULAR TACHYCARDIA (VT)/VENTRICULAR FIBRILLATION (VF). THE PATIENT CLINICALLY WORSENED AND REQUIRED VASOPRESSORS/INOTROPES. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE HAD OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION A3/A4: PATIENT GENDER AND WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22640163&pc=OZP | 2032227-2025-220601 | 2025-06-23 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-29 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Break | Hyperglycemia; Hypoglycemia; Cancer | P160017 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY AT THE HOSPITAL ON (B)(6) 2025. THE CUSTOMER WAS ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS CANCER AND DIABETES. THE CUSTOMER ALSO REPORTED SERIAL NUMBER WAS WORN OFF. THE CUSTOMER¿S BLOOD GLUCOSE WAS UNKNOWN. THE EVENT INVOLVED PRODUCT(S) MMT-332A, MMT-1884, MMT-397A. TROUBLESHOOTING WAS PERFORMED. IT WAS UNKNOWN THAT THE CUSTOMER WAS WEARING THE INSULIN PUMP AT THE TIME OF DEATH. THE CUSTOMER WAS NOT USING SENSORS. THE INSULIN PUMP WAS LAST WORN WAS UNKNOWN. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A, MMT-1884, MMT-397A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22640829&pc=PNJ | 9612164-2025-03676 | 2025-07-09 04:00 | Death | MEDTRONIC IRELAND | 2025-07-29 04:00 | PNJ | MICRA¿ AV2 | Appropriate Term/Code Not Available | Pericardial Effusion; Heart Failure/Congestive Heart Failure | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IT WAS FURTHER REPORTED THAT POST IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT EXPERIENCED PERICARDIAL EFFUSION.Event Description: IT WAS FURTHER NOTED THAT WHILE BLOOD PRESSURE STABILIZED AFTER DRAINAGE, THE PATIENT LATER PASSED AWAY DUE TO ACUTE HEART FAILURE.Event Description: THE PHYSICIAN'S OPINION WAS THAT THE ISSUE OCCURRED BEFORE THE PROCEDURE, IT WAS CONCLUDED THAT THERE IS NO CAUSAL RELATIONSHIP WITH THE MICRA IMPLANTATION PROCEDURE. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MC2AVR1] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [(B)(6)]). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22640831&pc=DYB | 9612164-2025-03677 | 2025-07-09 04:00 | Death | MEDTRONIC, INC. | 2025-07-29 04:00 | DYB | MICRA | Appropriate Term/Code Not Available | Pericardial Effusion; Heart Failure/Congestive Heart Failure | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IT WAS FURTHER REPORTED THAT POST IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT EXPERIENCED PERICARDIAL EFFUSION.Event Description: IT WAS FURTHER NOTED THAT WHILE BLOOD PRESSURE STABILIZED AFTER DRAINAGE, THE PATIENT LATER PASSED AWAY DUE TO ACUTE HEART FAILURE.Event Description: THE PHYSICIAN'S OPINION WAS THAT THE ISSUE OCCURRED BEFORE THE PROCEDURE, IT WAS CONCLUDED THAT THERE IS NO CAUSAL RELATIONSHIP WITH THE MICRA IMPLANTATION PROCEDURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22640884&pc=DQY | 9612164-2025-03680 | 2025-07-10 04:00 | Death | MEDTRONIC IRELAND | 2025-07-29 04:00 | DQY | TELESCOPE GUIDE EXTENSION CATHETER | Difficult to Remove; Detachment of Device or Device Component | Device Embedded In Tissue or Plaque | K183353 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PERCUTANEOUS CORONARY INTERVENTION (PCI) PROCEDURE INVOLVING ONE TELESCOPE GUIDE EXTENSION CATHETER (GEC) A DETACHMENT OF THE TIP OF THE DEVICE OCCURRED. THE GEC WAS INSPECTED PRIOR TO USE AND PREPPED PER IFU. THE LESION BEING TREATED HAD MODERATE TO SEVERE TORTUOSITY AND WAS SEVERELY CALCIFIED LOCATED IN THE MID CIRCUMFLEX (CX) ARTERY WITH 70% STENOSIS. THERE WAS POOR GUIDE SUPPORT WAS WITH ANGULATED CX. THE GEC WAS BEING USED TO FACILITATE NC BALLOON POSITIONING. IT WAS ATTEMPTED TO DELIVER THE STENT WITHOUT THE GEC HOWEVER THIS WAS UNSUCCESSFUL AND THE GEC WAS RE-DEPLOYED USING NC BALLOON ANCHOR TECHNIQUE TO ADVANCE. THE GEC WAS DELIVERED ACROSS THE STENOSIS WITHOUT DIFFICULTY FOLLOWING EXPANSION OF NC BALLOON AT NOMINAL PRESSURE IN THE HEAVILY CALCIFIED SEGMENT. MILD RESISTANCE WELL WITHIN NORMAL PRACTICE WAS EXPERIENCED ADVANCING THE GEC TO DELIVER THE GEC FOR ANCHOR SUPPORT WITH NO EXCESSIVE FORCE USED. THE STENT WAS THEN DELIVERED THROUGH THE GEC TO THE INTENDED DEPLOYMENT SITE WITHOUT RESISTANCE. IT WAS THEN ATTEMPTED TO WITHDRAW THE GEC OVER THE STENT IN ORDER TO DEPLOY THE STENT HOWEVER THE GEC WAS UNABLE TO BE WITHDRAWN. ATTEMPTS TO REMOVE THE STENT AND GEC EN BLOC FAILED. THERE WAS NO INTERACTION BETWEEN THE STENT AND THE GEC. A DETACHMENT OF THE SOFT POL YMER TIP OF THE GEC PROXIMAL TO THE RADIO-OPAQUE MARKER OCCURRED. THE STENT GUIDE AND WIRE WERE REMOVED EN BLOC. A NEW ACCESS WAS MADE AND THE GEC TIP WAS REWIRED WITH DIFFICULTY. MULTIPLE ATTEMPTS WERE MADE TO REMOVE THE GEC USING DISTAL NC BALLOON WHICH WAS INFLATED HOWEVER THIS WAS UNSUCCESSFUL. IT WAS UNABLE TO WIRE AROUND THE TIGHT STENOSIS PRESENT. THE PROCEDURE WAS ABANDONED. THE PATIENT WAS HAEMODYNAMICALLY STABLE, WITH NO ST ELEVATION, TIMI 3 FLOW IN VESSEL AT THE END OF THE PROCEDURE. THE PATIENT WAS TRANSFERRED TO THE SURGICAL THEATRE FOR CORONARY ARTERY BYPASS GRAFT (CABG) PROCEDURE TO REMOVE THE DETACHED TIP. CARDIAC SURGERY WAS WAS ALSO INDICATED FOR THE PATIENT DUE TO SEVERE LEFT ANTERIOR DESCENDING (LAD) ARTERY, CX AND RIGHT CORONARY ARTERY (RCA). THE PATIENT SUBSEQUENTLY DIED AS A RESULT OF POST-OPERATIVE COMPLICATIONS OF THE CABG SURGERY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22641986&pc=OZD | 1220648-2025-30798 | 2025-07-12 04:00 | Death | ABIOMED, INC. | 2025-07-29 04:00 | OZD | IMPELLA CP | Pumping Stopped | Great Vessel Perforation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND DURING THE IMPLANT, THERE WAS A MASSIVE RUPTURE OF ONE OF BIG ARTERIES. VASCULAR SURGERY ADVICE WAS TAKEN BUT THE PROCEDURE COULD NOT BE ACHIEVED DUE TO MASSIVE INSTABILITY AND ONGOING RESUSCITATION. ADDITIONALLY, THERE WAS AN ISSUE WITH THE PUMP UPON START UP. THE PUMP DID NOT START IN AUTO MODE BUT DIRECTLY ON A P-LEVEL. THE CONSOLE WAS RESTARTED BUT STILL THIS DID NOT RESOLVE THE ISSUE. THE CATHETER ALSO WENT FROM P5 TO P0 WITHOUT INTERACTION. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: ENTERING CARDIAC OUTPUT ¿MAINTAINING ACT AT OR ABOVE 250 SECONDS WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP.¿ SECTION: IMPELLA STOPPED ¿IF THE IMPELLA CATHETER HAS STOPPED SUDDENLY: 1. TRY TO RESTART THE CATHETER AT PREVIOUS P-LEVEL. 2. IF THE IMPELLA DOES NOT RESTART, TRY TO RESTART AT P-2. 3. IF THE IMPELLA DOES NOT RESTART OR STOPS AGAIN, WAIT 1 MINUTE AND TRY TO RESTART AGAIN. 4. IF THE IMPELLA RESTARTS, WEAN DOWN TO P-2 AS THE PATIENT CAN TOLERATE. UNDER THESE CIRCUMSTANCES, CATHETER FUNCTION IS NOT RELIABLE AND THE IMPELLA MAY STOP AGAIN. 5. IF THE IMPELLA DOES NOT RESTART, REMOVE THE IMPELLA FROM THE VENTRICLE AS SOON AS POSSIBLE TO AVOID AORTIC INSUFFICIENCY.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22642419&pc=MIH | 3007284313-2025-04152 | 2025-07-02 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-07-29 04:00 | MIH | GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Obstruction/Occlusion | P200030 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE FOR AN EMERGENT RUPTURED PARARENAL ABDOMINAL AORTIC ANEURYSM (AAA) UTILIZING A GORE® EXCLUDER® CONFORMABLE AAA ENDOPROSTHESIS, GORE® EXCLUDER® AAA ENDOPROSTHESIS AND GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX). REPORTEDLY, PHYSICIAN DECIDED TO TREAT WITH AN ARTIVION DEVICE AND USE VBX FOR THE BRANCHES AND AN EXCLUDER CONFORMABLE FOR THE ABDOMINAL BRIDGING STENTS. PHYSICIAN COULD NOT ACCESS THE RENAL ARTERIES DUE TO PATIENT'S ANATOMICAL CHALLENGES AS BOTH RENALS WERE SMALL AND HAD POOR PATENCY, THUS, DECIDED TO PUT THE EXCLUDER CONFORMABLE GRAFT OVER THE RENAL BRANCHES. PHYSICIAN WAS AWARE THAT THEY WOULD NEED TO COVER AT LEAST ONE RENAL (RIGHT) DUE TO THE ANATOMY BUT TRIED TO PRESERVE THE LEFT WITH THE BAIL OUT TO COVER AS DONE DURING THE PROCEDURE. THE ANGIO APPEARED TO SHOW THAT THEY HAVE SEALED THE AAA. THEY THEN HAD TO EVACUATE/SUCTION THE PATIENT'S ABDOMEN DUE TO BLOOD LOSS FILLING DURING THE PROCEDURE BUT WHEN THEY OPENED THE PATIENT'S ABDOMEN, THEY DISCOVERED THE AAA WAS NOT FULLY SEALED AND IT WAS STILL LEAKING BUT UNSURE OF THE ORIGIN. PHYSICIAN WAS UNABLE TO SEAL SURGICALLY AS THERE WAS TOO MUCH BLOOD LOSS AND THE PATIENT EXPIRED DURING PROCEDURE WITH PHYSICIAN STATING CAUSE OF DEATH BEING DUE TO THE UNSEALED RUPTURED AAA AND IS UNRELATED TO GORE GRAFTS. Manufacturer Narrative: W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22642431&pc=BZD | 2518422-2025-048389 | 2023-05-20 04:00 | Death | RESPIRONICS, INC. | 2025-07-29 04:00 | BZD | BIPAP S/T C SERIES | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A BIPAP S/T C SERIES CORE PACKAGE DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A PATIENT WITH A BIPAP S/T C SERIES CORE PACKAGE DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. PREVIOUSLY SUBMITTED REPORT WAS INCORRECTLY FILED WITH WRONG HEALTH IMPACT GRID. IN THIS REPORT, BOX H: HEALTH IMPACT GRID (1) HAS BEEN UPDATED AND CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22642573&pc=OZD | 1220648-2025-30806 | 2025-07-14 04:00 | Death | ABIOMED, INC. | 2025-07-29 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT EXPERIENCED LIMB ISCHEMIA AND SUBSEQUENTLY EXPIRED. THE PATIENT WAS ADMITTED WITH INCREASED SHORTNESS OF BREATH AND PALPITATIONS FOUND TO BE IN ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE. THE PATIENT WAS STARTED ON MEDICATION BY MOUTH FOR RATE CONTROL. THE FOLLOWING DAY THE PATIENT BEGAN TO HAVE MORE SHORTNESS OF BREATH AND REQUIRED INTUBATION AND CONTINUED TO DECOMPENSATE REQUIRING MULTIPLE VASOACTIVE DRIPS. THE PATIENT WAS TAKEN TO CATHETERIZATION LAB FOR AN IMPELLA CP PLACEMENT, WHICH WAS SUCCESSFULLY COMPLETED. SHOCK CALL WAS MADE, AND THE DECISION TO TRANSFER THE PATIENT TO ADVANCED HEART FAILURE MANAGEMENT FACILITY. THE NEXT DAY ON ARRIVAL, PULSES ARE ABSENT BILATERALLY AND ARTERIAL ULTRASOUND TO ASSESS FOR FLOW TO LOWER EXTREMITIES WERE ORDERED. EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS RECOMMENDED, AND THE PATIENT'S FAMILY DECLINED THE ECMO THERAPY. THE FAMILY OPTED FOR COMFORT MEASURES ONLY, AND THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22642938&pc=OZD | 1220648-2025-30808 | 2025-07-14 04:00 | Death | ABIOMED, INC. | 2025-07-29 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE IMPELLA CP WAS PLACED ON A PATIENT FOR PERSISTENT HYPOTENSION AND CARDIOGENIC SHOCK (CGS) SECONDARY TO RIGHT HEART FAILURE. DISCUSSION WAS HAD ABOUT THE DECREASED FLOWS, PHYSICIAN DECLINED TO INVESTIGATE FURTHER. PEEL AWAY SHEATH REMOVED AND REPOSITIONING SHEATH ADVANCED. FORWARD TENSION SUTURES APPLIED. DURING SUPPORT IT WAS REPORTED IMPELLA SITE CLEAN, DRY AND INTACT HOWEVER OTHER SITES WERE OOZING SIGNIFICANTLY AS WELL AS INCREASED DRAINAGE FROM STERNUM. 2 UNITS FRESH FROZEN PLASMA GIVEN AND 2 UNITS OF PLASMA GIVEN, OOZING SLOWED DOWN AS WELL AS DRAINAGE FROM WOUND VAC. AUTO TURNED OFF, SUCTION ALARMS CONTINUED UNTIL ON P4 FLOWING AT 1.8L, ECHO PERFORMED SHOWS IT MEASURING AT 3.4. LEFT VENTRICLE APEX HYPERTROPHIC, DISCUSSION HAD ABOUT FLOWS BEING LOWER THAN EXPECTED AND POTENTIALLY SHALLOWING IMPELLA TO SEE IF FLOWS IMPROVE IF SUCTION ALARMS CONTINUE. MULTIPLE SUCTION EVENTS OCCURRED THAT CORRECTED WITH REDUCING P-LEVEL AND STOPPED PULLING FLUID ON CONTINUOUS RENAL REPLACEMENT THERAPY. THEY CONTINUED TO HAVE ISSUES WITH SUCTION AND UNABLE TO ESCALATE P-LEVEL WITHOUT SUCTION. IMPELLA POSITION WAS 3.0 CM FROM THE ANNULUS IN-FLOW IN THE MID LEFT VENTRICLE SPACE. HE HAS HAD ONGOING OCCULT BLOOD LOSS, WE DID DISCUSS LABS MEASURING HEMOLYSIS IN THE SETTING OF ANURIA. HIS WHITE BLOOD CELLS REMAINED SIGNIFICANTLY ELEVATED ON BROAD SPECTRUM ANTIBIOTICS. HIS STERNAL WOUND HAD GROWN CANDIDA. SHOCK PICTURE WAS MIXED RIGHT SIDED CARDIOGENIC SHOCK WITH CONCURRENT SEPTIC SHOCK. THERE HAD ALSO BEEN A STEADY DECLINE IN PLATELETS. IMPELLA DRESSING CHANGED FOR OOZING. THE PATIENT WAS ACTIVELY DECOMPENSATING IN THE SETTING OF MAX INOTROPE AND PRESSOR SUPPORT. THE PATIENT CONTINUED TO DEMONSTRATE SEVERE OCCULT BLOOD LOSS, SEVERE ACIDOSIS, SEVERE RIGHT VENTRICLE DYSFUNCTION AS SUCH THE PATIENT PROGRESSED TO A SLOW CHEMICAL CODE PROGRESSING TO PULSELESS ELECTRICAL ACTIVITY (PEA) AND ADVANCED CARDIOVASCULAR LIFE SUPPORT ALGORITHMS. PHYSICIAN PLACED IMPELLA AT P-2 AND PULLED INTO THE DESCENDING AORTA PRIOR TO PULSELESS ELECTRICAL ACTIVITY, HE FELT THE IMPELLA MAY HAVE CONTRIBUTED TO ARRHYTHMIA AND POSSIBLY HEMOLYSIS. THEY DID NOT REQUEST TO SAVE THE DEVICE FOR RETURN FOR INTERNAL REVIEW. ULTIMATELY RESUSCITATION EFFORTS WERE SEEN AS FUTILE AND THE PATIENT WAS ALLOWED TO EXPIRE WITH FAMILY AT THE BEDSIDE. THE IMPELLA WAS UNPLUGGED AND TURNED OFF BY STAFF. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22643187&pc=DTZ | 8010762-2025-0000323 | 2025-07-19 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-07-29 04:00 | DTZ | TUBING SET | Partial Blockage; Infusion or Flow Problem; Pressure Problem | Insufficient Information |  |  | 1 | Event Description: THE EVENT OCCURRED IN USA. IT WAS REPORTED THAT THE HLS SET LOSS FLOW AND THAT THE DELTA PRESSURE WENT UP. THE CUSTOMER CONFIRMED THAT THE PATIENT DIED OF NON-CARDIAC ISSUES. THE AFFECTED PRODUCT WILL NOT BE RETURNED BY THE CUSTOMER. NEW INFORMATION RECEIVED ON (B)(6) 2025 THAT THE HLS SET WAS NOT REPLACED DURING TREATMENT SINCE IT WAS NOT A CARDIAC ISSUE PER ECHO. FURTHER CUSTOMER INFORMED SSU THAT CLOTS WERE OBSERVED IN THE PRODUCT. AS ANY PRESSURE READING ISSUE, CAN LEAD TO A PUMP STOP IF THE INTERVENTION IS SET BY THE USER, THEREFORE A REPORT IS REQUIRED. COMPLAINT ID (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THE DISPOSABLE WAS DISCARDED BY THE CUSTOMER. THEREFORE IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22643894&pc=OZP | 2032227-2025-220389 | 2025-06-21 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-07-29 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Coma; Loss of consciousness | P160017 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY AT THE HOSPITAL ON (B)(6) 2025. THE CUSTOMER WAS ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS DIABETIC COMA, HEART ISSUE AND THE CUSTOMER WAS UNCONSCIOUS. THE CUSTOMER¿S BLOOD GLUCOSE WAS UNKNOWN. THE EVENT INVOLVED PRODUCT(S) MMT-1884, MMT-332A, MMT-242A. THE CUSTOMER WAS WEARING THE INSULIN PUMP AT THE TIME OF DEATH. THE CUSTOMER WAS USING SENSORS. THE INSULIN PUMP WAS LAST WORN WAS UNKNOWN, PRIOR TO PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A, MMT-242A. MMT-1884 WAS REQUESTED AND CUSTOMER RESPONSE WAS THE DEVICE WILL BE RETURNED. THE CUSTOMER WILL DISCONTINUE USING THE DEVICE. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22644009&pc=LNH | 3002808157-2025-00019 | 2025-07-16 04:00 | Death | SIEMENS HEALTHINEERS AG | 2025-07-29 04:00 | LNH | MAGNETOM ESPREE | Use of Device Problem | Airway Obstruction | K041112 |  | 1 | Event Description: SIEMENS HEALTHINEERS BECAME AWARE OF A FATAL INCIDENT INVOLVING THE MAGNETOM ESPREE SYSTEM. THE INCIDENT REPORTEDLY OCCURRED ON (B)(6) 2025, WHEN A 61-YEAR-OLD MALE ENTERED AN MRI ROOM WITH A METALLIC OBJECT AROUND HIS NECK. THE OBJECT WAS ATTRACTED TO THE MRI¿S MAGNETIC FIELD, RESULTING IN ALLEGED INJURY/MEDICAL EVENT(S) THAT RESULTED IN HIS DEATH. Manufacturer Narrative: D4: DUE TO THE AGE OF THE REPORTED DEVICE, A UDI IS NOT AVAILABLE. E1: A REPORTING FACILITY CONTACT NAME WAS NOT PROVIDED FOR REPORTING. H3, H6: THE INVESTIGATION IS ONGOING; HOWEVER, IT IS CURRENTLY UNKNOWN IF SIEMENS WILL BE ABLE TO EVALUATE THE REPORTED SYSTEM. THE SYSTEM WAS BOUGHT USED AND WAS INSTALLED AT THE FACILITY BY A THIRD-PARTY AND SIEMENS DOES NOT HAVE A SERVICE CONTRACT WITH THE FACILITY. SIEMENS HAS REACHED OUT TO THE FACILITY BY TELEPHONE FOR ADDITIONAL INFORMATION AND LEFT A MESSAGE WITH CONTACT INFORMATION WITH THE RECEPTIONIST. A RETURN CALL FROM THE FACILITY MANAGER IS PENDING RECEIPT. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS PROVIDED UPON COMPLETION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22644025&pc=DTQ | 8010762-2025-0000324 | 2025-07-21 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-07-29 04:00 | DTQ | TUBING SET | Fluid/Blood Leak | Insufficient Information | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN CHINA DURING TREATMENT. IT WAS REPORTED THAT A PATIENT WAS CONNECTED TO THE ECMO DUE TO CARDIOGENIC DISEASE. A LEAKAGE ON THE QUICK VENT LUER CONNECTOR WAS REPORTED. THE CONNECTOR CAP WAS REPLACED BUT NOT SOLVED THE LEAKAGE. THE PATIENT PASSED WHILE NOT ON THE DEVICE. DUE TO THE LEAKAGE DURING USE AND THE HLS SET REPLACEMENT, A REPORT IS REQUIRED. COMPLAINT ID# (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE CHINESE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22644447&pc=OZD | 1220648-2025-30815 | 2025-07-15 04:00 | Death | ABIOMED, INC. | 2025-07-29 04:00 | OZD | IMPELLA CP | Material Split, Cut or Torn | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT AN IMPELLA CP WAS SUCCESSFULLY IMPLANTED IN A PATIENT HOWEVER, THE SLEEVE BECAME TORN AND LATER, THE PATIENT EXPIRED. Manufacturer Narrative: A2, A3, A4, AND A5 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22644492&pc=DTZ | 3013876692-2025-00060 | 2025-07-19 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-07-29 04:00 | DTZ | TUBING SET | Partial Blockage; Infusion or Flow Problem; Pressure Problem | Insufficient Information |  |  | 1 | Event Description: THE EVENT OCCURRED IN USA. IT WAS REPORTED THAT THE HLS SET LOSS FLOW AND THAT THE DELTA PRESSURE WENT UP. THE CUSTOMER CONFIRMED THAT THE PATIENT DIED OF NON-CARDIAC ISSUES. THE AFFECTED PRODUCT WILL NOT BE RETURNED BY THE CUSTOMER. NEW INFORMATION RECEIVED ON 2025-07-23 THAT THE HLS SET WAS NOT REPLACED DURING TREATMENT SINCE IT WAS NOT A CARDIAC ISSUE PER ECHO. FURTHER CUSTOMER INFORMED SSU THAT CLOTS WERE OBSERVED IN THE PRODUCT. AS ANY PRESSURE READING ISSUE, CAN LEAD TO A PUMP STOP IF THE INTERVENTION IS SET BY THE USER, THEREFORE A REPORT IS REQUIRED. COMPLAINT ID (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THE DISPOSABLE WAS DISCARDED BY THE CUSTOMER. THEREFORE, IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22644759&pc=DTQ | 3013876692-2025-00061 | 2025-07-21 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-07-29 04:00 | DTQ | TUBING SET | Fluid/Blood Leak | Insufficient Information | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN CHINA DURING TREATMENT. IT WAS REPORTED THAT A PATIENT WAS CONNECTED TO THE ECMO DUE TO CARDIOGENIC DISEASE. A LEAKAGE ON THE QUICK VENT LUER CONNECTOR WAS REPORTED. THE CONNECTOR CAP WAS REPLACED BUT NOT SOLVED THE LEAKAGE. THE PATIENT PASSED WHILE NOT ON THE DEVICE. DUE TO THE LEAKAGE DURING USE AND THE HLS SET REPLACEMENT, A REPORT IS REQUIRED. COMPLAINT ID#: (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE CHINESE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22644830&pc=OZD | 1220648-2025-30818 | 2025-07-18 04:00 | Death | ABIOMED, INC. | 2025-07-29 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hematoma; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT FOLLOWING IMPELLA 5.5 IMPLANT THE PATIENT EXPERIENCED SOME BLEEDING FROM THEIR MOUTH. THE PATIENT WAS TRANSFUSED TWO UNITS OF BLOOD WHILE PLATELET LEVELS BEGAN TO TREND DOWNWARD. THE FOLLOWING DAY, A CT SCAN WAS PERFORMED DUE TO ACUTE CHANGES IN THE PATIENT'S NEURO STATUS AND A INTRAPARENCHYMAL HEMATOMA WAS DISCOVERED. HEPARIN WAS DISCONTINUED AS FURTHER ACTIONS WERE BEING DISCUSSED. THE PATIENT'S FAMILY WAS CALLED IN RESPONSE TO THE HEMORRHAGIC STROKE AND THE DECISION WAS MADE TO WITHDRAW CARE. THE PATIENT PASSED AWAY SHORTLY AFTER. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: A5, ETHNICITY, IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22644971&pc=OZD | 1220648-2025-30819 | 2025-07-16 04:00 | Death | ABIOMED, INC. | 2025-07-29 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT FOLLOWING AN IMPELLA CP IMPLANT, THE PATIENT BEGAN TO DEVELOP HEMOLYSIS. FOUR UNITS OF BLOOD WERE GIVEN TO TREAT THE CONDITION. THE FOLLOWING DAY, THE PATIENT WAS DIAGNOSED WITH ISCHEMIA IN BILATERAL LIMBS. THE PUMP WAS REMOVED. A FASCIOTOMY WAS PERFORMED AND THE PATIENT WAS ESCALATED TO IMPELLA 5.5 SUPPORT. LATER THAT SAME DAY, THE PHYSICIAN BELIEVED THAT THE RIGHT LEG WAS DEAD. ADDITIONAL INTERVENTION WAS NOT PERFORMED. THE PROCEDURAL OUTCOME WAS THE PATIENT SURVIVED THE SURGERY.Event Description: THE COMPLAINANT REPORTED A PATIENT (UNKNOWN RACE) WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY IMPLANTED WITH AN IMPELLA CP DEVICE FOLLOWED BY AN IMPELLA RP FLEX DEVICE FOR BIPELLA MECHANICAL CIRCULATORY SUPPORT AND FOLLOWING BIPELLA IMPLANT, THE PATIENT BEGAN TO DEVELOP HEMOLYSIS. FOUR UNITS OF BLOOD WERE GIVEN TO TREAT THE CONDITION. THE FOLLOWING DAY, THE PATIENT WAS DIAGNOSED WITH ISCHEMIA IN BOTH EXTREMITIES. LACTIC REMAINS ELEVATED, CONTRIBUTING TO LIMB ISCHEMIA. THE IMPELLA CP PUMP WAS REMOVED, A FASCIOTOMY WAS PERFORMED, AND THE PATIENT WAS ESCALATED TO IMPELLA 5.5 SUPPORT, WHICH WAS PLACED WITH SOME DIFFICULTY. SUPPORT RAMPED UP AND POSITION WAS STABLE. LATER THAT SAME DAY, THE PHYSICIAN BELIEVED THAT THE RIGHT LEG WAS DEAD. ADDITIONAL INTERVENTION WAS NOT PERFORMED HOWEVER. ALSO, IT WAS NOTED THE BELLY KEPT GETTING BIGGER, A LOT OF BLOOD SEEN AROUND THE LIVER ON TRANSESOPHAGEAL ECHOCARDIOGRAM. IT WAS NOTED AT THAT SAID TIME, THERE WAS NO PLAN FOR ANY INTERVENTION. DUE TO THE PATIENT'S PROGNOSIS, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: GIVEN THE DETAILS OF THE EVENT, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE OUTCOME. HEMOLYSIS ALONG WITH LIMB ISCHEMIA PRESENT THAT PROGRESSED AND ALONG WITH VASCULAR SURGICAL REPAIR RELATED TO THE IMPELLA CP DEVICE REMOVAL. REFER MANUFACTURER REPORT NUMBER 1220648-2025-45796 FOR THE REPORT ON THE IMPELLA RP FLEX. THE EVENT WAS REVISITED, AND THE DEMISE OUTCOME DETERMINED TO BE ASSOCIATED WITH THE IMPELLA CP. CONSEQUENTLY, UPDATES WERE MADE TO THE FOLLOWING SECTION(S): B2: TO INCLUDE DISABILITY OR PERMANENT DAMAGE/DEATH/DATE OF DEATH. B5: TO INCLUDE BIPELLA SUPPORT AND DEMISE OUTCOME. H1: CHANGE TO DEATH TYPE OF REPORTABLE EVENT. H6: HEALTH EFFECT CLINICAL CODE(S)/IMPACT CODE(S)/DEVICE PROBLEM CODE(S)/COMPONENT CODE(S) REPOPULATED BASED ON EVENT WITH A DEMISE OUTCOME ASSOCIATION. H10: RELATED REPORT NUMBER. NOTE: H6 INVESTIGATION CODING AS SUBMITTED WITH INITIAL REPORT REMAINS UNCHANGED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645058&pc=DSQ | 2916596-2025-04996 | 2025-07-21 04:00 | Death | THORATEC CORPORATION | 2025-07-29 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-ORGAN COMPLICATIONS AND THE PATIENT WAS PALLIATED. THIS WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE PATIENT BECAME A DESTINATION THERAPY PATIENT AND WAS PALLIATED. THIS WAS NOT THOUGHT TO BE DEVICE RELATED AND THERE WERE NO CONCERNS WITH DEVICE FUNCTION. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645099&pc=OZD | 1220648-2025-30821 | 2024-11-24 05:00 | Death | ABIOMED, INC. | 2025-07-29 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND EXPERIENCED ACCESS SITE BLEEDING. PERCLOSE SUTURES WERE USED AND THE PUMP WAS EXPLANTED. LATER, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645240&pc=DSQ | 2916596-2025-05062 | 2025-07-15 04:00 | Death | THORATEC CORPORATION | 2025-07-29 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645470&pc=OZD | 1220648-2025-30824 | 2025-07-16 04:00 | Death | ABIOMED, INC. | 2025-07-29 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXHIBITED LOWER EXTREMITIES HAD BEEN DUSKY SINCE ADMISSION, AND THE PATIENT HAD BEEN COMPLAINING OF DISCOMFORT. ULTRASOUND AND DOPPLER CONFIRMED BLOOD FLOW, HOWEVER, BLOOD CULTURE RESULTS WERE POSITIVE. ULTIMATELY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: A5, ETHNICITY, IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS. ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645675&pc=DSQ | 2916596-2025-05061 | 2025-06-02 04:00 | Death | THORATEC CORPORATION | 2025-07-29 04:00 | DSQ | VENTRICULAR (ASSIST) BYPASS | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Encephalopathy | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FROM A COMBINATION OF PSEUDOMONAS INFECTION, ACUTE KIDNEY INJURY, ENCEPHALOPATHY, AND FAILURE TO THRIVE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENTS; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU AND THE HEARTMATE II PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, PUMP POCKET), RENAL DYSFUNCTION, OTHER NEUROLOGICAL EVENTS (NOT STROKE-RELATED), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS INFECTION AND NEUROLOGICAL DYSFUNCTION AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE II LVAS. SEVERAL SECTIONS OF THE HMII IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645744&pc=NIQ | 2124215-2025-49268 | 2025-07-07 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-29 04:00 | NIQ | PROMUS ELITE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Chest Pain; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH CORONARY ARTERY DISEASE AND UNDERWENT PERCUTANEOUS CORONARY INTERVENTION. THE TARGET LESIONS WERE LOCATED IN THE 80% STENOSED AND MODERATELY CALCIFIED PROXIMAL LEFT ANTERIOR DESCENDING (LAD) ARTERY AND LEFT CIRCUMFLEX ARTERY (LCX). A 20 X 3.50MM PROMUS ELITE DRUG-ELUTING STENT WAS SUCCESSFULLY DEPLOYED IN PROXIMAL LAD FOLLOWED BY A 2.75X28MM NON-BOSTON SCIENTIFIC STENT IN THE LCX. WHILE IN THE CORONARY CARE UNIT (CCU), THE PATIENT DEVELOPED SUDDEN HYPOTENSION AND CHEST PAIN, FOLLOWED SHORTLY BY CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS PERFORMED IMMEDIATELY; HOWEVER, THE PATIENT COULD NOT BE REVIVED AND WAS PRONOUNCED DECEASED. THE OFFICIAL CAUSE OF DEATH WAS CARDIAC ARREST. NO AUTOPSY WAS PERFORMED. Manufacturer Narrative: E1 INITIAL REPORTER FIRST NAME: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645889&pc=DSQ | 3007042319-2025-01533 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Disconnection; Use of Device Problem; Battery Problem | Cardiac Arrest | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE CONTROLLER EXHIBITED A CONTROLLER FAULT ALARM. THE PATIENT ATTEMPTED TO EXCHANGE THE CONTROLLER AND ARRESTED WHILE THE VENTRICULAR ASSIST DEVICE (VAD) DRIVELINE WAS DISCONNECTED. EMERGENCY MEDICAL SERVICES INITIATED CHEST COMPRESSIONS AND THE PATIENT WAS HOSPITALIZED. REVIEW OF LOG FILES DATA ALSO REVEALED THAT THE CONTROLLER HAD EXHIBITED A PRIOR LOSS OF POWER, WHICH WAS ATTRIBUTED TO AN ACCIDENTAL DOUBLE DISCONNECTION OF POWER SOURCES. THE CONTROLLER FAULT ALARM WAS ATTRIBUTED TO DEPLETION OF THE INTERNAL CONTROLLER BATTERY. THE CONTROLLER FAULT ALARM WAS PERMANENTLY SILENCED. IT WAS NOTED THAT THE VAD WAS OFF FOR APPROXIMATELY 50 MINUTES. THE PATIENT SUBSEQUENTLY EXPIRED. THE VAD REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) AND THE CONTROLLER WITH UNKNOWN SERIAL NUMBERS WERE NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE SERIAL NUMBER IS UNKNOWN. LOG FILES WERE NOT AVAILABLE FOR REVIEW. AS A RESULT, THE REPORTED EVENTS COULD NOT BE CONFIRMED. APPLICABLE RISK DOCUMENTATION, EXPERIENCE WITH EVENTS OF SIMILAR CIRCUMSTANCES, AND THE AVAILABLE INFORMATION WERE CONSIDERED; A POSSIBLE ROOT CAUSE OF THE REPORTED CONTROLLER FAULT ALARM EVENT MAY BE ATTRIBUTED, BUT NOT LIMITED, TO A REDUCED CHARGE CAPACITY OF THE INTERNAL NICKEL-METAL HYDRIDE (NIMH) BATTERY AND/OR A FAULTY INTERNAL BATTERY CHARGER INTEGRATED CIRCUIT. A POSSIBLE ROOT CAUSE OF THE REPORTED CONTROLLER LOSS OF POWER EVENT CAN BE ATTRIBUTED, BUT NOT LIMITED, TO THE REPORTED DISCONNECTION OF BOTH POWER SOURCES FROM THE CONTROLLER AS DESCRIBED IN THE EVENT DETAILS. BASED ON THE LIMITED AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED DEATH EVENT. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING CARDIAC ARREST. PER THE INSTRUCTIONS FOR USE, CARDIOPULMONARY ARREST AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGUL ANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM - UNKNOWN CONTROLLER. D9: NO. H3: NO. H4: MFG DATE: H5: NO. H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645890&pc=DSQ | 3007042319-2025-01539 | 2016-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645893&pc=DSQ | 3007042319-2025-01547 | 2021-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645895&pc=DSQ | 3007042319-2025-01563 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE SERIAL NUMBER WAS UNAVAILABLE AND THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. REVIEW OF LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. THERE IS INSUFFICIENT INFORMATION REGARDING THE REPORTED EVENT. AS A RESULT, THE REPORTED EVENT COULD NOT BE CONFIRMED AND A POSSIBLE CAUSE OF THE REPORTED EVENT COULD NOT BE ESTABLISHED. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645900&pc=DSQ | 3007042319-2025-01573 | 2018-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645901&pc=DSQ | 3007042319-2025-01565 | 2022-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645904&pc=DSQ | 3007042319-2025-01559 | 2021-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645907&pc=DSQ | 3007042319-2025-01567 | 2023-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645911&pc=DSQ | 3007042319-2025-01572 | 2023-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE SERIAL NUMBER WAS UNAVAILABLE AND THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. REVIEW OF LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. THERE IS INSUFFICIENT INFORMATION REGARDING THE REPORTED EVENT. AS A RESULT, THE REPORTED EVENT COULD NOT BE CONFIRMED AND A POSSIBLE CAUSE OF THE REPORTED EVENT COULD NOT BE ESTABLISHED. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645912&pc=DSQ | 3007042319-2025-01586 | 2022-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE SERIAL NUMBER WAS UNAVAILABLE AND THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. REVIEW OF LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. THERE IS INSUFFICIENT INFORMATION REGARDING THE REPORTED EVENT. AS A RESULT, THE REPORTED EVENT COULD NOT BE CONFIRMED AND A POSSIBLE CAUSE OF THE REPORTED EVENT COULD NOT BE ESTABLISHED. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645913&pc=DSQ | 3007042319-2025-01579 | 2023-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO END STAGE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645915&pc=DSQ | 3007042319-2025-01581 | 2019-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645916&pc=DSQ | 3007042319-2025-01580 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645918&pc=DSQ | 3007042319-2025-01598 | 2023-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO CONGESTIVE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO CONGESTIVE HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645922&pc=DSQ | 3007042319-2025-01604 | 2020-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645924&pc=DSQ | 3007042319-2025-01540 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645926&pc=DSQ | 3007042319-2025-01535 | 2017-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED CIRCULATORY ISSUE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) WITH UNKNOWN LOT NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE, AND THE LOT NUMBER IS UNKNOWN. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645930&pc=DSQ | 3007042319-2025-01549 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED MAJOR INFECTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE STERILITY CERTIFICATE COULD NOT BE PERFORMED SINCE THE SERIAL NUMBER IS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS INFECTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645933&pc=DSQ | 3007042319-2025-01548 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Bacterial Infection | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR A BACTERIAL BLOOD STREAM INFECTION REQUIRING INTRAVENOUS (IV) MEDICATION AND SURGICAL INTERVENTION. THE PATIENT SUBSEQUENTLY EXPIRED. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE STERILITY CERTIFICATE COULD NOT BE PERFORMED SINCE THE SERIAL NUMBER IS UNKNOWN. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645937&pc=DSQ | 3007042319-2025-01566 | 2022-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO END STAGE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645939&pc=DSQ | 3007042319-2025-01564 | 2022-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645940&pc=DSQ | 3007042319-2025-01558 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO END STAGE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645946&pc=DSQ | 3007042319-2025-01585 | 2022-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED CAUSES. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS UNSPECIFIED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645950&pc=DSQ | 3007042319-2025-01592 | 2019-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645952&pc=DSQ | 3007042319-2025-01595 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Liver Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO HEPATIC DYSFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEPATIC DYSFUNCTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645954&pc=DSQ | 3007042319-2025-01588 | 2019-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645958&pc=DSQ | 3007042319-2025-01615 | 2019-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED CAUSES. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS UNSPECIFIED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645960&pc=DSQ | 3007042319-2025-01582 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645963&pc=DSQ | 3007042319-2025-01602 | 2017-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED CAUSES. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS UNSPECIFIED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645965&pc=DSQ | 3007042319-2025-01600 | 2020-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645966&pc=DSQ | 3007042319-2025-01594 | 2022-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO END STAGE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645969&pc=DSQ | 3007042319-2025-01608 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO END STAGE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645970&pc=DSQ | 3007042319-2025-01616 | 2014-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645975&pc=DSQ | 3007042319-2025-01607 | 2022-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Cardiomyopathy | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO ISCHEMIC CARDIOMYOPATHY. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS ISCHEMIC CARDIOMYOPATHY. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645977&pc=DSQ | 3007042319-2025-01619 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645978&pc=DSQ | 3007042319-2025-01623 | 2019-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645984&pc=DSQ | 3007042319-2025-01630 | 2019-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645987&pc=DSQ | 3007042319-2025-01629 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645991&pc=DSQ | 3007042319-2025-01638 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645993&pc=DSQ | 3007042319-2025-01636 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22645998&pc=DSQ | 3007042319-2025-01646 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646011&pc=DSQ | 3007042319-2025-01657 | 2020-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO END STAGE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646012&pc=DSQ | 3007042319-2025-01658 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO RESPIRATORY FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO RESPIRATORY FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646016&pc=DSQ | 3007042319-2025-01666 | 2023-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646017&pc=DSQ | 3007042319-2025-01660 | 2015-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646026&pc=DSQ | 3007042319-2025-01670 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646028&pc=DSQ | 3007042319-2025-01668 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED CAUSES. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS UNSPECIFIED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646035&pc=DSQ | 3007042319-2025-01676 | 2020-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646036&pc=NGV | 2124215-2025-50080 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-29 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade |  | E2017028 | 3 | Event Description: THIS REPORT SUMMARIZES 3 DEATH EVENTS FOR CARDIAC TAMPONADE. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. AGE RANGE: 60 TO 88 YEARS. PATIENT SEX: 2 FEMALES; 1 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 8 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646043&pc=NGV | 2124215-2025-50079 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-29 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR UNEXPECTED MEDICAL INTERVENTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT, BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 229 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646044&pc=DSQ | 3007042319-2025-01686 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED CAUSES. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS UNSPECIFIED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646060&pc=DSQ | 3007042319-2025-01685 | 2017-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646067&pc=DSQ | 3007042319-2025-01695 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED MAJOR INFECTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE STERILITY CERTIFICATE COULD NOT BE PERFORMED SINCE THE SERIAL NUMBER IS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS INFECTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646073&pc=DSQ | 3007042319-2025-01700 | 2019-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO RIGHT HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO RIGHT HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646074&pc=DSQ | 3007042319-2025-01697 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO RIGHT HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO RIGHT HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646077&pc=DSQ | 3007042319-2025-01704 | 2022-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO END STAGE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646078&pc=OZD | 1220648-2025-30827 | 2025-07-16 04:00 | Death | ABIOMED, INC. | 2025-07-29 04:00 | OZD | IMPELLA CP | Failure to Advance | Ventricular Fibrillation; Asystole | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT DURING THE INITIAL IMPLANT ATTEMPT OF AN IMPELLA CP PUMP, THE PATIENT WENT ASYSTOLIC AND CODED. THE PATIENT RECEIVED CARDIOPULMONARY RESUSCITATION FOR OVER AN HOUR AND RECEIVED SEVERAL SHOCKS FOR VENTRICULAR FIBRILLATION. IT WAS BELIEVED BY THE PHYSICIAN THAT A POSSIBLE AORTIC ANEURYSM WAS PREVENTING ADVANCEMENT OF THE PUMP. THE PATIENT EVENTUALLY PASSED AWAY AFTER RETURN OF SPONTANEOUS CIRCULATION COULD NOT BE ACHIEVED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646080&pc=DSQ | 3007042319-2025-01707 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646082&pc=DSQ | 3007042319-2025-01705 | 2016-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646084&pc=DSQ | 3007042319-2025-01712 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646086&pc=DSQ | 3007042319-2025-01711 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646087&pc=DSQ | 3007042319-2025-01714 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO RESPIRATORY FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO RESPIRATORY FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646088&pc=DSQ | 3007042319-2025-01709 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO END STAGE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646090&pc=DSQ | 3007042319-2025-01722 | 2023-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646092&pc=DSQ | 3007042319-2025-01717 | 2022-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646094&pc=DSQ | 3007042319-2025-01720 | 2021-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO END STAGE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646095&pc=DSQ | 3007042319-2025-01723 | 2018-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO END STAGE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646103&pc=DSQ | 3007042319-2025-01731 | 2019-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646105&pc=DSQ | 3007042319-2025-01733 | 2018-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED CAUSES. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS UNSPECIFIED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646108&pc=DSQ | 3007042319-2025-01734 | 2020-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646109&pc=DSQ | 3007042319-2025-01732 | 2019-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646111&pc=DSQ | 3007042319-2025-01741 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED. THE PATIENT SUBSEQUENTLY EXPIRED DUE TO UNSPECIFIED CAUSES. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS UNSPECIFIED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646119&pc=DSQ | 3007042319-2025-01744 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO RIGHT HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO RIGHT HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646121&pc=DSQ | 3007042319-2025-01748 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO CONGESTIVE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO CONGESTIVE HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646122&pc=DSQ | 3007042319-2025-01745 | 2022-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO CONGESTIVE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO CONGESTIVE HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646125&pc=DSQ | 3007042319-2025-01751 | 2021-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646126&pc=DSQ | 3007042319-2025-01743 | 2018-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646129&pc=DSQ | 3007042319-2025-01747 | 2020-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO RIGHT HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO RIGHT HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646134&pc=DSQ | 3007042319-2025-01757 | 2018-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO RIGHT HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO RIGHT HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646135&pc=DSQ | 3007042319-2025-01756 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO CONGESTIVE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO CONGESTIVE HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646141&pc=DSQ | 3007042319-2025-01763 | 2019-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646144&pc=DSQ | 3007042319-2025-01764 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646146&pc=DSQ | 3007042319-2025-01759 | 2019-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646147&pc=DSQ | 3007042319-2025-01760 | 2022-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646150&pc=DSQ | 3007042319-2025-01765 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646151&pc=DSQ | 3007042319-2025-01768 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED CAUSES. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS UNSPECIFIED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646157&pc=DSQ | 3007042319-2025-01772 | 2022-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED CAUSES. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS UNSPECIFIED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646159&pc=DSQ | 3007042319-2025-01774 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Cardiomyopathy | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO ISCHEMIC CARDIOMYOPATHY. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS ISCHEMIC CARDIOMYOPATHY. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646163&pc=DSQ | 3007042319-2025-01775 | 2023-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646165&pc=DSQ | 3007042319-2025-01790 | 2023-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646177&pc=DSQ | 3007042319-2025-01794 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE SERIAL NUMBER WAS UNAVAILABLE AND THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. REVIEW OF LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. THERE IS INSUFFICIENT INFORMATION REGARDING THE REPORTED EVENT. AS A RESULT, THE REPORTED EVENT COULD NOT BE CONFIRMED AND A POSSIBLE CAUSE OF THE REPORTED EVENT COULD NOT BE ESTABLISHED. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646181&pc=DSQ | 3007042319-2025-01795 | 2021-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646182&pc=DSQ | 3007042319-2025-01796 | 2018-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646183&pc=DSQ | 3007042319-2025-01799 | 2021-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646184&pc=DSQ | 3007042319-2025-01798 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE SERIAL NUMBER WAS UNAVAILABLE AND THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS DUE TO AN UNSPECIFIED DEVICE MALFUNCTION. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. REVIEW OF LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE FOR ANALYSIS. THERE IS INSUFFICIENT INFORMATION REGARDING THE REPORTED EVENT. AS A RESULT, THE REPORTED EVENT COULD NOT BE CONFIRMED AND A POSSIBLE CAUSE OF THE REPORTED EVENT COULD NOT BE ESTABLISHED. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646186&pc=DSQ | 3007042319-2025-01813 | 2020-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646190&pc=DSQ | 3007042319-2025-01803 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED. THE PATIENT SUBSEQUENTLY EXPIRED DUE TO WITHDRAWAL OF SUPPORT. THE VENTRICULAR AS SIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER WITHDRAWAL OF SUPPORT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646194&pc=DSQ | 3007042319-2025-01815 | 2022-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Unspecified Nervous System Problem | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO UNSPECIFIED NEUROLOGICAL DYSFUNCTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646201&pc=DSQ | 3007042319-2025-01824 | 2016-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646202&pc=DSQ | 3007042319-2025-01817 | 2023-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Cardiomyopathy | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO ISCHEMIC CARDIOMYOPATHY. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS ISCHEMIC CARDIOMYOPATHY. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646204&pc=DSQ | 3007042319-2025-01820 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO CARDIAC ARRHYTHMIA. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER CARDIAC ARRHYTHMIA. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646205&pc=DSQ | 3007042319-2025-01818 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO CONGESTIVE HEART FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO CONGESTIVE HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646207&pc=DSQ | 3007042319-2025-01833 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Hemolysis; Renal Failure; Thrombosis/Thrombus | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR A SUSPECTED VENTRICULAR ASSIST DEVICE (VAD) THROMBUS, ACUTE STAGE THREE RENAL FAILURE AND ANURIA, HEMOLYSIS INCLUDING HEMOGLOBINURIA, AND ABNORMAL PUMP PARAMETERS REQUIRING INTRAVENOUS (IV) MEDICATION. THE VENTRICULAR ASSIST DEVICE (VAD) WAS DECOMMISSIONED, TURNED OFF, AND EXPLANTED. THE PATIENT SUBSEQUENTLY EXPIRED. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE CONTROLLER LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE. AS A RESULT, THE REPORTED ABNORMAL PUMP PARAMETERS COULD NOT BE CONFIRMED. THE REPORTED THROMBUS EVENT COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PUMP WAS DECOMMISSIONED, TURNED OFF, EXPLANTED, AND THE PATIENT SUBSEQUENTLY EXPIRED. BASED ON RISK DOCUMENTATION, MULTIPLE FACTORS MAY HAVE CONTRIBUTED TO THE ABNORMAL PUMP PARAMETERS EVENT INCLUDING BUT NOT LIMITED TO THROMBUS FORMATION/INGESTION, THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646216&pc=DSQ | 3007042319-2025-01835 | 2025-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Renal Failure; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR RIGHT HEART FAILURE, ACUTE STAGE TWO RENAL FAILURE, AND A BACTERIAL BLOOD STREAM INFECTION REQUIRING INTRAVENOUS (IV) MEDICATION. THE PATIENT SUBSEQUENTLY EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) WAS EXPLANTED. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE STERILITY CERTIFICATE COULD NOT BE PERFORMED SINCE THE SERIAL NUMBER IS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646218&pc=DSQ | 3007042319-2025-01834 | 2025-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Pneumonia; Gastrointestinal Hemorrhage | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR A GASTROINTESTINAL BLEED (GIB) AND PNEUMONIA REQUIRING INTRAVENOUS (IV) MEDICATION. THE PATIENT SUBSEQUENTLY EXPIRED. THE VENTRICULAR ASSIST DEVICE (VAD) WAS TURNED OFF, DECOMMISSIONED, AND EXPLANTED. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE STERILITY CERTIFICATE COULD NOT BE PERFORMED SINCE THE SERIAL NUMBER IS UNKNOWN. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER THE BLEEDING AND PNEUMONIA EVENT. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646224&pc=DSQ | 3007042319-2025-01841 | 2025-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Hemolysis; Thrombosis/Thrombus | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) EXHIBITED HIGH POWER WITH HIGH WATT ALARMS. A VAD THROMBUS WAS SUSPECTED. IT WAS NOTED THAT THE PATIENT WAS ON HOSPICE AND SUBSEQUENTLY EXPIRED DUE TO HEMOLYSIS. THE VAD REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE CONTROLLER LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE. AS A RESULT, THE REPORTED HIGH POWER EVENT COULD NOT BE CONFIRMED. THE REPORTED THROMBUS EVENT COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE RISK DOCUMENTATION, POSSIBLE ROOT CAUSES OF THE HIGH POWER EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING, BUT NOT LIMITED, TO EXTERNAL FACTORS SUCH AS THROMBUS FORMATION/INGESTION, INAPPROPRIATE PUMP ROTATIONAL SPEED, AND/OR PATIENT FACTORS. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEMOLYSIS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULIN ANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646225&pc=DSQ | 3007042319-2025-01845 | 2018-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Pneumonia; Cardiac Tamponade | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED PNEUMONIA REQUIRING INTRAVENOUS (IV) MEDICATION AND CARDIAC TAMPONADE THAT REQUIRED PER ICARDIAL DRAINAGE WITH A HEART CATH. THE PATIENT SUBSEQUENTLY EXPIRED. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE STERILITY CERTIFICATE COULD NOT BE PERFORMED SINCE THE SERIAL NUMBER IS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER PNEUMONIA AND CARDIAC TAMPONADE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646226&pc=DSQ | 3007042319-2025-01843 | 2025-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Intracranial Hemorrhage | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR A BACTERIAL BLOOD STREAM INFECTION AND A SUBARACHNOID HEMORRHAGE RESULTING IN MODERATE DISABILITY AND OVERT CENTRAL NERVOUS SYSTEM (CNS) INJURY THAT WAS DIAGNOSED USING COMPUTERIZED TOMOGRAPHY (CT). THE PATIENT SUBSEQUENTLY EXPIRED. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE STERILITY CERTIFICATE COULD NOT BE PERFORMED SINCE THE SERIAL NUMBER IS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER INFECTION AND SUBARACHNOID HEMORRHAGE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646235&pc=LYJ | 1644487-2025-10417 | 2025-07-02 04:00 | Death | HOUSTON | 2025-07-29 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED FROM A CORONER'S OFFICE THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS CURRENTLY UNKNOWN. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646247&pc=DSQ | 3007042319-2025-01847 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A CEREBRAL HEMORRHAGE RESULTING IN OVERT CENTRAL NERVOUS SYSTEM (CNS) INJURY THAT WAS DIAGNOSED USING COMPUTERIZED TOMOGRAPHY (CT). THE VENTRICULAR ASSIST DEVICE (VAD) WAS DECOMMISSIONED AND TURNED OFF. THE PATIENT SUBSEQUENTLY EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) WAS EXPLANTED. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646253&pc=DSQ | 3007042319-2025-01842 |  | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR A SUSPECTED VENTRICULAR ASSIST DEVICE (VAD) INFLOW CANNULA THROMBUS AND ABNORMAL PUMP PARAMETERS REQUIRING INTRAVENOUS MEDICATION. THE VAD WAS DEACTIVATED AND THE PATIENT SUBSEQUENTLY EXPIRED DUE TO END STAGE HEART FAILURE. THE VAD REMAINS IMPLANTED IN THE PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. REVIEW OF THE CONTROLLER LOG FILES COULD NOT BE CONDUCTED SINCE LOG FILES WERE NOT AVAILABLE. AS A RESULT, THE REPORTED ABNORMAL PUMP PARAMETERS COULD NOT BE CONFIRMED. THE REPORTED THROMBUS EVENT COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON RISK DOCUMENTATION, MULTIPLE FACTORS MAY HAVE CONTRIBUTED TO THE ABNORMAL PUMP PARAMETERS EVENT INCLUDING BUT NOT LIMITED TO THROMBUS FORMATION/INGESTION, THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO HEART FAILURE. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE C OURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646257&pc=DSQ | 3007042319-2025-01825 | 2020-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED DUE TO MULTISYSTEM ORGAN FAILURE. THE VENTRICULAR ASSIST DEVICE (VAD) REMAINS IN PATIENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646261&pc=DSQ | 3007042319-2025-01852 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-07-29 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED FOR CARDIAC ARRHYTHMIA. THE PATIENT SUBSEQUENTLY EXPIRED. THE VENTRICULAR ASSIST D EVICE (VAD) WAS EXPLANTED. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. Manufacturer Narrative: PRODUCT EVENT SUMMARY: ONE (1) PUMP WITH UNKNOWN SERIAL NUMBER WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE AND THE DEVICE SERIAL NUMBER WAS UNKNOWN. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER CARDIAC ARRHYTHMIA. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. THIS EVENT WAS REPORTED IN THE Q1 2025 INTERMACS DATA REGISTRY THAT TRACKS CLINICAL OUTCOMES OF PATIENTS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT. THE DATA REGISTRY DOES NOT CONTAIN DEVICE IDENTIFYING INFORMATION OR EVENT DATE AND THEREFORE CANNOT BE CORRELATED TO ANY PREVIOUSLY RECEIVED REPORT OF THE EVENT. BASED ON THE PROVIDED DATA, DEVICE ANALYSIS WILL NOT BE POSSIBLE AND NO FURTHER INFORMATION WILL LIKELY BE MADE AVAILABLE CONCERNING THE EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646469&pc=NPT | 2025587-2025-05326 | 2025-02-16 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Sepsis; Ulcer; Loss of consciousness; Gastrointestinal Hemorrhage | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT 12 DAYS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT PRESENTED WITH MELENA. ENDOSCOPIC EXAMINATION WAS PERFORMED THAT REVEALED A RECTAL ULCER, AND HEMOSTASIS WAS ACHIEVED. SUBSEQUENTLY, SEPSIS WAS SUSPECTED AS THE PATIENT DEVELOPED TEMPORARILY IMPAIRED CONSCIOUSNESS AND HYPOTENSION. THEREFORE, THE PATIENT WAS ADMINISTERED ANTIBIOTIC TREATMENT. APPROXIMATELY ONE MONTH FOLLOWING THE IMPLANT OF THE VALVE, A CARDIAC ARREST OCCURRED AND THE PATIENT DIED. THE CAUSE OF DEATH WAS REPORTED TO BE DUE TO A GASTROINTESTINAL HEMORRHAGE. PER THE PHYSICIAN, THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND VALVE OR IMPLANT PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646474&pc=NGV | 2124215-2025-49992 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-29 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR UNSPECIFIED INFECTION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646545&pc=DSQ | 2916596-2025-05068 | 2025-07-09 04:00 | Death | THORATEC CORPORATION | 2025-07-29 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Intracranial Hemorrhage; Sepsis; Liver Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO INTRAPARENCHYMAL HEMORRHAGE. IT WAS SAID THE DEVICE FUNCTIONED AS INTENDED.Event Description: IT WAS LATER REPORTED THAT THE INTRAPARENCHYMAL HEMORRHAGE WAS DUE TO CRITICALLY ILL-BACTEREMIA. IT WAS ALSO SAID THAT THE STROKE WAS THOUGHT TO BE RELATED TO CHANGES TO PATIENT CONDITION OR ANTICOAGULATION STATUS. THE PATIENT WAS SAID TO HAVE BEEN SUFFERING FROM ONGOING COMPLICATIONS OF SEPSIS-LIVER FAILURE. IT WAS ALSO SAID THAT THE PATIENT "STOPPED MOVING/ FOLLOWING COMMANDS". DEATH WAS SAID NOT TO HAVE BEEN CONSIDERED DEVICE RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. A SPECIFIC CAUSE FOR THE REPORT OF BACTEREMIA AND SEPSIS COULD NOT BE CONCLUSIVELY DETERMINED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS STROKE, BLEEDING, LOCALIZED INFECTION, SEPSIS, HEPATIC DYSFUNCTION, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. THIS SECTION ALSO INCLUDES INFORMATION REGARDING HOW TO PREVENT AND CONTROL INFECTION. THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. A, CONTAINS SEVERAL SECTIONS WITH INFORMATION ABOUT HOW TO PREVENT AND CONTROL INFECTION. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646564&pc=NGV | 2124215-2025-50096 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-29 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Pericardial Effusion |  | E2017028 | 3 | Event Description: THIS REPORT SUMMARIZES 3 DEATH EVENTS FOR PERICARDIAL EFFUSION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. OF THE REPORTED 3 SERIOUS INJURY PERICARDIAL EFFUSION EVENTS: 1 INVOLVED SURGICAL INTERVENTIONS, 3 INVOLVED CARDIAC TAMPONADE, 3 INVOLVED AN ADDITIONAL DEVICE (DRAINAGE), AGE RANGE: 60 TO 88 YEARS, PATIENT SEX: 2 FEMALES; 1 MALE. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 8 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646732&pc=NGV | 2124215-2025-50143 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-29 04:00 | NGV | WATCHMAN FLX? PRO | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | E2017028 | 1 | Event Description: THIS REPORT SUMMARIZES 1 DEATH EVENT FOR IMPLANT EMBOLIZATION. NOTE THAT MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER REPORTING GUIDELINES. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC RECEIVED NOTIFICATION OF EVENTS FOR THE WATCHMAN LAAC DEVICE REPORTED IN THE LEFT ATRIAL APPENDAGE OCCLUSION REGISTRY (LAAO) REGISTRY, WATCHMAN FLX PRO SURPASS CLINICAL STUDY TO ASSESS LONG-TERM SAFETY AND EFFECTIVENESS OUTCOMES ASSOCIATED WITH THE USE AND IMPLANTATION OF THE WATCHMAN FLX PRO LEFT ATRIAL APPENDAGE (LAA) CLOSURE DEVICE WITH DELIVERY SYSTEM IN A ROUTINE CLINICAL SETTING. PATIENT EVENTS WERE REPORTED AS EVENT TERMS WITH NO FURTHER DETAILED INFORMATION. MULTIPLE EVENT TERMS MAY APPLY TO A SINGLE PATIENT BUT ARE CAPTURED SEPARATELY UNDER THE TERMS OF THE SUMMARY REPORTING GUIDELINES. DATA INCLUDES NEWLY IMPLANTED PATIENTS AND FOLLOW UP OF PATIENTS INCLUDED IN THE PREVIOUS DATA SET. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, DATA IS ANONYMIZED BEFORE BEING TRANSMITTED TO BSC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE REGISTRY DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. Manufacturer Narrative: QUARTERLY REPORTING PERIOD: Q2 2025. AVERAGE TIME TO EVENT: 0 DAYS. DEVICE EVALUATION: UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). REGISTRY NAME: WATCHMAN FLX PRO DEVICE SURVEILLANCE POST APPROVAL ANALYSIS (WATCHMAN FLX PRO SURPASS).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22646995&pc=OZD | 1220648-2025-30833 | 2025-07-17 04:00 | Death | ABIOMED, INC. | 2025-07-29 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT FOLLOWING IMPELLA 5.5 IMPLANT, THE PURGE PRESSURE BEGAN TO CLIMB COINCIDING WITH A RISING LEFT VENTRICLE SIGNAL METRIC LINE. AN ECHOCARDOIGRAM WAS PERFORMED TO VERIFY THE POSITIONING OF THE DEVICE AND IT WAS DISCOVERED THAT THE ENTIRE LEFT VENTRICLE HAD BEEN CLOTTED OFF. THE DECISION WAS MADE TO END THE SUPPORTED PROCEDURE AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION: SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER: SECTION: DIRECT AORTIC INSERTION: SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." . IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22647085&pc=OZD | 1220648-2025-30834 | 2025-07-16 04:00 | Death | ABIOMED, INC. | 2025-07-29 04:00 | OZD | IMPELLA CP | Positioning Failure | Arrhythmia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND AFTER THE IMPLANT, THE PATIENT HAD MULTIPLE EPISODES OF VENTRICULAR TACHYCARDIA/VENTRICULAR FIBRILLATION WITH MULTIPLE SHOCKS. ADDITIONALLY, THE PUMP WAS SITTING A LITTLE DEEP AND THE DEVICE WAS REPOSITIONED HOWEVER, THE DOCTOR ELECTED TO LEAVE A LITTLE DEEP AS WAVEFORMS AND FLOWS ARE APPROPRIATE. THE PATIENT CONTINUED TO DETERIORATE REQUIRING MAXED OUT EPI, LEVO, AND VASO WITH BORDERLINE MAP. REMAINED VERY ACIDOTIC AND NOTE OXYGENATING WELL. MAP DROPPED INTO THE 40S TO 50S WITH SUCTION AT ALL REDUCED P LEVELS. THREADS PULSE NOTED SO INITIATED CARDIOPULMONARY RESUSCITATION. AFTER MULTIPLE ROUNDS THE FAMILY ELECTED TO STOP LIFE SAVING MEASURES AND THE PATIENT PASSED SHORTLY AFTER. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22647955&pc=NPT | 9617601-2025-00732 | 2025-07-21 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-07-29 04:00 | NPT | EVOLUT FX PLUS VALVE | Malposition of Device; Patient Device Interaction Problem | Foreign Body In Patient | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER IMPLANTATION OF THE TRANSCATHETER AORTIC VALVE EVFXPLUS-23 DURING A VALVE-IN-VALVE PROCEDURE, THE PATIENT EXPERIENCED OCCLUSION OF BOTH CORONARY ARTERIES ATTRIBUTED TO THE IMPLANTED VALVE. THE LEFT CORONARY ARTERY WAS FULLY BLOCKED DURING THE PROCEDURE, WHILE ONLY THE RIGHT CORONARY ARTERY HAD BEEN PROTECTED PRIOR TO STARTING THE TRANSCATHETER AORTIC VALVE IMPLANTATION. THE PATIENT PASSED AWAY FOLLOWING THE PROCEDURE.Event Description: IT WAS REPORTED THAT AFTER IMPLANTATION OF THE TRANSCATHETER AORTIC VALVE EVFXPLUS-23 DURING A VALVE-IN-VALVE PROCEDURE, THE PATIENT EXPERIENCED OCCLUSION OF BOTH CORONARY ARTERIES ATTRIBUTED TO THE IMPLANTED VALVE. THE LEFT CORONARY ARTERY WAS FULLY BLOCKED DURING THE PROCEDURE, WHILE ONLY THE RIGHT CORONARY ARTERY HAD BEEN PROTECTED PRIOR TO STARTING THE TRANSCATHETER AORTIC VALVE IMPLANTATION. THE PATIENT PASSED AWAY FOLLOWING THE PROCEDURE. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THIS WAS A VALVE-IN-VALVE CASE WITH A MEDTRONIC VALVE IMPLANTED WITHIN A 23MM NON-MEDTRONIC SURGICAL VALVE (SORIN CROWN). DURING PRE-PROCEDURAL COMPUTED TOMOGRAPHY IMAGING, THE VIRTUAL VALVE-TO-CORONARY (VTC) DISTANCE WAS CALCULATED AS 6MM ON THE LEFT WHILE THE VTC ON THE RIGHT WAS 3MM. THIS WAS WHY THE PHYSICIAN DECIDED TO INITIALLY PROTECT THE RIGHT CORONARY OSTIA ONLY AND NOT THE LEFT CORONARY OSTIA. AT THE START OF THE PROCEDURE, THE RIGHT CORONARY OSTIA WAS PROTECTED, BUT IT WAS THE LEFT CORONARY OSTIA THAT WAS COMPLETELY BLOCKED UPON VALVE DEPLOYMENT. THE MEDTRONIC VALVE WAS SNARED TO THE ASCENDING AORTA AND SOME PERFUSION RETURNED TO THE CORONARY ARTERIES. HOWEVER, WHEN A SECOND VALVE (EDWARDS) WAS FULLY IMPLANTED IT CAUSED A COMPLETE OCCLUSION OF THE CORONARY ARTERIES. THE ATTEMPTS WERE UNSUCCESSFUL AND THERE WAS NOT ENOUGH TIME TO SAVE THE PATIENT. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS OCCLUSION OF THE RIGHT AND LEFT CORONARY ARTERIES. THE PHYSICIAN ALSO NOTED THE MEDTRONIC DELIVERY CATHETER SYSTEM HAD NOTHING TO DO WITH THE PATIENT'S DEATH. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. H6. ADDITIONAL CODES. CORRECTED DATA: 2. OUTCOME ATTRIBUTED TO ADVERSE EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012869494); PRODUCT TYPE: 0195-HEART VALVES. PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012864343); PRODUCT TYPE: 0195-HEART VALVES. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22648080&pc=DSR | MW5173508 | 2023-07-01 04:00 | Death | CVRX, INC. | 2025-07-29 04:00 | DSR | BAROSTIM HEART IMPLANT DEVICE | Insufficient Information | Cardiac Arrest; Heart Failure/Congestive Heart Failure; Respiratory Arrest |  |  | 1 | Event Description: MY HUSBAND (B)(6), DOB (B)(6) 1961. HE WAS A PATIENT OF DR. (B)(6), OF (B)(6) FOR MANY WONDERFUL YEARS OF TREATMENT OVER 25 PLUS YEARS. HE WAS RELEASED FROM CARE FROM DR. (B)(6), AFTER HIS CONDITION WORSENED, I AM ASSUMING. HE THEN WAS SENT TO (B)(6) AT (B)(6). DR. (B)(6) HAD GREAT BEDSIDE MANNERS, SEEMED TO BE A GREAT DOCTOR AND WAS TRUSTED BY ME AND MY HUSBAND BECAUSE WE TRUSTED THE JUDGEMENT/ADVICE OF DR. (B)(6), MY HUSBAND DEVOTED CARE GIVER/DOCTOR AGAIN FOR OVER 20 PLUS YEARS. EVENTS LEADING UP TO THE DEATH OF MY HUSBAND ON (B)(6) OF 2023. MY HUSBAND (B)(6) WAS ADMITTED INTO THE HOSPITAL ON (B)(6) FOR THE IMPLANTATION OF THE BARAOSTIM DEVICE. WE WERE EXCITED ABOUT WHAT WE WERE TOLD THAT THIS DEVICE WOULD DO AS FAR AS GIVING HIM A BETTER QUALITY OF LIFE. WE WERE TOLD THOSE INDIVIDUALS THAT HAD BEEN CANDIDATES FOR HEART TRANSPLANT HAD BEEN TAKEN OFF THE TRANSPLANT LIST DO THEM RECEIVING THE BARAOSTIM. MY HUSBAND WAS SKEPTICAL BUT AFTER WE WERE TOLD OF THE POSSIBILITIES, WE THOUGHT IT WAS WELL WORTH IT TO HAVE THE PROCEDURE DON'T. MY HUSBAND, ON THE SURGERY DATE (B)(6) 2023, WAS EXTREMELY FULL OF FLUID. I THOUGHT THEY WOULD HAVE HELD OFF DUE TO HIM BEEN SO FULL OF FLUID. HE HAD A DEFIBRILLATOR AND A PACEMAKER AND WAS TAKEN OFF THE HEART TRANSPLANT DUE TO HIM HAVING TOO MANY OTHER CONDITIONS AND TOLD THAT A NEW HEART WOULD SUFFER THE SAME CONDITION AS THE OLD HEART. HE WAS NOT A CANDIDATE FOR THE BARAOSTIM ACCORDING TO YOUR FDA QUALIFICATION ATTACHED. I WANT SOMEONE TO TELL ME WHY MY HUSBAND THE LOVE OF MY LIFE WAS TREATED AS AN EXPERIMENT BECAUSE HE WAS CONSIDERED VERY ILL AND WAS ASSUMED TO BE ON THE LAST DAYS OF LIFE. THEY KNEW HOW SICK HE WAS AND THAT THE BARAOSTIM AND NOTHING ELSE COULD HAVE SAVED HIS LIFE. THE BARAOSTIM TOOK HIS LIFE. HE WAS ADMITTED AT (B)(6) HOSPITAL ON (B)(6) 2023 FULL OF FLUID FOR THE PROCEDURE. MY HUSBAND CAME OUT OF RECOVERY ON (B)(6), AFTER THE PROCEDURE AND STATED "BABY, I DON'T WANT TO GO HOME". HE SAID I KNOW THIS IS A ONE-DAY PROCEDURE, BUT I DON'T FEEL RIGHT AT ALL. THEY SENT HIM HOME ON (B)(6) STILL FULL OF FLUID. HE PASSED AWAY AT HOME AS SOON AS I GOT HIM IN THE BED RIGHT AFTER BEING DISMISSED FROM (B)(6). THIS WAS NOT REPORTED TO CVRX THE BARAOSTIM MAKER, WHICH WAS STATED TO ME BY MR. (B)(4), VICE PRESIDENT OF QUALITY ASSURANCE AND REGULATORY AFFAIRS. MR. (B)(4) ALSO STATED THAT HE WAS SURPRISED THAT IT WAS NOT BECAUSE THIS IS PROCEDURE TO LET THE MAKERS KNOW THE PROS AND CONS AND HOW WELL THE DEVICE IS DOING TO HELP HEART PTS. I SURELY WOULD THINK THAT A DEATH WOULD SURELY BEEN REPORTED BECAUSE THAT WAS THE ULTIMATE, BUT IT WAS NOT WHY BECAUSE THEY KNOW IT KILLED HIM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22648091&pc=NIQ | 2024168-2025-07850 | 2014-09-01 04:00 | Death | ABBOTT VASCULAR | 2025-07-29 04:00 | NIQ | XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P070015 |  | 1 | Event Description: THE STUDY CONDUCTED TO TO A ANALYZE THE IMPACT OF THE LEFT ANTERIOR DESCENDING CORONARY ARTERY (LAD) AS THE CULPRIT VESSEL OF THE ST-ELEVATION MYOCARDIAL INFARCTION (STEMI) ON VERY-TERM OUTCOME PATIENTS UNDERGOING PRIMARY PCI WITH XIENCE V STENTS AND MULTILINK VISION STENT ENROLLED IN THE EXAMINATION-EXTEND STUDY (10-YEAR FOLLOW-UP OF THE EXAMINATION TRIAL) . PATIENT'S WERE SEPARATED INTO GROUPS LAD-RELATED STEMI AND NON-LAD-RELATED STEMI. AT THE 0 TO 1 YEAR ANALYSIS BOTH GROUPS NOTED (DEATH, VENTRICULAR ARRHYTHMIAS, CARDIOGENIC SHOCK, STENT THROMBOSIS, TARGET LESION REVASCULARIZATION (TLR). AT THE 10 YEAR OUTCOME BOTH GROUPS NOTED (CARDIAC DEATH, MYOCARDIAL INFRACTION, REVASCULARIZATION, MYOCARDIAL INFRACTION, STENT THROMBOSIS. THERE WERE NO DIFFERENCES IN PATIENT-ORIENTED COMPOSITE ENDPOINT (POCE) BETWEEN LAD AS THE STEMI-RELATED CULPRIT VESSEL AND OTHER VESSELS AT THE 10 YEARS FOLLOW-UP. HOWEVER, IT WAS NOTED ALL CAUSE MORTALITY WAS MORE COMMON IN THE LAD-STEMI GROUP WITHIN THE FIRST YEAR. SPECIFIC PATIENT INFORMATION REMAINS UNKNOWN. ADDITIONAL INFORMATION CAN BE FOUND IN THE ATTACHED ARTICLE "LONG-TERM PROGNOSTIC IMPACT OF THE LEFT ANTERIOR DESCENDING CORONARY ARTERY AS THE STEMI-RELATED CULPRIT VESSEL: SUBANALYSIS OF THE EXAMINATION-EXTEND TRIAL". NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICES WERE NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBERS AND/OR LOT NUMBERS WERE NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT(S), AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE REPORTED PATIENT EFFECT OF DEATH, IS LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B3, B6, D6: DATES ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL DEVICE REFERENCED IN B5 IS FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. ATTACHMENT: ARTICLE TITLE: "LONG-TERM PROGNOSTIC IMPACT OF THE LEFT ANTERIOR DESCENDING CORONARY ARTERY AS THE STEMI-RELATED CULPRIT VESSEL: SUBANALYSIS OF THE EXAMINATION-EXTEND TRIAL."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22648594&pc=MAF | 2024168-2025-07853 | 2014-09-01 04:00 | Death | ABBOTT VASCULAR INC. | 2025-07-29 04:00 | MAF | MULTI-LINK VISION | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | P020047 |  | 1 | Event Description: THE STUDY CONDUCTED TO A ANALYZE THE IMPACT OF THE LEFT ANTERIOR DESCENDING CORONARY ARTERY (LAD) AS THE CULPRIT VESSEL OF THE ST-ELEVATION MYOCARDIAL INFARCTION (STEMI) ON VERY-TERM OUTCOME PATIENTS UNDERGOING PRIMARY PCI WITH XIENCE V STENTS AND MULTILINK VISION STENT ENROLLED IN THE EXAMINATION-EXTEND STUDY (10-YEAR FOLLOW-UP OF THE EXAMINATION TRIAL) . PATIENT'S WERE SEPARATED INTO GROUPS LAD-RELATED STEMI AND NON-LAD-RELATED STEMI. AT THE 0 TO 1 YEAR ANALYSIS BOTH GROUPS NOTED (DEATH, VENTRICULAR ARRHYTHMIAS, CARDIOGENIC SHOCK, STENT THROMBOSIS, TARGET LESION REVASCULARIZATION (TLR). AT THE 10 YEAR OUTCOME BOTH GROUPS NOTED (CARDIAC DEATH, MYOCARDIAL INFRACTION, REVASCULARIZATION, MYOCARDIAL INFRACTION, STENT THROMBOSIS. THERE WERE NO DIFFERENCES IN PATIENT-ORIENTED COMPOSITE ENDPOINT (POCE) BETWEEN LAD AS THE STEMI-RELATED CULPRIT VESSEL AND OTHER VESSELS AT THE 10 YEARS FOLLOW-UP. HOWEVER, IT WAS NOTED ALL CAUSE MORTALITY WAS MORE COMMON IN THE LAD-STEMI GROUP WITHIN THE FIRST YEAR. SPECIFIC PATIENT INFORMATION REMAINS UNKNOWN. ADDITIONAL INFORMATION CAN BE FOUND IN THE ATTACHED ARTICLE "LONG-TERM PROGNOSTIC IMPACT OF THE LEFT ANTERIOR DESCENDING CORONARY ARTERY AS THE STEMI-RELATED CULPRIT VESSEL: SUBANALYSIS OF THE EXAMINATION-EXTEND TRIAL". NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B3, B6, D6: DATES ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL DEVICE REFERENCED IN B5 IS FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. ATTACHMENT: ARTICLE TITLE: "LONG-TERM PROGNOSTIC IMPACT OF THE LEFT ANTERIOR DESCENDING CORONARY ARTERY AS THE STEMI-RELATED CULPRIT VESSEL: SUBANALYSIS OF THE EXAMINATION-EXTEND TRIAL". Manufacturer Narrative: THE DEVICES WERE NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBERS AND/OR LOT NUMBERS WERE NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT(S), AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE REPORTED PATIENT EFFECT OF DEATH, IS LISTED IN THE VISION CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22648859&pc=OZD | 1220648-2025-30837 | 2025-07-16 04:00 | Death | ABIOMED, INC. | 2025-07-29 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED AN IMPELLA CP (AFTER PLACEMENT OF ANTEGRADE PERFUSION SHEATH) FEMORALLY TO SUPPORT THE CRITICAL PATIENT. THE PATIENT HAD A RENAL HISTORY AND WAS ON EXTRACORPOREAL MEMBRANE OXYGENATION PRIOR TO IMPELLA CP PLACEMENT. THE IMPELLA CP WAS SUPPORTING WHEN ON THE SECOND DAY THERE WAS CONCERNS OF HEMOLYSIS AND LOW FLOW THAT WERE TREATED BY BLOOD PRODUCTS AND CONTINUOUS RENAL REPLACEMENT THERAPY. HOWEVER, EVENTUALLY CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22648930&pc=DSQ | 2916596-2025-05034 | 2025-07-18 04:00 | Death | THORATEC CORPORATION | 2025-07-29 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE VENTRICULAR ASSIST DEVICE (VAD) COORDINATOR DID NOT MENTION ANY DEVICE RELATED ISSUES LEADING TO THE DEATH OF THE PATIENT. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: THERE WAS NO MENTION OF ANY DEVICE RELATED ISSUES LEADING TO THE DEATH OF THE PATIENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTI ORGAN FAILURE AND DYSFUNCTION (INCLUDING RIGHT HEART FAILURE, RESPIRATORY FAILURE, HEPATIC DYSFUNCTION, AND RENAL DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22648966&pc=NPT | 2025587-2025-05327 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Pulmonary Dysfunction | P130021 | E2014038 | 37 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE PULMONARY. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED EIGHTY-TWO DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 57 TO (B)(6) AND OVER. THERE WERE 16 MALES AND 21 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 37 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22648972&pc=NPT | 2025587-2025-05328 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Pulmonary Dysfunction | P130021 | E2014038 | 10 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE PULMONARY. THE AVERAGE TIME TO EVENT WAS THIRTY-EIGHT DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 77 TO 88. THERE WERE 8 MALES AND 2 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 10 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22648983&pc=NPT | 2025587-2025-05329 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Insufficient Information | Pulmonary Dysfunction | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS PULMONARY. THE TIME TO EVENT WAS TWO HUNDRED SIXTY-NINE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 86. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22648985&pc=NPT | 2025587-2025-05330 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Insufficient Information | Pulmonary Dysfunction | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS PULMONARY. THE TIME TO EVENT WAS TWO HUNDRED FIFTY-THREE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 89. THE PATIENT WAS FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649009&pc=NPT | 2025587-2025-05332 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Cardiac Arrest | P130021 | E2014038 | 51 | Event Description: THE TYPES OF DEATHS BEING REPORTED ARE SUDDEN CARDIAC DEATHS AND DEATH DUE TO OTHER CARDIAC REASONS. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED FIFTY-THREE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 54 TO (B)(6) AND OVER. THERE WERE 29 MALES AND 22 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 51. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649014&pc=NPT | 2025587-2025-05333 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Cardiac Arrest | P130021 | E2014038 | 27 | Event Description: THE TYPES OF DEATHS BEING REPORTED ARE SUDDEN CARDIAC DEATHS AND DEATH DUE TO OTHER CARDIAC REASONS. THE AVERAGE TIME TO EVENT WAS EIGHTEEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 66 TO (B)(6) AND OVER. THERE WERE 11 MALES AND 16 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 27 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649034&pc=NPT | 2025587-2025-05334 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Insufficient Information | Cardiac Arrest | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED IS SUDDEN CARDIAC DEATH. THE TIME TO EVENT WAS THREE HUNDRED TWENTY-TWO DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 65. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649040&pc=NPT | 2025587-2025-05331 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Pulmonary Dysfunction | P130021 | E2014038 | 2 | Event Description: THE TYPE OF DEATHS BEING REPORTED ARE PULMONARY. THE AVERAGE TIME TO EVENT WAS SIXTY-SIX DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 63 TO 79. THERE WAS 1 MALE AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 2 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649054&pc=NPT | 2025587-2025-05335 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Cardiac Arrest | P130021 | E2014038 | 4 | Event Description: THE TYPES OF DEATHS BEING REPORTED ARE SUDDEN CARDIAC DEATHS AND DEATH DUE TO OTHER CARDIAC REASONS. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED NINETY-FIVE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 67 TO 87. THERE WERE 3 MALES AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 4 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649084&pc=NPT | 2025587-2025-05336 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 2 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE TRAUMA. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED FOURTEEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 78 TO 81. THERE WAS 1 MALE AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 2 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649103&pc=NPT | 2025587-2025-05337 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED WAS TRAUMA. THE TIME TO EVENT WAS ONE HUNDRED SIXTY-ONE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 78. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649107&pc=NPT | 2025587-2025-05338 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED WAS TRAUMA. THE TIME TO EVENT WAS TWO HUNDRED SIXTY-SEVEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 81. THERE WAS 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649124&pc=NPT | 2025587-2025-05340 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Myocardial Infarction | P130021 | E2014038 | 2 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE MYOCARDIAL INFARCTIONS. THE AVERAGE TIME TO EVENT WAS EIGHTEEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 75 TO 77. THERE WERE 2 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 2 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649151&pc=NPT | 2025587-2025-05339 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Myocardial Infarction | P130021 | E2014038 | 11 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE MYOCARDIAL INFARCTIONS. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED EIGHTY-NINE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 55 TO (B)(6) AND OVER. THERE WERE 4 MALES AND 7 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 11 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649162&pc=NPT | 2025587-2025-05341 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Myocardial Infarction | P130021 | E2014038 | 2 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE MYOCARDIAL INFARCTIONS. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED FORTY-NINE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 72 TO (B)(6) AND OVER. THERE WAS 1 MALE AND 1 FEMALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 2 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649209&pc=DXF | 2124215-2025-51150 | 2025-07-02 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-07-29 04:00 | DXF | VERSACROSS CONNECT LAAC ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Bradycardia; Cardiac Arrest; Low Blood Pressure/ Hypotension; Cardiac Tamponade; Pericardial Effusion | K150709 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED.
DURING A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A VERSACROSS CONNECT LAAC WAS SELECTED FOR USE. THE TRANSSEPTAL PUNCTURE WAS PERFORMED WITHOUT FLUOROSCOPY GUIDANCE. THE PUNCTURE WAS NOT DONE AT ITS MOST IDEAL POSITION. A 35MM WM FLX PRO DEVICE WAS OPENED AFTER CROSS VALIDATING APPENDAGE MEASUREMENTS FROM COMPUTED TOMOGRAPHY SCAN AND TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). THE PHYSICIANS STILL TRIED TO SEE IF USING TRUSTEER SHEATH WOULD COMPENSATE FOR THE INADEQUATE TRANSSEPTAL POSITION. IT WAS TRIED MULTIPLE TIMES TO POSITION THE WM FLX PRO DEVICE CORRECTLY IN THE APPENDAGE, BUT THE POSITION OF THE DEVICE WAS NOT ACCEPTABLE. HENCE, IT WAS DECIDED TO FULLY RE-CAPTURE THE WM FLX PRO 35MM DEVICE WITHIN THE TRUSTEER SHEATH AND REMOVE BOTH. IN REMOVING THE WATCHMAN SHEATH AND DEVICE, A PERICARDIAL EFFUSION FLAGGED. THE PHYSICIANS RIGHT AWAY CALLED FOR STERILE TRAYS AND ALL THE EQUIPMENT NEEDED FOR DRAINAGE OF PERICARDIAL EFFUSION. AT THIS POINT, ABOUT TWO TO THREE 60CC SYRINGES OF BLOOD HAD BEEN EXTRACTED. THE PHYSICIAN THEN DECIDED TO GO FOR A SECOND TRANSSEPTAL. THE SECOND TRANSEPTAL PUNCTURE SUCCESSFULLY. THE TRUSTEER SHEATH AND THE 35MM WM FLX PRO DEVICE WERE RE-INTRODUCED AND DEPLOYED IN THE PATIENT. AT THIS TIME, THE PHYSICIANS WERE PLEASED THAT THE IMPLANT OF THE WM FLX PRO HELPED 'BLOCK' THE LEAK. THE PATIENT EFFUSION WAS STARTING TO POOL AGAIN. THE PHYSICIAN DECIDED TO RELEASE THE 35MM CLOSURE DEVICE WHICH WAS WELL POSITIONED IN THE APPENDAGE. ADDITIONAL DRAINAGE WAS THEN PERFORMED, BUT THE PATIENT WENT INTO CARDIAC ARREST. THE NURSES PERFORMED HEART MASSAGE. THE PATIENT EXPERIENCED SEVERE BRADYCARDIA, AND HYPOTENSION (ARTERIAL TENSION LOW 16 HIGH 34). A DEFIBRILLATOR (200 JOULESX1) WAS USED AND LIFE CELL WAS USED TO GIVE ADDITIONAL BLOOD UNITS. A CODE WAS CALLED IN THE HOSPITAL TO ADVISE THE SURGERY TEAM. PROTAMINE WAS GIVEN TO THE PATIENT. THE PATIENT SEEMED TO STABILIZE SLIGHTLY, BUT THE EFFUSION PERSISTED. THE PHYSICIAN KEPT REMOVING 60CC SYRINGES OF BLOOD FROM THE PATIENT. THE PATIENT REMAINED IN THE INTERVENTIONAL CARDIOLOGY LAB UNDER ACTIVE RESUSCITATION AND WAS REPORTED TO HAVE DIED. THE PHYSICIAN SUSPECTED THAT THE INADEQUATE TRANSEPTAL LOCATION CAUSED THE WATCHMAN TRUSTEER ACCESS SYSTEM TO RUB AGAINST THE RIDGE AND POSSIBLY CAUSED A SMALL TEAR IN THE APPENDAGE. THERE WAS NO PERFORATION NOTED ON ANY IMAGING. THE CAUSE OF DEATH WAS DETERMINED TO BE TAMPONADE. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649332&pc=NPT | 2025587-2025-05342 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 5 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE CARDIOVASCULAR PROCEDURE RELATED. THE AVERAGE TIME TO EVENT WAS TWO DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 77 TO (B)(6) AND OVER. THERE WAS 1 MALE AND 4 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 5 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649396&pc=NPT | 2025587-2025-05344 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 1 | Event Description: THE TYPE OF DEATH BEING REPORTED WAS CARDIOVASCULAR PROCEDURE RELATED. THE TIME TO EVENT WAS ZERO DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT'S AGE WAS 71. THE PATIENT WAS MALE. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 1 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649397&pc=NPT | 2025587-2025-05348 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Heart Failure/Congestive Heart Failure | P130021 | E2014038 | 6 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE HEART FAILURE. THE AVERAGE TIME TO EVENT WAS EIGHTEEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 73 TO (B)(6) AND OVER. THERE WERE 3 MALES AND 3 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 6 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649412&pc=NPT | 2025587-2025-05343 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Insufficient Information | P130021 | E2014038 | 6 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE CARDIOVASCULAR PROCEDURE RELATED. THE AVERAGE TIME TO EVENT WAS TWO DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 63 TO (B)(6) AND OVER. THERE WERE 6 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 6 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649417&pc=NPT | 2025587-2025-05346 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Hemorrhage/Bleeding | P130021 | E2014038 | 6 | Event Description: THE TYPES OF DEATHS BEING REPORTED WERE HEMORRHAGIC AND CARDIOVASCULAR HEMORRHAGES. THE AVERAGE TIME TO EVENT WAS EIGHT DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 70 TO 86. THERE WERE 6 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 6 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649419&pc=NPT | 2025587-2025-05350 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information | Stroke/CVA | P130021 | E2014038 | 9 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE STROKE. THE AVERAGE TIME TO EVENT WAS EIGHTEEN DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 70 TO (B)(6) AND OVER. THERE WERE 4 MALES AND 5 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 9 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649427&pc=NPT | 2025587-2025-05345 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Hemorrhage/Bleeding | P130021 | E2014038 | 6 | Event Description: THE TYPES OF DEATHS BEING REPORTED WERE HEMORRHAGIC AND CARDIOVASCULAR HEMORRHAGES. THE AVERAGE TIME TO EVENT WAS NINETY-NINE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 76 TO 89. THERE WERE 3 MALES AND 3 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038. QUARTERLY REPORTING PERIOD: Q2 2025. TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 6. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649429&pc=NPT | 2025587-2025-05351 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Stroke/CVA | P130021 | E2014038 | 2 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE STROKE. THE AVERAGE TIME TO EVENT WAS SIXTY-SIX DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 75 TO 84. THERE WERE 2 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 2 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649438&pc=NPT | 2025587-2025-05347 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Heart Failure/Congestive Heart Failure | P130021 | E2014038 | 30 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE HEART FAILURE. THE AVERAGE TIME TO EVENT WAS ONE HUNDRED FIFTY-NINE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 62 TO (B)(6) AND OVER. THERE WERE 16 MALES AND 14 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 30 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649442&pc=NPT | 2025587-2025-05349 | 2025-05-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-07-29 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Stroke/CVA | P130021 | E2014038 | 18 | Event Description: THE TYPE OF DEATHS BEING REPORTED WERE STROKE. THE AVERAGE TIME TO EVENT WAS TWO HUNDRED AND ONE DAYS FOLLOWING THE IMPLANT PROCEDURE. THE PATIENT AGES IN THIS REPORT RANGE FROM 59 TO (B)(6) AND OVER. THERE WERE 10 MALES AND 8 FEMALES. THE CAUSE OF DEATH PROVIDED IS LIMITED IN NATURE AS IT IS PROVIDED VIA A THIRD-PARTY DATABASE. MEDTRONIC RECEIVED INFORMATION REGARDING PATIENT/DEVICE EVENTS VIA A THIRD-PARTY POST-IMPLANT DEVICE REGISTRY (THE SOCIETY OF THORACIC SURGEONS/AMERICAN COLLEGE OF CARDIOLOGY TRANSCATHETER VALVE THERAPY REGISTRY). THE INFORMATION IN THIS REPORT WAS PROVIDED TO MEDTRONIC IN A DE-IDENTIFIED FORMAT AND HAS BEEN ORGANIZED INTO SUMMARIES OF OBSERVATIONS RELATED TO PATIENT DEATHS. Manufacturer Narrative: TVT REGISTRY EXEMPTION NUMBER: E2014038 QUARTERLY REPORTING PERIOD: Q2 2025 TOTAL NUMBER OF EVENTS BEING SUMMARIZED: 18 UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED PATIENT DEMOGRAPHICS DETAILS AND LIMITED DETAILS WERE PROVIDED REGARDING THE ADVERSE EVENTS AND OUTCOMES. THE LISTED EVENT DATE IS THE DATE THE INFORMATION WAS RECEIVED BY MEDTRONIC. THE PATIENT INFORMATION INCLUDED IN SECTION A. IS AN AVERAGE OF THE DATA PROVIDED FOR THE EVENTS. A PRODUCT ANALYSIS WAS NOT ABLE TO BE PERFORMED AS NO PRODUCT WAS RETURNED. WITHOUT THE RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATION. THE VIEWS OR OPINIONS PRESENTED HERE DO NOT REPRESENT THOSE OF THE AMERICAN COLLEGE OF CARDIOLOGY FOUNDATION, THE SOCIETY OF THORACIC SURGEONS, OR THE STS/ACC TVT REGISTRY. THE INDUSTRY DATA FILE UPON WHICH THE ANALYSIS WAS PERFORMED IS A SUBSET OF THE FULL STS/ACC TVT REGISTRY DATA SUBMISSIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649869&pc=PNJ | 2017865-2025-96381 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-29 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22649873&pc=PNJ | 2017865-2025-96380 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-29 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT ATRIUM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22650514&pc=PNJ | 2017865-2025-96390 | 2025-07-09 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-29 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Activation, Positioning or Separation Problem | Arrhythmia; Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR A SCHEDULED PROCEDURE. DURING THE PROCEDURE, THE RIGHT VENTRICULAR LEADLESS PACEMAKER (VLP) WAS FIXATED AND THE THRESHOLD REMAINED ELEVATED OVER TIME AND DID NOT DECREASE. REDOCKING WAS REPEATED THREE TIMES AND DURING THE FOURTH PRE-MAPPING, WHILE REMOVING THE SLEEVE, THE VLP ADVANCED UNINTENTIONALLY. A CARDIAC TAMPONADE HAD OCCURRED AND A PERICARDIOCENTESIS WAS PERFORMED. BLOOD PRESSURE INCREASED AND THE PERFORATION SITE WAS TREATED BY THORACOTOMY. IT WAS THEN LATER CONFIRMED THAT THE PATIENT DECEASED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22650547&pc=PNJ | 2017865-2025-96389 | 2025-07-09 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-07-29 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Capturing Problem | Arrhythmia; Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR A SCHEDULED PROCEDURE. DURING THE PROCEDURE, THE RIGHT VENTRICULAR LEADLESS PACEMAKER (VLP) WAS FIXATED AND THE THRESHOLD REMAINED ELEVATED OVER TIME AND DID NOT DECREASE. REDOCKING WAS REPEATED THREE TIMES AND DURING THE FOURTH PRE-MAPPING, WHILE REMOVING THE SLEEVE, THE VLP ADVANCED UNINTENTIONALLY. A CARDIAC TAMPONADE HAD OCCURRED AND A PERICARDIOCENTESIS WAS PERFORMED. BLOOD PRESSURE INCREASED AND THE PERFORATION SITE WAS TREATED BY THORACOTOMY. IT WAS THEN LATER CONFIRMED THAT THE PATIENT DECEASED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22650592&pc=LWQ | 2135147-2025-04170 | 2025-07-04 04:00 | Death | ABBOTT MEDICAL | 2025-07-29 04:00 | LWQ | MASTERS SERIES HEMODYNAMIC PLUS VALVE | Detachment of Device or Device Component | Shock | P810002 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM MASTERS SERIES HEMODYNAMIC PLUS VALVE WAS SELECTED FOR PROCEDURE. THE PATIENT HAD A PREOPERATIVE LEFT VENTRICULAR EJECTION FRACTION (LVEF) OF 27%. IT WAS NOTED DURING PROCEDURE THAT A LEAFLET OF THE VALVE BECOME DISLODGED WHILE USING THE VALVE ROTATOR. THE VALVE HAD ALREADY BEEN PLACED AND FIXATED WITH STITCHES TO THE AORTIC RING. THERE WAS PROPER ALIGNMENT OF THE VALVE AND EXCESSIVE FORCE WAS NOT USED. THE VALVE WAS IN A CLOSED POSITION WHEN IT WAS ROTATED, AND THE HOLDER HANDLE WAS NOT FULLY INSERTED INTO THE ORIFICE AT A 90-DEGREE ANGLE. THE DISLODGED LEAFLET CAME OFF AS ONE PIECE AND WAS SUCCESSFULLY RECOVERED FROM THE PATIENT. THE PHYSICIAN DID NOT ATTEMPT TO REINSERT THE LEAFLET. INSTRUMENTS THAT ENCOUNTERED THE VALVE AT THE TIME OF DISLODGMENT INCLUDED THE VALVE ROTATOR AND A SILICONE PENCIL. NO RESISTANCE WAS REPORTED DURING ROTATION. NO OTHER TOOLS BESIDES THE HOLDER HANDLE WERE USED TO POSITION OR ROTATE THE VALVE. THE DECISION WAS MADE TO REMOVE THE 25MM MASTERS SERIES HEMODYNAMIC PLUS VALVE. A NEW 23MM MASTERS SERIES HEMODYNAMIC PLUS VALVE HAD TO BE PREPARED FOR IMPLANT AND RESULTED IN PROLONGED THE AORTIC CLAMPING AND EXTRACORPOREAL CIRCULATION TIMES. THIS DIRECTLY AFFECTED THE PATIENT'S CONDITION. THE PATIENT ULTIMATELY DIED ON (B)(6) 2025, WITH THE CAUSE OF DEATH REPORTED AS CARDIOGENIC SHOCK. Manufacturer Narrative: THE DEVICE IS NOT RETURNING FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22653564&pc=NIQ | 2124215-2025-46133 | 2024-11-13 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-07-29 04:00 | NIQ | SYNERGY XD | Positioning Failure; Device Dislodged or Dislocated | Foreign Body In Patient; Vascular Dissection |  |  | 1 | Event Description: IT WAS REPORTED VIA USER FACILITY TGA REPORT (B)(4), STENT DISLODGMENT OCCURRED REQUIRING ADDITIONAL INTERVENTION, AND THE PATIENT DIED. THE PATIENT UNDERWENT A STAGED PERCUTANEOUS CORONARY INTERVENTION (PCI) VIA THE LEFT RADIAL ARTERY. DURING THE PROCEDURE, A 2.50 X 12MM SYNERGY XD DRUG-ELUTING STENT FAILED TO EXPAND DESPITE ATTEMPTS USING TWO DIFFERENT INFLATION DEVICES. UPON REMOVAL, THE STENT DELIVERY SYSTEM WAS FOUND TO BE BROKEN. FURTHER ASSESSMENT CONFIRMED THAT THE STENT HAD DETACHED AND REMAINED WITHIN THE VESSEL. RETRIEVAL ATTEMPTS USING A SNARE WERE UNSUCCESSFUL, LEADING TO THE DEPLOYMENT OF AN ADDITIONAL STENT TO SECURE THE DETACHED STENT AGAINST THE VESSEL WALL. ADDITIONALLY, A DISSECTION OF THE OBTUSE MARGINAL BRANCH OF THE LEFT CIRCUMFLEX CORONARY ARTERY WAS NOTED, HOWEVER, NO FURTHER DETAILS REGARDING THIS COMPLICATION WERE PROVIDED. IT WAS ALSO REPORTED THAT THE PATIENT PASSED AWAY, THOUGH NO ADDITIONAL DETAILS WERE PROVIDED REGARDING THE CIRCUMSTANCES OF DEATH. Manufacturer Narrative: B2: DATE OF DEATH WAS NOT PROVIDED. GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22655465&pc=LDF | 2182208-2025-03746 | 2025-06-11 04:00 | Death | MEDTRONIC, INC. | 2025-07-29 04:00 | LDF | ELECTRODE, PACEMAKER, TEMPORARY | Pacing Problem; Use of Device Problem | Arrhythmia; Cardiac Arrest; Ventricular Fibrillation | K963898 |  | 1 | Event Description: IT WAS REPORTED THAT WHILE THE EXTERNAL PULSE GENERATOR (EPG) WAS IN USE, THE PATIENT EXPERIENCED AN R WAVE ON T WAVE EVENT, WHICH PRECIPITATED A VENTRICULAR FIBRILLATION (VF) ARREST. THE PATIENT WAS RESUSCITATED BUT ASPIRATED DURING RE-INTUBATION. IT WAS NOTED THAT HE SUBSEQUENTLY DEVELOPED ADDITIONAL COMPLICATIONS AND, UNFORTUNATELY, PASSED AWAY ON POSTOPERATIVE DAY SIX. IT WAS ALSO NOTED THAT THE PATIENT HAD BEEN ALERT, EXTUBATED, AND AMBULATORY ON POSTOPERATIVE DAY ONE FOLLOWING OPEN-HEART SURGERY, WHICH INCLUDED AN AORTIC VALVE REPLACEMENT AND ANEURYSM REPAIR. NO FURTHER INFORMATION WAS AVAILABLE THE EPG IS EXPECTED TO BE RETURNED FOR EVALUATION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22660598&pc=DSY | 1721504-2025-00216 | 2022-03-01 05:00 | Death | MERIT MEDICAL SYSTEMS INC. | 2025-07-29 04:00 | DSY | HERO® | Obstruction of Flow | Liver Damage/Dysfunction; Thrombosis/Thrombus; Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem | K213845 |  | 1 | Event Description: A LITERATURE ARTICLE DESCRIBES A 40-YEAR-OLD FEMALE WITH END-STAGE RENAL DISEASE AND A COMPLEX VASCULAR HISTORY UNDERWENT SUCCESSFUL IMPLANTATION OF A SUPER HERO GRAFT IN 2021. DESPITE CORRECT IMPLANTATION AND MULTIPLE SUCCESSFUL THROMBECTOMIES OVER THE FOLLOWING MONTHS, THE PATIENT EXPERIENCED REPEATED THROMBOTIC OCCLUSIONS OF THE GRAFT. AFTER RECEIVING A KIDNEY TRANSPLANT AND DISCONTINUING WARFARIN, DEVICE REMOVAL WAS PLANNED BUT NOT IMMEDIATELY EXECUTED. IN MARCH 2022, THE PATIENT PRESENTED WITH ACUTE ABDOMINAL SYMPTOMS AND UNDERWENT AN URGENT CHOLECYSTECTOMY. POSTOPERATIVELY, SHE DEVELOPED UNEXPLAINED SHOCK, AND IMAGING REVEALED EXTENSIVE THROMBOSIS OF THE INFERIOR VENA CAVA AND SUPRAHEPATIC VEINS CONSISTENT WITH BUDD-CHIARI SYNDROME. DESPITE URGENT TRANSFEMORAL THROMBECTOMY, PARTIAL THROMBUS REMAINED, AND THE PATIENT DEVELOPED MULTI-ORGAN FAILURE, INCLUDING DIALYSIS-DEPENDENT RENAL FAILURE, HEPATIC DYSFUNCTION, AND RESPIRATORY COMPROMISE. SHE DIED 10 DAYS AFTER ICU ADMISSION. Manufacturer Narrative: THE SUSPECT DEVICE WAS NOT RETURNED FOR EVALUATION. THE COMPLAINT COULD NOT BE CONFIRMED. THE ROOT CAUSE COULD NOT BE DETERMINED. A REVIEW OF THE DEVICE HISTORY AND COMPLAINT DATABASE COULD NOT BE PERFORMED SINCE THE LOT NUMBER WAS NOT PROVIDED.