## May'25
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22033332&pc=DSP | 2249723-2025-0002232 | 2025-04-16 04:00 | Death | DATASCOPE CORP. | 2025-05-16 04:00 | DSP | CS300 INTRA-AORTIC BALLOON PUMP, SPANISH, 110V | Unexpected Shutdown | Insufficient Information | K063525 |  | 1 | Event Description: IT WAS REPORTED THAT CS300 INTRA-AORTIC BALLOON PUMP (IABP) SHUT DOWN DURING THE PATIENT'S TRANSFER FROM THE CORONARY ICU TO THE CT SCAN. EVENT OCCURRED ON (B)(6) 2025, 4:30 PM. THE PATIENT DIED, ALTHOUGH IT IS UNKNOWN IF IT WAS DUE TO THE EQUIPMENT. Manufacturer Narrative: DUE TO CHARACTER LIMITATION IN E1, FULL EVENT SITE NAME IS (B)(6). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22034539&pc=NVZ | 3008973940-2025-03478 |  | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-05-16 04:00 | NVZ | ADAPTA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. FOLLOW-UP WAS ATTEMPTED TO OBTAIN THE RELEVANT INFORMATION; HOWEVER, A CAUSE OF DEATH COULD NOT BE IDENTIFIED. THIS EVENT IS BEING CONSERVATIVELY REPORTED IN AN ABUNDANCE OF CAUTION IN RESPONSE TO THE ADVISORY DESCRIBING THE POTENTIAL FOR DEVICE CIRCUIT ERROR TO OCCUR WHILE THE DEVICE IS PROCESSING AN ATRIAL-SENSED EPISODE. THIS PATIENT WAS CONFIRMED TO BE PROGRAMMED TO A PACING MODE SUSCEPTIBLE TO THE DESCRIBED ADVISORY ISSUE. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. ANALYSIS OF THE DEVICE REVEALED NORMAL BATTERY DEPLETION. THE BATTERY INDICATOR SIGNIFYING THAT IT IS TIME FOR DEVICE REPLACEMENT OCCURRED AFTER EXPLANT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22034660&pc=MIH | 9612164-2025-02438 | 2025-05-07 04:00 | Death | MEDTRONIC IRELAND | 2025-05-16 04:00 | MIH | ENDURANT II ILIAC STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hematoma; Hemorrhage/Bleeding; Necrosis; Sepsis; Rupture; Obstruction/Occlusion | P100021 |  | 1 | Event Description: AN ENDURANT IIS STENT GRAFT SYSTEM WAS IMPLANTED AS PART OF THE ADVANCE CLINICAL STUDY FOR THE TREATMENT OF ABDOMINAL AORTIC ANEURYSM. IT WAS REPORTED THAT THE LIMBS ETLW1620C93EE (B)(6) AND ETEW2020C82EE (B)(6) WERE IMPLANTED IN THE LEFT ILIAC AND THE LIMB ETLW1620C93EE (B)(6) WAS IMPLANTED IN THE RIGHT ILIAC. IT WAS REPORTED THAT AN INTRAOPERATIVE VESSEL-RUPTURE OF THE PELVIS-AXIS IN THE LEFT ILIAC ARTERY OCCURRED DURING THE SURGERY FOLLOWING IMPLANTATION OF THE EVAR-PROSTHESIS. THE PATIENT DEVELOPED A RETROPERITONEAL HEMATOMA, WHICH TRIGGERED A SECONDARY INTERVENTION. AN OPEN SURGICAL CONVERSION WITH REMOVAL OF THE EVAR DEVICE WAS PERFORMED ON THE FOLLOWING DAY. A BLOOD TRANSFUSION WAS ADMINISTERED, AND THE PATIENT UNDERWENT CARE IN THE INTENSIVE CARE UNIT (ICU) REQUIRING PROLONGED HOSPITALIZATION. THE PATIENT EXPIRED 3 DAYS POST INDEX PROCEDURE. IT WAS NOTED THAT TACHYARRHYTHMIA AND SIGMOID NECROSIS (COLON) OCCURRED, WHICH LED TO SEPSIS. THE CAUSE OF DEATH WAS ATTRIBUTED TO SEPSIS. THE SITE ASSESSED THE ILIAC RUPTURE AS HAVING A CAUSAL RELATIONSHIP TO THE STUDY DEVICE, INDEX PROCEDURE AND NOT RELATED TO AN UNDERLYING CONDITION OR DISEASE. THE RETROPERITONEAL HEMATOMA AND THE SIGMOID NECROSIS WERE ASSESSED BY THE SITE AS NOT RELATED TO THE DEVICE AND HAVING A CAUSAL RELATIONSHIP TO THE PROCEDURE (SECONDARY INTERVENTION) AND AN UNDERLYING CONDITION OR DISEASE. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: ETEW2020C82EE, SERIAL/LOT #: (B)(6), UBD: 05-DEC-2026, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5; ADDITIONAL INFORMATION RECEIVED: THE RETROPERITONEAL HEMATOMA AND SIGMOID NECROSIS WERE ASSESSED AS HAVING A CAUSAL RELATIONSHIP TO THE INDEX PROCEDURE. MEDICATION ( METOPROLOL AND AMIODARON ) WERE GIVEN TO TREAT THE TACHYARRHYTHMIA. THE PATIENT ALSO EXPERIENCED CLINICALLY SIGNIFICANT INCREASED SERUM CREATININE LEVELS ON THE SAME DATE AS DEATH. THE SPONSOR ASSESSED THIS AS HAVING A CAUSAL RELATIONSHIP TO THE INDEX PROCEDURE , NOT RELATED TO THE DEVICE. THE TACHYARRHYTHMIA WAS ASSESSED BY THE SPONSOR AS HAVING A CAUSAL RELATIONSHIP TO THE INDEX PROCEDURE, POSSIBLY RELATED TO THE STUDY DEVICE , NOT RELATED TO ANY UNDERLYING CONDITION. THE INCREASED SERUM CREATININE WAS ASSESSED BY THE SITE AS HAVING A CAUSAL RELATIONSHIP TO THE PROCEDURE, NOT RELATED TO THE STUDY DEVICE OR AN UNDERLYING DISEASE. THE TACHYARRHYTHMIA WAS ASSESSED BY THE SITE AS NOT RELATED TO THE STUDY DEVICE/STUDY PROCEDURE OR AN UNDERLYING DISEASE. THE SITE REPORTED THAT THE LEFT DISTAL ILIAC LIMB WAS INVOLVED IN THE ILIAC RUPTURE AND THE RUPTURE OCCURRED POSSIBLY DUE TO ILIAC CALCIFICATION IN THE CONTEXT OF BALLOONING OF DISTAL SEALING ZONE. IT WAS CONFIRMED THAT ALL STENT GRAFTS WERE EXPLANTED, EXCEPT FOR THE SUPRARENAL STENTS, WHICH REMAINED IN SITU. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5: ADDITIONAL INFORMATION RECEIVED: IT WAS REPORTED THAT DURING THE INDEX PROCEDURE, THERE WAS A INTRAOPERATIVE THROMBOSIS OF BOTH LIMBS DUE TO DISTAL ATHEROSCLEROTIC DISEASE WHICH RESULTED IN A GRAFT THROMBECTOMY AND A PATCHING OF THE BILATERAL COMMON FEMORAL ARTERIES BEING PERFORMED. THIS RESOLVED THE EVENT. THE THROMBOSIS OF BOTH LIMBS WAS ASSESSED BY THE SITE AS HAVING A CAUSAL RELATIONSHIP TO THE PROCEDURE BUT NOT RELATED TO THE STUDY DEVICE OR AN UNDERLYING CONDITION OR DISEASE. *CORRECTION (MISSED INFORMATION FROM PREVIOUS REPORT PREVIOUS REPORT): THE RETROPERITONEAL HEMATOMA AND SIGMOID NECROSIS WERE ASSESSED AS NOT RELATED TO AN UNDERLYING CONDITION OR DISEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED; THE SITE HAS UPDATED THE RELATEDNESS ASSESSMENT OF THE TACHYARRHYTHMIA AS HAVING A CAUSAL RELATIONSHIP WITH AN UNDERLYING CONDITION OR DISEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE SITE RELATED ASSESSMENT OF THE TACHYARRHYTHMIA HAS BEEN UPDATED TO NOT RELATED TO AN UNDERLYING CONDITION OR DISEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATE TO H6; CODING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THE SIGMOID NECROSIS EVENT WAS TREATED WITH A SIGMOID COLOSCOPY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22034674&pc=MIH | 9612164-2025-02439 | 2025-05-07 04:00 | Death | MEDTRONIC IRELAND | 2025-05-16 04:00 | MIH | ENDURANT IIS BIFURCATED STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hematoma; Hemorrhage/Bleeding; Necrosis; Sepsis; Rupture; Obstruction/Occlusion | P100021 |  | 1 | Event Description: AN ENDURANT IIS STENT GRAFT SYSTEM WAS IMPLANTED AS PART OF THE ADVANCE CLINICAL STUDY FOR THE TREATMENT OF ABDOMINAL AORTIC ANEURYSM. IT WAS REPORTED THAT THE LIMBS ETLW1620C93EE (B)(6) AND ETEW2020C82EE (B)(6) WERE IMPLANTED IN THE LEFT ILIAC AND THE LIMB ETLW1620C93EE (B)(6) WAS IMPLANTED IN THE RIGHT ILIAC. IT WAS REPORTED THAT AN INTRAOPERATIVE VESSEL-RUPTURE OF THE PELVIS-AXIS IN THE LEFT ILIAC ARTERY OCCURRED DURING THE SURGERY FOLLOWING IMPLANTATION OF THE EVAR-PROSTHESIS. THE PATIENT DEVELOPED A RETROPERITONEAL HEMATOMA, WHICH TRIGGERED A SECONDARY INTERVENTION. AN OPEN SURGICAL CONVERSION WITH REMOVAL OF THE EVAR DEVICE WAS PERFORMED ON THE FOLLOWING DAY. A BLOOD TRANSFUSION WAS ADMINISTERED, AND THE PATIENT UNDERWENT CARE IN THE INTENSIVE CARE UNIT (ICU) REQUIRING PROLONGED HOSPITALIZATION. THE PATIENT EXPIRED 3 DAYS POST INDEX PROCEDURE. IT WAS NOTED THAT TACHYARRHYTHMIA AND SIGMOID NECROSIS (COLON) OCCURRED, WHICH LED TO SEPSIS. THE CAUSE OF DEATH WAS ATTRIBUTED TO SEPSIS. THE SITE ASSESSED THE ILIAC RUPTURE AS HAVING A CAUSAL RELATIONSHIP TO THE STUDY DEVICE, INDEX PROCEDURE AND NOT RELATED TO AN UNDERLYING CONDITION OR DISEASE. THE RETROPERITONEAL HEMATOMA AND THE SIGMOID NECROSIS WERE ASSESSED BY THE SITE AS NOT RELATED TO THE DEVICE AND HAVING A CAUSAL RELATIONSHIP TO THE PROCEDURE (SECONDARY INTERVENTION) AND AN UNDERLYING CONDITION OR DISEASE. Manufacturer Narrative: B5; ADDITIONAL INFORMATION RECEIVED: THE RETROPERITONEAL HEMATOMA AND SIGMOID NECROSIS WERE ASSESSED AS HAVING A CAUSAL RELATIONSHIP TO THE INDEX PROCEDURE. MEDICATION (METOPROLOL AND AMIODARON) WERE GIVEN TO TREAT THE TACHYARRHYTHMIA. THE PATIENT ALSO EXPERIENCED CLINICALLY SIGNIFICANT INCREASED SERUM CREATININE LEVELS ON THE SAME DATE AS DEATH. THE SPONSOR ASSESSED THIS AS HAVING A CAUSAL RELATIONSHIP TO THE INDEX PROCEDURE, NOT RELATED TO THE DEVICE. THE TACHYARRHYTHMIA WAS ASSESSED BY THE SPONSOR AS HAVING A CAUSAL RELATIONSHIP TO THE INDEX PROCEDURE, POSSIBLY RELATED TO THE STUDY DEVICE, NOT RELATED TO ANY UNDERLYING CONDITION. THE INCREASED SERUM CREATININE WAS ASSESSED BY THE SITE AS HAVING A CAUSAL RELATIONSHIP TO THE PROCEDURE, NOT RELATED TO THE STUDY DEVICE OR AN UNDERLYING DISEASE. THE TACHYARRHYTHMIA WAS ASSESSED BY THE SITE AS NOT RELATED TO THE STUDY DEVICE/STUDY PROCEDURE OR AN UNDERLYING DISEASE. THE SITE REPORTED THAT THE LEFT DISTAL ILIAC LIMB WAS INVOLVED IN THE ILIAC RUPTURE AND THE RUPTURE OCCURRED POSSIBLY DUE TO ILIAC CALCIFICATION IN THE CONTEXT OF BALLOONING OF DISTAL SEALING ZONE. IT WAS CONFIRMED THAT ALL STENT GRAFTS WERE EXPLANTED, EXCEPT FOR THE SUPRARENAL STENTS, WHICH REMAINED IN SITU. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5: ADDITIONAL INFORMATION RECEIVED: IT WAS REPORTED THAT DURING THE INDEX PROCEDURE, THERE WAS A INTRAOPERATIVE THROMBOSIS OF BOTH LIMBS DUE TO DISTAL ATHEROSCLEROTIC DISEASE WHICH RESULTED IN A GRAFT THROMBECTOMY AND A PATCHING OF THE BILATERAL COMMON FEMORAL ARTERIES BEING PERFORMED. THIS RESOLVED THE EVENT. THE THROMBOSIS OF BOTH LIMBS WAS ASSESSED BY THE SITE AS HAVING A CAUSAL RELATIONSHIP TO THE PROCEDURE BUT NOT RELATED TO THE STUDY DEVICE OR AN UNDERLYING CONDITION OR DISEASE. *CORRECTION (MISSED INFORMATION FROM PREVIOUS REPORT): THE RETROPERITONEAL HEMATOMA AND SIGMOID NECROSIS WERE ASSESSED AS NOT RELATED TO AN UNDERLYING CONDITION OR DISEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED; THE SITE HAS UPDATED THE RELATEDNESS ASSESSMENT OF THE TACHYARRHYTHMIA AS HAVING A CAUSAL RELATIONSHIP WITH AN UNDERLYING CONDITION OR DISEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE SITE RELATED ASSESSMENT OF THE TACHYARRHYTHMIA HAS BEEN UPDATED TO NOT RELATED TO AN UNDERLYING CONDITION OR DISEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATE TO H6; CODING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THE SIGMOID NECROSIS EVENT WAS TREATED WITH A SIGMOID COLOSCOPY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: ETLW1620C93EE, SERIAL/LOT#: (B)(6), UBD: 16-JUL-2025, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22034682&pc=DTB | 3008973940-2025-03479 |  | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-05-16 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. FOLLOW-UP WAS ATTEMPTED TO OBTAIN THE RELEVANT INFORMATION; HOWEVER, A CAUSE OF DEATH COULD NOT BE IDENTIFIED. THIS EVENT IS BEING CONSERVATIVELY REPORTED IN AN ABUNDANCE OF CAUTION IN RESPONSE TO THE ADVISORY DESCRIBING THE POTENTIAL FOR DEVICE CIRCUIT ERROR TO OCCUR WHILE THE DEVICE IS PROCESSING AN ATRIAL-SENSED EPISODE. THIS PATIENT WAS CONFIRMED TO BE PROGRAMMED TO A PACING MODE SUSCEPTIBLE TO THE DESCRIBED ADVISORY ISSUE. Manufacturer Narrative: CORRECTED: H3 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22034706&pc=DTB | 2649622-2025-13401 |  | Death | MPRI | 2025-05-16 04:00 | DTB | CAPSURE SENSE LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P830061 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. FOLLOW-UP WAS ATTEMPTED TO OBTAIN THE RELEVANT INFORMATION; HOWEVER, A CAUSE OF DEATH COULD NOT BE IDENTIFIED. THIS EVENT IS BEING CONSERVATIVELY REPORTED IN AN ABUNDANCE OF CAUTION IN RESPONSE TO THE ADVISORY DESCRIBING THE POTENTIAL FOR DEVICE CIRCUIT ERROR TO OCCUR WHILE THE DEVICE IS PROCESSING AN ATRIAL-SENSED EPISODE. THIS PATIENT WAS CONFIRMED TO BE PROGRAMMED TO A PACING MODE SUSCEPTIBLE TO THE DESCRIBED ADVISORY ISSUE. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22034710&pc=OZD | 1220648-2025-28470 | 2025-04-18 04:00 | Death | ABIOMED, INC. | 2025-05-16 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT HAD AN IMPELLA CP PLACED ALONG WITH EXTRACORPOREAL MEMBRANE OXYGENATION FOR CARDIOMYOPATHY AND VENTRICULAR TACHYCARDIA (VT) STORM. THE PATIENT WAS SEEN TO BE THROMBOSED AT THE ATRIOVENTRICULAR VALVE AND AORTA. HOWEVER, THE PATIENT EXPIRED. THE CAUSE OF DEATH IS NOT YET DETERMINED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF THROMBUS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. G2 REPORT SOURCE SELECTION FOREIGN WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28470. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION, SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22034923&pc=BZD | 2518422-2025-105748 | 2025-03-25 04:00 | Death | RESPIRONICS, INC. | 2025-05-16 04:00 | BZD | REMSTAR AUTOA-FLEX | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K113068 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REMSTAR AUTOA-FLEX DEVICE HAS PASSED AWAY AND THE DAUGHTER WANTED TO RETURN THE DEVICE. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22034970&pc=OZD | 1220648-2025-28473 | 2025-04-25 04:00 | Death | ABIOMED, INC. | 2025-05-16 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT SUFFERED A MASSIVE STROKE DURING IMPELLA 5.5 SUPPORT. THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: STROKE/CVA: "STROKES/CVAS CAN OCCUR DURING OR AFTER IMPELLA SUPPORT, AND CAN BE EITHER ISCHEMIC OR HEMORRHAGIC. TO MAKE A DETERMINATION AS TO THE CAUSE OF THE STROKE, THE FOLLOWING CLINICAL INFORMATION MUST BE CONSIDERED: -PATIENT'S MEDICAL HISTORY, INCLUDING ANY PREVIOUS STROKES, CARDIOVASCULAR DISEASE, OR RISK FACTORS SUCH AS HYPERTENSION, DIABETES, OR ATRIAL FIBRILLATION. -TIMING OF THE STROKE IN RELATION TO IMPELLA SUPPORT (DURING OR POST-IMPLANT). -DETAILED DESCRIPTION OF THE SYMPTOMS EXPERIENCED BY THE PATIENT. -NEUROLOGICAL EXAMINATION FINDINGS AND ANY FOCAL DEFICITS OBSERVED. -DIAGNOSTIC IMAGING RESULTS, SUCH AS CT SCAN OR MRI OF THE BRAIN, TO IDENTIFY THE TYPE AND LOCATION OF THE STROKE (ISCHEMIC OR HEMORRHAGIC). -LABORATORY TEST RESULTS, INCLUDING COAGULATION PROFILES AND CARDIAC MARKERS. -ANY RELEVANT PROCEDURAL OR DEVICE-RELATED FACTORS, SUCH AS DURATION AND SETTINGS OF IMPELLA SUPPORT, PRESENCE OF EMBOLIC PROTECTION DEVICES, OR ANY DOCUMENTED PROCEDURAL COMPLICATIONS. -RESPONSE TO ANY PREVIOUS ANTICOAGULATION OR ANTIPLATELET THERAPIES. -EVALUATION OF POTENTIAL ALTERNATIVE CAUSES OF STROKE, SUCH AS ARRHYTHMIAS, EMBOLIC SOURCES, OR UNDERLYING VASCULAR PATHOLOGY. AS THIS CLINICAL INFORMATION WAS NOT PROVIDED, THE TRUE ROOT CAUSE OF THE STROKE/CVA COULD NOT BE DETERMINED.¿ Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22035008&pc=LWQ | 2025587-2025-03720 | 1993-02-18 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-16 04:00 | LWQ | HALL EASY FIT | Insufficient Information | Insufficient Information | P790018 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT 32 YEARS AND 2 MONTHS POST IMPLANT OF THIS HALL MECHANICAL VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE HALL MECHANICAL VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22035095&pc=KPR | MW5170457 | 2025-05-10 04:00 | Death | UNK | 2025-05-16 04:00 | KPR | X-RAY | Defective Device | Insufficient Information |  |  | 1 | Event Description: (B)(6) ANIMAL SERVICES, IS SAID TO HAVE XRAY MACHINES, HOWEVER THEY HAVE STATED THAT THEIR MACHINE HAS BEEN UNFUNCTIONAL FOR WEEKS, DESPITE THAT, MDAS HAS MISDIAGNOSED SEVERAL DOGS LEADING TO DEATH. THEY DO NOT HAVE A FUNCTIONAL XRAY MACHINE BUT HAVE EUTHANIZED SEVERAL DOGS STATING THEY SAW RADIOGRAPHS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22035220&pc=OZD | 1220648-2025-28475 | 2025-04-25 04:00 | Death | ABIOMED, INC. | 2025-05-16 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A 51-YEAR-OLD MALE PATIENT WITH KNOWN AORTIC ANEURYSM WAS TRANSFERRED FOR ADVANCED CARDIAC THERAPIES AND TAKEN TO THE OR FOR IMPELLA 5.5 PLACEMENT, ECMO DECANNULATION, INTRA-AORTIC BALLOON PUMP (IABP) REMOVAL, AND TRACHEOSTOMY. IT WAS NOTED THAT THE IMPELLA 5.5 PLACEMENT WAS SUCCESSFULLY PLACED. VA ECMO WAS DECANNULATED SURGICALLY REPAIRED AND A WOUND VAC INSERTED.THE LEFT GROIN IABPWAS TURNED OFF, REMOVED, AND SITE PRECLOSED. THE PATIENT WAS THEN PREPPED FOR TRACHEOSTOMY. UPON TRACH EXCHANGE THERE WAS AN ABSENCE OF VENTILATION FOR 10 SECONDS. PATIENT WENT INTO VENTRICULAR TACHYCARDIA (VT), HYPOTENSIVE AND CONSEQUENTLY DEFIBRILLATED. VASOPRESSIN 2 UNITS WERE ADMINISTERED WITH COMPRESSIONS AND PATIENT RECOVERED. IMPELLA TURNED BACK UP TO P5. DURING SUPPORT IT WAS REPORTED THAT THE PATIENT WAS IN/OUT OF SUPRAVENTRICULAR TACHYCARDIA (SVT). IN ADDITION, THE PATIENT WAS RECANNULATED AFTER HAVING ARRHYTHMIAS AND DROP IN MIXED VENOUS. THE PATIENT HAD PARADOXICAL BREATHING AND WAS NOT FOLLOWING COMMANDS. A HEAD CT REVEALED DIFFUSE ENCEPHALOPATHY. THE PATIENT'S NEURO STATUS WAS NOT CHANGING AND EVENTUALLY CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22035259&pc=NGV | 22035259 | 2025-04-16 04:00 | Death | ABBOTT | 2025-05-16 04:00 | NGV | ABBOTT | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade |  |  | 1 | Event Description: ON (B)(6) 2025 A 74-YEAR-OLD PATIENT HAD A LEFT ATRIAL APPENDAGE CLOSURE WITH AN AMULET CLOSURE DEVICE. THE PATIENT WAS STABLE AFTER THE PROCEDURE AND DISCHARGED THE SAME DAY. THE PATIENT EXPIRED LATER THAT NIGHT. THE PRESUMED CAUSE OF DEATH WAS EROSION OF THE LEFT ATRIAL APPENDAGE ON THE PULMONARY ARTERY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22035626&pc=BTR | 3012307300-2025-05793 | 2025-03-21 04:00 | Death | SMITHS HEALTHCARE MANUFACTURING S.A. DE C.V. | 2025-05-16 04:00 | BTR | PORTEX ENDOTRACHEAL TUBE, SILICONIZED | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Hydrocephalus | K790312 |  | 1 | Event Description: IT WAS REPORTED THAT AN PRETERM INFANT REQUIRED INTUBATION FOR ONGOING RESPIRATORY MANAGEMENT. DURING A PLANNED UPSIZING OF THE ENDOTRACHEAL (ET) TUBE, THE INFANT HAD A 3RD EPISODE OF A TYPICAL INTUBATION FINDING AND REQUIRED USE OF A RECALLED SIZE 2.0 ET TUBE. THE TUBE WAS USED IN THE KNOWLEDGE THAT IT HAD BEEN RECALLED BUT WAS USED AS A LIFESAVING INTERVENTION. THE INFANT WAS SUBSEQUENTLY TRANSFERRED TO ANOTHER NEONATAL UNIT FOR EAR, NOSE, AND THROAT (ENT) REVIEW, AND THE ET TUBE WAS STRAIGHTFORWARD, AND A SUCTION CATHETER COULD BE PASSED THROUGH THE ET TUBE TO PROVIDE ET SUCTION. THE INFANT HAD A SIZE 2.0 ENT TUBE INSERTED IN THE EARLY HOURS ON (B)(6) 2025, (B)(6) 2025 AS AN EMERGENCY PROCEDURE PRIOR TO THE ABOVE INTUBATION. IT WAS REPORTED TO BE UNKNOWN WHAT BRAND OF ENT TUBE THAT WAS, AND IT WAS NOT POSSIBLE TO A SUCTION CATHETER VIA THAT ET TUBE. THE INFANT HAD BILATERAL GRADE 4 INTRAVENTRICULAR HEMORRHAGES (IVHS) WITH EARLY HYDROCEPHALUS EVIDENT ON SCAN. CARE WAS RE-ORIENTED ON (B)(6) 2025, AND THE INFANT DIED. Manufacturer Narrative: H3, H6: NO PRODUCT OR PICTURES WERE RETURNED. THE REPORTED COMPLAINT COULD NOT BE CONFIRMED, AND A PROBABLE CAUSE COULD NOT BE DETERMINED. IF THE PRODUCT IS RETURNED THIS COMPLAINT WILL BE REOPENED FOR FURTHER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22036010&pc=OZD | 1220648-2025-28480 | 2025-04-26 04:00 | Death | ABIOMED, INC. | 2025-05-16 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WAS ON IMPELLA CP SUPPORT AND HAD MULTIPLE ARRHYTHMIAS AND RAPIDLY DECOMPENSATED REQUIRING AN INCREASE OF PRESSOR SUPPORT. THE PATIENT NEEDED ESCALATION OF CARE. HOWEVER, THE FAMILY CHOSE TO TRANSITION TO COMFORT MEASURES AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22036294&pc=NPT | 2135147-2025-02565 | 2025-04-22 04:00 | Death | ABBOTT MEDICAL | 2025-05-16 04:00 | NPT | NAVITOR VISION | Migration or Expulsion of Device; Improper or Incorrect Procedure or Method; Incomplete Coaptation; Material Deformation; Central Regurgitation | Cardiac Arrest; Aortic Valve Insufficiency/ Regurgitation | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 29MM NAVITOR VISION VALVE WAS SELECTED FOR IMPLANTATION IN A PATIENT WITH AORTIC STENOSIS. THE VALVE WAS SIZED BASED ON COMPUTED TOMOGRAPHY MEASUREMENTS. THE NATIVE VALVE WAS PRE-DILATED USING A 23MM NON-ABBOTT BALLOON. DURING THE INITIAL ATTEMPT TO DEPLOY THE VALVE, IT WAS OBSERVED TO BE SHALLOW ON THE LEFT CORONARY CUSP (LCC). A SUCCESSFUL PARTIAL RECAPTURE AND REPOSITIONING WERE PERFORMED, AND THE VALVE WAS DEPLOYED AT AN APPROXIMATE DEPTH OF 4MM TO 4.5MM BELOW THE ANNULUS. THERE WERE NO DEPLOYMENT OR RETRACTION ISSUES NOTED. THE VALVE WAS NEVER RE-SHEATHED MORE THAN TWICE DURING THE PROCEDURE. THE AORTIC ANGLE WAS 59 DEGREES. WHILE IN THE CUSP OVERLAP VIEW (COV), THE INFLOW EDGE OF THE CONSTRAINED VALVE WAS WITHIN THE CAPSULE POSITIONED AT THE BASE OF AORTIC ANNULUS WHILE THE PIGTAIL POSITIONED WAS CONFIRMED TO BE AT THE BOTTOM OF THE NON-CORONARY CUSP. POST-IMPLANT, THERE WAS NO ENTANGLEMENT WITH THE LARGE FLEXNAV DELIVERY SYSTEM AND VALVE NOTED. HOWEVER, IT WAS NOTED VIA ECHOCARDIOGRAM THAT THERE WAS SMALL SPOT WHERE THE FRAME OF THE VALVE APPEARED TO BE INDENTED. A TRANSTHORACIC ECHOCARDIOGRAM (TTE) REVEALED CENTRAL AORTIC REGURGITATION. THE PATIENT WAS INTUBATED, AND A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) PROBE WAS PLACED TO CONFIRM THE AORTIC REGURGITATION (AR). THE TEE INDICATED THAT THE LEAFLET ON THE NON-CORONARY CUSP WAS IMMOBILE. IT'S UNCERTAIN WHAT WAS PINNING THE LEAFLET. AN ATTEMPT WAS MADE TO MOBILIZE THE LEAFLET USING THE 23MM NON-ABBOTT BALLOON, BUT THIS WAS UNSUCCESSFUL. THE PHYSICIAN THEN ATTEMPTED A VALVE-IN-VALVE (VIV) PROCEDURE WITH A 29MM NON-ABBOTT VALVE. HOWEVER, AFTER TWO POSITIONING ATTEMPTS, THE 29MM NON-ABBOTT DEVICE PULLED ON THE NAVITOR VALVE, CAUSING IT TO EMBOLIZE INTO THE SINUSES. THE MIGRATED VALVE DID NOT APPEAR TO BLOCK ANY CORONARY ARTERIES. EFFORTS TO SNARE THE NAVITOR VALVE WITH A LASSO AND BRING IT ABOVE THE SINOTUBULAR JUNCTION (STJ) WERE UNSUCCESSFUL. SEVERE AR WAS NOTED, AND THE PATIENT CODED, NECESSITATING EMERGENT CARDIOTHORACIC SURGERY. THE 29MM NAVITOR VISION VALVE WAS NOT DAMAGED BY THE SNARED OR THE NON-ABBOTT DEVICE USED IN THE VIV. UNFORTUNATELY, THE PATIENT PASSED AWAY LATER THAT EVENING IN THE INTENSIVE CARE UNIT DUE TO HEMODYNAMIC COMPROMISE DUE TO SEVERE AORTIC REGURGITATION. Manufacturer Narrative: AN EVENT OF MIGRATION, SEVERE AORTIC REGURGITATION AND PATIENTS DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE NOT RETURNED FOR ANALYSIS. IT WAS INDICATED THAT THE VALVE WAS PLACED AT A DEPTH OF 4 MM TO 4.5 MM BELOW THE ANNULUS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. THE CAUSE OF THE REPORTED MIGRATION APPEARS TO BE UNDER PROCEDURAL CONDITIONS(NON-ABBOTT DEVICE PULLED ON THE NAVITOR VALVE). THE CAUSE OF AORTIC REGURGITATION APPEARS TO BE RELATED TO PROCEDURAL CONDITIONS, POSSIBLY DUE TO MIGRATION OR THE SNARING OF THE MIGRATED VALVE. THE CAUSE OF DEATH COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED UNEXPECTED MEDICAL INTERVENTIONS AND SURGERY WERE RESULTS OF CASE-SPECIFIC CIRCUMSTANCES, AS THE VALVE WAS SNARED, POST-IMPLANT VALVULOPLASTY, ANOTHER DEVICE WAS IMPLANTED VALVE-IN-VALVE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. PLEASE NOTE THE PER THE INSTRUCTION FOR USE (IFU), (B)(6) REVISION D, ¿POST-IMPLANTATION PRECAUTIONS: ONCE THE VALVE IS FULLY DEPLOYED, REPOSITIONING AND RETRIEVAL OF THE VALVE IS NOT POSSIBLE. ATTEMPTED RETRIEVAL (E.G., USE OF A GUIDEWIRE, SNARE, OR FORCEPS) MAY CAUSE AORTIC ROOT, CORONARY ARTERY, AND/OR MYOCARDIAL DAMAGE.¿ THERE IS NO INDICATION THAT THE DEVIATION CAUSED ANY PATIENT HARM, A LETTER WILL NOT BE SENT. FROM MEDICAL REVIEW: "ATTEMPT TO IMPLANT A MEDTRONIC VALVE-IN-VALVE PROCEDURE WAS NOT SUCCESSFUL AND CAUSED EMBOLIZATION OF THE NAVITOR VALVE OUT OF THE ANNULUS. THE PATIENT PASSED AWAY LATER THE SAME DAY BUT NO ADDITIONAL INFORMATION WAS PROVIDED. NO FURTHER ANALYSIS CAN BE PROVIDED WITH THE LIMITED INFORMATION PROVIDED. IF ADDITIONAL INFORMATION AND IMAGING IS PROVIDED, AN ADDENDUM WILL BE PROVIDED TO THIS MEDICAL REVIEW." Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22036361&pc=DSQ | 2916596-2025-03031 | 2024-11-05 05:00 | Death | THORATEC CORPORATION | 2025-05-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. NO FURTHER DETAILS WERE PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22036536&pc=NVN | 2017865-2025-53937 | 2025-04-28 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-16 04:00 | NVN | TENDRIL STS | Signal Artifact/Noise | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD UNEXPECTEDLY PASSED AWAY. THE PATIENT¿S RIGHT VENTRICULAR (RV) HAD NOISE, AND NOISE REVERSION PACING WAS WORKING APPROPRIATELY. THE PHYSICIAN HAD PLANNED ON REPLACING THE RV LEAD DURING A GENERATOR CHANGE. THE CAUSE OF DEATH IS NOT KNOWN. THE PATIENT WAS NOT ON REMOTE MONITORING, AND THE DEVICE HAD NOT BEEN CHECKED POST-MORTEM. NO FURTHER INFORMATION WAS REPORTED. Manufacturer Narrative: D4 - PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER CANNOT BE PROVIDED AS A PRODUCT IDENTIFIER COULD NOT BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22036838&pc=DXF | 2124215-2025-31557 | 2023-06-22 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-05-16 04:00 | DXF | VERSACROSS ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | K150709 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A STROKE. THE PATIENT PASSED AWAY. PER LITERATURE REVIEW, IT WAS REPORTED THAT: PER LITERATURE REVIEW, THIS IS A SINGLE-CENTER RETROSPECTIVE STUDY OF CONSECUTIVE PATIENTS WHO UNDERWENT MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR (MTEER) WITH THE PASCAL (NON-BOSTON SCIENTIFIC DEVICE). TRANSSEPTAL PUNCTURE WAS UNDERTAKEN WITH EITHER A MECHANICAL NEEDLE OR THE VERSACROSS WIRE. THE PRIMARY ENDPOINTS WERE SUCCESS OF TRANSSEPTAL PUNCTURE (TSP) AND SUCCESSFUL DELIVERY OF THE EDWARDS SHEATH ON THE CHOSEN DELIVERY WIRE. SECONDARY ENDPOINTS INCLUDED NUMBER OF WIRES USED, TAMPONADE RATE, INTERVAL FROM FEMORAL VENOUS ACCESS TO TSP AND FIRST PASCAL DEVICE DEPLOYMENT, PROCEDURAL DEATH, AND STROKE. BETWEEN MARCH 2021 AND OCTOBER 2022, 33 CONSECUTIVE M-TEER CASES WERE IDENTIFIED. THIRTY-THREE CONSECUTIVE PATIENTS (10 WITH MECHANICAL NEEDLE, 23 WITH VERSACROSS) WHO UNDERWENT MTEER WITH THE EDWARDS PASCAL DEVICE WERE IDENTIFIED. ALL PATIENTS HAD SUCCESSFUL TSP. IN THE VERSACROSS ARM, THE RADIOFREQUENCY WIRE WAS USED SUCCESSFULLY FOR DELIVERY OF THE SHEATH IN ALL CASES. THERE WERE NO CASES OF PERICARDIAL EFFUSION/TAMPONADE IN EITHER ARM. INTERVAL FROM FEMORAL VENOUS ACCESS TO TSP AND TO DEPLOYMENT OF THE FIRST PASCAL DEVICE WAS SHORTER WITH THE VERSACROSS SYSTEM. SIGNIFICANTLY FEWER WIRES WERE USED WITH VERSACROSS. IN THE MECHANICAL NEEDLE GROUP, SIGNIFICANTLY MORE GUIDEWIRES WERE REQUIRED IN THE PROCEDURE. THERE WERE NO PROCEDURAL STROKES OR DEATHS. ONE LATE STROKE OCCURRED IN THE VERSACROSS GROUP AT 36 HOURS POST MTEER. THE SAME PATIENT SUBSEQUENTLY DIED BEFORE DISCHARGE. THE EMBOLIC STROKE WAS NOT DEEMED TO BE RELATED TO THE VERSACROSS DEVICE AND WAS ATTRIBUTED TO ATRIAL FIBRILLATION, A DIFFICULT PROCEDURE WITH MULTIPLE REPOSITIONINGS, AND THE PASCAL DEVICE ITSELF. THERE WERE NO PROCEDURAL DEATHS OR STROKES IN EITHER GROUP. IN CONCLUSION, VERSACROSS APPEARS A SAFE AND EFFECTIVE METHOD OF TSP AND FOR DELIVERY OF THE 22FR SHEATH FOR M-TEER WITH PASCAL. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS. ARTICLE CITATION: DOSHI, S. N., SAVVOULIDIS, P., MECHERY, A., LAWTON, E., AND NADIR, M. A. (2023). VERSACROSS TRANSSEPTAL SYSTEM FOR MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR WITH THE PASCAL REPAIR PLATFORM. STRUCTURAL HEART, 7(6), 100203. HTTPS://DOI.ORG/10.1016/J.SHJ.2023.100203. Manufacturer Narrative: THE FIELD B3: DATE OF EVENT WAS UPDATED WITH ESTIMATED DATE AS 22JUN2023, SINCE AOI HAS NO SPECIFIC PROCEDURE DATE, DATE OF WHEN AOI WAS PUBLISHED WAS USED AS EVENT DATE. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. THIS WAS NOT AVAILABLE BECAUSE THIS WAS RELATED TO A LITERATURE STUDY. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. ARTICLE CITATION: DOSHI, S. N., SAVVOULIDIS, P., MECHERY, A., LAWTON, E., AND NADIR, M. A. (2023). VERSACROSS TRANSSEPTAL SYSTEM FOR MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR WITH THE PASCAL REPAIR PLATFORM. STRUCTURAL HEART, 7(6), 100203. HTTPS://DOI.ORG/10.1016/J.SHJ.2023.100203.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22037024&pc=DSQ | 2916596-2025-02985 | 2025-04-08 04:00 | Death | THORATEC CORPORATION | 2025-05-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22037070&pc=LOX | 9612164-2025-02447 | 2025-04-06 04:00 | Death | MEDTRONIC MEXICO | 2025-05-16 04:00 | LOX | EUPHORA RX | Deflation Problem; Entrapment of Device; Difficult to Remove; Detachment of Device or Device Component | Device Embedded In Tissue or Plaque; Insufficient Information | K143480 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING A EUPHORA BALLOON RX, THE DEVICE BECAME STUCK IN THE INFLATED POSITION WITHIN THE LEFT MAIN TO CIRCUMFLEX ARTERY. THE COMPLICATION AROSE DURING FINAL POST-DILATATION OF STENTS IN THE LEFT MAIN CORONARY ARTERY EXTENDING INTO THE LEFT CIRCUMFLEX ARTERY, WHEN A 3.5X15 MM NON-MEDTRONIC NC BALLOON FRACTURED. THE DISTAL (MONORAIL) OPAQUE TIP SEPARATED FROM THE HYPOTUBE AND REMAINED WITHIN THE CORONARY ARTERY, STUCK IN THE PREVIOUSLY IMPLANTED STENT IN THE LEFT MAIN TO THE CIRCUMFLEX ARTERY. MULTIPLE UNSUCCESSFUL ATTEMPTS WERE MADE TO RETRIEVE THE FRACTURED NON-MEDTRONIC BALLOON. DURING EFFORTS TO RETRIEVE THE NON-MEDTRONIC DEVICE, A EUPHORA BALLOON RX WAS INTRODUCED AND SUBSEQUENTLY BECAME STUCK IN THE SAME SEGMENT, ALSO IN THE INFLATED POSITION. AN ATTEMPT WAS MADE TO DEFLATE THE EUPHORA BALLOON BY CUTTING THE SHAFT, BUT THIS WAS UNSUCCESSFUL. SEVERAL ADDITIONAL RETRIEVAL EFFORTS FAILED. THE PATIENT WAS ULTIMATELY REFERRED FOR SURGICAL INTERVENTION FOR POTENTIAL DEVICE RETRIEVAL AND CORONARY ARTERY BYPASS GRAFTING (CABG). Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LESION BEING TREATED WAS VERY CALCIFIED, AT A 90 DEGREE ANGLE, BIFURCATED AND HAD 90% STENOSIS. IT WAS NOT DIFFICULT TO REMOVE THE PROTECTIVE SHEATH OR THE PACKAGING STYLETTE. THE EUPHORA DEVICE WAS INSPECTED BEFORE USE WITH NO ISSUES NOTED. RESISTANCE WAS NOT NOTED WHILE ADVANCING THE DEVICE TO THE LESION. A CONCENTRATION OF 30/50 CONTRAST/SALINE WAS USED. THE PREVIOUSLY IMPLANTED STENTS WERE NON-MEDTRONIC DEVICES. THERE WERE NO DIFFICULTIES NOTED DURING INFLATION OF THE DEVICE. FOUR INFLATIONS WERE PERFORMED PRIOR TO THE DEFLATION DIFFICULTIES. INFLATION PRESSURES OF 15, 19, 18, AND 12 ATM WERE APPLIED, RESPECTIVELY, FOR EACH INFLATION. THE DEVICE WAS MOVED OR REPOSITIONED WHILE INFLATED. THE BALLOON FAILED TO DEFLATE. THE DETACHED PORTION OF THE EUPHORA DEVICE WAS NOT SUCCESSFULLY REMOVED FROM THE PATIENT. THE CABG WAS SUCCESSFULLY PERFORMED BUT WITH THE BALLOON STILL IN THE PATIENT ANATOMY. THE PATIENT LATER PASSED AWAY. DEVICE LOT NUMBER PROVIDED. PATIENT WEIGHT PROVIDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: ANNEX D CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22037140&pc=DSQ | 3007042319-2025-01156 | 2025-05-05 04:00 | Death | HEARTWARE, INC. | 2025-05-16 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Stopped; Use of Device Problem; Pumping Problem | Stroke/CVA; Dyspnea; Renal Failure; Numbness; Thromboembolism; Convulsion/Seizure; Syncope/Fainting | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED WITH A STROKE, PRESENTING WITH SYMPTOMS OF SYNCOPE, RIGHT ARM NUMBNESS AND TINGLING, AND SHORTNESS OF BREATH. IMAGING REVEALED A PARIETAL LOBE INFARCT, BILATERAL PULMONARY EMBOLI, A SPLENIC INFARCT, AND RIGHT RADIAL AND ULNAR ARTERY THROMBOSIS CONFIRMED BY COMPUTERIZED TOMOGRAPHY (CT). THE PATIENT WAS ON ANTICOAGULATION THERAPY, BUT IT WAS SUBTHERAPEUTIC WITH AN INTERNATIONAL NORMALIZED RATIO (INR) OF 1.7. IT WAS NOTED THAT EARLIER THAT WEEK THE PATIENT HAD HAD THEIR COUMADIN HELD. A HEPARIN DRIP WAS ADMINISTERED AS PART OF THE MEDICAL INTERVENTION. IT WAS REPORTED THAT THE PATIENT REMAINED NEUROLOGICALLY INTACT. REVIEW OF LOG FILES DATA REVEALED THAT THE VENTRICULAR ASSIST DEVICE (VAD) HAD EXHIBITED MULTIPLE MOTOR START EVENTS WITH STARTING PARAMETERS OUTSIDE OF THE TYPICAL RANGE. IT WAS REPORTED THAT THE CAUSE OF THE MOTOR STARTS WAS SUSPECTED TO BE DUE TO ACCIDENTAL DOUBLE POWER SOURCE DISCONNECTIONS BY THE PATIENT. THE VAD AND CONTROLLER REMAIN IN USE.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS DISCHARGED TO REHAB STILL NEUROLOGICALLY INTACT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS READMITTED WITH ALTERED MENTAL STATUS AND SEIZURES REQUIRING SEDATION AND INTUBATION FOR AIRWAY PROTECTION. UPON ADMISSION, THE PATIENT WAS SEVERELY HYPOGLYCEMIC AND HAD LACTIC ACIDOSIS AND A SUPRATHERAPEUTIC INTERNATIONAL NORMALIZED RATIO (INR). THE PATIENT'S ANTICOAGULATION WAS HELD. THE INITIAL COMPUTERIZED TOMOGRAPHY (CT) WAS NEGATIVE FOR ACUTE HEMORRHAGE. REVIEW OF THE LOG FILES SHOWED A TREND DOWN IN FLOW AND INCREASED PULSATILITY WITH CONSISTENT SUCTION. THE PATIENT DEVELOPED TEA COLORED URINE AND HAD AN ELEVATED LACTATE DEHYDROGENASE (LDH). THE INR RETURNED TO NORMAL AND A HEPARIN DRIP WAS STARTED. IT WAS ALSO, REPORTED THAT THE WARFARIN AND HEPARIN WAS STOPPED DUE TO RETROPERITONEAL BLEED. THE PATIENT HAD ACUTE KIDNEY INJURY AND RETROPERITONEAL BLEED THIS HOSPITALIZATION. THE PATIENT WAS EXTUBATED AND PLACED ON COMFORT MEASURES PER HER FAMILY REQUEST. CONSEQUENTLY, THE PATIENT EXPIRED.Event Description: IT WAS NOTED THAT THE VAD START-UPS ASSOCIATED WITH THE SUSPECTED DOUBLE DISCONNECTS OF THE CONTROLLER HAD STARTING PARAMETERS WITHIN THE TYPICAL EXPECTED RANGE, NOT OUTSIDE THE RANGE. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR A CORRECTION TO SECTIONS B5 AND H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION SUMMARY. ADDITIONAL PRODUCT: SERIAL#: (B)(6), H3: YES, H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) AND ONE (1) CONTROLLER (B)(6) WERE NOT RETURNED FOR EVALUATION. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICE ISSUE. LOG FILE ANALYSIS REVEALED A SLIGHT DECREASE IN ESTIMATED FLOWS, WITH AN ASSOCIATED INCREASE IN PULSATILITY, STARTING ON (B)(6) 2025. REVIEW OF THE EVENT LOG FILE REVEALED A CONTROLLER POWER UP WITH ASSOCIATED PUMP START EVENT LOGGED ON 20/APR/2025 AT 18:03:49, INDICATING A LOSS OF POWER TO THE CONTROLLER. THE DATA POINT RECORDED PRIOR TO THE LOSS OF POWER REVEALED THAT (B)(6) WAS CONNECTED TO POWER PORT ONE (1) WITH 54% RELATIVE STATE OF CHARGE (RSOC) AND THAT (B)(6) WAS CONNECTED TO POWER PORT TWO (2) WITH 23% RSOC. THE DATA POINT RECORDED AFTER THE LOSS OF POWER REVEALED THAT (B)(6) WAS CONNECTED TO POWER PORT ONE AND THERE WAS NO CONNECTION IN POWER PORT TWO. NO ANOMALIES WERE RECORDED LEADING UP TO THE EVENT. THE CONTROLLER WAS WITHOUT POWER FOR 14 SECONDS. AS A RESULT, THE REPORTED LOW FLOWS, CHANGES IN PULSATILITY AND LOSS OF POWER WERE ABLE TO BE CONFIRMED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED LOW FLOW EVENT. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE REPORTED LOW FLOWS AND SUCTION MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. PER THE INSTRUCTIONS FOR USE, POSSIBLE ROOT CAUSES OF CHANGES IN PULSATILITY CAN BE ATTRIBUTED TO PATIENT CONDITIONS INCLUDING LEFT VENTRICULAR CONTRACTILITY, RIGHT HEART FUNCTION AND LEFT VENTRICULAR AFTERLOAD, WHICH CAN BE AFFECTED BY PUMP ROTATIONAL SPEED. A POSSIBLE ROOT CAUSE OF THE LOSS OF POWER CAN BE ATTRIBUTED TO THE DOUBLE DISCONNECTION OF POWER SOURCES AS DESCRIBED IN THE EVENT DETAILS. CAPA: PR00551638 INVESTIGATED CONTROLLER LOSSES OF POWER. EVEN THOUGH THIS CAPA IS NOW CLOSED, (B)(6) FALLS IN SCOPE OF THIS CAPA. PER THE INSTRUCTIONS FOR USE SYNCOPE, RENAL DYSFUNCTION, NEUROLOGICAL DYSFUNCTION, BLEEDING, THROMBOSIS, AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. BASED ON REVIEW OF PAST ADVERSE EVENTS, THIS PATIENT HAD A HISTORY OF THROMBUS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION. INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED AS ADDITIONAL INFORMATION HAS BEING RECEIVED FOR THIS EVENT AND SECTIONS HAVE BEEN UPDATED. ADDITIONAL PRODUCT: SERIAL# (B)(6). H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM - CONTROLLER 2.0 D4: MODEL #: 1420 / CATALOG #: 1420 / EXPIRATION DATE: 31-MAY-2024 / SERIAL #: (B)(6) D9: NO H3: NO H4: MFG DATE: 15-MAY-2023 H5: NO H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22037197&pc=DSQ | 2916596-2025-03021 | 2025-04-28 04:00 | Death | THORATEC CORPORATION | 2025-05-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIOGENIC SHOCK. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE PATIENT OUTCOME, AS WELL AS A DIRECT CORRELATION TO HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22037219&pc=NGV | 2124215-2025-31747 | 2025-04-18 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-16 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT DEATH OCCURRED. A CONCOMITANT LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 31MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED. SIX (6) DAYS POST INDEX PROCEDURE, THE PATIENT WAS DISCHARGED TO HOSPICE AND THE FOLLOWING DAY, THE PATIENT DIED. THE CAUSE OF DEATH REMAINS UNKNOWN, AND NO FURTHER DETAILS WERE PROVIDED.Event Description: IT WAS REPORTED THAT DEATH OCCURRED. A CONCOMITANT LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 31MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED. SIX (6) DAYS POST INDEX PROCEDURE, THE PATIENT WAS DISCHARGED TO HOSPICE AND THE FOLLOWING DAY, THE PATIENT DIED. THE CAUSE OF DEATH REMAINS UNKNOWN, AND NO FURTHER DETAILS WERE PROVIDED. IT WAS FURTHER REPORTED THE CONCOMITANT PROCEDURE WAS A PULSE-FIELD ABLATION (PFA) CASE USING FARAPULSE PFA. THERE WERE NO DIFFICULTIES DURING THE PFA ABLATION PORTION OR THE LAA CLOSURE PORTION OF THE CONCOMITANT PROCEDURE. Manufacturer Narrative: B5: INFORMATION ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22037523&pc=NGV | 2124215-2025-31361 | 2025-05-13 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-16 04:00 | NGV | SYSTEM, APPENDAGE CLOSURE, LEFT ATRIAL | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: REPORTED VIA PATIENT CALL CENTER. IT WAS REPORTED VIA THE WATCHMAN PATIENT CALL CENTER THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED, AND A WATCHMAN LAA CLOSURE DEVICE WAS IMPLANTED. A SIBLING OF THE PATIENT REPORTED TO THE WATCHMAN PATIENT CALL CENTER THAT THE WATCHMAN KILLED THEIR BROTHER. THERE WERE NO SPECIFIC DETAILS RELATED TO THE PATIENT OR PATIENT DEATH PROVIDED. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: B3 DATE OF EVENT : AWARE DATE WAS USED AS EVENT DATE AS ACTUAL EVENT DATE WAS NOT KNOWN. D1: BRAND NAME IS BLANK BECAUSE THE DEVICE IS KNOWN TO BE A WATCHMAN LAA CLOSURE DEVICE, BUT THE EXACT BRAND NAME IS UNKNOWN. D4 UNIQUE IDENTIFIER (UDI) #: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22037723&pc=OZD | 1220648-2025-28493 | 2025-04-29 04:00 | Death | ABIOMED, INC. | 2025-05-16 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Ischemia; Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT FOLLOWING IMPELLA CP IMPLANT, DISTAL PERFUSION TO THE IMPELLA LEG WAS ABSENT. A DISTAL PERFUSION SHEATH WAS PLACED TO RESTORE FLOW TO THE LIMB. AS SUPPORT CONTINUED, ONGOING SUCTION ALARMS APPEARED. AN ULTRASOUND SHOWED A LARGE PERICARDIAL EFFUSION AND A PERFORATION WAS SUSPECTED. UPON PULLING THE IMPELLA BACK, THE PIGTAIL APPEARED TO BE CAUGHT ON A STRUCTURE AND STRAIGHTENED. WHEN IT CAME FREE, THE EFFUSION CONTINUED TO REACCUMULATE AT A FASTER RATE. THE FAMILY MADE THE DECISION TO WITHDRAW CARE AND THE PATIENT PASSED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: VENTRICLE PERFORATION: CLINICAL DETAILS NOTE THAT A LARGE PERICARDIAL EFFUSION WAS FOUND AFTER THE PATIENT HAD GOTTEN MORE UNSTABLE WITH VERY LOW PULSATILITY, DECREASED FLOWS, AND SUCTION ALARMS. IT IS UNKNOWN WHEN THE EFFUSION BEGAN. THE EFFUSION KEPT REACCUMULATING EVEN AFTER DRAINAGE, SIGNALING A RV OR LV PERFORATION. THE IMPELLA PIGTAIL APPEARED TO BE CAUGHT ON A STRUCTURE, WHICH BECAME FREE AFTER PULLING THE PUMP BACK. THE TIMING OF THE REPOSITIONING IS UNKNOWN. THE EFFUSION REACCUMULATED AT A FASTER RATE AFTER THIS AND FLOWS BEGAN DECREASING AGAIN. CLINICAL DETAILS NOTED THAT IT WAS DIFFICULT TO RULE OUT THE IMPELLA AS THE CAUSE OF THE PERFORATION SINCE IMPELLA INSERTION WAS THE ONLY PROCEDURE PERFORMED THIS DAY; HOWEVER, CLINICAL DETAIL ALSO NOTES THE INFILTRATIVE MYOCARDITIS MAY HAVE MADE THE LV TISSUE MORE FRIABLE. DATA LOG REVIEW SHOWED NO POSITIONING ALARMS. THE RETURNED PRODUCT DID NOT HAVE ANY CANNULA KINKS. THE ROOT CAUSE OF THE VENTRICLE PERFORATION WAS NOT DETERMINED SINCE INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. LOW PUMP FLOW: CLINICAL DETAILS NOTE THE PATIENT HAD GOTTEN MORE UNSTABLE WITH VERY LOW PULSATILITY, DECREASED FLOWS, AND SUCTION ALARMS BEFORE A PERICARDIAL EFFUSION WAS FOUND. FLOWS CONTINUED DECREASING WITH MORE SUCTION ALARMS AS THE EFFUSION KEPT REACCUMULATING. THE IMPELLA PIGTAIL APPEARED TO BE CAUGHT ON A STRUCTURE, WHICH BECAME FREE AFTER PULLING THE PUMP BACK. DATA LOG REVIEW SHOWED LOW PUMP FLOW DURING THE ENTIRE CASE AT DIFFERENT P-LEVELS. IT ALSO SHOWED DOWNWARD TRENDING PLACEMENT SIGNAL AND DECREASING MOTOR CURRENT WITH STABLE P-LEVEL. THERE WERE SEVERAL SUCTION ALARMS BUT NO POSITIONING ALARMS OR MOTOR CURRENT SPIKES. THE RETURNED PRODUCT DID NOT HAVE ANY CANNULA KINKS. THERE WAS SMALL BIOMATERIAL IN THE PURGE GAP. LOW PUMP FLOW WAS NOT REPRODUCED WITH THE RETURNED PRODUCT. THE ROOT CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY PATIENT CONDITION RELATED AS THE PATIENT WAS UNSTABLE DURING SUPPORT WITH VERY LOW PULSATILITY AND PERICARDIAL EFFUSION BASED ON CLINICAL DETAILS AND DATA LOG REVIEW SHOWED DOWNWARD TRENDING IN PLACEMENT SIGNAL. LIMB ISCHEMIA - REVERSIBLE: CLINICAL DETAILS STATE DISTAL PERFUSION TO IMPELLA SIDE LEG WAS ABSENT CONFIRMED BY ULTRASOUND. THE TREATMENT TEAM DECIDED TO PLACE DISTAL PERFUSION SHEATHS, WHICH RESTORED FLOW. LIMB ISCHEMIA CAN OCCUR DURING IMPELLA SUPPORT. WHEN MAKING A DETERMINATION AS TO THE CAUSE OF THE LIMB ISCHEMIA, THE FOLLOWING CLINICAL INFORMATION IS NECESSARY: PATIENTS PRE-MORBID CONDITION AND PERI-PROCEDURAL CONDITION. ANY CONCOMITANT MEDICATION. VASCULAR SIZE OR VARIATIONS THEREOF. DETAILED DESCRIPTION OF THE SYMPTOMS EXPERIENCED BY THE PATIENT. PHYSICAL EXAMINATION FINDINGS, INCLUDING PULSE ASSESSMENT AND VISUAL EXAMINATION OF THE AFFECTED LIMB. DIAGNOSTIC IMAGING RESULTS, SUCH AS DOPPLER ULTRASOUND, ANGIOGRAPHY, OR MAGNETIC RESONANCE ANGIOGRAPHY. LABORATORY TEST RESULTS, INCLUDING BLOOD TESTS FOR MARKERS OF INFLAMMATION OR CLOTTING DISORDERS. ANY RELEVANT INFORMATION ABOUT RISK FACTORS FOR PERIPHERAL ARTERIAL DISEASE, SUCH AS SMOKING, DIABETES, OR HYPERTENSION. PATIENT'S RESPONSE TO PREVIOUS TREATMENTS OR INTERVENTIONS FOR VASCULAR CONDITIONS WITH A RETURNED IMPELLA, THE MAX OD COULD BE ASSESSED TO ENSURE IT IS WITHIN SPECIFICATION. THE PRODUCT COULD ALSO BE EVALUATED FOR ANY BURRS OR SHARP EDGES. ADDITIONALLY, THE CLINICAL INFORMATION ABOVE WOULD NEED TO BE CONSIDERED IN THE CONTEXT OF THE RETURNED PRODUCT. TO DETERMINE THE TRUE ROOT CAUSE OF LIMB ISCHEMIA, THE CLINICAL INFORMATION MENTIONED ABOVE WOULD NEED TO BE ASSESSED IN CONJUNCTION WITH EVIDENCE FROM THE RETURNED DEVICE. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22037878&pc=NPW | 2015691-2025-03899 | 2025-04-30 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-16 04:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Malposition of Device | Cardiac Perforation | P230013 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION OF AN EVOQUE PROCEDURE IN TRICUSPID POSITION WHERE THE WIRE PRESSURE WAS USED TO ACHIEVE PROPER HEIGHT FOR TIMEOUT. WHEN THE CAPSULE WAS RETRACTED TO EXPOSE THE ANCHORS THERE WERE NO ISSUES. AT APPROXIMATELY 75 DEGREES OF ANCHOR RELEASE, THE WIRE JUMPED FORWARD DRAMATICALLY INTO THE RVOT. IT WAS AT THIS POINT THAT IMAGING LOOKED FOR, AND FOUND, A SIGNIFICANT EFFUSION. THE REMAINDER OF THE PROCEDURE WAS ACCELERATED AND THE VALVE WAS DEPLOYED AT APPROPRIATE HEIGHT. INITIAL INTERROGATION OF THE VALVE REVEALED NO PVL AND TRIVIAL CENTRAL LEAK WITH THE WIRE STILL ACROSS. AT THIS POINT THE IMAGING WAS ABANDONED TO GUIDE LIFE SAVING EFFORTS. EFFORTS WERE UNSUCCESSFUL. Manufacturer Narrative: THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED/ADDED: B4, G3, G6, H2, H6 AND H11. ADDITIONAL TYPE OF INVESTIGATION CODES THAT WERE UNABLE TO BE ADDED IN H6: LABELING REVIEW. INSUFFICIENT INFORMATION AVAILABLE. THE COMPLAINT FOR DELIVERY SYSTEM AND/OR GUIDEWIRE PUSHED INTO VENTRICULAR WALL WAS CONFIRMED WITH OTHER EMPIRICAL EVIDENCE BASED ON INFORMATION GATHERED FROM THE EDWARDS CLINICAL SPECIALIST. THE KEY CONTRIBUTING FACTORS SUPPORTING THIS COMPLAINT INCLUDE SUBOPTIMAL WIRE PLACEMENT, THE USE OF WIRE PRESSURE TO GAIN HEIGHT, AND THE BEHAVIOR OF WIRE DURING ANCHOR RELEASE. NO MANUFACTURING NON-CONFORMITIES WERE ABLE TO BE IDENTIFIED. AVAILABLE INFORMATION SUGGESTS THAT THE BEHAVIOR OF WIRE DURING ANCHOR RELEASE LED TO THE PERICARDIAL EFFUSION, WHICH WAS CONFIRMED BY PHYSICIAN AND THE INTERVENTIONS ATTEMPTED TO STABILIZE THE PATIENT. SINCE NO MANUFACTURING NON-CONFORMANCES WERE FOUND AND NO LABELING, TRAINING, OR IFU DEFICIENCIES WERE IDENTIFIED, NO PREVENTATIVE OR CORRECTIVE ACTIONS WERE REQUIRED. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22037973&pc=QEW | 3020347218-2025-00030 | 2025-04-18 04:00 | Death | INARI MEDICAL, INC. | 2025-05-16 04:00 | QEW | CLOTTRIEVER BOLD CATHETER | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism | K230494 |  | 1 | Event Description: A 72-YEAR-OLD FEMALE UNDERWENT A THROMBECTOMY WITH INARI DEVICES TO ADDRESS HER DEEP VEIN THROMBOSIS. THE PATIENT HAD CLOT RANGING FROM THE FEMORAL VEIN INTO THE INFERIOR VENA CAVA (IVC). THE PLAN WAS TO PERFORM THE PROCEDURE WITHOUT EMBOLIC PROTECTION. WIRE POSITIONING WAS OBTAINED WITHOUT ISSUE AND A CLOTTRIEVER BOLD CATHETER WAS ADVANCED INTO PLACE AFTER A 13 FR CLOTTRIEVER SHEATH WAS INSERTED. 3 PASSES WERE PERFORMED THEN THE VITALS MACHINE STOPPED PICKING UP THE PATIENT'S OXYGEN SATURATION. THE PATIENT THEN BECAME TACHYCARDIC AND BEGAN HAVING PREMATURE VENTRICULAR CONTRACTIONS. THE PROCEDURE WAS ABORTED IN ATTEMPTS TO STABILIZE THE PATIENT. A RAPID RESPONSE WAS INITIATED, AND A CODE WAS CALLED SOON AFTER. CARDIOPULMONARY RESUSCITATION WAS CONDUCTED FOR 45 MINUTES BEFORE THE PATIENT'S FAMILY DECIDED TO DISCONTINUE IT. THE PATIENT EXPIRED. FOLLOW-UP INFORMATION REVEALED THAT THE PATIENT'S INITIAL COMPUTED TOMOGRAPHY SCAN SHOWED A SMALL DISTAL PULMONARY EMBOLISM, THE PHYSICIAN DECIDED NOT TO INTERVENE. THE PHYSICIAN ALSO INDICATED THAT THE PATIENT'S DETERIORATION COULD POTENTIALLY BE THE RESULT OF SEDATION AND/OR CLOT BEING SENT TO THE PULMONARY ARTERIES. Manufacturer Narrative: THE INARI DEVICES WERE DISCARDED BY THE USER FACILITY AND ARE NOT AVAILABLE FOR EVALUATION. THE DEVICE IDENTIFIERS WERE NOT PROVIDED; THE LOT HISTORY RECORDS OF ALL POTENTIAL LOTS SHIPPED TO THE FACILITY WERE REVIEWED. THERE WERE NO DISCREPANCIES OR UNUSUAL FINDINGS. THERE WAS NO REPORT OF ANY DEVICE MALFUNCTION. BASED ON THE INFORMATION REVIEWED, THERE IS NO EVIDENCE TO INDICATE THE PRESENCE OF A POTENTIAL QUALITY ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. THE CASE WAS REVIEWED BY INARI'S CHIEF MEDICAL OFFICER, WHO CONCLUDED THAT THE PATIENT'S DEATH MAY HAVE BEEN THE RESULT OF INADVERTENT CLOT EMBOLISM. DISTAL EMBOLIZATION OF BLOOD CLOTS IS IDENTIFIED IN THE DEVICE LABELING AS POSSIBLE ADVERSE EVENTS/COMPLICATIONS. THE MANUFACTURER'S REFERENCE #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22038210&pc=PJF | 3003925919-2025-00002 | 2025-04-02 04:00 | Death | CERUS CORPORATION | 2025-05-16 04:00 | PJF | INTERCEPT BLOOD SYSTEM FOR PLATELETS | Microbial Contamination of Device | Sepsis | BP140143 |  | 1 | Event Description: SEPTIC REACTION. INITIAL INFORMATION RECEIVED: 18-APR-2025, 21-APR-2025, 23-APR-2025, 28-APR-2025, 06-MAY-2025. INITIAL ADVERSE EVENT REPORT: ON 18-APR-2025, CERUS RECEIVED A SERIOUS ADVERSE EVENT REPORT OF SEPTIC REACTION. THIS ADVERSE EVENT WAS REPORTED TO CERUS BY THE VICE PRESIDENT OF QUALITY ASSURANCE AT (B)(6). MEDICAL HISTORY AND MEDICATIONS: THIS REPORT INVOLVED A 34-YEAR-OLD FEMALE WITH MEDICAL HISTORY OF ACUTE MYELOID LEUKEMIA (AML) STATUS POST STEM CELL TRANSPLANT ON (B)(6) 2024, NON-COVID CORONAVIRUS, AND NEUTROPENIA. THE PATIENT'S PAST TRANSFUSION HISTORY WAS NOT REPORTED. CONCOMITANT MEDICATIONS WERE MEROPENUM, LINEZOLILD, BACTRIM, VALACYCLOVIR, POSACONAZOLE, AND LEVAQUIN. ADVERSE EVENT DESCRIPTION: PATIENT WAS ADMITTED TO (B)(6) ON (B)(6) 2025 FOR PERIRECTAL ABSCESS TREATMENT. MEROPENUM AND LINEZOLILD WERE STARTED ON ADMISSION, BACTRIM WAS ADDED ON (B)(6) 2025. PROPHYLACTIC ANTIMICROBIALS VALACYCLOVIR, POSACONAZOLE AND LEVAQUIN WERE ALSO STARTED ON (B)(6) 2025. BLOOD CULTURE PERFORMED ON ADMISSION WAS NEGATIVE; NO WOUND CULTURES WERE DONE DURING HOSPITALIZATION. PRE-TRANSFUSION, THE PATIENT TESTED NEGATIVE CMV, EBV, HHV6, HEP C AND ASPERGILLUS. PICC LINE BLOOD CULTURE WAS NEGATIVE. ON (B)(6) 2025, PATIENT'S URINE CULTURE WAS NEGATIVE. THE PATIENT RECEIVED RBC & PLATELET TRANSFUSIONS ON (B)(6) 2025 WITHOUT INCIDENT. ON (B)(6) 2025 23:45H, THE PATIENT RECEIVED 140 ML OF A 4-DAY-OLD IMPLICATED PSORALEN TREATED PLATELETS (PTP) THROUGH THE PICC LINE FOR PANCYTOPENIA [DONATION IDENTIFICATION NUMBER (DIN): (B)(6), BLOOD COMPONENT PRODUCT CODE: E8331V00]. SHE DEVELOPED BACK PAIN, CHILLS, RIGORS, HEADACHE, HYPOTENSION, TACHYCARDIA, TACHYPNEA AND FEVER APPROXIMATELY ONE HOUR INTO THE TRANSFUSION. SHE WAS TRANSFERRED TO THE INTENSIVE CARE UNIT FOR A HIGHER LEVEL OF CARE AND A CENTRAL LINE WAS PLACED TO ALLOW ADVANCED CARE IN THE SETTING OF NEW ONSET SHOCK. THE PATIENT'S PRE-TRANSFUSION (23:45H) / DURING TRANSFUSION / POST-TRANSFUSION (00:45H) VITAL SIGNS WERE: BODY TEMPERATURE: 36.6 °C / NOT REPORTED / 37.7 °C. RESPIRATORY RATE: 15 BREATHS PER MINUTE/ NOT REPORTED / 20 BREATHS PER MINUTE. BLOOD PRESSURE: 100/62 MMHG / NOT REPORTED / 65/30 MMHG. HEART RATE: 77 BPM / NOT REPORTED / 143 BPM. THE IMPLICATED PTP AND PATIENT'S POST-TRANSFUSION (COLLECTED THROUGH THE NEWLY PLACED CENTRAL LINE) SAMPLES WERE GRAM STAINED AND CULTURED. BOTH SAMPLES WERE POSITIVE WITH GRAM NEGATIVE RODS, IDENTIFIED AS KLEBSIELLA PNEUMONIAE. PER HOSPITAL MICROBIOLOGY LAB, THESE WERE THE ONLY SAMPLES THAT TESTED POSITIVE FOR KLEBSIELLA PNEUMONIAE ON THAT DAY. ENVIRONMENTAL CONTAMINATION WAS RULED OUT BY THE MICROBIOLOGY LAB. THE PATIENT DIED ON (B)(6) 2025, 26 DAYS AFTER THE SEPTIC TRANSFUSION EVENT AND CAUSE OF DEATH WAS REPORTED AS RESPIRATORY FAILURE. THE SEPTIC TRANSFUSION EVENT WAS NOT IMPLICATED IN THE CAUSE OF DEATH. AUTOPSY WILL NOT BE PERFORMED. BLOOD DONOR(S) INFORMATION: DEMOGRAPHIC DETAILS OF THE DONOR OF THE IMPLICATED PTP WERE NOT REPORTED. THE DONOR WAS A REPEAT DONOR WITH FOUR PLATELET DONATIONS IN 2025. THIS DONATION WAS COLLECTED AT (B)(6) AND PROCESSED IN (B)(6). AT THE TIME OF DONATION, IODINE WAS USED AS ANTI-SEPTIC AGENT TO CLEAN THE VENIPUNCTURE AREA AND THE DONOR HAD NO SCARRING REPORTED ON ANTECUBITAL AREA. THERE WERE NO DEVIATIONS NOTED DURING THE DONATION PROCESS. IMPLICATED PATHOGEN REDUCED BLOOD COMPONENT(S) COLLECTION AND MANUFACTURING INFORMATION: ON (B)(6) 2025 AT 10:46H, (B)(6) CENTER COMPLETED COLLECTION OF AN APHERESIS PLATELET SUSPENDED IN 100% PLASMA (364 ML) USING THE TRIMA APHERESIS INSTRUMENT. ON (B)(6) 2025 AT 10:23H, THE IMPLICATED PTP WAS TREATED WITH INTERCEPT BLOOD SYSTEM FOR PLATELETS PROCESSING SET (LOT#: CE24I12A01) AND WAS ILLUMINATED WITH INTERCEPT ILLUMINATOR (SERIAL#: (B)(6). THE IMPLICATED PTP WAS INTERCEPT TREATED AND STORED AT THE SAME MANUFACTURING FACILITY IN GAINESVILLE FLORIDA. THE IMPLICATED PTP WAS TRANSPORTED VIA COURIER FROM THE BLOOD CENTER TO THE HOSPITAL, WHICH IS APPROXIMATELY 15 MINUTES APART. CO-COMPONENT(S) INFORMATION: THERE WERE NO CO-COMPONENTS. ADVERSE EVENT INVESTIGATIONS: (B)(6) INVESTIGATION: THE IMPLICATED PTP PASSED VISUAL INSPECTION AT THE TIME OF RELEASE FOR TRANSFUSION. THE IMPLICATED PTP WAS WRAPPED AND SENT IMMEDIATELY TO THE HOSPITAL TRANSFUSION SERVICE AFTER RECOGNITION OF THE ADVERSE EVENT. THE HOSPITAL TRANSFUSION SERVICE PERFORMED PRESSURE TESTING OF THE IMPLICATED PTP BAG USING THE PLASMA EXPRESSOR AND FOUND NO OBVIOUS LEAKS. MULTIPLE REPEAT BLOOD CULTURES WERE PERFORMED AFTER THE REPORTED ADVERSE EVENT ON 02-APR-2025, 04-APR-2025, 11-APR-2025, 22-APR-2025 AND 26-APR-2025, RESULTS WERE ALL NEGATIVE. URINE CULTURE ON (B)(6) 2025 WAS NEGATIVE. SPUTUM AND BRONCHOALVEOLAR LAVAGE CULTURES WERE POSITIVE FOR CANDIDA GLABRATA ON (B)(6) 2025. NON-COVID-19 CORONAVIRUS PCR WAS POSITIVE ON (B)(6) 2025. (B)(6) INVESTIGATION: THERE WERE NO DEVIATIONS ASSOCIATED WITH THE PATHOGEN REDUCTION PROCESS OR STERILE CONNECT DEVICE. THERE WAS NO VISIBLE DAMAGE TO THE IMPLICATED PTP BAG. THE IMPLICATED CAD BAG HAS BEEN DISCARDED. THE DONOR RETURNED FOR ANOTHER PLATELET DONATION ON (B)(6) 2025 AND SAMPLES WERE CULTURED IN BOTH AEROBIC AND ANAEROBIC BOTTLES USING BACT-ALERT WITH NEGATIVE RESULTS. CERUS INVESTIGATION: BATCH RECORD REVIEW: A BATCH RECORD REVIEW FOR INTERCEPT BLOOD SYSTEM FOR PLATELETS PROCESSING SET (PRODUCT CODE#: INT2240B / LOT#: CE24I12A01; EXPIRATION DATE: 04-DEC-2025) WAS PERFORMED. THE REVIEW OF THE BATCH RECORDS FOR THE FINISHED GOODS DID NOT INDICATE ANY OUT-OF-LIMITS RESULTS DURING TESTING AND CONTROLS. THE PRODUCT MET ALL QUALITY REQUIREMENTS AND SPECIFICATIONS AT THE TIME OF RELEASE. THIS BATCH WAS RELEASED ON 30-DEC-2024. NO OTHER COMPLAINTS HAVE BEEN RECEIVED FOR THIS LOT. THE TOTAL BATCH SIZE WAS (B)(4) PROCESSING SETS. PATHOGEN INACTIVATION INVESTIGATION: THE KLEBSIELLA PNEUMONIAE CULTURE ISOLATES FROM THE PATIENT AND IMPLICATED PTP HAVE BEEN RECEIVED BY THE CERUS MICROBIOLOGY TEAM. PATHOGEN INACTIVATION STUDY IS PENDING. FRESENIUS KABI (FK) INVESTIGATION: THE IMPLICATED PTP BAG WILL BE SENT TO FK FOR INVESTIGATION OF LEAKS AND DAMAGE. ADVERSE EVENT ASSESSMENT: 1.~SEPTIC REACTION [PT: TRANSMISSION OF AN INFECTIOUS AGENT VIA PRODUCT] THE REPORTER ASSESSED THE EVENT OF SEPTIC REACTION AS DEATH IN SEVERITY AND SERIOUSNESS. POSSIBLE IN RELATION TO PSORALEN TREATED PLATELETS AND EXCLUDED IN RELATION TO INTERCEPT BLOOD SYSTEM FOR PLATELETS. THE CERUS MEDICAL REVIEWER CONSIDERED THE EVENT AS SERIOUS DUE TO DEATH. POSSIBLE IN RELATION TO PSORALEN TREATED PLATELETS AND INTERCEPT BLOOD SYSTEM FOR PLATELETS. THIS PATIENT WAS NEUTROPENIC AND ADMITTED TO THE HOSPITAL FOR PERIRECTAL ABSCESS TREATMENT. WHILE THE BLOOD CULTURE ON ADMISSION/PRE-TRANSFUSION WAS NEGATIVE, A WOUND CULTURE WAS NOT DONE DURING THE HOSPITALIZATION. THE CLINICAL PRESENTATION OF THE ADVERSE EVENT IS CONSISTENT WITH A SEPTIC TRANSFUSION REACTION, ESPECIALLY GIVEN THE KLEBSIELLA PNEUMONIAE CULTURE RESULTS OBTAINED FROM BOTH THE PATIENT'S POST-TRANSFUSION SAMPLE AND THE IMPLICATED PTP. GIVEN THE POSITIVE GRAM STAIN RESULT, WHICH SUGGESTS A HIGH BACTERIAL LOAD, IT IS POSSIBLE THAT VISUAL INSPECTION MIGHT HAVE PREVENTED THE RELEASE OF THIS UNIT IF IT WAS CONTAMINATED. THERE WERE NO VISUAL INDICATIONS AT THE TIME OF RELEASE AND TRANSFUSION THAT THE IMPLICATED PTP MAY BE CONTAMINATED. FINALLY, THE INTERCEPT BLOOD SYSTEM FOR PLATELETS CAN EFFECTIVELY INACTIVATE KLEBSIELLA PNEUMONIAE BY >6.2 LOG, WHICH WOULD RAISE THE POSSIBILITY OF POST-INTERCEPT TREATMENT CONTAMINATION RATHER THAN A FAILURE OF THE TECHNOLOGY TO INACTIVATE THE CONTAMINATING BACTERIA. PATHOGEN INACTIVATION STUDY ON THE KLEBSIELLA PNEUMONIAE CULTURE ISOLATES WILL BE CONDUCTED TO VERIFY THE ACTUAL LOG INACTIVATION USING THE INTERCEPT BLOOD SYSTEM FOR PLATELETS. FURTHERMORE, IMPLICATE PTP BAG INTEGRITY INVESTIGATION WILL BE COMPLETED BY FK TO RULE OUT THE POSSIBILITY OF POST-INTERCEPT TREATMENT CONTAMINATION. FINALLY, ALTHOUGH THE PATIENT EVENTUALLY PASSED AWAY 26 DAYS AFTER THE SEPTIC TRANSFUSION EVENT, HER UNDERLYING MEDICAL CONDITION, OPPORTUNISTIC PNEUMONIA AND RESULTANT RESPIRATORY FAILURE WAS CONSIDERED THE CAUSE OF DEATH. NUMEROUS BLOOD CULTURES PERFORMED AFTER THE INITIAL POSITIVE POST-TRANSFUSION BLOOD CULTURE, WERE NEGATIVE, WHICH SUGGEST THAT THE SYSTEMIC INFECTION WAS EFFECTIVELY TREATED. THE CERUS MEDICAL REVIEWER CONSIDERED THE LISTEDNESS OF THIS EVENT IN RELATION TO THE PSORALEN TREATED PLATELETS AND INTERCEPT BLOOD SYSTEM FOR PLATELETS AS LISTED BASED ON APPLICABLE CERUS REFERENCE SAFETY INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22038268&pc=OZD | 1220648-2025-28499 | 2025-04-26 04:00 | Death | ABIOMED, INC. | 2025-05-16 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT SUFFERED A STROKE DURING IMPELLA 5.5 SUPPORT. NO INTERVENTION WAS REQUIRED TO MANAGE THE CONDITION. THREE DAYS LATER, THE FAMILY OPTED TO PURSUE COMFORT CARE AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)." SECTION: WARNINGS, CAUTIONS AND PRECAUTIONS: "TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22038482&pc=OZD | 1220648-2025-28500 | 2025-04-25 04:00 | Death | ABIOMED, INC. | 2025-05-16 04:00 | OZD | IMPELLA CP | Positioning Failure | Hemolysis; Renal Failure | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD HEMOLYSIS WITH WORSENING ACUTE KIDNEY INJURY. THE PATIENT HAD ELEVATED LACTATE DEHYDROGENASE. THE PUMP WAS REPOSITIONED DUE TO BEING IN CONTACT WITH WALL OF LEFT VENTRICLE. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS, RENAL FAILURE, AND POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. HEMOLYSIS: THE ROOT CAUSE OF THE HEMOLYSIS WAS MOST LIKELY IMPROPER POSITIONING SINCE THE HEMOLYSIS WAS SUCCESSFULLY MANAGED BY REPOSITIONING THE PUMP PER CLINICAL DETAILS. RENAL FAILURE: THE CAUSE OF THE RENAL FAILURE WAS NOT DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. POSITIONING ISSUES: THE ROOT CAUSE OF THE POSITIONING ISSUES WAS MOST LIKELY PATIENT CONDITION RELATED AS THE PATIENT HAD AN INCREDIBLY SMALL VENTRICLE PER CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22038499&pc=DSQ | 2916596-2025-03163 | 2025-04-17 04:00 | Death | THORATEC CORPORATION | 2025-05-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Fluid/Blood Leak; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Coagulation Disorder; Hemorrhage/Bleeding; Hypersensitivity/Allergic reaction; Renal Failure; Liver Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT ELECTIVE HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION WITH CORONARY ARTERY BYPASS GRAFT SURGERY (CABG) X 1, TRICUSPID ANNULOPLASTY AND LEFT ATRIAL APPENDAGE (LAA) LIGATION. FOLLOWING 20 MINUTES OF REPERFUSION, THE PATIENT'S RIGHT VENTRICLE (RV) APPEARED SLUGGISH AND CENTRIMAG RIGHT VENTRICULAR ASSIST DEVICE (RVAD) WAS INITIATED. MULTIPLE CONFIGURATIONS OF RVAD WERE ATTEMPTED BUT IT FAILED TO IMPROVE DRAINAGE. THE CIRCUIT WAS THEN CONVERTED TO EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WITH TEMPORARY CHEST CLOSURE IN VIEW OF PROLONGED SURGERY TIME. POST-OPERATIVELY, THE PATIENT WAS SUPPORTED ON HIGH INOTROPES. THE PATIENT HAD PERSISTENT HEMODYNAMIC INSTABILITY DUE TO ONGOING BLEEDING DESPITE MAXIMAL INOTROPIC REQUIREMENT. THEY DEVELOPED SEVERE ACIDOSIS, VASOPLEGIA AND WERE COAGULOPATHIC DESPITE MASSIVE CORRECTION AND BLOOD TRANSFUSION. THE PATIENT WAS BROUGHT BACK TO THE OPERATING ROOM FOR A RELOOK AND HEMOSTASIS. THERE WAS ONLY A SMALL OOZE FROM RIGHT ATRIAL (RA) CANNULA SITE NOTED WITH NO OTHER OBVIOUS SURGICAL BLEED OR CLOTS/TAMPONADE. THE PATIENT REMAINED COAGULOPATHIC WITH SEVERE VASOPLEGIA AND SEVERE THIRD SPACING AND CAPILLARY LEAK RESULTING IN INABILITY TO SUSTAIN LVAD AND ECMO FLOWS. TH ECMO FLOWS WERE SUBOPTIMAL <2 LITER PER MINUTE (L/MIN) AND LVAD FLOWS <1 L/MIN. THE PATIENT WAS PERSISTENTLY ACIDOTIC DESPITE RESUMING CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT). THE DECISION WAS MADE FOR BEDSIDE REOPENING IN INTENSIVE CARE UNIT (ICU). THERE WERE NO CLOTS/TAMPONADE NOR OBVIOUS BLEEDING POINT NOTED, AND THE WOUND WAS CLOSED AGAIN. THE PATIENT CONTINUED TO DECLINE DESPITE MAXIMAL ACTIVE TREATMENT. IN VIEW OF POOR PROGNOSIS, A FAMILY CONFERENCE WAS HELD AND THE DECISION MADE TO TERMINATE LVAD AND ECMO SUPPORT. THE PATIENT PASSED AWAY ON (B)(6) 2025 FOLLOWING WITHDRAWAL OF LIFE SUPPORT. THE CAUSE OF DEATH WAS THOUGHT TO BE DILATED CARDIOMYOPATHY, SEVERE VASOPLEGIA LIKELY FROM LIVER FAILURE, AND A SEVERE REACTION TO PROTAMINE. TH DEVICE WAS NOT EXPLANTED.Event Description: THE CAUSE OF DEATH WAS CONFIRMED TO BE DILATED CARDIOMYOPATHY. THE DEATH WAS NOT DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. IT WAS NOTED THAT ALL DEATHS WITHIN 30 DAYS POST-SURGERY ARE MADE A CORONER'S CASE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME, AS WELL AS A DIRECT CORRELATION TO HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS INVESTIGATION. FURTHER, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW HAZARD ALARMS. ALTHOUGH A SPECIFIC CAUSE FOR THE ALARMS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS INVESTIGATION, THE ACCOUNT REPORTS THAT THE PATIENT¿S SEVERE VASOPLEGIA, SEVERE THIRD SPACING, AND CAPILLARY LEAK RESULTED IN AN INABILITY TO SUSTAIN LEFT VENTRICULAR ASSIST DEVICE (LVAD) AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) FLOWS. THE CONTROLLER EVENT LOG FILE CONTAINED RELEVANT EVENTS FROM (B)(6) 2025 AND CAPTURED LOW SPEED ADVISORY ALARMS THROUGHOUT THE DURATION OF THE LOG FILE DUE TO THE FIXED SPEED BEING SET BELOW THE LOW-SPEED LIMIT. FURTHER, THE ESTIMATED FLOW RANGED FROM 0.0-0.5 LPM THROUGHOUT THE FILE, RESULTING IN A CONTINUOUS LOW FLOW HAZARD ALARM. AN INTENTIONAL PUMP STOP WAS CAPTURED FROM 22:10:55 THROUGH 22:46:47, BOTH POWER CABLES WERE DISCONNECTED AT 22:45:57, AND THE DRIVELINE WAS DISCONNECTED AT 22:46:45. THESE EVENTS APPEARED CONSISTENT WITH THE REPORTED WITHDRAWAL OF SUPPORT. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE FIXED SPEED PRIOR TO THE INTENTIONAL PUMP STOP. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING RESPIRATORY FAILURE, RIGHT HEART FAILURE, HEPATIC DYSFUNCTION, AND DEATH. SECTION 1 OF THE IFU ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿HEARTMATE TOUCH COMMUNICATION SYSTEM¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LPM, A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. THIS SECTION ALSO OUTLINES SITUATIONS THAT CAN RESULT IN A LOW FLOW ALARM AND FURTHER EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU, UNDER "RIGHT HEART FAILURE", DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RVAD PLACEMENT. SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING," AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿ OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22039047&pc=DSQ | 2916596-2025-03168 | 2025-04-22 04:00 | Death | THORATEC CORPORATION | 2025-05-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO UNKNOWN ETIOLOGY. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6) , AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22039057&pc=LZO | 1818910-2025-08045 | 2024-05-09 04:00 | Death | DEPUY ORTHOPAEDICS INC US | 2025-05-16 04:00 | LZO | UNKNOWN HIP FEMORAL CONSTRUCT ACTIS | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Cardiac Arrest; High Blood Pressure/ Hypertension |  |  | 1 | Event Description: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: LEITNER L, SCHITZ F, SADOGHI P, PUCHWEIN P, HOLINKA J, LEITHNER A, KALCHER E. TREATMENT OF FEMORAL NECK FRACTURES USING ACTIS STEM: COMPLICATION RATE IN 188 UNCEMENTED HEMIARTHROPLASTIES. ARCH ORTHOP TRAUMA SURG. 2024 JUN;144(6):2831-2838. DOI: 10.1007/S00402-024-05352-Z. EPUB 2024 MAY 9. PMID: 38722364; PMCID: PMC11211130. OBJECTIVE/METHODS/STUDY DATA: THIS STUDY AIMED TO COMPARE OUR RESULTS TO EARLIER PUBLISHED DATA ON CEMENTED HA, WE WANTED TO EVALUATE WHETHER THIS TECHNIQUE REMAINS A REASONABLE TREATMENT OPTION DESPITE CONTRADICTORY RECOMMENDATIONS, OR WHETHER WE SHOULD SWITCH TO A CEMENTED SYSTEM. BETWEEN JANUARY 2017 AND SEPTEMBER 2022, 1,194 PATIENTS TREATED WITH FEMORAL NECK FRACTURE IN OUR TWO STUDY CENTERS (MEDICAL UNIVERSITY OF GRAZ, STYRIA, AUSTRIA AND HANUSCH KRANKENHAUS, VIENNA, AUSTRIA), WERE RETRIEVED FROM OUR DATA SYSTEM. FROM THESE CASES, 188 PATIENTS (15.7%) TREATED WITH HA WITH AN UNCEMENTED, COLLARED WEDGE IMPLANT (ACTIS®, DEPUY SYNTHES,) WERE IDENTIFIED, AND THEREFORE RETROSPECTIVELY INCLUDED IN THIS STUDY. THERE WERE 122 FEMALE AND 66 MALE WITH A MEAN AGE OF 83.1 YEARS. THE MEAN FOLLOW - UP WAS UNKNOWN. LOT, MODEL AND CATALOG NUMBER ARE NOT AVAILABLE, BUT THE SUSPECTED DEPUY SYNTHES DEVICE(S) POSSIBLY ASSOCIATED WITH REPORTED ADVERSE EVENTS: DEPUY ACTIS® UNCEMENTED STEM WITH BIPOLAR HEAD. MANUFACTURER OF THE ACETABULAR CUP/ LINER ARE NOT MENTIONED. ADVERSE EVENT(S) AND PROVIDED INTERVENTIONS POSSIBLY ASSOCIATED WITH UNK HIP FEMORAL CONSTRUCT ACTIS (QTY.2) 2 REPORTED DEATHS OCCURRED WITHIN 48 HOURS OF THE PATIENTS' SURGICAL PROCEDURES; ONE PATIENT HAD A CARDIAC ARREST, AND THE OTHER HAD A PULMONARY EMBOLISM AND HYPERTENSION. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF THE INFORMATION IS UNKNOWN, NOT AVAILABLE OR DOES NOT APPLY, THE SECTION/FIELD OF THE FORM IS LEFT BLANK. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. H3, H6: NO DEVICE ASSOCIATED WITH THIS REPORT WAS RECEIVED FOR EXAMINATION. THIS COMPLAINT WAS OPENED TO DOCUMENT COMPLAINTS DERIVED THROUGH A JOURNAL ARTICLE REVIEW. FOLLOW-UPS WERE DONE TO TRY AND OBTAIN ADDITIONAL INFORMATION FROM THE AUTHOR OF THE JOURNAL ARTICLE. NO FURTHER INFORMATION WAS RECEIVED. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS WILL BE CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF THE POST-MARKET SURVEILLANCE. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. H3, H6: THE DEVICE LOT IS UNKNOWN, THEREFORE A DEVICE HISTORY REVIEW COULD NOT BE PERFORMED. IF THE LOT/SERIAL NUMBER BECOMES AVAILABLE, THE RECORD WILL BE RE-ASSESSED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22039077&pc=LZO | 1818910-2025-08046 | 2024-05-09 04:00 | Death | DEPUY ORTHOPAEDICS INC US | 2025-05-16 04:00 | LZO | UNKNOWN HIP ACETABULAR CONSTRUCT | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Cardiac Arrest; High Blood Pressure/ Hypertension |  |  | 1 | Event Description: THIS COMPLAINT IS FROM A LITERATURE SOURCE. THE FOLLOWING LITERATURE CITE HAS BEEN REVIEWED: LEITNER L, SCHITZ F, SADOGHI P, PUCHWEIN P, HOLINKA J, LEITHNER A, KALCHER E. TREATMENT OF FEMORAL NECK FRACTURES USING ACTIS STEM: COMPLICATION RATE IN 188 UNCEMENTED HEMIARTHROPLASTIES. ARCH ORTHOP TRAUMA SURG. 2024 JUN;144(6):2831-2838. DOI: 10.1007/S00402-024-05352-Z. EPUB 2024 MAY 9. PMID: 38722364; PMCID: PMC11211130. OBJECTIVE/METHODS/STUDY DATA: THIS STUDY AIMED TO COMPARE OUR RESULTS TO EARLIER PUBLISHED DATA ON CEMENTED HA, WE WANTED TO EVALUATE WHETHER THIS TECHNIQUE REMAINS A REASONABLE TREATMENT OPTION DESPITE CONTRADICTORY RECOMMENDATIONS, OR WHETHER WE SHOULD SWITCH TO A CEMENTED SYSTEM. BETWEEN JANUARY 2017 AND SEPTEMBER 2022, 1,194 PATIENTS TREATED WITH FEMORAL NECK FRACTURE IN OUR TWO STUDY CENTERS (MEDICAL UNIVERSITY OF GRAZ, STYRIA, AUSTRIA AND HANUSCH KRANKENHAUS, VIENNA, AUSTRIA), WERE RETRIEVED FROM OUR DATA SYSTEM. FROM THESE CASES, 188 PATIENTS (15.7%) TREATED WITH HA WITH AN UNCEMENTED, COLLARED WEDGE IMPLANT (ACTIS®, DEPUY SYNTHES,) WERE IDENTIFIED, AND THEREFORE RETROSPECTIVELY INCLUDED IN THIS STUDY. THERE WERE 122 FEMALE AND 66 MALE WITH A MEAN AGE OF 83.1 YEARS. THE MEAN FOLLOW - UP WAS UNKNOWN. LOT, MODEL AND CATALOG NUMBER ARE NOT AVAILABLE, BUT THE SUSPECTED DEPUY SYNTHES DEVICE(S) POSSIBLY ASSOCIATED WITH REPORTED ADVERSE EVENTS: DEPUY ACTIS® UNCEMENTED STEM WITH BIPOLAR HEAD. MANUFACTURER OF THE ACETABULAR CUP/ LINER ARE NOT MENTIONED. ADVERSE EVENT(S) AND PROVIDED INTERVENTIONS POSSIBLY ASSOCIATED WITH UNK HIP ACETABULAR CONSTRUCT (QTY.2) 2 REPORTED DEATHS OCCURRED WITHIN 48 HOURS OF THE PATIENTS' SURGICAL PROCEDURES; ONE PATIENT HAD A CARDIAC ARREST, AND THE OTHER HAD A PULMONARY EMBOLISM AND HYPERTENSION. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF THE INFORMATION IS UNKNOWN, NOT AVAILABLE OR DOES NOT APPLY, THE SECTION/FIELD OF THE FORM IS LEFT BLANK. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. H3, H6: NO DEVICE ASSOCIATED WITH THIS REPORT WAS RECEIVED FOR EXAMINATION. THIS COMPLAINT WAS OPENED TO DOCUMENT COMPLAINTS DERIVED THROUGH A JOURNAL ARTICLE REVIEW. FOLLOW-UPS WERE DONE TO TRY AND OBTAIN ADDITIONAL INFORMATION FROM THE AUTHOR OF THE JOURNAL ARTICLE. NO FURTHER INFORMATION WAS RECEIVED. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS WILL BE CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF THE POST-MARKET SURVEILLANCE. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. H3, H6: THE DEVICE LOT IS UNKNOWN; THEREFORE, A DEVICE HISTORY REVIEW COULD NOT BE PERFORMED. IF THE LOT/SERIAL NUMBER BECOMES AVAILABLE, THE RECORD WILL BE RE-ASSESSED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22039095&pc=DSQ | 2916596-2025-02990 | 2025-04-28 04:00 | Death | THORATEC CORPORATION | 2025-05-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST, WHICH WAS NOT CONSIDERED DEVICE RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED OR RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22039172&pc=NPT | 2025587-2025-03741 | 2025-01-08 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-16 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22039176&pc=DSQ | 2916596-2025-02983 | 2025-01-24 05:00 | Death | THORATEC CORPORATION | 2025-05-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Fall; Retroperitoneal Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT HAD A NON-SYNCOPAL FALL AT HOME FOR WHICH THEY DID NOT SEEK CARE. A FEW DAYS LATER, THEY PRESENTED TO THE EMERGENCY DEPARTMENT ON (B(6) 2025 FOR FEELING UNWELL. INTERNATIONAL NORMALIZED RATIO (INR) AT THAT TIME WAS 3.5. IMAGING REVEALED A MASSIVE RETROPERITONEAL BLEED. THE PATIENT DECOMPENSATED SHORTLY AFTER THIS AND WAS PLACED IN THE INTENSIVE CARE UNIT (ICU) WHERE THEY ULTIMATELY PASSED AWAY ON (B)(6) 2025. THE FALL WAS MECHANICAL. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE LVAS WAS NOT RETURNED FOR EVALUATION. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22039199&pc=DSQ | 2916596-2025-03032 | 2025-04-17 04:00 | Death | THORATEC CORPORATION | 2025-05-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE INTRACRANIAL BLEEDING (ICB). THE OUTCOME WAS NOT CONSIDERED LEFT VENTRICULAR ASSIST DEVICE (LVAD) OR LVAD THERAPY RELATED. THE PUMP WAS WORKING AS EXPECTED. NO PUMP RELATED ISSUES WERE PROVIDED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED.Event Description: THE PATIENT REMAINED INTUBATED AFTER LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION. THE PATIENT DEVELOPED PUPILLARY DIFFERENCES AND DIFFERENCES IN NEAR INFRARED SPECTROSCOPY LEVELS. A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED, AND AN INTRACRANIAL BLEED WAS CONFIRMED. NO IMAGES WERE PROVIDED BY THE HOSPITAL. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22040406&pc=MFA | 3007284006-2025-00100 | 2025-05-08 04:00 | Death | SPECTRANETICS CORPORATION | 2025-05-16 04:00 | MFA | GLIDELIGHT LASER SHEATH | Adverse Event Without Identified Device or Use Problem | Sepsis | P960042 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) AND A RIGHT ATRIAL (RA) LEAD DUE TO BACTEREMIA. AT THE START OF THE PROCEDURE, A LARGE VEGETATION ON THE RV LEAD WAS DETECTED VIA TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). SPECTRANETICS LLD LEAD LOCKING DEVICES (LLDS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 14F GLIDELIGHT LASER SHEATH ON THE RV LEAD, PROGRESS WAS MADE TO THE TRICUSPID VALVE, WHERE THE VEGETATION WAS ENCOUNTERED. ADVANCEMENT CONTINUED, BUT THE VEGETATION BROKE APART AND RELEASED INTO THE BLOOD STREAM. WHILE SWITCHING TO A SPECTRANETICS 13F TIGHTRAIL ROTATING DILATOR SHEATH, THE PATIENT¿S BLOOD PRESSURE DROPPED, AND RESCUE EFFORTS BEGAN, INCLUDING CPR AND STERNOTOMY. THE RV LEAD WAS REMOVED POST-STERNOTOMY; HOWEVER, THE PATIENT DID NOT SURVIVE. POST-PROCEDURE, THE PHYSICIAN STATED THAT THE DEATH WAS NOT RELATED TO THE EXTRACTION, BUT RATHER DUE TO SEPSIS. THIS REPORT CAPTURES THE 14F GLIDELIGHT IN USE WHEN THE VEGETATION BROKE FREE FROM THE RV LEAD, REQUIRING INTERVENTION, BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF THE SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: A2) PATIENT DATE OF BIRTH - NOT AVAILABLE FROM FACILITY. A3B) PATIENT GENDER - NOT AVAILABLE FROM FACILITY. THE LOT NUMBER WAS NOT PROVIDED, THUS THE FOLLOWING INFORMATION IS UNKNOWN: DEVICE SERIAL NUMBER, UNIQUE ID, EXPIRATION DATE, AND MANUFACTURE DATE. H3) THE GLIDELIGHT WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) MIGRATION OF VEGETATION AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE GLIDELIGHT DEVICE. THE IFU WARNS "WHEN MARKED CALCIFICATION THAT MOVES WITH THE LEAD TO BE EXTRACTED IS SEEN ON FLUOROSCOPY, PARTICULARLY IN THE ATRIUM, THE AVAILABILITY OF IMMEDIATE SURGICAL ASSISTANCE IS PARAMOUNT IF A PROBLEM PRESENTS ITSELF AS A RESULT OF THE EXTRACTION PROCEDURE." SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22023626&pc=QLO | 3014015528-2025-00016 |  | Death | EXTHERA MEDICAL CORPORATION | 2025-05-15 04:00 | QLO | SERAPH 100 | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: THIS INFORMATION WAS REPORTED PUBLICLY BY THE NEW YORK TIMES. SEE CARRERAI, JOHN. "A START-UP CLAIMED ITS DEVICE COULD CURE CANCER. THEN PATIENTS BEGAN DYING." THE NEW YORK TIMES, (B)(6) 2025. PATIENT P6, NOT OTHERWISE IDENTIFIED, DIED FOLLOWING TREATMENT. OVERALL THE ARTICLE INDICATES THAT P6 WAS ONE OF SEVERAL PATIENTS SEEKING OFF-LABEL TREATMENT FOR CANCER. THE ARTICLE STATES THAT THESE TREATMENTS TOOK PLACE OUTSIDE THE UNITED STATES, IN ANTIGUA. FURTHER INVESTIGATION INDICATES THAT TREATMENT AND ANY SUBSEQUENT ADVERSE EVENT WOULD HAVE OCCURRED SOME TIME IN 2024, DUE TO THE TIMING OF THE ARTICLE AND THE AVAILABILITY OF TREATMENT IN ANTIGUA. Manufacturer Narrative: THIS WAS STATED IN THE NYT, SEE CARREYROU, JOHN. "A START-UP CLAIMED ITS DEVICE COULD CURE CANCER. THEN PATIENTS BEGAN DYING." THE NEW YORK TIMES, (B)(6) 2025. THE MANUFACTURER FOLLOWED UP WITH THE FACILITY IN ANTIGUA WHERE THE CLAIMED ADVERSE EVENTS TOOK PLACE AND WITH THE TREATING PHYSICIAN. THE FACILITY HAS STATED THAT IT WAS NOT AWARE OF ANY ADVERSE EVENTS. THE PHYSICIAN HAS NOT RESPONDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22023627&pc=QLO | 3014015528-2025-00006 | 2024-04-18 04:00 | Death | EXTHERA MEDICAL CORPORATION | 2025-05-15 04:00 | QLO | SERAPH 100 | Lack of Effect | Cancer |  |  | 1 | Event Description: THIS INFORMATION WAS REPORTED PUBLICLY BY THE NEW YORK TIMES. SEE CARREYROU, JOHN. "A START-UP CLAIMED ITS DEVICE COULD CURE CANCER. THEN PATIENTS BEGAN DYING." THE NEW YORK TIMES, JANUARY 23, 2025. AS REPORTED IN NYT ARTICLE, PATIENT P1 (IDENTIFIED AS MALE, 55, DIAGNOSED WITH LATE-STAGE METASTATIC ESOPHAGEAL CANCER) SAW PROGRESSION OF HIS CANCER AFTER TREATMENT. PER THE ARTICLE, "TUMORS IN HIS LIVER, ADRENAL GLANDS, BONES, AND SOFT TISSUES HAD MULTIPLIED AND GROWN." HIS WIFE CLAIMED THAT THE TREATMENT ACCELERATED MR. HUDLOW'S CANCER PROGRESSION. DAY OF INCIDENT UNKNOWN BUT ASSUMED TO BE BETWEEN THE DAY HE ARRIVED AT THE MAYO CLINIC AND THE DAY THE HE LEFT TO GO HOME (APRIL 10TH TO APRIL 16TH 2024). HE DIED ON APRIL 18TH 2024, APPROXIMATELY 1-2 WEEKS AFTER HIS LAST TREATMENT. THE ARTICLE STATES THAT SERAPH 100 TREATMENTS TOOK PLACE OUTSIDE THE UNITED STATES, IN ANTIGUA. Manufacturer Narrative: THIS WAS STATED IN THE NYT, SEE CARREYROU, JOHN. "A START-UP CLAIMED ITS DEVICE COULD CURE CANCER. THEN PATIENTS BEGAN DYING." THE NEW YORK TIMES, JANUARY 23, 2025. THE MANUFACTURER FOLLOWED UP WITH THE FACILITY IN ANTIGUA WHERE THE CLAIMED ADVERSE EVENTS TOOK PLACE AND WITH THE TREATING PHYSICIAN. THE FACILITY HAS STATED THAT IT WAS NOT AWARE OF ANY ADVERSE EVENTS. THE PHYSICIAN HAS NOT RESPONDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22023628&pc=QLO | 3014015528-2025-00009 | 2024-04-24 04:00 | Death | EXTHERA MEDICAL CORPORATION | 2025-05-15 04:00 | QLO | SERAPH 100 | Lack of Effect | Insufficient Information |  |  | 1 | Event Description: THIS INFORMATION WAS REPORTED PUBLICLY BY THE NEW YORK TIMES. SEE (B)(6). "A START-UP CLAIMED ITS DEVICE COULD CURE CANCER. THEN PATIENTS BEGAN DYING." THE NEW YORK TIMES, (B)(6) 2025. AS REPORTED IN NYT ARTICLE, PATIENT P2 (IDENTIFIED AS MALE, AGE UNKNOWN, DIAGNOSED WITH COLON CANCER) DIED ON (B)(6), APPROXIMATELY 2 WEEKS AFTER HIS LAST TREATMENT. THE ARTICLE STATES THAT THESE TREATMENTS TOOK PLACE OUTSIDE THE UNITED STATES, IN (B)(6). Manufacturer Narrative: THIS WAS STATED IN THE NYT, SEE (B)(6). "A START-UP CLAIMED ITS DEVICE COULD CURE CANCER. THEN PATIENTS BEGAN DYING." THE NEW YORK TIMES, (B)(6) 2025. THE MANUFACTURER FOLLOWED UP WITH THE FACILITY IN ANTIGUA WHERE THE CLAIMED ADVERSE EVENTS TOOK PLACE AND WITH THE TREATING PHYSICIAN. THE FACILITY HAS STATED THAT IT WAS NOT AWARE OF ANY ADVERSE EVENTS. THE PHYSICIAN HAS NOT RESPONDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22024346&pc=DYB | 9612164-2025-02405 | 2025-04-21 04:00 | Death | MEDTRONIC IRELAND | 2025-05-15 04:00 | DYB | MICRA | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiac Perforation; Vascular Dissection | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IT WAS NOTED THAT DURING THE IMPLANTATION OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT EXPERIENCED BLOOD LOSS, DISSECTION AND CARDIAC PERFORATION. IT WAS FURTHER REPORTED THAT THE MICRA INTRODUCER APPEARED TO "JUMP FORWARD". THE LEADLESS IPG WAS ATTEMPTED/NOT USED AND REMOVED.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IT WAS NOTED THAT DURING THE IMPLANTATION OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT EXPERIENCED FEMORAL COMPLICATIONS RESULTING IN DISSECTION AND BLOOD LOSS. IT WAS FURTHER REPORTED THAT THE MICRA INTRODUCER APPEARED TO "JUMP FORWARD". ONCE THE INTRODUCER WAS REMOVED ARTERIAL BLOOD WAS PUMPING OUT OF THE GROIN. MANUAL PRESSURE WAS ATTEMPTED, A FEMORAL ARTERY COMPRESSION DEVICE WAS APPLIED, BUT HEMOSTASIS WAS NOT MAINTAINED. THE VASCULAR SURGEON WAS CALLED IN AND THE PATIENT WAS TRANSPORTED TO THE SURGICAL SUITE. HOWEVER, THE PATIENT PASSED AWAY FROM FEMORAL COMPLICATIONS. Manufacturer Narrative: CORRECTED: B5; H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22024506&pc=OZD | 1220648-2025-28419 | 2025-05-02 04:00 | Death | ABIOMED, INC. | 2025-05-15 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND BRIDGE TO TRANSPLANT. APPROXIMATELY 13 DAYS AFTER START OF THE SUPPORT, PLACEMENT SIGNAL LOW ALARMS WITH SUCTION WERE ENCOUNTERED. THE ALARMS WERE NOTED TO HAVE SELF-RESOLVED. AN ULTRASOUND WAS COMPLETED AT BEDSIDE AND REVEALED THE IMPELLA 5.5 WAS MALPOSITIONED. AN ECHOCARDIOGRAM MEASURED THE CANNULA AT 6.7 CENTIMETERS FROM THE AORTIC VALVE ANNULUS. THE SURGEON REPOSITIONED WITH ECHOCARDIOGRAM AT BEDSIDE WITH CANNULA MEASUREMENT NOW 5 CENTIMETERS FROM THE AORTIC VALVE ANNULUS. THE PATIENT CONTINUED ON IMPELLA MCS AND WAS NOTED FOR POSSIBLE DURABLE VENTRICULAR ASSIST DEVICE PLACEMENT WORK UP. APPROXIMATELY 10 DAYS LATER, PLACEMENT ALARMS WERE TRIGGERED. AN ECHOCARDIOGRAM WAS PERFORMED, AND THE DEVICE WAS MEASURED AT 6.58 CENTIMETERS. THE PATIENT WAS TAKEN TO INTERVENTIONAL RADIOLOGY FOR A PROCEDURE. THE IMPELLA SUPPORT WAS ON P-2 AND AWAITING REPOSITIONING. IT WAS LATER REPORTED THAT THE PATIENT EXPERIENCED A RETROPERITONEAL BLEED AND HAD GONE TO INTERVENTIONAL RADIOLOGY FOR CAUTERIZATION OF THE BLEED NOTED TO BE UNRELATED TO THE IMPELLA 5.5 DEVICE. THE PATIENT WAS LETHARGIC AND PLACED ON A VENTILATOR OVER THAT WEEKEND. THE PATIENT WAS NOTED ON SEVERAL DRIPS. LACTIC HAD A DOWNWARD TREND AND THE LATEST ECHOCARDIOGRAM MEASURED 5.4 CENTIMETERS FROM THE AORTIC VALVE ANNULUS. A COUPLE OF DAYS LATER IT WAS NOTED THE PATIENT'S CONDITION WAS NOTED TO BE "ACTIVELY" DECLINING. A DECISION WAS MADE WITH THE HEART FAILURE TEAM AND FAMILY TO WITHDRAW CARE, AND THE PATIENT WOULD EXPIRE AFTER APPROXIMATELY 27 TOTAL DAYS ON IMPELLA MCS. Manufacturer Narrative: AS IT APPEARS HEPARIN WAS IN USE FOR THE IMPELLA DEVICE (AND NOTED SUBSEQUENTLY PAUSED), BLEEDING IS BEING REPORTED ACCORDINGLY (EVEN THOUGH IT WAS NOTED THE RETROPERITONEAL BLEED IS UNRELATED TO THE IMPELLA) AS IT UNKNOWN IF THE USE OF THE HEPARIN ADVERSELY IMPACTED AND/OR EXACERBATED THE SITUATION. ADDITIONALLY, GIVEN THE PATIENT¿S DEMISED OUTCOME, THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR GIVEN THE EVENT. THE IMPELLA DEVICE WAS NOT RETURNED, AND THE INVESTIGATION HAS BEEN COMPLETED. THE ROOT CAUSE OF VASCULAR DAMAGE PERIPHERAL VESSEL COULD NOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22024709&pc=NPT | 2015691-2025-03850 | 2025-04-18 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-15 04:00 | NPT | EDWARDS COMMANDER DELIVERY | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140031 |  | 1 | Event Description: PER REPORT RECEIVED FROM ITALY, IT WAS A CASE OF AN IMPLANT OF A 23MM SAPIEN 3 ULTRA VALVE IN AORTIC POSITION BY TRANSFEMORAL APPROACH. DURING THE FINAL ECHO CHECK, A PERICARDIAL EFFUSION WAS OBSERVED THAT WAS IMMEDIATELY TREATED WITH A PERICARDIAL PUNCTURE TO DRAIN THE EFFUSION. THE CAUSE OF THE PERICARDIAL EFFUSION WAS SUSPECTED TO BE RELATED TO THE NON-EDWARDS GUIDEWIRE PERFORATING THE VENTRICLE. THE PATIENT LEFT THE ROOM IN STABLE CONDITION BUT PASSED AWAY ON POD5. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR VENTRICULAR PERFORATION DURING A TVR PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS INDICATE PROCEDURAL FACTORS (NON-EDWARDS GUIDEWIRE PERFORATING THE VENTRICLE) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22025372&pc=OZP | 2032227-2025-175253 | 2025-04-16 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-15 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Myocardial Infarction; Diabetic Ketoacidosis; Multiple Organ Failure; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2025, AND THE CAUSE OF DEATH WAS HEART FAILURE AND DIABETIC KETOACIDOSIS. THE CUSTOMER REPORTED A BLOOD GLUCOSE VALUE OF 682 MG/DL. THE CUSTOMER REPORTED THE HYPERGLYCEMIA AND DIABETIC KETOACIDOSIS WAS TREATED WITH THE DISCONTINUE USE OF THE INSULIN DELIVERY SYSTEM AND HOSPITALIZATION, AND THE MYOCARDIAL INFARCTION AND ORGAN FAILURE WERE TREATED WITH HOSPITALIZATION. THE EVENT INVOLVED PRODUCTS MMT-399A, MMT-7040MA, MMT-7841ZN, MMT-1884, AND MMT-332A. TROUBLESHOOTING WAS PERFORMED AND FOUND THAT THE DOCUMENT OF ANY HEALTH ISSUES OR ILLNESS THAT MAY HAVE CONTRIBUTED OR LED UP TO PASSING WAS THE HEART ATTACK AND ORGAN FAILURE. THE CUSTOMER WAS STATED THAT THE PUMP WAS NOT WORN AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-399A. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040MA. NO PRODUCT RETURN IS REQUIRED FOR MMT-7841ZN. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22025582&pc=LOD | 3006946279-2025-00056 | 2025-04-01 04:00 | Death | BIOMET FRANCE S.A.R.L. | 2025-05-15 04:00 | LOD | OPTIPAC 80 REFOB BONE CMT R-3 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | EXEMPT |  | 1 | Event Description: A STUDY REPORTED A CLINICAL ADVERSE EVENT FOR DEATH, THE CAUSE OF WHICH IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE.Event Description: NO FURTHER EVENT INFORMATION AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: (B)(4). THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL AND/OR CORRECTED INFORMATION. THE FOLLOWING SECTIONS WERE UPDATED: B4, B5, G3, G6, H2, H3, H6, H11. NO PRODUCT WAS RETURNED OR PICTURES PROVIDED; A PRODUCT EVALUATION COULD NOT BE PERFORMED. A RETAINED SAMPLE OF THE BATCH HAS BEEN TESTED IN THE LABORATORY UNDER STANDARDIZED CONDITIONS. NO UNUSUAL CEMENT PASTE TEMPERATURE HAS BEEN NOTICED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS CONFIRMED NO ABNORMALITIES OR DEVIATIONS. A REVIEW OF THE RAW MATERIAL CERTIFICATE CONFIRMED NO ABNORMALITIES OR DEVIATIONS THAT COULD BE RELATED TO THE REPORTED EVENT. REVIEW OF THE COMPLAINT HISTORY FOUND NO ADDITIONAL RELATED COMPLAINTS FOR THIS ITEM AND THE REPORTED PART AND LOT COMBINATION. THERE IS INSUFFICIENT INFORMATION PROVIDED REGARDING THE CLINICAL SIGNS, HEALTH IMPACT, AND NATURE OF THE ADVERSE EVENT. AS A RESULT, IT IS NOT POSSIBLE TO DETERMINE WHETHER THE EVENT CONSTITUTES A SERIOUS INJURY. FURTHERMORE, THE ROLE OF CEMENT IN THE REPORTED INCIDENT IS UNCLEAR. SPECIFIC DETAILS ARE LACKING CONCERNING HOW, WHEN, AND UNDER WHAT CIRCUMSTANCES THE DEATH OCCURRED, MAKING IT DIFFICULT TO ASSESS WHETHER AND TO WHAT EXTENT THE CEMENT MAY HAVE CONTRIBUTED TO THE FATAL OUTCOME. THEREFORE, THE DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED. IF ANY FURTHER INFORMATION IS FOUND WHICH WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS. Manufacturer Narrative: (B)(4). MDR REPORT #: MW5168977. MDR REPORT KEY: (B)(4). D4: THIS DEVICE IS SOLD OUTSIDE THE US AND THEREFORE NO GUDID INFORMATION EXISTS. THIS DEVICE IS CONSIDERED SIMILAR TO (B)(4). G2 ¿ UNKNOWN WHICH COUNTRY THE EVENT OCCURRED IN. G4: THE REPORTED PRODUCT IS NOT SOLD IN THE US, THE PRE-MARKET SUBMISSION NUMBER FOR THE SIMILAR PRODUCT SOLD IN THE US IS K171540. THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MEDWATCH 3500A WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22025780&pc=MIH | 9612164-2025-02406 | 2025-05-13 04:00 | Death | MEDTRONIC IRELAND | 2025-05-15 04:00 | MIH | ENDURANT IIS BIFURCATED STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Rupture | P100021 |  | 1 | Event Description: AN ENDURANT IIS STENT GRAFT SYSTEM WAS IMPLANTED DURING THE EMERGENCY ENDOVASCULAR TREATMENT OF A RUPTURED ABDOMINAL AORTIC ANEURYSM. IT WAS REPORTED DURING THE INDEX PROCEDURE THAT A A SECONDARY RUPTURE OCCURRED, UNRELATED TO THE DEVICE, LEADING TO A DECLINE IN THE PATIENT'S VITAL SIGNS (SPECIFIC VALUES NOT PROVIDED). THE EVAR PROCEDURE WAS COMPLETED SUCCESSFULLY WITHOUT ENDOLEAK, USING THE ENDURANT IIS DEVICE IN COMBINATION WITH OTHER NON-MDT DEVICES. FOLLOWING EVAR, THE PATIENT UNDERWENT AN ABDOMINAL SURGERY TO RELIEVE INTRAPERITONEAL PRESSURE. DESPITE CLOSURE OF THE ABDOMEN AND ADMINISTRATION OF FLUIDS, THE PATIENT DIED IN THE INTENSIVE CARE UNIT (ICU) POSTOPERATIVELY. PER THE PHYSICIAN THE CAUSE OF DEATH WAS NOT DETERMINED, EXCESSIVE INTRAOPERATIVE BLEEDING IS SUSPECTED. THE DEATH WAS NOT ATTRIBUTED TO THE DEVICE. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME ENDURANT II ILIAC STENT GRAFT; PRODUCT ID ETLW1613C93EJ (SERIAL: (B)(6)); PRODUCT TYPE: 0180-AORTIC STENT GRAFT; IMPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22026261&pc=LWS | 9614453-2025-01939 | 2025-04-10 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-05-15 04:00 | LWS | COBALT¿ DR MRI SURESCAN¿ | Inappropriate/Inadequate Shock/Stimulation | Shock from Patient Lead(s) | P980016 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT¿S DAUGHTER THAT THE PATIENT HAD EXPIRED, AND THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS NOT RELATED TO THE PATIENT¿S DEATH. THE DAUGHTER STATED THE DEVICE WAS BEEPING AT THE PATIENT¿S FUNERAL SERVICE AND THE FUNERAL HOME USED A MAGNET ON THE DEVICE, BUT THE BEEPING CONTINUED PERIODICALLY. THE PATIENT THOUGHT THE DEVICE WAS PREVIOUSLY TURNED OFF AT THE HOSPITAL WHEN THE PATIENT WAS PRONOUNCED DECEASED, AND THE DEVICE SHOCKED THE PATIENT SEVERAL TIMES THE DAY BEFORE THE PATIENT'S DEATH WHICH WAS "UNNECESSARY." Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22026387&pc=MNS | 2518422-2025-105609 | 2025-04-23 04:00 | Death | RESPIRONICS, INC. | 2025-05-15 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF A DREAMSTATION AUTO CPAP. IT HAS BEEN REPORTED THE PATIENT HAS SADLY PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22027062&pc=DSQ | 2916596-2025-02975 | 2025-04-24 04:00 | Death | THORATEC CORPORATION | 2025-05-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. THE DEVICE WILL NOT BE EXPLANTED AND WILL OR BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM INSTRUCTIONS FOR USE IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22027579&pc=DSQ | 2916596-2025-02991 | 2025-04-18 04:00 | Death | THORATEC CORPORATION | 2025-05-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AT HOME ON HOSPICE. THE DEVICE OPERATED AS EXPECTED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. ADDITIONAL INFORMATION RECEIVED STATED THAT THE REASON FOR THE PATIENT BEING PLACED ON HOSPICE CARE WAS MULTIFACTORIAL. THE PATIENT HAD END STAGE HEART FAILURE, HISTORY OF STROKES WITH SIGNIFICANT DEFICITS, FAILURE TO THRIVE, AND PROSTATE CANCER WITH LONG TERM CATHETER. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM INSTRUCTIONS FOR USE IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION,¿ OUTLINES POTENTIAL ADVERSE EVENTS, INCLUDING STROKE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INR (INTERNATIONAL NORMALIZED RATIO) RANGE. ADDITIONALLY, SECTION 6 LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, AND OUTLINES INDICATIONS OF THROMBUS AS WELL AS HOW TO RESPOND TO SUCH EVENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION D4: LOT NUMBER WAS NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22027592&pc=OZD | 1220648-2025-28435 | 2025-04-10 04:00 | Death | ABIOMED, INC. | 2025-05-15 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Thrombocytopenia | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD A IMPELLA CP PATIENT TRANSFERRED TO THE TERTIARY CENTER FOR MORE CARE AND ESCALATION TO THE IMPELLA 5.5 PUMP. WEEKS LATER THE LOQI REPORTED UPON 2 ADVERSE EVENTS. THERE WAS THROMBOCYTOPENIA THAT REQUIRED TREATMENT WITH TRANSFUSION DUE TO PLATELET DROP. ADDITIONALLY THERE WAS ANEMIA/BLOOD LOSS AFTER THE SHEATH/PUMP EXCHANGE. THE PATIENT HAD MORE TRANSFUSION AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF THROMBOCYTOPENIA AND ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF ACCESS SITE BLEEDING WAS MOST LIKELY USE ISSUE SINCE THE PERCLOSE FAILED. THE ROOT CAUSE OF THROMBOCYTOPENIA CANNOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22027666&pc=OZD | 1220648-2025-28437 | 2025-04-24 04:00 | Death | ABIOMED, INC. | 2025-05-15 04:00 | OZD | IMPELLA 5.5 | Failure to Advance; Patient-Device Incompatibility | Arrhythmia; Stroke/CVA; Ischemia; Vascular Dissection; Thrombocytopenia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT EXPERIENCED MULTIPLE COMPLICATIONS DURING IMPELLA 5.5 SUPPORT. UPON THE INITIAL ATTEMPTED DELIVERY, IT WAS NOTED THAT THE J-WIRE AND GLIDE CATHETER WERE UNABLE TO ADVANCE PASSED THE INNOMINATE ARTERY. AN ANGIOGRAM REVEALED A DISSECTION WAS PRESENT AND THE DEVICES WERE REMOVED. A STERNOTOMY WAS PERFORMED AND A GRAFT TO THE SUBCLAVIAN ARTERY WAS PLACED TO BYPASS THE DISSECTION. IMPELLA IMPLANT WAS THEN SUCCESSFUL, HOWEVER, THE PATIENT LOST DISTAL FLOW DOWN THEIR RIGHT ARM FROM THE SUBCLAVIAN DOWN. SUPPORT CONTINUED DESPITE THE NOTED CONDITION AND THREE DAYS LATER, THE PATIENT SUFFERED A COUPLE RUNS OF VENTRICULAR TACHYCARDIA. CARDIOVERSION WAS REQUIRED TO TREAT THE CONDITION. THE FOLLOWING DAY, THE PATIENT WAS DETERMINED TO BE HIT POSITIVE AND HEPARIN WAS PAUSED DUE TO OOZING AND GASTROINTESTINAL BLEED CONCERNS. LATER THAT SAME DAY, THE PATIENT'S COMPUTED TOMOGRAPHY SCAN RESULTS SHOWED THAT THEY HAD SUFFERED AN EMBOLIC STROKE. THE FAMILY OPTED TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) - ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS - ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22027679&pc=DSQ | 2916596-2025-03068 | 2025-03-12 04:00 | Death | THORATEC CORPORATION | 2025-05-15 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE SYSTEMIC INFECTION. THE INFECTION WAS TREATED AS PER THE STANDARD INFECTION TREATMENT. THE OUTCOME WAS NOT CONSIDERED AS DEVICE THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22027754&pc=DXY | 2182208-2025-02379 | 2025-03-03 05:00 | Death | MEDTRONIC, INC. | 2025-05-15 04:00 | DXY | MEDTRONIC ICD | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE OUTCOMES WITH A SECOND-GENERATION INTERCOSTAL EXTRAVASCULAR LEAD PAIRED WITH STANDARD, TRANSVENOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) OR IMPLANTABLE PULSE GENERATOR (IPG). THE AUTHORS DESCRIBED ONE PATIENT DEATH DUE TO A STROKE WITHIN THIRTY DAYS OF THE INTERCOSTAL EXTRAVASCULAR LEAD REMOVAL PROCEDURE. THERE WAS ONE PATIENT WHO HAD THEIR SYSTEM REMOVED DUE TO INFECTION ON POSTOPERATIVE DAY SIX AND ANOTHER WITH SINUS NODE DYSFUNCTION AT THE INTERCOSTAL EXTRAVASCULAR LEAD REMOVAL VISIT. THERE WERE TWO DEVICES WHICH INAPPROPRIATELY DETECTED EPISODES DUE TO INCORRECT RHYTHM DISCRIMINATION. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/57 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: EVALUATION OF A SECOND-GENERATION INTERCOSTAL EXTRAVASCULAR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR LEAD WITH A P ECTORAL PULSE GENERATOR FOR SENSING, DEFIBRILLATION, AND ANTI-TACHYCARDIA PACING. EUROPACE. 2025. 27, EUAF044. DOI: 10.1093/EUROPACE/EUAF044 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22027763&pc=DXY | 2182208-2025-02380 | 2025-03-03 05:00 | Death | MEDTRONIC, INC. | 2025-05-15 04:00 | DXY | MEDTRONIC LEAD | Adverse Event Without Identified Device or Use Problem | Stroke/CVA |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE OUTCOMES WITH A SECOND-GENERATION INTERCOSTAL EXTRAVASCULAR LEAD PAIRED WITH STANDARD, TRANSVENOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) OR IMPLANTABLE PULSE GENERATOR (IPG). THE AUTHORS DESCRIBED ONE PATIENT DEATH DUE TO A STROKE WITHIN THIRTY DAYS OF THE INTERCOSTAL EXTRAVASCULAR LEAD REMOVAL PROCEDURE. THERE WAS ONE PATIENT WHO HAD THEIR SYSTEM REMOVED DUE TO INFECTION ON POSTOPERATIVE DAY SIX AND ANOTHER WITH SINUS NODE DYSFUNCTION AT THE INTERCOSTAL EXTRAVASCULAR LEAD REMOVAL VISIT. THERE WERE TWO DEVICES WHICH INAPPROPRIATELY DETECTED EPISODES DUE TO INCORRECT RHYTHM DISCRIMINATION. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/57 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: EVALUATION OF A SECOND-GENERATION INTERCOSTAL EXTRAVASCULAR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR LEAD WITH A PECTORAL PULSE GENERATOR FOR SENSING, DEFIBRILLATION, AND ANTI-TACHYCARDIA PACING. EUROPACE. 2025. 27, EUAF044. DOI: 10.1093/EUROPACE/EUAF044. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22027795&pc=DTZ | 3013876692-2025-00039 |  | Death | MAQUET CARDIOPULMONARY GMBH | 2025-05-15 04:00 | DTZ | HLS SET | Pressure Problem | Insufficient Information | K112360 |  | 1 | Event Description: IT WAS ORIGINALLY REPORTED THAT DURING PATIENT TREATMENT ISSUES WERE PRESENT WITH ERRONEOUS PRESSURE READINGS (ARTERIAL, INTERNAL AND DELTA) FROM AN HLS-SET. THE AFFECTED HLS SET WAS DISCARDED. ON 2025-04-16 THE INFORMATION WAS RECEIVED THAT THE FAMILY HAD DECIDED TO DISCONTINUE TREATMENT. HOWEVER, THERE IS NO INFORMATION ON DEATH OR SERIOUS INJURY. AS A PRESSURE READING ISSUE CAN LEAD TO A PUMP STOP, IF THE INTERVENTION IS SET BY THE USER. COMPLAINT ID #(B)(4). NEW INFORMATION WAS ON 05/14/2025 BY THE SSU/ CUSTOMER, THAT THE INVOLVED PATIENT PASSED AWAY. THE PATIENT¿S FAMILY DECIDED TO DISCONTINUE THE TREATMENT, RESULTING IN THE PATIENT PASSING AWAY.Event Description: COMPLAINT ID # (B)(4). Manufacturer Narrative: NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THEREFORE, IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER. Manufacturer Narrative: THE EVENT OCCURRED IN THE US. IT WAS REPORTED THAT DURING PATIENT TREATMENT ISSUES WERE PRESENT WITH ERRONEOUS PRESSURE READINGS (ARTERIAL, INTERNAL AND DELTA) FROM AN HLS-SET. NEW INFORMATION HAS BEEN RECEIVED ON 2025-05-14 BY THE SSU/CUSTOMER, THAT THE INVOLVED PATIENT PASSED AWAY. THE PATIENT¿S FAMILY DECIDED TO DISCONTINUE THE TREATMENT, RESULTING IN THE PATIENT PASSING AWAY. AS THE PATIENT PASSED AWAY, A REPORT IS REQUIRED. THE AFFECTED PRODUCT WAS NOT AVAILABLE FOR TECHNICAL INVESTIGATION AS IT WAS DISCARDED BY THE CUSTOMER. THUS, THE EXACT ROOT CAUSE REMAINS UNKNOWN. HOWEVER, A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-06-02 WITH FOLLOWING CONCLUSION: "THE REPORTED INCIDENT INVOLVES INACCURATE PRESSURE READINGS FROM A PRE-PRIMED HLS SET USED DURING ECMO THERAPY. THE CHIEF OF PERFUSION AT THE TREATING INSTITUTION, SUGGESTED DURING A CONFERENCE THAT FLUID MAY HAVE ENTERED THE WHITE CONNECTOR OF THE HLS SET PRIOR TO THERAPY, POSSIBLY CONTRIBUTING TO THE MALFUNCTION. THE TEAM REPORTEDLY ATTEMPTED TO CLEAN THE AREA BEFORE USE. ECMO THERAPY BEGAN ON (B)(6) 2025. NO ISSUES WERE NOTED DURING INITIAL PRIMING WITH PLASMALYTE, AND NO DEPOSITS OR CRYSTAL FORMATIONS WERE SEEN AT THE SENSOR PORT WHICH MAY BE IN CONFLICT WITH THE PREVIOUS SUGGESTION OF FLUID INTRUSION IN THE PORT THE ABSENCE OF VISIBLE DEPOSITS OR RESIDUE AT THE TIME OF SETUP MAY SUGGEST THAT INTRUSION MAY HAVE OCCURRED AFTER ECMO INITIATION AS INTRUSION WAS NOT IDENTIFIED DURING INITIAL DEVICE INSPECTION. WHILE THE CIRCUIT APPEARED CLEAN AT INITIATION, THE DISCOVERY OF BLOOD UPON DISASSEMBLY RAISES QUESTIONS ABOUT WHETHER CONTAMINATION MAY HAVE DEVELOPED SOMETIME DURING SUPPORT. THE REPORTED PRESSURE ISSUE WAS FIRST OBSERVED ON (B)(6), RESULTING IN THE REPLACEMENT OF THE SENSOR PROBE (CABLE) ON (B)(6). THE REPORTED BEHAVIOR SUGGESTS THE ROOT CAUSE MAY LIE IN EITHER THE CABLE ITSELF OR THE INTERNAL SENSOR PORT OF THE HLS SET. CLOTS WERE NOT INITIALLY OBSERVED BUT WERE LATER FOUND NEAR THE POST-OXYGENATOR MEMBRANE. LAB VALUES ON FEBRUARY 1 INDICATED STABLE PARAMETERS. THE HLS SET WAS NOT REPLACED, BUT THE SENSOR CABLE WAS EXCHANGED TWICE. ELEVATED PINT PRESSURES REPORTEDLY CORRELATED WITH PHYSICAL DISTURBANCES OF THE CABLE, SUCH AS DURING BLOOD SAMPLING. NO REDUCTION IN BLOOD FLOW WAS OBSERVED. CLOTS WERE NOT INITIALLY OBSERVED BUT WERE LATER IDENTIFIED NEAR THE POST-OXYGENATOR MEMBRANE AT THE END OF THERAPY. LABORATORY VALUES OBTAINED ON FEBRUARY 1 SHOWED A HEMOGLOBIN LEVEL OF 8.6 G/DL, HEMATOCRIT OF 28%, D-DIMER CONCENTRATION OF 12.21 MICROGRAM/ML, PLATELET COUNT OF 126 × 1 BILLION/L, AND A PTT OF 26.7 SECONDS. WHILE THESE VALUES APPEARED RELATIVELY STABLE FOR AN ECMO PATIENT, THE MARKEDLY ELEVATED D-DIMER LEVEL MAY SUGGEST EARLY THROMBOTIC ACTIVITY WITHIN THE CIRCUIT, DESPITE THE ABSENCE OF VISIBLE CLOTS AT THAT TIME (1, 2). ACCORDING TO RAJSIC ET AL. (2022) (3), THE RECOMMENDED THERAPEUTIC RANGE FOR THE PREVENTION OF VENOUS THROMBOEMBOLISM IS 1.5 TO 2.5 TIMES THE PATIENT¿S BASELINE APTT (40¿80 SECONDS). A PTT OF 26.7 SECONDS MAY HAVE CONTRIBUTED TO THROMBUS FORMATION WITHIN THE ECMO CIRCUIT. THE HLS SET ITSELF WAS NOT REPLACED DURING THE ECMO RUN, BUT THE SENSOR CABLE WAS EXCHANGED TWICE. ELEVATED PINT PRESSURES WERE OBSERVED IN CORRELATION WITH PHYSICAL DISTURBANCES OF THE CABLE, SUCH AS DURING PRE- AND POST-OXYGENATOR BLOOD SAMPLING. DESPITE THESE ANOMALIES, NO REDUCTION IN BLOOD FLOW WAS REPORTED, AND THE CIRCUIT CONTINUED TO PROVIDE EFFECTIVE SUPPORT. IN CONCLUSION, A PRODUCT MALFUNCTION INVOLVING THE INTERNAL PRESSURE SENSOR OF THE HLS SET WAS LIKELY, AS EVIDENCED BY PERSISTENT INACCURATE PRESSURE READINGS DESPITE CABLE REPLACEMENT. HOWEVER, THIS MALFUNCTION DID NOT RESULT IN A LOSS OF ECMO SUPPORT OR REDUCTION IN BLOOD FLOW. THAT SAID, REDUNDANT PRESSURE MONITORING DURING WAS, ASSUMEDLY, PRESENT VIA THE PART PRESSURE FROM THE HLS DISPOSABLE. NOTABLY, THE FIRST PRESSURE ANOMALY WAS REPORTED ON FEBRUARY 1¿NINE DAYS AFTER ECMO INITIATION ON (B)(6)¿INDICATING THAT THE HLS SET INITIALLY PERFORMED AS EXPECTED. THE SUBSEQUENT DISCOVERY OF BLOOD RESIDUES ON THE DEVICE UPON DISASSEMBLY SUGGESTS THAT CONTAMINATION OF THE INTERNAL PRESSURE SENSOR PORT MAY HAVE OCCURRED WHILE THERAPY, POTENTIALLY CONTRIBUTING TO THE OBSERVED BEHAVIOR. THE PATIENT¿S CLINICAL COURSE ULTIMATELY ENDED WITH WITHDRAWAL OF CARE AT THE FAMILY¿S REQUEST. WHILE THROMBOTIC EVENTS MAY HAVE CONTRIBUTED TO OVERALL MORBIDITY, THE DECISION TO WITHDRAW CARE APPEARS TO BE INFLUENCED BY BROADER CLINICAL OR ETHICAL CONSIDERATIONS. BASED ON AVAILABLE INFORMATION, THE PATIENT¿S EXPIRATION WAS NOT INFLUENCED BY THE DEVICE MALFUNCTION. THE FOLLOWING IS A QUOTE FROM THE INITIATING COMPLAINT INDICATING APPARENT SYNCHRONY WITH THE ABOVE OPINION ON DEVICE INFLUENCE: "ECLS COORDINATOR AT (B)(6) HOSPITAL, MENTIONED THEY ARE HAVING ISSUES WITH ERRONEOUS PRESSURE READINGS FROM AN HLS ON A PATIENT. IT IS NOT NEGATIVELY IMPACTING THE PATIENT OR LEVEL OF CARE, NO PATIENT HARM." ACCORDING TO THE INSTRUCTION FOR USE (HLS SET, CHAPTER "PREPARATION AND INSTALLATION") THE PRESSURE SENSORS HAVE TO BE CHECKED BEFORE PRIMING. IN GENERAL, IT SHOULD BE NOTED THAT PRE-PRIMING CAN HAVE A NEGATIVE IMPACT ON THE SENSOR FUNCTION AND THE ASSOCIATED PRESSURE VALUES AND CAN LEAD TO THE OUTPUT OF UNEXPECTED / IMPLAUSIBLE PRESSURE VALUES SHORTLY BEFORE USE, AS INDICATED IN CHAPTER "PRIMING THE SYSTEM", WHERE IT SAYS "ONLY FILL THE SYSTEM SHORTLY BEFORE USE AND IN ACCORDANCE WITH THE PRIMING INSTRUCTIONS. CALIBRATION VALUES ARE LOST WHEN THE DRIVE UNIT IS SWITCHED OFF". FURTHER THE CARDIOHELP HAS A FLOW/BUBBLE SENSOR AND A VENOUS PROBE TO MEASURE AND CONTROL THE BLOOD FLOW AND PARAMETERS. IF THE MEASURED VALUES ARE ABOVE HIGH LIMIT OR BELOW LOW LIMIT OF THE SET LIMITS THE SYSTEM GENERATES A VISUAL AND ACOUSTICAL ALARM. IN ADDITION, WITHIN THE CHAPTER ¿PRIMING THE SYSTEM¿ IT IS WRITTEN NOT TO PRE-PRIME THE CIRCUIT TOO EARLY. BASED ON THE RESULTS THE REPORTED FAILURE "PRESSURE ISSUES" COULD NOT BE CONFIRMED. THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON 2025-06-04. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARD TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED FAILURE. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THEREFORE, IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22027931&pc=OZD | 1220648-2025-28441 | 2025-01-28 05:00 | Death | ABIOMED, INC. | 2025-05-15 04:00 | OZD | IMPELLA 5.5 | Infusion or Flow Problem; Material Split, Cut or Torn | Arrhythmia; Bacterial Infection; Cardiac Arrest; Fungal Infection | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT EXPERIENCED CARDIAC ARREST DURING IMPELLA 5.5 SUPPORT. CARDIOPULMONARY RESUSCITATION WAS PERFORMED AND BOTH EPI AND SODIUM BICARB WERE GIVEN TO MANAGE THE CONDITION. AS SUPPORT CONTINUED, THE PATIENT DEVELOPED AN INFECTION OF UNKNOWN ORIGIN. BOTH BACTEREMIA AND CANDIDEMIA WERE PRESENT AT THE TIME OF THE INFECTION. THE PATIENT WAS GIVEN IV MEROPENEM, IV MICAFUNGIN, AND SERIAL BLOOD CULTURES TO MONITOR/TREAT THE INFECTION. IN ADDITION, THE CENTRAL LINES AND ARTERIAL LINES WERE REMOVED. DESPITE THE NOTED INTERVENTION, INFECTIONS OF VARIOUS DEGREES REMAINED INTERMITTENT THROUGHOUT SUPPORT. IT WAS ADDITIONALLY REPORTED THAT THE ANTI-CONTAMINATION SLEEVE WAS TORN. THE SLEEVE WAS SECURED VIA TAPE AND SUPPORT REMAINED UNINTERRUPTED. FORTY-SIX DAYS INTO SUPPORT THE PATIENT HAD RUN OF VENTRICULAR TACHYCARDIA (VT). THE PATIENT RECEIVED ONE SHOCK IN RESPONSE TO THE VT. A WEEK LATER, THE PATIENT AGAIN HAD A SHORT RUN OF VT TREATED WITH A SHOCK. ROUGHLY 108 DAYS INTO SUPPORT, A SPIKE IN MOTOR CURRENT WAS ENCOUNTERED AND SATURATED FLOWS WERE IDENTIFIED. IN RESPONSE TO THE MALFUNCTION, THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & CP WITH SMARTASSIST SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT, INFECTION/SEPSIS, PRODUCT DAMAGE (SLEEVE DAMAGE), AND SATURATED FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. PRODUCT DAMAGE (SLEEVE DAMAGE): DUE TO LIMITED CLINICAL DETAILS AND LACK OF PRODUCT RETURNED, THE ROOT CAUSE OF THE PRODUCT DAMAGE (SLEEVE DAMAGE) CANNOT BE DETERMINED. SATURATED FLOW: THERE WAS A SHIFT IN THE PS AS WELL AS THE PURGE PRESSURE. THE ROOT CAUSE OF THE SATURATED FLOW IS MOST LIKELY DUE TO INGESTED BIOMATERIAL BASED ON DATA LOG ANALYSIS AND THE CLINICAL DETAILS PROVIDED. THROMBUS: FAILURE MODE ADDED BASED ON DATA ANALYSIS. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE THROMBUS WAS NOT DETERMINED. INFECTION/SEPSIS: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE THROMBUS WAS NOT DETERMINED. VT/VF: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE THROMBUS WAS NOT DETERMINED. B2 REQUIRED INTERVENTION WAS INADVERTENTLY NOT SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28441.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22028030&pc=LWS | 2649622-2025-13285 | 2025-05-08 04:00 | Death | MPRI | 2025-05-15 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem |  | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS BEING EXPLANTED DUE TO A BACTEREMIA INFECTION WITH VEGETATION NOTED ON BOTH THE RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEADS. DURING THE LASER LEAD EXTRACTION, VEGETATION DISLODGED AND THE PATIENT DECOMPENSATED. THE PATIENT PASSED AWAY AFTER THE CHEST WAS CRACKED. Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: 559445 LEAD IMPLANTED: (B)(6) 2005. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22028060&pc=LWS | 3004209178-2025-08457 | 2025-05-08 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-15 04:00 | LWS | EVERA MRI XT DR SURESCAN | Adverse Event Without Identified Device or Use Problem |  | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS BEING EXPLANTED DUE TO A BACTEREMIA INFECTION WITH VEGETATION NOTED ON BOTH THE RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEADS. DURING THE LASER LEAD EXTRACTION, VEGETATION DISLODGED AND THE PATIENT DECOMPENSATED. THE PATIENT PASSED AWAY AFTER THE CHEST WAS CRACKED. Manufacturer Narrative: CONTINUATION OF D10: 559445 LEAD IMPLANTED: (B)(6) 2005. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22028067&pc=OZD | 1220648-2025-28442 | 2025-04-28 04:00 | Death | ABIOMED, INC. | 2025-05-15 04:00 | OZD | IMPELLA 5.5 | Malposition of Device | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH CARDIOMYOPATHY AND AN EJECTION FRACTION OF 20% WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED RUNS OF VENTRICULAR TACHYCARDIA AND AS WELL AS DEVICE MALPOSITIONING. THE PATIENT WOULD SUBSEQUENT EXPIRE NOTED TO BE INELIGIBLE FOR ANY ADVANCE THERAPIES. THE PATIENT, WITH MULTIPLE HOSPITALIZATIONS, WAS BEING EVALUATED FOR HEART FAILURE AND THE DECISION WAS MADE FOR AN IMPELLA 5.5 DEVICE PLACEMENT AND LEFT VENTRICULAR ASSIST DEVICE WORKUP. THE IMPELLA 5.5 DEVICE WAS IMPLANTED WITHOUT INCIDENT, AND THE PATIENT WAS TRANSFERRED TO THE UNIT ON IMPELLA MCS. TWO DAYS AFTER START OF THE MCS, THE MEDICAL TEAM BECAME AWARE OF PINK URINE POST BIVALIRUDIN INITIATION. THE NEED FOR ECHOCARDIOGRAM WAS DISCUSSED TO CONFIRM IMPELLA 5.5 DEVICE PLACEMENT, AND THE RESULTS REVEALED THE PUMP WAS SLIGHTLY TOWARDS MITRAL STRUCTURE BUT NOT INTERACTING. FURTHER NOTED WAS THE PATIENT EXPERIENCED RUNS OF VENTRICULAR TACHYCARDIA REQUIRING FIVE SHOCKS OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR. THE FOLLOWING DAY, THE PATIENT BECAME DELIRIOUS FROM SUSPECTED ETHYL ALCOHOL WITHDRAWAL. LABS AND HEMODYNAMICS WERE STABLE, AND THE PINK TINGED URINE SELF-RESOLVED THE DAY PRIOR. THE NEXT DAYS, THE PATIENT CONTINUED TO BE DELIRIOUS AND CONFUSED AND PULLED OUT THE SWAN CATHETER. ADDITIONALLY, THERE WAS OCCASIONAL SUCTION WHEN THE PATIENT WAS THRASHING IN THE BED. A NOSEBLEED WAS PRESENT AND BIVALIRUDIN WAS PLACED ON HOLD. A COMPUTED TOMOGRAPHY SCAN COMPLETED ONE WEEK PRIOR REVEALED NOTHING REMARKABLE. THE PSYCHIATRIST NOTED THE PATIENT MAY BE HAVING SOME TYPE OF PSYCHOSIS OR DELIRIUM. ANTICOAGULATION REMAINED ON HOLD DUE TO NOSE BLEEDING. THE PATIENT WAS NOTED AS STABLE AND THE PHYSICIAN WANTED ADDITIONAL TIME TO MANAGE. HOWEVER, THE FAMILY WITHDREW CARE DUE TO NO CHANGE IN THE PATIENT'S NEUROLOGICAL STATUS AND INELIGIBILITY FOR ADVANCED THERAPIES. THE PATIENT WOULD SUBSEQUENTLY EXPIRE AFTER A TOTAL OF APPROXIMATELY 16 DAYS ON IMPELLA MCS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VT/VF/AF/AT AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. LOW PUMP FLOW: STARTING THREE DAYS AFTER IMPLANTATION ON CONTINUING ON AND OFF THROUGHOUT CASE, SUCTION EVENTS OCCUR. BLOOD VOLUME WAS GIVEN FOR SOME BUT NOT ALL SUCTION REPORTS. THE PUMP WAS NOT RETURNED. DATA LOGS WERE RECOVERED AND REVIEW CONFIRMED SUCTION EVENTS OCCUR WITHOUT CASE. SUCTION EVENTS RESOLVE BY REDUCING P-LEVEL. THE CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY PATIENT CONDITION RELATED DUE TO SUCTION EVENTS OCCURRING AROUND RUNS OF LOW FLOW THAT SEEM TO RESOLVE THE PROBLEM ALONG WITH BLOOD VOLUME GIVEN. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22028070&pc=DTB | 2649622-2025-13288 | 2025-05-08 04:00 | Death | MPRI | 2025-05-15 04:00 | DTB | CAPSURE SP NOVUS LEAD | Adverse Event Without Identified Device or Use Problem |  | P850089 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS BEING EXPLANTED DUE TO A BACTEREMIA INFECTION WITH VEGETATION NOTED ON BOTH THE RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEADS. DURING THE LASER LEAD EXTRACTION, VEGETATION DISLODGED AND THE PATIENT DECOMPENSATED. THE PATIENT PASSED AWAY AFTER THE CHEST WAS CRACKED. Manufacturer Narrative: CONTINUATION OF D10: 6935M62, LEAD, IMPLANTED: (B)(6) 2018. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22028337&pc=NGV | 2124215-2025-31097 | 2025-03-18 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-15 04:00 | NGV | WATCHMAN FLX LEFT ATRIAL APPENDAGE CLOSURE DEVICE WITH DELIVERY SYSTEM | Failure to Seal | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: REPORTED VIA JOURNAL ARTICLE. IT WAS REPORTED THAT THE DEVICE DID NOT SEAL, AND DEATHS OCCURRED. THE PURPOSE OF THIS STUDY WAS TO ASSESS THE POSSIBLE CLINICAL IMPLICATIONS FOR A FABRIC LEAK AFTER WATCHMAN FLX DEVICE IMPLANTATION. THE RESULTS OF A 45-DAY COMPUTED TOMOGRAPHY (CT) SCAN ON PATIENT WITH A WATCHMAN FLX DEVICE SURVEILLANCE WERE REPORTED AND FOLLOWING 1 YEAR CT SCAN IN THOSE WHO HAD RESIDUAL LEFT ATRIAL APPENDAGE (LAA) PATENCY IN 105 CONSECUTIVE PATIENTS FEBRUARY 2021 TO MARCH 2022. A TOTAL OF 14 (13.3%) CASES OF DISTAL LAA PATENCY WERE REPORTED, WHICH WAS DEFINED AS "MACCHIATO SIGN" BUT IT IS NOW WELL KNOWN AS WATCHMAN FLX "FABRIC LEAK" (FL). IT IS DEFINED AS THE DISTAL OPACIFICATION OF THE LAA DESPITE THE ABSENCE OF VISIBLE COMMUNICATION BETWEEN THE LEFT ATRIUM AND LAA AND IS ATTRIBUTABLE TO A TILTED DEVICE WHICH RESULTS IN LEAKAGE WITH BLOOD FLOW INTO THE LAA BENEATH THE FABRIC MEMBRANE. THE CLINICAL SIGNIFICANCE OF FABRIC LEAK IS CURRENTLY UNKNOWN. THEREFORE, WE DECIDED TO INVESTIGATE AT THE LONGEST AVAILABLE FOLLOW UP THE RATE OF CLINICAL EVENTS FOR THE THREE GROUPS OF OUR STUDY (COMPLETELY SEALED [CS] [N = 70], COMPLETELY PATENT [CP] [N = 21] AND FABRIC LEAK (FL) [N = 14]). THE PRIMARY COMPOSITE ENDPOINT OF STROKE, SYSTEMIC EMBOLISM (SE) OR TRANSIENT ISCHEMIC ATTACK (TIA) ALL CAUSE DEATH AND NONPROCEDURAL MAJOR BLEEDING WERE REPORTED USING THE KAPLAN MEIER ESTIMATE CURVES TO PRESENT THE CLINICAL OUTCOMES OVER TIME, AND CURVES WERE COMPARED USING THE LOG RANK TEST. FOR THE PRIMARY COMPOSITE ENDPOINT, HRS WERE COMPUTED FOR THE COMPARISON BETWEEN GROUPS. P VALUES < 0.05 WERE CONSIDERED THE THRESHOLD FOR STATISTICAL SIGNIFICANCE. OVER A MEDIAN FOLLOW-UP OF 3.1 YEARS, THE PRIMARY COMPOSITE OF STROKE, SYSTEMIC EMBOLISM (SE), OR TRANSIENT ISCHEMIC ATTACK (TIA) WERE HIGHER IN THE FL AND CP GROUPS COMPARED TO THE CS GROUP. COMPLETELY SEALED (CS) GROUP (70 PATIENTS): STROKE, SE, OR TIA: 2.8% NON-PROCEDURAL MAJOR BLEEDING: 5.7% ALL-CAUSE DEATH: 7.2% COMPLETELY PATENT (CP) GROUP (21 PATIENTS): STROKE, SE, OR TIA: 14.28% NON-PROCEDURAL MAJOR BLEEDING: 4.76% ALL-CAUSE DEATH: 9.5% FABRIC LEAK (FL) GROUP (14 PATIENTS): STROKE, SE, OR TIA: 14.28% NON-PROCEDURAL MAJOR BLEEDING: 7.14% ALL-CAUSE DEATH: 14.3% OF NOTE, ONE (1) CASE OF STROKE OCCURRED 99 DAYS AFTER THE INDEX PROCEDURE IN A PATIENT WITH FL AT 45 DAYS AND SUBSEQUENT COMPLETE SEALING OF THE LAA AT 1 YEAR CT SCAN. Manufacturer Narrative: B3: ESTIMATED AS 03/18/2025 AS THE PUBLISHED DATE FOR THE ARTICLE IS NOTED AS 18 MARCH 2025. D4 MODEL NUMBER, D4 LOT NUMBER, D4 EXPIRATION DATE, D4 UNIQUE IDENTIFIER (UDI) # ARE UNKNOWN AS THE INFORMATION WAS NOT PROVIDED IN THE LITERATURE ARTICLE. LITERATURE CITATION: FRAZZETTO, M., YOON, S.H., & FILBY, S.J. (2025). CLINICAL IMPLICATIONS FOR FABRIC LEAK AFTER WATCHMAN FLX DEVICE IMPLANTATION: A STUDY UPDATE. CATHETERIZATION AND CARDIOVASCULAR INTERVENTIONS. HTTPS://DOI.ORG/10.1002/CCD.31494
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22028783&pc=PJF | 3003925919-2025-00001 | 2024-03-30 04:00 | Death | CERUS CORPORATION | 2025-05-15 04:00 | PJF | INTERCEPT BLOOD SYSTEM FOR PLATELETS | Activation, Positioning or Separation Problem | Hypoxia; Hypervolemia | BP140143 |  | 1 | Event Description: FATAL TRANSFUSION REACTION. REPORT TYPE: LITERATURE. INITIAL INFORMATION RECEIVED: 15-APR-2025, 17-APR-2025, 06-MAY-2025. INITIAL ADVERSE EVENT REPORT: ON 15-APR-2025, CERUS RECEIVED A SERIOUS ADVERSE EVENT REPORT OF A FATAL TRANSFUSION REACTION. CERUS' CHIEF SCIENTIFIC OFFICER DISCOVERED THIS ADVERSE EVENT WHILE REVIEWING THE INTERNATIONAL SOCIETY OF BLOOD TRANSFUSION (ISBT) 2025 ABSTRACTS. CERUS WAS REFERRED BY THE ABSTRACT AUTHOR, THE MEDSTAR GEORGETOWN UNIVERSITY HOSPITAL TRANSFUSION SERVICE MEDICAL DIRECTOR, TO THE BLOOD SUPPLIER. CERUS MADE CONTACT WITH MEDSTAR GEORGETOWN UNIVERSITY HOSPITAL'S BLOOD PROVIDER, AMERICAN RED CROSS (ARC), TO OBTAIN MORE INFORMATION. MEDICAL HISTORY AND MEDICATIONS: THIS REPORT INVOLVED A 65-YEAR-OLD MALE PATIENT WITH MEDICAL HISTORY OF PRIMARY MYELOFIBROSIS STATUS POST HEMATOPOIETIC STEM CELL TRANSPLANT AND HEMORRHAGIC CYSTITIS. THE PATIENT EXPERIENCED A FEBRILE NON-HEMOLYTIC TRANSFUSION REACTION ON 24-MAR-2024 AFTER RECEIVING 51 ML OF RED BLOOD CELLS. CONCOMITANT MEDICATIONS INCLUDED ACETAMINOPHEN, MAALOX, ATOVAQUONE, LOMOTIL, TUMS, CALCIUM GLUCONATE, NEXIUM, LIDOCAINE 2% TOPICAL, LOPERAMIDE, AND BUMEX. ADVERSE EVENT DESCRIPTION: ON 29-MAR-2024, THE PATIENT RECEIVED 3 UNITS OF PSORALEN TREATED PLATELETS (PTP) AND 1 UNIT OF IRRADIATED RBC AT MEDSTAR GEORGETOWN UNIVERSITY HOSPITAL, WASHINGTON D.C. ON 30-MAR-2024, HE RECEIVED ONE UNIT OF PTP STARTING AT 00:20H AND 1 UNIT OF IRRADIATED RBC STARTING AT 02:06H. SOMETIME AROUND THE EARLY MORNING OF 30-MAR-2024, PATIENT BECAME TACHYCARDIC (100-139 BMP). LASIX 40 MG IV WAS ADMINISTERED, AND IV FLUIDS WERE DISCONTINUED. CXR PERFORMED AT 09:59H SHOWED DIFFUSE INTERSTITIAL OPACITIES, WHICH MAY REFLECT PULMONARY EDEMA VS VIRAL INFECTION, A CONSISTENT FINDING SINCE 27-MAR-2024. OVER THE COURSE OF THE DAY, HE BECAME HYPOTENSIVE, TACHYPNEIC (RR 40) AND DEVELOPED ABUNDANT AND INTRACTABLE VOMITING WITH ALTERED MENTAL STATUS AT 11:18H. PATIENT WAS INTUBATED AND TRANSFERRED TO THE INTENSIVE CARE UNIT AND PRONOUNCED DEAD AT 12:01H. A REVIEW OF THE PATIENT'S TRANSFUSION HISTORY DURING THIS HOSPITALIZATION SHOWED A TOTAL OF 5 PTP AND 5 RBC BETWEEN 16-MAR-2024 AND 28-MAR-2024. MEDSTAR GEORGETOWN UNIVERSITY HOSPITAL TRANSFUSION SERVICE MEDICAL DIRECTOR SUSPECTED VOLUME OVERLOAD AND ELECTROLYTE ABNORMALITY AS THE POTENTIAL CAUSE OF THE ADVERSE EVENT ASSOCIATED WITH THE TRANSFUSIONS. TRANSFUSED BLOOD COMPONENTS WITHIN 24 HOURS OF THE EVENT OF FATAL TRANSFUSION REACTION PSORALEN TREATED PLATELETS #1: 29-MAR-2024 AT 08:59H - 252 ML (UNIT W185324801937002; PRODUCT CODE E8343V00), #2: 29-MAR-2024 AT 09:52H - 346 ML (UNIT W185324805528009; PRODUCT CODE E8342V00), #3: 29-MAR-2024 AT 21:41H - 331 ML (UNIT W185324808518002; PRODUCT CODE E8342V00), #4: 30-MAR-2024 AT 00:20H - 267 ML (UNIT W185324808535000; PRODUCT CODE E8341V00). RED BLOOD CELLS #1: 29-MAR-2025 AT 14:14H - 281 ML (UNIT W185324000875008; PRODUCT CODE E0686V00), #2: 30-MAR-2024 AT 02:06H - 290 ML (UNIT W185324026024008; PRODUCT CODE E0686V00). BLOOD DONOR(S) INFORMATION: THREE MALE AND ONE FEMALE DONOR. IMPLICATED PATHOGEN REDUCED BLOOD COMPONENT(S) COLLECTION AND MANUFACTURING INFORMATION: NO DETAILS PROVIDED. ADVERSE EVENT INVESTIGATIONS: MEDSTAR GEORGETOWN UNIVERSITY HOSPITAL INVESTIGATION: THE PATIENT'S FLUID BALANCE DURING THE 24 HOURS BEFORE HIS DEATH WAS 2,721 ML POSITIVE, WITH TRANSFUSIONS ACCOUNTING FOR 1,767 ML. THE PATIENT HAD NO KNOWN ALLERGY HISTORY AND HAD PREVIOUSLY TOLERATED 3 SINGLE DOSES OF PTP IN ONE DAY WITHOUT REACTION. ARC INVESTIGATION: FOUR PLATELET DONORS WERE IMPLICATED WITH THE PTP TRANSFUSED WITHIN 24 HOURS OF DEATH. ALL 4 DONORS WERE CONSIDERED LOW RISK FOR TRALI: 3 WERE MALES AND 1 WAS A FEMALE WITH NEGATIVE HLA/HNA ANTIBODIES. CERUS INVESTIGATION: AT THE TIME OF THIS REPORT, THE CERUS INVESTIGATION IS PENDING AND WILL BE UPDATED ONCE RECEIVED. ADVERSE EVENT ASSESSMENT: 1.~FATAL TRANSFUSION REACTION [PT: TRANSFUSION REACTION]: THE REPORTER ASSESSED THE EVENT AS DEATH IN SEVERITY, SERIOUS DUE TO BEING A DEATH EVENT, POSSIBLE IN RELATION TO PSORALEN TREATED PLATELETS. CAUSALITY TO INTERCEPT BLOOD SYSTEM FOR PLATELETS DEVICE WAS NOT REPORTED. REVIEW OF THE ARC INVESTIGATION INDICATED THERE WAS CONCERN FOR TRANSFUSION RELATED ACUTE LUNG INJURY (TRALI) AND/OR TRANSFUSION ASSOCIATED CIRCULATORY OVERLOAD (TACO), GIVEN THE PATIENT'S RESPIRATORY DECLINE, DIFFUSE INTERSTITIAL OPACITIES ON CXR, LARGE VOLUME OF TRANSFUSIONS DURING HOSPITALIZATION (9 UNITS OF PTP; 7 UNITS OF RBC) AND POSITIVE FLUID STATUS OF 2,721 ML WITHIN 24 HOURS OF DEATH. THE CERUS MEDICAL REVIEWER CONSIDERED THE EVENT AS SERIOUS DUE TO BEING A DEATH EVENT. EXCLUDED IN RELATION TO PSORALEN TREATED PLATELETS AND INTERCEPT BLOOD SYSTEM FOR PLATELETS, AND COMPATIBLE WITH TACO. CONFLICTING INFORMATION WAS PRESENTED IN THE ISBT ABSTRACT BY THE AUTHOR. IN THE ABSTRACT, THE AUTHOR STATED THAT CXR WAS NEGATIVE FOR EVIDENCE OF TACO OR TRALI AND THAT THE BLOOD SUPPLIER (I.E., ARC) CLASSIFIED THIS ADVERSE EVENT AS AN ALLERGIC REACTION. ALTHOUGH THE AUTHOR DID ACKNOWLEDGE THAT THE PATIENT'S REACTION MAY HAVE BEEN RELATED TO FLUID OVERLOAD, THE AUTHOR ALSO RAISED THE POSSIBILITY OF PSORALEN TOXICITY OR REACTION TO AN UNKNOWN ALLERGEN. GIVEN THE CLINICAL PRESENTATION THE MOST LIKELY CAUSE OF THIS PATIENT'S RESPIRATORY DISTRESS WAS TACO SINCE TRALI WAS RULED OUT WHEN ALL IMPLICATED PTP DONORS WERE DETERMINED TO BE LOW TRALI RISK. PSORALEN TOXICITY HAS NOT BEEN DESCRIBED IN NONCLINICAL STUDIES AT THE CONCENTRATIONS FOUND IN PSORALEN-TREATED PLATELETS, EVEN WITH MULTIPLE TRANSFUSION EVENTS WITHIN 24 HOURS. FURTHERMORE, CXR DEMONSTRATED PERSISTENCE OF DIFFUSE INTERSTITIAL OPACITIES, WHICH MAY REFLECT PULMONARY EDEMA VS VIRAL INFECTION, SINCE 27-MAR-2024 (3 DAYS PRIOR TO THE PATIENT'S DEATH). THE PATIENT'S VOLUME OVERLOAD WAS CONTRIBUTED BY THE TRANSFUSION OF PTP, RBC AND INFUSION AND IV FLUIDS, THUS, THE ADVERSE EVENT OF TACO CANNOT BE SOLELY ATTRIBUTED TO PTP. FINALLY, TACO IS A WELL-RECOGNIZED ADVERSE EVENT ASSOCIATED WITH ANY BLOOD TRANSFUSION BECAUSE THE UNDERLYING PATHOLOGY IS VOLUME OVERLOAD. LISTEDNESS ASSESSMENT IS NOT APPLICABLE BECAUSE THIS EVENT IS NOT RELATED TO THE PSORALEN TREATED PLATELETS OR INTERCEPT BLOOD SYSTEM FOR PLATELETS DEVICE.Event Description: CERUS RECEIVED ADDITIONAL FOLLOW UP INFORMATION ON THIS EVENT ON 21-MAY-2025. CERUS INVESTIGATION: (NEW INFORMATION) BATCH RECORD REVIEW. BATCH RECORD REVIEWS FOR THE FOLLOWING INTERCEPT BLOOD SYSTEM FOR PLATELETS PROCESSING SETS WERE PERFORMED: 1.~LOT # CE23J16L71; EXPIRATION DATE: 26-MAY-2024. A.~THE REVIEW OF THE BATCH RECORDS FOR THE FINISHED GOODS DID NOT INDICATE ANY OUT-OF-LIMITS RESULTS DURING TESTING AND CONTROLS. THE PRODUCT MET ALL QUALITY REQUIREMENTS AND SPECIFICATIONS AT THE TIME OF RELEASE. THIS BATCH WAS RELEASED ON 29-DEC-2023. NO OTHER COMPLAINTS HAVE BEEN RECEIVED FOR THIS LOT CE23J16L71. THE TOTAL BATCH SIZE WAS (B)(4) PROCESSING SETS. 2.~LOT # CE23J18L72; EXPIRATION DATE: 25-MAY-2024. A.~THE REVIEW OF THE BATCH RECORDS FOR THE FINISHED GOODS DID NOT INDICATE ANY OUT-OF-LIMITS RESULTS DURING TESTING AND CONTROLS. THE PRODUCT MET ALL QUALITY REQUIREMENTS AND SPECIFICATIONS AT THE TIME OF RELEASE. THIS BATCH WAS RELEASED ON 26-DEC-2023. THREE LEAK COMPLAINTS WERE FOUND ASSOCIATED WITH LOT CE23J18L72 FROM ARC CONSIGNEES. ALL THREE LEAKS OCCURRED DURING STORAGE AND WERE DISCARDED. THE CAUSE OF THE LEAKS WAS ATTRIBUTED TO HANDLING. THE TOTAL BATCH SIZE WAS (B)(4) PROCESSING SETS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22028868&pc=OZD | 1220648-2025-28446 | 2025-04-24 04:00 | Death | ABIOMED, INC. | 2025-05-15 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THERE WERE CONCERNS FOR THE PATIENT'S HEMOLYSIS. ADDITIONALLY, THE PATIENT HAD OOZING AT THE ACCESS SITE WITH A HEMATOMA. MANUAL PRESSURE WAS HELD AND IT IS UNKNOWN IF BLOOD PRODUCTS WERE GIVEN. REPORTABLE DUE TO PATIENT'S POSSIBLE HEMOLYSIS AND ACCESS SITE ISSUES. FURTHEMORE, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22028945&pc=NVZ | 2017865-2025-53821 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-15 04:00 | NVZ | ELLIPSE VR | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P910023 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S WIFE THAT THE PATIENT DIED OF A CARDIAC ARREST. THE WIFE WAS CONCERNED THE PATIENT'S TRANSMITTER ALLEGEDLY DID NOT ALERT OF A CARDIAC ARREST AND THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS DEFECTIVE. INFORMATION RECEIVED SUGGESTS THE PATIENT EXPERIENCED A FALL BACKWARDS AS HE GOT UP ON (B)(6) 2025. 911 WAS CALLED AND THE PATIENT WAS TRANSPORTED TO THE EMERGENCY ROOM AND ADMITTED TO THE INTENSIVE CARE UNIT (ICU) WHERE HE PASSED AWAY (B)(6) 2025. INFORMATION RECEIVED FROM THE FOLLOWING CLINIC INDICATES THE PATIENT'S LAST IN PERSON FOLLOW UP WITH THEIR PHYSICIAN WAS (B)(6) 2024, AND SYSTEM INTERROGATION WAS NORMAL. THE PHYSICIAN TRAVELLED OUT OF TOWN FOR THE VISIT. THE PATIENT WAS DUE FOR A YEARLY IN PERSON VISIT (B)(6) 2025 BUT DID NOT PRESENT FOR THE APPOINTMENT. THE PATIENT WAS ALSO BEING MONITORED REMOTELY AT THE FACILITY'S REMOTE MONITORING CLINIC. THE LAST REMOTE TRANSMISSION RECEIVED AND CHECKED BY THE PHYSICIAN FROM THE REMOTE MONITORING CLINIC WAS DATED (B)(6) 2024. THE PATIENT¿S PRESENTING RHYTHM WAS VENTRICULAR FIBRILLATION WITH VENTRICULAR TACHYCARDIA (VF/VT), BATTERY LONGEVITY 6.5 YRS, THE DEVICE/SYSTEM FUNCTION WAS STABLE PER PHYSICIAN NOTES IN (B)(6) 2024, AND THE NEXT REMOTE TRANSMISSION FOLLOW UP WAS SCHEDULED TO BE 91 DAYS FROM THIS REVIEW. NO OTHER TRANSMISSION WAS RECEIVED FOR REVIEW. THE DEVICE REMAINS IMPLANTED AND WILL NOT BE RETURNED. THE PATIENT WAS NOTED TO HAVE DIABETES AS WELL AS OTHER CARDIAC ISSUES UNRELATED TO THE SYSTEM FOR WHICH THEY WERE RECEIVING TREATMENT FROM OTHER CARDIOLOGISTS AS WELL AS OTHER CO-MORBIDITIES. THE CAUSE OF DEATH LISTED ON THE DEATH CERTIFICATE WAS (1) CARDIAC ARREST (2) ENCEPHALOPATHY. THE MANNER OF DEATH IS RECORDED AS NATURAL CAUSES. THE PATIENT'S RELATIVE DISPUTES THE CAUSE OF DEATH. Manufacturer Narrative: FOLLOWING PHYSICIAN: DR. (B)(6). HOSPITAL WHERE PATIENT DIED: (B)(6) MEDICAL CENTER, (B)(6). Manufacturer Narrative: THE PROVIDED SESSION RECORD FROM MARCH 4TH WAS REVIEWED BY TECHNICAL SERVICES AND IT WAS OBSERVED THAT THE DEVICE WAS PERFORMING PER PROGRAMMED SETTINGS. SESSION RECORDS AROUND THE TIME OF PATIENT DEATH WERE UNABLE TO BE OBTAINED. THE DEVICE HISTORY RECORD REVIEW WAS PERFORMED, AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22029043&pc=DSQ | 2916596-2025-03250 | 2025-05-03 04:00 | Death | THORATEC CORPORATION | 2025-05-15 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY SECONDARY TO AN UNKNOWN CAUSE, NATURAL DEATH. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE WAS WORKING AS EXPECTED FOR THE WHOLE TIME OF SUPPORT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6) , AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22029357&pc=NPT | 2015691-2025-03861 | 2025-04-23 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-15 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Rupture | P140031 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN ITALY. AS REPORTED, THIS WAS A CASE OF AN IMPLANT OF A 26MM SAPIEN 3 ULTRA IN AORTIC POSITION BY TRANSFEMORAL APPROACH. A BALLOON VALVULOPLASTY WAS PERFORMED USING A 22 MM BALLOON AND THEN THE VALVE WAS IMPLANTED. AFTER THE DEPLOYMENT, THE PATIENT WENT INTO CARDIAC ARREST. AFTER MULTIPLE ATTEMPTS, THE TEAM DECIDED TO CONVERT THE PROCEDURE TO SURGERY. AN ANNULAR RUPTURE WAS OBSERVED AND IT WAS DECIDED TO EXPLANT THE VALVE AND IMPLANT A 21 MM EDWARDS INSPIRIS SURGICAL VALVE. EVEN SO THE PATIENT WAS NOT RESPONDING AND PASSED AWAY. AS PER MEDICAL OPINION, THE ROOT CAUSE OF THE ANNULAR RUPTURE WAS SEVERE CALCIFICATION. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THAT PATIENT RELATED FACTORS (TYPE 1 BICUSPID VALVE WITH R-L RAPHE AND SEVERE CALCIFICATIONS) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22029543&pc=DSQ | 2916596-2025-02896 | 2025-05-15 04:00 | Death | THORATEC CORPORATION | 2025-05-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Hypoxia; Tachycardia; Respiratory Failure; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER.Event Description: THE PATIENT WAS ADMITTED FROM BIR TO 10R WITH PNEUMONIA AND A RIGHT INGUINAL HERNIA (REDUCED). ON (B)(6) 2024, THE PATIENT HAD RECURRENT VENTRICULAR TACHYCARDIA AND PULSELESS ELECTRICAL ACTIVITY ARREST SECONDARY TO HYPOXIA VS BRADYCARDIA. THEY WERE TRANSFERRED TO 4R INTENSIVE CARE UNIT (ICU) AND THEY WERE INTUBATED. THEY WERE EXTUBATED ON (B)(6) 2025 WITH LEFT VENTRICULAR ASSIST DEVICE (LVAD) SPEED BACK TO 5200 REVOLUTIONS PER MINUTE (RPM). ON (B)(6) 2024, THE PATIENT HAD INCREASED WORK OF BREATHING ON HIGH FREQUENCY OSCILLATION AND BIPAP. THE PATIENT WAS BRIEFLY HYPOTENSIVE AND BRADYCARDIC AROUND 18:30 WHICH SELF RESOLVED. THEY WERE ON BIPAP THEN. AROUND 20:30, THE LVAD HAD LOW FLOW ALARMS WITH FLOWS AT 1.5 LITERS PER MINUTE (LPM). THEIR MEAN ARTERIAL PRESSURES (MAP) WERE 30-40 AND THEY WERE LETHARGIC. EPINEPHRINE, NOREPINEPHRINE, AND VASO WAS STARTED. THEY WERE LATER INTUBATED. THEY CODED 3 TIMES AND THEN THE LVAD STARTED FLOWING AGAIN. PULSELESS ELECTRICAL ACTIVITY WITH NO VENTRICULAR TACHYCARDIA/VENTRICULAR FIBRILLATION. THIS WAS THOUGHT TO BE DUE TO A RESPIRATORY CAUSE WITH ISTAT SHOWING HYPERCAPNIA AND INITIALLY HYPOXIA WHICH IMPROVED WITH INTUBATION AND VENTILATION. THE PATIENT GOT BRADYCARDIC TO 30S, ALTHOUGH THEY HAD AN ICD (ST. JUDE), BUT IT WAS NOT PACING AT ALL. IT WAS DECIDED TO REPLACE SWAN WITH PACING SWAN DUE TO STILL BEING BRADYCARDIC AFTER HYPOXIA IMPROVED. THE PACING SWAN WAS PLACED AND CAPTURING WELL. BRONCHOSCOPE WAS DONE AND FOUND SIGNIFICANT THICK SECRETION. LATER DURING THE NIGHT, THE PATIENT CODED AGAIN, WITH LOW FLOWS AND MAPS AT 30S TO UNDETECTABLE. CARDIOPULMONARY RESUSCITATION (CPR) WAS CONTINUED FOR A FEW ROUNDS, AND THEY INTERMITTENTLY REGAINED FLOW. BEDSIDE POINT-OF-CARE ULTRASOUND SHOWED RIGHT VENTRICLE DILATION (PRE-EXISTING) WITH RIGHT VENTRICLE DYSFUNCTION. THERE WAS NO PERICARDIAL EFFUSION. THE PATIENT'S ABDOMEN WAS MORE DISTENDED. ULTRASOUND FOUND ASCITES RULED OUT HEMOPERITONEUM. HEMATOCRIT FROM ISTAT WAS 15 (FROM 23-24 PREVIOUSLY). IT WAS DISCUSSED TO POTENTIALLY USE VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO)/ RIGHT VENTRICLE SUPPORT, BUT THEY ALSO ADDRESSED ACTIVE BLEEDING AT THE TIME. NEEDED FOR EX-LAP GIVEN HEMOPERITONEUM. THE PATIENT'S SPOUSE ASKED FOR COMFORT CARE. CODE STATUS CHANGED TO DNAR-AND. COMFORT MEDICATION WAS GIVEN. LVAD ON EXTENDED ALARM SILENCE AND THEN OFF. PATIENT PASSED AT 01.12AM. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED, AND IT OPERATED AS EXPECTED. NO DEVICES WOULD RETURN. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE REPORTED LOW FLOW ALARMS COULD NOT BE CONFIRMED AS NO LOG FILES WERE PROVIDED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS ADVERSE EVENTS, INCLUDING ARRHYTHMIA, RIGHT HEART FAILURE, RESPIRATORY FAILURE, BLEEDING, AND DEATH, WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THE IFU ALSO LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THIS DOCUMENT ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 1 ALSO PROVIDES AN EXPLANATION OF EACH OF THE PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4, ¿SYSTEM MONITOR,¿ EXPLAINS THAT THE LOW FLOW HAZARD ALARM OCCURS WHEN PUMP FLOW IS LESS THAN 2.5 LPM. A 10-SECOND DELAY IS IMPOSED BETWEEN THE DETECTION OF THE LOW FLOW STATUS AND THE ACTIVATION OF THE ASSOCIATED AUDIO AND VISUAL INDICATORS ON THE SYSTEM CONTROLLER. CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7, ¿ALARMS AND TROUBLESHOOTING,¿ EXPLAINS SYSTEM ALARMS AND THE RECOMMENDED ACTIONS ASSOCIATED WITH THEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: OTHER ANNEX E CODES: HEART E06: ARRHYTHMIA E0601: BRADYCARDIA E060104. MENTAL, EMOTIONAL, AND BEHAVIOURAL DISORDERS E02: LETHARGY E0204. METABOLISM AND NUTRITION E12: ELECTROLYTE IMBALANCE E1202: BLOOD BICARBONATE INCREASED E120207. HEART E06: HEART FAILURE/CONGESTIVE HEART FAILURE E0611. GASTROINTESTINAL SYSTEM E10: ABDOMINAL DISTENTION E1001. GASTROINTESTINAL SYSTEM E10: ASCITES E1004. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22029544&pc=DSQ | 2916596-2025-02886 | 2024-11-02 04:00 | Death | THORATEC CORPORATION | 2025-05-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Tachycardia; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FROM VENTRICULAR TACHYCARDIA STORM AND RIGHT HEART FAILURE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED OR RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN DETAILS REGARDING THE ONSET OF THE VENTRICULAR TACHYCARDIA AND RIGHT HEART FAILURE; HOWEVER, NO RESPONSE WAS RECEIVED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS (INCLUDING DEATH, CARDIAC ARRYTHMIA, AND RIGHT HEART FAILURE) THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. SECTION 6, UNDER ¿RIGHT HEART FAILURE¿, STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22029604&pc=OZD | 1220648-2025-28455 | 2025-04-25 04:00 | Death | ABIOMED, INC | 2025-05-15 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP SUFFERED FROM VENTRICULAR TACHYCARDIA. DESPITE BEING DEFIBRILLATED, THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22029829&pc=DSQ | 2916596-2025-02920 | 2024-12-01 05:00 | Death | THORATEC CORPORATION | 2025-05-15 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemorrhage/Bleeding; Unspecified Infection; Ischemia; Multiple Organ Failure; Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE ARTICLE IMPACT OF LEFT VENTRICULAR SIZE ON OUTCOMES FOLLOWING HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION: ANALYSIS OF THE EUROPEAN REGISTRY FOR PATIENTS WITH MECHANICAL CIRCULATORY SUPPORT (EUROMACS) SOUGHT TO ASSESS PREIMPLANTATION RISK FACTORS THAT INFLUENCE POST- LEFT VENTRICULAR ASSIST DEVICE (LVAD) OUTCOMES IN PATIENTS WITH SMALLER LEFT VENTRICULAR (LV) SIZE. THE STUDY PERFORMED AN OBSERVATIONAL, RETROSPECTIVE ANALYSIS OF 834 ADULT (=18 YEARS) PATIENTS WITH CARDIAC FROM THE EUROMACS REGISTRY WHO UNDERWENT A CONTINUOUS-FLOW, LONG-TERM CENTRIFUGAL HEARTMATE 3 (HM3) LVAD IMPLANTATION BETWEEN AUGUST 2014 AND FEBRUARY 2023. PATIENTS ON ISOLATED LVAD SUPPORT WITH AN HM3 DEVICE WITH A MEDIAN FOLLOW-UP OF 605 DAYS (IQR: 239¿1,210) WERE INCLUDED, ADULTS WITH ISOLATED HM3 WITHOUT CONCOMITANT VALVE SURGERY WITH DATA ON PREOPERATIVE LEFT VENTRICULAR END-DIASTOLIC DIAMETER (LVEDD) AND AN AVAILABLE FOLLOW-UP WERE INCLUDED AND PATIENTS WITH PRIOR STERNOTOMIES AND TEMPORARY MECHANICAL CIRCULATORY SUPPORT BEFORE LVAD IMPLANTATION WERE EXCLUDED. PATIENT CHARACTERISTICS: THREE LVEDD GROUPS COMPRISED THE INVESTIGATIVE COHORT; 1) LESS THAN OR EQUAL TO 65 (N = 251), 2) 65¿80 (N = 441), AND 3) GREATER THAN OR EQUAL TO 80 MM (N = 142) HAD FOLLOW-UP DURATIONS OF 534 (IQR: 211¿1,198), 663 (IQR: 232¿1,210), AND 673 DAYS (IQR: 320¿1,214), RESPECTIVELY. MEAN AGE OF THE STUDY COHORT WAS 54.4 ± 12.4 YEARS, WITH A MEAN BSA OF 2.0 ± 0.23 M2 AND A MEAN BODY MASS INDEX (BMI) OF 27.3 ± 5.0 KG/M2; 87.9% WERE MALE, 40% HAD ISCHEMIC CARDIOMYOPATHY AND 43.8% HAD DILATED CARDIOMYOPATHY. THE GROUP WITH AN LVEDD LESS THAN OR EQUAL TO 65 MM WAS SIGNIFICANTLY OLDER (56.1 ± 11.0 VS. 53.8 ± 13.0 VS. 53.0 ± 12.3 YEARS; P = 0.015), HAD A HIGHER PREVALENCE OF FEMALES (17.9% VS. 11.6% VS. 4.9%; P < 0.001) HAD A SMALLER BODY SURFACE AREA (BSA) (1.95 ± 2.03 VS. 2.01 ± 0.22 VS. 2.08 ± 0.22; P < 0.01) AND A LOWER BMI (26.6 ± 4.8 VS. 27.4 ± 5.1 VS. 28.2 ± 5.0; P = 0.01), COMPARED WITH THE GROUPS WITH AN LVEDD OF 65¿80 AND GREATER THAN OR EQUAL TO 80 MM, RESPECTIVELY. SURVIVAL AND ADVERSE EVENTS: PATIENTS WITH AN LVEDD LESS THAN OR EQUAL TO 65 MM EXHIBITED SIGNIFICANTLY LOWER IN-HOSPITAL SURVIVAL (88.1% VS. 94.2% VS. 96.4%; P = 0.003). THE HIGHEST SURVIVAL IN THE GROUP WITH AN LVEDD GREATER THAN OR EQUAL TO 80 MM (75.7% VS. 84.6% VS. 89.5%; P = 0.003. HOWEVER, MULTIVARIABLE ANALYSIS OF LVEDD AS A CONTINUOUS PARAMETER SHOWED AN APPROXIMATELY 1% INCREASE IN 1 YEAR SURVIVAL PER 1 MM OF PREOPERATIVE LVEDD (OR: 1.02, 95% CI: [1.00¿1.04], P = 0.039). PATIENTS WITH AN LVEDD LESS THAN OR EQUAL TO 65 MM HAD A SIGNIFICANTLY HIGHER RISK OF POSTOPERATIVE RIGHT HEART FAILURE (RHF) (RR: 1.54, 95% CI: [1.00¿2.36], P = 0.049) BUT A LOWER RISK FOR CARDIAC ARRHYTHMIA INCLUDING SUSTAINED VENTRICULAR TACHYCARDIA (RR: 0.64, 95% CI: [0.45¿0.90], P = 0.012 AND RR: 0.50, 95% CI: [0.27¿0.89], P = 0.024), RESPECTIVELY (THE MIDDLE COHORT SERVES AS REFERENCE, WITH NO SIGNIFICANT DIFFERENCE BETWEEN THE LARGER COHORTS). THE RISK FOR OTHER POSTOPERATIVE COMPLICATIONS IS REPORTED IN TABLE 1 OF THE ATTACHED ARTICLE AND INCLUDES: CARDIAC ARRHYTHMIA, SUSTAINED VT, RIGHT HEART FAILURE, INFECTION, DEVICE THROMBOSIS, DEVICE MALFUNCTION, STROKE, GI BLEEDING. A FULL LIST OF PRIMARY CAUSES OF DEATH BETWEEN THE MATCHED COHORTS IS REPORTED IN TABLE 2 OF THE ATTACHED ARTICLE AND INCLUDES: BLEEDING, CANCER, CARDIOPULMONARY FAILURE, CARDIOVASCULAR, CEREBROVASCULAR ACCIDENT, INFECTION, INTESTINAL ISCHEMIA, LUNG FAILURE, MULTIORGAN FAILURE, OTHER CAUSE OF DEATH, RIGHT HEART FAILURE, SEPSIS, STROKE: INTRACEREBRAL BLEEDING, SUDDEN UNEXPLAINED DEATH, SUICIDE AND UNKNOWN CAUSE OF DEATH. CONCLUSION: DESPITE SIGNIFICANT ADVANCEMENTS IN LVAD THERAPY, PATIENTS WITH HF WITH A SMALLER LV SIZE MIGHT SUSTAIN INFERIOR SURVIVAL AND CONTINUE TO DEMONSTRATE AN INCREASED RISK OF POSTOPERATIVE RHF COMPARED TO THOSE WITH LARGER LVS. IN THIS REGARD, THERE REMAINS AN UNMET NEED FOR TECHNOLOGICAL ADVANCEMENTS TO IMPROVE SURVIVAL WITH LVAD IN PATIENTS WITH SMALL LV DIMENSIONS. BEYOND THIS, SURGICAL IMPLANTATION TECHNIQUE, PUMP FLOW OPTIMIZATION, AND POSTOPERATIVE PATIENT CARE NEED TO BE CAREFULLY EVALUATED IN PATIENTS WITH A SMALL LVEDD SCHEDULED FOR LVAD IMPLANTATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, PUMP POCKET), SEPSIS, STROKE, MULTI ORGAN FAILURE, RIGHT HEART FAILURE, RESPIRATORY FAILURE, BLEEDING, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, "PATIENT CARE AND MANAGEMENT" (UNDER "RIGHT HEART FAILURE"), DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THIS SECTION ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. NERSESIAN, G., JAISWAL, A., STEIN, J., PILLAI, A., DE BY, T. M., POTAPOV, E., & BARAN, D. A. (2024). IMPACT OF LEFT VENTRICULAR SIZE ON OUTCOMES FOLLOWING HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION: ANALYSIS OF THE EUROPEAN REGISTRY FOR PATIENTS WITH MECHANICAL CIRCULATORY SUPPORT (EUROMACS). ASAIO JOURNAL. HTTPS://DOI.ORG/10.1097/MAT.0000000000002370. DEPARTMENT OF CARDIOTHORACIC AND VASCULAR SURGERY, DEUTSCHES HERZZENTRUM DER CHARITÉ, BERLIN, GERMANY; ¿DEUTSCHES HERZZENTRUM DER CHARITÉ (GERMAN CENTRE FOR CARDIOVASCULAR RESEARCH), PARTNER SITE BERLIN, BERLIN, GERMANY; ¿DEPARTMENT OF CARDIOLOGY, HARTFORD HEALTHCARE HEART AND VASCULAR INSTITUTE, HARTFORD, CONNECTICUT; §THE EUROPEAN REGISTRY FOR PATIENTS WITH MECHANICAL CIRCULATORY SUPPORT, THE EUROPEAN ASSOCIATION FOR CARDIO THORACIC SURGERY, WINDSOR, UNITED KINGDOM; AND ¶CLEVELAND CLINIC FLORIDA, HEART, VASCULAR AND THORACIC INSTITUTE, ADVANCED HEART FAILURE PROGRAM, WESTON, FLORIDA. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22029935&pc=DSQ | 2916596-2025-03003 | 2025-04-26 04:00 | Death | THORATEC CORPORATION | 2025-05-15 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Tachycardia; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE PATIENT'S CAUSE OF DEATH WAS MULTISYSTEM ORGAN FAILURE. AFTER IMPLANT, THE PATIENT LEFT THE OPERATING ROOM WITH THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) AND CENTRIMAG RIGHT VENTRICULAR ASSIST DEVICE (RVAD). ON (B)(6) 2025, THE PATIENT RETURNED TO THE OPERATING ROOM FOR WEANING THE RVAD. THEY DID NOT TOLERATE RVAD WEAN AND WERE PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SUPPORT WITH REMOVAL OF THE RVAD. HE STARTED ON CONTINUOUS VENO-VENOUS HEMODIALYSIS (CVVHD) ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT UNDERWENT LOWER EXTREMITY FASCIOTOMY (RLE) ARTERIOGRAM WITH THROMBECTOMY AND STENTING OF RIGHT SUPERFICIAL FEMORAL ARTERY (R SFA) WITH DEVELOPMENT OF RLE COMPARTMENT SYNDROME REQUIRING BEDSIDE RLE FASCIOTOMY. DURING THEIR STAY, THE PATIENT EXPERIENCED MULTIPLE EPISODES OF VENTRICULAR TACHYCARDIA (VT), REQUIRING NUMEROUS SHOCKS. ON (B)(6) 2025, THE FAMILY REQUESTED TO WITHDRAW CARE, INITIATING COMFORT CARE, AFTER A FAMILY MEETING. PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTI ORGAN FAILURE, THROMBOEMBOLISM, CARDIAC ARRHYTHMIA, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AND THROMBOEMBOLISM AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. THIS SECTION, UNDER ¿ANTICOAGULATION¿, ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INR RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22029957&pc=LWS | 2649622-2025-13339 | 2025-02-20 05:00 | Death | MPRI | 2025-05-15 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Under-Sensing | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE ARRYTHMIA WAS BELOW THE LOWEST DETECTION ZONE AND THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING OF THE VENTRICULAR FIBRILLATION (VF) THAT WAS REPORTED TO HAVE CONTRIBUTED TO THE FAILED RESCUE SHOCKS AND THE PATIENT DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22030136&pc=OZD | 1220648-2025-28460 | 2025-04-25 04:00 | Death | ABIOMED, INC | 2025-05-15 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation; Embolism/Embolus | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA 5.5 BEGAN COMPLAINING OF CHEST PAIN AND PROFOUND SHORTNESS OF BREATH, THEN WENT INTO VENTRICULAR FIBRILLATION. UNSUCCESSFUL ATTEMPTS WERE MADE TO CONVERT THE SINUS RHYTHM AND CPR WAS INITIATED. THE DECISION WAS EVENTUALLY MADE TO HALT RESUSCITATION EFFORTS AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST SYSTEM LISTS VENTRICULAR ARRYTHMIA AS AN ADVERSE EVENT THAT OCCURRED WHILE IN USE WITH THE IMPELLA DEVICE. IT STATES THAT "ONLY A FRACTION OF THESE RATES WERE ATTRIBUTED TO THE IMPELLA DEVICE AND THE EVENTS RESOLVED WITHOUT RESIDUAL EFFECT IN MOST OF THE CASES, UNLESS THE EVENT OF DEATH OCCURRED. OVERALL, THE RESULTS DID NOT SHOW ANY EVIDENCE OF INCREASED MORBIDITY ASSOCIATED WITH THE IMPELLA SUPPORT IN CARDIOMYOPATHY, MYOCARDITIS, AND PERIPARTUM CARDIOMYOPATHY (PPCM) PATIENTS." Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT AND EMBOLISM HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. G1 REPORTING CONTACT FACILITY NAME WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28460. H6 HEALTH EFFECT - CLINICAL CODE 4438 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28460.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22030576&pc=MVK | 3008642652-2025-04874 | 2025-04-23 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-15 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 09:05:07 ON (B)(6) 2025. AT 16:09:38, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS TACHYCARDIA @ 110 BPM WITH TALL T WAVES. AT 16:10:29, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. DOUBLE COUNTING CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS TACHYCARDIA @ 110 BPM WITH TALL T WAVES. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA @ 100 BPM WITH TALL T WAVES AND MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 16:51:59 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22030579&pc=MVK | 3008642652-2025-04875 | 2025-03-27 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-15 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THIRTEEN APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 14:49:47 ON (B)(6) 2025. AT 05:46:01 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. FROM 05:46:38 TO 06:06:29, THE PATIENT RECEIVED TWELVE APPROPRIATE TREATMENTS IN RESPONSE TO VF. THE POST SHOCK RHYTHM OF THESE TREATMENTS WAS SINUS RHYTHM @20-70 BPM. AT 06:07:03, THE PATIENT RECEIVED THE THIRTEENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF WITH MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. THE ELECTRODE BELT WAS DISCONNECTED AT 08:11:14 ON (B)(6) 2025. Manufacturer Narrative: INCOMING TESTING OF THE ELECTRODE BELT HAS BEEN COMPLETED. THE ELECTRODE BELT ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22031079&pc=NPT | 2025587-2025-03713 | 2025-04-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-15 04:00 | NPT | EVOLUT FX PLUS VALVE | Mechanics Altered; Patient Device Interaction Problem; Material Split, Cut or Torn; Central Regurgitation | Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER VALVE PROCEDURE INVOLVING THE EVFXPLUS-26, FOLLOWING THE PLACEMENT OF THE FIRST VALVE, SEVERE AORTIC INSUFFICIENCY (AI) WAS OBSERVED. IT WAS THOUGHT THAT THE SEVERE AI WAS DUE TO A TORN OR STUCK LEAFLET. A POST-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED TO ADDRESS THE AI. DESPITE THE BAV, THE AI DID NOT IMPROVE. A SECOND VALVE WAS IMPLANTED IN THE PATIENT AS CARDIOPULMONARY RESUSCITATION WAS PERFORMED. ULTIMATELY, THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: L-EVOLUTFX-2329 (LOT: 0012347822); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; PRODUCT ID: D-EVOLUTFX-2329 (LOT: 0012370579); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; PRODUCT ID: EVFXPLUS-26 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22031084&pc=NPT | 2025587-2025-03714 | 2025-04-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-15 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem; Central Regurgitation | Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER VALVE PROCEDURE INVOLVING THE EVFXPLUS-26, FOLLOWING THE PLACEMENT OF THE FIRST VALVE, SEVERE AORTIC INSUFFICIENCY (AI) WAS OBSERVED. IT WAS THOUGHT THAT THE SEVERE AI WAS DUE TO A TORN OR STUCK LEAFLET. A POST-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED TO ADDRESS THE AI. DESPITE THE BAV, THE AI DID NOT IMPROVE. A SECOND VALVE WAS IMPLANTED IN THE PATIENT AS CARDIOPULMONARY RESUSCITATION WAS PERFORMED. ULTIMATELY, THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: L-EVOLUTFX-2329 (LOT: 0012347822); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; PRODUCT ID: D-EVOLUTFX-2329 (LOT: 0012370579); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE PRODUCT ID EVFXPLUS-26 (K055267); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22031085&pc=BZD | 2518422-2025-105711 | 2025-04-21 04:00 | Death | RESPIRONICS, INC. | 2025-05-15 04:00 | BZD | REMSTAR PRO C-FLEX+ SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K113068 |  | 1 | Event Description: THE REPORTER CALLED IN STATING THE DEVICE IS NO LONGER AVAILABLE FOR RETURN. THE REPORTER STATED THE PATIENT EXPIRED 3 YEARS AGO (EXACT DATE OF DEATH NOT REPORTED). THE CAUSE OF DEATH WAS NOT REPORTED. FOLLOW-UP EFFORTS HAVE BEEN MADE TO ACQUIRE INFORMATION SURROUNDING THE PATIENT'S DEATH. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22031137&pc=FRN | 3005053560-2025-00003 | 2025-04-03 04:00 | Death | IRADIMED CORPORATION | 2025-05-15 04:00 | FRN | MRIDIUM | Improper or Incorrect Procedure or Method | Insufficient Information | K143369 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A POTENTIAL UNDER-INFUSION EVENT. ACCORDING TO BIOMED, THE MRI TECHNICIAN OBSERVED A BUBBLE OF AIR IN THE IV LINE. HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED REGARDING WHETHER ANY ALARMS OCCURRED DURING THE EVENT. PER BIOMED THE PATIENT CODED DURING THE EVENT. Manufacturer Narrative: THE DEVICE HISTORY LOG SUGGESTED THAT THE USER LIKELY HAD ISSUES PRIMING THE IV SET LINE. A "BUBBLE DETECTOR ACTIVITY" WAS LOGGED, INDICATING THAT THE BUBBLE DETECTOR WAS ACTIVE AND DETECTED A BUBBLE BELOW THE ALARM THRESHOLD IMMEDIATELY AFTER THE INFUSION STARTED. SUBSEQUENTLY, A "BUBBLE DETECTED" ALARM OCCURRED SHORTLY AFTER THE INFUSION BEGAN. THIS WAS CONSISTENT WITH THE EVENT DESCRIPTION PROVIDED BY THE BIOMED, WHERE THE USER OBSERVED AIR IN THE IV SET WHEN THE "BUBBLE DETECTED" ALARM WAS TRIGGERED. THE PRESENCE OF THE "BUBBLE DETECTED" ALARM, OCCURRING AFTER THE INFUSION PARAMETERS WERE SET AND IMMEDIATELY AFTER THE INFUSION BEGAN, STRONGLY INDICATED THAT AIR WAS PRESENT IN THE IV LINE AT THE START OF THE INFUSION. PROPER PRIMING OF THE IV SET IS CRUCIAL TO REMOVE ALL AIR FROM THE TUBING BEFORE CONNECTING IT TO THE PATIENT AND INITIATING THE INFUSION. THE OPERATOR'S MANUAL AND THE IV SET POUCH LABEL CONTAINED INSTRUCTIONS ON HOW TO PROPERLY PRIME THE IV SET. THE DEVICE HISTORY LOG SHOWED THAT THE USER OPENED THE DOOR AFTER THE BUBBLE ALARM, POTENTIALLY ATTEMPTING TO TROUBLESHOOT THE ISSUE. ALL ALARMS WERE SILENCED BY THE USER BEFORE THE INFUSION WAS CANCELED. A BUBBLE ALARM IN THIS CONTEXT IS A COMMON SIGN THAT THE PRIMING STEP MAY HAVE BEEN INSUFFICIENT, MISSED, OR THAT THE USER EXPERIENCED DIFFICULTY PRIMING AND SETTING UP THE LINE. THIS COULD HAVE LED TO AIR BEING DETECTED BY THE PUMP AND MAY HAVE RESULTED IN A DELAY OF THERAPY. THE CUSTOMER CONFIRMED THAT THE PATIENT CODED DURING THE EVENT AND THAT MEDICAL INTERVENTION WAS REQUIRED. THEREFORE, IRADIMED IS REPORTING THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22031335&pc=NAY | 2955842-2025-20615 | 2025-04-09 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-05-15 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Exsanguination; Insufficient Information | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI ASSISTED NEPHRECTOMY, THERE WAS AN UNSPECIFIED SURGICAL COMPLICATION RESULTING IN MAJOR HEMORRHAGING. THE SURGEON REPORTED THAT THE BLEEDING REQUIRED A CONVERSION TO OPEN LAPAROTOMY AND AN OPEN THORACOTOMY. DESPITE ALL RESUSCITATIVE EFFORTS THE PATIENT EXPIRED IN THE OPERATING ROOM. THE SURGEON REPORTED THAT THE SURGICAL COMPLICATION WAS NOT RELATED TO THE ROBOTIC SYSTEM AND THERE WAS NO MALFUNCTION OF THE SYSTEM OR INSTRUMENTS. THE STATED CAUSE OF DEATH WAS MAJOR HEMORRHAGE. ADDITIONAL EVENT INFORMATION WAS REQUESTED BUT WAS NOT PROVIDED BY THE SURGEON.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: THE FOLLOWING CONCOMITANT DEVICES WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, FENESTRATED BIPOLAR FORCEPS, PROGRASP FORCEPS, VESSEL SEALER EXTEND. A SITE REVIEW FOUND THAT NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THE PRODUCTS USED. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICES USED DURING THE EVENT FOUND NO RELATED NON-CONFORMANCES. DIGITAL SIGNAL PROCESSOR (DSP) LOGS SHOW THAT THE VESSEL SEALER EXTEND (VSE) INSTRUMENT WAS INSTALLED 3 TIMES AND PASSED HOMING 3 TIMES. ALL CUT COMPLETE AND SEAL EVENTS WERE NOTED WITH NO ERRORS. THE MASTER SUPERVISORY CONTROLLER (MSC) LOGS DID NOT SHOW ANY VSE RELATED ERRORS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT, WITH A SIGNIFICANT COMORBID DIAGNOSIS OF XANTHOGRANULOMATOUS PYELONEPHRITIS, WAS UNDERGOING A LEFT NEPHRECTOMY WHEN SIGNIFICANT BLEEDING OCCURRED. WHAT BLED AND HOW THEY GOT INTO THAT BLEEDING WAS NOT PROVIDED. THE PROCEDURE WAS CONVERTED TO AN OPEN LAPAROTOMY AND ALSO REQUIRED A THORACOTOMY WHICH SUGGESTS THEY HAD TO EMERGENTLY GET CONTROL OF THE AORTA. HOW THESE EVENTS UNFOLDED ARE NOT PROVIDED, HOWEVER, THE PATIENT EVENTUALLY DIED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22032686&pc=EOQ | 2955842-2025-20837 | 2025-04-18 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-05-15 04:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT'S BLOOD PRESSURE DROPPED SIGNIFICANTLY, CAUSING HEMODYNAMIC INSTABILITY LEADING TO DEATH. THE PROCEDURE INVOLVED ION NAVIGATION AND RADIAL ENDOBRONCHIAL ULTRASOUND (REBUS), IT WAS INTERRUPTED BEFORE THE BIOPSY COULD BE PERFORMED. APPROXIMATELY 10-15 MINUTES INTO NAVIGATION, THE ANESTHESIOLOGIST INFORMED THE PROCEDURE TEAM OF THE PATIENT¿S ¿DANGEROUSLY LOW¿ BLOOD PRESSURE. THE ION SYSTEM WAS UNDOCKED, AND CARDIOPULMONARY RESUSCITATION WAS INITIATED. THE PATIENT WAS PRONOUNCED DEAD APPROXIMATELY 45 MINUTES LATER. THERE WAS NO REPORTED DEVICE MALFUNCTION ASSOCIATED WITH THE EVENT. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE PHYSICIAN WERE MADE; NO RESPONSE HAS BEEN RECEIVED. Manufacturer Narrative: A REVIEW OF THE ION SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND NO RELATED SYSTEM ERRORS OCCURRED DURING THE PROCEDURE THAT WERE RELEVANT TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT UNDERWENT AN ION BRONCHOSCOPY, AND DURING NAVIGATION AND REBUS IMAGING, PRIOR TO BIOPSY, THE PATIENT DEVELOPED HYPOTENSION. EFFORTS AT RESUSCITATION WERE UNSUCCESSFUL AND THE PATIENT DIED. THERE WERE NO REPORTED MALFUNCTIONS OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. NO FURTHER DATA IS AVAILABLE DESPITE MULTIPLE EFFORTS. BASED ON THE AVAILABLE DATA THE EVENTS WERE PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. --FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. --FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. --KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. --BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22032689&pc=EOQ | 2955842-2025-20836 | 2025-04-17 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-05-15 04:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Hemorrhage/Bleeding; Insufficient Information | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, A PATIENT EXPERIENCED SIGNIFICANT BLEEDING, PULMONARY EMBOLISM, CARDIAC ARREST AND DEATH. THE INTUITIVE ENDOLUMINAL SALES MANAGER (ESM) WAS INFORMED BY HOSPITAL PERSONNEL THAT DURING THE BIOPSY PROCEDURE, A THIRD PARTY CRYOPROBE WAS USED TO COLLECT THE BIOPSY SAMPLES; THE BIOPSY SPECIMEN APPEARED SUGGESTIVE OF BLOOD VESSEL INVOLVEMENT. FOLLOWING THE BIOPSY, THE PATIENT HAD SIGNIFICANT BLEEDING, PROMPTING A BRONCHOSCOPIC AIRWAY INSPECTION AND SUCTIONING OF THE BLEEDING. CONE BEAM CT SCAN SHOWED A PULMONARY EMBOLISM. SHORTLY THEREAFTER, THE PATIENT SUFFERED CARDIAC ARREST AND REQUIRED RESUSCITATION, INCLUDING TRANSFUSION OF BLOOD PRODUCTS. THE PATIENT WAS STABILIZED; HOWEVER, EXPIRED AT AN UNSPECIFIED LATER TIME. THE ESM WAS INFORMED THAT THE ADVERSE EVENTS WERE NOT ATTRIBUTED TO THE ION SYSTEM, AND NO DEVICE MALFUNCTION WAS IDENTIFIED. THE COMPLICATIONS WERE BELIEVED TO BE RELATED TO THE CRYOPROBE BIOPSY THAT INDICATED VESSEL INVOLVEMENT. ADDITIONAL INFORMATION WAS REQUESTED FROM THE RESPIRATORY MEDICINE PHYSICIAN WHO STATED THE FACTS PROVIDED WERE INCORRECT, AND THAT THE MATTER IS CONFIDENTIAL AND INTERNAL REVIEW IS IN PROGRESS. NO CLARIFICATION OF THE INCORRECT FACTS WAS PROVIDED. Manufacturer Narrative: SYSTEM LOGS ARE NOT AVAILABLE FOR REVIEW. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT UNDERWENT AN ION LUNG BIOPSY WHERE A CRYO-PROBE WAS USED TO OBTAIN BIOPSY SAMPLES. DURING THE PROCEDURE, AFTER BIOPSIES WERE OBTAINED, SIGNIFICANT BLEEDING WAS NOTED AND A PULMONARY EMBOLISM WAS DIAGNOSED ON CONE BEAM CT. THE PATIENT FURTHER DECOMPENSATED WITH A CARDIAC ARREST. DESPITE AGGRESSIVE INTERVENTIONS AND EFFORTS AT RESUSCITATION, THE PATIENT ULTIMATELY DIED. BASED ON THE AVAILABLE DATA THE BLEEDING, PULMONARY EMBOLISM, AND DEATH POST BIOPSY WERE LIKELY PROCEDURE RELATED. THERE IS NO COMPELLING EVIDENCE THAT THE EVENTS WERE DEVICE RELATED. OF NOTE, THE PHYSICIAN REPORTED DISCREPANCIES IN THE DETAILS BUT DECLINED TO PROVIDE FURTHER INFORMATION. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW-RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50 ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE-CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF ANY BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF MORE SEVERE CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. THE RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). THE PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). THE RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). THERE WERE NO REPORTED EVENTS OF VENOUS THROMBOEMBOLISM (VTE) INCLUDING DEEP VENOUS THROMBOSIS OR PULMONARY EMBOLISM. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS AND NO VTE. HOWEVER, IT IS POSSIBLE THAT IMMOBILITY DURING GENERAL ANESTHESIA AND BLEEDING WITH SUBSEQUENT CLOTTING ASSOCIATED WITH LUNG BIOPSIES MAY CONFER SOME DEGREE OF INCREASED RISK FOR VTE EVENTS. LIKELY OF MORE RELEVANCE, IT IS ESTIMATED THAT CANCER CONFERS A 4-7X HIGHER RISK FOR VTE. THE HIGHER RATE OF VTE HAS BEEN REPORTED TO RANGE BETWEEN 3-13.9% IN PATIENTS WITH LUNG CANCER. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. BLOM JW, DOGGEN CJ, OSANTO S, ROSENDAAL FR. MALIGNANCIES, PROTHROMBOTIC MUTATIONS, AND THE RISK OF VENOUS THROMBOSIS. JAMA. 2005. BAGCHI A, KHAN MS, SARASWAT A, ANSARI A, NAI Q, IYER V, HAMOUDA D, KHUDER S, VERGHESE C. INCREASED INCIDENCE OF THROMBOTIC COMPLICATIONS WITH NON-SMALL CELL LUNG CANCER NECESSITATES CONSIDERATION OF PROPHYLACTIC ANTICOAGULATION IN YOUNG INDIVIDUALS. CUREUS. 2021. CHEW HK, DAVIES AM, WUN T, HARVEY D, ZHOU H, WHITE RH. THE INCIDENCE OF VENOUS THROMBOEMBOLISM AMONG PATIENTS WITH PRIMARY LUNG CANCER. J THROMB HAEMOST. 2008. CONNOLLY GC, DALAL M, LIN J, KHORANA AA. INCIDENCE AND PREDICTORS OF VENOUS THROMBOEMBOLISM (VTE) AMONG AMBULATORY PATIENTS WITH LUNG CANCER. LUNG CANCER. 2012.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22016886&pc=PNJ | 9612164-2025-02390 | 2025-04-22 04:00 | Death | MEDTRONIC IRELAND | 2025-05-14 04:00 | PNJ | MICRA AV | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Stroke/CVA | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE THE INTRODUCER WAS INSERTED AND THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) DELIVERY SYSTEM (DS) WAS INSERTED THROUGH THE INTRODUCER. THE PATIENT STOPPED RESPONDING AND THE PHYSICIAN STOPPED THE PROCEDURE. RESCUE EFFORTS WERE MADE WITH CARDIOPULMONARY RESUSCITATION, DEFIBRILLATION AND MASSAGE PERFORMED, HOWEVER THE PATIENT HAD A STROKE AND PASSED AWAY. THE LEADLESS IPG DS AND INTRODUCER WERE REMOVED. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE FULL DELIVERY SYSTEM WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. THE ANALYST NOTED THE FULL DELIVERY SYSTEM WAS RETURNED WITH THE DEVICE INTACT IN THE DEVICE CUP. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [PRODUCT ID-DELSYS]. D9: YES, RETURN DATE: 10-JUN-2025. H3: YES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22016892&pc=DYB | 9612164-2025-02391 | 2025-04-22 04:00 | Death | MEDTRONIC IRELAND | 2025-05-14 04:00 | DYB | MICRA | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Stroke/CVA | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE THE INTRODUCER WAS INSERTED AND THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) DELIVERY SYSTEM (DS) WAS INSERTED THROUGH THE INTRODUCER. THE PATIENT STOPPED RESPONDING AND THE PHYSICIAN STOPPED THE PROCEDURE. RESCUE EFFORTS WERE MADE WITH CARDIOPULMONARY RESUSCITATION, DEFIBRILLATION AND MASSAGE PERFORMED, HOWEVER THE PATIENT HAD A STROKE AND PASSED AWAY. THE LEADLESS IPG DS AND INTRODUCER WERE REMOVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE INTRODUCER WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. THE ANALYST NOTED THE INTRODUCER WAS RETURNED WITH THE DILATOR FULLY INSERTED IN THE SHEATH OF THE INTRODUCER. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22016978&pc=QFG | 3004464228-2025-20970 | 2025-04-22 04:00 | Death | INSULET CORPORATION | 2025-05-14 04:00 | QFG | OMNIPOD 5 AUTOMATED INSULIN DELIVERY SYSTEM | Insufficient Information | Insufficient Information | K203768 |  | 1 | Event Description: A HEALTH CARE PROFESSIONAL (HCP) HAS CALLED INSULET ON (B)(6) 2025 AT 3.14PM TO REPORT AN OMNIPOD (SPECIFIC PRODUCT WAS NOT REPORTED) INSULIN PUMP USER DEATH. THE HCP'S PATIENT REPORTED TO THEM THAT THEY KNEW THAT SOMEONE HAD DIED WHILE USING THE OMNIPOD SYSTEM. THE EXACT DATE OF THE EVENT WAS NOT PROVIDED. THEY ALLEGE THAT "THE PERSONAL DIABETES MANAGER (PDM) GAVE THE ORDER TO GIVE ALL THE INSULIN IN ONE TIME TO THE PATIENT AND THAT IS WHY THE PATIENT DIED". IT IS UNKNOWN WHICH OMNIPOD DEVICE THE PATIENT WAS USING AT THE TIME THEY PASSED AWAY. THE HCP WAS NOT WILLING TO PROVIDE THEIR PATIENT'S DETAILS OR ANY DETAILS OF THE USER THAT HAS PASSED AWAY. THEREFORE, AT THIS TIME, NO FURTHER INFORMATION ON THE ¿SYSTEM ERROR¿ REPORTED BY THE HCP PATIENT CAN BE OBTAINED. THE INVESTIGATION IS ONGOING AND INSULET IS ATTEMPTING TO RECOVER ADDITIONAL DETAILS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE HCP DID NOT HAVE ANY DETAILS RELATED TO THE DEVICE, EVENT DETAILS, OR PATIENT IDENTIFICATION. THE HCP DECLINED WHEN ASKED FOR THE REPORTER CONTACT INFORMATION. IF NEW INFORMATION IS RECEIVED RELATED TO THIS REPORT IT WILL BE REASSESSED AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE INVESTIGATION IS ONGOING AND INSULET IS ATTEMPTING TO RECOVER ADDITIONAL DETAILS, INCLUDING PRODUCT NAME, LOT NUMBER AND SERIAL NUMBER. IF ADDITIONAL INFORMATION IS RECEIVED, INSULET WILL SUBMIT A SUPPLEMENTAL REPORT AND CONDUCT ALL POSSIBLE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22017031&pc=PNJ | 9612164-2025-02392 | 2025-04-25 04:00 | Death | MEDTRONIC IRELAND | 2025-05-14 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) DELIVERY SYSTEM (DS), THE PATIENT EXPERIENCED CARDIAC TAMPONADE, CARDIAC PERFORATION AND PERICARDIAL EFFUSION. A PERICARDIOCENTESIS WAS PERFORMED WITH OVER ONE LITER OF BLOOD REMOVED. THE PATIENT PASSED AWAY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE FULL DELIVERY SYSTEM WAS RETURNED WITH THE DEVICE INTACT IN THE DEVICE CUP AND NO ANOMALIES WERE FOUND. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6)). D9: YES, RETURN DATE: 2025-05-22 H3: <YES> DEV RTN TO MFR? YES, EVAL SUMM SUBMITTED? YES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ID# MC2AVR1-DELSYS CORRECTED: ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1 -DELSYS] (SERIAL: [(B)(6)]). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22017364&pc=MNS | 2518422-2025-105595 | 2025-05-12 04:00 | Death | RESPIRONICS, INC. | 2025-05-14 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE THIRD-PARTY SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND IDENTIFIED THE FIRST TEST FAILED DUE TO A PRESSURE-RELATED ISSUE, SO THE BLOWER WAS CHANGED. FOLLOWING THE REPLACEMENT, THE DEVICE UNDERWENT TESTING AND SUCCESSFULLY PASSED. NO ERRORS WERE OBSERVED DURING THE EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22017594&pc=MNS | 2518422-2025-105596 | 2025-05-12 04:00 | Death | RESPIRONICS, INC. | 2025-05-14 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE THIRD-PARTY SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND NO ERRORS. THE DEVICE WAS TESTED AND PASSED ALL TESTS. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22017605&pc=OZP | 2032227-2025-174472 | 2025-04-16 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-14 04:00 | OZP | PUMP MMT-1884 MM780G MG | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER HAD PASSED AWAY, AND LEADING UP TO THE EVENT WERE THE COMPLICATIONS FROM DIABETES. THE EVENT INVOLVED THE PRODUCTS UNOMEDICAL AND UNK_NGP. TROUBLESHOOTING WAS PARTIALLY PERFORMED. IT WAS UNKNOWN THAT THE INSULIN PUMP WAS USED WITHIN 48 HOURS. IT WAS ALSO UNKNOWN THAT THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR AN UNKNOWN NGP.Event Description: UPDATED SUMMARY: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER HAD PASSED AWAY DUE TO COMPLICATIONS FROM DIABETES. THE EVENT INVOLVED THE PRODUCTS MMT-397A, MMT-332A AND MMT-1884. TROUBLESHOOTING WAS PARTIALLY PERFORMED. IT WAS UNKNOWN WHETHER THE INSULIN PUMP WAS USED WITHIN 48 HOURS. IT WAS ALSO UNKNOWN WHETHER THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-397A. NO PRODUCT RETURN IS REQUIRED FOR AN MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: ADDITIONAL INFORMATION HAS BEEN RECEIVED REGARDING THE PRODUCT MATERIAL CHANGE WHICH WAS NOT INCLUDED IN THE INITIAL REPORT. SO, THE CORRECT PRODUCT MATERIAL HAS BEEN CHANGED FROM UNK_NGP TO MMT-1884 AS PER NEW ADDITIONAL INFORMATION AND UPDATED IN SECTIONS B5, D1/D2A/D2B, D3, D4 AND D10 (UPDATED THE CONCOMITANT PRODUCTS) WITH THIS REPORT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22017734&pc=JXG | 9612501-2025-01283 | 2025-03-28 04:00 | Death | MEDTRONIC DOMINICANA | 2025-05-14 04:00 | JXG | UNKNOWN VALVE/SHUNT | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: MICHAEL J. STUART, ANNABELLE M. HARBISON, TIMOTHY RUDER, JOSHUA HACKNEY, NORMAN MA, ROBERT A. J. CAMPBELL, AMELIA J. JARDIM, AND LIAM G. COULTHARD. THE RISKS OF OMITTING ROUTINE POSTOPERATIVE IMAGING AFTER INSERTION OF NEW VENTRICULAR SHUNTS IN CHILDREN: A 10-YEAR RETROSPECTIVE COHORT STUDY. JOURNAL OF NEUROSURGERY PEDIATRICS (2025). DOI: 10.3171/2024.12.PEDS24496. OBJECTIVE VENTRICULOPERITONEAL SHUNT INSERTION IS THE PRIMARY TREATMENT FOR PEDIATRIC HYDROCEPHALUS AND IS ONE OF THE MOST COMMON PROCEDURES PERFORMED IN PEDIATRIC NEUROSURGERY. HOWEVER, THE VALUE OF ROUTINE POSTOPERATIVE IMAGING REMAINS CONTROVERSIAL. ALTHOUGH SOME EVIDENCE SUGGESTS THAT ROUTINE POSTOPERATIVE IMAGING PROVIDES A LOW YIELD AFTER SHUNT REVISION, ITS ROLE FOLLOWING INITIAL SHUNT INSERTION IN CHILDREN HAS NOT BEEN REPORTED. ADDITIONALLY, THE ROLE OF ROUTINE POSTOPERATIVE IMAGING AS A BASELINE STUDY IN ALTERING THE RATE OF LONGER-TERM SHUNT REVISION RATES REMAINS UNEXPLORED. METHODS A RETROSPECTIVE REVIEW WAS CONDUCTED USING A PROSPECTIVELY MAINTAINED SURGICAL DATABASE, ENCOMPASSING 10 YEARS OF CONSECUTIVE CASES FROM A QUATERNARY PEDIATRIC NEUROSURGICAL CENTER SERVING A POPULATION OF 5.3 MILLION. STATEWIDE ELECTRONIC MEDICAL RECORDS INCLUDING IMAGING AND MORTALITY DATA WERE REVIEWED. ONLY CASES INVOLVING THE INSERTION OF A NEW VENTRICULAR SHUNT SYSTEM WERE INCLUDED. BASED ON LONGSTANDING PHYSICIAN PREFERENCES, APPROXIMATELY HALF THE NEUROSURGEONS IN THE DEPARTMENT DID NOT ROUTINELY PERFORM POSTOPERATIVE IMAGING DURING THE STUDY PERIOD. THE PRIMARY OUTCOMES WERE SHUNT-RELATED MORTALITY DURING THE STUDY PERIOD, SHUNT REVISION OR MORTALITY WITHIN 7 DAYS OF SURGERY, OR HOSPITAL DISCHARGE, WITH TIME TO FIRST REVISION AS A SECONDARY OUTCOME. RESULTS DURING THE STUDY PERIOD 1485 SHUNT PROCEDURES WERE PERFORMED, OF WHICH 427 INVOLVED THE PLACEMENT OF A NEW VENTRICULOPERITONEAL SHUNT SYSTEM. THE MEAN AGE OF THE PATIENTS WAS 5.2 YEARS (RANGE PREMATURE TO 18 YEARS). A TOTAL OF 206/427 (48%) UNDERWENT ROUTINE POSTOPERATIVE IMAGING WITH MR OR CT, AND 153/427 (36%) UNDERWENT ROUTINE POSTOPERATIVE SHUNT SERIES RADIOGRAPHS. DURING THE STUDY PERIOD, 2 ABDOMINAL RADIOGRAPHS IDENTIFIED PREPERITONEAL PLACEMENT; HOWEVER, IN BOTH CASES A SYMPTOMATIC SUBCUTANEOUS FLUID COLLECTION WAS PRESENT (2/427, 0.5%). POSTOPERATIVE CT OR MR STUDIES IDENTIFIED 1 CASE OF EXTRAVENTRICULAR CATHETER PLACEMENT AND 5 CASES OF A KINK ADJACENT TO THE VALVE (6/427, 1.4%), OF WHICH 4 CASES WERE DETECTED ON ROUTINE IMAGING (4/427, 0.9%). NO SIGNIFICANT DIFFERENCES WERE DETECTED BETWEEN PATIENTS WITH OR WITHOUT IMAGING IN 7-DAY POSTOPERATIVE (14 [7%] VS 8 [4%], P = 0.25) OR 7-DAY POSTDISCHARGE (2 [1%] VS 1 [<(><<)> 1%], P = 0.96) RATES OF NONINFECTIVE SHUNT REVISION, SHUNT-RELATED MORTALITY (4 [2%] VS 5 [2%], RESPECTIVELY, P = 0.8) OR KAPLAN-MEIER LONG-TERM SHUNT SURVIVAL CURVES (ALL P > 0.05). CONCLUSIONS ROUTINE POSTOPERATIVE IMAGING AFTER PEDIATRIC SHUNT INSERTION RARELY IDENTIFIES CLINICALLY SIGNIFICANT ISSUES, AND CONSI DERATION COULD BE GIVEN TO ITS OMISSION. REPORTED EVENTS: - NO STATISTICALLY SIGNIFICANT DIFFERENCES IN SHUNT-RELATED MORTALITY DURING THE STUDY PERIOD OCCURRED IN THE GROUPS WHO DID OR DID NOT UNDERGO ROUTINE POSTOPERATIVE AXIAL IMAGING (4 [2%] VS 5 [2%], P = 0.8) OR RADIOGRAPHS (2 [1%] VS 7 [3%], P = 0.39). IT WAS INDICATED THAT MEDTRONIC DELTA AND STRATA VALVES WERE USED IN THE LITERATURE ARTICLE, HOWEVER, THE ARTICLE DOES NOT INDICATED WHICH FAILURES AND ADVERSE EVENTS ARE ASSOCIATED WITH WHAT VALVES. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22017907&pc=BZD | 2518422-2025-105597 | 2025-04-03 04:00 | Death | RESPIRONICS, INC. | 2025-05-14 04:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE STATED SHE RECEIVED A LETTER REGARDING THE PATIENT'S DEVICE, BUT THE PATIENT HAS PASSED AWAY AND NO LONGER NEEDS A REPLACEMENT DEVICE OR ANYTHING. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22018270&pc=DSQ | 2916596-2025-02845 | 2020-10-18 04:00 | Death | THORATEC CORPORATION | 2025-05-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED IN THE RESEARCH ARTICLE ¿CHARACTERIZING SURVIVAL FOR PATIENTS SUPPORTED WITH INOTROPES AFTER THE 2018 DONOR ALLOCATION RESTRUCTURING: A UNOS DATABASE ANALYSIS,¿ THAT THE HEARTMATE 3 MAY BE ASSOCIATED WITH STROKE, DEATH, AND TRANSPLANT. A RETROSPECTIVE COHORT STUDY THAT UTILIZED THE UNITED NETWORK FOR ORGAN SHARING (UNOS) TRANSPLANT REGISTRY TO IDENTIFY ALL ADULT PATIENTS WHO WERE WAITLISTED WHILE SUPPORTED WITH INOTROPES WAS PERFORMED. PATIENTS LISTED FOR HEART TRANSPLANT BETWEEN 18JULY2014 AND 28JUL2016 (PRE POLICY REVISION) AND BETWEEN 18OCT2018 AND 18OCT2020 (POST POLICY REVISION) WERE IDENTIFIED FROM THE UNOS DATABASE. PATIENTS WITH RIGHT SIDED VENTRICULAR ASSIST DEVICE WERE EXCLUDED FROM THE STUDY, AND PATIENTS WITH ADDITIONAL CIRCULATORY SUPPORT WERE ALSO EXCLUDED. SUB-ANALYSES EXAMINED TRENDS IN DEVICE PROGRESSION WHERE PATIENTS LISTED ON INOTROPES WERE LATER TRANSPLANTED ON INOTROPES AND/OR ON EXTRACORPOREAL MEMBRANOUS OXYGENATOR (ECMO), DURABLE LEFT VENTRICULAR ASSIST DEVICE (LVAD- HVAD, HEARTMATE 2 AND HEARTMATE 3), TEMPORARY MECHANICAL CIRCULATORY SUPPORT (TMCS), OR INTRA-AORTIC BALLOON PUMP (IABP). SURVIVAL DATA POST-TRANSPLANT WERE CALCULATED AND PLOTTED. PATIENTS FOR WHOM PRIORITY STATUS WAS NOT PROVIDED OR COULD NOT BE DEDUCED FROM THE REGISTRY WERE EXCLUDED FROM THE STUDY. PATIENTS NOT LISTED ON EITHER INOTROPES OR ANOTHER DEVICE AT THE TIME OF LISTING OR TRANSPLANTATION WERE REMOVED. A TOTAL OF 3,189 PATIENTS WERE WAITLISTED WHILE SUPPORTED WITH INOTROPES (1,408 PRE POLICY, 1,781 POST POLICY). AFTER THE POLICY CHANGE, PATIENTS TENDED TO HAVE A LOWER CARDIAC OUTPUT (CO) AND A HIGHER MEAN PULMONARY CAPILLARY WEDGE PRESSURE (PCWP) (3.9L/MIN [3.2, 4.8] VS. 3.8 L/MIN [3.1,4.7], P=0.04 AND 21MMHG [16,26] VS. 22 [16,27], P =0.02, RESPECTIVELY) AT THE TIME OF LISTING. ADDITIONALLY, FOLLOWING THE POLICY CHANGE, MORE RECIPIENTS WERE LIKELY TO BE ON A VENTILATOR AT THE TIME OF LISTING (1.5% VS.3.3%, P <0.01, RESPECTIVELY). THE CHANGE IN THE ALLOCATION POLICY LED TO MORE PATIENTS BEING TRANSPLANTED AS WELL AS SHORTER WAITLIST TIMES (78.9% VS. 89.8%, P<0.01, RESPECTIVELY). COMPARED TO THE PERIOD PRIOR TO THE POLICY CHANGE, LESS PATIENTS WERE TRANSPLANTED ON INOTROPES (72.6% VS. 42.6%, P < 0.01) OR ON AN LVAD (14.9% VS. 3.4%, P < 0.01). MEDIAN FOLLOW-UP TIME FOR PATIENTS WAITLISTED PRE AND POST POLICY DID NOT DIFFER (2.9 YEARS [1.6, 3.5] VS. 2.8 YEARS [2, 3.1], P < 0.01). THERE WERE NO DIFFERENCES IN STROKE RATES BETWEEN THE TWO TIME PERIODS FOR WAITLISTED PATIENTS (2.1% VS. 2.8%, P = 0.22). FOLLOWING THE CHANGE TO THE ALLOCATION POLICY, THERE WAS NO DIFFERENCE IN THE NUMBER OF PATIENTS REMOVED FROM THE WAITLIST, BUT OF THOSE REMOVED, MORE WERE REMOVED DUE TO TRANSPLANTATION. FURTHERMORE, MORTALITY DECREASED FOR WAITLISTED PATIENTS. FOR PATIENTS TRANSPLANTED WITH HIGHER FORMS OF CIRCULATORY SUPPORT, SURVIVAL DID NOT DIFFER BETWEEN POLICY PERIODS. SINCE THE 2018 UNOS POLICY CHANGE, VARIOUS BENEFITS HAVE BEEN DEMONSTRATED INCLUDING DECREASED WAITLIST MORTALITY, MORE EQUITABLE DISTRIBUTION OF ORGANS TO HIGHER PRIORITY PATIENTS ACROSS WIDER GEOGRAPHICAL AREAS, AND INCREASED TRANSPLANTATION RATE. ALLOCATION POLICY REVISIONS HAVE CONTRIBUTED TO INCREASED UTILIZATION OF TEMPORARY SUPPORT (ECMO, TMCS, AND IABP) AND DECREASED UTILIZATION OF OTHERS SUCH AS DURABLE LVADS. ADDITIONALLY, REVISIONS HAVE LED TO IMPROVED SURVIVAL AND INCREASED TRANSPLANTATION FOR PATIENTS WAITLISTED ON INOTROPES, YET SIMILAR SURVIVAL FOR EACH INDIVIDUAL FORM OF TEMPORARY SUPPORT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B2: SPECIFIC DEATH DATE IS DOCUMENTED AS UNKNOWN. SECTION B3: DATE OF EVENT IS APPROXIMATE AS THE DATA WERE COLLECTED BETWEEN 18JUL2014 AND 18OCT2020. AUTHOR INFORMATION: PRITTING, C., CHEN, J., EKIYOYO, A., WARNER, E., BRAILOVSKY, Y., TCHANTCHALEISHVILI, V., & RAJAPREYAR, I. (2025). CHARACTERIZING SURVIVAL FOR PATIENTS SUPPORTED WITH INOTROPES AFTER THE 2018 DONOR ALLOCATION RESTRUCTURING: A UNOS DATABASE ANALYSIS. CLINICAL TRANSPLANTATION, 39(3). HTTPS://DOI.ORG/10.1111/CTR.70105 DEPARTMENT OF SURGERY, BAYHEALTH MEDICAL CENTER, DOVER, DELAWARE, USA; DIVISION OF CARDIAC SURGERY, THOMAS JEFFERSON UNIVERSITY, PHILADELPHIA, PENNSYLVANIA, USA ; DIVISION OF CARDIOLOGY, THOMAS JEFFERSON UNIVERSITY, PHILADELPHIA, PENNSYLVANIA; USA DIVISION OF CARDIOLOGY, HEART AND VASCULAR INSTITUTE, CLEVELAND CLINIC, CLEVELAND, OHIO, USA; DIVISION OF CARDIOLOGY, TUFTS UNIVERSITY, BOSTON, MASSACHUSETTS, USA CORRESPONDENCE: ERIC WARNER (ERICDWARNER1124@GMAIL.COM). MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22018288&pc=DSQ | 3007042319-2025-01145 | 2025-05-10 04:00 | Death | HEARTWARE, INC. | 2025-05-14 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Intracranial Hemorrhage; Thrombosis/Thrombus; Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE (VAD) PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH IS UNKNOWN. THERE WERE NO REPORTED PATIENT SYMPTOMS OR COMPLICATIONS ASSOCIATED WITH THIS EVENT. THE DEVICE IS IMPLANTED AND IS NOW OUT OF SERVICE.Event Description: IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS DUE TO THE PATIENT HAD BASILAR TIP THROMBUS. STATUS POST (S/P) ATTEMPTED THROMBECTOMY AND DEVELOPED INTRACEREBRAL HEMORRHAGE (ICH) WHICH LED TO CATASTROPHIC SUBARACHNOID HEMORRHAGE (SAH) WITH BRAIN STEM COMPRESSION. INTERNATIONAL NORMALIZED RATIO (INR) REVERSED BUT WAS NOT ABLE TO CONTROL BLEEDING. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINED IMPLANTED. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO COMPLICATIONS FROM A THROMBECTOMY FOR BASILAR TIP THROMBUS. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, BLEEDING, NEUROLOGICAL DYSFUNCTION, THROMBUS, AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION WAS RECEIVED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22018358&pc=OZD | 1220648-2025-28389 | 2025-04-23 04:00 | Death | ABIOMED, INC | 2025-05-14 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXHIBITED BLEEDING FROM THE FEMORAL SITE. THE CARE TEAM APPLIED PRESSURE AND STOPPED HEPARIN. THE FAMILY MADE CHOICE TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: A4 AND A5, RACE, ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22018427&pc=DYB | 2017865-2025-53699 | 2025-04-24 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-14 04:00 | DYB | AVEIR¿ INTRODUCER KIT ¿ 50 CM | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K211989 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC FOR AVEIR LEADLESS PACEMAKER SYSTEM IMPLANT. DURING POST-OPERATION RECOVERY, IT WAS NOTED THAT THE PATIENT SUFFERED EXCESSIVE BLEEDING FROM THE INTRODUCER PUNCTURE SITE. THE PATIENT DECEASED DUE TO BLOOD LOSS. Manufacturer Narrative: CORRECTIONS: REPORT TYPE, REPORT SUBTYPE, AND H1.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22018646&pc=FTL | 1213643-2025-090207 | 2016-11-22 05:00 | Death | DAVOL INC., SUB. C.R. BARD, INC. -1213643 | 2025-05-14 04:00 | FTL | PERFIX PLUG | Defective Device; Patient Device Interaction Problem | Abscess; Adhesion(s); Unspecified Infection; Impaired Healing | K922916 |  | 1 | Event Description: PER INFORMATION PROVIDED BY PATIENT'S ATTORNEY: (B)(6) 2011 - PATIENT WAS DIAGNOSED WITH INCISIONAL UMBILICAL HERNIA THEREBY UNDERWENT REPAIR WITH IMPLANT OF NON-BARD/BD PROCEED MESH. PER OP NOTES, ¿THE UMBILICAL DEFECT WAS CLOSED THEN SYNTHETIC MESH WAS PLACED, SECURED USING SUTURES AND TACKS.¿ (B)(6) 2011 - PATIENT UNDERWENT WOUND EXPLORATION AND REMOVAL OF INFECTED MESH. PER OP NOTES, ¿THE PURULENT FRAGMENTED MESH THAT BECAME LOOSE IN SUBCUTANEOUS TISSUES WAS REMOVED.¿ (B)(6) 2016 - PATIENT WAS DIAGNOSED WITH PANTALOON RIGHT DIRECT AND INDIRECT INGUINAL HERNIAS THEREBY UNDERWENT OPEN REPAIR WITH IMPLANT OF TWO PERFIX PLUG (DEVICE #1 & #2). PER OP NOTES, ¿THE PANTALOON RIGHT DIRECT AND INDIRECT INGUINAL HERNIAS WAS REPAIRED USING TWO PERFIX PLUG (DEVICE #1 & DEVICE #2). AN OVERLAY PATCH AS WELL AS PLACED AND SECURED USING SUTURES.¿ (B)(6) 2016 - PATIENT WAS DIAGNOSED WITH CHRONIC NON-HEALING WOUND AND INFECTED MESH THEREBY UNDERWENT REPAIR WITH MESH EXPLANT. PER OP NOTES, ¿THE PURULENT MATERIAL FROM CHRONIC GROIN ABSCESS CAVITY WAS ASPIRATED. THE 2 MESH PLUGS WITH COVERING MESH (DEVICE #1 & #2) ADHERENT TO SURROUNDING STRUCTURE WERE REMOVED.¿ Manufacturer Narrative: ROOT CAUSE IS INCONCLUSIVE; NO CONCLUSION CAN BE MADE AS TO THE EXTENT TO WHICH THE IMPLANT MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S POSTOPERATIVE COURSE. ATTORNEY ALLEGES WRONGFUL DEATH OF THE PATIENT. NO INFORMATION WAS PROVIDED REGARDING PATIENT¿S DEATH IN THE MEDICAL RECORDS PROVIDED TO DATE. NO AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. NOTE, THE MANUFACTURING DATE (15-SEP-2015) IS CONSIDERED TO BE A BEST ESTIMATE. THIS EMDR REPRESENTS THE PERFIX PLUG (DEVICE #2). ADDITIONAL SUPPLEMENTAL EMDR'S WERE SUBMITTED TO REPRESENT THE VENTRALEX AND PERFIX PLUG (DEVICE #1). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22018651&pc=DSQ | 2916596-2025-02934 | 2025-01-23 05:00 | Death | THORATEC CORPORATION | 2025-05-14 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DIE TO SYSTEMIC INFECTION. THE SOURCE OF INFECTION WAS NOT PROVIDED. THE PATIENT WAS TREATED WITH THE STANDARD INFECTION TREATMENT USED BY THE HOSPITAL. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED AND THE DEVICE OPERATED AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22018838&pc=DSQ | 2916596-2025-02910 | 2022-10-01 04:00 | Death | THORATEC CORPORATION | 2025-05-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RETROSPECTIVE COHORT STUDY "IMPACT OF CARDIAC RESYNCHRONIZATION THERAPY ON VENTRICULAR ARRHYTHMIAS AND SURVIVAL AFTER DURABLE LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION" MAY BE ASSOCIATED WITH VENTRICULAR ARRHYTHMIAS, INFECTION, HEMATOMA, RESPIRATORY FAILURE, RIGHT VENTRICULAR ASSIST DEVICE IMPLANTATION, RIGHT VENTRICLE FAILURE, AND DEATH. THIS SINGLE-CENTER RETROSPECTIVE COHORT STUDY EVALUATED THE IMPACT OF POSTOPERATIVE MANAGEMENT OF CARDIAC RESYNCHRONIZATION THERAPY (CRT) ON CLINICAL OUTCOMES FOLLOWING LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION. ALL PATIENTS 18 YEARS OLD AND ABOVE WHO WERE IMPLANTED AT THE STUDY HOSPITAL WERE CONSIDERED FOR THE STUDY. A TOTAL OF 789 PATIENTS UNDERWENT LVAD IMPLANTATION, FROM 01OCT2006 TO 01OCT2022, INCLUDING 195 PATIENTS (24.7%) WITH PREOPERATIVE CRT. THE HEARTMATE 3 (HM3) WAS THE MOST COMMON DEVICE IMPLANTED (41.4%), WHEREAS 34.6% OF PATIENTS RECEIVED A HEART WARE VENTRICULAR ASSIST DEVICE (HVAD) DEVICE AND 23.8% OF PATIENTS RECEIVED A HEARTMATE II (HMII). OF NOTE, PATIENTS WITH PREOPERATIVE CRT IN WHOM CHARACTERIZATION OF POSTOPERATIVE DEVICE MANAGEMENT WAS NOT POSSIBLE WERE EXCLUDED FROM ANALYSIS. THE MEAN PATIENT AGE AT LVAD IMPLANTATION WAS 62.9 YEARS. THE MAJORITY OF PATIENTS WERE MEN (73.3%). AFTER LVAD IMPLANTATION, 42 HM3 PATIENTS HAD CRT PROGRAMMED ¿OFF¿ AND 74 HM3 PATIENTS HAD CRT PROGRAMMED ¿ON.¿ PATIENTS WITH PREOPERATIVE CRT WHO DID NOT HAVE DATA DETAILING POSTOPERATIVE CRT PROGRAMMING WERE EXCLUDED FROM THE ANALYSIS (N = 36). MAINTAINING CRT ¿ON¿ WAS ASSOCIATED WITH A SIGNIFICANTLY GREATER RISK OF MORTALITY COMPARED TO PROGRAMMING CRT ¿OFF¿ AFTER LVAD IMPLANTATION AFTER ADJUSTING FOR AGE, SEX, AND RACE (SDHR = 1.54; 1.06¿2.37 95% CI; P = 0.036). PATIENTS WITH CRT ¿OFF¿ HAD SIGNIFICANTLY LONGER MEDIAN SURVIVAL TIME ON LVAD SUPPORT COMPARED TO PATIENTS WITH CRT ¿ON¿ (970 DAYS [462¿1,821] VS. 496 DAYS [208¿953]; P = 0.001). HEART TRANSPLANTATION OCCURRED EQUALLY IN PATIENTS WITH CRT ¿ON¿ AND CRT ¿OFF¿ (13.5% VS. 14.1%; P = 0.613) AFTER SIMILAR WAIT TIMES (659 DAYS [386¿1,406] VS. 548 DAYS [247¿698]; 0.375) (THERE WAS NO SIGNIFICANT DIFFERENCE BETWEEN INCIDENCE OF VENTRICULAR ARRHYTHMIAS IN PATIENTS WITH AND WITHOUT POSTOPERATIVE CRT ¿ON¿ (35.1% VS. 48.2%; P = 0.095). OF THE 85 PATIENTS WITH CRT TURNED ¿OFF¿ FOLLOWING LVAD IMPLANTATION, 18 PATIENTS WERE DEPENDENT ON RV-ONLY PACING AND RV PACED >95% OF THE TIME. COMPARED TO PATIENTS WITH RV-ONLY PACING, BIVENTRICULAR PACING (LV LEAD ¿ON¿) WAS ASSOCIATED WITH A HIGHER RISK OF MORTALITY AFTER ADJUSTING FOR AGE, SEX, AND RACE (SDHR = 2.61; 1.19¿5.75 95% CI; P = 0.017). THE PRIMARY CLINICAL OUTCOMES OF THE STUDY WERE TRANSPLANT FREE SURVIVAL AND INCIDENCE OF POSTOPERATIVE VENTRICULAR ARRHYTHMIAS. TRANSPLANT-FREE SURVIVAL WAS DEFINED AS TIME FROM INDEX LVAD IMPLANTATION TO TRANSPLANTATION, LVAD EXPLANT, DEATH, OR MOST RECENT FOLLOW-UP. IT WAS NOTED THAT IN TOTAL 67 PATIENTS EXPERIENCED A VENTRICULAR ARRHYTHMIA, 26 PATIENTS WERE TRANSPLANTED, 80 PATIENTS PASSED AWAY, OF WHICH 35 PATIENTS PASSED AWAY DUE TO AN UNSPECIFIED CARDIOVASCULAR OUTCOME, WHILE 45 PATIENTS PASSED AWAY DUE TO AN UNSPECIFIED NON-CARDIOVASCULAR OUTCOME. IT WAS NOTED THAT AN UNKNOWN NUMBER OF PATIENTS EXPERIENCED RESPIRATORY FAILURE OR REQUIRED RIGHT VENTRICULAR ASSIST DEVICES. ADDITIONALLY, IT WAS NOTED THAT 25.5% OF PATIENTS IN THE CRT OFF CATEGORY HAD COMPLETE HEART BLOCK. PROCEDURE RELATED HEMATOMA¿S, INFECTION, AND RIGHT VENTRICLE FAILURE WERE ALSO REPORTED TO HAVE OCCURRED, THOUGH THE AMOUNT OF THESE EVENTS WAS NOT PROVIDED, HOWEVER IT WAS REPORTED AT LEAST ONE PATIENT REQUIRED A WASHOUT DUE TO A HEMATOMA, AND ANOTHER PATIENT EXPERIENCED A BLOOD STREAM INFECTION. FINALLY, IT WAS REPORTED THAT SOME PATIENTS EXPERIENCED POSTOPERATIVE SHOCK AND REQUIRED VASOPRESSOR MEDICATIONS. IN CONCLUSION, MAINTAINING CRT ¿ON¿ AFTER LVAD IMPLANTATION, AND POSTOPERATIVE MANAGEMENT OF CRT WAS NOT ASSOCIATED WITH INCIDENCE OF VENTRICULAR ARRHYTHMIAS POST IMPLANT, BUT WAS ASSOCIATED WITH INCREASED RISK OF MORTALITY FOLLOWING LVAD IMPLANTATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO HEARTMATE 3 LVAD¿S WERE RETURNED FOR THIS EVALUATION. THE DEVICE SERIAL NUMBERS WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B2: SPECIFIC DEATH DATE IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01OCT2022 AS PATIENTS WERE IMPLANTED BETWEEN 01OCT2006 AND 01OCT2022. OATES, C.P.; LAWRENCE, L.L.; BIGHAM, G.E.; MEDA, N.S.; BASYAL, B; ET AL (2025). IMPACT OF CARDIAC RESYNCHRONIZATION THERAPY ON VENTRICULAR ARRHYTHMIAS AND SURVIVAL AFTER DURABLE LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION. ASAIO JOURNAL, 71(2), 157-163. HTTP://DX.DOI.ORG/10.1097/MAT.0000000000002279. THE MEDSTAR HEART AND VASCULAR INSTITUTE, MEDSTAR WASHINGTON HOSPITAL CENTER, WASHINGTON, DC, UNITED STATES, GEORGETOWN UNIVERSITY SCHOOL OF MEDICINE, WASHINGTON, DC, UNITED STATES, DEPARTMENT OF CARDIOLOGY, HELMSLEY ELECTROPHYSIOLOGY CENTER, ICAHN SCHOOL OF MEDICINE AT MOUNT SINAI, NEW YORK, NY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22018875&pc=DSQ | 2916596-2025-02952 | 2025-04-13 04:00 | Death | THORATEC CORPORATION | 2025-05-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE DEVICE WOULD NOT BE EXPLANTED AND RETURNED FOR EVALUATION. Manufacturer Narrative: SECTION B3: DATE OF DEATH CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22018977&pc=DSQ | 2916596-2025-02936 | 2025-04-05 04:00 | Death | THORATEC CORPORATION | 2025-05-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO NATURAL CAUSES, DEMENTIA ON HOSPICE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR DEVICE THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED AND THE DEICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22018985&pc=KDI | 3003464075-2025-00071 | 2025-04-09 04:00 | Death | NXSTAGE MEDICAL, INC. | 2025-05-14 04:00 | KDI | NXSTAGE SYSTEM ONE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K230632 |  | 1 | Event Description: A REPORT WAS RECEIVED ON (B)(6) 2025 FROM THE HOME THERAPY NURSE (HTN) OF A 56-YEAR-OLD MALE PATIENT WITH A MEDICAL HISTORY INCLUDING CONGESTIVE HEART FAILURE, MULTIPLE MYELOMA, NEOPLASM/TUMORS AND END STAGE RENAL DISEASE, WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2025. ADDITIONAL INFORMATION WAS RECEIVED ON 17 APR 2025 FROM THE HTN WHO STATED CAUSE OF DEATH IS UNKNOWN, NO FURTHER DETAILS WERE PROVIDED. Manufacturer Narrative: THERE WAS NO DEVICE MALFUNCTION. THE CYCLER WAS RECEIVED FOR EVALUATION AND SUCCESSFULLY PASSED TESTING. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. AVAILABLE LOG FILES WERE RETRIEVED AND ANALYZED WHICH SHOWED DEVICE PERFORMANCE WAS WITHOUT DEFICIENCY AND WAS UNREMARKABLE. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22019293&pc=OUR | 3012495575-2025-00003 | 2025-04-11 04:00 | Death | SURGENTEC, LLC | 2025-05-14 04:00 | OUR | TILINK-P SCREW 12MM X 25MM | No Apparent Adverse Event | Hemorrhage/Bleeding | K242141 |  | 1 | Event Description: THE CASE WAS ON (B)(6) 2025 AT 7:30 AM. REP STATED THAT IT WAS A ROUTINE CASE. REP ALSO STATED THERE WAS MINOR BLEEDING THAT OCCURRED DURING THE SURGERY, AND THEY HAD IT UNDER CONTROL IN LESS THAN A MINUTE. REP ALSO STATED THAT NO SURGENTEC INSTRUMENTS WERE FAULTY OR HAD ANY ISSUES PRIOR TO OR DURING THE CASE. THE PATIENT DID NOT STAY UNDER ANESTHESIA LONGER THAN ORIGINALLY ANTICIPATED. THE REP THEN STATED THAT THE PATIENT STOOD UP 6.5 HOURS POST-OP AND EXPERIENCED A VASCULAR ISSUE. THE PATIENT THEN ULTIMATELY PASSED AWAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22019564&pc=DSQ | 2916596-2025-02818 | 2025-04-01 04:00 | Death | THORATEC CORPORATION | 2025-05-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Tachycardia; Cardiogenic Shock; Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ABSTRACT TITLED ¿A CASE OF FATAL RECALCITRANT POLYMORPHIC VENTRICULAR TACHYCARDIA STORM POSTLEFT VENTRICULAR ASSIST DEVICE IMPLANTATION: LESSONS LEARNED¿ IDENTIFYING THAT HEARTMATE 3 (HM3) MAY BE RELATED TO COMPLICATED POST-OPERATIVE COURSE WITH CARDIOGENIC SHOCK, AND ACUTE KIDNEY INJURY, AND DEATH. THIS IS A CASE STUDY OF A 73-YEAR-OLD PATIENT WITH STAGE 3B CHRONIC KIDNEY DISEASE, SOLITARY KIDNEY, CORONARY ARTERY DISEASE WITH BYPASS AND ANGIOPLASTY; AND STAGE D HEART FAILURE UNDERWENT SUCCESSFUL HM3 LEFT VENTRICULAR ASSIST DEVICE (LVAD) PLACEMENT. THE PATIENT'S POST-OPERATIVE COURSE WAS COMPLICATED BY CARDIOGENIC SHOCK AND ACUTE KIDNEY INJURY DUE TO CARDIO-RENAL SYNDROME REQUIRING CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT). ON DAY 26, THE PATIENT BEGAN MANIFESTING RECURRENT EPISODES OF NEW ONSET SUSTAINED VENTRICULAR TACHYCARDIA (VT) REQUIRING MULTIPLE DIRECT CURRENT CARDIOVERSIONS. PATIENT FAILED AMIODARONE. MORE THAN 10GM LOAD WAS ALREADY GIVEN. RESULTANT QT PROLONGATION LED TO LIDOCAINE INFUSION¿THIS WAS STOPPED DUE TO SEIZURES. MEXILETINE AND QUINIDINE DID NOT SUPPRESS VT. RIGHT VENTRICULAR FUNCTION WAS SUPPORTED WITH INHALED EPOPROSTENOL AND LOW DOSE DOBUTAMINE. GIVEN INCESSANT VT, THE PATIENT WAS INTUBATED AND DEEPLY SEDATED FOR SYMPATHETIC DRIVE SUPPRESSION. ECHOCARDIOGRAM RULED OUT CANNULA MALPOSITION¿LVAD SPEED ADJUSTED TO PREVENT SUCTION EVENTS. LEFT STELLATE GANGLION BLOCK FAILED. EVEN AFTER SUBSTRATE MODIFICATION TARGETING AREAS OF FOCAL ABNORMAL VENTRICULAR ACTIVATION IN LEFT VENTRICLE, INCLUDING DISTAL PURKINJE AND INFERO-SEPTAL REGIONS, INCESSANT VT CONTINUED. OVERDRIVE PACING TO SUPPRESS TRIGGERING PREMATURE VENTRICULAR CONTRACTIONS AFTER DUAL CHAMBER EXTERNALIZED PACEMAKER IMPLANTATION WAS UNSUCCESSFUL. AFTER FAMILY DISCUSSION, COMFORT CARE WAS INITIATED AND THE PATIENT PASSED AWAY.Event Description: ADDITIONAL INFORMATION RECEIVED STATED THAT THE REPORTED EVENTS WERE NOT DEVICE RELATED. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS SAME AS PUBLISHED DATE, 01APR2025, SINCE THE DATE OF IMPLANT OR DATA COLLECTION WAS NOT PROVIDED. AUTHOR INFORMATION: DHRUVI PATEL, AND AMIT BADIYE. JOURNAL OF THE AMERICAN COLLEGE OF CARDIOLOGY, SUPPL. SUPPLEMENT 85.12: 4027. HTTP://DX.DOI.ORG/10.1016/S0735-1097(25)04511-5. EASTERN VIRGINIA MEDICAL SCHOOL, NORFOLK, VA, USA. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: CORRECTED DATA: H6 - RESPIRATORY FAILURE CODE REMOVED. SECTION E1: CUSTOMER SITES WERE (B)(6). MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6, "PATIENT CARE AND MANAGEMENT", LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22019744&pc=BZD | 2518422-2025-105605 | 2025-04-08 04:00 | Death | RESPIRONICS, INC. | 2025-05-14 04:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED THAT A REP DREAMSTATION AUTO CPAP, DOM - RECRT DEVICE PATIENT HAS PASSED AWAY. IT WAS NOT SUGGESTED THAT THE DEVICE CAUSED THE DEATH. THE PATIENT'S WIFE CALLED REGARDING A LETTER SHE HAD RECEIVED REQUESTING THE RETURN OF HIS RECALLED DEVICE. SHE REPORTED THAT THE PATIENT HAD PASSED AWAY AT THE BEGINNING OF FEBRUARY. SHE INDICATED THAT THE RECALLED DEVICE HAD BEEN DISCARDED, BUT THAT SHE STILL POSSESSED HIS REPLACEMENT DEVICE AND WISHED TO RETURN IT. MORE DETAILS ABOUT THE REPORTED DEATH ARE NEEDED. THE MANUFACTURER'S INVESTIGATION IS CURRENTLY ONGOING, AND ONCE IT IS COMPLETED, A FOLLOW-UP REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22019927&pc=DSQ | 2916596-2025-03179 | 2025-04-29 04:00 | Death | THORATEC CORPORATION | 2025-05-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: SECTION A: PATIENT AGE/DATE OF BIRTH AND WEIGHT COULD NOT BE PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS UNKNOWN WHETHER THE PUMP WILL BE EXPLANTED OR RETURNED. IT WAS COMMUNICATED THAT DUE TO HOSPITAL POLICY, NO ADDITIONAL INFORMATION WILL BE PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22019935&pc=MNT | 2518422-2025-105612 | 2024-11-01 04:00 | Death | RESPIRONICS, INC. | 2025-05-14 04:00 | MNT | BIPAP A40 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K121623 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY IN (B)(6) 2024. THERE WAS NO INFORMATION PROVIDED AS TO THE CAUSE OF THE DEATH. THE DEVICE WAS DISPOSED OF BY THE PATIENT'S FAMILY AND IS UNABLE TO BE RETURNED. A CLINICAL ASSESSMENT WAS PERFORMED AND THE WAS ASSESSED AS FOLLOWS. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED. BASED ON AVAILABLE INFORMATION AT THIS TIME, THE ALLEGED HARM DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22020199&pc=DSQ | 3007042319-2025-01146 | 2025-01-01 05:00 | Death | HEARTWARE, INC. | 2025-05-14 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Obstruction of Flow; Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke; Ischemia Stroke; Thrombosis/Thrombus | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LEFT VENTRICULAR ASSIST DEVICE (VAD) PATIENTS WHO UNDERWENT INVASIVE PERCUTANEOUS OUTFLOW GRAFT STUDIES. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN, BUT SOME OF THE PATIENTS HAD UNDERGONE OUTFLOW GRAFT STENTING FOR STENOSIS AND ONE PATIENT HAD EXPERIENCED AN ISCHEMIC STROKE AND PUMP THROMBOSIS. PATIENTS WHO UNDERWENT THE STUDY WERE ADMITTED FOR REPETITIVE SUCTION EVENTS, HIGH POWER OR LOW FLOW ALARMS, VOLUME OVERLOAD, ACUTE CORONARY SYNDROME POST VAD IMPLANTATION, AND CONCERNS OF STROKE. THERE WERE PATIENTS WHO EXPERIENCED ISCHEMIC OR HEMORRHAGIC STROKES. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/61 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: PERCUTANEOUS ASSESSMENT AND INTERVENTION OF OUTFLOW GRAFT STENOSIS IN LEFT VENTRICULAR ASSIST DEVICE PATIENTS¿A COHORT STUDY. ASAIO JOURNAL. 2025. 1-8. DOI: 10.1097/MAT.0000000000002417 ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ OUTFLOW GRAFT D4: MODEL #: UNK / CATALOG #: UNK / EXPIRATION DATE: UNK / SERIAL OR LOT#: UNK D10: NO H4: MFG DATE: UNK H5: YES D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22020220&pc=NPT | 2025587-2025-03667 | 2025-04-22 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-14 04:00 | NPT | EVOLUT FX DCS | Difficult to Advance; Insufficient Information; Patient Device Interaction Problem | Ischemia; Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF A TRANSCATHETER VALVE, THE DELIVERY CATHETER SYSTEM (DCS) WAS APPROACHED AT AN INLINE, BUT IT WAS DIFFICULT TO PASS THE DCS THROUGH THE AORTIC ARCH. SUBSEQUENTLY, AN 18 FRENCH (FR) LONG SHEATH WAS USED AND IT WAS DECIDED TO STOP ADVANCEMENT AT THE DESCENDING AORTA RATHER THAN ADVANCING THROUGH THE ARCH. UPON ADVANCING THE ARCH, THERE WAS RESISTANCE FELT AND THE ORIENTATION OF THE DCS WAS ADJUSTED AND THE VALVE WAS ABLE TO BE ADVANCED PASSED THE ARCH. FOLLOWING THE IMPLANT OF THE VALVE, A DIFFERENCE IN BLOOD PRESSURE BETWEEN THE ASCENDING AORTA AND LOWER LIMBS WAS OBSERVED, AND A DISSECTION WAS SUSPECTED. AN INTRAVASCULAR ULTRASOUND (IVUS) AND CONTRAST IMAGING WERE PERFORMED THAT SHOWED A TYPE B DISSECTION EXTENDING FROM LEFT SUBCLAVIAN ARTERY TO THE DISTAL ENTRY POINT. IT WAS REPORTED THAT THE PATIENT'S ABDOMINAL BLOOD FLOW WAS NORMAL, AND BLOOD FLOW TO THE LUMEN WAS "SECURED", THEREFORE, THE DISSECTION WAS MONITORED. FOLLOWING THE WITHDRAWAL OF THE SHEATH, AN ANGIOGRAM WAS PERFORMED THAT SHOWED A MILD INTIMAL FLAP DISSECTION AT THE BIFURCATION OF THE COMMON ILIAC ARTERY (CIA) FROM THE ABDOMINAL AORTA, BUT THERE WERE NO ISSUES WITH THE PATIENT'S BLOOD FLOW.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF THE DISSECTION WAS DUE TO CALCIFICATION IN THE AORTIC ARCH AND INTERFERENCE WITH THE DCS. PER THE PHYSICIAN, THE DCS COULD NOT BE RULED OUT AS A CONTRIBUTING FACTOR TO THE DISSECTION. 1 DAY FOLLOWING THE IMPLANT OF THE VALVE, A THORACIC ENDOVASCULAR AORTIC REPAIR (TEVAR) WAS PERFORMED. SUBSEQUENTLY, NON-OCCLUSIVE MESENTERIC ISCHEMIA (NOMI) OCCURRED. 16 DAYS FOLLOWING THE IMPLANT OF THE VALVE, TOTAL RESECTION WAS PERFORMED, AND IT WAS NOTED THAT THE PATIENT HAD A DNR (DO NOT RESUSCITATE), AND SUBSEQUENTLY DIED. THE CAUSE OF DEATH WAS NOMI. PER THE PHYSICIAN, THE VALVE AND THE VALVE IMPLANT PROCEDURE DID NOT CAUSE OR CONTRIBUTE TO THE DEATH. Manufacturer Narrative: UPDATED DATA: B2. B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. CONTINUATION OF D10: PRODUCT ID {EVOLUTFX-26}; PRODUCT LOT/SERIAL NUMBER {B)(6}; PRODUCT TYPE: {0195-HEARTVALVES}}; IMPLANT DATE {(B)(6) 2025}; EXPLANT DATE {N/A}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22020238&pc=DSQ | 3007042319-2025-01148 | 2015-03-19 04:00 | Death | HEARTWARE, INC. | 2025-05-14 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT DIED DUE TO UNKNOWN CAUSES. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINS IMPLANTED. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS WAS NOT PERFORMED SINCE LOG FILES WERE NOT AVAILABLE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH IS UNKNOWN. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22020561&pc=NIQ | 2124215-2025-31052 | 2025-05-08 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-14 04:00 | NIQ | PROMUS PREMIER | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH CHEST PAIN AND DISCOMFORT. DIAGNOSTIC TESTING RESULTED IN A DIAGNOSIS OF UNSTABLE ANGINA AND ANGIOGRAPHY REVEALED 95% STENOSIS OF THE LEFT ANTERIOR DESCENDING ARTERY (LAD). DURING PERCUTANEOUS CORONARY INTERVENTION (PCI), THE TARGET LESION WAS CROSSED WITH A NON-BSC GUIDEWIRE AND PRE-DILATED WITH A 2.75 X 15 MM BALLOON. 2.75 X 38 MM AND 3.00 X 20 MM PROMUS PREMIER STENTS WERE DEPLOYED AND THE LESION WAS THEN POST-DILATED WITH A 2.75 X 15 MM NON-COMPLIANT BALLOON. THE PATIENT WAS TRANSFERRED TO THE ICU FOR MEDICAL FOLLOW UP. WITHIN 24 HOURS OF THE PCI, THE PATIENT DEVELOPED VENTRICULAR TACHYCARDIA. DEFIBRILLATION WAS PERFORMED AND MEDICAL THERAPY WAS INITIATED, BUT THE PATIENT WENT INTO CARDIAC ARREST AND DIED.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH CHEST PAIN AND DISCOMFORT. DIAGNOSTIC TESTING RESULTED IN A DIAGNOSIS OF UNSTABLE ANGINA AND ANGIOGRAPHY REVEALED 95% STENOSIS OF THE LEFT ANTERIOR DESCENDING ARTERY (LAD). DURING PERCUTANEOUS CORONARY INTERVENTION (PCI), THE TARGET LESION WAS CROSSED WITH A NON-BSC GUIDEWIRE AND PRE-DILATED WITH A 2.75 X 15 MM BALLOON. 2.75 X 38 MM AND 3.00 X 20 MM PROMUS PREMIER STENTS WERE DEPLOYED AND THE LESION WAS THEN POST-DILATED WITH A 2.75 X 15 MM NON-COMPLIANT BALLOON. THE PATIENT WAS TRANSFERRED TO THE ICU FOR MEDICAL FOLLOW UP. WITHIN 24 HOURS OF THE PCI, THE PATIENT DEVELOPED VENTRICULAR TACHYCARDIA. DEFIBRILLATION WAS PERFORMED AND MEDICAL THERAPY WAS INITIATED, BUT THE PATIENT WENT INTO CARDIAC ARREST AND DIED. Manufacturer Narrative: CORRECTED H6 IMPACT CODES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22020563&pc=DSQ | 2916596-2025-02973 | 2025-04-22 04:00 | Death | THORATEC CORPORATION | 2025-05-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hypoxia; Respiratory Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS HYPOXIC RESPIRATORY FAILURE. THE PATIENT WAS ON HOSPICE BECAUSE OF CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD) AND END STAGE HEART FAILURE (HF). THE PATIENT'S FAMILY AND MEDICAL TEAM MUTUALLY DECIDED BASED ON CLINICAL OUTLOOK AND QUALITY OF LIFE THAT HOSPICE WAS THE BEST OPTION. THE PUMP OPERATED AS EXPECTED. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS REPORTED THAT HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATION FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22020570&pc=NIQ | 2124215-2025-31047 | 2025-05-08 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-14 04:00 | NIQ | PROMUS PREMIER | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH CHEST PAIN AND DISCOMFORT. DIAGNOSTIC TESTING RESULTED IN A DIAGNOSIS OF UNSTABLE ANGINA AND ANGIOGRAPHY REVEALED 95% STENOSIS OF THE LEFT ANTERIOR DESCENDING ARTERY (LAD). DURING PERCUTANEOUS CORONARY INTERVENTION (PCI), THE TARGET LESION WAS CROSSED WITH A NON-BSC GUIDEWIRE AND PRE-DILATED WITH A 2.75 X 15 MM BALLOON. 2.75 X 38 MM AND 3.00 X 20 MM PROMUS PREMIER STENTS WERE DEPLOYED AND THE LESION WAS THEN POST-DILATED WITH A 2.75 X 15 MM NON-COMPLIANT BALLOON. THE PATIENT WAS TRANSFERRED TO THE ICU FOR MEDICAL FOLLOW UP. WITHIN 24 HOURS OF THE PCI, THE PATIENT DEVELOPED VENTRICULAR TACHYCARDIA. DEFIBRILLATION WAS PERFORMED AND MEDICAL THERAPY WAS INITIATED, BUT THE PATIENT WENT INTO CARDIAC ARREST AND DIED.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH CHEST PAIN AND DISCOMFORT. DIAGNOSTIC TESTING RESULTED IN A DIAGNOSIS OF UNSTABLE ANGINA AND ANGIOGRAPHY REVEALED 95% STENOSIS OF THE LEFT ANTERIOR DESCENDING ARTERY (LAD). DURING PERCUTANEOUS CORONARY INTERVENTION (PCI), THE TARGET LESION WAS CROSSED WITH A NON-BSC GUIDEWIRE AND PRE-DILATED WITH A 2.75 X 15 MM BALLOON. 2.75 X 38 MM AND 3.00 X 20 MM PROMUS PREMIER STENTS WERE DEPLOYED AND THE LESION WAS THEN POST-DILATED WITH A 2.75 X 15 MM NON-COMPLIANT BALLOON. THE PATIENT WAS TRANSFERRED TO THE ICU FOR MEDICAL FOLLOW UP. WITHIN 24 HOURS OF THE PCI, THE PATIENT DEVELOPED VENTRICULAR TACHYCARDIA. DEFIBRILLATION WAS PERFORMED AND MEDICAL THERAPY WAS INITIATED, BUT THE PATIENT WENT INTO CARDIAC ARREST AND DIED. Manufacturer Narrative: CORRECTED H6 IMPACT CODES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22020847&pc=DSQ | 3007042319-2025-01149 | 2025-01-01 05:00 | Death | HEARTWARE, INC. | 2025-05-14 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING LEFT VENTRICULAR ASSIST DEVICE (VAD) PATIENTS WITH SLEEVE GASTRECTOMY COMPARED TO THOSE WITH A VAD ONLY. THE AUTHORS DESCRIBED PATIENT DEATHS; THE CAUSES OF DEATH WERE FATAL BLEEDING EVENTS AND OTHER UNKNOWN CAUSES WHICH OCCURRED WITHIN THIRTY DAYS OF IMPLANT. THERE WERE PATIENTS WHO EXPERIENCED MAJOR AND MINOR BLEEDING EVENTS, AND THROMBOEMBOLIC EVENTS WHICH INCLUDED GASTROINTESTINAL BLEED (GIB), SUBTHERAPEUTIC INTERNATIONAL NORMALIZED RATIO (INR), INTRACRANIAL HEMORRHAGES, VENOUS THROMBOEMBOLISMS, ARTERIAL THROMBOEMBOLISMS, AND PUMP THROMBOSIS. MAJOR BLEEDING EVENTS REQUIRED TRANSFUSIONS OF TWO OR MORE UNITS OF PACKED RED BLOOD CELLS (PRBC). THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/45 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: EVALUATION OF WARFARIN DOSE AND INR TIME IN THERAPEUTIC RANGE IN LEFT VENTRICULAR ASSIST DEVICE PATIENTS WITH SLEEVE GASTRECTOMY. THE INTERNATIONAL JOURNAL OF ARTIFICIAL ORGANS. 2025. VOL. 48(2) 77¿83. DOI: 10.1177/03913988251317820. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE NECESSARY INFORMATION IS UNAVAILABLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22021075&pc=DSQ | 2916596-2025-03166 | 2025-04-22 04:00 | Death | THORATEC CORPORATION | 2025-05-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE (MSOF). THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND THE DEVICE WAS NOT EXPLANTED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS REPORTED THAT HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22021129&pc=DSQ | 2916596-2025-02774 | 2025-04-17 04:00 | Death | THORATEC CORPORATION | 2025-05-14 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22021240&pc=DSQ | 2916596-2025-03010 | 2025-04-28 04:00 | Death | THORATEC CORPORATION | 2025-05-14 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED BY THE CUSTOMER. THE CUSTOMER DID NOT PROVIDE INFORMATION LEADING TO THE PATIENT DEATH. IT WAS REPORTED THAT THE PATIENT'S DEATH WAS NOT DEVICE OR THERAPY RELATED. IT WAS REPORTED THAT THE DEVICE WORKED AS EXPECTED. THE DEVICE WAS NOT BE EXPLANTED AND WAS NOT RETURNED FOR EVALUATION. THE CUSTOMER DID NOT PROVIDE ANY INFORMATION LEADING TO THE DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6) AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22021265&pc=NPT | 2025587-2025-03677 | 2025-04-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-14 04:00 | NPT | EVOLUT FX DCS | Difficult to Advance; Patient Device Interaction Problem |  | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-34 VALVE, THE VALVE WAS DEPLOYED LOW DURING THE FIRST DEPLOYMENT. THE VALVE WAS RECAPTURED AND DEPLOYED A SECOND TIME WHEN AN INFOLDING OF THE VALVE WAS OBSERVED WHICH NECESSITATED THE REMOVAL AND REPLACEMENT WITH ANOTHER FX34 VALVE. THE PATIENT, WHO WAS DESCRIBED AS VERY FRAGILE, UNDERWENT SURGERY ON A TABLE SHORTLY AFTER THE SECOND VALVE WAS IMPLANTED. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL ISSUE DUE TO AN AORTIC DISSECTION CAUSED BY THE PROCEDURE. UNFORTUNATELY, THE PATIENT DID NOT SUPPORT THE SURGERY AND DIED DURING THE NIGHT. ADDITIONAL INFORMATION WAS RECEIVED WHICH CONFIRMED THAT THE PREVIOUSLY REPORTED PERICARDIAL ISSUE WAS A PERICARDIAL EFFUSION. THE TIMING OF THE AORTIC DISSECTION WAS UNCERTAIN AS IT WAS NOT NOTED UNTIL THE FINAL ANGIOGRAM. PER THE PHYSICIAN, THE CAUSE OF THE DISSECTION WAS THE PATIENT'S AORTIC ANGULATION OF 27° ALONG WITH DIFFICULTY ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) IN THE AORTIC ARCH AND THE ASCENDING AORTA. ADDITIONALLY, THE FIRST AND THE SECOND DCS REQUIRED PRESSURE TO ADVANCE. FOR THE SECOND VALVE IMPLANT, A NON-MEDTRONIC GUIDEWIRE WAS PUT IN THE PIGTAIL CATHETER TO HELP THE VALVE CROSS THE AORTIC ARCH. INITIALLY, DRAINAGE WAS INTENDED FOR THE PERICARDIAL EFFUSION, BUT THEN WAS CHANGED TO SURGICAL INTERVENTION. THE PATIENT'S TORTUOUS ANATOMY AND CALCIFIED ANNULUS CONTRIBUTED TO THE DISSECTION. PER THE PHYSICIAN, THE PATIENT'S UNDERLYING MEDICAL HISTORY INCLUDING LUNG CANCER AND HAVING ONLY ONE LUNG, ALONG WITH THE ACCUMULATION OF MULTIPLE ANATOMICAL DIFFICULTIES, VALVE INFOLDING, AND THE IN AND OUT WITH BOTH DCS, ALL CONTRIBUTED TO THE PATIENT'S DEATH.Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE EVOLUTFX-34 VALVE, THE VALVE WAS DEPLOYED LOW DURING THE FIRST DEPLOYMENT. THE VALVE WAS RECAPTURED AND DEPLOYED A SECOND TIME WHEN AN INFOLDING OF THE VALVE WAS OBSERVED WHICH NECESSITATED THE REMOVAL AND REPLACEMENT WITH ANOTHER FX34 VALVE. THE PATIENT, WHO WAS DESCRIBED AS VERY FRAGILE, UNDERWENT SURGERY ON A TABLE SHORTLY AFTER THE SECOND VALVE WAS IMPLANTED. AN ECHOCARDIOGRAM REVEALED A PERICARDIAL ISSUE DUE TO AN AORTIC DISSECTION CAUSED BY THE PROCEDURE. UNFORTUNATELY, THE PATIENT DID NOT SUPPORT THE SURGERY AND DIED DURING THE NIGHT. ADDITIONAL INFORMATION WAS RECEIVED WHICH CONFIRMED THAT THE PREVIOUSLY REPORTED PERICARDIAL ISSUE WAS A PERICARDIAL EFFUSION. THE TIMING OF THE AORTIC DISSECTION WAS UNCERTAIN AS IT WAS NOT NOTED UNTIL THE FINAL ANGIOGRAM. PER THE PHYSICIAN, THE CAUSE OF THE DISSECTION WAS THE PATIENT'S AORTIC ANGULATION OF 27° ALONG WITH DIFFICULTY ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) IN THE AORTIC ARCH AND THE ASCENDING AORTA. ADDITIONALLY, THE FIRST AND THE SECOND DCS REQUIRED PRESSURE TO ADVANCE. FOR THE SECOND VALVE IMPLANT, A NON-MEDTRONIC (LUNDERQUIST) GUIDEWIRE WAS PUT IN THE PIGTAIL CATHETER TO HELP THE VALVE CROSS THE AORTIC ARCH. INITIALLY, DRAINAGE WAS INTENDED FOR THE PERICARDIAL EFFUSION, BUT THEN WAS CHANGED TO SURGICAL INTERVENTION. THE PATIENT'S TORTUOUS ANATOMY AND CALCIFIED ANNULUS CONTRIBUTED TO THE DISSECTION. PER THE PHYSICIAN, THE PATIENT'S UNDERLYING MEDICAL HISTORY INCLUDING LUNG CANCER AND HAVING ONLY ONE LUNG, ALONG WITH THE ACCUMULATION OF MULTIPLE ANATOMICAL DIFFICULTIES, VALVE INFOLDING, AND THE IN AND OUT WITH BOTH DCS, ALL CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: UPDATED DATA: D. SUSPECT MEDICAL DEVICE (SECOND DCS USED IN THE CASE): EXPIRATION DATE; LOT #; FULL UDI # H4. DEVICE MFG DATE H6. EVAL CODE METHOD. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED H.6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10. PRODUCT ID D-EVOLUTFX-34; PRODUCT LOT NUMBER 0012423694; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID EVOLUTFX-34; PRODUCT SERIAL NUMBER (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE PRODUCT ID EVOLUTFX-34; PRODUCT SERIAL NUMBER (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. H6. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22021293&pc=DSR | 3007972010-2025-00043 | 2025-04-13 04:00 | Death | CVRX, INC. | 2025-05-14 04:00 | DSR | BAROSTIM NEO2 | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Tachycardia; Unspecified Kidney or Urinary Problem | P180050 |  | 1 | Event Description: A BAROSTIM SYSTEM WAS IMPLANTED ON (B)(6) 2025. THE PATIENT HAD BEEN HOSPITALIZED SINCE MID-FEBRUARY WITH A HEART-FAILURE RELATED ADMISSION, WAS PRESCRIBED BAROSTIM, AND WAS DISCHARGED HOME ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT PRESENTED TO THE EMERGENCY ROOM WITH HYPOTENSION, A NEED FOR DIALYSIS, AND VENTRICULAR TACHYCARDIA. INTERVENTIONS WERE PERFORMED, BUT THE PATIENT PASSED AWAY THE SAME DAY. Manufacturer Narrative: THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE AT THIS TIME, AND THE REPORTED DEVICE WAS NOT RETURNED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED EVENT COULD NOT BE CONCLUSIVELY DETERMINED. THE DEVICE HISTORY AND STERILIZATION RECORD FOR THIS DEVICE SERIAL NUMBER HAS BEEN REVIEWED. NO ISSUES OR DISCREPANCIES WERE NOTED DURING THIS REVIEW THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. THE DEVICE MET MATERIAL, ASSEMBLY, AND QUALITY CONTROL REQUIREMENTS. CVRX ID# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22021577&pc=BZD | 2518422-2025-105650 | 2022-07-13 04:00 | Death | RESPIRONICS, INC. | 2025-05-14 04:00 | BZD | DS2ADV AUTO CPAP W/HUMID CELL/BT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT NEVER RECEIVED THEIR DREAMSTATION 2 REPLACEMENT DEVICE. THE PATIENT REPORTEDLY PASSED AWAY, AND THE REPLACEMENT DEVICE IS NO LONGER NEEDED. THE AFFECTED DEVICE WAS DISCARDED/RECYCLED AT THEIR LOCAL COURTHOUSE. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22021594&pc=OZD | 1220648-2025-28412 | 2025-04-26 04:00 | Death | ABIOMED, INC. | 2025-05-14 04:00 | OZD | IMPELLA CP | Positioning Failure; Restricted Flow rate | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD VENTRICULAR FIBRILLATION AND NEEDED TO BE DEFIBRILLATION WITH AMIODARONE /LIDOCAINE STARTED. THE PUMP WAS ALSO REPOSITIONED DURING SUPPORT. FURTHERMORE, THERE WERE SUCTION ALARMS THAT WERE SILENCED AND THE PUMP WAS SET TO P2. LATER, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: PUMP WAS NOT RETURNED FOR INVESTIGATION. DATA LOG WAS NOT PROVIDED FOR THIS CASE RUN. THE CAUSE OF THE LOW PUMP FLOW ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND SUFFICIENT CLINICAL INFORMATION WAS PROVIDED. THE CAUSE OF THE POSITIONING ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED. CARDIAC ARRYTHMIA CAN BE REPORTED DURING IMPELLA SUPPORT. TO MAKE A DETERMINATION AS TO THE CAUSE OF THE CARDIAC ARRYTHMIA, THE FOLLOWING CLINICAL INFORMATION MUST BE CONSIDERED: PATIENT'S MEDICAL HISTORY, INCLUDING ANY PRE-EXISTING CARDIAC CONDITIONS, PREVIOUS EPISODES OF ARRHYTHMIAS, OR STRUCTURAL HEART DISEASE. TIMING OF THE ARRYTHMIA EVENT IN RELATION TO IMPELLA SUPPORT (DURING OR POST-IMPLANT) ELECTROCARDIOGRAM (EKG) RECORDINGS OF THE ARRHYTHMIA EPISODES. EVALUATION OF ANY UNDERLYING ELECTRICAL DISTURBANCES, SUCH AS QT PROLONGATION OR ELECTROLYTE IMBALANCES. ASSESSMENT OF HEMODYNAMIC INSTABILITY DURING THE ARRHYTHMIA, INCLUDING BLOOD PRESSURE AND CARDIAC OUTPUT MEASUREMENTS. PRESENCE OF OTHER POTENTIALLY CONTRIBUTORY FACTORS, SUCH AS MEDICATION CHANGES, ISCHEMIA, OR ELECTROLYTE ABNORMALITIES. RESPONSE TO ANY ANTIARRHYTHMIC TREATMENTS OR INTERVENTIONS DURING THE EVENT. ANY DOCUMENTED DEVICE-RELATED ISSUES OR COMPLICATIONS DURING IMPELLA SUPPORT. EVALUATION OF POTENTIAL TRIGGERS OR PRECIPITATING FACTORS, SUCH AS CATHETER MANIPULATION, REPERFUSION INJURY, OR INCREASED MYOCARDIAL OXYGEN DEMAND. REVIEW OF ANY CONCOMITANT MEDICATIONS THAT MAY INFLUENCE CARDIAC ELECTROPHYSIOLOGY. WITH A RETURNED IMPELLA, THE DEVICE CAN BE INSPECTED FOR ANY BURRS OR ROUGH EDGES THAT MAY CONTACT THE HEART WALL. TO DETERMINE THE TRUE ROOT CAUSE OF THE VT/VF, THE CLINICAL INFORMATION MENTIONED ABOVE WOULD NEED TO BE ASSESSED IN CONJUNCTION WITH EVIDENCE FROM THE RETURNED DEVICE. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22021974&pc=OZD | 1220648-2025-28417 | 2025-04-08 04:00 | Death | ABIOMED, INC. | 2025-05-14 04:00 | OZD | IMPELLA CP | Defective Device | Hemolysis; Hemorrhage/Bleeding; Ischemia; Sepsis; Thrombocytopenia; Mitral Valve Insufficiency/ Regurgitation | P140003 |  | 1 | Event Description: IT WAS REPORTED THROUGH ABIOMED'S LONG-TERM OUTCOME AND OUALITY IMPELLA REGISTRY DATABASE MULTIPLE ADVERSE EVENTS WITH POSSIBLE/PROBABLE/DEFINITE RELATIONSHIP TO THE USE OF THE IMPELLA CP. THE PATIENT EXPERIENCED THROMBOCYTOPENIA. THEREFORE, THE HEPARIN DRIP WAS CHANGED TO ARGATROBAN. THE PATIENT EXPERIENCED LIMB ISCHEMIA, ACCESS SITE BLEEDING, AND HEMOLYSIS. THE PUMP SHEATH WAS REPLACED AND WAS TRANSFUSED SIX UNITS OF PACKED RED BLOOD CELLS. THE PATIENT EXPERIENCED MITRAL VALVE REGURGITATION AND WAS TREATED WITH NOREPIPHRINE, VASOPRESSIN, NITROPRUSSIDE, AND MILRINONE. THE PATIENT BECAME SEPTIC AND WAS TREATED WITH VANCOMYCIN AND CEFEPIME. THE FAMILY DECIDED TO WITHDRAW LIFE SUPPORT AND THE PATIENT EXPIRED. Manufacturer Narrative: B.2 EXACT DATE OF DEATH IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: IMPELLA CATHETER IN PAPILLARY MUSCLE ¿IF THE INLET AREA OF THE IMPELLA CATHETER IS TOO CLOSE TO OR ENTANGLED IN THE PAPILLARY MUSCLE AND/OR SUBANNULAR STRUCTURES SURROUNDING THE MITRAL VALVE, IT CAN AFFECT MITRAL VALVE FUNCTION AND NEGATIVELY IMPACT CATHETER FLOW.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION. ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22022665&pc=NPU | 2015691-2025-03831 | 2025-04-25 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-05-14 04:00 | NPU | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER VALVE | Migration or Expulsion of Device; Perivalvular Leak | Insufficient Information | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), REPORTED THAT AN OFF LABEL NATIVE MITRAL VALVE WITH SEVERE HEMOLYSIS PROCEDURE. A 29MM SAPIEN 3 ULTRA RESILIA +7ML HAD SEVERE PVL NOTED AFTER DEPLOYMENT. THE TEAM STARTED TO PREP SECOND VALVE AND FIRST VALVE EMBOLIZED ATRIAL. NO VALVE WAS IMPLANTED. IMPLANT PATIENT REGISTRY SENT A REPORT THAT THE PATIENT PASSED AWAY. NO OTHER INFORMATION PROVIDED. EXACT DATE OF DEATH IS UNKNOWN.Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), REPORTED THAT AN OFF LABEL NATIVE MITRAL VALVE WITH SEVERE HEMOLYSIS PROCEDURE. A 29MM SAPIEN 3 ULTRA RESILIA +7ML HAD SEVERE PVL NOTED AFTER DEPLOYMENT. THE TEAM STARTED TO PREP SECOND VALVE AND FIRST VALVE EMBOLIZED ATRIAL. THE PATIENT WAS NO VALVE WAS IMPLANTED. IMPLANT PATIENT REGISTRY SENT A REPORT THAT THE PATIENT PASSED AWAY. NO OTHER INFORMATION PROVIDED. UPON FOLLOW UP, THE FCS ADVISED THIS WAS A VALVE IN NATIVE MITRAL WITH MITRAL CLIPS IN PLACE THAT WAS HAVING SEVERE HEMOLYSIS AND WAS A LAST DITCH TO TRY AND RESOLVE IT AS THEY WERE NOT A SURGICAL CANDIDATE . THE VALVE EMBOLIZED ATRIAL AFTER ABOUT 10 MINUTES POST DEPLOYMENT WE ATTEMPTED TO SNARE IT FOR SEVERAL HOURS WITHOUT SUCCESS. THE PATIENT LEFT THE OR ALIVE.Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), REPORTED THAT AN OFF LABEL NATIVE MITRAL VALVE WITH SEVERE HEMOLYSIS PROCEDURE. A 29MM SAPIEN 3 ULTRA RESILIA +7ML HAD SEVERE PVL NOTED AFTER DEPLOYMENT. THE TEAM STARTED TO PREP SECOND VALVE AND FIRST VALVE EMBOLIZED ATRIAL. NO VALVE WAS IMPLANTED. IMPLANT PATIENT REGISTRY SENT A REPORT THAT THE PATIENT PASSED AWAY. NO OTHER INFORMATION PROVIDED. UPON FOLLOW UP, THE FCS ADVISED THIS WAS A VALVE IN NATIVE MITRAL WITH MITRAL CLIPS IN PLACE THAT WAS HAVING SEVERE HEMOLYSIS AND WAS A LAST DITCH TO TRY AND RESOLVE IT AS THEY WERE NOT A SURGICAL CANDIDATE. THE VALVE EMBOLIZED ATRIAL AFTER ABOUT 10 MINUTES POST DEPLOYMENT WE ATTEMPTED TO SNARE IT FOR SEVERAL HOURS WITHOUT SUCCESS. THE PATIENT LEFT THE OR ALIVE. UPON FOLLOW UP, AN OPERATIVE NOTE FROM THE TMVR PROCEDURE WAS PROVIDED. THE PATIENT PRESENTED WITH MITRAL REGURGITATION. UNDER ANESTHESIA, THE PATIENT HAD A 29MM SAPIEN S3UR IN THE MITRAL POSITION VIA RIGHT FEMORAL VEIN APPROACH. THE IMPLANT WAS A SUCCESS. WHILE EVALUATING THE VALVE ON TEE, THE VALVE EJECTED INTO THE LA ACUTELY AND THE PATIENT BECAME UNSTABLE. THE EMBOLIZED VALVE WAS PULLED THROUGH THE LA INTO THE RA. THE PATIENT WAS STABILIZED. THE NATIVE VALVE HAD RESIDUAL FUNCTION TO KEEP THE PATIENT SOMEWHAT STABLE. THERE WAS NO SURGICAL OPTION. THE PATIENT HAD PROFOUND HEMOLYSIS AND WAS CONTINUOUSLY BECOMING ANEMIC, THROMBOCYTOPENIC, AND GETTING PROGRESSIVE GANGRENE DISTALLY. THE PATIENT WAS SENT TO THE ICU FOR CONTINUED MANAGEMENT. THE FAMILY'S PLAN WAS TO MOVE TOWARDS COMFORT CARE/HOSPICE IF THIS INTERVENTION DID NOT WORK. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION OF THE DEVICE, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, LOSS OF PACING CAPTURE, A NARROW SINOTUBULAR JUNCTION (IN AORTIC POSITION PROCEDURES), RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE LANDING ZONE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. WITH THE LIMITED INFORMATION PROVIDED, THE CAUSE OF THE EMBOLIZATION IS UNKNOWN BUT MAY BE RELATED TO THE MECHANISMS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. THE DEVICE WAS USED IN OFF-LABEL IMPLANTATION IN THE MITRAL POSITION. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO NATIVE MITRAL PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO MEDICAL RECORD REVIEW, SECTIONS G6, H2, B5, H6: DEVICE CODE, TYPE OF INVESTIGATION ADDITIONAL CODE, AND H11 HAS BEEN ADDED. PER THE INSTRUCTIONS FOR USE (IFU), VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION OF THE DEVICE, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, LOSS OF PACING CAPTURE, A NARROW SINOTUBULAR JUNCTION (IN AORTIC POSITION PROCEDURES), RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE LANDING ZONE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. WITH THE LIMITED INFORMATION PROVIDED, THE CAUSE OF THE EMBOLIZATION WAS CAUSED BY SEVERE PVL ON THE DEPLOYED VALVE AND/OR MAY BE RELATED TO THE MECHANISMS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. PER THE INSTRUCTIONS FOR USE (IFU), PARAVALVULAR LEAK (PVL) IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE/PRE-STENT AND THE LANDING ZONE, AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE. SOME PVL IS NOT UNCOMMON POST-DEPLOYMENT. MANY CASES ARE MILD TO MODERATE, AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS. OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION. THE PATIENT SCREENING MANUAL AND THE PROCEDURE DIDACTIC IDENTIFY SEVERAL PROCEDURAL AND ANATOMICAL FACTORS WHICH COULD CONTRIBUTE TO PVL, INCLUDING DEVICE MALPOSITION, INACCURATE MEASUREMENT OF THE LANDING ZONE, UNEVEN DISTRIBUTION OF CALCIUM ON THE LANDING ZONE, BULKY OR SEVERE CALCIFICATION, AN ELLIPTICAL ANNULUS SHAPE, AND VALVE UNDER-SIZING. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. THE OFF-LABEL IMPLANTATION IN THE MITRAL POSITION COULD HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). THE DEVICE TRAINING MANUALS INSTRUCTS THE OPERATOR ON PROPER IMAGING SCREENING REQUIREMENTS, INCLUDING THE USE OF GOOD QUALITY ECHOCARDIOGRAPHY AND/OR COMPUTED TOMOGRAPHY (CT) TO APPROPRIATELY MEASURE THE LANDING ZONE, ASSESS THE CONTENT AND DISTRIBUTION OF CALCIUM, AND OTHER PATIENT ANATOMICAL FACTORS. MULTIPLE IMAGING MODALITIES SHOULD BE CONSIDERED DURING VALVE SIZE SELECTION. CONTRAINDICATIONS, IMPORTANT CONSIDERATIONS WHEN ASSESSING THE VALVE, AND CHOOSING THE PROPER TRANSCATHETER HEART VALVE ARE ALSO DISCUSSED. THE DEVICE TRAINING MANUALS ALSO INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING PATIENTS SCREENING AND PROCEDURAL CONSIDERATIONS. DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES ARE ALSO INCLUDED. THE DEVICE WAS USED IN OFF-LABEL IMPLANTATION IN THE MITRAL POSITION. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO NATIVE MITRAL PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ADDITIONAL INFORMATION RECEIVED, SECTIONS G6, H2, B5 UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22023511&pc=BZD | 2518422-2025-105664 | 2024-09-04 04:00 | Death | RESPIRONICS, INC. | 2025-05-14 04:00 | BZD | DREAMSTATION HUMIDIFIER, GB-RENT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF A DREAMSTATION AUTO CPAP. THE PATIENT'S DAUGHTER REPORTED THE PATIENT HAS SADLY PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THIS NOTIFICATION IS CURRENTLY UNDER REVIEW FOLLOWING A REPORT INVOLVING A USER OF THE DREAMSTATION HUMIDIFIER, MODEL GB-RENT. THE PATIENT'S DAUGHTER REPORTED THAT THE PATIENT HAS, UNFORTUNATELY, PASSED AWAY. THE CAUSE OF DEATH WAS NOT DISCLOSED. AN INVESTIGATION BY THE MANUFACTURER IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. PLEASE NOTE THAT A CORRECTION HAS BEEN MADE TO THE ASSOCIATED 510(K) NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22007842&pc=LKK | 3004209178-2025-08194 | 2025-04-17 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-13 04:00 | LKK | SYNCHROMED II | Adverse Event Without Identified Device or Use Problem | Purulent Discharge; Erythema; Unspecified Infection; Muscular Rigidity; Pain; Twitching; Swelling/ Edema | P860004 |  | 1 | Event Description: INFORMATION WAS RECEIVED FROM A FRIEND/FAMILY MEMBER REGARDING A PATIENT WHO WAS RECEIVING MORPHINE DRUG AND BACLOFEN VIA AN IMPLANTABLE PUMP FOR INTRACTABLE SPASTICITY INDICATIONS FOR USE. IT WAS REPORTED THAT THE PATIENT HAD THE PUMP FILLED ON MONDAY THE (B)(6), AND BY THURSDAY THE (B)(6), AN INFECTION WAS DISCOVERED AROUND THE PUMP. THE PATIENT REPRESENTATIVE (REP) STATED THAT ON SATURDAY, (B)(6) 2025, THEY HAD TO LIFE FLIGHT THE PATIENT TO THE (B)(6). ON SUNDAY (B)(6) 2025, AND THEY REMOVED THE PUMP AND THE CATHETER DUE TO THE INFECTION. THE PATIENT WAS ON ANTIBIOTICS, AND THEY WERE TRYING TO FIND A BALANCE BETWEEN THE BACLOFEN/MORPHINE THAT WAS IN THE PUMP TO COVER THE SPASTICITY/JERKS AND PAIN WITH ORAL MEDICATIONS. THE REP STATED THAT THE PATIENT HAS BEEN IN A STATE OF SEDATION SINCE THE SURGERY. SINCE THEY WILL NOT BE ABLE TO PUT IN ANOTHER PUMP UNTIL AFTER THE INFECTION IS CLEARED, THE REP WAS ASKING IF A MEDTRONIC PUMP COULD BE USED OUTSIDE THE BODY WITH EPIDURAL TYPE CATHETER.Event Description: ADDITIONAL INFORMATION RECEIVED FROM A HEALTHCARE PROFESSIONAL (HCP) REPORTED THAT AT A RECENT REFILL ON (B)(6) 2024 THE PATIENT COMPLAINED OF SWELLING OVER THE PUMP SITE. THEY HAD DENIED ANY SYSTEMIC SYMPTOMS. THE PATIENT WENT TO THE EMERGENCY ROOM ON (B)(6) 2025 AND 95 CC OF PURULENT DRAINAGE HAD BEEN DRAINED WITH SURROUNDING ERYTHEMA. IT HAD BEEN DECIDED TO REMOVE THE IMPLANTED SYSTEM DUE TO INFECTION ON (B)(6) 2025. NO FURTHER COMPLICATIONS WERE NOTED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED FROM A HEALTHCARE PROVIDER (HCP) AND IT WAS REPORTED THE PATIENT WAS DECEASED. ADDITIONAL INFORMATION WAS RECEIVED FROM A HEALTHCARE PROVIDER (HCP) ON 2025-06-05 AND IT WAS REPORTED THAT THE PATIENT'S DEATH WAS RELATED TO THE DEVICE AND/OR THERAPY. IT WAS NOTED THE PATIENT HAD COMPLICATIONS WITH INFECTION STATUS POST PUMP REFILL THAT LED TO PATIENT EXPIRATION. ADDITIONAL INFORMATION WAS RECEIVED AGAIN FROM A HEALTHCARE PROVIDER (HCP) ON 2025-06-20 AND IT WAS REPORTED THAT THE PATIENT WAS IN AN OUT OF AREA HOSPITAL WHEN THE DEVICE WAS REMOVED SO THE DEVICE STATUS WAS UNKNOWN. Manufacturer Narrative: B2: EXACT DATE OF DEATH IS UNKNOWN; DATE OF DEATH IS APPROXIMATE BASED ON THE INFORMATION REPORTED. H6. CORRECTION: PATIENT CODE: E2338 IS ALSO APPLICABLE TO THIS EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22008663&pc=OZD | 1220648-2025-28341 | 2025-04-30 04:00 | Death | ABIOMED, INC. | 2025-05-13 04:00 | OZD | IMPELLA CP-JP | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT ECPELLA WAS USED FOR VENTRICULAR STORM (VT) IN HYPERTROPHIC CARDIOMYOPATHY (HCM) PATIENT. THE PATIENT EXPIRED DESPITE TREATMENT. THE CAUSE OF DEATH WAS FATAL ARRHYTHMIA AND MULTIORGAN FAILURE. THE DOCTOR COMMENTED THAT THERE WAS NO CAUSAL RELATIONSHIP WITH THE IMPELLA. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA USED AS AN ASSOCIATED FACTOR WITH THE EVENT AND FOR THIS REASON THIS EVENT WILL BE REPORTED. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. B3 DATE OF EVENT WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28341. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28341 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22009153&pc=OZD | 1220648-2025-28347 | 2025-04-21 04:00 | Death | ABIOMED, INC. | 2025-05-13 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THERE WAS SUCTION AT P6. THE PUMP WAS LOWERED TO P5. ADDITIONALLY, THE PATIENT HAD A COMPUTED TOMOGRAPHY AND ELECTROENCEPHALOGRAM AND IT SHOWED NO BRAIN ACTIVITY AND BRAIN HERNIATION. IT IS UNKNOWN IF THE IMPELLA CAUSED OR CONTRIBUTED TO THE BRAIN INJURIES. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22009338&pc=LWS | 2124215-2025-30663 | 2025-04-25 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-13 04:00 | LWS | RESONATE EL ICD DR | Failure to Capture; Pacing Problem; Failure to Convert Rhythm; Inaccurate Synchronization; Under-Sensing | Insufficient Information | P960040/S385 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) PASSED AWAY (B)(6) 2025 ON A PLANE WHILE TRAVELING TO HAVE A VENTRICULAR TACHYCARDIA (TV) ABLATION PROCEDURE. REVIEW OF DEVICE DATA BY BOSTON SCIENTIFIC TECHNICAL SERVICES INDICATED THERE WERE MULTIPLE EPISODES RECORDED IN THE DEVICE MEMORY ON THE DATE THE PATIENT PASSED. THESE EPISODES NOTED PREMATURE VENTRICULAR CONTRACTIONS TRIGGERING VT WHICH WAS SUBSEQUENTLY TREATED WITH ANTI-TACHYCARDIA PACING (ATP). TWO OTHER EPISODES NOTED UNDERSENSING CAUSING INAPPROPRIATE PACING AS WELL AS RHYTHM ACCELERATION OF VT. LOSS OF CAPTURE WAS ALSO NOTED TOWARDS THE END OF THE FINAL EPISODE RECORDED, LIKELY FROM THE PATIENT'S HEART DYING. ALL EVIDENCE INDICATES THE ICD REMAINS IN SITU. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22009397&pc=LWP | 2017865-2025-53502 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-13 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO THEIR PACEMAKER SYSTEM. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO PATIENT EXPIRED. THE DEVICE WAS RECEIVED WITH A BATTERY LEVEL WITHIN THE NORMAL RANGE OF OPERATION. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. AFTER ALL ELECTRICAL TESTS PERFORMED, INCLUDING THERMAL AND MECHANICAL STRESS TESTING THE DEVICE EXHIBITS NORMAL CHARACTERISTICS. LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY. NO ANOMALIES WERE FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22009591&pc=DSQ | 3007042319-2025-01136 | 2025-03-19 04:00 | Death | HEARTWARE, INC. | 2025-05-13 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Gastritis; Hemorrhage/Bleeding; Intracranial Hemorrhage; Unspecified Infection; Septic Shock; Ulcer; Hematemesis; Melena | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING GASTROINTESTINAL BLEEDS (GIB) AND RISK SCORES IN PATIENTS WITH A LEFT VENTRICULAR ASSIST DEVICE (VAD). PATIENTS WERE STUDIED IN GROUPS OF GIB OR NO GIB WITH MANY PATIENTS EXPERIENCING MULTIPLE MAJOR ADVERSE EVENTS. THE AUTHORS DESCRIBED PATIENT DEATHS; THE CAUSES OF DEATH WERE INFECTION, BLEEDING EVENTS, SEPTIC SHOCK, INFECTIVE PNEUMONITIS, INTRACRANIAL HEMORRHAGE, AND HEMORRHAGIC SHOCKS SECONDARY TO GIB. GIB WAS DEFINED AS OVERT BLEEDING WITHIN THE GASTROINTESTINAL TRACT (MELENA, HEMATOCHEZIA, HEMATEMESIS, AND COFFEE-GROUND EMESIS). GIB ETIOLOGIES INCLUDED ARTERIOVENOUS MALFORMATION (AVM), GASTRIC ULCER, COLON DIVERTICULOSIS, AND EROSIVE GASTRITIS, AND MANY WERE TREATED WITH TRANSFUSIONS. IN ADDITION TO GIB, THERE WERE PATIENTS WHO EXPERIENCED ISCHEMIC STROKES, VAD THROMBOSIS, INTRACRANIAL BLEEDING, HEMOTHORAX, NASOPHARYNGEAL HEMORRHAGE, COLOSTOMY BLEEDING, DRIVELINE INFECTIONS, SEPSIS, WOUND DEHISCENCE, PNEUMONIA, PLEURAL EFFUSION, ACUTE RENAL FAILURE, TRANSIENT ISCHEMIC ATTACK (TIA), ACUTE CHOLECYSTITIS, PANCREATITIS, PSEUDOMEMBRANOUS COLITIS, COGNITIVE IMPAIRMENT, PNEUMOTHORAX, PNEUMOPERITONEUM, VENTRICULAR ARRHYTHMIAS, ATRIAL ARRHY THMIAS, VASCULAR ACCESS BLEEDING, NASOPHARYNGEAL HEMORRHAGES, AND HOSPITALIZATION FOR RIGHT VENTRICULAR (RV) FAILURE AND UNKNOWN VAD FAILURES. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/58 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: GASTROINTESTINAL BLEEDING DURING LONG-TERM LEFT VENTRICULAR ASSIST DEVICE SUPPORT: EXTERNAL VALIDATION OF UTAH BLEEDING RISK SCORE. JOURNAL OF CARDIOVASCULAR DEVELOPMENT AND DISEASE. 2025. 12, 105. DOI: 10.3390/JCDD12030105 D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22009797&pc=MUE | MW5170248 | 2024-10-13 04:00 | Death | UNK | 2025-05-13 04:00 | MUE | MAMMOGRAPH-DIGITAL | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: PATIENT DECEASED PER FAMILY MEMBER, DEATH DATE WAS UPDATED IN ELECTRONIC MEDICAL RECORD. NO INFORMATION REGARDING CAUSE OF DEATH IS AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22009814&pc=LJE | MW5170249 | 2025-03-03 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-13 04:00 | LJE | BOSTON SCIENTIFIC PERCUFLEX URINARY DIVERSION STENT SET 7 FR X80CM URETERAL STEN | Fracture; Difficult to Remove | Sepsis; Foreign Body In Patient |  |  | 1 | Event Description: FRACTURED SINGLE-J STENT UNABLE TO BE REMOVED AT THE TIME OF STENT CHANGE RESULTING IN RETAINED FRAGMENT WITH SUBSEQUENT SEPSIS, AND DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22009888&pc=MKJ | 1220908-2025-01605 | 2025-04-19 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-05-13 04:00 | MKJ | X SERIES | Failure to Deliver Shock/Stimulation | Cardiac Arrest | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE FAILED TO DISCHARGE. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS RECEIVED THE PRODUCT AND WILL BE PROVIDING A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS COMPLETED. Manufacturer Narrative: THE DEVICE WAS RECEIVED AT ZOLL MEDICAL CORPORATION FOR EVALUATION. REVIEW OF THE LOG CONFIRMED CHARGES OCCURRED, BUT DEFIBRILLATION WAS NOT DELIVERED DUE TO A PADS LEAD FAULT. THE ISSUE APPEARS TO BE RELATED TO IMPROPER CONNECTION OF CPR PAD ACCESSORIES, SPECIFICALLY, THE CPR ADAPTER WAS LIKELY NOT CONNECTED TO THE MULTIFUNCTION CABLE (MFC). CABLE ID CHANGES OBSERVED IN THE ACTIVITY LOG SUPPORT THIS CONCLUSION. THE DEVICE AND ACCESSORIES WERE RETURNED, TESTED, AND FOUND FULLY FUNCTIONAL. ALL TESTS, INCLUDING BENCH TESTING AND DEFIB CYCLING PASSED WITH NO FAULT FOUND. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22010299&pc=OZD | 1220648-2025-28336 | 2025-04-16 04:00 | Death | ABIOMED, INC. | 2025-05-13 04:00 | OZD | IMPELLA 5.5 | Delivered as Unsterile Product | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT WAS ADMITTED TO ER WITH ABDOMEN PAIN AND CHEST PAIN WITH AN INFERIOR MYOCARDIAL INFARCTION. THE PATIENT WAS NOTED TO BE CONFUSED ON ARRIVAL WITH HYPOXIA AND INTUBATED. GIVEN PROFOUND ACIDOSIS AND COMORBITIES WITH 3 PRESSERS AND NOT A SURGICAL OR EXTRACORPOREAL MEMBRANE OXYGENATION CANDIDATE. PATIENT PROCEEDED TO CATH LAB. RIGHT HEART CATH WAS COMPLETED CONFIRMING CARDIOGENIC SHOCK AND LEFT VENTRICLE GRAM REVEALED VSD. PATIENT TRANSFERRED FOR FURTHER MEDICAL TREATMENT. THE COMPLAINANT REPORTED THAT THE PACKAGING OF AN INTRODUCER OPENED IRREGULARLY RESULTING IN THE SIDEARM GETTING CONTAMINATED. A SECOND INTRODUCER KIT WAS OPENED FROM A DIFFERENT 5.5 KIT FOR USE. FOLLOWING A SUCCESSFUL IMPLANT, A WOUND VAC WAS PLACED AT THE ACCESS SITE FOR BLEEDING. UPON ARRIVAL TO ICU PATIENT CONTINUED TO DECLINE. LACTATE CONTINUED TO RISE, GREATER THAN 20. SHE WAS ANURIC AND WOULD NOT TOLERATE CONTINUOUS RENTAL REPLACEMENT THERAPY. PATIENT WAS ALSO NOT A CANDIDATE FOR EXTRACORPOREAL MEMBRANE OXYGENATION DUE TO BODY MASS INDEX AND GROIN ACCESS. AFTER FURTHER DISCUSSION WITH FAMILY, DECISION WAS MADE TO WITHDRAW CARE ONCE FAMILY WAS ABLE TO ARRIVE AT BEDSIDE. WITHDRAW OF CARE OCCURRED AND PATIENT EXPIRED QUICKLY AFTER. Manufacturer Narrative: THE MEDICAL SAFETY OFFICER REVIEWED AND DETERMINED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA AS AN ASSOCIATION FACTOR TO THE PATIENT'S OUTCOME. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22010387&pc=MIH | 9612164-2025-02366 | 2025-02-27 05:00 | Death | MEDTRONIC IRELAND | 2025-05-13 04:00 | MIH | ENDURANT II STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Ischemia; Necrosis | P100021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ¿ RUPTURED ABDOMINAL AORTIC ANEURYSM TREATED USING INFRARENAL-OCCLUSION-BALLOON TECHNIQUE WITH ENDURANT IIS¿. AN ENDURANT IIS BIFURCATE STENT GRAFT WAS IMPLANTED IN THE ENDOVASCULAR TREATMENT OF A 61MM RUPTURED ABDOMINAL AORTIC ANEURYSM ON AN UNKNOWN DATE OVER AN EIGHT YEAR PERIOD. DURING THE PROCEDURE AN OCCLUSIVE BALLOON WAS INFLATED JUST BELOW THE RENAL ARTERY TO HELP AVOID DAMAGE TO THE NATIVE AORTA AND RISK OF ORGAN ISCHEMIA. INTRA-BLADDER PRESSURE (IBP) WAS USED AS A PROXY FOR INTRA-ABDOMINAL PRESSURE TO DIAGNOSE ACUTE COMPARTMENT SYNDROME (ACS). DURING THE PROCEDURE THE PATIENTS IBP WAS > 30CM SO A LAPAROTOMY FOR DECOMPRESSION WAS PERFORMED. THE PATIENT SURVIVED THE PROCEDURE BUT DIED ON POST-OPERATIVE DAY 3 FROM INTESTINAL NECROSIS. THE DEATH WAS NOT ATTRIBUTED TO THE DEVICE. NO ADDITIONAL CLINICAL "SEQUELAE" WERE PROVIDED. Manufacturer Narrative: MEDTRONIC RECEIVED THE FOLLOWING INFORMATION FROM A JOURNAL ARTICLE ENTITLED; ¿RUPTURED ABDOMINAL AORTIC ANEURYSM TREATED USING INFRARENAL-OCCLUSION-BALLOON TECHNIQUE WITH ENDURANT IIS¿. AZUMA S, ASADA Y, SHIMADA R, NAKAMURA S ANNALS OF VASCULAR SURGERY - BRIEF REPORTS AND INNOVATIONS 5 (2025) 100373 HTTPS://DOI.ORG/10.1016/J.AVSURG.2025.100373. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22010447&pc=PNJ | 9612164-2025-02361 | 2024-11-08 05:00 | Death | MEDTRONIC IRELAND | 2025-05-13 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Dyspnea | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED ONE DAY POST IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. ON THE SAME DAY AS THE IMPLANT THE PATIENT REPORTED DYSPNEA IN THE WARD, LEADING TO AN URGENT COMPUTERIZED TOMOGRAPHY (CT) SCAN. A CT SIGN IN THE BASILAR ARTERY WAS OBSERVED, LEADING TO A SUSPICION OF CEREBRAL INFARCTION. AT THE TIME OF THE CHECK IT WAS NOTED THAT THE IPG EXHIBITED GOOD DATA. THE FOLLOWING DAY, THE PATIENT LOST CONSCIOUSNESS AND PASSED AWAY IN THE INTENSIVE CARE UNIT (ICU). CAUSE OF DEATH WAS PROVIDED AS NON-CARDIAC DEATH. BASILAR ARTERY INFARCTION. THE IPG REMAINS IN THE PATIENT.Event Description: T WAS FURTHER REPORTED THAT SPONTANEOUS BREATHING WAS ACHIEVED, AND CIRCULATION STABILIZED THROUGH RESPIRATORY MANAGEMENT, BUT CONSCIOUSNESS DID NOT RETURN, LEADING TO POOR PUPIL RESPONSE. THE PATIENT'S MEDICATION WAS ADJUSTED. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [(B)(6)). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22010681&pc=HRS | 0001032347-2025-00202 | 2024-06-17 04:00 | Death | BIOMET MICROFIXATION | 2025-05-13 04:00 | HRS | UNK STERNALOCK PLATE | Insufficient Information | Insufficient Information | NI |  | 10 | Event Description: IT WAS REPORTED IN A JOURNAL ARTICLE FROM PERIOPERATIVE & CRITICAL CARE: RESEARCH, WHICH STUDIED CONTRIBUTIONS OF RIGID PLATE FIXATION ON POSTOPERATIVE PAIN, OPIOID USE, AND OTHER STERNOTOMY OUTCOMES. TEN PATIENTS HAD DEATH OCCUR WITHIN 90 DAYS OF OPERATION DUE TO AN UNKNOWN REASON. ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE AT THIS TIME.Event Description: NO FURTHER EVENT INFORMATION IS AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL INFORMATION. THE FOLLOWING SECTIONS WERE UPDATED: B3, B4, B5, G3, G6, H2, H6, H11. D4: ATTEMPTS HAVE BEEN MADE TO GATHER ALL PRODUCT IDENTIFICATION INFORMATION AND NO FURTHER INFORMATION HAS BEEN PROVIDED. A JOURNAL ARTICLE REVIEWED 608 MEDIAN STERNOTOMY PATIENTS- 59.2% WERE ISOLATED CORONARY ARTERY BYPASS GRAFTING, 7.7% WERE ISOLATED VALVE PROCEDURES, AND THE REST WERE MIXED/CONCOMITANT PROCEDURES. OF THE COMPLEX CASES, TEN PATIENTS EXPIRED LESS THAN NINETY DAYS POST-OPERATIVE WITH A STUDY POPULATION OF A MEAN AGE OF 65.7 YEARS AT TIME OF SURGERY. IT WAS REPORTED THAT THE STUDY HAD A SUBSTANTIAL NUMBER OF HIGH RISK CARDIAC PROCEDURES WITH A MORTALITY RATE AT OR BELOW EXPECTED VALUES AND THAT THE PROCEDURES WERE COMPLETED WITHOUT DEVICE RELATED COMPLICATIONS. AS THE COMPLAINT INDICATES, THE PATIENT'S EXPIRED WITHIN NINETY DAYS OF THE PROCEDURE WITHOUT ALLEGATION AGAINST THE DEVICE. ROOT CAUSE OF THE REPORTED EVENT IS NOT DEVICE RELATED. IF ANY FURTHER INFORMATION IS FOUND WHICH WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS. Manufacturer Narrative: (B)(4). THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MEDWATCH 3500A WILL BE SUBMITTED. GERDISCH, MARC, JOHNS, CHANICE M, BARKSDALE, ANDREW, PARIKSHAK, MANESH. RIGID STERNAL FIXATION AND ENHANCED RECOVERY FOR OPIOID-FREE ANALGESIA AFTER CARDIAC SURGERY. ANN THORAC SURG. 2024;118(4):931-939. HTTPS://DOI.ORG/10.1016/J.ATHORACSUR.2024.06.032.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22010725&pc=MVK | 3008642652-2025-04725 | 2025-04-14 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-13 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FOUR INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 00:09:15 ON (B)(6) 2025. AT 05:17:49, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION/TACTILE ARTIFACT. AT 05:19:03, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. AT 05:22:38, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH MOTION/TACTILE ARTIFACT. AT 05:23:15, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION/TACTILE ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION/TACTILE ARTIFACT. AT 05:23:42, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION/TACTILE ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION/TACTILE ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 05:25:07 ON (B)(6) 2025. Manufacturer Narrative: INCOMING TESTING OF THE ELECTRODE BELT HAS BEEN COMPLETED. THE ELECTRODE BELT ECG ACQUISITION AND PULSE DELIVERY CIRCUITRY WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE MONITOR HAS BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22010864&pc=DSQ | 2916596-2025-02979 | 2025-04-04 04:00 | Death | THORATEC CORPORATION | 2025-05-13 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO AN UNKNOWN ETIOLOGY. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME. IT WAS REPORTED THE HEARTMATE 3 LVAS WAS NOT EXPLANTED FOR EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), (REV. C) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22010870&pc=MVK | 3008642652-2025-04727 | 2025-04-18 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-13 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE APPROPRIATE TREATMENTS AND TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 09:05:18 ON (B)(6) 2025. AT 12:07:21, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 260 BPM. AT 12:07:56, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 290 BPM. POST SHOCK RHYTHM WAS SVT @ 190 BPM. AT 12:08:30, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS NSVT. AT 12:09:04, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS NSVT. AT 12:10:26, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 260 BPM WITH TACTILE ARTIFACT. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 50 BPM WITH TACTILE ARTIFACT AND PAC¿S. AT 13:12:03, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 200 BPM WITH MOTION ARTIFACT. AT 13:12:38, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 200 BPM WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS SVT/SINUS TACHYCARDIA FROM 160-110 BPM WITH NSVT. THE ELECTRODE BELT WAS DISCONNECTED AT 13:13:07 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22010908&pc=MVK | 3008642652-2025-04726 | 2025-04-16 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-13 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 22:35:52 ON (B)(6) 2025. AT 18:09:27 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 150 BPM WITH MOTION ARTIFACT. AT 18:10:37, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 160 BPM. POST SHOCK RHYTHM WAS ASYSTOLE WITH NSVT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 18:11:05, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC AND NSVT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 70 BPM DEGRADING TO ASYSTOLE WITH CPR/MOTION ARTIFACT/INTERMITTENT CARDIAC ACTIVITY. THE ELECTRODE BELT WAS DISCONNECTED AT 18:11:23 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22011131&pc=DSQ | 2916596-2025-03076 | 2025-04-29 04:00 | Death | THORATEC CORPORATION | 2025-05-13 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE REASON WAS UNKNOWN. THE DEATH WAS NOT DEVICE RELATED. Manufacturer Narrative: SECTION A3: PATIENT GENDER WAS NOT PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE PROVIDED INFORMATION INDICATED THE PATIENT EXPIRED. ALTHOUGH THE CAUSE OF DEATH WAS REPORTED AS UNKNOWN, THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22011173&pc=MIH | 9612164-2025-02365 | 2025-02-27 05:00 | Death | MEDTRONIC IRELAND | 2025-05-13 04:00 | MIH | ENDURANT II STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Sepsis; Peritonitis | P100021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ¿ RUPTURED ABDOMINAL AORTIC ANEURYSM TREATED USING INFRARENAL-OCCLUSION-BALLOON TECHNIQUE WITH ENDURANT IIS¿. AN ENDURANT IIS BIFURCATE STENT GRAFT WAS IMPLANTED IN THE ENDOVASCULAR TREATMENT OF A 80MM RUPTURED ABDOMINAL AORTIC ANEURYSM ON AN UNKNOWN DATE BETWEEN 2016 AND 2024. DURING THE PROCEDURE AN OCCLUSIVE BALLOON WAS INFLATED JUST BELOW THE RENAL ARTERY TO HELP AVOID DAMAGE TO THE NATIVE AORTA AND RISK OF ORGAN ISCHEMIA. INTRA-BLADDER PRESSURE (IBP) WAS USED AS A PROXY FOR INTRA-ABDOMINAL PRESSURE TO DIAGNOSE ACUTE COMPARTMENT SYNDROME (ACS). DURING THE PROCEDURE THE PATIENTS IBP WAS > 30CM SO A LAPAROTOMY FOR DECOMPRESSION WAS PERFORMED. THE PATIENT SURVIVED THE PROCEDURE BUT DIED ON POST-OPERATIVE DAY 28 FROM SEPSIS DUE TO PERITONITIS . THE DEATH WAS NOT ATTRIBUTED TO THE DEVICE. NO ADDITIONAL CLINICAL SEQUALAE WERE PROVIDED. Manufacturer Narrative: MEDTRONIC RECEIVED THE FOLLOWING INFORMATION FROM A JOURNAL ARTICLE ENTITLED; ¿RUPTURED ABDOMINAL AORTIC ANEURYSM TREATED USING INFRARENAL-OCCLUSION-BALLOON TECHNIQUE WITH ENDURANT IIS¿. AZUMA S, ASADA Y, SHIMADA R, NAKAMURA S ANNALS OF VASCULAR SURGERY - BRIEF REPORTS AND INNOVATIONS 5 (2025) 100373 HTTPS://DOI.ORG/10.1016/J.AVSURG.2025.100373. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22011307&pc=HRS | 0001032347-2025-00203 |  | Death | BIOMET MICROFIXATION | 2025-05-13 04:00 | HRS | UNK STERNALOCK PLATE | Insufficient Information | Insufficient Information | NI |  | 9 | Event Description: IT WAS REPORTED IN A JOURNAL ARTICLE FROM PERIOPERATIVE & CRITICAL CARE: RESEARCH, WHICH STUDIED CONTRIBUTIONS OF RIGID PLATE FIXATION ON POSTOPERATIVE PAIN, OPIOID USE, AND OTHER STERNOTOMY OUTCOMES. NINE PATIENTS HAD DEATH OCCUR WITHIN 90 DAYS OF OPERATION DUE TO AN UNKNOWN REASON. IT IS NOT KNOWN WHETHER THIS WAS A DEVICE RELATED COMPLICATION. ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE AT THIS TIME.Event Description: NO FURTHER EVENT INFORMATION IS AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: (B)(4). THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MEDWATCH 3500A WILL BE SUBMITTED. GERDISCH, MARC, JOHNS, CHANICE M, BARKSDALE, ANDREW, PARIKSHAK, MANESH. RIGID STERNAL FIXATION AND ENHANCED RECOVERY FOR OPIOID-FREE ANALGESIA AFTER CARDIAC SURGERY. ANN THORAC SURG. 2024;118(4):931-939. HTTPS://DOI.ORG/10.1016/J.ATHORACSUR.2024.06.032. Manufacturer Narrative: THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL INFORMATION. THE FOLLOWING SECTIONS WERE UPDATED: B3, B4, B5, G3, G6, H2, H6, H11. D4: ATTEMPTS HAVE BEEN MADE TO GATHER ALL PRODUCT IDENTIFICATION INFORMATION AND NO FURTHER INFORMATION HAS BEEN PROVIDED. NO PRODUCT WAS RETURNED, OR PHOTOS PROVIDED; VISUAL AND DIMENSIONAL EVALUATIONS COULD NOT BE PERFORMED. LOT IDENTIFICATION IS NECESSARY FOR REVIEW OF DEVICE HISTORY RECORDS; LOT IDENTIFICATION WAS NOT PROVIDED. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED. THE REPORTED EVENT IS UNCONFIRMED. IF ANY FURTHER INFORMATION IS FOUND WHICH WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22011338&pc=OZD | 1220648-2025-28358 | 2025-04-21 04:00 | Death | ABIOMED, INC. | 2025-05-13 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND THE FOLLOWING DAY THERE WERE FREQUENT SUCTION ALARMS. VOLUME WAS PROVIDED TO THE PATIENT AND IMPELLA MCS CONTINUED. APPROXIMATELY THREE DAYS LATER, THE PATIENT EXPERIENCED RECURRENT VENTRICULAR TACHYCARDIA, WENT ASYSTOLE, BRADY, WHICH WAS RESOLVE WITH EIGHT SHOCKS. AMINO AND LIDOCAINE WERE ADMINISTERED. THE PATIENT AS GIVEN PACKED RED BLOOD CELLS AND PLATELETS AFTER LABS. THE FOLLOWING DAY THE PATIENT'S OVERALL CONDITION WAS NOTED AS UNCHANGED BUT STABLE. THE PATIENT WAS A FULL CODE; HOWEVER, THE CARE TEAM CAPPED CARE WITH NO PLANS FOR FURTHER INTERVENTIONS. TWO DAYS LATER ATTEMPTED UPDATE WAS UNSUCCESSFUL AS THE PATIENT WAS NOT IN THE DESIGNATED ROOM ON THE INTENSIVE CARE UNIT, AND THE IMPELLA PUMP WAS NOT STREAMING. STATUS OF THE PATIENT IS UNKNOWN.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED REGARDING THE PATIENT OUTCOME; DEATH. B2 REVISED CHANGED LIFE THREATENING TO DEATH AND DATE OF DEATH ADDED. B5 UPDATED WITH ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE D6B EXPLANT DATE ADDED. H1 TYPE OF EVENT REVISED TO DEATH. H6 HEALTH EFFECT - IMPACT CODE 1802 ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22011472&pc=NPT | 2025587-2025-03630 | 2025-03-25 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-13 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem; Central Regurgitation | Atrial Fibrillation; Heart Failure/Congestive Heart Failure; Mitral Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT ON THE SAME DAY AS THE IMPLANT OF A TRANSCATHETER VALVE IN A PATIENT WITH PRE-EXISTING CONGESTIVE HEART FAILURE (CHF), MITRAL REGURGITATION AND ATRIAL FIBRILLATION WERE OBSERVED, AND THE PATIENT'S CHF WORSENED. SUBSEQUENTLY, THE PATIENT DIED TO DUE CHF. PER THE PHYSICIAN, THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND DEVICE OR IMPLANT PROCEDURE. Manufacturer Narrative: UPDATED DATA: D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. CONTINUATION OF D10: PRODUCT ID {D-EVOLUTFX-2329}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEARTVALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22011691&pc=DSQ | 2916596-2025-02737 | 2025-04-16 04:00 | Death | THORATEC CORPORATION | 2025-05-13 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Unspecified Infection; Sepsis; Respiratory Acidosis; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure; Respiratory Insufficiency | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE DISTRIBUTOR UPDATED THAT DURING A FOLLOW UP ECHOCARDIOGRAM (ECHO) A STRANGE OBJECT IN THE LEFT VENTRICLE IN CLOSE PROXIMITY OF THE INFLOW CANNULA WAS SEEN. THE HOSPITAL WAS REQUESTING LOG FILE ANALYSIS TO CONFIRM POTENTIAL PUMP THROMBOSIS AND ROTOR CONDITION STATUS. CLINICAL CONDITION AND OTHER INFORMATION WERE REQUESTED TO BE COMPLETED. THERE WERE NO LOW FLOW ALARMS. THE LOG FILES WERE UNABLE TO BE PARSED AS SUBMITTED. IT WAS REQUESTED THAT THE LOG FILES BE DOWNLOADED AGAIN AND RE SENT. THE LOG FILES WERE THEN REVIEWED AND SHOWED POWER INCREASING ALONG WITH ROTOR NOISE AND DISPLACEMENT. THIS WAS SUSPECT OF THROMBUS. THE PUMP TEMPERATURE WAS NORMAL AT THE TIME. THE PUMP LOG FILE DATES WERE NOTED TO ALL BE FROM 2024 WHICH WAS SAID TO BE ODD. IT WAS NOTED THAT IN THE OPINION OF THE SITE DISSEMINATED INTRAVASCULAR COAGULATION (DIC) IN THE HYPERCOAGULABLE PHASE, TRIGGERED BY SEPSIS (DEEP DRIVELINE INFECTION) CONTRIBUTED TO THE EVENT. THERE WERE NO RECENT CHANGES TO PUMP PARAMETERS. DUE TO THE CONFIRMED DIAGNOSIS OF PUMP THROMBOSIS AND THE PATIENT'S CONDITION (ON MECHANICAL VENTILATION IN THE ICU), COMPUTED TOMOGRAPHY (CT) IMAGING WAS NOT PERFORMED. THERE WAS NO HEMATURIA OR OTHER SIGNS OF HEMOLYSIS. THE PATIENT DEVELOPED WORSENING SYMPTOMS OF HEART FAILURE AND RESPIRATORY INSUFFICIENCY AS WELL AS METABOLIC ACIDOSIS. THE PATIENT WAS SUBSEQUENTLY PLACED ON MECHANICAL VENTILATION. THE PATIENT RECEIVED ANTICOAGULATION THERAPY WITH HEPARIN. THE INTERNATIONAL NORMALIZED RATIO (INR) WAS MAINTAINED AT 2.4. THE PATIENT WAS UNDER CLOSE OBSERVATION AND THEY WERE ADDED TO THE URGENT TRANSPLANT WAITING LIST.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE WOUND CULTURE DONE ON (B)(6) 2022 SHOWED CULTURES OF BURKHOLDERIA CEPACIA ¿ 104 CFU, STAPHYLOCOCCUS AUREUS (MRSA) ¿ 104 CFU (MFR. REPORT # 2916596-2022-12807). INFECTION WAS TREATED WITH ANTIBACTERIAL THERAPY WITH SOME TEMPORARY SUCCESS. THERE WAS PROGRESSION INTO A DEEP DRIVELINE INFECTION. ON (B)(6) 2022 SURGICAL INCISION AND DRAINAGE OF THE ABSCESS WAS DONE IN THE LOWER THIRD OF THE STERNUM ALONG THE DRIVELINE PATHWAY. WITHIN THE 3 YEARS ADDITIONAL SURGICAL REVISIONS AND DEBRIDEMENTS WERE CARRIED OUT (MFR. REPORT #S 2916596-2022-14910, 2916596-2023-02374). THE PATIENT WAS PLACED ON THE URGENT TRANSPLANT LIST AS THE INFECTION HADN'T RESOLVED. THE RESPIRATORY INSUFFICIENCY WAS MULTIFACTORIAL, PRIMARILY DUE TO PROGRESSIVE HEART FAILURE ON THE BACKGROUND OF PUMP THROMBOSIS, ACCOMPANIED BY METABOLIC ACIDOSIS AND SYSTEMIC INFLAMMATORY RESPONSE ASSOCIATED WITH SEPSIS. THE PATIENT HAD SYMPTOMS OF PROGRESSIVE DYSPNEA, TACHYPNEA, HYPOXIA AND SIGNS OF OTHER RESPIRATORY DISTRESS. THEY DEVELOPED SEVERE METABOLIC ACIDOSIS AND REQUIRED INITIATION OF MECHANICAL VENTILATION. THE PATIENT WAS TREATED WITH SUPPLEMENTAL OXYGEN, FOLLOWED BY ENDOTRACHEAL INTUBATION AND MECHANICAL VENTILATION FROM THE PROGRESSIVE RESPIRATORY FAILURE. DESPITE THE SUPPORTIVE THERAPY THE RESPIRATORY INSUFFICIENCY CONTINUED AS IT WAS SECONDARY TO PUMP THROMBOSIS AND SYSTEMIC COMPLICATION OF SEPSIS. THIS LEAD TO THE PATIENT PASSING AWAY ON (B)(6) 2025. Manufacturer Narrative: SECTION E1: CUSTOMER SITE WAS NATIONAL RESEARCH CARDIA(B)(6), KAZAKHSTAN. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: ADDITIONAL INFO IN H6 HEALTH EFFECT CLINICAL CODES: 1918 - HYPOXIA, 1816 - DYSPNEA. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22012113&pc=OZD | 1220648-2025-28362 | 2025-04-24 04:00 | Death | ABIOMED, INC. | 2025-05-13 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate; Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT ON IMPELLA CP-EXTRACORPOREAL MEMBRANE OXYGENATION SUPPORT WAS BEING ESCALATED TO AN IMPELLA 5.5 FOR MORE SUPPORT WHEN THE TEAM OBSERVED A LEFT ATRIUM CLOT. THE IMPELLA 5.5 WAS STILL PLACED BUT, WHEN THE FLOWS WERE NOT AS EXPECTED THE LEFT VENTRICLE AND PUMP WERE ASSESSED AND WERE SEEN TO BE FULL OF THROMBUS BURDEN. THE IMPELLA 5.5 WAS REMOVED AND FLUSHED. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION AND WAS SHOCKED AND THE CLOTTING SITUATION PROMPTED THE TEAM TO WITHDRAW/STOP SUPPORT AND CARE. THE PATIENT EXPIRED. THE EUROPEAN TEAM COMMUNICATED THE PUMP WAS NOT TO BE THE CAUSE. Manufacturer Narrative: THE INVESTIGATION OF THROMBUS, LOW PUMP FLOW, AND VT/VF HAS BEEN COMPLETED. LOW FLOW: REPORTED PUMP UNABLE TO GENERATE FLOW WHEN STARTED. PUMP WAS REMOVED AND NO THROMBUS WAS NOTED ON THE PUMP AND RAN WELL IN WATER. IMPELLA 5.5 WAS RETURNED AND NO DAMAGES OR ABNORMALITIES WERE FOUND. DATA LOGS REVIEWED AND NO MOTOR CURRENT (MC) SPIKE, PLACEMENT SIGNAL TRENDS OR SUCTION ALARM OBSERVED. PUMP WAS UNABLE TO RUN WITH FLOWS AT ALMOST ZERO DURING INITIATION OF PUMP RUN. THE ROOT CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY PATIENT CONDITION RELATED SINCE PUMP WAS UNABLE TO RUN IN THE PATIENT BUT RAN WITHOUT ISSUE IN WATER AFTER IT WAS EXPLANTED AND ALSO PATIENT HAD VOLUME ISSUES WITH HUGE THROMBI BLOCKING AND COMPLETELY FILLED IN LEFT VENTRICLE (LV). THROMBUS: REPORTED LARGE THROMBUS IN LEFT VENTRICLE. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE THROMBUS WAS NOT DETERMINED VT/VF: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. B1 PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28362. H6 MEDICAL DEVICE PROBLEM CODE 1248 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28362. Manufacturer Narrative: THE IMPELLA PUMP AND PURGE CASSETTE WERE RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH.". IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22012750&pc=DSQ | 2916596-2025-03065 | 2025-04-17 04:00 | Death | THORATEC CORPORATION | 2025-05-13 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE ANY ADDITIONAL INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22012751&pc=HRS | 0001032347-2025-00208 | 2024-07-14 04:00 | Death | BIOMET MICROFIXATION | 2025-05-13 04:00 | HRS | UNK STERNALOCK PLATE | Insufficient Information | Insufficient Information | NI |  | 8 | Event Description: IT WAS REPORTED IN A JOURNAL ARTICLE FROM PERIOPERATIVE & CRITICAL CARE: RESEARCH, WHICH STUDIED CONTRIBUTIONS OF RIGID PLATE FIXATION ON POSTOPERATIVE PAIN, OPIOID USE, AND OTHER STERNOTOMY OUTCOMES. EIGHT PATIENTS HAD DEATH OCCUR WITHIN 90 DAYS OF OPERATION DUE TO AN UNKNOWN REASON. IT IS NOT KNOWN WHETHER THIS WAS A DEVICE RELATED COMPLICATION. ATTEMPTS HAVE BEEN MADE AND ADDITIONAL INFORMATION ON THE REPORTED EVENT IS UNAVAILABLE AT THIS TIME.Event Description: NO FURTHER EVENT INFORMATION IS AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: (B)(4). THE DEVICE WILL NOT BE RETURNED FOR ANALYSIS; HOWEVER, AN INVESTIGATION OF THE REPORTED EVENT IS IN PROGRESS. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL MEDWATCH 3500A WILL BE SUBMITTED. GERDISCH, MARC, JOHNS, CHANICE M, BARKSDALE, ANDREW, PARIKSHAK, MANESH. RIGID STERNAL FIXATION AND ENHANCED RECOVERY FOR OPIOID-FREE ANALGESIA AFTER CARDIAC SURGERY. ANN THORAC SURG. 2024;118(4):931-939. HTTPS://DOI.ORG/10.1016/J.ATHORACSUR.2024.06.032. Manufacturer Narrative: THIS FOLLOW-UP REPORT IS BEING SUBMITTED TO RELAY ADDITIONAL INFORMATION. THE FOLLOWING SECTIONS WERE UPDATED: B3, B4, B5, G3, G6, H2, H6, H11. D4: ATTEMPTS HAVE BEEN MADE TO GATHER ALL PRODUCT IDENTIFICATION INFORMATION AND NO FURTHER INFORMATION HAS BEEN PROVIDED. PRODUCT HAS NOT BEEN RETURNED. NO PRODUCT EVALUATION WAS ABLE TO BE COMPLETED. NO FURTHER ACTIONS WILL BE INITIATED AS A RESULT OF REPORTED EVENT. ROOT CAUSE OF THE REPORTED EVENT IS NOT DEVICE RELATED. IT WAS REPORTED THAT 8 PATIENTS IN THE STUDY EXPIRED WITHIN 90 DAYS POSTOPERATIVELY. PATIENTS IN THE STUDY WERE UNDERGOING HIGH RISK CARDIAC SURGERIES AND CLOSED WITH STERNAL FIXATION DEVICES. THESE DEVICES ARE USED AS ADJUNCT THERAPY AND PLACED AT THE END OF THE PROCEDURE TO SUPPORT PROPER BONE FIXATION AND HEALING. THE AUTHOR CONFIRMED THAT NO DEVICE RELATED COMPLICATIONS OCCURRED. THEREFORE, PATIENT OUTCOMES ARE ATTRIBUTED TO THE CARDIAC PROCEDURE OR UNDERLYING PATIENT CARDIAC COMORBIDITIES. IF ANY FURTHER INFORMATION IS FOUND WHICH WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22012793&pc=DSQ | 2916596-2025-02760 | 2022-01-01 05:00 | Death | THORATEC CORPORATION | 2025-05-13 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿CLINICAL OUTCOMES OF HEARTMATE 3 LVAD RECIPIENTS INELIGIBLE FOR HEART TRANSPLANTATION DUE TO SOCIAL FACTORS¿ THAT THE HEARTMATE 3 COULD BE ASSOCIATED WITH GASTROINTESTINAL (GI) BLEEDING, STROKE, DRIVELINE INFECTIONS, AND DEATH. A COHORT OF SUBJECTS WHO HAD UNDERGONE HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION WERE PREDOMINANTLY MALE (94%), AND DIVERSE (62.6% WHITE, 31.2% BLACK, 6.2%HISPANIC). THE MEAN AGE OF THE SUBJECTS WAS 56.5 ± 8.8 YEARS, AND THEY HAD BMI OF 29.0 ± 4.1, WITH AN AVERAGE FOLLOW-UP TIME OF 737.3 ± 478 DAYS FROM LVAD IMPLANTATION. THE MEAN INTERMACS PROFILE WAS 1.875 ± 0.96, WHERE 50%,12.5%, 37.5% OF SUBJECTS WERE INTERMACS PROFILE I, II, III RESPECTIVELY. COMPLICATIONS INCLUDED GI BLEEDING (18.75%), STROKE (12.50%), AND INFECTION (31.25%). THE 1-AND 4-YEAR SURVIVAL WAS 93.75% AND 87.5% RESPECTIVELY. THE AVERAGE HOSPITAL READMISSION PER SUBJECT OVER THE FOLLOW UP PERIOD WAS 8.2. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBERS WERE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DATE OF EVENT IS APPROXIMATE AS THE DATA WAS COLLECTED BETWEEN 2019 AND 2022. DEVICE WAS IMPLANTED AT TIME OF EVENT. AUTHOR INFORMATION: CHIA, R., GIYANANI, N., KAZEMIAN, P., ICE, D., ROSS, R., MOSHIYAKHOV, M., & BARN, K. (2025). CLINICAL OUTCOMES OF HEARTMATE 3 LVAD RECIPIENTS INELIGIBLE FOR HEART TRANSPLANTATION DUE TO SOCIAL FACTORS. JOURNAL OF CARDIAC FAILURE, 31(1), 231. HTTPS://DOI.ORG/10.1016/J.CARDFAIL.2024.10.132. (B)(6). NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RESEARCH ARTICLE ¿CLINICAL OUTCOMES OF HEARTMATE 3 LVAD RECIPIENTS INELIGIBLE FOR HEART TRANSPLANTATION DUE TO SOCIAL FACTORS¿ WAS RECEIVED. THE STUDY SOUGHT TO EVALUATE THE CLINICAL OUTCOMES OF HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE (LVAD) RECIPIENTS WHO WERE INELIGIBLE FOR TRANSPLANTATION DUE TO SOCIAL CRITERIA, EMPHASIZING THEIR PROMISING POTENTIAL AS A COMPARABLE ALTERNATIVE TO HEART TRANSPLANTATION. THE STUDY CONCLUDED THAT FOR PATIENTS INELIGIBLE FOR HEART TRANSPLANTATION, LVADS REPRESENT A VIABLE, LIFE-EXTENDING ALTERNATIVE. THE DATA SUGGESTS LVAD THERAPY AS A POTENTIALLY COMPARABLE OPTION TO HEART TRANSPLANTATION, PARTICULARLY WITH ADVANCEMENTS IN DEVICE TECHNOLOGY AND PATIENT MANAGEMENT STRATEGIES. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE HEARTMATE 3 DEVICE SERIAL NUMBERS AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE HEARTMATE 3 LVAS IFU, REVISION D, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22012899&pc=OZD | 1220648-2025-28369 | 2025-04-27 04:00 | Death | ABIOMED, INC. | 2025-05-13 04:00 | OZD | IMPELLA CP | Malposition of Device | Cardiac Arrest; Hemolysis; Renal Failure | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND EXHIBITED A DROP IN P LEVELS DURING PUMP REPOSITIONING. THE PATIENT BECAME NON-PULSITILE AND A CODE BLUE WAS INITIATED. THE PATIENT WAS DEFIBRILLATED AND RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED. REPOSITIONING OF THE IMPLELLA WAS ATTEMPTED TO RETRACT THE CATHETER AS AN ECHOCARDIOGRAM SHOWED SHALLOW PLACEMENT BUT THE DECISION WAS MADE TO LEAVE THE IMPELLA CP AT FIVE CENTIMETERS WITHIN A DILATED LEFT VENTRICLE. THE PATIENT EXHIBITED DARK COLORED URINE AND WAS PLACED ON CONTINUOUS DIALYSIS. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES, VF/VT/AF/AT, HEMOLYSIS, AND RENAL FAILURE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL AND THE DATA LOG ANALYSIS, THE ROOT CAUSE OF THE HEMOLYSIS IS MOST LIKELY DUE TO IMPROPER POSITIONING. DUE TO LACK OF CLINICAL DETAILS PROVIDED, IT IS UNKNOWN OF WHY THE PUMP CAME OUT OF POSITION. THE CAUSE OF THE RENAL FAILURE AND VT/VF WERE NOT DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. B1 PRODUCT PROBLEM SELECTION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28369.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22013049&pc=LWP | 2017865-2025-53510 | 2025-04-21 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-13 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR MAGNETIC RESONANCE IMAGING (MRI) SCAN. THE MRI STAFF REQUESTED THE PULSE GENERATOR REPROGRAMMING FOR MRI COMPATIBILITY. THE DEVICE WAS SET TO VENTRICULAR PACING ONLY MODE AT 85 BEATS PER MINUTE PROVIDING 100% PACING SUPPORT DURING THE SCAN. DURING THE MRI THE PATIENT BECAME UNSTABLE, CODED, AND RESUSCITATION WAS ATTEMPTED. THE PATIENT LATER EXPIRED; THERE IS NO ALLEGATION AT THIS TIME THAT THE THAT THE GENERATOR OR RIGHT VENTRICULAR (RV) LEAD MALFUNCTION CONTRIBUTED TO THE PATIENT¿S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22013716&pc=MKJ | 3030677-2025-001327 | 2025-05-08 04:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-05-13 04:00 | MKJ | PHILIPS HEARTSTART FRX | Failure of Device to Self-Test | Cardiac Arrest | P180028 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT THE DEVICE EXHIBITED AN ERROR DURING A PATIENT USE EVENT. THE PATIENT DID NOT SURVIVE. Manufacturer Narrative: UPDATED REPORTER ADDRESS AND NAME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22013752&pc=MVK | 3008642652-2025-04785 | 2025-01-17 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-13 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 18:33:23 ON (B)(6) 2025. AT 07:56:21 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE TRANSITIONING TO VT FROM 130-180 BPM WITH VARYING AMPLITUDES AND CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. THE DEVICE PROPERLY DETECTED VT. VARYING AMPLITUDES PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE DEVICE WAS SHUT DOWN AT 18:51:09 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22014319&pc=LWS | 2649622-2025-13008 | 2025-04-19 04:00 | Death | MPRI | 2025-05-13 04:00 | LWS | SPRINT QUATTRO SECURE MRI SURESCAN | Under-Sensing | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE ARRYTHMIA WAS BELOW THE LOWEST DETECTION ZONE AND THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING OF THE VENTRICULAR FIBRILLATION (VF) THAT WAS REPORTED TO HAVE CONTRIBUTED TO THE FAILED RESCUE SHOCKS AND THE PATIENT DEATH. IT WAS NOTED THAT THE RHYTHM WAS LIKELY AGONAL AND THERE WAS NO ATRIAL MECHANISM, THE VENTRICULAR RATE WAS SLOW AND DISORGANIZED WITH OCCASIONAL UNDERSENSING. Manufacturer Narrative: CONTINUATION OF D10: DTPC2D1 CRT-D IMPLANT DATE: (B)(6) 2023, 419688 LEAD IMPLANT DATE: (B)(6) 2011. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22014410&pc=DSQ | 2916596-2025-03133 | 2021-01-01 05:00 | Death | THORATEC CORPORATION | 2025-05-13 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿HEMOCOMPATIBILITY- RELATED ADVERSE EVENTS ASSOCIATED WITH OR WITHOUT LOW MOLECULAR WEIGHT HEPARIN BRIDGING IN OUTPATIENTS WITH A HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE¿, THAT THE HM3 MAY BE ASSOCIATED WITH DEATH, BLEEDING, AND THROMBOTIC EVENTS. A SINGLE- CENTER, RETROSPECTIVE, ELECTRONIC MEDICAL RECORD REVIEW OF ALL ADULT PATIENTS WITH A HEARTMATE 3 IMPLANTED FROM 1JAN2015 TO 1JAN 2021 WAS PERFORMED. PATIENTS WERE CATEGORIZED INTO TWO GROUPS: THOSE WHO RECEIVED OUTPATIENT BRIDGING WITH ENOXAPARIN FOR INR VALUES = 1.7 AND THOSE WHO DID NOT. INCLUDED PATIENTS WERE REQUIRED TO BE ON HM3 SUPPORT FOR AT LEAST 18 MONTHS AND WERE EXCLUDED IF THEY DIED PRIOR TO INDEX DISCHARGE, STOPPED WARFARIN FOR > 30 CONSECUTIVE DAYS OF THE FOLLOW- UP PERIOD, HAD AN INR GOAL OTHER THAN 2¿3 OR 2¿2.5, HAD NO RECORDED INR FOR > 90 CONSECUTIVE DAYS, OR HAD A DIAGNOSIS OF HEPARIN-INDUCED THROMBOCYTOPENIA. THE PRIMARY OUTCOME WAS A COMPOSITE OF BLEEDING AND THROMBOTIC EVENTS (HEMOCOMPATIBILITY- RELATED ADVERSE EVENTS [HRAES]) BETWEEN PATIENTS WHO WERE BRIDGED WITH ENOXAPARIN AND THOSE WHO WERE NOT, RANGING IN SEVERITY FROM NUISANCE TO FATAL BLEEDING. BRIDGING OPPORTUNITIES (BOS) WERE DEFINED AS INRS = 1.7. EACH PATIENT COULD CONTRIBUTE MORE THAN ONE BO DURING THE FOLLOW- UP PERIOD. HEMOCOMPATIBILITY- RELATED ADVERSE EVENTS WERE COLLECTED FOR 30 DAYS FOLLOWING THE END OF EACH BO. ONE HUNDRED TWENTY PATIENTS WERE EVALUATED. FIFTY- ONE PATIENTS WERE EXCLUDED, PRIMARILY DUE TO DEATH PRIOR TO INDEX HOSPITAL DISCHARGE OR HAVING THEIR HM3 FOR < 18 MONTHS. THERE WAS AN AVERAGE OF 13 (RANGE 2¿33) BOS PER PATIENT. OF THE INCLUDED PATIENTS, 15 RECEIVED ENOXAPARIN BRIDGING DURING = 1 BO, AND 64 DID NOT. OF THE 1036 TOTAL BOS, ONLY 39 (3.76%) WERE BRIDGED WITH ENOXAPARIN. IN THE BRIDGED GROUP, THERE WERE 39 BOS RESULTING IN NO HRAES, AS COMPARED TO 997 BOS RESULTING IN 12 HRAES (1.20%) IN THE NON- BRIDGED GROUP. OF THE 12 HRAES IN THE NON- BRIDGED GROUP, ONLY 2 (0.2%) WERE THROMBOTIC, BOTH BEING MINOR VENOUS THROMBOEMBOLISMS (VTES). NO EPISODES OF CONFIRMED OR SUSPECTED PUMP THROMBOSIS OR STROKES OCCURRED IN EITHER GROUP. A TOTAL OF 34 BLEEDING EVENTS OF ANY SEVERITY OCCURRED IN THE NON- BRIDGED GROUP (0.034 EVENTS PER BO) AS COMPARED TO 2 IN THE BRIDGED GROUP (0.051 EVENTS PER BO). MAJOR BLEEDING EVENTS IN THE NON- BRIDGED GROUP WERE EITHER GASTROINTESTINAL OR DRIVELINE SITE BLEEDING. BOTH BLEEDING EVENTS IN THE BRIDGED GROUP WERE CLASSIFIED AS TYPE 1 (MINOR BLEEDING). ACROSS BOTH GROUPS, THERE WERE A TOTAL OF 38 HRAES AND A TOTAL OF 0.32 HRAE EVENTS PER PATIENT- YEAR. IN SUMMARY, THE EXCEEDINGLY LOW THROMBOTIC EVENT RATE FOUND IN PATIENTS WITH A HM3 AND LOW RATES OF BACKGROUND ASPIRIN USE WHO WERE NOT BRIDGED WITH ENOXAPARIN DURING PERIODS OF LOW INR CALLS INTO SIGNIFICANT QUESTION THE MERITS OF BRIDGING PATIENTS WITH ENOXAPARIN AS A ROUTINE PRACTICE. Manufacturer Narrative: SECTION B3: DATE OF EVENT IS APPROXIMATE (01JAN2021) AS THE DATA WERE COLLECTED BETWEEN 01JAN2015 TO 01JAN 2021. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. AUTHOR INFORMATION: JAMES, J. R., FULLER, M., HUNT, M. H., POWELL, E., LU, G., LI, Q., OMARY, A., & HOLLIS, I. B. (2025). HEMOCOMPATIBILITY-RELATED ADVERSE EVENTS ASSOCIATED WITH OR WITHOUT LOW MOLECULAR WEIGHT HEPARIN BRIDGING IN OUTPATIENTS WITH A HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE. ARTIFICIAL ORGANS. HTTPS://DOI.ORG/10.1111/AOR.14986. DEPARTMENT OF PHARMACY, UNIVERSITY OF NORTH CAROLINA HOSPITALS, CHAPEL HILL, NORTH CAROLINA, USA ¦ DEPARTMENT OF PHARMACY, WAKEMED HEALTH & HOSPITALS, RALEIGH, NORTH CAROLINA, USA ¦ UNIVERSITY OF NORTH CAROLINA ESHELMAN SCHOOL OF PHARMACY, CHAPEL HILL, NORTH CAROLINA, USA ¦ DEPARTMENT OF PHARMACY, NOVANT HEALTH PRESBYTERIAN MEDICAL CENTER, HUNTERSVILLE, NORTH CAROLINA, USA ¦ DEPARTMENT OF BIOSTATISTICS, UNIVERSITY OF NORTH CAROLINA GILLINGS SCHOOL OF GLOBAL PUBLIC HEALTH, CHAPEL HILL, NORTH CAROLINA, USA. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS IN USE WERE NOT PROVIDED. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22014873&pc=NIK | 3004209178-2025-08287 |  | Death | MEDTRONIC EUROPE SARL | 2025-05-13 04:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 23 DAYS FOLLOWING IMPLANT OF THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) THE PATIENT EXPIRED. ADDITIONAL INFORMATION HAS BEEN REQUESTED AND HAS NOT YIELDED ANY RESULTS. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22014965&pc=LWS | 2649622-2025-13023 | 2024-11-19 05:00 | Death | MPRI | 2025-05-13 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED POSSIBLE OVERSENSING. IT WAS FURTHER REPORTED THAT THE PATIENT PASSED APPROXIMATELY FIVE MONTHS LATER. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22015001&pc=QZK | 2017233-2025-06181 | 2025-04-11 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-05-13 04:00 | QZK | GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Aneurysm | P230023 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED: ON (B)(6) 2025, THE PATIENT UNDERWENT TREATMENT OF AN ANASTOMOTIC AORTIC ANEURYSM IN ZONE 5 USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE). THE PATIENT TOLERATED THE PROCEDURE. REPORTEDLY THE PATIENT HAD TWO PREVIOUS OPEN ANEURYSM REPAIRS, OPEN ARCH AND THORACOABDOMINAL. THERE WAS DEGENERATION OF THE NATIVE AORTA AT THE ANASTOMOSIS SITE. IT WAS REPORTED THAT ON (B)(6) 2025, THE PATIENT WENT INTO SPONTANEOUS PULSELESS ELECTRICAL ACTIVITY (PEA) ARREST AND EXPIRED. THE PHYSICIAN IS NOT ATTRIBUTING THE DEATH TO THE GORE DEVICE BUT STATED THE PEA ARREST MAY HAVE BEEN BROUGHT ON FROM THE ANESTHESIA. Manufacturer Narrative: PATIENT MEDICATIONS: NORVASC, ASPIRIN, CHOLECALCIFEROL, CATAPRES, PEPCID, AND LOPRESSOR. ACCORDING TO THE GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) INSTRUCTIONS FOR USE (IFU), COMPLICATIONS ASSOCIATED WITH USE OF THE GORE® TAG® THORACIC ENDOPROSTHESIS MAY INCLUDE BUT ARE NOT LIMITED TO DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY HAVE NOT BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22015277&pc=MHX | 8030229-2025-04950 | 2024-07-03 04:00 | Death | NIHON KOHDEN CORPORATION | 2025-05-13 04:00 | MHX | PU-681RA | No Audible Alarm; Protective Measures Problem | Arrhythmia; Asystole | K102376 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT THE PATIENT PASSED AWAY WHILE BEING MONITORED AT THE CENTRAL NURSE'S STATION (CNS). Manufacturer Narrative: DETAILS OF COMPLAINT: THE CUSTOMER REPORTED THAT THE PATIENT PASSED AWAY WHILE BEING MONITORED AT THE CENTRAL NURSE'S STATION (CNS). THEY REQUESTED AN INVESTIGATION ON THE LOGS FROM THE DEVICES IN USE AT THE TIME OF DEATH. IT WAS FOUND THAT AN ASYSTOLE OCCURRED AT THE RELEVANT TIME. INVESTIGATION SUMMARY: THE DEVICE LOGS AND ECG PRINTS WERE REVIEWED BY NKC UNDER (B)(4). IT WAS FOUND THAT THE DEVICE ALARMED APPROPRIATELY FOR ASYSTOLE AT THE TIME OF THE PATIENT EVENT, AND THAT THE USER SILENCED THE ALARM. NK WILL CONTINUE TO MONITOR AND TREND. ADDITIONAL DEVICE INFORMATION: D10 CONCOMITANT MEDICAL DEVICE: THE FOLLOWING DEVICE WAS USED IN CONJUNCTION WITH THE CNS: TELEMETRY TRANSMITTER: MODEL #: ZM-531PA SERIAL #: (B)(6) DEVICE MANUFACTURER DATA: 09/25/2020 UNIQUE IDENTIFIER (UDI) #: (B)(4) RETURNED TO NIHON KOHDEN: NA.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22015369&pc=DSQ | 2916596-2025-02962 | 2025-04-10 04:00 | Death | THORATEC CORPORATION | 2025-05-13 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CAUSE OF DEATH WAS REPORTED AS UNKNOWN AND WHETHER THE PATIENT¿S OUTCOME WAS CONSIDERED TO BE DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22015383&pc=DSQ | 2916596-2025-02898 | 2025-04-05 04:00 | Death | THORATEC CORPORATION | 2025-05-13 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER.Event Description: THE PATIENT WAS ON HOSPICE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22015729&pc=NPT | 2025587-2025-03646 | 2025-05-12 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-13 04:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Activation; Activation Failure; Patient Device Interaction Problem | Atrial Fibrillation; Low Blood Pressure/ Hypotension; Tachycardia; Ventricular Fibrillation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS 29 MM TRANSCATHETER BIOPROSTHETIC VALVE, A PRE-DILATION WAS PERFORMED WITH A 25 MM N ON-MEDTRONIC BALLOON. THE PATIENT WAS HYPOTENSIVE AND IN FIRST DEGREE ATRIO-VENTRICULAR (AV) BLOCK SUPPORTED BY ANESTHESIA. THE DELIVERY CATHETER SYSTEM (DCS) WAS INSERTED AND UPON THE FIRST DEPLOYMENT WITH RAPID PACING, THE VALVE DID NOT REACH THE LEFT CORONARY CUSP (LCC). THE VALVE WAS FULLY RECAPTURED FOR MECHANICAL EXPANSION UNDER RAPID PACING. THE VALVE WAS DEPLOYED A SECOND TIME TO 80% AND THE PATIENT WAS UNSTABLE WITH PERSISTENT HYPOTENSION. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. A 34 MM VALVE WAS REQUESTED TO BE OPENED AND PREPARED. THE PATIENT FLUCTUATED BETWEEN ATRIAL FIBRILLATION (AFIB), VENTRICULAR TACHYCARDIA, AND VENTRICULAR FIBRILLATION DURING THE MULTIPLE ROUNDS OF CPR. DUE TO THE PATIENT INSTABILITY, THE DECISION WAS MADE TO DEPLOY THE 29 MM VALVE. THE VALVE WAS COMPLETELY DEPLOYED IN A 2 CUSP VIEW AND CPR RESUMED. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) CONFIRMED THE VALVE WAS WELL SEATED WITH NO LEFT VENTRICLE PERFORATION AND NO AORTIC TEAR. MULTIPLE ROUNDS OF CPR CONTINUED. THE PATIENT WAS IN PULSELESS ELECTRICAL ACTIVITY (PEA) AND DIED.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS 29 MM TRANSCATHETER BIOPROSTHETIC VALVE, A PRE-DILATION WAS PERFORMED WITH A 25 MM N ON-MEDTRONIC BALLOON. THE PATIENT WAS HYPOTENSIVE AND IN FIRST DEGREE ATRIO-VENTRICULAR (AV) BLOCK SUPPORTED BY ANESTHESIA. THE DELIVERY CATHETER SYSTEM (DCS) WAS INSERTED AND UPON THE FIRST DEPLOYMENT WITH RAPID PACING, THE VALVE DID NOT REACH THE LEFT CORONARY CUSP (LCC). THE VALVE WAS FULLY RECAPTURED FOR MECHANICAL EXPANSION UNDER RAPID PACING. THE VALVE WAS DEPLOYED A SECOND TIME TO 80% AND THE PATIENT WAS UNSTABLE WITH PERSISTENT HYPOTENSION. CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. A 34 MM VALVE WAS REQUESTED TO BE OPENED AND PREPARED. THE PATIENT FLUCTUATED BETWEEN ATRIAL FIBRILLATION (AFIB), VENTRICULAR TACHYCARDIA, AND VENTRICULAR FIBRILLATION DURING THE MULTIPLE ROUNDS OF CPR. DUE TO THE PATIENT INSTABILITY, THE DECISION WAS MADE TO DEPLOY THE 29 MM VALVE. THE VALVE WAS COMPLETELY DEPLOYED IN A 2 CUSP VIEW AND CPR RESUMED. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) CONFIRMED THE VALVE WAS WELL SEATED WITH NO LEFT VENTRICLE PERFORATION AND NO AORTIC TEAR. MULTIPLE ROUNDS OF CPR CONTINUED. THE PATIENT WAS IN PULSELESS ELECTRICAL ACTIVITY (PEA) AND DIED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE VALVE DID NOT REACH THE LCC DURING FIRST DEPLOYMENT BECAUSE THE VALVE HAD 7% OVERSIZING AND REQUIRED MECHANICAL EXPANSION. AFTER BEING RECAPTURED, THE VALVE FULLY EXPANDED WITH DEPLOYMENT. PER THE PHYSICIAN, THE PATIENT WAS ALREADY HYPOTENSIVE AFTER THE PRE-DILATATION BUT CONTINUED TO DECOMPENSATE THROUGHOUT THE PROCEDURE. THE PHYSICIAN DID NOT ATTRIBUTE THE DCS AS A CONTRIBUTING CAUSE TO THE DEATH BUT MULTIPLE RECAPTURES COULD HAVE POTENTIALLY CONTRIBUTED TO THE CAUSE OF DECOMPENSATION ALONG WITH A MEDICAL HISTORY OF HEART FAILURE, PULMONARY EDEMA, LEFT VENTRICULAR DYSFUNCTION, AND LOW FLOW LOW GRADIENT. PER THE PHYSICIAN, IF THE PATIENT HAD NOT HAD THE TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) THEY WOULD NOT HAVE DIED. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: E VFXPLUS-29, SERIAL/LOT #: (B)(6), UBD: 06-DEC-2026, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED: B1. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22001474&pc=MNS | 2518422-2025-105480 | 2025-05-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-12 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22001513&pc=OZD | 1220648-2025-28291 | 2024-11-27 05:00 | Death | ABIOMED, INC | 2025-05-12 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Anemia; Hematoma; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: IT WAS REPORTED THROUGH ABIOMED'S LOQI STUDY DATABASE THAT A PATIENT IMPLANTED WITH AN IMPELLA 5.5 EXHIBITED ANEMIA. THE RELATIONSHIP IS NOTED AS "POSSIBLY" RELATED TO THE 5.5 PUMP SUPPORT. LOQI NOTED: THAT HEMOGLOBIN DROPPED FROM 14.7 TO 11.1. IT WAS FOUND TO HAVE HEMATOMA AND LIKELY INITIAL CAUSE OF ACUTE ANEMIA. ANEMIA DID NOT IMPROVE POST HEMATOMA EVACUATIONS. PATIENT ALSO HAD RECURRENT GI BLEEDING SECONDARY TO PERFORATED DUODENUM. PATIENT HAD MULTIPLE BLOOD PRODUCTS THROUGHOUT HOSPITAL STAY. THE PATIENT EXPIRED. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER, ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: ACCESS SITE BLEEDING: THE ROOT CAUSE OF ACCESS SITE BLEEDING CANNOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. VASCULAR DAMAGE PERIPHERAL VESSEL: THE ROOT CAUSE OF VASCULAR DAMAGE PERIPHERAL VESSEL CANNOT BE DETERMINED SINCE THE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS WERE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22001710&pc=MSX | 2221819-2025-00006 | 2025-04-12 04:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-05-12 04:00 | MSX | BENEVISION DISTRIBUTED MONITORING SYSTEM (DMS) | No Device Output | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, THE DATA FROM A PATIENT CONNECTED TO A TD60 TELEMETRY UNIT (SERIAL NUMBER (B)(6)) ID (B)(6) WAS LOST. THERE WAS NO POWER OUTAGE DURING THIS TIME. THE CUSTOMER IS REQUESTING TO EVALUATE IF THE DATA WAS DELETED OR IF THE PATIENT WAS NOT ADMITTED/DISCHARGED PROPERLY. THE PATIENT EXPIRED.Event Description: IT WAS REPORTED THAT ON APRIL 12, 2025, THE DATA FROM A PATIENT CONNECTED TO A TD60 TELEMETRY UNIT (SERIAL NUMBER (B)(6)) ID 5N TELE 2 IN ROOM 502 WAS LOST. THERE WAS NO POWER OUTAGE DURING THIS TIME. THE CUSTOMER IS REQUESTING TO EVALUATE IF THE DATA WAS DELETED OR IF THE PATIENT WAS NOT ADMITTED/DISCHARGED PROPERLY. THE PATIENT EXPIRED. Manufacturer Narrative: SYSTEM LOGS WERE COLLECTED BY MINDRAY FIELD SERVICE REPRESENTATIVE. UNDER INVESTIGATION. Manufacturer Narrative: MINDRAY CONDUCTED A REVIEW OF THE BENEVISION CMS LOGS FOR PATIENTS ADMITTED BETWEEN (B)(6) 2025 AND DISCHARGED ON (B)(6) 2025. THE REVIEW FOCUSED ON PATIENT INFORMATION DELETION RECORDS, BED NUMBERS, AND EQUIPMENT SERIAL NUMBERS, AND DID NOT IDENTIFY ANY INFORMATION RELATED TO THE REPORTED OCCURRENCE INVOLVING THE PATIENT WITH ID 5N TELE 2 IN ROOM 502. IT WAS DETERMINED THAT THE TD60 TELEMETRY DEVICE WAS LIKELY NOT CONNECTED TO THE CENTRAL STATION. AS NO LOGS FOR THE TD60 WERE AVAILABLE AFTER MULTIPLE ATTEMPTS, THE REPORTED ISSUE COULD NOT BE CONFIRMED. MINDRAY BENEVISION CMS WITH TD60 TELEMETRY FUNCTIONED ACCORDING TO SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22001715&pc=BZD | 2518422-2025-105482 | 2025-05-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-12 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22001758&pc=OZD | 1220648-2025-28294 | 2025-04-20 04:00 | Death | ABIOMED, INC. | 2025-05-12 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND HAD A PULMONARY HEMORRHAGE DURING A BRONCHOSCOPY. HEPARIN DRIP WAS HELD TO MANAGE THE CONDITION. TWO DAYS LATER, SUCTION ALARMS WERE ENCOUNTERED AT HIGHER P-LEVELS. HEPARIN ADMINISTRATION WAS INCREASED IN RESPONSE TO THE LOW FLOW, HOWEVER, THE PATIENT HAD A DO NOT RESUSCITATE ORDER AND PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. VASCULAR DAMAGE - PERIPHERAL VESSEL: IT WAS STATED THAT THERE WAS A PULMONARY HEMORRHAGE FOUND DURING A BRONCHOSCOPY. DUE TO LACK OF CLINICAL DETAILS, THE ROOT CAUSE OF THE VASCULAR DAMAGE - PERIPHERAL VESSEL CANNOT BE DETERMINED. LOW PUMP FLOW: ON 4/22 THERE WAS SUCTION ALARMS AND LOWER FLOWS. THERE WAS A TREND IN THE MOTOR CURRENT AND PLACEMENT SIGNAL. THE CLINICAL STATED THAT HEPARIN WAS ADDED TO INCREASE THE RESPONSE TO THE LOW PUMP FLOW BUT THE PATIENT WAS MADE DO NOT RESUSCITATE (DNR) AND PASSED AWAY THE FOLLOWING DAY. BASED ON THE CLINICAL DETAILS AND DATA LOG ANALYSIS, THE ROOT CAUSE OF THE LOW PUMP FLOW IS MOST LIKELY DUE TO THE PATIENTS CONDITION. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22002054&pc=LWQ | 2135147-2025-02495 | 2025-04-11 04:00 | Death | ABBOTT MEDICAL | 2025-05-12 04:00 | LWQ | REGENT HEART VALVE WITH FLEX CUFF | Physical Resistance/Sticking | Perforation | P810002 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A REGENT IMPLANT PROCEDURE OCCURRED. DURING THE PROCEDURE, THE CHIEF SURGEON SUTURED AN ARTIFICIAL BLOOD VESSEL AND ENLARGED THE ROOT OF THE AORTIC VALVE ANNULUS. SUBSEQUENTLY, THE CHIEF SURGEON MEASURED THE SIZE OF THE ANNULUS USING STANDARD METHODS AND DECIDED TO IMPLANT THE 23MM REGENT MECHANICAL VALVE. AFTER SUTURING, IT WAS FOUND THAT THE VALVE LEAFLETS COULD NOT BE FULLY OPENED. THE CHIEF SURGEON REPLACED THE VALVE WITH ANOTHER BRAND OF MECHANICAL VALVE TO COMPLETE THE IMPLANTATION. THERE WERE SOME CORRELATIONS WITH HEMODYNAMIC INSTABILITY AND PROCEDURAL DELAY DUE TO THE LEAFLET ISSUE. AFTER THE SURGERY, ON THE SAME DAY, THE PATIENT DIED DUE TO A RUPTURED HEART. THE PHYSICIAN DOESN'T BELIEVE THE PATIENT OR USER EXPERIENCED ADVERSE HEALTH CONSEQUENCES DUE TO THE PERFORMANCE OF THE PRODUCT. THEY BELIEVE THE PATIENT'S POSTOPERATIVE DEATH IS RELATED TO THE NON-ABBOTT VALVE, THE PROCEDURE, OR THE PATIENT'S COMORBIDITIES.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A REGENT IMPLANT PROCEDURE OCCURRED. DURING THE PROCEDURE, THE CHIEF SURGEON SUTURED AN ARTIFICIAL BLOOD VESSEL AND ENLARGED THE ROOT OF THE AORTIC VALVE ANNULUS. SUBSEQUENTLY, THE CHIEF SURGEON MEASURED THE SIZE OF THE ANNULUS USING STANDARD METHODS AND DECIDED TO IMPLANT THE 23MM REGENT MECHANICAL VALVE. AFTER SUTURING, IT WAS FOUND THAT THE VALVE LEAFLETS COULD NOT BE FULLY OPENED. THE CHIEF SURGEON REPLACED THE VALVE WITH ANOTHER BRAND OF MECHANICAL VALVE TO COMPLETE THE IMPLANTATION. THERE WERE SOME CORRELATIONS WITH HEMODYNAMIC INSTABILITY AND PROCEDURAL DELAY DUE TO THE LEAFLET ISSUE. AFTER THE SURGERY, ON THE SAME DAY, THE PATIENT DIED DUE TO A RUPTURED HEART. THE PHYSICIAN DOESN'T BELIEVE THE PATIENT OR USER EXPERIENCED ADVERSE HEALTH CONSEQUENCES DUE TO THE PERFORMANCE OF THE PRODUCT. THEY BELIEVE THE PATIENT'S POSTOPERATIVE DEATH IS RELATED TO THE NON-ABBOTT VALVE, THE PROCEDURE, OR THE PATIENT'S COMORBIDITIES. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT OF PERFORATION, VALVE LEAFLETS NOT OPENING AND CLOSING DURING IMPLANT AND PATIENT DEATH WAS REPORTED. THE DEVICE WAS RETURNED TO ABBOTT FOR INVESTIGATION. NO ANOMALIES WERE FOUND WITH THE VALVE LEAFLETS WITH BOTH LEAFLETS ABLE TO FULLY OPEN AND CLOSE WITHOUT RESISTANCE. THE VALVE WAS ABLE TO BE ROTATED WITHOUT DIFFICULTY. HYDRODYNAMIC EXAMINATION INDICATED THE VALVE MET FUNCTIONAL SPECIFICATION. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, INCLUDING FUNCTIONAL TESTING, AND THE PRODUCT MET ALL SPECIFICATIONS. THE EXACT CAUSE OF THE INCOMPLETE COAPTATION COULD NOT BE CONCLUSIVELY DETERMINED. THE CAUSE OF PATIENT'S POSTOPERATIVE DEATH IS RELATED TO THE NON-ABBOTT VALVE, THE PROCEDURE, OR THE PATIENT'S COMORBIDITIES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. FROM MEDICAL REVIEW: THE CAUSE OF DEATH IS PROCEDURE RELATED DUE TO A RUPTURED HEART WHICH MOST LIKELY CAUSED SIGNIFICANT BLEEDING WITH THE NEED FOR EMERGENCY INTERVENTION AND DIFFICULTY RESCUING THE PATIENT WHICH ULTIMATELY LED TO THE DEATH. BLEEDING FROM THE AORTIC SUTURE LINE FOLLOWING AN AORTIC ANNULAR ENLARGEMENT PROCEDURE IS A KNOWN POTENTIAL COMPLICATION. ALTHOUGH ADDITIONAL SURGERY TIME WAS REQUIRED TO REPLACE THE REGENT 23MM MHV, IT IS UNLIKELY TO HAVE CAUSED THE REPORTED CARDIAC RUPTURE, PARTICULARLY SINCE THE REGENT VALVE WAS REMOVED AND A NON-ABBOTT MHV WAS SUBSEQUENTLY IMPLANTED. THE DIFFICULTY IN THE OPENING OF THE REGENT 23MM VALVE LEAFLETS MOST LIKELY WAS A RESULT OF OVERSIZING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22002068&pc=OZD | 1220648-2025-28296 | 2025-02-16 05:00 | Death | ABIOMED, INC. | 2025-05-12 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A 57-YEAR-OLD FEMALE ON A IMPELLA 5.5 FOR CARDIOMYOPATHY. DURING SUPPORT THE PATIENT HAD HEMOLYZED LABS AND A SUCTION EVENT WHICH THEY GAVE VOLUME. IT WAS NOTED THAT THE PATIENT WAS NOT A CANDIDATE FOR TRANSPLANT. CARE WAS ULTIMATELY WITHDRAWN AND THE PATIENT EXPIRED. THE PATIENT WAS ON IMPELLA SUPPORT FOR A TOTAL OF 2,288.40 HOURS. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. THE PATIENT'S OUTCOME WAS REVIEWED BY ABIOMED'S MEDICAL SAFETY OFFICER (MSO) AND IT WAS DETERMINED THAT THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22002116&pc=MNS | 2518422-2025-105484 | 2025-04-22 04:00 | Death | RESPIRONICS, INC. | 2025-05-12 04:00 | MNS | BIPAP AUTOSV ADVANCED | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K090539 |  | 1 | Event Description: THE MANUFACTURER WAS INFORMED THAT A PATIENT USING A BIPAP AUTOSV ADV DEVICE HAD PASSED AWAY SEVERAL YEARS AGO. THE PATIENT¿S WIFE REQUESTED TO STOP RECEIVING CORRESPONDENCE AND INDICATED THAT A REPLACEMENT DEVICE WAS NOT NEEDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT¿S DEATH, AND NO FURTHER DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22002150&pc=NGV | 2124215-2025-30060 | 2025-04-21 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-12 04:00 | NGV | WATCHMAN FLX? PRO | Use of Device Problem; Migration | Stroke/CVA | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED. A 31MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS IMPLANTED. THE CLOSURE DEVICE DID NOT MEET POSITION, ANCHOR, SEAL AND SIZE (PASS) CRITERIA BUT THE PHYSICIAN MADE THE DECISION TO RELEASE THE CLOSURE DEVICE. THE PATIENT WAS DISCHARGED. THE PATIENT HAD A ROUTINE FOLLOW UP TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) AND IMAGING SHOWED THE IMPLANTED CLOSURE DEVICE MIGRATED FROM ITS FINAL POSITION. A COMPUTED TOMOGRAPHY (CT) SCAN AND POSSIBLE CLOSURE DEVICE EXPLANT WERE DISCUSSED. PRIOR TO THE CT SCAN THE PATIENT REQUIRED HOSPITALIZATION TO ADDRESS A SUSPECTED CEREBROVASCULAR EVENT. THE PATIENT LATER PASSED AWAY DUE TO THIS EVENT. Manufacturer Narrative: B2: DATE OF DEATH IS AN APPROXIMATE DATE AS THE ACTUAL DATE OF DEATH IS NOT KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22002225&pc=OZD | 1220648-2025-28297 | 2025-04-20 04:00 | Death | ABIOMED, INC. | 2025-05-12 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia; Hemorrhagic Stroke | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT AN IMPELLA CP WAS IMPLANTED IN A PATIENT POST EXTRACORPOREAL MEMBRANE OXYGENATION CANNULATION. DURING SUPPORT, THE PATIENT EXHIBITED BLEEDING, SUSPECTED TO BE FROM CARDIOPULMONARY RESUSCITATION AND MASSIVE TRANSFUSION PROTOCOL WAS GIVEN. IN ADDITION, DOPPLER PULSES COULD NOT BE FOUND ON THE IMPELLA RIGHT SIDE. THE PATIENT RECEIVED ONE UNIT OF PACKED RED BLOOD CELLS. THE PATIENT THEN SUFFERED A HEMORRHAGIC STROKE. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22002257&pc=BZD | 2518422-2025-105487 | 2025-04-22 04:00 | Death | RESPIRONICS, INC. | 2025-05-12 04:00 | BZD | DREAMSTATION BIPAP PRO | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS INFORMED THAT A PATIENT USING A DREAMSTATION BIPAP PRO DEVICE HAS PASSED AWAY. THE PATIENT¿S HUSBAND CALLED IN REQUESTING A RETURN LABEL TO RETURN THE RECALLED DEVICE. NO FURTHER DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RECEIVED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS INFORMED THAT A PATIENT USING A DREAMSTATION BIPAP PRO DEVICE HAD PASSED AWAY. THE PATIENT¿S HUSBAND CALLED IN REQUESTING A RETURN LABEL TO RETURN THE RECALLED DEVICE. NO FURTHER DETAILS WERE PROVIDED. THE DEVICE HAS BEEN RETURNED FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE THIRD-PARTY SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND NO ERRORS OR DEVICE PROBLEMS. THERE WAS NO EVIDENCE OF FOAM PARTICLES OR DEGRADATION. THE DEVICE WAS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22002273&pc=QZI | 2124215-2025-30121 | 2025-04-11 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-12 04:00 | QZI | FARAWAVE NAV | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Dyspnea; Hemorrhage/Bleeding; Vomiting |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A PULSED FIELD ABLATION (PFA) PROCEDURE WAS COMPLETED SUCCESSFULLY WITH A FARAWAVE CATHETER. THE CATHETER WAS INSERTED THROUGH THE FARADRIVE SHEATH USING A VERSACROSS CONNECT DILATOR AND PLACED INTO THE LEFT ATRIUM WITHOUT ISSUES. USING THE OPAL MAPPING SYSTEM A LEFT ATRIAL ANATOMY AND VOLTAGE MAP WERE CREATED. THE ABLATION WAS PERFORMED IN THE LEFT SUPERIOR PULMONARY VEIN (LSPV), LEFT INFERIOR PULMONARY VEIN (LIPV), RIGHT SUPERIOR PULMONARY VEIN (RSPV) AND RIGHT INFERIOR PULMONARY VEIN (RIPV) WITHOUT ANY ISSUE. AFTER CONFIRMING THE PULMONARY VEIN ISOLATION, THE DEVICE WAS REMOVED FROM THE PATIENT. THE PATIENT WAS KEPT FOR OVERNIGHT OBSERVATION. AROUND THE MIDNIGHT, THE PATIENT CALLED THE NURSE COMPLAINING ABOUT SHORTNESS OF BREATH AND SOON AFTER, BEGAN VOMITING BLOOD. SUBSEQUENTLY, THE PATIENT ASPIRATED BLOOD AND WENT INTO CARDIAC ARREST. RESUSCITATION MANEUVERS WERE PERFORMED BUT WITHOUT SUCCESS. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22002275&pc=QZI | 2124215-2025-30096 | 2025-03-04 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-12 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22002294&pc=OZD | 1220648-2025-28298 | 2025-04-13 04:00 | Death | ABIOMED, INC. | 2025-05-12 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Hemorrhage/Bleeding; Sepsis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE IMPELLA CP WAS PLACED VIA AXILLARY ARTERY AFTER 5.5 ATTEMPT WAS UNSUCCESSFUL. IN ICU, BLEEDING WAS NOTED FROM AXILLARY SURGICAL SITE. THE DRESSING WAS CHANGED AND ONE UNIT OF BLOOD WAS GIVEN. ON DAY SIX OF IMPELLA SUPPORT, THE PATIENT WENT INTO ATRIAL FIBRILLATION WITH RVR AND THE PATIENT SEEMED VERY SEPTIC AND NOT A SURGICAL CANDIDATE. IT WAS FURTHER REPORTED THAT THE FAMILY DECIDED TO WITHDREW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ THE PATIENT'S OUTCOME WAS REVIEWED BY MEDICAL SAFETY OFFICER. ALTHOUGH CARE WAS WITHDRAW, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE (RELATED COMPLICATIONS: ACCESS SITE BLEEDING/BLEEDING REQUIRING TRANSFUSION, ARRHYTHMIC EVENT AND SEPSIS) AS AN ASSOCIATED FACTOR TO THE OVERALL OUTCOME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22002515&pc=OZD | 1220648-2025-28300 | 2025-04-18 04:00 | Death | ABIOMED, INC. | 2025-05-12 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA 5.5 EXHIBITED PERSISTENT BLEEDING. THE PUMP REMAINED ON FOR SUPPORT FOR WEEKS. THE CHOICE WAS THEN MADE TO WITHDRAW CARE AND PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22002536&pc=NGV | 2124215-2025-29883 | 2025-04-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-12 04:00 | NGV | WATCHMAN FLX? PRO | Device Dislodged or Dislocated | Low Blood Pressure/ Hypotension; Syncope/Fainting; Unspecified Tissue Injury | P130013 |  | 1 | Event Description: IT WAS REPORTED DEVICE EMBOLIZATION OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 27MM WATCHMAN FLX PRO LAA CLOSURE DEVICE WAS IMPLANTED ON (B)(6) 2025. THE PATIENT WAS DISCHARGED. THE PATIENT CAME RETURNED TO THE HOSPITAL ON (B)(6) 2025 WITH SYNCOPE SYMPTOMS WHICH INCLUDED WEAK, LIGHTHEADED/DIZZINESS AND HYPOTENSION. ULTRASOUND IMAGING REVEALED THAT THE 27MM WATCHMAN IMPLANT HAD EMBOLIZED INTO THE LEFT VENTRICLE OUTFLOW TRACT (LVOT) AND MITRAL VALVE CAUSING DAMAGE TO THE MITRAL VALVE. CARDIOTHORACIC (CT) SURGERY WAS COMPLETED ON (B)(6) 2025 TO REMOVE THE DEVICE. THE DEVICE WAS FOUND TO BE TANGLED IN THE CHORDS AND WAS CAREFULLY TEASED OUT. ABRASIONS WERE NOTED ON THE CHORDS, BUT NO LARGE CHORDS WERE TORN. THE PATIENT PASSED AWAY THREE DAYS LATER FROM POST-SURGICAL COMPLICATIONS ON (B)(6) 2025. THE OFFICIAL CAUSE OF DEATH WAS NOT KNOWN. Manufacturer Narrative: MEDIA EVALUATED BY BSC MEDICAL SAFETY: FIVE (5) VIDEO LOOPS (3 PROCEDURAL FLUOROSCOPY AND 2 ECHOS) AND SEVEN (7) ECHOS STILL IMAGES WERE SUBMITTED FOR REVIEW. THE LEFT ATRIAL APPENDAGE (LAA) OSTIUM DIAMETER WAS MEASURED AT 1.5CM AND THE DEPLOYED DEVICE ABOUT 2CM. ONE (1) FLUOROSCOPY LOOP WITH CONTRAST SHOWED A SIGNIFICANT INDENTATION OF THE DEVICE AT THE OSTIUM LEVEL WITH AT LEAST 1/3 OF DEVICE OUTSIDE THE LAA. THIS DEVICE INDENTATION DUE TO OSTIUM ANATOMY WAS ALSO SEEN IN AN ECHO IMAGE AND THE ECHO IMAGING ALSO SHOWED A PROXIMAL DEVICE POSITION (IN SOME IMAGES THE DEVICE WAS MORE THAN 1/3 OUTSIDE OF THE LAA). IN SEVERAL ECHO IMAGES, THE ANCHOR ENGAGEMENT SEEMED SUBOPTIMAL. THERE WAS NO EFFICIENT TUG TEST SEEN IN THE FLUOROSCOPY IMAGING. ONE ECHO IMAGE AND ONE VIDEO LOOP SHOWED THE EMBOLIZED DEVICE IN THE LEFT VENTRICLE (LV), ENTANGLED IN THE ANTERIOR LEAFLET OF THE MITRAL VALVE UNDER THE LEFT VENTRICULAR OUTFLOW TRACT (LVOT). IN CONCLUSION, IT IS POSSIBLE THAT THE PROXIMAL POSITION AND SUBOPTIMAL ANCHOR ENGAGEMENT CONTRIBUTED TO THE DEVICE EMBOLIZATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22002652&pc=OZD | 1220648-2025-28301 | 2025-04-20 04:00 | Death | ABIOMED, INC. | 2025-05-12 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Malposition of Device | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT UNDERGOING HIGH-RISK PERCUTANEOUS CORONARY INTERVENTION IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT EXPERIENCED LIMB ISCHEMIA, POSITIONING ISSUE, AND INTERMITTENT LOW PUMP FLOW AND WOULD EVENTUALLY EXPIRE. THE PATIENT WAS CONSULTED FOR CORONARY ARTERY BYPASS GRAFT SURGERY AND WAS PROMPTLY TURNED DOWN. AFTER THE PHYSICIANS COLLABORATED IT WAS DETERMINED THAT PROTECTED (PCI) WITH IMPELLA WAS THE BEST COURSE OF ACTION FOR THE PATIENT. THE IMPELLA WAS SUCCESSFULLY PLACED, AND EXTENSIVE INTERVENTION OF THE LEFT ANTERIOR DESCENDING (4 DRUG-ELUTING STENTS) AND LEFT MAIN (1 DRUG-ELUTING STENT) ARTERIES WAS COMPLETED WITH "SEMI" GOOD RESULT. THERE WERE MANY EPISODES OF NO CORONARY RE-FLOW THAT RESULTED IN VENTRICULAR FIBRILLATION X3 WITH SUCCESSFUL RETURN OF SPONTANEOUS CIRCULATION (NO CARDIOPULMONARY RESUSCITATION WAS NOTED). PATIENT WAS VERY HYPOTENSIVE POST INTERVENTION. FLUIDS AND LEVOPHED WERE STARTED WHILE ON SUPPORT. THE GROIN WAS WARM, DRY AND NOT BLEEDING. THE PATIENT WAS TRANSPORTED BACK TO THE UNIT IN "VERY" CRITICAL BUT STABLE CONDITION. THE FOLLOWING DAY, THE PATIENT WAS TRANSFERRED TO A DIFFERENT FACILITY. THE PATIENT¿S RIGHT LEG (IMPELLA LEG) WAS SEVERELY MOTTLED AND COLD TO THE TOUCH, NO PULSES. THE LEFT LEG WAS COOL TO TOUCH AND HAD A FEMORAL PULSE. UPON ARRIVAL TO THE RECEIVING FACILITY'S UNIT, THERE WAS A FEMSTOP ON THE RIGHT LEG THAT WAS SWIFTLY REMOVED. BEDSIDE ECHOCARDIOGRAM SHOWED THE IMPELLA AT 5.1CM FROM THE AORTIC VALVE. THE FELLOW DID NOT FEEL COMFORTABLE ADJUSTING TO PULL BACK EVEN THOUGH IMPELLA WAS INTERMITTENTLY HAVING SUCTION AGAINST THE SEPTAL WALL. REEVALUATION WAS PLANNED FOR THE NEXT MORNING. THE ECHOGRAM ALSO SHOWED NEW BI-VENTRICULAR FAILURE. THE RIGHT VENTRICLE WAS HARDLY MOVING, AND THE LEFT VENTRICLE WAS NOT MOVING AT ALL. THE PATIENT HAD SOME ASYSTOLIC EPISODES FOR 30-45 SECONDS BUT MAINTAINED A MEAN ARTERIAL PRESSURE WITH THE IMPELLA. DISCUSSION REGARDING ESCALATION TO AN IMPELLA 5.5 WAS PENDING WITH CARDIOTHORACIC SURGERY. HOWEVER, THE PATIENT WOULD EVENTUALLY EXPIRE THIS SAME DAY. CARE WAS WITHDRAWN. Manufacturer Narrative: PRODUCT WAS NOT RETURNED FOR REVIEW. DATA REVIEW: DATA LOGS ONLY AVAILABLE FOR FIRST DAY OF SUPPORT. SUCTION ALARMS WERE REPORTED ON SECOND DAY OF SUPPORT. DATA LOGS SHOWED ABOUT 3 HOURS INTO SUPPORT, FLOW DROPPED SLIGHTLY FROM 3.5 L/MIN TO 3.1 L/MIN THAT TRIGGERED AUTO SUCTION RESPONSE & SUCTION ALARMS. NO POSITIONING ALARM WAS TRIGGERED ON THE LOGS. CLINICAL STATES BEDSIDE ECHO SHOWED IMPELLA AT 5.1CM FROM THE AORTIC VALVE. FELLOW DID NOT FEEL COMFORTABLE ADJUSTING TO PULL BACK EVEN THOUGH IMPELLA WAS INTERMITTENTLY HAVING SUCTION AGAINST THE SEPTAL WALL. ECHO ALSO SHOWED NEW BI-V FAILURE WITH THE RV HARDLY MOVING AND THE LV IS NOT MOVING AT ALL. LOW FLOWS: THE ROOT CAUSE OF LOW FLOW WAS MOST LIKELY PATIENT CONDITION SINCE ECHO SHOWED BI-VENTRICLE FAILURE WITH THE RV HARDLY MOVING AND THE LV NOT MOVING AT ALL. POSITIONING ISSUE: THE ROOT CAUSE OF POSITIONING ISSUE WAS UNABLE TO BE DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED FOR REVIEW. LIMB ISCHEMIA: ISCHEMIA OCCURRED AFTER PATIENT WAS TRANSFERRED TO NEW HOSPITAL ON THE 2ND DAY OF IMPELLA SUPPORT. ECHO FOUND PUMP WAS IN IMPROPER POSITION WITH SUCTION ALARMS & LACK OF MOVING ON BI-VENTRICLES. PUMP THEN WAS EXPLANTED AS THE FAMILY WITHDREW THE CARE. LIMB ISCHEMIA CAN OCCUR DURING IMPELLA SUPPORT. WHEN MAKING A DETERMINATION AS TO THE CAUSE OF THE LIMB ISCHEMIA, THE FOLLOWING CLINICAL INFORMATION IS NECESSARY: PATIENTS PRE-MORBID CONDITION AND PERI-PROCEDURAL CONDITION; ANY CONCOMITANT MEDICATION; VASCULAR SIZE OR VARIATIONS THEREOF; DETAILED DESCRIPTION OF THE SYMPTOMS EXPERIENCED BY THE PATIENT; PHYSICAL EXAMINATION FINDINGS, INCLUDING PULSE ASSESSMENT AND VISUAL EXAMINATION OF THE AFFECTED LIMB; DIAGNOSTIC IMAGING RESULTS, SUCH AS DOPPLER ULTRASOUND, ANGIOGRAPHY, OR MAGNETIC RESONANCE ANGIOGRAPHY; LABORATORY TEST RESULTS, INCLUDING BLOOD TESTS FOR MARKERS OF INFLAMMATION OR CLOTTING DISORDERS; ANY RELEVANT INFORMATION ABOUT RISK FACTORS FOR PERIPHERAL ARTERIAL DISEASE, SUCH AS SMOKING, DIABETES, OR HYPERTENSION; PATIENT'S RESPONSE TO PREVIOUS TREATMENTS OR INTERVENTIONS FOR VASCULAR CONDITIONS; WITH A RETURNED IMPELLA, THE MAX OD COULD BE ASSESSED TO ENSURE IT IS WITHIN SPECIFICATION. THE PRODUCT COULD ALSO BE EVALUATED FOR ANY BURRS OR SHARP EDGES. ADDITIONALLY, THE CLINICAL INFORMATION ABOVE WOULD NEED TO BE CONSIDERED IN THE CONTEXT OF THE RETURNED PRODUCT. TO DETERMINE THE TRUE ROOT CAUSE OF LIMB ISCHEMIA, THE CLINICAL INFORMATION MENTIONED ABOVE WOULD NEED TO BE ASSESSED IN CONJUNCTION WITH EVIDENCE FROM THE RETURNED DEVICE. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. Manufacturer Narrative: CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. HOWEVER, GIVEN THE SEVERITY OF THE ADVERSE EVENT, SUBOPTIMAL POSITIONING AND INTERMITTENT LOW PUMP FLOW POTENTIALLY COMPROMISING HEMODYNAMIC SUPPORT, THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22002795&pc=DSQ | 2916596-2025-02815 | 2025-04-17 04:00 | Death | THORATEC CORPORATION | 2025-05-12 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT HAD A CARDIAC ARREST THAT LED TO DEATH. Manufacturer Narrative: SECTION E1: REPORT CONTACT INFORMATION WAS NOT AVAILABLE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE PATIENT'S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED, AND THE DEVICE OPERATED AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED, AND THE PUMP WAS NOT EXPLANTED FOR EVALUATION. NO FURTHER INFORMATION WAS PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22003001&pc=MOZ | MW5170192 |  | Death | BIO-TECHNOLOGY GENERAL (ISRAEL) LTD. | 2025-05-12 04:00 | MOZ | EUFLEXXA | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: EVENT: (B)(6) AT HCP REPORTED THE PT HAS PASSED AWAY. FREQ: MD TO INJECT 20 MG INTRA-ARTICULAR INTO THE LEFT KNEE EVERY 7 DAYS, FOR 3 WEEKS. REF REPORTS: MW5170193 AND MW5170194.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22003020&pc=MOZ | MW5170193 |  | Death | BIO-TECHNOLOGY GENERAL (ISRAEL) LTD. | 2025-05-12 04:00 | MOZ | EUFLEXXA | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: EVENT: (B)(6) AT HCP REPORTED THE PT HAS PASSED AWAY. FREQ: MD TO INJECT 20 MG INTRA-ARTICULAR INTO THE LEFT KNEE EVERY 7 DAYS, FOR 3 WEEKS. REF REPORTS: MW5170192 AND MW5170194.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22003021&pc=MOZ | MW5170194 |  | Death | BIO-TECHNOLOGY GENERAL (ISRAEL) LTD. | 2025-05-12 04:00 | MOZ | EUFLEXXA | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: EVENT: (B)(6) AT HCP REPORTED THE PT HAS PASSED AWAY. FREQ: MD TO INJECT 20 MG INTRA-ARTICULAR INTO THE LEFT KNEE EVERY 7 DAYS, FOR 3 WEEKS. REF REPORTS: MW5170192 AND MW5170193.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22003092&pc=BZD | 2518422-2025-044736 | 2022-05-06 04:00 | Death | RESPIRONICS, INC. | 2025-05-12 04:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22003111&pc=MUE | MW5170195 | 2024-10-13 04:00 | Death | UNK | 2025-05-12 04:00 | MUE | MAMMOGRAPHY-DIGITAL | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: PATIENT DECEASED PER FAMILY MEMBER; DEATH DATE WAS UPDATED IN ELECTRONIC MEDICAL RECORD. NO INFORMATION REGARDING CAUSE OF DEATH IS AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22003181&pc=MIH | MW5170196 | 2025-01-07 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-05-12 04:00 | MIH | TEVAR GRAFT STENT | Patient Device Interaction Problem | Perforation of Vessels |  |  | 1 | Event Description: WAS UNDERGOING PLACEMENT OF AN ENDOVASCULAR AORTIC GRAFT (TEVAR) FOR A THORACIC ANEURYSM. THE GRAFT TRANSECTED THE ILIAC ARTERY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22003268&pc=NPT | 2025587-2025-03599 | 2025-04-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-12 04:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Positioning; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Insufficient Information; Activation Failure; Patient Device Interaction Problem | Ventricular Fibrillation; Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH SEVERE CALCIFICATION IN BOTH THE SIGMOID AND TRICUSPID AREAS, AFTER FULL RELEASE OF THE FIRST EVOLUT, THE PLACEMENT DEPTH WAS NOTED TO BE SHALLOW. THE LOWER END WAS ANCHORED TO THE ANNULUS; HOWEVER, POOR EXPANSION WAS OBSERVED ON THE NON-CORONARY CUSP (NCC) SIDE DUE TO CALCIFICATION. DURING WITHDRAWAL OF THE DELIVERY CATHETER SYSTEM (DCS), NOSE CONE CENTERING WAS INITIALLY CONFIRMED. AS THE DEVICE WAS WITHDRAWN, HOWEVER, THE NOSE CONE INTERACTED WITH THE VALVE FRAME CAUSING THE VALVE TO DISLODGE. THE DISLODGE WAS THOUGHT TO HAVE BEEN INFLUENCED BY THE EXPANSION ISSUE. THERE WAS ALSO A CONSIDERATION THAT THE VALVE SIZE MIGHT NOT HAVE BEEN APPROPRIATE. THE INTERVENTION PLAN INCLUDED PLACING A SECOND VALVE. THE SECOND VALVE WAS POSITIONED SLIGHTLY DEEPER THAN THE FIRST HAD BEEN, WHEN INITIALLY DEPLOYED. THIS STILL RESULTED IN A SHALLOW PLACEMENT UPON FULL RELEASE. OF NOTE, EXTREMELY DEEP PLACEMENT WAS AVOIDED DUE TO OVERLAP RISKS WITH THE FIRST DISLODGED VALVE AND CONCERN FOR ATRIO-VENTRICULAR BLOCK. ADJUSTMENTS HAD BEEN MADE TO REACH APPROXIMATELY 5 MM ON BOTH THE NCC AND LEFT CORONARY CUSP (LCC) AT 80% DEPLOYMENT. FOLLOWING PLACEMENT, THE DCS NOSE CONE WAS CAREFULLY CENTERED AND AN ATTEMPT TO WITHDRAW THE DCS WAS INITIATED. THE VALVE WAS GRADUALLY RAISED AS THE DCS WAS WITHDRAWN AND IT BECAME EVIDENT THAT THE VALVE'S POSITION WAS BECOMING SHALLOW. DESPITE ATTEMPTS TO CHANGE THE WITHDRAWAL STRATEGY, SUCH AS SWITCHING OUT THE LEFT VENTRICLE WIRE, THE VALVE WAS STILL DISLODGED DURING REMOVAL. IT WAS CONSIDERED THAT SNARING THE FIRST VALVE DOWN TO THE DESCENDING AORTA MIGHT CREATE SPACE FOR PLACING A THIRD EVOLUT, BUT THE DIFFICULTY AND RISK OF MOVING IT DOWN LED TO A CONVERSION TO NON-MEDTRONIC DEVICE. DURING THE PLACEMENT OF THE THIRD VALVE (SAPIEN 23MM), THERE WAS RESISTANCE WHEN PASSING THE THIRD VALVE THROUGH THE FIRST AND SECOND DISLODGED VALVES. DESPITE THIS, THE THIRD VALVE SUCCESSFULLY CROSSED OVER THE AORTIC VALVE AND WAS PLACED. DURING PLACEMENT, HOWEVER, CORONARY ARTERY DISSECTION WAS NOTED AND VENTRICULAR FIBRILLATION (VF) OCCURRED. CARDIOPULMONARY RESUSCITATION AND DIRECT CURRENT SHOCK WERE ADMINISTERED WHILE EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS INITIATED TO MAINTAIN HEMODYNAMICS. EVEN AFTER TREATING THE CORONARY ARTERY DISSECTION THAT CAUSED THE VF, THE HEART WAS "NOT FUNCTIONING". CARDIAC FUNCTION RESUMED AFTER AN AORTOGRAM WAS TAKEN, POSSIBLY DUE TO CHANGES IN VASOPRESSOR MEDICATION. JUST BEFORE EXITING THE PROCEDURE ROOM, PERICARDIAL EFFUSION WAS OBSERVED. UPON INVESTIGATION, IT WAS FOUND THAT AT SOME POINT, THE LOWER END OF THE FIRST EVOLUT HAD PIERCED NEAR THE SINOTUBULAR JUNCTION, CAUSING AN ASCENDING AORTIC DISSECTION. A DRAINAGE KIT WAS USED TO MANAGE THE DISSECTION. THE PATIENT EXITED THE PROCEDURE WITH ECMO STILL IN PLACE AND WITH THE ASCENDING AORTIC DISSECTION HAVING BEEN TEMPORARILY PRESERVED. AT SOME POINT WITHIN THE THREE DAYS FOLLOWING PROCEDURE, THE PATIENT DIED. THE DATE OF DEATH IS UNKNOWN. THE CAUSE OF DEATH IS UNDER INVESTIGATION.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH SEVERE CALCIFICATION IN BOTH THE SIGMOID AND TRICUSPID AREAS, AFTER FULL RELEASE OF THE FIRST EVOLUT, THE PLACEMENT DEPTH WAS NOTED TO BE SHALLOW. THE LOWER END WAS ANCHORED TO THE ANNULUS; HOWEVER, POOR EXPANSION WAS OBSERVED ON THE NON-CORONARY CUSP (NCC) SIDE DUE TO CALCIFICATION. DURING WITHDRAWAL OF THE DELIVERY CATHETER SYSTEM (DCS), NOSE CONE CENTERING WAS INITIALLY CONFIRMED. AS THE DEVICE WAS WITHDRAWN, HOWEVER, THE NOSE CONE INTERACTED WITH THE VALVE FRAME CAUSING THE VALVE TO DISLODGE. THE DISLODGE WAS THOUGHT TO HAVE BEEN INFLUENCED BY THE EXPANSION ISSUE. THERE WAS ALSO A CONSIDERATION THAT THE VALVE SIZE MIGHT NOT HAVE BEEN APPROPRIATE. THE INTERVENTION PLAN INCLUDED PLACING A SECOND VALVE. THE SECOND VALVE WAS POSITIONED SLIGHTLY DEEPER THAN THE FIRST HAD BEEN, WHEN INITIALLY DEPLOYED. THIS STILL RESULTED IN A SHALLOW PLACEMENT UPON FULL RELEASE. OF NOTE, EXTREMELY DEEP PLACEMENT WAS AVOIDED DUE TO OVERLAP RISKS WITH THE FIRST DISLODGED VALVE AND CONCERN FOR ATRIO-VENTRICULAR BLOCK. ADJUSTMENTS HAD BEEN MADE TO REACH APPROXIMATELY 5 MM ON BOTH THE NCC AND LEFT CORONARY CUSP (LCC) AT 80% DEPLOYMENT. FOLLOWING PLACEMENT, THE DCS NOSE CONE WAS CAREFULLY CENTERED AND AN ATTEMPT TO WITHDRAW THE DCS WAS INITIATED. THE VALVE WAS GRADUALLY RAISED AS THE DCS WAS WITHDRAWN AND IT BECAME EVIDENT THAT THE VALVE'S POSITION WAS BECOMING SHALLOW. DESPITE ATTEMPTS TO CHANGE THE WITHDRAWAL STRATEGY, SUCH AS SWITCHING OUT THE LEFT VENTRICLE WIRE, THE VALVE WAS STILL DISLODGED DURING REMOVAL. IT WAS CONSIDERED THAT SNARING THE FIRST VALVE DOWN TO THE DESCENDING AORTA MIGHT CREATE SPACE FOR PLACING A THIRD EVOLUT, BUT THE DIFFICULTY AND RISK OF MOVING IT DOWN LED TO A CONVERSION TO NON-MEDTRONIC DEVICE. DURING THE PLACEMENT OF THE THIRD VALVE, THERE WAS RESISTANCE WHEN PASSING THE THIRD VALVE THROUGH THE FIRST AND SECOND DISLODGED VALVES. DESPITE THIS, THE THIRD VALVE SUCCESSFULLY CROSSED OVER THE AORTIC VALVE AND WAS PLACED. DURING PLACEMENT, HOWEVER, CORONARY ARTERY DISSECTION WAS NOTED AND VENTRICULAR FIBRILLATION (VF) OCCURRED. CARDIOPULMONARY RESUSCITATION AND DIRECT CURRENT SHOCK WERE ADMINISTERED WHILE EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS INITIATED TO MAINTAIN HEMODYNAMICS. EVEN AFTER TREATING THE CORONARY ARTERY DISSECTION THAT CAUSED THE VF, THE HEART WAS "NOT FUNCTIONING". CARDIAC FUNCTION RESUMED AFTER AN AORTOGRAM WAS TAKEN, POSSIBLY DUE TO CHANGES IN VASOPRESSOR MEDICATION. JUST BEFORE EXITING THE PROCEDURE ROOM, PERICARDIAL EFFUSION WAS OBSERVED. UPON INVESTIGATION, IT WAS FOUND THAT AT SOME POINT, THE LOWER END OF THE FIRST EVOLUT HAD PIERCED NEAR THE SINOTUBULAR JUNCTION, CAUSING AN ASCENDING AORTIC DISSECTION. A DRAINAGE KIT WAS USED TO MANAGE THE DISSECTION. THE PATIENT EXITED THE PROCEDURE WITH ECMO STILL IN PLACE AND WITH THE ASCENDING AORTIC DISSECTION HAVING BEEN TEMPORARILY PRESERVED. AT SOME POINT WITHIN THE THREE DAYS FOLLOWING PROCEDURE, THE PATIENT DIED. THE DATE OF DEATH IS UNKNOWN. THE CAUSE OF DEATH IS UNDER INVESTIGATION. ADDITIONAL INFORMATION WAS RECEIVED NOTING THAT THE SUSPECTED CAUSE OF THE DISSECTION WAS THE FIRST VALVE. IT IS THOUGHT THAT THIS OCCURRED WHEN THE SECOND OR THIRD VALVE WAS PLACED. IT WAS NECESSARY TO PASS THESE VALVES THROUGH THE FIRST VALVE, WHICH REMAINED IN THE BODY. ALTHOUGH THE FIRST VALVE HAD BEEN SNARED, IT COULD NOT BE COMPLETELY FIXED. THE FIRST VALVE WOULD "RISE AND FALL". THE INSERTION OF THE NEW VALVES MAY HAVE CAUSED THE FIRST VALVE TO PIERCE INTO THE SINOTUBULAR JUNCTION. PER THE PHYSICIAN, THE DCS LIKELY DID NOT CONTRIBUTE TO THE DISSECTION AND IT IS HIGHLY LIKELY THAT THE ASCENDING AORTIC DISSECTION CONTRIBUTED TO THE CAUSE OF DEATH. Manufacturer Narrative: CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: E VOLUTFX-26, SERIAL/LOT #: (B)(6), UBD: 11-SEP-2026, UDI#: (B)(4). SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2. ADDED DATE OF DEATH. DATE IS APPROXIMATE: MONTH AND YEAR VALID. B5. ADDED SECOND PARAGRAPH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22003288&pc=NPT | 2025587-2025-03600 | 2025-04-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-12 04:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Positioning; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Insufficient Information; Activation Failure; Patient Device Interaction Problem | Ventricular Fibrillation; Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH SEVERE CALCIFICATION IN BOTH THE SIGMOID AND TRICUSPID AREAS, AFTER FULL RELEASE OF THE FIRST EVOLUT, THE PLACEMENT DEPTH WAS NOTED TO BE SHALLOW. THE LOWER END WAS ANCHORED TO THE ANNULUS; HOWEVER, POOR EXPANSION WAS OBSERVED ON THE NON-CORONARY CUSP (NCC) SIDE DUE TO CALCIFICATION. DURING WITHDRAWAL OF THE DELIVERY CATHETER SYSTEM (DCS), NOSE CONE CENTERING WAS INITIALLY CONFIRMED. AS THE DEVICE WAS WITHDRAWN, HOWEVER, THE NOSE CONE INTERACTED WITH THE VALVE FRAME CAUSING THE VALVE TO DISLODGE. THE DISLODGE WAS THOUGHT TO HAVE BEEN INFLUENCED BY THE EXPANSION ISSUE. THERE WAS ALSO A CONSIDERATION THAT THE VALVE SIZE MIGHT NOT HAVE BEEN APPROPRIATE. THE INTERVENTION PLAN INCLUDED PLACING A SECOND VALVE. THE SECOND VALVE WAS POSITIONED SLIGHTLY DEEPER THAN THE FIRST HAD BEEN, WHEN INITIALLY DEPLOYED. THIS STILL RESULTED IN A SHALLOW PLACEMENT UPON FULL RELEASE. OF NOTE, EXTREMELY DEEP PLACEMENT WAS AVOIDED DUE TO OVERLAP RISKS WITH THE FIRST DISLODGED VALVE AND CONCERN FOR ATRIO-VENTRICULAR BLOCK. ADJUSTMENTS HAD BEEN MADE TO REACH APPROXIMATELY 5 MM ON BOTH THE NCC AND LEFT CORONARY CUSP (LCC) AT 80% DEPLOYMENT. FOLLOWING PLACEMENT, THE DCS NOSE CONE WAS CAREFULLY CENTERED AND AN ATTEMPT TO WITHDRAW THE DCS WAS INITIATED. THE VALVE WAS GRADUALLY RAISED AS THE DCS WAS WITHDRAWN AND IT BECAME EVIDENT THAT THE VALVE'S POSITION WAS BECOMING SHALLOW. DESPITE ATTEMPTS TO CHANGE THE WITHDRAWAL STRATEGY, SUCH AS SWITCHING OUT THE LEFT VENTRICLE WIRE, THE VALVE WAS STILL DISLODGED DURING REMOVAL. IT WAS CONSIDERED THAT SNARING THE FIRST VALVE DOWN TO THE DESCENDING AORTA MIGHT CREATE SPACE FOR PLACING A THIRD EVOLUT, BUT THE DIFFICULTY AND RISK OF MOVING IT DOWN LED TO A CONVERSION TO NON-MEDTRONIC DEVICE. DURING THE PLACEMENT OF THE THIRD VALVE (SAPIEN 23MM), THERE WAS RESISTANCE WHEN PASSING THE THIRD VALVE THROUGH THE FIRST AND SECOND DISLODGED VALVES. DESPITE THIS, THE THIRD VALVE SUCCESSFULLY CROSSED OVER THE AORTIC VALVE AND WAS PLACED. DURING PLACEMENT, HOWEVER, CORONARY ARTERY DISSECTION WAS NOTED AND VENTRICULAR FIBRILLATION (VF) OCCURRED. CARDIOPULMONARY RESUSCITATION AND DIRECT CURRENT SHOCK WERE ADMINISTERED WHILE EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS INITIATED TO MAINTAIN HEMODYNAMICS. EVEN AFTER TREATING THE CORONARY ARTERY DISSECTION THAT CAUSED THE VF, THE HEART WAS "NOT FUNCTIONING". CARDIAC FUNCTION RESUMED AFTER AN AORTOGRAM WAS TAKEN, POSSIBLY DUE TO CHANGES IN VASOPRESSOR MEDICATION. JUST BEFORE EXITING THE PROCEDURE ROOM, PERICARDIAL EFFUSION WAS OBSERVED. UPON INVESTIGATION, IT WAS FOUND THAT AT SOME POINT, THE LOWER END OF THE FIRST EVOLUT HAD PIERCED NEAR THE SINOTUBULAR JUNCTION, CAUSING AN AS CENDING AORTIC DISSECTION. A DRAINAGE KIT WAS USED TO MANAGE THE DISSECTION. THE PATIENT EXITED THE PROCEDURE WITH ECMO STILL IN PLACE AND WITH THE ASCENDING AORTIC DISSECTION HAVING BEEN TEMPORARILY PRESERVED. AT SOME POINT WITHIN THE THREE DAYS FOLLOWING PROCEDURE, THE PATIENT DIED. THE DATE OF DEATH IS UNKNOWN. THE CAUSE OF DEATH IS UNDER INVESTIGATION.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE IN A PATIENT WITH SEVERE CALCIFICATION IN BOTH THE SIGMOID AND TRICUSPID AREAS, AFTER FULL RELEASE OF THE FIRST EVOLUT, THE PLACEMENT DEPTH WAS NOTED TO BE SHALLOW. THE LOWER END WAS ANCHORED TO THE ANNULUS; HOWEVER, POOR EXPANSION WAS OBSERVED ON THE NON-CORONARY CUSP (NCC) SIDE DUE TO CALCIFICATION. DURING WITHDRAWAL OF THE DELIVERY CATHETER SYSTEM (DCS), NOSE CONE CENTERING WAS INITIALLY CONFIRMED. AS THE DEVICE WAS WITHDRAWN, HOWEVER, THE NOSE CONE INTERACTED WITH THE VALVE FRAME CAUSING THE VALVE TO DISLODGE. THE DISLODGE WAS THOUGHT TO HAVE BEEN INFLUENCED BY THE EXPANSION ISSUE. THERE WAS ALSO A CONSIDERATION THAT THE VALVE SIZE MIGHT NOT HAVE BEEN APPROPRIATE. THE INTERVENTION PLAN INCLUDED PLACING A SECOND VALVE. THE SECOND VALVE WAS POSITIONED SLIGHTLY DEEPER THAN THE FIRST HAD BEEN, WHEN INITIALLY DEPLOYED. THIS STILL RESULTED IN A SHALLOW PLACEMENT UPON FULL RELEASE. OF NOTE, EXTREMELY DEEP PLACEMENT WAS AVOIDED DUE TO OVERLAP RISKS WITH THE FIRST DISLODGED VALVE AND CONCERN FOR ATRIO-VENTRICULAR BLOCK. ADJUSTMENTS HAD BEEN MADE TO REACH APPROXIMATELY 5 MM ON BOTH THE NCC AND LEFT CORONARY CUSP (LCC) AT 80% DEPLOYMENT. FOLLOWING PLACEMENT, THE DCS NOSE CONE WAS CAREFULLY CENTERED AND AN ATTEMPT TO WITHDRAW THE DCS WAS INITIATED. THE VALVE WAS GRADUALLY RAISED AS THE DCS WAS WITHDRAWN AND IT BECAME EVIDENT THAT THE VALVE'S POSITION WAS BECOMING SHALLOW. DESPITE ATTEMPTS TO CHANGE THE WITHDRAWAL STRATEGY, SUCH AS SWITCHING OUT THE LEFT VENTRICLE WIRE, THE VALVE WAS STILL DISLODGED DURING REMOVAL. IT WAS CONSIDERED THAT SNARING THE FIRST VALVE DOWN TO THE DESCENDING AORTA MIGHT CREATE SPACE FOR PLACING A THIRD EVOLUT, BUT THE DIFFICULTY AND RISK OF MOVING IT DOWN LED TO A CONVERSION TO NON-MEDTRONIC DEVICE. DURING THE PLACEMENT OF THE THIRD VALVE, THERE WAS RESISTANCE WHEN PASSING THE THIRD VALVE THROUGH THE FIRST AND SECOND DISLODGED VALVES. DESPITE THIS, THE THIRD VALVE SUCCESSFULLY CROSSED OVER THE AORTIC VALVE AND WAS PLACED. DURING PLACEMENT, HOWEVER, CORONARY ARTERY DISSECTION WAS NOTED AND VENTRICULAR FIBRILLATION (VF) OCCURRED. CARDIOPULMONARY RESUSCITATION AND DIRECT CURRENT SHOCK WERE ADMINISTERED WHILE EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS INITIATED TO MAINTAIN HEMODYNAMICS. EVEN AFTER TREATING THE CORONARY ARTERY DISSECTION THAT CAUSED THE VF, THE HEART WAS "NOT FUNCTIONING". CARDIAC FUNCTION RESUMED AFTER AN AORTOGRAM WAS TAKEN, POSSIBLY DUE TO CHANGES IN VASOPRESSOR MEDICATION. JUST BEFORE EXITING THE PROCEDURE ROOM, PERICARDIAL EFFUSION WAS OBSERVED. UPON INVESTIGATION, IT WAS FOUND THAT AT SOME POINT, THE LOWER END OF THE FIRST EVOLUT HAD PIERCED NEAR THE SINOTUBULAR JUNCTION, CAUSING AN ASCENDING AORTIC DISSECTION. A DRAINAGE KIT WAS USED TO MANAGE THE DISSECTION. THE PATIENT EXITED THE PROCEDURE WITH ECMO STILL IN PLACE AND WITH THE ASCENDING AORTIC DISSECTION HAVING BEEN TEMPORARILY PRESERVED. AT SOME POINT WITHIN THE THREE DAYS FOLLOWING PROCEDURE, THE PATIENT DIED. THE DATE OF DEATH IS UNKNOWN. THE CAUSE OF DEATH IS UNDER INVESTIGATION. ADDITIONAL INFORMATION WAS RECEIVED NOTING THAT THE SUSPECTED CAUSE OF THE DISSECTION WAS THE FIRST VALVE. IT IS THOUGHT THAT THIS OCCURRED WHEN THE SECOND OR THIRD VALVE WAS PLACED. IT WAS NECESSARY TO PASS THESE VALVES THROUGH THE FIRST VALVE, WHICH REMAINED IN THE BODY. ALTHOUGH THE FIRST VALVE HAD BEEN SNARED, IT COULD NOT BE COMPLETELY FIXED. THE FIRST VALVE WOULD "RISE AND FALL". THE INSERTION OF THE NEW VALVES MAY HAVE CAUSED THE FIRST VALVE TO PIERCE INTO THE SINOTUBULAR JUNCTION. PER THE PHYSICIAN, THE DCS LIKELY DID NOT CONTRIBUTE TO THE DISSECTION AND IT IS HIGHLY LIKELY THAT THE ASCENDING AORTIC DISSECTION CONTRIBUTED TO THE CAUSE OF DEATH. Manufacturer Narrative: UPDATED DATA: B2. ADDED DATE OF DEATH. DATE IS APPROXIMATE: MONTH AND YEAR VALID. B5. ADDED SECOND PARAGRAPH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: D-EVOLUTFX-2329 (LOT: 0012585684); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID: L-EVOLUTFX-2329 (LOT: 0012289396); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID: L-EVOLUTFX-2329 (LOT: 0012508016); PRODUCT TYPE: 0195-HEART VALVES. SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME: EVOLUT FX VALVE; PRODUCT ID: EVOLUTFX-26 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: ON (B)(6) 2025; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22003323&pc=DSQ | 2916596-2025-03120 | 2025-04-30 04:00 | Death | THORATEC CORPORATION | 2025-05-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RESPIRATORY FAILURE. THE OUTCOME WAS NOT CONSIDERED LEFT VENTRICULAR ASSIST DEVICE (LVAD) OR LVAD THERAPY RELATED. THE PUMP FUNCTIONED AS INTENDED. Manufacturer Narrative: SECTION A: PATIENT AGE/DATE OF BIRTH, SEX, AND WEIGHT COULD NOT BE PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE AND DEATH, WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22003705&pc=PYX | 1220648-2025-28311 | 2025-04-21 04:00 | Death | ABIOMED, INC. | 2025-05-12 04:00 | PYX | IMPELLA RP FLEX | Restricted Flow rate | Insufficient Information | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AN IMPELLA RP FLEX WAS PLACED FOR ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK AND WITH GOOD POSITION ON FLUOROSCOPY BUT COULD NOT GET THE IMPELLA FLOW ABOVE ONE LITER PER MINUTE. AS THE IMPELLA RP FLEX WAS PULLED BACK SLIGHTLY, FLOW IMPROVED TO LESS THAN TWO LITERS PER MINUTE ON P5. THE PATIENT CONTINUED TO DETERIORATE, AND THE FAMILY MADE THE PATIENT DO NOT RESUSCITATE AND DECIDED TO WITHDRAW CARE. THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW HAS BEEN COMPLETED. LOW PUMP FLOW: CLINICAL STATES THAT THE PUMP WAS GETTING LOW FLOWS UPON ACTIVATION. PUMP WAS REPOSITIONED AND CONSOLE WAS CHANGED AND ADEQUATE FLOWS WERE OBSERVED AT P5. PUMP WAS RETURNED FOR INVESTIGATION. THE PUMP HAD NO BIOMATERIAL, KINK OR ANY PHYSICAL OR MECHANICAL DEFECTS. LOW PUMP FLOW WAS NOT REPRODUCED. THE PUMP WORKED AS PER THE SPECIFICATION. DATA LOGS WERE INVESTIGATED. LOW FLOWS WERE CONFIRMED. THE PUMP WAS ATTEMPTING TO RUN AT HIGHER P-LEVELS BUT WAS GETTING FLOW OF BETWEEN 1L/MIN TO 2L/MIN. THE EXPECTED FLOW AT P7 IS 2.7L/MIN TO 3.4L/MIN. NO SUCTION ALARMS. CONSOLE WAS REPLACED AND SHOWED NO LOW FLOW. THE PUMP HAD FLOWS WITHIN THE RANGE ON P5. NO SUCTION ALARMS PRESENT. PLACEMENT SIGNAL WAS SEEN SLIGHTLY DRIFTING UP. THEREFORE, BASED ON THE PRODUCT INVESTIGATION, DATA LOG REVIEW AND CLINICAL INFORMATION, THE ROOT CAUSE OF THE LOW PUMP FLOW ISSUE WAS MOST LIKELY IMPROPER POSITIONING OF THE DEVICE SINCE AFTER CHANGING THE CONSOLE AND REPOSITIONING THE PUMP, FLOWS WERE WITHIN THE RANGE. D9 DEVICE RETURNED TO MANUFACTURER DATE ADDED G4 PMA/510(K)NUMBER WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER: 1220648-2025-28311. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22003945&pc=OZD | 1220648-2025-28313 | 2025-04-21 04:00 | Death | ABIOMED, INC. | 2025-05-12 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WAS ON IMPELLA 5.5 SUPPORT HAD A RUN OF VENTRICULAR TACHYCARDIA (VT) AND WAS SHOCKED TO GET BACK INTO NORMAL SINUS RHYTHM. THE FOLLOWING DAY, THE PATIENT WAS REPORTED TO CONTINUE TO GO INTO VT WHEN MOVED. LIDOCAINE WAS STARTED FOR VT. IT WAS FURTHER REPORTED THAT A DECISION WAS MADE TO WITHDREW CARE AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. B2 REQUIRED INTERVENTION SELECTION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28313.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22004109&pc=MCX | 3004742232-2025-00136 | 2025-04-28 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-05-12 04:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hypoxia; Perforation of Vessels; Obstruction/Occlusion; Pericardial Effusion | P130005 |  | 1 | Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT AN ECCENTRIC LEFT CORONARY ARTERY (LCX) #11 LESION WITH 90% STENOSIS, HEAVY CALCIFICATION, AND MODERATE TORTUOSITY. THREE LOW SPEED TREATMENTS WERE PERFORMED. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED AFTER EACH TREATMENT TO OBSERVE THE LESION. AFTER THE THIRD TREATMENT, A PERFORATION WAS OBSERVED AT THE BRANCH SECTION. THE PHYSICIAN ATTEMPTED TO USE A GRAFMASTER, HOWEVER, A GUIDE WIRE AT THE LCX HAD BEEN REMOVED BY THAT TIME. THE GRAFTMASTER WAS IMPLANTED TOWARDS THE LEFT ANTERIOR DESCENDING ARTERY (LAD) AND BLOOD FLOW WAS STOPPED TOWARDS THE LCX. PERICARDIAL DRAINAGE WAS PERFORMED AND AN INTRA-AORTIC BALLOON PUMPING WAS PERFORMED. A VA-ECMO WAS IMPLANTED AND THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU). THE DAY AFTER THE PROCEDURE, THE PATIENT'S HEART RATE AND OXYGEN SATURATION (SPO2) WAS NOTED TO BE DECREASING, THEREFORE, AN EMERGENCY PROCEDURE WAS PERFORMED. ULTRASOUND CONFIRMED PERICARDIAL EFFUSION, THEREFORE PERICARDIAL DRAINAGE WAS PERFORMED. CAG WAS PERFORMED AND CONFIRMED A PERFORATION ON COLLATERAL IN THE LCX. A COIL EMBOLIZATION WAS PERFORMED, BUT THE PATIENT EXPIRED.Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT AN ECCENTRIC LEFT CORONARY ARTERY (LCX) #11 LESION WITH 90% STENOSIS, HEAVY CALCIFICATION, AND MODERATE TORTUOSITY. THREE LOW SPEED TREATMENTS WERE PERFORMED. INTRAVASCULAR ULTRASOUND (IVUS) WAS USED AFTER EACH TREATMENT TO OBSERVE THE LESION. AFTER THE THIRD TREATMENT, A PERFORATION WAS OBSERVED AT THE BRANCH SECTION. THE PHYSICIAN ATTEMPTED TO USE A GRAFMASTER, HOWEVER, A GUIDE WIRE AT THE LCX HAD BEEN REMOVED BY THAT TIME. THE GRAFTMASTER WAS IMPLANTED TOWARDS THE LEFT ANTERIOR DESCENDING ARTERY (LAD) AND BLOOD FLOW WAS STOPPED TOWARDS THE LCX. PERICARDIAL DRAINAGE WAS PERFORMED AND AN INTRA-AORTIC BALLOON PUMPING WAS PERFORMED. A VA-ECMO WAS IMPLANTED AND THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU). THE DAY AFTER THE PROCEDURE, THE PATIENT'S HEART RATE AND OXYGEN SATURATION (SPO2) WAS NOTED TO BE DECREASING, THEREFORE, AN EMERGENCY PROCEDURE WAS PERFORMED. ULTRASOUND CONFIRMED PERICARDIAL EFFUSION, THEREFORE PERICARDIAL DRAINAGE WAS PERFORMED. CAG WAS PERFORMED AND CONFIRMED A PERFORATION ON COLLATERAL IN THE LCX. A COIL EMBOLIZATION WAS PERFORMED, BUT THE PATIENT EXPIRED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED PATIENT ARRHYTHMIA, DEATH, HOSPITALIZATION, HYPOXIA, OBSTRUCTION/OCCLUSION, PERFORATION OF VESSELS, PERICARDIAL EFFUSION, SURGICAL INTERVENTION, AND UNEXPECTED MEDICAL INTERVENTION APPEARS TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE IT IS LIKELY THAT THE REPORTED PATIENT ARRHYTHMIA, DEATH, HOSPITALIZATION, HYPOXIA, OBSTRUCTION/OCCLUSION, PERFORATION OF VESSELS, PERICARDIAL EFFUSION, SURGICAL INTERVENTION, AND UNEXPECTED MEDICAL INTERVENTION ARE A RESULT OF THE PATIENT¿S DISEASE SEVERITY COMBINED WITH USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. H6: CODES 4118, 3233, AND 11 NO LONGER APPLY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22004242&pc=DSQ | 2916596-2025-02822 | 2025-03-28 04:00 | Death | THORATEC CORPORATION | 2025-05-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS UNEXPECTEDLY FOUND DEAD ON (B)(6) 2025 BY A FAMILY MEMBER. WHEN FOUND, NOTHING WAS ATTACHED TO THE CORD THAT COMES OUT OF THE ABDOMEN. HOWEVER, NO DRIVELINE DISCONNECT ALARM WAS SEEN IN THE HISTORY. THE LOG FILE SHOWED NO EXTERNAL POWER EVENTS ON (B)(6) 2025 AT 23:40:06. A BACKUP BATTERY FAULT WAS ALSO NOTED THAT MAY HAVE BEEN DUE TO NO EXTERNAL POWER TO CHANGE THE BATTERY. THE PUMP WAS RUNNING AT THIS TIME ON THE BACKUP BATTERY. ON (B)(6) 2025 AT 02:09:52 AN INTENTIONAL PUMP STOP OCCURRED DUE TO DUE TO THE BACKUP BATTERY BEING DEPLETED OF POWER. DRIVELINE DISCONNECT ALARMS WERE NOT SEEN IN THE HISTORY. ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE OF DEATH WAS UNCLEAR. IT WAS ALSO UNCLEAR IF THE DEATH WAS CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2025. A COMPLETE SYSTEM SHUTDOWN DUE TO FULL BATTERY DEPLETION WAS CAPTURED (REFER TO MFR# 2916596-2025-02821). PRIOR TO THE PUMP STOP, THE DEVICE APPEARED TO BE OPERATING AS EXPECTED AT THE FIXED SPEED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22004332&pc=PNJ | 9612164-2025-02345 | 2025-04-17 04:00 | Death | MEDTRONIC IRELAND | 2025-05-12 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiomyopathy; Unspecified Infection; Septic Shock; Loss of consciousness | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT PROCEDURE OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), PATIENT HAD DEVELOPED PNEUMONIA. THE LEADLESS IPG WAS IMPLANTED SUCCESSFULLY WITHOUT ANY COMPLICATIONS. THE FOLLOWING DAY, DURING DIALYSIS, THE PATIENT EXPERIENCED CARDIAC ARREST AND LOSS OF CONSCIOUSNESS. TAKOTSUBO CARDIOMYOPATHY WAS SUSPECTED. EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS INTRODUCED. FURTHER, TWO DAYS LATER, IT WAS NOTED THAT THE PATIENT DEVELOPED AN INFECTION DUE TO THE COAGULATION OF THE BLOOD. CARDIAC ARREST WAS NOTED WITH MULTIPLE ORGAN FAILURE. EVENTUALLY, THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22004345&pc=DSQ | 2916596-2025-03057 | 2025-04-09 04:00 | Death | THORATEC CORPORATION | 2025-05-12 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Hemorrhagic Stroke; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 AFTER HAVING HAD A HEMORRHAGIC STROKE ON (B)(6) 2025. THE SITE TOOK THE PATIENT IN FOR A CRANIOTOMY AND EVACUATED THE CLOT AND STOPPED THE BLEEDING AND THEY WERE INITIALLY CONFIDENT THAT IT HAD WORKED. THE SEPSIS WAS ONGOING FROM THE PATIENT'S DRIVELINE AND ANTIBIOTICS WERE GIVEN THAT THE PATIENT WAS SENSITIVE TO. ON (B)(6) 2025 BRAIN STEM DEATH WAS DECLARED, THE PUMP WAS TURNED OFF, AND LOG FILES WERE OBTAINED.Event Description: IT WAS COMMUNICATED ON 09 JUN 2025 THAT THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED, AND THE DEVICE OPERATED AS EXPECTED. NO PRODUCTS WERE TO RETURN FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: A2 - PATIENT AGE - CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. B ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, PUMP POCKET), SEPSIS, STROKE, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THIS SECTION ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22004755&pc=DXY | 2182208-2025-02297 | 2024-08-21 04:00 | Death | MEDTRONIC, INC. | 2025-05-12 04:00 | DXY | MEDTRONIC ICD | Inappropriate/Inadequate Shock/Stimulation | Shock from Patient Lead(s) |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING SHOCK THERAPY IN PATIENTS WITH AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD). PATIENTS WITH AN ICD WERE FOLLOWED FOR APPROPRIATE AND INAPPROPRIATE INCIDENTS AND RECURRENT SHOCKS. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO EXPERIENCED INAPPROPRIATE SHOCKS DUE TO ATRIAL FIBRILLATION (AF)/ATRIAL FLUTTER, LEAD-RELATED ISSUES, ICD-RELATED ISSUES, AND T-WAVE OVERSENSING (TWOS). PATIENTS ALSO HAD DEVICE RELATED COMPLICATIONS WHICH INCLUDED SYSTEMIC INFECTIONS/ENDOCARDITIS, POCKET INFECTION, EROSION, SENSING/PACING FAILURES, AND PREMATURE BATTERY DEPLETION. THERE WERE ALSO LEAD-RELATED COMPLICATIONS WHICH INCLUDED DISLODGEMENT, HIGH THRESHOLDS, OVERSENSING, UNDERSENSING, HIGH IMPEDANCE, INSULATION FAILURES/CONDUCTION BREAKS, CARDIAC PERFORATION, EXTRACARDIAC STIMULATION, AND DEVICE-LEAD INTERFACE PROBLEMS. PATIENTS WERE TREATED WITH MEDICATION CHANGES, ABLATIONS, REPROGRAMMING, OR REPLACEMENTS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/62 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: INCIDENCE AND RISK FACTORS FOR FIRST AND RECURRENT ICD SHOCK THERAPY IN PATIENTS WITH AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR. JOURNAL OF INTERVENTIONAL CARDIAC ELECTROPHYSIOLOGY (2025) 68:125¿139. DOI: 10.1007/S10840-024-01873-0 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22004760&pc=DXY | 2182208-2025-02298 | 2024-08-21 04:00 | Death | MEDTRONIC, INC. | 2025-05-12 04:00 | DXY | MEDTRONIC LEAD | Inappropriate/Inadequate Shock/Stimulation | Shock from Patient Lead(s) |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING SHOCK THERAPY IN PATIENTS WITH AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD). PATIENTS WITH AN ICD WERE FOLLOWED FOR APPROPRIATE AND INAPPROPRIATE INCIDENTS AND RECURRENT SHOCKS. THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO EXPERIENCED INAPPROPRIATE SHOCKS DUE TO ATRIAL FIBRILLATION (AF)/ATRIAL FLUTTER, LEAD-RELATED ISSUES, ICD-RELATED ISSUES, AND T-WAVE OVERSENSING (TWOS). PATIENTS ALSO HAD DEVICE RELATED COMPLICATIONS WHICH INCLUDED SYSTEMIC INFECTIONS/ENDOCARDITIS, POCKET INFECTION, EROSION, SENSING/PACING FAILURES, AND PREMATURE BATTERY DEPLETION. THERE WERE ALSO LEAD-RELATED COMPLICATIONS WHICH INCLUDED DISLODGEMENT, HIGH THRESHOLDS, OVERSENSING, UNDERSENSING, HIGH IMPEDANCE, INSULATION FAILURES/CONDUCTION BREAKS, CARDIAC PERFORATION, EXTRACARDIAC STIMULATION, AND DEVICE-LEAD INTERFACE PROBLEMS. PATIENTS WERE TREATED WITH MEDICATION CHANGES, ABLATIONS, REPROGRAMMING, OR REPLACEMENTS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/62 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: INCIDENCE AND RISK FACTORS FOR FIRST AND RECURRENT ICD SHOCK THERAPY IN PATIENTS WITH AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR. JOURNAL OF INTERVENTIONAL CARDIAC ELECTROPHYSIOLOGY (2025) 68:125¿139. DOI: 10.1007/S10840-024-01873-0 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22004899&pc=DSQ | 2916596-2025-02768 | 2025-04-16 04:00 | Death | THORATEC CORPORATION | 2025-05-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Disconnection; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Air Embolism; Stroke/CVA; Cardiogenic Shock; Heart Failure/Congestive Heart Failure; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. ADDITIONAL INFORMATION WAS NOT PROVIDED. THE EVENT LOG FILES APPEARED TO SHOW AN ONSET OF LOW FLOW CAUSING MULTIPLE LOW FLOW HAZARD ALARMS ON (B)(6) 2025. THE PUMP DRIVELINE APPEARED TO HAVE BEEN DISCONNECTED A SHORT TIME LATER.Event Description: THE CAUSE OF DEATH WAS RIGHT VENTRICLE FAILURE AND STROKES FROM AIR EMBOLISM, DISTRIBUTIVE SHOCK ALONG WITH THE CARDIOGENIC SHOCK. Manufacturer Narrative: B7 - OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS: CORRECTED. E2 - HEALTH PROFESSIONAL: CORRECTED. E3 - OCCUPATION: CORRECTED. G2 - REPORT SOURCE: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THE ALARMS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE SUBMITTED CONTROLLER EVENT AND PERIODIC LOG FILES ARE REVIEWED UNDER THE INVESTIGATION OF THE SYSTEM CONTROLLER. OF NOTE, LOW FLOW ALARMS ACTIVATED ON (B)(6) 2025 AT 16:48:17 DUE TO THE CALCULATED FLOW DECREASING BELOW THE LOW FLOW THRESHOLD OF 2.5 LITERS PER MINUTE. ON (B)(6) 2025 AT 16:52:12, AN LVAD OFF ALARM ACTIVATED ASSOCIATED WITH AN INTENTIONAL PUMP STOP FAULT. SHORTLY AFTER THE INTENTIONAL PUMP STOP THE DRIVELINE WAS DISCONNECTED AT 17:01:30 WHICH APPEARS CONSISTENT WITH TERMINATION OF SUPPORT. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE DEVICE APPEARED TO OPERATE AS INTENDED AT THE SET SPEED. HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH, RIGHT HEART FAILURE, AND STROKE. SECTION 5 OF THE IFU, ¿SURGICAL PROCEDURES¿ INCLUDES INFORMATION FOR PRIMING THE PUMP AS WELL AS DE-AIRING THE PUMP. SECTION 5 WARNS THAT ALL ENTRAPPED AIR MUST BE REMOVED FROM THE PUMP/SEALED OUTFLOW GRAFT ASSEMBLY BLOOD PATH TO MINIMIZE THE RISK OF AIR EMBOLUS. TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) SHOULD BE UTILIZED TO MONITOR FOR AIR EMBOLI. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR,¿ EXPLAINS THAT THE LOW FLOW HAZARD ALARM OCCURS WHEN PUMP FLOW IS LESS THAN 2.5 LPM. A 10-SECOND DELAY IS IMPOSED BETWEEN THE DETECTION OF THE LOW FLOW STATUS AND THE ACTIVATION OF THE ASSOCIATED AUDIO AND VISUAL INDICATORS ON THE SYSTEM CONTROLLER. CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING,¿ EXPLAINS ALL SYSTEM ALARMS AND THE RECOMMENDED ACTIONS ASSOCIATED WITH THEM. THE HEARTMATE 3 LVAS PATIENT HANDBOOK ALSO OUTLINES ALL SYSTEM CONTROLLER ALARMS AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22005106&pc=OZD | 1220648-2025-28321 | 2025-04-20 04:00 | Death | ABIOMED, INC. | 2025-05-12 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Hemolysis; Hemorrhage/Bleeding; Renal Failure | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THE 75-YEAR-OLD FEMALE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THERE WAS ACCESS SITE BLEEDING. SITE WAS REDRESSED; BETTER ANGLE MATCH AND FORWARD TENSION WAS APPLIED. HEPARIN WAS WITHHELD DUE TO THE BLEEDING AND FEMSTOP WAS APPLIED. ADDITIONALLY, THE PATIENT HAD A SUSTAINED RUN OF ATRIAL FIBRILLATION (AF). THE PATIENT WAS TREATED WITH METOPROLOL. SUPPORT CONTINUED. ADDITIONALLY, THE PATIENT DEVELOPED HEMOLYSIS ON SUPPORT. THE HEMOLYSIS LED TO ORGAN DAMAGE AND CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) WAS STARTED. BLOOD PRODUCTS WERE GIVEN. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING, HEMOLYSIS, VT/VF/AF/AT, AND RENAL FAILURE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING - MAJOR WAS MOST LIKELY DUE TO PATIENT MOVEMENT AS EVIDENCED BY CLINICAL DETAIL OF HEMATOMA DEVELOPING AFTER MOVING PATIENT LEG THE ROOT CAUSE OF THE HEMOLYSIS, RENAL FAILURE, AND AF WAS NOT DETERMINED AS NO PRODUCT OR DATA LOGS WERE RETURNED FOR ANALYSIS AND LIMITED CLINICAL INFORMATION RELATED TO FAILURE WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22005454&pc=BZD | 2518422-2025-105503 | 2025-04-28 04:00 | Death | RESPIRONICS, INC. | 2025-05-12 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K113068 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT OF THE REMSTAR AUTO A-FLEX DEVICE PASSED AWAY. ACCORDING TO THE PATIENT'S NIECE, THE PATIENT PASSED AWAY IN (B)(6) 2024 AND SHE WOULD LIKE TO STOP GETTING LETTERS. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22005715&pc=NKM | 2135147-2025-02509 | 2018-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-05-12 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "PLASMA VOLUME STATUS PREDICTING CLINICAL OUTCOMES IN PATIENTS UNDERGOING TRANSCATHETER EDGE-TO-EDGE MITRAL VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE PROGNOSTIC SIGNIFICANCE OF PLASMA VOLUME STATUS (PVS) IN THESE PATIENTS. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT PREOPERATIVE PVS IS A STRONG INDEPENDENT PROGNOSTIC MARKER IN PATIENTS UNDERGOING M-TEER, CORRELATING WITH INCREASED RISK OF MORTALITY AND HF HOSPITALIZATION. PVS MAY PROVIDE VALUABLE CLINICAL INSIGHTS FOR PATIENT STRATIFICATION AND MANAGEMENT STRATEGIES IN M-TEER PATIENTS.. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MASANORI YAMAMOTO AT NAGOYA HEART CENTER INSTITUTION WITH CORRESPONDING EMAIL MASA-NORI@NMS.AC.JP; YAMAMOTO@HEART-CENTER.OR.JP]. THE TIME FRAME OF THE STUDY WAS APRIL 2018 TO JUNE 2023. A TOTAL OF 3,763 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 3,763 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DYSLIPIDEMIA, DIABETES, ATRIAL FILIATION, CORONARY ARTERY DISEASE, PRIOR STROKE. (B)(6), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DYSLIPIDEMIA, DIABETES, ATRIAL FILIATION, CORONARY ARTERY DISEASE, PRIOR STROKE. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22005854&pc=MKJ | 1220908-2025-01558 |  | Death | ZOLL MEDICAL CORPORATION | 2025-05-12 04:00 | MKJ | X SERIES | Unable to Obtain Readings | Cardiac Arrest | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A PATIENT IN CARDIAC ARREST (AGE UNKNOWN AND GENDER: FEMALE) THE DEVICE WAS UNABLE TO OBTAIN AN ECG SIGNAL VIA ELECTRODE PADS. THE COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. BASED ON LOG REVIEW, ECG WAS AVAILABLE IF THE USER SWITCHED FROM LEAD II TO PADS VIEW. LEAD II HAD AN INTERMITTENT SIGNAL SUGGESTING COUPLING WAS A FACTOR, BUT PAD VIEW HAD A STABLE SIGNAL THROUGHOUT THE CASE. THE X SERIES DEVICE REQUIRES MANUAL SELECTION OF PADS VIEW AND DOES NOT SWITCH AUTOMATICALLY WHEN PADS ARE CONNECTED. THE DEVICE WAS FUNCTIONING PROPERLY, AND NO FAULTS WERE FOUND. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS FOR REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22005893&pc=NPT | 2025587-2025-03614 | 2024-05-10 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-12 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Bacterial Infection; Fever; Unspecified Infection; Sepsis | P130021 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC A SEVERE HEMORRHAGE AT THE RIGHT FEMORAL ACCESS SITE WAS NOTED. BLEEDING PERSISTED DESPITE THE USE OF 1 PERCLOSE AND MANUAL PRESSURE. SUBSEQUENTLY, THE HEMORRHAGE WAS STOPPED SURGICALLY. APPROXIMATELY TWO AND A HALF MONTHS LATER, THE PATIENT DEVELOPED A FEVER OF AN UNKNOWN CAUSE AND WAS HOSPITALIZED FOR UNSPECIFIED ¿DETAILED EXAMINATION AND TREATMENT¿. PER THE PHYSICIAN, THE VALVE AND PROCEDURE WERE UNRELATED TO THE FEVER. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN; THE DELIVERY CATHETER SYSTEM AND THE PATIENT'S "SLIGHTLY MEANDERING" ANATOMY COULD HAVE CONTRIBUTED TO THE INJURY. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE PATIENT DEVELOPED SEPTICEMIA SEPSIS AND SUBSEQUENTLY DIED THREE AND A HALF MONTHS AFTER THE VALVE IMPLANT. THE CAUSE OF THE SEPTICEMIA, SEPSIS, AND DEATH WERE UNKNOWN; HOWEVER, THE PHYSICIAN INDICATED NEITHER THE VALVE NOR THE VALVE IMPLANT PROCEDURE WERE CONTRIBUTING FACTORS. THE CAUSE OF DEATH WAS SEPSIS INFECTION. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22005938&pc=DYB | 2124215-2025-29369 | 2025-04-28 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-12 04:00 | DYB | ISLEEVE 14F | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Perforation of Vessels; Thrombosis/Thrombus |  |  | 1 | Event Description: IT WAS REPORTED THAT VESSEL TRAUMA AND PATIENT DEATH OCCURRED. THE PATIENT EXPERIENCED WORSENING COUGH AND GENERAL WEAKNESS AND WAS ADMITTED TO THE HOSPITAL FOR HEART FAILURE MANAGEMENT. IMAGING IDENTIFIED BILATERAL PLEURAL EFFUSION AND PNEUMONIA. TWO (2) DAYS POST HOSPITAL ADMITTANCE, TRANSTHORACIC ECHOCARDIOGRAPHY (TTE) SHOWED SEVERE AORTIC STENOSIS WITH AN EJECTION FRACTION OF 52 PERFECT. NINE (9) DAYS POST HOSPITAL ADMITTANCE A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WAS PERFORMED. VASCULAR ACCESS WAS OBTAINED VIA A MILDLY TORTUOUS TRANSFEMORAL APPROACH. THE NATIVE AORTIC ANNULUS MEASURED 24.6MM IN DIAMETER. A 14F ISLEEVE INTRODUCER SHEATH WAS PLACED AND SAFARI2 GUIDEWIRE WAS ADVANCED INTO POSITION. BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED WITH TWO (2) INFLATIONS OF A 23MM NON-BOSTON SCIENTIFIC (BSC) BALLOON CATHETER. AN ACURATE NEO2 VALVE SIZE MEDIUM WAS PREPARED AND LOADED ONTO AN ACURATE NEO2 TRANSFEMORAL (TF) DELIVERY SYSTEM (DS). THE ACURATE NEO2 DS WAS ADVANCED INTO POSITION, COMMISSURAL ALIGNMENT WAS ACHIEVED, AND THE ACURATE NEO2 VALVE WAS IMPLANTED AT THE INTENDED LOCATION. POST-DILATATION BAV WAS PERFORMED WITH ONE (1) INFLATION OF A NON-BSC BALLOON CATHETER. POST-PROCEDURE THE PATIENT WAS TRANSFERRED TO THE CRITICAL CARE UNIT DUE TO HYPOTENSION. LOWER ABDOMINAL DISTENSION WAS IDENTIFIED AND COMPUTED TOMOGRAPHY (CT) SHOWED A RUPTURE OF THE RIGHT EXTERNAL ILIAC ARTERY. ANGIOPLASTY USING A BOVINE PATCH WAS PERFORMED. THE PATIENT RETURNED TO THE CRITICAL CARE UNIT AND METABOLIC ACIDOSIS, INCREASED LACTIC ACID, ANURIA, AND PERSISTENT BLEEDING OCCURRED. BLOOD TRANSFUSIONS WERE ADMINISTERED, HEMOGLOBIN LEVELS WERE MONITORED, THE PATIENT WAS INTUBATED, AND CONTINUOUS RENAL REPLACEMENT THERAPY WITH CONSERVATIVE MANAGEMENT WAS PERFORMED. THE SYSTEMIC CONDITION OF THE PATIENT WORSENED. THE PATIENT HAD AN ADVANCED DIRECTIVE INDICATING LIFE SUSTAINING TREATMENT INCLUDING CARDIOPULMONARY RESUSCITATION WAS NOT TO BE PERFORMED. THE PATIENT DIED WITHOUT LIFE-SUSTAINING TREATMENT IN ACCORDANCE WITH THE ADVANCED DIRECTIVE.Event Description: IT WAS REPORTED THAT VESSEL TRAUMA AND PATIENT DEATH OCCURRED. THE PATIENT EXPERIENCED WORSENING COUGH AND GENERAL WEAKNESS AND WAS ADMITTED TO THE HOSPITAL FOR HEART FAILURE MANAGEMENT. IMAGING IDENTIFIED BILATERAL PLEURAL EFFUSION AND PNEUMONIA. TWO (2) DAYS POST HOSPITAL ADMITTANCE, TRANSTHORACIC ECHOCARDIOGRAPHY (TTE) SHOWED SEVERE AORTIC STENOSIS WITH AN EJECTION FRACTION OF 52 PERFECT. NINE (9) DAYS POST HOSPITAL ADMITTANCE A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WAS PERFORMED. VASCULAR ACCESS WAS OBTAINED VIA A MILDLY TORTUOUS TRANSFEMORAL APPROACH. THE NATIVE AORTIC ANNULUS MEASURED 24.6MM IN DIAMETER. A 14F ISLEEVE INTRODUCER SHEATH WAS PLACED AND SAFARI2 GUIDEWIRE WAS ADVANCED INTO POSITION. BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED WITH TWO (2) INFLATIONS OF A 23MM NON-BOSTON SCIENTIFIC (BSC) BALLOON CATHETER. AN ACURATE NEO2 VALVE SIZE MEDIUM WAS PREPARED AND LOADED ONTO AN ACURATE NEO2 TRANSFEMORAL (TF) DELIVERY SYSTEM (DS). THE ACURATE NEO2 DS WAS ADVANCED INTO POSITION, COMMISSURAL ALIGNMENT WAS ACHIEVED, AND THE ACURATE NEO2 VALVE WAS IMPLANTED AT THE INTENDED LOCATION. POST-DILATATION BAV WAS PERFORMED WITH ONE (1) INFLATION OF A NON-BSC BALLOON CATHETER. POST-PROCEDURE THE PATIENT WAS TRANSFERRED TO THE CRITICAL CARE UNIT DUE TO HYPOTENSION. LOWER ABDOMINAL DISTENSION WAS IDENTIFIED AND COMPUTED TOMOGRAPHY (CT) SHOWED A RUPTURE OF THE RIGHT EXTERNAL ILIAC ARTERY. ANGIOPLASTY USING A BOVINE PATCH WAS PERFORMED. THE PATIENT RETURNED TO THE CRITICAL CARE UNIT AND METABOLIC ACIDOSIS, INCREASED LACTIC ACID, ANURIA, AND PERSISTENT BLEEDING OCCURRED. BLOOD TRANSFUSIONS WERE ADMINISTERED, HEMOGLOBIN LEVELS WERE MONITORED, THE PATIENT WAS INTUBATED, AND CONTINUOUS RENAL REPLACEMENT THERAPY WITH CONSERVATIVE MANAGEMENT WAS PERFORMED. THE SYSTEMIC CONDITION OF THE PATIENT WORSENED. THE PATIENT HAD AN ADVANCED DIRECTIVE INDICATING LIFE SUSTAINING TREATMENT INCLUDING CARDIOPULMONARY RESUSCITATION WAS NOT TO BE PERFORMED. THE PATIENT DIED WITHOUT LIFE-SUSTAINING TREATMENT IN ACCORDANCE WITH THE ADVANCED DIRECTIVE. IT WAS FURTHER REPORTED DURING PLACEMENT OF THE NON-BSC ACCESS SITE CLOSURE DEVICE, THE PATIENT BECAME HYPOTENSIVE. MEDICATION WAS ADMINISTERED AND BLOOD PRESSURE RECOVERED. WHEN REMOVING THE 14F ISLEEVE INTRODUCER SHEATH MULTIPLE THROMBI WERE PRESENT ON THE INTRODUCER SHEATH. DURING CLOSURE OF THE FEMORAL ACCESS SITE A CONTRAST ASSESSMENT WAS NOT PERFORMED AND THE EXACT TIMING OF THE ILIAC RUPTURE WAS UNABLE TO BE IDENTIFIED ALTHOUGH IT WAS SUSPECTED THE RUPTURE OCCURRED DURING ADVANCEMENT OR WITHDRAWAL OF THE 14F ISLEEVE INTRODUCER SHEATH. THE OFFICIAL CAUSE OF DEATH WAS DETERMINED TO BE THE ILIAC RUPTURE. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED WITH AN UPDATE TO SECTIONS: B2 DATE OF DEATH, B5 DESCRIBE EVENT OR PROBLEM, H6 PATIENT CODES, H6 EVALUATION METHOD CODES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22006894&pc=QGM | 3027664504-2025-00007 | 2025-04-22 04:00 | Death | HISTOSONICS, INC. | 2025-05-12 04:00 | QGM | EDISON SYSTEM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Urinary Retention; Renal Impairment | K233466 |  | 1 | Event Description: A 81-YEAR-OLD MALE PATIENT WITH ADVANCED METASTATIC CHOLANGIOCARCINOMA RECEIVED REPEAT HISTOTRIPSY TREATMENT (INITIAL TREATMENT (B)(6) 2025, TOTAL PTV 49.0 CC) ON (B)(6) 2025, TO A SINGLE LESION IN SEGMENT V, WITH TWO PARTIALLY OVERLAPPING PLANNED TREATMENT VOLUMES (PTVS) REPORTED AS PTV 1: TOTAL PTV = 31 CC, PTV 2: TOTAL PTV = 29.5CC, FOR A TOTAL PTV OF 60.5 CC. THE PATIENT WAS DISCHARGED POST-PROCEDURE WITHOUT INCIDENT AND ULTIMATELY RETURNED HOME. ON POST-OPERATIVE DAY 3 ((B)(6) 2025), HE DEVELOPED ANURIA, AND HIS CREATININE ROSE TO 7. THE TREATING PHYSICIAN REPORTED THAT THE PATIENT UNDERWENT HEMODIALYSIS AND HAD SOME URINE OUTPUT, WITH A DIAGNOSIS OF ACUTE TUBULAR NECROSIS (ATN) ATTRIBUTED TO A COMBINATION OF ANESTHESIA, HISTOTRIPSY, AND NSAID USE, WITH POSSIBLE CONTRIBUTION FROM ACUTE INTERSTITIAL NEPHRITIS (AIN). THE PATIENT REMAINED DIALYSIS-DEPENDENT, BUT ON-GOING HYPOTENSION COMPLICATED MANAGEMENT. THE FAMILY ULTIMATELY ELECTED TO WITHDRAW CARE AND THE PATIENT DIED ON (B)(6) 2025. NO DEVICE MALFUNCTIONS OR OTHER NOTEWORTHY EVENTS OCCURRED DURING THE CASE. Manufacturer Narrative: NO DEVICE MALFUNCTIONS OR OTHER NOTEWORTHY EVENTS OCCURRED DURING THE CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22006995&pc=NVN | 2017865-2025-53503 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-12 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO THEIR PACEMAKER SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22006999&pc=NVN | 2017865-2025-53506 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-12 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO THEIR PACEMAKER SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22007071&pc=NVN | 2017865-2025-53511 | 2025-04-21 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-12 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR MAGNETIC RESONANCE IMAGING (MRI) SCAN. THE MRI STAFF REQUESTED THE PULSE GENERATOR REPROGRAMMING FOR MRI COMPATIBILITY. THE DEVICE WAS SET TO VENTRICULAR PACING ONLY MODE AT 85 BEATS PER MINUTE PROVIDING 100% PACING SUPPORT DURING THE SCAN. DURING THE MRI THE PATIENT BECAME UNSTABLE, CODED, AND RESUSCITATION WAS ATTEMPTED. THE PATIENT LATER EXPIRED; THERE IS NO ALLEGATION AT THIS TIME THAT THE THAT THE GENERATOR OR RIGHT VENTRICULAR (RV) LEAD MALFUNCTION CONTRIBUTED TO THE PATIENT¿S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22007263&pc=KGG | 3008114965-2025-00387 | 2025-04-02 04:00 | Death | MEDOS INTERNATIONAL SARL | 2025-05-12 04:00 | KGG | TRUFILL | Improper or Incorrect Procedure or Method; Adverse Event Without Identified Device or Use Problem | Ischemia | P990040 |  | 1 | Event Description: THE EVENT WAS REPORTED VIA THE VOLUNTARY MEDWATCH (MW5169480). THE PATIENT WAS ADMITTED TO INTERVENTIONAL RADIOLOGY (IR) FOR COILING AND GLUE PLACEMENT TO TREAT AN ENDOLEAK. THE PATIENT UNDERWENT THE ENDOVASCULAR EMBOLIZATION PROCEDURE WHERE THE TRUFILL® N-BCA LIQUID EMBOLIC SYSTEM-2 GRAM KIT (631400 / LOT# UNKNOWN) WAS USED. AFTER THE PROCEDURE, THE PATIENT HAD DISCHARGE ORDERS BUT STARTED TO EXPERIENCE ACUTE ABDOMINAL PAIN AND THE PATIENT LOST THE PULSE IN HIS LEFT FOOT. THE PATIENT WAS ADMITTED TO THE CRITICAL CARE UNIT (CCU) AND CT ANGIOGRAM SHOWED HIGH-DENSITY MATERIAL IN THE RIGHT HEPATIC, DISTAL SPLENIC ARTERIES, AND DISTAL SUPERIOR MESENTERIC ARTERY (SMA). THE PATIENT WAS TAKEN TO THE OPERATING ROOM (OR) EMERGENTLY FOR AORTOGRAM, SMA STENT, LEFT POPLITEAL CUT DOWN, AND TIBIAL THROMBECTOMY, LEFT LOWER EXTREMITY ANGIOGRAM, LEFT DORSALIS PEDIS CUT DOWN, AND PEDAL THROMBECTOMY. IT WAS ALSO REPORTED THAT ON AN UNSPECIFIED / UNKNOWN DATE, THE PATIENT ALSO HAD AN EXPLORATORY LAPAROTOMY WITH RIGHT HEMICOLECTOMY. SURGICAL FINDINGS INCLUDE AN ISCHEMIC ILEUM WITH A WELL-PERFUSED JEJUNUM AND COLON. ON ANOTHER UNSPECIFIED / UNKNOWN DATE, THE PATIENT WAS RETURNED TO THE OR FOR A SECOND LOOK. THE FINDINGS INDICATED THAT THE MAJORITY OF THE SMALL BOWELS ARE NONVIABLE, AND THE GALLBLADDER IS NECROTIC. THERE ARE LESS THAN 40 CM OF VIABLE BOWELS, AND THIS AREA IS PUNCTUATED WITH SMALL PATCHES OF ISCHEMIA. IT WAS THE UNDERSTANDING OF THE NURSING STAFF THAT THE GLUE TRAVELED THROUGH THE VASCULAR SYSTEM WHICH RESULTED IN THE ISCHEMIC LEG AND NECROTIC BOWELS. IT WAS REPORTED THAT WHEN THE PHYSICIAN SPOKE WITH THE FAMILY, HE MENTIONED THAT HE WAS UNSURE HOW THIS COULD HAVE HAPPENED. PALLIATIVE CARE WAS CONSULTED AND THE PATIENT¿S FAMILY DECIDED TO TRANSITION THE PATIENT TO COMFORT CARE. THE PATIENT PASSED AWAY (THE DATE OF DEATH WAS UNSPECIFIED / UNKNOWN). Manufacturer Narrative: MANUFACTURER¿S REF. NO: (B)(4). INFORMATION REGARDING PATIENT IDENTIFIER, DATE OF BIRTH, GENDER, AND WEIGHT, WERE NOT PROVIDED. SECTION B.2: DATE OF DEATH: THE DATE OF DEATH IS NOT KNOWN. SECTION D.4: THE EXPIRATION DATE OF THE DEVICE IS NOT KNOWN AS THE DEVICE LOT NUMBER IS NOT AVAILABLE / NOT REPORTED. THE FULL UDI IS NOT AVAILABLE WITHOUT THE EXPIRATION DATE. SECTION H.4: THE DEVICE MANUFACTURE DATE IS NOT KNOWN AS THE DEVICE LOT NUMBER IS NOT AVAILABLE / NOT REPORTED. BASED ON COMPLAINT INFORMATION, THE DEVICE REMAINS IMPLANTED AND IS THUS NOT AVAILABLE FOR EVALUATION. THE DEVICE LOT NUMBER WAS NOT AVAILABLE. THE DEVICE HISTORY RECORD (DHR) REVIEW COULD NOT BE PERFORMED WITHOUT THE LOT NUMBER. NON-TARGET EMBOLIZATION IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH TRUFILL N-BCA LIQUID EMBOLIC AGENT AND IS LISTED IN THE INSTRUCTIONS FOR USE (IFU) AS SUCH. THERE WAS NO ALLEGED QUALITY ISSUE RELATED TO THE USED DEVICE, AS THE DEVICE PERFORMED AS INTENDED. AS EXPLAINED IN THE REFERENCED MEDICAL LITERATURE, ISCHEMIC COMPLICATIONS THAT OCCUR IMMEDIATELY AFTER ENDOVASCULAR AORTIC REPAIR (EVAR) CAN BE DUE TO CLOT FORMATION OR CLOT EMBOLIZATION INTO AORTIC SIDE BRANCHES AND INCLUDE COLONIC, RENAL, AND PELVIC ISCHEMIA. IN PATIENTS WITH CONCURRENT SMALL BOWEL ISCHEMIA, THROMBUS OR ATHEROMA IN THE ANEURYSM NECK OR ILIAC ARTERIES WERE FOUND TO BE DISLODGED. IT IS THOUGHT THAT DURING THE PROCEDURE, DEBRIS CAN BE FLUSHED UPSTREAM INTO THE SUPERIOR MESENTERIC ARTERY BY BLOOD FLOW TURBULENCE AND MECHANICAL MANIPULATION. THE MICROEMBOLI MAY BE FLUSHED INTO RENAL ARTERIES, THE INFERIOR MESENTERIC ARTERY, THE HYPOGASTRIC ARTERY, AND DISTAL VESSELS, LEADING TO ACUTE RENAL INSUFFICIENCY, COLON ISCHEMIA, AND ISCHEMIC LOWER EXTREMITIES, AS WELL. PATTERNS OF SEGMENTAL OR PATCHY ISCHEMIA SUPPORT THE CONCEPT THAT MICROEMBOLIZATION HAS AN IMPORTANT ROLE IN ISCHEMIC COMPLICATIONS AFTER EVAR. THE EVENT MAY HAVE BEEN RELATED TO CLOT FORMATION/MICROEMBOLIZATION DURING THE PROCEDURE, AS THERE WAS NO INDICATION THAT NON-TARGET EMBOLIZATION OCCURRED DURING THE USE OF TRUFILL N-BCA. HOWEVER, THE CORRELATING RELATIONSHIP BETWEEN THE EVENT OF THE USED TRUFILL N-BCA AS A CAUSAL OR CONTRIBUTING FACTOR CANNOT RULED OUT ENTIRELY. ADDITIONALLY, THE NURSING STAFF FELT THE EVENT WAS POSSIBLE ATTRIBUTED TO GLUE TRAVELING THROUGH THE VASCULAR SYSTEM AND CAUSING THE ISCHEMIC LEG AND NECROTIC BOWEL. BASED ON THIS INFORMATION AND THE INABILITY TO RULE OUT THE DEVICE AS A CONTRIBUTING FACTOR, THIS EVENT MEETS USA FDA REPORTING CRITERIA UNDER 21 CFR 803 WITH A CLASSIFICATION OF ¿DEATH.¿ THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. REFERENCES: MALEUX G, KOOLEN M, HEYE S. COMPLICATIONS AFTER ENDOVASCULAR ANEURYSM REPAIR. SEMIN INTERVENT RADIOL. 2009 MAR;26(1):3-9. DOI: 10.1055/S-0029-1208377. PMID: 21326525; PMCID: PMC3036452. ZHANG WW, KULAYLAT MN, ANAIN PM, DOSLUOGLU HH, HARRIS LM, CHERR GS, DAYTON MT, DRYJSKI ML. EMBOLIZATION AS CAUSE OF BOWEL ISCHEMIA AFTER ENDOVASCULAR ABDOMINAL AORTIC ANEURYSM REPAIR. J VASC SURG. 2004 NOV;40(5):867-72. DOI: 10.1016/J.JVS.2004.08.054. PMID: 15557898. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. THE MANUFACTURER WILL SUBMIT A SUPPLEMENTAL REPORT IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT. ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22007294&pc=BZD | 2518422-2025-105520 | 2025-05-07 04:00 | Death | RESPIRONICS, INC. | 2025-05-12 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF A DREAMSTATION AUTO CPAP. THE SPOUSE OF THE PATIENT REPORTED THE PATIENT HAS SADLY PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21999495&pc=LWQ | 1649833-2025-00014 | 2025-04-14 04:00 | Death | ON-X LIFE TECHNOLOGIES, INC. | 2025-05-10 04:00 | LWQ | ONX MITRAL STANDARD 25 | Adverse Event Without Identified Device or Use Problem | Mitral Valve Insufficiency/ Regurgitation | P000037 |  | 1 | Event Description: ACCORDING TO THE INITIAL REPORT RECEIVED VIA EMAIL ON 15APRIL2025 FROM ARTIVION SALES MANAGER NORTH INDIA, (B)(6)., MITRAL VALVE ONXM-25 WAS IMPLANTED. VALVE STOPPED FUNCTIONING AS PER DR. (B)(6) AND PATIENT VITALS STARTED GETTING DETERIORATED THEN REDO SURGERY WAS DONE WITH ANOTHER ONXM-25 MITRAL VALVE. ADDITIONAL INFORMATION RECEIVED RELAYED A 49-YEAR-OLD FEMALE HAD THE ONXM-25 VALVE IMPLANTED ON (B)(6) 2025 DUE TO RHD WITH SEVERE MITRAL REGURGITATION AND MODERATE MITRAL STENOSIS. MODERATE AORTIC REGURGITATION WITH ATRIAL FIBRILATION WITH SMALL AORTIC ROOT WITH AF WITH NYHA CLASS IV WAS ALSO PRESENT. AVR WITH AORTIC ROOT ENLARGEMENT BY NICK'S PROCEDURE WITH MVR WERE PERFORMED ON (B)(6) 2025. SURGERY WAS UNEVENTFUL, TEE SHOWED NORMAL PROSTHESIS FUNCTION. PATIENT WAS EXTUBATED NEXT MORNING AND WAS BEHAVING NORMAL. THE PATIENT DEVELOPED CARDIAC DYSFUNCTION ON 1ST POD AND WAS MANAGED; ACCORDINGLY, TEE WAS DONE WHICH SHOWED REGURGITATION ACROSS MITRAL PROSTHESIS. A RE-DO MVR WAS DONE ON (B)(6) 2025 DUE TO POSSIBLE PROSTHESIS DYSFUNCTION OR PARAVALVULAR LEAK WITH NEW PROSTHESIS IMPLANTED. THE PATIENT DEVELOPED SEVERE LEFT VENTRICULAR DYSFUNCTION AND WAS MANAGED WITH IABP AND IONOTROPIC SUPPORT. PATIENT HAD PERRSISTENT LOW CARDIAC OUTPUT SYNDROME AND DEVELOPED MULTIORGAN FAILURE. THE PATIENT DID NOT SURVIVE DESPITE BEST EFFORTS. SEE ATTACHMENT TITLED: MEDICAL RECORDS FOR ADDITIONAL INFORMATION. THE SURGEON IS A FREQUENT USER OF ON-X VALVES IMPLANTING 6-7/MONTH. NO ADDITIONAL INFORMATION FORTHCOMING. THIS INVESTIGATION IS RELEGATED TO ONXM-25 SERIAL NUMBER: (B)(6) WITH THE ALLEGATION THAT THE VALVE STOPPED FUNCTIONING.Event Description: ACCORDING TO THE INITIAL REPORT RECEIVED VIA EMAIL ON 15APRIL2025 FROM (B)(6), MITRAL VALVE ONXM-25 WAS IMPLANTED. VALVE STOPPED FUNCTIONING AS PER DR. (B)(6) AND PATIENT VITALS STARTED GETTING DETERIORATED THEN REDO SURGERY WAS DONE WITH ANOTHER ONXM-25 MITRAL VALVE. ADDITIONAL INFORMATION RECEIVED RELAYED A 49 YEAR-OLD FEMALE HAD THE ONXM-25 VALVE IMPLANTED ON (B)(6) 2025 DUE TO RHD WITH SEVERE MITRAL REGURGITATION AND MODERATE MITRAL STENOSIS. MODERATE AORTIC REGURGITATION WITH ATRIAL FIBRILATION WITH SMALL AORTIC ROOT WITH AF WITH NYHA CLASS IV WAS ALSO PRESENT. AVR WITH AORTIC ROOT ENLARGEMENT BY NICK'S PROCEDURE WITH MVR WERE PERFORMED ON (B)(6) 2025. SURGERY WAS UNEVENTFUL, TEE SHOWED NORMAL PROSTHESIS FUNCTION. PATIENT WAS EXTUBATED NEXT MORNING AND WAS BEHAVING NORMAL. THE PATIENT DEVELOPED CARDIAC DYSFUNCTION ON 1ST POD AND WAS MANAGED ACCORDINGLY, TEE WAS DONE WHICH SHOWED REGURGITATION ACROSS MITRAL PROSTHESIS. A RE-DO MVR WAS DONE ON (B)(6) 2025 DUE TO POSSIBLE PROSTHESIS DYSFUNCTION OR PARAVALVULAR LEAK WITH NEW PROSTHESIS IMPLANTED. THE PATIENT DEVELOPED SEVERE LEFT VENTRICULAR DYSFUNCTION AND WAS MANAGED WITH IABP AND IONOTROPIC SUPPORT. PATIENT HAD PERRSISTENT LOW CARDIAC OUTPUT SYNDROME AND DEVELOPED MULTIORGAN FAILURE. THE PATIENT DID NOT SURVIVE DESPITE BEST EFFORTS. SEE ATTACHMENT TITLED: MEDICAL RECORDS FOR ADDITIONAL INFORMATION. THE SURGEON IS A FREQUENT USER OF ON-X VALVES IMPLANTING 6-7/MONTH. NO ADDITIONAL INFORMATION FORTHCOMING. Manufacturer Narrative: THE MANUFACTURING RECORDS FOR ONXM-25 SN (B)(6) WERE REVIEWED, AND IT WAS CONFIRMED THAT ALL RECORDS WERE CONTROLLED, AVAILABLE FOR REVIEW, AND MET ALL SPECIFICATIONS PER THE DEVICE MASTER RECORD. ALL LOTS PASSED FUNCTIONAL TESTING AND MET RELEASE SPECIFICATIONS. AN ONXM-25 SN (B)(6) WAS IMPLANTED ON (B)(6) 2025 IN A 49-YEAR-OLD FEMALE PATIENT IN INDIA. THE PATIENT DEVELOPED COMPLICATIONS DESCRIBED BY THE SURGEON AS ¿THE VALVE STOPPED FUNCTIONING¿ AND THE PATIENTS VITAL SIGNS DETERIORATED, AND A REDO SURGERY WAS UNDERTAKEN. ADDITIONAL INFORMATION RECEIVED RELAYED A ¿49-YEAR-OLD FEMALE HAD THE ONXM-25 VALVE IMPLANTED ON (B)(6) 2025 DUE TO RHD [RHEUMATIC HEART DISEASE] WITH SEVERE MITRAL REGURGITATION AND MODERATE MITRAL STENOSIS. MODERATE AORTIC REGURGITATION WITH ATRIAL FIBRILLATION WITH SMALL AORTIC ROOT WITH NYHA CLASS IV WAS ALSO PRESENT. AVR [ AORTIC VALVE REPLACEMENT] WITH AORTIC ROOT ENLARGEMENT BY NICK'S PROCEDURE WITH MVR [MITRAL VALVE REPLACEMENT] WERE PERFORMED ON (B)(6) 2025. SURGERY WAS UNEVENTFUL, TEE SHOWED NORMAL PROSTHESIS FUNCTION. PATIENT WAS EXTUBATED NEXT MORNING AND WAS BEHAVING NORMAL. THE PATIENT DEVELOPED CARDIAC DYSFUNCTION ON 1ST POD AND WAS MANAGED ACCORDINGLY, TEE WAS DONE WHICH SHOWED REGURGITATION ACROSS MITRAL PROSTHESIS. A RE-DO MVR WAS DONE ON (B)(6) 2025 (2 DAYS POST IMPLANT) DUE TO POSSIBLE PROSTHESIS DYSFUNCTION OR PARAVALVULAR LEAK WITH NEW PROSTHESIS IMPLANTED. THE PATIENT DEVELOPED SEVERE LEFT VENTRICULAR DYSFUNCTION AND WAS MANAGED WITH IABP AND IONOTROPIC SUPPORT. PATIENT HAD PERSISTENT LOW CARDIAC OUTPUT SYNDROME AND DEVELOPED MULTIORGAN FAILURE. THE PATIENT DID NOT SURVIVE DESPITE BEST EFFORTS.¿ MEDICAL RECORDS AND VIDEO WERE PROVIDED AND REVIEWED; THE VALVE WAS NOT RETURNED FOR EVALUATION. MANUFACTURING RECORDS WERE REVIEWED AND SHOWED NO PROCESSING ISSUES. THE INSTRUCTIONS FOR USE (IFU) FOR THE ON-X VALVE IDENTIFY PARAVALVULAR LEAK (PVL) AND DEATH AS POTENTIAL COMPLICATIONS FOLLOWING PROSTHETIC VALVE REPLACEMENT, WHICH MAY NECESSITATE REOPERATION OR EXPLANTATION. HISTORICALLY, THE INCIDENCE OF MAJOR PVL HAS BEEN REPORTED AT 0.3% PER PATIENT-YEAR FOR RIGID HEART VALVE SUBSTITUTES, AS OUTLINED IN ISO 5840-2:2021(E). ADDITIONALLY, PUBLISHED DATA ESTIMATE AN OPERATIVE MORTALITY RATE OF 4.5% FOLLOWING MITRAL VALVE REPLACEMENT SURGERY (BOWDISH ET AL.). THE EXPLANTING SURGEON IDENTIFIED PARAVALVULAR LEAK (PVL) AS THE IMMEDIATE REASON FOR VALVE EXPLANTATION. HOWEVER, THE LIMITED DETAILS AVAILABLE HINDER A DEFINITIVE ASSESSMENT OF THE UNDERLYING CAUSE OF THE PVL. KNOWN CONTRIBUTORS TO PVL INCLUDE SUTURE DEHISCENCE, ANNULAR CALCIFICATION OR TISSUE FRIABILITY, INFECTION, AND INTRAOPERATIVE TECHNICAL FACTORS. BASED ON THE INFORMATION CURRENTLY AVAILABLE, IT IS NOT POSSIBLE TO DETERMINE WHETHER THE VALVE ITSELF PLAYED A ROLE IN THE DEVELOPMENT OF THE PVL THAT LED TO EXPLANTATION AND SUBSEQUENT PATIENT DEATH. THE EXPLANTING SURGEON IDENTIFIED PARAVALVULAR LEAK (PVL) AS THE IMMEDIATE REASON FOR VALVE EXPLANTATION. HOWEVER, THE LIMITED DETAILS AVAILABLE HINDER A DEFINITIVE ASSESSMENT OF THE UNDERLYING CAUSE OF THE PVL. BASED ON THE INFORMATION CURRENTLY AVAILABLE, IT IS NOT POSSIBLE TO DETERMINE WHETHER THE VALVE ITSELF PLAYED A ROLE IN THE DEVELOPMENT OF THE PVL THAT LED TO EXPLANTATION AND SUBSEQUENT PATIENT DEATH. THERE IS NO INDICATION THAT AN ERROR OR DEFICIENCY OCCURRED AT ARTIVION- FORMERLY CRYOLIFE/JOTEC AND THE IFU ADEQUATELY COMMUNICATES RISK. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION. Manufacturer Narrative: THIS INVESTIGATION IS CURRENTLY ONGOING. ANY ADDITIONAL INFORMATION WILL BE PROVIDED IN THE FOLLOW-UP REPORT. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21992565&pc=BZD | 2518422-2025-044693 | 2024-02-21 05:00 | Death | RESPIRONICS, INC. | 2025-05-09 04:00 | BZD | REMSTAR AUTO A-FLEX | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT DUE TO CANCER. THERE WAS A REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21992854&pc=OZD | 1220648-2025-28242 | 2025-04-18 04:00 | Death | ABIOMED, INC. | 2025-05-09 04:00 | OZD | IMPELLA 5.5 | Positioning Failure; Restricted Flow rate | Hemorrhagic Stroke | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED AN IMPELLA 5.5 PUMP WAS SUPPORTING A PATIENT ADMITTED IN AND AWAITING MULTI-ORGAN TRANSPLANT. DURING THE MONTH-LONG SUPPORT, THERE WERE POSITIONAL ALARMS AND THE PUMP WAS BEING SHALLOW. THE PUMP WAS REPOSITIONED. THE TEAM ALSO NOTED SUCTION DURING THE SUPPORT. IT WAS LATER REPORTED THAT CARE WOULD BE WITHDRAWN DUE TO A HEMORRHAGIC CEREBROVASCULAR ACCIDENT. ONCE CARE WAS WITHDRAWN, THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES, LOW PUMP FLOW, AND STROKE/CVA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. POSITIONING ISSUES: THE CAUSE OF THE POSITIONING ISSUES WAS MOST LIKELY PATIENT CONDITION RELATED TO THE PATIENT'S SMALL LV. LOW PUMP FLOW: CLINICAL REPORTS SUCTION ALARMS ABOVE P-6, PATIENT HAD SMALL LV AND UNABLE TO REPOSITIONING AND OPTIMIZE POSITION. PUMP LEFT SHALLOW AND SUCTION REPORTED WITH NO RESOLVE. DATA LOG REVIEWED: NO MOTOR CURRENT (MC) SPIKE OBSERVED OUTSIDE OF P-LEVEL CHANGES. FEW RESOLVED POSITION UNKNOWN ALARMS. MANY SUCTIONS ALARMS OCCURRED DURING THE CASE. THE CAUSE OF THE LOW FLOW WAS MOST LIKELY IMPROPER POSITIONING OF THE DEVICE AS SUCTION ALARMS WERE REPORTED AND THE PUMP WAS STATED TO BE SHALLOW. STROKE/CVA: AS NECESSARY CLINICAL INFORMATION WAS NOT PROVIDED, THE TRUE ROOT CAUSE OF THE STROKE/CVA COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21993078&pc=OZD | 1220648-2025-28247 | 2025-04-18 04:00 | Death | ABIOMED, INC. | 2025-05-09 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A 62 YEAR OLD MALE PATIENT WHO HAD A CP PUMP PLACEMENT FOR SUPPORT DUE TO MYOCARDIAL INFARCTION. THE PATIENT HAD HEMOLYSIS OBSERVED AFTER INSERTION OF THE FOLEY CATHETER. AFTER OVER 16 HOURS OF SUPPORT THE PATIENT WAS MADE A DO NOT RESUSCITATION (DNR) AND EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21993491&pc=FPO | 0001831750-2025-00695 | 2024-03-03 05:00 | Death | STRYKER MEDICAL-KALAMAZOO | 2025-05-09 04:00 | FPO | UNKNOWN_MEDICAL_PRODUCT | Device Dislodged or Dislocated | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS BEING TRANSPORTED IN AN AMBULANCE WHEN IT WAS INVOLVED IN A CRASH. THE DEVICE BROKE FREE, CONTRIBUTING TO THE INJURIES AND EVENTUAL DEATH OF THE PATIENT.Event Description: NO NEW INFORMATION. Manufacturer Narrative: IT WAS REPORTED BY THE CUSTOMER THAT AN UNSPECIFIED AMBULANCE FASTENER WAS IN AN AMBULANCE ACCIDENT. THE COT AND LOCK DOWN SYSTEM BROKE FREE, CONTRIBUTING TO THE INJURIES AND THE PATIENT EXPIRED. FURTHER INFORMATION SURROUNDING THE INCIDENT, DETAILS OF THE MALFUNCTION, AND DEVICE INVOLVED WERE NOT AVAILABLE AS THE ALLEGED ISSUE HAS ON-GOING LEGAL MATTERS WITH THE CUSTOMER. THE ISSUE WAS RESOLVED FOR THE CUSTOMER BY DOCUMENTING THIS REPORT. IF FURTHER INFORMATION BECOMES AVAILABLE THROUGH CORPORATE LITIGATION, THIS RECORD MAY BE REOPENED AND UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21993679&pc=GKT | 3004548776-2025-00256 | 2025-03-30 04:00 | Death | FENWAL INTERNATIONAL INC. | 2025-05-09 04:00 | GKT | PLASMACELL XI DISPOSABLE SET WITH SPIKESMART | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | BK160028 |  | 1 | Event Description: THE DONOR WAS A 22-YEAR-OLD MALE WHO COMPLETED AN UNCOMPLICATED DONATION ON (B)(6) 2025. THE CENTER WAS NOTIFIED ON 04/09/2025 BY THE DONOR'S HEALTHCARE PROVIDER THAT THE DONOR HAD DIED ON (B)(6) 2025. THE DONOR HAD DONATED TWO TIMES AT THE DONOR CENTER SINCE HIS FIRST DONATION ON (B)(6) 2025. REVIEW OF THE ELECTRONIC RECORDS FROM THE DONOR'S TWO DONATIONS, INCLUDING THE DONOR'S NEW APPLICANT DONOR ELIGIBILITY ASSESSMENT FROM (B)(6) 2025, REVEALED THAT THE ONLY MEDICAL CONDITION REPORTED BY THE DONOR WAS ANXIETY FOR WHICH HE HAD BEEN PRESCRIBED CLONAZEPAM TO USE AS NEEDED. HE REPORTED NO OTHER MEDICAL CONDITIONS, MEDICATIONS, OR ALLERGIES. THE HANDS-ON PHYSICAL EXAM FROM THIS DATE WAS UNREMARKABLE OTHER THAN AN ABDOMINAL SCAR FROM A REPORTED APPENDIX REMOVAL IN 2015. THE DONOR COMPLETED SUCCESSFUL DONATIONS (B)(6) 2025 AND (B)(6) 2025 WITH NO ASSOCIATED DONOR ADVERSE EVENTS OR RBC LOSSES. THE DONOR ATTEMPTED A DONATION ON (B)(6) 2025 BUT WAS DEFERRED DUE TO AN UNACCEPTED SCREENING PULSE READING OF 101. A SERUM PROTEIN ELECTROPHORESIS COLLECTED ON (B)(6) 2025 WAS ACCEPTABLE. THE DONATION ON (B)(6) 2025 BEGAN AT 11:31 AM FOLLOWING A SUCCESSFUL VENIPUNCTURE IN THE RIGHT ARM. THE TOTAL COLLECTION VOLUME WAS 898 ML AND THE PLASMA COLLECTION VOLUME WAS 802 ML. THE DONOR RECEIVED 500 ML OF PROCEDURAL SALINE AT THE END OF THE PROCEDURE. NO APHERESIS COMPLICATIONS OR DONOR ADVERSE EVENTS WERE NOTED DURING OR AFTER THE DONATION. THE DONOR CENTER LEARNED OF THE DONOR'S DEATH ON (B)(6) 2025 WHEN THEY RECEIVED A CALL FROM A NURSE NAMED AT A MEDICAL OFFICE IN (B)(6), IN REPORTING THAT THE DONOR HAD DIED ON (B)(6) 2025. THE NURSE INDICATED THAT DONOR WAS FOUND AT HIS RESIDENCE SITTING IN A CHAIR, DECEASED AND THAT AN AUTOPSY WAS BEING PERFORMED. THE (B)(6) IN CORONER'S OFFICE WAS CONTACTED AND THEY CONFIRMED THEIR OFFICE IS INVESTIGATING THE CASE. THE REPRESENTATIVE CONFIRMED THE DATE OF DEATH AS (B)(6) 2025 AND THE TIME OF DEATH AS 2:24 PM. AN AUTOPSY HAS BEEN PERFORMED AND TOXICOLOGY SAMPLES HAVE BEEN COLLECTED. HOWEVER, A FINAL CAUSE AND MANNER OF DEATH ARE PENDING THE RESULTS. THE REPRESENTATIVE INDICATED THAT THE DONOR WAS FOUND DECEASED AT HIS RESIDENCE WITH NO EVIDENCE OF FOUL PLAY OR DRUG PARAPHERNALIA AT THE SCENE. THE ONLY KNOWN MEDICAL HISTORY WAS A DIAGNOSIS OF ANXIETY FOR WHICH THE DONOR WAS PRESCRIBED CLONAZEPAM AS NEEDED. THE BATCH RECORD FOR THE AURORA XI DISPOSABLE KIT LOT FA24K14033 WAS REVIEWED. NO EXCEPTIONS WERE GENERATED THAT COULD CLASSIFY AS A POSSIBLE ROOT CAUSE OF THIS DEFECT. THE FINISHED GOOD LOT HAS PASSED ALL SAMPLING ACCEPTANCE CRITERIA FOR ALL TESTS PERFORMED INCLUDING IN-PROCESS TESTING AND PRODUCT TESTING. FURTHER INVESTIGATION IS PENDING. GIVEN THAT THE DONOR DIED APPROXIMATELY 27 HOURS AFTER THE FINAL UNCOMPLICATED DONATION AND THE CAUSE OF THE DEATH IS CURRENTLY UNKNOWN, WE ARE REPORTING THIS INCIDENT CONSERVATIVELY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21993711&pc=DWF | 9612164-2025-02312 | 2025-03-31 04:00 | Death | MEDTRONIC MEXICO S. DE R.L. DE CV | 2025-05-09 04:00 | DWF | CUSTOM TUBING PACK | Material Separation | Cardiac Arrest; Congenital Defect/Deformity | K171979 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING USE OF A CUSTOM TUBING PACK, IT WAS REPORTED THAT THE STOPCOCK ON THE VENOUS SIDE OF THE BRIDGE BECAME LOOSE. THE LOOSE COMPONENT DID NOT FALL INTO THE PATIENT. THE PATIENT CODED WHEN THE STOPCOCK, INCLUDING THE ATTACHED TUBING, DISCONNECTED. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED AND THE PATIENT WAS REVIVED. THEY REPLACED THE EXTRA CORPOREAL MEMBRANE OXYGENATION (ECMO) CIRCUIT. ACCORDING TO THE CLINICIAN, A SPECIALIST'S HOSPITAL BADGE CAUGHT ONTO THE STOPCOCK ON THE VENOUS SIDE OF THE BRIDGE AND IT POPPED OFF AS A RESULT. THE TUBING WAS ATTACHED WITH IT. THE PATIENT WAS VERY SICK FROM THE OUTSET. THE PATIENT IS CURRENTLY STILL ON ECMO AND NOT DOING WELL. THE PATIENT EXPERIENCED A CONGENITAL ANOMALY. Manufacturer Narrative: ADDITIONAL INFORMATION B5: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE PATIENT PASSED AWAY. CORRECTION D2 (PRODUCT CODE), D2.1 (COMMON DEVICE NAME), D3 (MFR NAME) & D3.6 (POSTAL CODE): THESE FIELDS HAVE BEEN UPDATED. CORRECTION ADDITIONAL CODES IMG (ANNEX G/COMPONENT): THIS CODE HAS BEEN UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B.5. MEDTRONIC RECEIVED INFORMATION THAT DURING USE OF A CUSTOM TUBING PACK, DURING A PROCEDURE ON A PATIENT WITH A CONGENITAL ANOMALY, IT WAS REPORTED THAT THE STOPCOCK ON THE VENOUS SIDE OF THE BRIDGE BECAME LOOSE. THE LOOSE COMPONENT DID NOT FALL INTO THE PATIENT. THE PATIENT CODED WHEN THE STOPCOCK, INCLUDING THE ATTACHED TUBING, DISCONNECTED. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED AND THE PATIENT WAS REVIVED. THEY REPLACED THE EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) CIRCUIT. ACCORDING TO THE CLINICIAN, A SPECIALIST¿S HOSPITAL BADGE CAUGHT ONTO THE STOPCOCK ON THE VENOUS SIDE OF THE BRIDGE AND IT POPPED OFF AS A RESULT AND THE TUBING WAS ATTACHED WITH IT. THE PATIENT WAS VERY SICK FROM THE OUTSET. THE PATIENT IS CURRENTLY STILL ON ECMO AND NOT DOING WELL. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THERE IS NO ALLEGATION THAT THE TUBING PACK CONTRIBUTED TO PATIENT'S DEATH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21993744&pc=OZD | 1220648-2025-28255 | 2025-05-02 04:00 | Death | ABIOMED, INC. | 2025-05-09 04:00 | OZD | IMPELLA CP | Deformation Due to Compressive Stress | Hemorrhage/Bleeding; Great Vessel Perforation; Aortic Valve Insufficiency/ Regurgitation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA CP AND DURING THE IMPLANT, THE PLACEMENT WIRE COULD NOT BE REMOVED AND FORMED A LOOP IN THE DESCENDING AORTA. THE WIRE WAS REMOVED PULLING IT FROM THE RADIAL ARTERY USING A SNARE CATHETER. HOWEVER, THE BLOOD VESSEL WAS DAMAGED BY THE 18-INCH WIRE, AND THE PATIENT EXPERIENCED INTRA-ABDOMINAL BLEEDING, SO A STENT GRAFT WAS PLACED. ADDITIONALLY, THE PATIENT HAD SEVERE AORTIC REGURGITATION (AR) AND HEMODYNAMIC COLLAPSE OCCURRED. THE PUMP WAS REPOSITIONED; HOWEVER, THIS DID NOT IMPROVE THE AR. THE PATIENT'S OUTCOME IS UNKNOWN.Event Description: ADDITIONALLY, THE PATIENT EXPIRED ON SUPPORT. THE CAUSE OF DEATH WAS DETERMINED TO BE CIRCULATORY FAILURE DUE TO ABDOMINAL AORTIC INJURY AND ABDOMINAL HEMORRHAGE. HOWEVER, IT IS UNKNOWN IF THE IMPELLA WAS RELATED TO THE CAUSE OF DEATH. Manufacturer Narrative: B2: DEATH AND DATE OF DEATH ADDED BASED ON ADDITIONAL INFORMATION RECEIVED REGARDING THE PATIENT OUTCOME. B5: UPDATED WITH ADDITIONAL INFORMATION REGARDING THE PATIENT OUTCOME. H1: TYPE OF REPORT REVISED BASED ON ADDITIONAL INFORMATION RECEIVED REGARDING THE PATIENT OUTCOME. H6: HEALTH EFFECT - IMPACT CODE 1802 DEATH ADDED. HEART VALVE DAMAGE: AFTER THE INTRODUCTION OF IMPELLA, SEVERE AORTIC REGURGITATION OCCURRED AND THE PATIENT'S HEMODYNAMICS COLLAPSED. THE DEVICE WAS SWITCHED TO VA-ECMO AND REPOSITIONED, BUT THE PATIENT'S CONDITION DID NOT IMPROVE. CP EXPLANTED AND AR WAS ALLEVIATED. THE PUMP WAS NOT RETURNED. THE CAUSE OF THE AORTIC REGURGITATION WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS. PRODUCT DAMAGE (GUIDEWIRE DAMAGE - GREAT VESSEL): FOLLOWING INITIAL PLACEMENT OF THE PUMP, PLACEMENT WIRE COULD NOT BE REMOVED AND FORMED A LOOP IN THE DESCENDING AORTA. THE 18-INCH WIRE WAS REMOVED BY PULLING IT FROM THE RADIAL ARTERY USING A SNARE CATHETER. HOWEVER, THE BLOOD VESSEL WAS DAMAGED BY THE 18-INCH WIRE, AND THE PATIENT EXPERIENCED INTRAABDOMINAL BLEEDING, SO A STENT GRAFT WAS PERFORMED. NO PRODUCT WAS RETURNED. THE CAUSE OF THE PRODUCT DAMAGE (GUIDEWIRE DAMAGE - GREAT VESSEL) WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION. SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ SECTION: IMPELLA CP WITH SMARTASSIST CATHETER INSERTION ¿IMPELLA CATHETER PERFORMANCE WILL BE COMPROMISED IF CORRECT PLACEMENT CANNOT BE CONFIRMED. WHILE OTHER IMAGING TECHNIQUES, SUCH AS TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE), PORTABLE C-ARM FLUOROSCOPY, OR CHEST X-RAY CAN HELP CONFIRM THE POSITION OF THE IMPELLA CATHETER AFTER PLACEMENT, THESE METHODS DO NOT ALLOW VISUALIZATION OF THE ENTIRE CATHETER ASSEMBLY AND ARE INADEQUATE FOR RELIABLY PLACING THE IMPELLA CATHETER ACROSS THE AORTIC VALVE¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER. IMPELLA CATHETER INLET TO THE AORTA, ¿IF THE INLET AREA OF THE IMPELLA CATHETER IS TOO CLOSE TO OR ENTANGLED IN THE PAPILLARY MUSCLE AND/OR SUBANNULAR STRUCTURES SURROUNDING THE MITRAL VALVE, IT CAN AFFECT MITRAL VALVE FUNCTION AND NEGATIVELY IMPACT CATHETER FLOW.¿ SECTION: UNDERSTANDING AND MANAGING IMPELLA CATHETER POSITION ALARMS, ¿IF THE IMPELLA CATHETER IS COMPLETELY OUT OF THE VENTRICLE, DO NOT ATTEMPT TO REPOSITION THE CATHETER ACROSS THE VALVE WITHOUT A GUIDEWIRE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21994087&pc=MSX | 3009156722-2025-00006 | 2025-04-12 04:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-05-09 04:00 | MSX | BENEVISION DISTRIBUTED MONITORING SYSTEM (DMS) | No Device Output | Insufficient Information | K220058 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, THE DATA FROM A PATIENT CONNECTED TO A TD60 TELEMETRY UNIT (SERIAL NUMBER (B)(6)) ID (B)(6) WAS LOST. THERE WAS NO POWER OUTAGE DURING THIS TIME. THE CUSTOMER IS REQUESTING TO EVALUATE IF THE DATA WAS DELETED OR IF THE PATIENT WAS NOT ADMITTED/DISCHARGED PROPERLY. THE PATIENT EXPIRED.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, THE DATA FROM A PATIENT CONNECTED TO A TD60 TELEMETRY UNIT (SERIAL NUMBER (B)(6)) ID 5N TELE 2 IN ROOM 502 WAS LOST. THERE WAS NO POWER OUTAGE DURING THIS TIME. THE CUSTOMER IS REQUESTING TO EVALUATE IF THE DATA WAS DELETED OR IF THE PATIENT WAS NOT ADMITTED/DISCHARGED PROPERLY. THE PATIENT EXPIRED. Manufacturer Narrative: MINDRAY CONDUCTED A REVIEW OF THE BENEVISION CMS LOGS FOR PATIENTS ADMITTED BETWEEN (B)(6) 2025 AND DISCHARGED ON (B)(6) 2025. THE REVIEW FOCUSED ON PATIENT INFORMATION DELETION RECORDS, BED NUMBERS, AND EQUIPMENT SERIAL NUMBERS, AND DID NOT IDENTIFY ANY INFORMATION RELATED TO THE REPORTED OCCURRENCE INVOLVING THE PATIENT WITH ID 5N TELE 2 IN ROOM 502. IT WAS DETERMINED THAT THE TD60 TELEMETRY DEVICE WAS LIKELY NOT CONNECTED TO THE CENTRAL STATION. AS NO LOGS FOR THE TD60 WERE AVAILABLE AFTER MULTIPLE ATTEMPTS, THE REPORTED ISSUE COULD NOT BE CONFIRMED. MINDRAY BENEVISION CMS WITH TD60 TELEMETRY FUNCTIONED ACCORDING TO SPECIFICATIONS. Manufacturer Narrative: SYSTEM LOGS WERE COLLECTED BY MINDRAY FIELD SER VICE REPRESENTATIVE. UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21994418&pc=NKM | 2135147-2025-02440 | 2012-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-05-09 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock | P100009 |  | 1 | Event Description: THE ARTICLE "IMPACT OF MITRAL VALVE TRANSCATHETER EDGE-TO-EDGE REPAIR ON HAEMODYNAMIC PARAMETERS IN CARDIOGENIC SHOCK" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE HEMODYNAMIC AND ECHOCARDIOGRAPHIC PARAMETERS IN PATIENTS WITH CARDIOGENIC SHOCK UNDERGOING TRANSCATHETER EDGE-TO-EDGE REPAIR (TEER). DEVICES MENTIONED INCLUDE MITRACLIP AND EDWARDS PASCAL. THE ARTICLE CONCLUDED THAT TEER LEADS TO FAVORABLE HEMODYNAMIC CHANGES IN PATIENTS WITH CS.. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MICHAL DROPPA, DEPARTMENT OF CARDIOLOGY AND ANGIOLOGY, UNIVERSITY HOSPITAL OF TÜBINGEN, TÜBINGEN, GERMANY WITH CORRESPONDING EMAIL *MICHAL.DROPPA@MED.UNI-TUEBINGEN.DE*. ] THE TIME FRAME OF THE STUDY WAS JANUARY 2012 TO NOVEMBER 2023. A TOTAL OF 25 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 22 RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 65 AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED DIABETES, HYPERTENSION, DYSLIPIDEMIA, ATRIAL FIBRILLATION, CORONARY ARTERY DISEASE, PRIOR MYOCARDIAL INFARCTION, PRIOR PERCUTANEOUS INTERVENTION, PRIOR CORONARY ARTERY BYPASS GRAFT, SECONDARY MITRAL REGURGITATION, MIXED MITRAL REGURGITATION, ISCHEMIC CARDIOMYOPATHY, NON-ISCHEMIC CARDIOMYOPATHY, VASOPRESSORS, INOTROPICS, MECHANICAL CARDIAC SUPPORT, HEART FAILURE. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, CARDIOGENIC SHOCK. Manufacturer Narrative: THE DEVICE WILL NOT BE RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP PROCEDURE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, HYPERTENSION, DYSLIPIDEMIA, ATRIAL FIBRILLATION, CORONARY ARTERY DISEASE, PRIOR MYOCARDIAL INFARCTION, PRIOR PERCUTANEOUS INTERVENTION, PRIOR CORONARY ARTERY BYPASS GRAFT, SECONDARY MITRAL REGURGITATION, MIXED MITRAL REGURGITATION, ISCHEMIC CARDIOMYOPATHY, NON-ISCHEMIC CARDIOMYOPATHY, VASOPRESSORS, INOTROPICS, MECHANICAL CARDIAC SUPPORT, HEART FAILURE. COMPLICATIONS REPORTED INCLUDED CARDIOGENIC SHOCK AND DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21994511&pc=PYX | 1220648-2025-28263 | 2025-04-22 04:00 | Death | ABIOMED, INC. | 2025-05-09 04:00 | PYX | IMPELLA RP FLEX | Restricted Flow rate; Pumping Stopped | Embolism/Embolus; Insufficient Information | P170011 |  | 1 | Event Description: IT WAS REPORTED THAT AN IMPLANTED IMPELLA RP FLEX PUMP EXPERIENCED A TEMPORARY STOP. UPON RESTART, THE PUMP PERFORMED AT LOWER FLOWS. IT WAS RECOMMENDED TO REPLACE THE PUMP, HOWEVER, THE DOCTOR CONTINUED SUPPORT. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION OF TEMPORARY PUMP STOP AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. TEMPORARY PUMP STOP: DATA LOG REVIEW SHOWED SOME MOTOR CURRENT SPIKES IN THE TIME LEADING UP TO THE PUMP STOP. AFTER THE PUMP WAS RESTARTED, THERE WERE LOWER PUMP FLOWS AND MOTOR CURRENT DISTURBANCES. PURGE PRESSURE ALSO BEGAN TO RISE A FEW MINUTES BEFORE THE PUMP STOP. THE ROOT CAUSE OF THE TEMPORARY PUMP STOP WAS MOST LIKELY BIOMATERIAL INGESTION SINCE THERE WERE MOTOR CURRENT SPIKES LEADING UP TO THE PUMP STOP WITH MOTOR SPEED DEVIATIONS AND REDUCED FLOWS AFTERWARD. LOW PUMP FLOW: AS MENTIONED ABOVE, DATA LOG REVIEW SHOWED THE PUMP HAD LOWER FLOWS AFTER BEING RESTARTED. THERE WERE MOTOR CURRENT SPIKES LEADING UP TO THE PUMP STOP AND A SLIGHT RISE IN PURGE PRESSURE. PURGE FLUID WAS CHANGED FROM BICARB TO HEPARIN A FEW HOURS AFTER THE PUMP STOP. THE LOW PUMP FLOW ABRUPTLY RESOLVED ABOUT 6 HOURS AFTER THE PUMP STOP. THE ROOT CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY BIOMATERIAL INGESTION SINCE THERE WERE MOTOR CURRENT SPIKES AND A PUMP STOP PRIOR TO THE DROP IN FLOWS. THROMBUS: FAILURE MODE ADDED BASED ON INVESTIGATION RESULTS. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE THROMBUS WAS NOT DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21994817&pc=CBK | 2518422-2025-105412 | 2024-12-16 05:00 | Death | RESPIRONICS, INC. | 2025-05-09 04:00 | CBK | TRILOGY EVO | Adverse Event Without Identified Device or Use Problem | Unspecified Respiratory Problem | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING DISTRIBUTOR WOULD LIKE THIS UNIT RETURNED FOR AN EVALUATION TO CONFIRM THAT THERE WAS NO ISSUE WITH THIS UNIT PRIOR TO SETTING IT UP ON ANOTHER PATIENT, AS THIS PATIENT PASSED AWAY. THERE WAS NO SPECIFIC ISSUE REPORTED WITH THE UNIT. DISTRIBUTOR REP STATED THAT A WELFARE CHECK WAS DONE ON PATIENT, AND THEY WERE FOUND TO HAVE ALREADY PASSED AWAY. THAT IS ALL THAT IS KNOWN ABOUT THE INCIDENT. INFORMATION THAT HAS BEEN REPORTED BY (B)(6), RESPIRATORY THERAPIST FOR INTEGRIS HEALTH: WHAT IS THE DATE AND TIME OF THE OCCURRENCE? - UNKNOWN FOR SURE, BELIEVES (B)(6) 2024. IS THERE AN ALLEGATION OF DEVICE MALFUNCTION? - NO, NOT TO HIS KNOWLEDGE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION ALLEGING DISTRIBUTOR WOULD LIKE THIS UNIT RETURNED FOR AN EVALUATION TO CONFIRM THAT THERE WAS NO ISSUE WITH THIS UNIT PRIOR TO SETTING IT UP ON ANOTHER PATIENT, AS THIS PATIENT PASSED AWAY. THERE WAS NO SPECIFIC ISSUE REPORTED WITH THE UNIT. DISTRIBUTOR REP STATED THAT A WELFARE CHECK WAS DONE ON PATIENT, AND THEY WERE FOUND TO HAVE ALREADY PASSED AWAY. THAT IS ALL THAT IS KNOWN ABOUT THE INCIDENT. INFORMATION THAT HAS BEEN REPORTED BY BILL BRAZIL, RESPIRATORY THERAPIST FOR INTEGRIS HEALTH: WHAT IS THE DATE AND TIME OF THE OCCURRENCE? - UNKNOWN FOR SURE, BELIEVES 16-DEC-2024. IS THERE AN ALLEGATION OF DEVICE MALFUNCTION? - NO, NOT TO HIS KNOWLEDGE. THE DEVICE WAS RECEIVED BY THE MANUFACTURER'S PRODUCT INVESTIGATION LAB (PIL) FOR EVALUATION. PIL RETRIEVED AND REVIEWED THE EVENT LOG FROM THE DEVICE AND NOTED THE SOFTWARE VERSION WAS 1.05.10. PIL NOTED THAT, PER THE LOG, THERE WAS NO ACTIVITY LOGGED AFTER 12/08/2024 UNTIL 05/07/2025. PIL POWERED ON THE DEVICE AND STARTED THERAPY WITH A TEST LUNG, UTILIZING THE EXISTING DEVICE SETTINGS. PIL RAN THERAPY FOR APPROXIMATELY 5 HOURS AND THE DEVICE OPERATED WITHOUT ISSUE. PIL THEN POST TESTED THE DEVICE ON THE PIL TRILOGY EVO MULTIFUNCTIONAL TEST STATION AND THE DEVICE PASSED ALL TESTING. PIL WAS UNABLE TO CONFIRM A FAILURE OF THE TRILOGY EVO, USA. PIL CONCLUDED THAT THE TRILOGY EVO, USA OPERATES AS DESIGNED. CODING HAS BEEN UPDATED ACCORDINGLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21994998&pc=DRM | 3005445717-2025-00004 | 2025-04-11 04:00 | Death | JOLIFE AB - 3005445717 | 2025-05-09 04:00 | DRM | LUCAS 3 CHEST COMPRESSION SYSTEM | Adverse Event Without Identified Device or Use Problem | Liver Laceration(s) | K173553 |  | 1 | Event Description: THE CUSTOMER CONTACTED STRYKER TO REPORT A PATIENT DEATH. USE OF THE DEVICE MAY HAVE CONTRIBUTED BUT IT IS UNKNOWN HOW THE INJURY CONTRIBUTED.Event Description: THE CUSTOMER CONTACTED STRYKER TO REPORT A PATIENT DEATH. USE OF THE DEVICE MAY HAVE CONTRIBUTED BUT IT IS UNKNOWN HOW THE INJURY CONTRIBUTED. Manufacturer Narrative: STRYKER CONTINUES TO INVESTIGATE THE REPORTED FAILURE AND WILL SUBMIT A SUPPLEMENTAL REPORT ON THIS EVENT TO THE FDA AS PROVIDED BY 21 CFR 803.56. Manufacturer Narrative: PARTIAL PATIENT INFORMATION WAS RECEIVED, SECTION A HAS BEEN UPDATED. STRYKER EVALUATED THE CUSTOMER'S DEVICE AND WAS NOT ABLE TO VERIFY OR DUPLICATE THE REPORTED ISSUE. AFTER OBSERVING PROPER DEVICE OPERATION THROUGH FUNCTIONAL AND PERFORMANCE TESTING, THE DEVICE WAS RETURNED TO THE CUSTOMER FOR USE. THE CAUSE OF THE REPORTED ISSUE COULD NOT BE DETERMINED. STRYKER PERFORMED A CLINICAL REVIEW OF THE REPORTED EVENT AND DETERMINED THE DEVICE CONTRIBUTION TO THE PATIENT OUTCOME IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21995053&pc=OZD | 1220648-2025-28267 | 2025-04-14 04:00 | Death | ABIOMED, INC. | 2025-05-09 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A HEMATOMA DEVELOPED AT THE IMPELLA CP ACCESS SITE FOLLOWING REMOVAL OF THE PEEL AWAY SHEATH. STEADY BLEEDING OCCURRED AT ALL ACCESS SITES WHICH THE PHYSICIAN ATTRIBUTED TO DISSEMINATED INTRAVASCULAR COAGULATION. COAGULATION FACTORS AND ONE UNIT OF BLOOD WERE GIVEN TO THE PATIENT TO TREAT THE CONDITION. SUPPORT WAS CONTINUED WITH THE IMPLANTED PUMP FOLLOWING THE INTERVENTION. HOWEVER, WHEN BLEEDING WAS CONTINUED FROM OTHER SITE, THE DECISION WAS ULTIMATELY MADE TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THE MEDICAL SAFETY OFFICER REVIEWED AND DETERMINED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME. Manufacturer Narrative: THE BLEEDING WAS NOT RESOLVED AND THE HEMOSTASIS WAS NOT ACHIEVED AS PER THE CLINICAL DETAILS. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING WAS DETERMINED TO BE PATIENT CONDITION BECAUSE OF THE BLEEDING FROM ALL ACCESS SITES. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: "ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA." SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: "REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE." SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21995186&pc=DSQ | 2916596-2025-03143 | 2025-04-10 04:00 | Death | THORATEC CORPORATION | 2025-05-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Cardiogenic Shock; Respiratory Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO BIVENTRICULAR HEART FAILURE, CARDIOGENIC SHOCK, ACUTE RESPIRATORY FAILURE, OLIGURIC ACUTE RENAL FAILURE. IT WAS REPORTED THAT THE PUMP FUNCTIONED AS INTENDED AND WAS NOT EXPLANTED. Manufacturer Narrative: SECTION A4: PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED MULTI ORGAN FAILURE AND CARDIOGENIC SHOCK AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (INCLUDING RIGHT HEART FAILURE, RESPIRATORY FAILURE, AND RENAL DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, UNDER ¿RIGHT HEART FAILURE¿, ALSO STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21995206&pc=LYJ | 1644487-2025-10105 | 2025-04-15 04:00 | Death | LIVANOVA USA, INC. | 2025-05-09 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. NO FURTHER DETAILS WERE PROVIDED. AN OBITUARY WAS FOUND WHICH NOTES THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 CFR PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21995207&pc=CBK | 2518422-2025-105421 | 2025-05-06 04:00 | Death | RESPIRONICS, INC. | 2025-05-09 04:00 | CBK | TRILOGY EVO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT DURING CLINICAL AND THERAPEUTIC USE, THE TRILOGY EVO VENTILATOR WAS ALARMING LOW INSPIRATORY PRESSURE AND LOW MINUTE VOLUME. THE PATIENT WAS REMOVED FROM THE DEVICE AND MANUAL VENTILATION INITIATED BY THE PATIENT'S FATHER. AIRWAY SUCTIONING WAS CONDUCTED WITH NO RESISTANCE FOUND AND THE PATIENT WAS NOTED TO HAVE BEEN UNRESPONSIVE DURING EMERGENCY INTERVENTIONS. EMERGENCY MEDICAL SERVICES (EMS) WAS CONTACTED AND THE PATIENT TRANSFERRED TO AN INSTITUTIONAL FACILITY WHERE THEY WERE PRONOUNCED DECEASED. THE PATIENT'S CAUSE OF DEATH WAS BELIEVED TO BE RESPIRATORY FAILURE; HOWEVER, A DEFINITE CAUSE IS PENDING A PATIENT AUTOPSY. FURTHER RELEVANT INFORMATION PROVIDED STATED THAT THE PATIENT WAS BEING MONITORED VIA A PULSE OXIMETER (MAKE/MODEL UNSPECIFIED) WHICH WAS NOT WORKING DURING THE TIME OF THE EVENT IN QUESTION. DEVICE THERAPY SETTINGS WERE REPORTED AS: MODE S/T AVAPS, IPAP MIN-MAX 24-46 CMH2O, EPAP 12 CMH2O, RATE 20 BPM, TIDAL VOLUME 80 ML, I-TIME 0.7 SECONDS, TMIN-TMAX 0.3-0.9 SECONDS, RISE 2. BASED ON THIS INFORMATION, NO CLEAR ALLEGATION OF DEVICE MALFUNCTION OR FAILURE TO PERFORM TO MANUFACTURER DECLARED SPECIFICATIONS HAS BEEN ALLEGED. CAUSE AND/OR CONTRIBUTION OF THE DEVICE TO THE PATIENT OUTCOME CANNOT BE CONFIRMED NOR REFUTED. FURTHER GOOD FAITH EFFORTS ARE REQUIRED TO DETERMINE IF A DEVICE ALLEGATION IS PRESENT AND TO PROVIDE FURTHER CLINICAL ASSESSMENT OF THE QUALITY RECORD. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21995248&pc=PYX | 1220648-2025-28268 | 2025-04-16 04:00 | Death | ABIOMED, INC. | 2025-05-09 04:00 | PYX | IMPELLA RP FLEX | Restricted Flow rate; Malposition of Device | Hemolysis | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT, POST-OPERATIVE AORTIC AND MITRAL VALVE REPLACEMENT SURGERY FIVE DAYS PRIOR AND WAS IMPLANTED WITH AN IMPELLA RP FLEX DEVICE VIA THE RIGHT FEMORAL VEIN FOR BIPELLA MECHANICAL CIRCULATORY SUPPORT. PRIOR TO BIPELLA SUPPORT, HEPARIN WAS STOPPED DUE TO A NOSEBLEED. THEREAFTER, APPROXIMATELY TWO DAYS AFTER START OF THE BIPELLA SUPPORT, THE PATIENT'S URINE WAS NOTED TO BE PORT WINE COLOR. HAPTOGLOBIN WAS LOW AND LABS WERE HEMOLYZING. THE FOLLOWING DAY, AN ASSESSMENT WAS MADE FOR FURTHER PATIENT BLEEDING (OUTCOME OF THE ASSESSMENT UNNOTED). HEMOLYSIS WAS STILL PRESENT, HOWEVER, WITH LOW HAPTOGLOBIN, ELEVATED LACTATE DEHYDROGENASE, AND RED URINE. ADDITIONALLY, THE IMPELLA RP FLEX FLOWS WERE NOT WITHIN NORMAL LIMITS FOR GIVEN THE PERFORMANCE LEVEL, AND THERE WAS AN ABRUPT CHANGE IN MOTOR CURRENT/FLOWS NOTED AT 06:00 IN THE MORNING. PREVIOUS POSTERIAL NASAL BLEEDING LED TO CLOTS, RESPIRATORY FAILURE, AND REINTUBATION, AND THE TEAM WAS, THEREFORE, HESITANT TO RESTART ANTICOAGULATION. THE NEXT DAY, THE PATIENT VOMITED AND ASPIRATED OVERNIGHT REQUIRING REINTUBATION. AN OROGASTRIC TUBE WAS PLACED AND SET TO SUCTION WITH 500 MILLILITERS OF DARK RED OUTPUT. THREE UNITS OF PACKED RED BLOOD CELLS AND TWO UNITS OF FRESH FROZEN PLASMA WERE GIVEN. THE PATIENT LOST PULSES IN RIGHT LOWER EXTREMITY, SYSTEMIC HEPARIN DRIP WAS INITIATED. THE IMPELLA RP FLEX FLOWS INCREASED FROM 1.8 LITERS PER MINUTE (L/MIN) TO 2.4 L/MIN. THE PATIENT HAD BEEN STARTED ON CONTINUOUS RENAL REPLACEMENT THERAPY. ON CHEST X-RAY, THE IMPELLA RP FLEX OUTLET APPEARED DEEP IN LEFT PULMONARY ARTERY. THE PHYSICIAN WAS NOTIFIED AND BIPELLA SUPPORT WAS CONTINUED. THE FOLLOWING DAY, THE PATIENT WAS TAKEN TO THE OPERATING ROOM FOR VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) CANNULATION AND IMPELLA RP FLEX EXPLANT. THE IMPELLA 5.5 REMAINED IMPLANTED AND IN PLACE AS VERIFIED BY ECHOCARDIOGRAM. HOWEVER, THREE DAYS LATER THE PATIENT WOULD EXPIRE. CARE WAS WITHDRAWN. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS, POSITIONING ISSUES, AND LOW PUMP FLOW HAS BEEN COMPLETED. DATA REVIEW: DATA LOGS CONFIRMED THE LOW FLOW AND CLINICAL NARRATIVE. LOGS SHOWED AFTER PUMP STARTED AND RUN FOR 49 HOURS, MOTOR CURRENT (MC) SPIKED FOLLOWED LOW FLOW THAT TRIGGERED AUTO SUCTION RESPONSE AND SUCTION ALARMS. THIS EVENT REMAINED TO THE REST OF THE CASE. NO POSITIONING ALARM WAS TRIGGERED DURING SUPPORT. PRODUCT ANALYSIS: PUMP WAS RETURNED IN DAMAGED CONDITION (CATHETER WAS CUT OFF AND& ONLY CANNULA WAS RETURNED FOR ANALYSIS). VISUAL INSPECTION FOUND BIOMATERIAL INSIDE THE PUMP HOUSING AND WRAPPED AROUND THE IMPELLER. NO OTHER ISSUES WERE FOUND ON THE REST OF THE PUMP. CONCLUSION: THE ROOT CAUSE OF HEMOLYSIS WAS BIOMATERIAL INGESTION. THE ROOT CAUSE OF THE LOW FLOW WAS BIOMATERIAL INGESTION. THE CAUSE OF THE POSITIONING ISSUE WAS NOT DETERMINED DUE TO INSUFFICIENT CLINICAL INFORMATION. Manufacturer Narrative: MEDICAL SAFETY REVIEWED THE EVENT AND NOTED THE PATIENT EXPERIENCED ADVERSE EVENTS WITH THE IMPELLA 5.5 AND IS A SERIOUS INJURY. HOWEVER, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA RP FLEX AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME (DEMISE) WITH THE REPORT OF INADEQUATE PUMP FLOW AND SUBOPTIMAL POSITIONING IN ADDITION TO AN ADVERSE EVENT. THE IMPELLA RP FLEX IS ONE OF TWO DEVICES ASSOCIATED WITH THE EVENT. NOTE THE FOLLOWING: REFER TO INITIAL MEDICAL DEVICE REPORT FOR THE IMPELLA 5.5-SERIAL NUMBER (B)(6) WITH DATE OF EVENT ON 11-APR-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). THE IMPELLA RP FLEX PUMP WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21995315&pc=NKM | 2135147-2025-02453 | 2015-11-01 04:00 | Death | ABBOTT MEDICAL | 2025-05-09 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "VALIDATION OF THE COAPT RISK SCORE IN POLISH PATIENTS UNDERGOING TRANSCATHETER EDGE-TOEDGE REPAIR OF SEVERE, FUNCTIONAL MITRAL REGURGITATION: A MULTICENTER, OBSERVATIONAL STUDY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE A COMPOSITE OF ALL-CAUSE MORTALITY AND HFH AT TWO-YEAR FOLLOW-UP AND EVALUATED IN THE OVERALL COHORT AND SEPARATELY FOR COAPT-ELIGIBLE AND -NON-ELIGIBLE PATIENTS. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT IN POLISH PATIENTS FULFILLING COAPT CRITERIA, THE COAPT RISK SCORE HAS MODERATE PREDICTIVE VALUE FOR POST-PROCEDURAL OUTCOMES. IN COAPT NON-ELIGIBLE PATIENTS, NOVEL TOOLS ARE REQUIRED TO PREDICT OUTCOMES. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS KAROLINA JASINSKA-GNIADZIK AT MEDICAL UNIVERSITY OF WARSAW INSTITUTION WITH CORRESPONDING EMAIL KAROLINA.JASINSKA-GNIADZIK@WUM.EDU.PL]. THE TIME FRAME OF THE STUDY WAS NOVEMBER 2015 TO FEBRUARY 2023. A TOTAL OF 225 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 225 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPOTENSION, DIABETES, PRIOR ISCHEMIC STROKE. (B)(4), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPOTENSION, DIABETES, PRIOR ISCHEMIC STROKE. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21995482&pc=NPT | 2015691-2025-03703 | 2025-04-17 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-09 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Difficult or Delayed Positioning | Foreign Body In Patient | P140031 |  | 1 | Event Description: AS REPORTED BY A FIELD CLINICAL SPECIALIST (FCS), DURING A TAVR PROCEDURE WITH A 23MM SAPIEN 3 ULTRA RESILIA (S3UR) VALVE IN A PRE-EXISTING SURGICAL VALVE THE IMPLANTER DECIDED TO DO A UNICORN PROCEDURE ON THE LEFT CUSP LEAFLET. THE LEFT CORONARY LEAFLET OF THE SURGICAL VALVE WAS PUNCTURED SUCCESSFULLY WITH AN RFA WIRE, AND THE LEAFLET WAS BALLOONED WITH 3MM AND 12MM NON-EDWARDS BALLOONS. A 23 MM S3UR WAS PREPPED WITH +2 CC VOLUME AND WAS THEN ADVANCED TO THE SURGICAL VALVE. AT THIS POINT, THE VALVE WOULD NOT ADVANCE THROUGH THE PERFORATED LEAFLET, IMPLANTER TRIED VARIOUS TECHNIQUES, USING HIGH FLEX AND NO FLEX, ADDING 2CC TO THE COMMANDER BALLOON BUT WAS UNABLE TO CROSS. THE PATIENT BECAME HYPOTENSIVE DUE TO THE AI FROM THE PERFORATED LEAFLET. A DECISION WAS THEN MADE TO PULL BACK THE SAPIEN VALVE AND DEPLOY IN THE THORACIC AORTA. A NEW 23 MM S3UR WAS THEN PREPPED, ADVANCED AND DEPLOYED INSIDE SURGICAL VALVE IN A NORMAL FASHION, AND A ROOT ANGIOGRAM WAS TAKEN TO CONFIRM CORONARY FLOW. THE PATIENT WAS STABLE AND TRANSFERRED TO ICU AFTER THE PROCEDURE. THERE WAS NO DAMAGE NOTED TO ANY EDWARDS DEVICES. PER ADDITIONAL INFORMATION FROM THE FCS, "THIS PATIENT PASSED TWO DAYS AFTER TAVR. PATIENT WAS INTUBATED IN THE ICU AND RECOVERING FROM THE PROCEDURE, BUT THE FAMILY AND PATIENT DECIDED TO WITHDRAW TREATMENT, ALSO, THERE WAS A DISSECTION NOTED IN THE ASCENDING AORTA." THE DATE AND CAUSE OF THE AORTIC DISSECTION AND DEATH ARE UNKNOWN AT THIS TIME. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ADDITIONAL INFORMATION PROVIDED. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED. B.2 H.1 AND H.6 IMPACT CODE. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR ADDITIONAL INFORMATION FROM A PRODUCT INVESTIGATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: B4, G3, G6, H2, AND H6. THE COMPLAINT FOR DIFFICULTY OR INABILITY TO CROSS NATIVE ANNULUS AND/OR BIOPROSTHESIS AND SUBSEQUENT AORTIC DISSECTION WAS UNABLE TO BE CONFIRMED AS NO DEVICE/ RELEVANT IMAGERY WAS PROVIDED FOR EVALUATION. THE REPORTED EVENT DOES NOT ALLEGE A MALFUNCTION THAT COULD BE RELATED TO AN EDWARDS MANUFACTURING DEFICIENCY AND/OR ONE WAS NOT CONFIRMED THROUGH INVESTIGATION. A REVIEW OF THE IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. FURTHERMORE, NO ABNORMALITIES WERE REPORTED DURING DEVICE UNPACKING OR PREPARATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. PER REVIEW, NO EVIDENCE OF PRODUCT NON-CONFORMANCES OR LABELING/IFU INADEQUACIES WERE IDENTIFIED IN THE EVALUATION. AS REPORTED, ''DURING A TAVR PROCEDURE WITH A 23MM SAPIEN 3 ULTRA RESILIA (S3UR) VALVE IN A PRE-EXISTING SURGICAL VALVE THE IMPLANTER DECIDED TO DO A UNICORN PROCEDURE ON THE LEFT CUSP LEAFLET. THE LEFT CORONARY LEAFLET OF THE SURGICAL VALVE WAS PUNCTURED SUCCESSFULLY WITH AN RFA WIRE, THE LEAFLET WAS BALLOONED WITH 3 MM AND 12 MM NON-EDWARDS BALLOONS. A 23 MM S3UR PREPPED WITH +2 CC VOLUME AND WAS THEN ADVANCED TO THE SURGICAL VALVE. AT THIS POINT, THE VALVE WOULD NOT ADVANCE THROUGH THE PERFORATED LEAFLET, IMPLANTER TRIED VARIOUS TECHNIQUES, USING HIGH FLEX AND NO FLEX, ADDING 2CC TO THE COMMANDER BALLOON BUT WAS UNABLE TO CROSS.'' IN THIS CASE, IT SHOULD BE NOTED THAT THIS WAS AN OFF-LABEL PROCEDURE. IT IS POSSIBLE THAT THE CRIMPED VALVE APICES INTERACTED WITH THE PUNCTURED LEAFLET, INHIBITING THE DELIVERY SYSTEM FROM CROSSING. WHILE A DEFINITIVE ROOT CAUSE IS UNKNOWN, AVAILABLE INFORMATION SUGGESTS THAT ABNORMAL USE (OFF LABEL PROCEDURE) MAY HAVE CONTRIBUTED TO THE DIFFICULTY CROSSING. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ASCENDING AORTIC DISSECTION OR RUPTURE MAY OCCUR WHEN MULTIPLE ATTEMPTS ARE MADE TO CROSS THE STENOTIC NATIVE VALVE, AND/OR WHEN EXCESSIVE FORCE IS USED. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). THE THV TRAINING MANUALS GUIDE TO FACILITATE SAFE CROSSING OF THE NATIVE OR FAILING VALVE, INCLUDING CAMERA PROJECTIONS, HANDLING DURING ADVANCEMENT, AND TROUBLESHOOTING TECHNIQUES IF DIFFICULTY IS ENCOUNTERED. AS STATED, EXCESSIVE FORCE SHOULD NOT BE USED WHEN THE DEVICE HAS DIFFICULTY CROSSING THE STENOTIC VALVE. ADDING TENSION TO THE WIRE, PULLING BACK THE SYSTEM TO RE-ORIENT THE VALVE, AS NEEDED, AND TORQUING THE FLEX CATHETER MAY HELP SOLVE THE PROBLEM. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PROCEDURAL FACTORS (FORCED MANEUVERING OF THE ORIGINAL COMMANDER IN ATTEMPT TO THE CROSS) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21995742&pc=LYJ | 1644487-2025-10108 | 2025-03-12 04:00 | Death | HOUSTON | 2025-05-09 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | Convulsion, Clonic | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO SUDEP. THE PATIENT HAD A VERY GOOD RESPONSE TO THE VNS AND THE PHYSICIAN DID NOT BELIEVE THE VNS WAS RELATED TO THE PATIENT'S PASSING. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 CFR PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21995755&pc=DQY | 1220452-2025-00007 | 2024-10-31 04:00 | Death | MEDTRONIC, INC | 2025-05-09 04:00 | DQY | LAUNCHER UNK GUIDE CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade; Cardiogenic Shock; Asystole | K132673 |  | 1 | Event Description: A CASE STUDY WAS SUBMITTED FOR REVIEW TITLED "THREE FATAL CONSEQUENCES OF ONE GUIDING CATHETER MANEUVER DURING PRIMARY PERCUTANEOUS ANGIOPLASTY IN ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION." IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO AN EMERGENCY DEPARTMENT DUE TO SEVERE CHEST PAIN. CORONARY ANGIOGRAPHY WAS PERFORMED, WHICH SHOWED NO SIGNIFICANT LESIONS IN THE LEFT CORONARY ARTERY AND A SUBTOTAL (95%) STENOSIS IN THE MIDDLE SEGMENT OF THE RIGHT CORONARY ARTERY (RCA) WITH (TIMI) 3 FLOW. DUE TO A TORTUOUS COURSE OF THE RCA, A MEDTRONIC 6F AMPLATZ AL 1.0 GUIDING CATHETER WAS USED TO PROVIDE SUFFICIENT SUPPORT. AFTER DELIVERING THE CATHETER TO THE RCA, INITIAL CONTRAST ADMINISTRATION EXPOSED A SEVERE, EXTENSIVE DISSECTION AND TOTAL OCCLUSION OF THE RCA WITH TIMI 0 FLOW. ATTEMPTS WERE MADE TO USE NON-MEDTRONIC GUIDEWIRES TO PASS THE DISSECTION, BUT IT COULD NOT BE PASSED. THE PATIENT DEVELOPED CARDIOGENIC SHOCK CAUSED BY THE AORTIC DISSECTION AS WELL AS CARDIAC TAMPONADE CONFIRMED ON ANGIOGRAPHY. PERICARDIOCENTESIS AND DRAINAGE OF 100 ML OF BLOOD RESTORED HEMODYNAMIC STABILITY. PERCUTANEOUS CORONARY INTERVENTION WAS STOPPED, AND THE PATIENT WAS TRANSFERRED TO AN INTENSIVE CARE UNIT. PATIENT'S CONDITION GRADUALLY WORSENED DESPITE VASOPRESSOR INFUSION. DUE TO HEMODYNAMIC INSTABILITY, A TRANSFER TO A CARDIAC SURGERY CENTER WAS IMPOSSIBLE, AND THE PATIENT WAS DISQUALIFIED AFTER A CONSULTATION. SUBSEQUENT ECHOCARDIOGRAPHY SHOWED SEVERELY HYPOKINETIC RIGHT VENTRICLE AND NO SIGNIFICANT TAMPONADE. AFTER SEVERAL HOURS, CARDIAC ARREST OCCURRED AS A CONSEQUENCE OF PULSELESS ELECTRICAL ACTIVITY, FOLLOWED BY ASYSTOLE, WHICH LED TO THE PATIENT¿S DEATH DESPITE RESUSCITATION Manufacturer Narrative: CASE STUDY: "THREE FATAL CONSEQUENCES OF ONE GUIDING CATHETER MANEUVER DURING PRIMARY PERCUTANEOUS ANGIOPLASTY IN ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION." YEAR: 2024 AUTHOR: JAN ROCZNIAK, STANISLAW BARTUS, MICHAL CHYRCHEL REF: DOI: 10.20452/PAMW.16881. B 2.3: DATE OF PUBLICATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: ANNEX D CODE. CORRECTION: ANNEX F CODE. JOURNAL: POLISH ARCHIVES OF INTERNAL MEDICINE ISSUE: 12 PMID: 39498945. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21995768&pc=DSQ | 2916596-2025-02719 | 2025-04-15 04:00 | Death | THORATEC CORPORATION | 2025-05-09 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. IT WAS UNKNOWN IF THIS WAS DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED AND THE PUMP WAS NOT EXPLANTED. NO FURTHER INFORMATION WAS AVAILABLE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21995784&pc=OZD | 1220648-2025-28270 | 2025-04-19 04:00 | Death | ABIOMED, INC. | 2025-05-09 04:00 | OZD | IMPELLA CP | Positioning Failure; Restricted Flow rate | Hemolysis; Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA CP WAS SUPPORTING A PATIENT WITH MYOCARDIAL INFARCTION. NO PRIOR CARDIAC HISTORY WAS NOTED. DURING SUPPORT, IT WAS NOTED THAT THE PATIENT HAD LITTLE TO NO PULSATILITY. IN ADDITION, THE PATIENT'S CONDITION CONTINUED TO WORSEN. MULTIPLE DRIPS WERE ADDED. THE PATIENT EXPERIENCED CONTINUOUS SUCTION ALARMS AND THE PUMP WAS REPOSITIONED. URINE OUTPUT DROPPED OFF. LABS WERE HEMOLYZING SO THERE WAS INCOMPLETE LAB DATA. THE DECISION WAS MADE TO ESCALATE TO AN IMPELLA 5.5. THE IMPELLA CP WAS EXPLANTED AND AN IMPELLA 5.5 WAS SUCCESSFULLY IMPLANTED. HOWEVER, THE PATIENT CODED AND RECEIVED SEVERAL SHOCKS. THE PATIENT EXPIRED ON IMPELLA 5.5 SUPPORT. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW, POSITIONING ISSUES, AND HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. POSITIONING ISSUES: THE MORNING AFTER DEVICE IMPLANTATION THE NIGHT BEFORE, THE DEVICE WAS REPOSITIONED AT BEDSIDE DUE TO SUCTION ALARMS, BUT NO CLINICAL INFORMATION WAS PROVIDED AS TO THE CAUSE OF THE POSITIONING. THE PUMP WAS NOT RETURNED. THE CAUSE OF THE POSITIONING ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS. LOW PUMP FLOW: THE PUMP WAS NOT RETURNED. DATA LOG REVIEW SHOWED FLOW TRENDS DOWN TO 2.5L/MIN AT P-7 WHERE EXPECTED FLOWS RANGE FROM 2.9-3.3L/MIN. THIS OCCURRED DURING NUMEROUS SUCTION ALARMS. REPOSITIONING WAS REPORTED TO HAVE OCCUR NUMEROUS TIMES WITH UNKNOWN RESOLVE, HOWEVER THE EXACT TIMING OF THESE REPOSITIONS ARE NOT KNOWN. MC CURRENT SPIKE OCCURS AT 0650 AND IS FOLLOWED BY INCREASE AND ELEVATED MC AND PUMP FLOWS, EVEN THOUGH THE PUMP FLOWS STILL REMAIN WITHIN THE P-6 FLOW RANGE OF 2.5-2.9L/MIN. THE REPOSITIONING OF THE PUMP, DETERIORATING PATIENT CONDITION / WORSENING HEART FUNCTION, AND MC SPIKE ARE ALL POTENTIAL CONTRIBUTING FACTORS, AND WITHOUT THE RETURNED PRODUCT AND FURTHER CLINICAL DETAILS ONE ROOT CAUSE CANNOT BE DETERMINED. THE CAUSE OF THE LOW PUMP FLOW WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS. HEMOLYSIS: DATA LOG REVIEW SHOWED NO PLACEMENT SIGNAL TRENDS. SUCTION ALARMS FOUND IN DATA LOGS. MC CURRENT SPIKE OCCURS AT 0650 AND IS FOLLOWED BY INCREASE AND ELEVATED MC AND PUMP FLOWS, EVEN THOUGH THE PUMP FLOWS STILL REMAIN WITHIN THE P-6 FLOW RANGE OF 2.5-2.9L/MIN. THE REPOSITIONING OF THE PUMP, DETERIORATING PATIENT CONDITION / WORSENING HEART FUNCTION, AND MC SPIKE ARE ALL POTENTIAL CONTRIBUTING FACTORS, AND WITHOUT THE RETURNED PRODUCT AND FURTHER CLINICAL DETAILS ONE ROOT CAUSE CANNOT BE DETERMINED. THE CAUSE OF THE HEMOLYSIS WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL INFORMATION. H6 HEALTH EFFECT - CLINICAL CODE 1942 AND COMPONENT CODE 3038 WERE ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28270. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21995983&pc=PYX | 1220648-2025-28272 | 2025-04-18 04:00 | Death | ABIOMED, INC. | 2025-05-09 04:00 | PYX | IMPELLA RP FLEX | Adverse Event Without Identified Device or Use Problem | Hemolysis; Thrombosis/Thrombus | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED A MALE PATIENT WAS PLACED ON BIPELLA SUPPORT. THE PATIENT LABS WERE HEMOLYZING. A DECISION WAS MADE TO REMOVE THE IMPELLA RP FLEX. UPON REMOVAL, A CLOT WAS NOTED AT THE PIGTAIL OF THE DEVICE. FURTHER INFORMATION STATED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21996179&pc=OZD | 1220648-2025-28274 | 2025-04-18 04:00 | Death | ABIOMED, INC. | 2025-05-09 04:00 | OZD | IMPELLA CP | Positioning Failure | Hemolysis; Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED A IMPELLA CP TO SUPPORT A PATIENT ADMITTED IN WITH ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK. IT WAS REPORTED THAT AN ECHOCARDIOGRAM SHOWED THE IMPELLA CP IN POOR POSITION. IT WAS REPORTED THAT THERE WAS SIGNS OF HEMOLYSIS WITH THE PATIENTS PLASMA FREE HEMOGLOBIN REPORTED TO BE 210. DECISION WAS MADE FOR CARE TO BE WITHDRAWN AND THE PATIENT EXPIRED.Event Description: ADDITIONALLY, THE PATIENT WAS PERI PROCEDURALLY SHOCKED TWICE FOR VENTRICULAR TACHYCARDIA (VT). Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES, HEMOLYSIS, AND VT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. B3 DATE OF EVENT WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28274. B5 ADDITIONAL INFORMATION THAT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28274. H6 HEALTH EFFECT - CLINICAL CODE 2095 AND HEALTH EFFECT - IMPACT CODE 4647 WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28274.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21996193&pc=NPT | 2025587-2025-03571 | 2025-02-05 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-09 04:00 | NPT | EVOLUT FX PLUS VALVE | Device Markings/Labelling Problem; Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER VALVE, THE INCORRECT VALVE SERIAL NUMBER WAS PROVIDED ON THE PATIENT'S REGISTRATION CARD. SUBSEQUENTLY, THE SERIAL NUMBER WAS CORRECTED; HOWEVER, A NEW PATIENT REGISTRATION CARD WAS NEVER SENT, AS THE PATIENT DIED DUE TO AN UNKNOWN REASON. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21996303&pc=OZD | 1220648-2025-28276 | 2025-04-19 04:00 | Death | ABIOMED, INC. | 2025-05-09 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT AN IMPLANTED IMPELLA CP PUMP IN A 68-YEAR-OLD FEMALE PATIENT WAS UNABLE TO GET FLOWS ABOVE .4 L/MIN. IT WAS NOTED THAT ACTIVE CPR WAS BEING PERFORMED DUE TO COMPLETE CARDIAC STANDSTILL LEADING UP TO AND THROUGHOUT THE EVENT. THE PUMP WAS REPLACED, BUT THE PATIENT LATER PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO DISASSOCIATE THE IMPELLA WITH THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE LOW PUMP FLOW ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED BASED ON THE DATA LOG INVESTIGATION SHOWING LOW FLOWS FROM START OF CASE AND PLACEMENT SIGNAL (PS) TRENDING DOWN, AND CLINICAL REPORTS OF PATIENT'S DETERIORATING CONDITION REQUIRING EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO)..
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21996739&pc=LWS | 2649622-2025-12699 | 2025-04-11 04:00 | Death | MPRI | 2025-05-09 04:00 | LWS | SPRINT QUATTRO SECURE MRI SURESCAN | Electrical /Electronic Property Problem; Over-Sensing | Cardiac Arrest; Fall; Paralysis; Ventricular Fibrillation; Vertebral Fracture | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND DOWN IN CARDIAC ARREST. THE PATIENT APPEARED TO BE IN VENTRICULAR FIBRILLATION (VF) AND CARDIOPULMONARY RESUSCITATION WITH MULTIPLE SHOCKS WAS PERFORMED BY EMERGENCY MEDICAL SERVICES. A LEAD INTEGRITY ALERT (LIA) TRIGGERED ON THE RIGHT VENTRICULAR (RV) LEAD DUE TO HIGH RATE NON-SUSTAINED EPISODES AND SHORT VENTRICULAR INTERVALS. OVERSENSING WAS ALSO NOTED. THERE WAS ALSO OVERSENSING AND UNDERSENSING ON THE RIGHT ATRIAL (RA) LEAD. THE HEALTH CARE PROFESSIONAL STATED, ¿THE DEVICE DIDN'T WORK TO SAVE THE PATIENT AND I FEEL LIKE IT WAS DUE TO THE LIA¿ AND "THE PATIENT IS ON HOSPICE AND GOING TO DIE BECAUSE OF THIS. PATIENT FRACTURED C4/C5 AND IS A QUADRIPLEGIC. THE DEVICE COULD HAVE BROUGHT THEM BACK AND AVOIDED ALL OF THIS.¿ NO INFORMATION WAS PROVIDED RELATED TO THE CONTEXT OF THE CERVICAL FRACTURE. THE EPISODES WERE REVIEWED WHICH NOTED THE SHORT VENTRICULAR INTERVALS WERE DUE TO TRUE EPISODES. ADDITIONALLY, THE DEVICE DID NOT TREAT THE PATIENT¿S ARRHYTHMIA DUE TO NEVER REACHING THE DETECTION OR THERAPY ZONES. THE PATIENT PASSED AWAY DUE TO MULTIPLE BLUNT FORCE INJURIES.Event Description: IT WAS FURTHER REPORTED, VIA FOLLOW UP, THAT THE PATIENT HAD A FALL DURING THEIR CARDIAC ARREST EPISODE, WHICH RESULTED IN THEIR NECK BREAKING. Manufacturer Narrative: CONTINUATION OF D10: 4568-53 LEAD, IMPLANTED (B)(6) 2003. 419388 LEAD, IMPLANTED (B)(6) 2003. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: H6 (IME/ANNEX E) E0602. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21996761&pc=DSQ | 2916596-2025-02946 | 2025-04-20 04:00 | Death | THORATEC CORPORATION | 2025-05-09 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Renal Impairment | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO A MASSIVE CEREBRAL HEMORRHAGE. THE CAUSE OF THE CEREBRAL HEMORRHAGE WAS NOT REPORTED BUT WAS NOT DUE TO MECHANICAL FALL OR A DEVICE RELATED REASON. THE PATIENT'S DEATH WAS RELATED TO A HEALTH CONDITION. THEY WERE INITIALLY TREATED FOR RENAL INSUFFICIENCY, BUT LATER ON EXPIRED DUE TO CEREBRAL HEMORRHAGE. THE DEVICE WAS WORKING AS INTENDED DURING TIME OF SUPPORT. IT WAS REPORTED THAT NO, THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION.Event Description: IT WAS REPORTED THAT THE PATIENT'S RENAL INSUFFICIENCY WAS NOT RELATED TO DEVICE OR DEVICE THERAPY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE REPORTED EVENTS, AS WELL AS A DIRECT CORRELATION TO HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS INVESTIGATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING STROKE, RENAL DYSFUNCTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21996771&pc=MNT | 2518422-2025-044714 | 2025-04-10 04:00 | Death | RESPIRONICS, INC. | 2025-05-09 04:00 | MNT | RESPIRONICS | Failure to Transmit Record |  | K102985 |  | 1 | Event Description: PHILIPS RECEIVED A COMPLAINT ON THE V60 VENTILATOR INDICATING THAT THERE WAS A COMMUNICATION FAILURE BETWEEN THE VENTILATOR AND THE CENTRAL MONITORING STATION. THE DEVICE WAS REPORTED TO BE IN USE AT THE TIME OF THE REPORTED PROBLEM. A PATIENT WAS ADMITTED FOR TREATMENT OF COVID-19. THE PATIENT WAS CONNECTED TO THE V60 VENTILATOR. THE VENTILATOR WAS CONNECTED TO A PHILIPS PATIENT MONITOR VIA A NETWORK CABLE TO MONITOR THE VENTILATION PARAMETERS AT A CENTRAL MONITORING STATION. THE NETWORK CABLE USED WAS FOUND TO HAVE A DAMAGED OR BROKEN FRONT LOCKING "NUT". THIS WAS NOT NOTICED AT FIRST BY THE USERS. AS A RESULT, THE VENTILATION DATA WAS NOT TRANSMITTED TO THE CENTRAL MONITORING SYSTEM, BUT THE DATA TRANSMISSION FAILURE INITIALLY WENT UNNOTICED. ADDITIONALLY, THE SPO2 FINGER SENSOR DETACHED FROM THE PATIENT MONITOR IN USE. PO2 SENSOR DISCONNECT IS NOT CONSIDERED TO BE RELATED TO THE V60 AND IS NOT CONSIDERED AN ADDITIONAL ALLEGATION AGAINST A PHILIPS PRODUCT, BUT IT WAS ENTERED INTO THE SUMMARY SO THAT ALL OF THE INFORMATION KNOWN PERTAINING TO THE SITUATION IS COMMUNICATED. A CRITICAL ALARM WAS TRIGGERED AT SOME POINT AFTER; IT IS UNCLEAR WHEN. THE CUSTOMER STATED THAT THE PATIENT'S CONDITION WAS ALREADY LIFE THREATENING. LIFE SAVING MEASURES WERE ATTEMPTED, BUT THE PATIENT DIED. THIS INVESTIGATION IS ONGOING. Manufacturer Narrative: E: (B)(6) GERMANY PHONE:(B)(6) Manufacturer Narrative: THE REMOTE SERVICE ENGINEER (RSE) STATED IN A GOOD FAITH EFFORT (GFE) RESPONSE THAT THE CUSTOMER COMMUNICATED "THE SMALL PLASTIC TAB FOR LOCKING THE RJ45 PLUG IS BROKEN OFF, SO THE CABLE COULD SLIP OUT UNNOTICED. IT WAS CONNECTED TO THE V60." THIS SMALL PLASTIC TAB IS PART OF THE RJ45 CABLE, NOT THE V60 DEVICE. BASED ON THE INFORMATION PROVIDED, IT APPEARS THAT THE CAUSE FOR THE DATA COMMUNICATION ISSUE IS THE DAMAGE TO THE NON-PHILIPS RJ45 CABLE CAUSING THE CABLE TO BE INSECURELY CONNECTED. ADDITIONAL GFE'S ARE BEING COMPLETED TO CONFIRM THAT THE DEVICE SUCCESSFULLY PASSES A PERFORMANCE VERIFICATION TEST (PVT) TO CONFIRM THE DEVICE IS FUNCTIONING AS INTENDED. THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21996798&pc=NPT | 2025587-2025-03572 | 2024-10-13 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-09 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS AND 18 DAYS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: UPDATED D.4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21996821&pc=DSQ | 2916596-2025-02495 | 2025-01-23 05:00 | Death | THORATEC CORPORATION | 2025-05-09 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Hemorrhagic Stroke; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿THE IMPACT OF PRE-TRANSPLANT VENTRICULAR ASSIST DEVICE SUPPORT IN PEDIATRIC PATIENTS WITH END-STAGE HEART FAILURE ON THE OUTCOMES OF HEART TRANSPLANTATION¿¿A SINGLE CENTER EXPERIENCE¿¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH CEREBRAL BLEEDING, MULTIORGAN FAILURE, AND DEATH. THIS RETROSPECTIVE SINGLE-CENTER COHORT STUDY EVALUATED 41 VAD PATIENTS WITH END-STAGE HEART FAILURE TRANSPLANTED BETWEEN OCT 1988 AND JULY 2024 (TWO OF WHICH WERE HEARTMATE 3 LVAD PATIENTS) FOR THE EFFECT OF VAD SUPPORT ON POST-TRANSPLANT ADVERSE EVENTS SUCH AS BLEEDING, EFFUSION, RESTERNOTOMY, PRIMARY GRAFT FAILURE, RE-TRANSPLANTATION, ARRHYTHMIA, AND STROKE. THE PRIMARY ENDPOINT WAS MORTALITY AT FOLLOW-UP TIME, AND THE SECONDARY ENDPOINT WAS ONE-YEAR RETRANSPLANT-FREE SURVIVAL. THE PATIENTS WERE DIVIDED INTO GROUP 1 (VAD PATIENTS) AND GROUP 2 (NON-VAD PATIENTS). THE OBJECTIVE OF THE STUDY WAS THE ESTIMATION OF THE IMPACT OF ISCHEMIC TIME ON THE DEVELOPMENT OF EARLY GRAFT FAILURE AND THE IDENTIFICATION OF STATISTICALLY SIGNIFICANT DIFFERENCES BETWEEN THE GROUPS OF ¿ALIVE¿ (DEFINED AS SURVIVAL AT FOLLOW-UP TIME) AND ¿MORTALITY¿ (RETRANSPLANT FREE SURVIVAL DEFINED AS DEATH OR RETRANSPLANTATION WITHIN 1-YEAR POST-TRANSPLANT). 12 OF THE PATIENTS EXPIRED WHILE AWAITING TRANSPLANTATION, 6 OF WHICH DIED FROM CEREBRAL BLEEDING AND 6 FROM MULTIORGAN FAILURE. THE CUMULATIVE SURVIVAL RATE AT FOLLOW-UP PERIOD WAS 67±10% IN GROUP 1 VS. 60±6% IN GROUP 2(P=0.769). THE MEAN WAITING TIME (DAYS) ON THE LIST WAS 205±155 IN GROUP 1 AND 119±69 IN GROUP 2 (P=0.002). THE MEAN LENGTH OF HOSPITAL STAY (DAYS) WAS 214±209 IN GROUP 1 AND 128±91 DAYS IN GROUP 2. EARLY PRIMARY-GRAFT FAILURE WAS 10% IN GROUP 1 AND 13% IN GROUP 2. ODDS RATIO [OR] IS AS FOLLOWS: 1.992, 95% CONFIDENCE INTERVAL [CI]: 0.983¿1.007, P=0.266, AORTIC CLAMP TIME PER MINUTES: OR: 1.008, 95% CI: (0.997¿1.019), P=0.164, HLM TIME PER MINUTES: OR: 0.996, 95% CI: (0.991¿1.001), P=0.146, OPERATION TIME PER MINUTES: OR: 1.000, 95% CI: (0.995¿1.004), P=0.861. THE MEAN TIME TO MORTALITY (YEARS) WAS 5.4±5 FOR GROUP 1 AND 7.5±7 FOR GROUP 2. THE DETAILS OF AGE(YEARS), WEIGHT(KG), AND BMI (KG/M2) WERE 9.3±8.2, 35±30, AND 17±5 FOR GROUP 1, RESPECTIVELY (P-VALUES WERE 0.970, 0.706, AND 0.589, RESPECTIVELY). THE 1-YEAR RETRANSPLANT-FREE SURVIVAL RATE WAS EVALUATED USING A LOG-RANK TEST (P=0.219) AND IN GROUP 1 TWO PATIENTS DIED AT 1-YEAR POSTTRANSPLANT. THE CUMULATIVE SURVIVAL RATE AT THE FOLLOW-UP TIME POINT DID NOT DEMONSTRATE A STATISTICALLY SIGNIFICANT DIFFERENCE BETWEEN THE TWO GROUPS (P=0.769). THE 95% CONFIDENCE INTERVAL (CI) DID NOT EXHIBIT SIGNIFICANT VARIATION ACROSS ALL OBSERVED VARIABLES. THE PROVISION OF PRE-HTX VAD SUPPORT DOES NOT HAVE AN ADVERSE EFFECT ON THE SHORT- AND LONG-TERM SURVIVAL OF PEDIATRIC PATIENTS UNDERGOING HTX. A HIGHER MORTALITY RATE WAS OBSERVED AMONG CHILDREN UNDER THREE MONTHS OF AGE WITH CONGENITAL HEART DISEASE. THE PATIENTS WHO RECEIVED VAD SUPPORT WERE IN A CRITICAL CONDITION AND REQUIRED MORE ECMO SUPPORT. THE RESULTS DEMONSTRATED A STATISTICALLY SIGNIFICANT CORRELATION BETWEEN PROLONGED WAITING TIMES AND LENGTH OF HOSPITAL STAY IN GROUP 1. MORE HOMOGENEOUS AND ADEQUATELY POWERED COHORTS ARE NEEDED TO BETTER UNDERSTAND THE IMPACT OF VAD SUPPORT ON POSTTRANSPLANT OUTCOMES. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING STROKE, MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: THE REPORT WAS ORIGINALLY SUBMITTED ON 08MAY2025 AND RECEIVED A 2ND ACKNOWLEDGEMENT. THE 3RD ACKNOWLEDGEMENT FAILED DUE TO A CORRECTION NEEDED FOR REPORTER ADDRESS IN SECTION E1 EXCEEDING THE 60-CHARACTER LIMIT. SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS THE DATE OF PUBLICATION (23JAN2025). AUTHOR INFORMATION: ROSENTHAL LL, GRINNINGER C, ULRICH SM, DALLA POZZA R, HAAS NA, BRENNER P, SCHMOECKEL M, MICHEL S, HAGL C AND HÖRER J (2025) THE IMPACT OF PRE-TRANSPLANT VENTRICULAR ASSIST DEVICE SUPPORT IN PEDIATRIC PATIENTS WITH END-STAGE HEART FAILURE ON THE OUTCOMES OF HEART TRANSPLANTATION¿¿A SINGLE CENTER EXPERIENCE¿. HTTPS://DOI.ORG/10.3389/FCVM.2025.1515218 LUDWIG MAXIMILIAN UNIVERSITY MUNICH NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21997014&pc=MVK | 3008642652-2025-04685 | 2025-03-24 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-09 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FIVE APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 13:35:37 ON (B)(6) 2025. AT 13:37:12, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. AT 13:42:07, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 13:43:22, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VF WITH CPR/MOTION ARTIFACT. AT 13:54:28, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. AT 13:54:58, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 13:55:23, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF WITH CPR/MOTION ARTIFACT. AT 13:57:26, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VF. THE ELECTRODE BELT WAS DISCONNECTED AT 13:57:45 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21997176&pc=QZI | 2124215-2025-30007 | 2025-04-12 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-09 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Ventricular Fibrillation; Respiratory Arrest |  |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER THE PATIENT PASSED AWAY. SIX DAYS AFTER THE PFA PROCEDURE THE PATIENT WAS BROUGHT TO THE HOSPITAL BY EMERGENCY MEDICAL STAFF (EMS) AFTER THEY EXPERIENCED CARDIAC ARREST AND RESPIRATORY FAILURE. WHEN EMS REACHED THE PATIENT, THEY WERE ALREADY ON THE GROUND AND IN ASYSTOLE. THE PATIENT WAS GIVEN EPINEPHRINE WHICH PUT HIM IN A JUNCTIONAL RHYTHM. FURTHER EPINEPHRINE LED TO VENTRICULAR FIBRILLATION. HE WAS GIVEN AMIODARONE BY EMS AND CARDIOVERSION WAS ATTEMPTED TWICE. BY THE TIME THE PATIENT ARRIVED AT THE HOSPITAL THEY WERE EXPERIENCING PULSELESS ELECTRICAL ACTIVITY. RESUSCITATION WAS ATTEMPTED FOR A LITTLE OVER AN HOUR BUT WAS ULTIMATELY UNSUCCESSFUL. THE PATIENT WAS COOL TO THE TOUCH AND HIS PUPILS WERE FIXED AND DILATED. THE CATHETER WAS DISPOSED BY THE FACILITY AND WILL NOT BE RETURNED AS THE DEVICE PERFORMED AS EXPECTED DURING THE PROCEDURE AND THERE WERE NO REPORTS OF ANY PERFORMANCE CONCERNS. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21997179&pc=NPT | 2025587-2025-03578 | 2025-04-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-09 04:00 | NPT | EVOLUT FX DCS | Difficult to Advance; Appropriate Term/Code Not Available; Patient Device Interaction Problem | Stroke/CVA | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO A TRANSCATHETER AORTIC VALVE PROCEDURE INVOLVING THE MEDTRONIC TRANSCATHETER AORTIC VALVE EVFXPLUS-29, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING A 23 MM NON-MEDTRONIC BALLOON DUE TO CALCIFICATION. A 29 FX+ VALVE WAS IMPLANTED AT 3MM ON THE NON-CORONARY CUSP (NCC) AND 6MM ON THE LEFT CORONARY CUSP (LCC). MODERATE TO SEVERE PARAVALVULAR LEAK (PVL) WAS OBSERVED NECESSITATING A POST-IMPLANT BAV. A POST-BAV WAS ATTEMPTED WITH A 24 MM NON-MEDTRONIC BALLOON, BUT THE VALVE MOVED ABOVE THE ANNULUS DUE TO THE TEMPORARY PACEMAKER BEING TURNED OFF BEFORE THE BALLOON WAS DEFLATED. MISCOMMUNICATION BETWEEN THE IMPLANTING PHYSICIAN DOCTOR AND REGISTERED NURSE LED TO PACING BEING STOPPED PREMATURELY, WHICH CONTRIBUTED TO THE VALVE MOVING ABOVE THE ANNULUS, RESULTING IN SEVERE PVL. A SECOND VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE INSERTED INTO THE PATIENT, BUT THE DCS NOSECONE PUSHED THE FIRST VALVE FURTHER INTO THE ANNULUS. THE VALVE'S NEW DEPTH WAS 16MM ON BOTH THE NCC AND LCC. THE FIRST VALVE WAS SNARED TO A HIGHER POSITION, RESULTING IN MODERATE PVL. THE IMPLANTING PHYSICIAN ATTEMPTED A BALLOON DILATATION TO ANCHOR THE FIRST VALVE IN PLACE, BUT THIS LED TO LOSS OF PACING CAPTURE AND CAUSED THE VALVE TO MOVE ABOVE THE SINOTUBULAR JUNCTION (STJ). THE FIRST VALVE WAS SNARED AGAIN TO A HIGHER POSITION IN THE ASCENDING AORTA, AND MULTIPLE ATTEMPTS TO CROSS THE FIRST VALVE WITH THE SECOND VA LVE WERE UNSUCCESSFUL. THE PROCEDURE WAS ABORTED FOR PATIENT SAFETY, BUT THE SNARE COULD NOT BE RELEASED FROM AROUND THE PADDLE ON THE 29 VALVE, POSSIBLY DUE TO A "KNOT". THE PATIENT WAS TAKEN TO THE OPERATING ROOM FOR SNARE REMOVAL AND EVALUATION OF THE NATIVE AORTIC VALVE, WITH A SURGICAL AORTIC VALVE REPLACEMENT (AVR) PERFORMED. THE PATIENT PASSED AWAY FROM A STROKE ON 9 DAYS LATER.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER AND A MEDTRONIC GUIDEWIRE WAS USED DURING THE PROCEDURE. PER THE PHYSICIAN, THE STROKE AND DEATH WERE NOT RELATED TO THE DELIVERY CATHETER SYSTEM (DCS), AND THE VALVE COULD BE RULED OUT AS A CONTRIBUTING FACTOR TO THE DEATH BECAUSE AN AORTIC VALVE REPLACEMENT WAS PERFORMED AND THE TRANSCATHETER VALVE WAS EXPLANTED PRIOR TO THE DEATH.Event Description: IT WAS REPORTED THAT PRIOR TO A TRANSCATHETER AORTIC VALVE PROCEDURE INVOLVING THE MEDTRONIC TRANSCATHETER AORTIC VALVE EVFXPLUS-29, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING A 23 MM NON-MEDTRONIC BALLOON DUE TO CALCIFICATION. A 29 FX+ VALVE WAS IMPLANTED AT 3MM ON THE NON-CORONARY CUSP (NCC) AND 6MM ON THE LEFT CORONARY CUSP (LCC). MODERATE TO SEVERE PARAVALVULAR LEAK (PVL) WAS OBSERVED NECESSITATING A POST-IMPLANT BAV. A POST-BAV WAS ATTEMPTED WITH A 24 MM NON-MEDTRONIC BALLOON, BUT THE VALVE MOVED ABOVE THE ANNULUS DUE TO THE TEMPORARY PACEMAKER BEING TURNED OFF BEFORE THE BALLOON WAS DEFLATED. MISCOMMUNICATION BETWEEN THE IMPLANTING PHYSICIAN DOCTOR AND REGISTERED NURSE LED TO PACING BEING STOPPED PREMATURELY, WHICH CONTRIBUTED TO THE VALVE MOVING ABOVE THE ANNULUS, RESULTING IN SEVERE PVL. A SECOND VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE INSERTED INTO THE PATIENT, BUT THE DCS NOSECONE PUSHED THE FIRST VALVE FURTHER INTO THE ANNULUS. THE VALVE'S NEW DEPTH WAS 16MM ON BOTH THE NCC AND LCC. THE FIRST VALVE WAS SNARED TO A HIGHER POSITION, RESULTING IN MODERATE PVL. THE IMPLANTING PHYSICIAN ATTEMPTED A BALLOON DILATATION TO ANCHOR THE FIRST VALVE IN PLACE, BUT THIS LED TO LOSS OF PACING CAPTURE AND CAUSED THE VALVE TO MOVE ABOVE THE SINOTUBULAR JUNCTION (STJ). THE FIRST VALVE WAS SNARED AGAIN TO A HIGHER POSITION IN THE ASCENDING AORTA, AND MULTIPLE ATTEMPTS TO CROSS THE FIRST VALVE WITH THE SECOND VA LVE WERE UNSUCCESSFUL. THE PROCEDURE WAS ABORTED FOR PATIENT SAFETY, BUT THE SNARE COULD NOT BE RELEASED FROM AROUND THE PADDLE ON THE 29 VALVE, POSSIBLY DUE TO A "KNOT". THE PATIENT WAS TAKEN TO THE OPERATING ROOM FOR SNARE REMOVAL AND EVALUATION OF THE NATIVE AORTIC VALVE, WITH A SURGICAL AORTIC VALVE REPLACEMENT (AVR) PERFORMED. THE PATIENT PASSED AWAY FROM A STROKE ON 9 DAYS LATER. ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER AND A MEDTRONIC GUIDEWIRE WAS USED DURING THE PROCEDURE. PER THE PHYSICIAN, THE STROKE AND DEATH WERE NOT RELATED TO THE DELIVERY CATHETER SYSTEM (DCS), AND THE VALVE COULD BE RULED OUT AS A CONTRIBUTING FACTOR TO THE DEATH BECAUSE AN AORTIC VALVE REPLACEMENT WAS PERFORMED AND THE TRANSCATHETER VALVE WAS EXPLANTED PRIOR TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE GUIDEWIRE DID NOT CONTRIBUTE TO THE DISLODGE, STROKE, OR DEATH. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVFXPLUS-29 (J375814); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. ADDED THIRD PARAGRAPH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW: ONE MEDIA FILE WAS PROVIDED FOR REVIEW OF THE EVENT DESCRIPTION ABOVE. ONLY ONE MEDIA FILE WAS PROVIDED SHOWING A VALVE IMPLANTED AND POSITIONED AT THE ANNULAR LEVEL AND SIGNIFICANT AORTIC REGURGITATION/PARAVALVULAR LEAK PRIOR TO THE POST IMPLANT DILATATION. THERE IS NO OTHER IMAGING AVAILABLE TO REVIEW THEREFORE THIS REVIEW IS INCONCLUSIVE. THE PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. UPDATED DATA: B5. D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21997182&pc=NPT | 2025587-2025-03577 | 2025-04-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-09 04:00 | NPT | EVOLUT FX PLUS VALVE | Perivalvular Leak; Device Dislodged or Dislocated; Patient Device Interaction Problem | Stroke/CVA; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO A TRANSCATHETER AORTIC VALVE PROCEDURE INVOLVING THE MEDTRONIC TRANSCATHETER AORTIC VALVE EVFXPLUS-29, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING A 23 MM NON-MEDTRONIC BALLOON DUE TO CALCIFICATION. A 29 FX+ VALVE WAS IMPLANTED AT 3MM ON THE NON-CORONARY CUSP (NCC) AND 6MM ON THE LEFT CORONARY CUSP (LCC). MODERATE TO SEVERE PARAVALVULAR LEAK (PVL) WAS OBSERVED NECESSITATING A POST-IMPLANT BAV. A POST-BAV WAS ATTEMPTED WITH A 24 MM NON-MEDTRONIC BALLOON, BUT THE VALVE MOVED ABOVE THE ANNULUS DUE TO THE TEMPORARY PACEMAKER BEING TURNED OFF BEFORE THE BALLOON WAS DEFLATED. MISCOMMUNICATION BETWEEN THE IMPLANTING PHYSICIAN DOCTOR AND REGISTERED NURSE LED TO PACING BEING STOPPED PREMATURELY, WHICH CONTRIBUTED TO THE VALVE MOVING ABOVE THE ANNULUS, RESULTING IN SEVERE PVL. A SECOND VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE INSERTED INTO THE PATIENT, BUT THE DCS NOSECONE PUSHED THE FIRST VALVE FURTHER INTO THE ANNULUS. THE VALVE'S NEW DEPTH WAS 16MM ON BOTH THE NCC AND LCC. THE FIRST VALVE WAS SNARED TO A HIGHER POSITION, RESULTING IN MODERATE PVL. THE IMPLANTING PHYSICIAN ATTEMPTED A BALLOON DILATATION TO ANCHOR THE FIRST VALVE IN PLACE, BUT THIS LED TO LOSS OF PACING CAPTURE AND CAUSED THE VALVE TO MOVE ABOVE THE SINOTUBULAR JUNCTION (STJ). THE FIRST VALVE WAS SNARED AGAIN TO A HIGHER POSITION IN THE ASCENDING AORTA, AND MULTIPLE ATTEMPTS TO CROSS THE FIRST VALVE WITH THE SECOND VALVE WERE UNSUCCESSFUL. THE PROCEDURE WAS ABORTED FOR PATIENT SAFETY, BUT THE SNARE COULD NOT BE RELEASED FROM AROUND THE PADDLE ON THE 29 VALVE, POSSIBLY DUE TO A "KNOT". THE PATIENT WAS TAKEN TO THE OPERATING ROOM FOR SNARE REMOVAL AND EVALUATION OF THE NATIVE AORTIC VALVE, WITH A SURGICAL AORTIC VALVE REPLACEMENT (AVR) PERFORMED. THE PATIENT PASSED AWAY FROM A STROKE ON 9 DAYS LATER.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER AND A MEDTRONIC GUIDEWIRE WAS USED DURING THE PROCEDURE. PER THE PHYSICIAN, THE STROKE AND DEATH WERE NOT RELATED TO THE DELIVERY CATHETER SYSTEM (DCS), AND THE VALVE COULD BE RULED OUT AS A CONTRIBUTING FACTOR TO THE DEATH BECAUSE AN AORTIC VALVE REPLACEMENT WAS PERFORMED AND THE TRANSCATHETER VALVE WAS EXPLANTED PRIOR TO THE DEATH.Event Description: IT WAS REPORTED THAT PRIOR TO A TRANSCATHETER AORTIC VALVE PROCEDURE INVOLVING THE MEDTRONIC TRANSCATHETER AORTIC VALVE EVFXPLUS-29, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED USING A 23 MM NON-MEDTRONIC BALLOON DUE TO CALCIFICATION. A 29 FX+ VALVE WAS IMPLANTED AT 3MM ON THE NON-CORONARY CUSP (NCC) AND 6MM ON THE LEFT CORONARY CUSP (LCC). MODERATE TO SEVERE PARAVALVULAR LEAK (PVL) WAS OBSERVED NECESSITATING A POST-IMPLANT BAV. A POST-BAV WAS ATTEMPTED WITH A 24 MM NON-MEDTRONIC BALLOON, BUT THE VALVE MOVED ABOVE THE ANNULUS DUE TO THE TEMPORARY PACEMAKER BEING TURNED OFF BEFORE THE BALLOON WAS DEFLATED. MISCOMMUNICATION BETWEEN THE IMPLANTING PHYSICIAN DOCTOR AND REGISTERED NURSE LED TO PACING BEING STOPPED PREMATURELY, WHICH CONTRIBUTED TO THE VALVE MOVING ABOVE THE ANNULUS, RESULTING IN SEVERE PVL. A SECOND VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE INSERTED INTO THE PATIENT, BUT THE DCS NOSECONE PUSHED THE FIRST VALVE FURTHER INTO THE ANNULUS. THE VALVE'S NEW DEPTH WAS 16MM ON BOTH THE NCC AND LCC. THE FIRST VALVE WAS SNARED TO A HIGHER POSITION, RESULTING IN MODERATE PVL. THE IMPLANTING PHYSICIAN ATTEMPTED A BALLOON DILATATION TO ANCHOR THE FIRST VALVE IN PLACE, BUT THIS LED TO LOSS OF PACING CAPTURE AND CAUSED THE VALVE TO MOVE ABOVE THE SINOTUBULAR JUNCTION (STJ). THE FIRST VALVE WAS SNARED AGAIN TO A HIGHER POSITION IN THE ASCENDING AORTA, AND MULTIPLE ATTEMPTS TO CROSS THE FIRST VALVE WITH THE SECOND VALVE WERE UNSUCCESSFUL. THE PROCEDURE WAS ABORTED FOR PATIENT SAFETY, BUT THE SNARE COULD NOT BE RELEASED FROM AROUND THE PADDLE ON THE 29 VALVE, POSSIBLY DUE TO A "KNOT". THE PATIENT WAS TAKEN TO THE OPERATING ROOM FOR SNARE REMOVAL AND EVALUATION OF THE NATIVE AORTIC VALVE, WITH A SURGICAL AORTIC VALVE REPLACEMENT (AVR) PERFORMED. THE PATIENT PASSED AWAY FROM A STROKE ON 9 DAYS LATER. ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER AND A MEDTRONIC GUIDEWIRE WAS USED DURING THE PROCEDURE. PER THE PHYSICIAN, THE STROKE AND DEATH WERE NOT RELATED TO THE DELIVERY CATHETER SYSTEM (DCS), AND THE VALVE COULD BE RULED OUT AS A CONTRIBUTING FACTOR TO THE DEATH BECAUSE AN AORTIC VALVE REPLACEMENT WAS PERFORMED AND THE TRANSCATHETER VALVE WAS EXPLANTED PRIOR TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE GUIDEWIRE DID NOT CONTRIBUTE TO THE DISLODGE, STROKE, OR DEATH. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-2329 (UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. ADDED THIRD PARAGRAPH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW: ONE MEDIA FILE WAS PROVIDED FOR REVIEW OF THE EVENT DESCRIPTION ABOVE. ONLY ONE MEDIA FILE WAS PROVIDED SHOWING A VALVE IMPLANTED AND POSITIONED AT THE ANNULAR LEVEL AND SIGNIFICANT AORTIC REGURGITATION/PARAVALVULAR LEAK PRIOR TO THE POST IMPLANT DILATATION. THERE IS NO OTHER IMAGING AVAILABLE TO REVIEW THEREFORE THIS REVIEW IS INCONCLUSIVE. THE PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. UPDATED DATA: B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21997195&pc=OZD | 1220648-2025-28281 | 2025-04-19 04:00 | Death | ABIOMED, INC. | 2025-05-09 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT FOLLOWING IMPELLA CP IMPLANT, THE PUMP ENCOUNTERED IMMEDIATE SUCTION. VOLUME WAS GIVEN, BUT THE FLOW REMAINED LOW AND THE SUCTION ALARMS RETURNED. THE STAFF BELIEVED THAT FURTHER CARE WAS FUTILE AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. DATA LOGS WERE INVESTIGATED AND LOW PUMP FLOW AND SUCTION ALARM WAS CONFIRMED AT CASE START. THE PLACEMENT SIGNAL WAS OBSERVED TO BE MARGINALLY DRIFTING DOWN. THEREFORE, AS PER THE CLINICAL INFORMATION AND DATA LOG REVIEW, THE ROOT CAUSE OF THE LOW PUMP FLOW ISSUE WAS MOST LIKELY SPONTANEOUS VENTRICULAR RUPTURE PATIENT CONDITION RELATED SINCE LOW FLOWS WERE OBSERVED SINCE PUMP START EVEN AFTER GIVING VOLUME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21997201&pc=LYJ | 1644487-2025-10112 | 2025-04-16 04:00 | Death | LIVANOVA USA, INC. | 2025-05-09 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED FROM THE PHYSICIAN THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. VNS PROGRAMMING DATA WAS SENT IN TO ASSESS DEVICE FUNCTION. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 CFR PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21997227&pc=NKM | 2135147-2025-02487 | 2021-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-05-09 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Incomplete Coaptation | Insufficient Information | P100009 |  | 1 | Event Description: THE ARTICLE "ANATOMICAL SUITABILITY AND CLINICAL OUTCOMES OF TRANSCATHETER EDGE-TO-EDGE REPAIR IN CHINESE PATIENTS WITH SIGNIFICANT MITRAL REGURGITATION: A REAL-WORLD ECHOCARDIOGRAPHIC STUDY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE ANATOMICAL SUITABILITY AND CLINICAL EFFICACY OF MITRAL VALVE TRANSCATHETER EDGE TO EDGE REPAIR (M TEER) IN CHINESE PATIENTS WITH MODERATE TO SEVERE OR SEVERE MITRAL REGURGITATION (MR). DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT M TEER SUCCESS RATES IN CHINESE PATIENTS ARE COMPARABLE TO WESTERN COUNTRIES, BUT FURTHER RESEARCH IS NEEDED TO EXPLORE POTENTIAL ETHNIC DIFFERENCES. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS PROFESSOR XIANGBIN PAN AT NATIONAL CENTER FOR CARDIOVASCULAR DISEASE, CHINA & FUWAI HOSPITAL, CHINESE ACADEMY OF MEDICAL SCIENCES & PEKING UNION MEDICAL COLLEGE, BEIJING, CHINA WITH CORRESPONDING EMAIL *PANXIANGBIN@FUWAIHOSPITAL.ORG*.] THE TIME FRAME OF THE STUDY WAS JANUARY 2021 TO FEBRUARY 2024. SAY NOT CONFIRMED OR REPORTED IF APPLICABLE. A TOTAL OF 481 PATIENTS WERE INCLUDED IN THIS STUDY, WITH 48% (205) OF THOSE RECEIVING M-TEER. OF WHICH X RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 68.81 YEARS, MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED CORONARY HEART DISEASE, HYPERTENSION, DIABETES, ATRIAL FIBRILLATION, CHRONIC LUNG DISEASE, RENAL INSUFFICIENCY, REFLUX ESOPHAGITIS, HEART FAILURE, CPT, INTRA-AORTIC BALLOON PUMP, CONGENITAL HEART DISEASE, HYPERTROPHIC CARDIOMYOPATHY, DILATED CARDIOMYOPATHY, ALCOHOLIC CARDIOMYOPATHY, MALIGNANT TUMOR SURGERY, PERMANENT PACEMAKER, INTRA-CARDIAC DEVICE, PRIOR CARDIAC PERCUTANEOUS INTERVENTION, PRIOR OPEN-HEART SURGERY. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE HOSPITALIZATION, RECURRENT MR, SINGLE LEAFLET DEVICE ATTACHMENT, CONVERSION TO SURGICAL INTERVENTION, PERICARDIAL EFFUSION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING CORONARY HEART DISEASE, HYPERTENSION, DIABETES, ATRIAL FIBRILLATION, CHORNIC LUNG DISEASE, RENAL INSUFFICIENCY, REFLUX ESOPHAGITIS, HEART FAILURE, CPT, INTRA-AORTIC BALLOON PUMP, CONGENITAL HEART DISEASE, HYPERTROPHIC CARDIOMYOPATHY, DILATED CARDIOMYOPATHY, ALCOHOLIC CARDIOMYOPATHY, MALIGNANT TUMOR SURGERY, PERMANENT PACEMAKER, INTRA-CARDIAC DEVICE, PRIOR CARDIAC PERCUTANEOUS INTERVENTION, PRIOR OPEN HEART SURGERY. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE HOSPITALIZATION, RECURRENT MR, SINGLE LEAFLET DEVICE ATTACHMENT, CONVERSION TO SURGICAL INTERVENTION, PERICARDIAL EFFUSION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. THE ADDITIONAL MALFUNCTION AND PATIENT EFFECTS REFERENCED IN B5 WILL BE FILED UNDER SEPARATE MEDWATCH REPORT NUMBERS. B3: THE EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: ANATOMICAL SUITABILITY AND CLINICAL OUTCOMES OF TRANSCATHETER EDGE-TO-EDGE REPAIR IN CHINESE PATIENTS WITH SIGNIFICANT MITRAL REGURGITATION: A REAL-WORLD ECHOCARDIOGRAPHIC STUDY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21997263&pc=FRN | 1416980-2025-02774 |  | Death | BAXTER HEALTHCARE CORPORATION | 2025-05-09 04:00 | FRN | NOVUM IQ | Filling Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K211122 |  | 1 | Event Description: IT WAS REPORTED DURING A LEVOPHED INFUSION THE NOVUM IQ LARGE VOLUME PUMP OVER-INFUSED THE ¿ENTIRE BAG¿ RESULTING IN THE PATIENT¿S DEATH. IN THE MEDWATCH, THE REPORTER STATED, ¿IF THE TUBING TERMINATING AT THE PATIENT RECEIVES JUST THE RIGHT AMOUNT OF TENSION IT CAUSES A SIPHONING EFFECT WHERE THE ENTIRE BAG OF MEDICATION WILL RAPIDLY INFUSE INTO THE PATIENT, WHILE IT APPEARS THAT THE MEDICATION IS INFUSING APPROPRIATELY ON THE PUMP SCREEN.¿ FURTHERMORE, THE REPORTER ADDED ¿THERE ARE NO SAFETY MEASURES IN EFFECT AT THIS TIME TO PREVENT THIS FROM HAPPENING AGAIN." THE RESPONDENT REQUESTED TO REMAIN ANONYMOUS; THEREFORE, FOLLOW UP INFORMATION WAS NOT ABLE TO BE OBTAINED. NO FURTHER INFORMATION WAS AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI) #: THE SERIAL NUMBER (10) IS UNKNOWN; THEREFORE, THE UDI NUMBER ONLY WILL CONTAIN THE DEVICE IDENTIFIER (01). THE DEVICE WAS NOT RETURNED, AND THE SERIAL NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. FURTHERMORE, THE RESPONDENT REQUESTED TO REMAIN ANONYMOUS; THEREFORE, BAXTER IS UNABLE TO REQUEST THE DEVICE FOR EVALUATION. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21997374&pc=DSQ | 2916596-2025-02813 | 2025-04-18 04:00 | Death | THORATEC CORPORATION | 2025-05-09 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO BIVENTRICULAR (BIV) FAILURE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. Manufacturer Narrative: SECTION D4: EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING IF THE DEVICE OPERATED AS EXPECTED; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21997484&pc=DSQ | 2916596-2025-02823 | 2025-04-21 04:00 | Death | THORATEC CORPORATION | 2025-05-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER: (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21997927&pc=DSQ | 3007042319-2025-01128 | 2025-04-04 04:00 | Death | HEARTWARE, INC. | 2025-05-09 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Power Problem; Material Split, Cut or Torn | Depression | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH A HISTORY OF MAJOR DEPRESSIVE EPISODES CUT THE DRIVELINE CABLE OF THEIR VENTRICULAR ASSIST DEVICE (VAD) WITH SUICIDAL INTENT. THE PATIENT SUBSEQUENTLY PASSED AWAY.Event Description: IT WAS FURTHER REPORTED THAT A CONTROLLER WAS RETURNED THE MANUFACTURER. MANUFACTURER'S ANALYSIS SUBSEQUENTLY REVEALED AN ENERGY STORAGE PROBLEM. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE DRIVELINE CABLE ASSOCIATED WITH THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE DRIVELINE CABLE ASSOCIATED WITH THE VAD (B)(6) MET ALL REQUIREMENTS FOR RELEASE. RAW LOGS WERE NOT AVAILABLE FOR ANALYSIS. AS A RESULT, THE REPORTED "CUT DRIVELINE" EVENT COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. INFORMATION RECEIVED FROM THE SITE INDICATED THAT THE PATIENT PASSED AWAY DUE TO SUICIDE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE MOST LIKELY ROOT CAUSE OF THE DRIVELINE DAMAGE MAY BE ATTRIBUTED TO THE REPORTED CUTTING OF THE DRIVELINE BY THE PATIENT, AS DESCRIBED IN THE EVENT DETAILS. PER THE INSTRUCTIONS FOR USE, PSYCHIATRIC EPISODES AND DEATH ARE A KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVEN T INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR DEVICE EVALUATION. PRODUCT EVENT SUMMARY: A SEGMENT OF THE DRIVELINE CABLE ASSOCIATED WITH (B)(6), TWO (2) CONTROLLERS ((B)(6)), THREE (3) CONTROLLER AC ADAPTERS ((B)(6)), FOUR (4) BATTERIES ((B)(6)), TWO (2) CONTROLLER DC ADAPTERS ((B)(6)), ONE (1) BATTERY CHARGER ((B)(6)), AND ONE (1) BATTERY CHARGER AC ADAPTER ((B)(6)) WERE RETURNED FOR EVALUATION. NO PERFORMANCE ALLEGATIONS WERE MADE AGAINST THE TWO (2) CONTROLLERS ((B)(6)), THREE (3) CONTROLLER AC ADAPTERS ((B)(6)), FOUR (4) BATTERIES ((B)(6)), TWO (2) CONTROLLER DC ADAPTERS ((B)(6)), ONE (1) BATTERY CHARGER ((B)(6)), AND ONE (1) BATTERY CHARGER AC ADAPTER ((B)(6)) WERE RETURNED FOR EVALUATION. REVIEW OF (B)(6) SERVICE HISTORY RECORDS INDICATED THAT THE DEVICE WAS PREVIOUSLY SERVICED, AND THE REPAIR ACTION WAS VERIFIED. REVIEW OF (B)(6) MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED FOR THE REMAINING DEVICES SINCE NO PERFORMANCE ALLEGATIONS WERE MADE AGAINST THEM. VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICES IN RELATION TO THE REPORTED EVENT. FUNCTIONAL TESTING OF THE DRIVELINE SEGMENT REVEALED THAT THE DRIVELINE PASSED THE CONTINUITY TEST AND LOCKING MECHANISM TEST. VISUAL INSPECTION OF THE RETURNED SEGMENT OF DRIVELINE CABLE REVEALED A CUT IN THE OUTER SHEATH, DISCOLORATION OF THE OUTER SHEATH AND DAMAGE TO THE STRAIN RELIEF. FAILURE ANALYSIS OF (B)(6) REVEALED THAT THE RETURNED CONTROLLER PASSED VISUAL INSPECTION. FUNCTIONAL TESTING REVEALED THAT THE NO POWER ALARM DID NOT SOUND WHEN BOTH POWER SOURCES WERE DISCONNECTED FROM THE CONTROLLER, INDICATING AN ISSUE WITH THE INTERNAL BATTERY. INTERNAL INSPECTION REVEALED THAT THE INTERNAL NICKEL-METAL HYDRIDE (NIMH) BATTERY, WHICH POWERS THE NO POWER ALARM, WAS SLIGHTLY SWOLLEN AND ONE CELL WAS SHORTED. AFTER THE INTERNAL BATTERY WAS REPLACED, THE CONTROLLER PERFORMED AS INTENDED. LOG FILE ANALYSIS ASSOCIATED WITH (B)(6) REVEALED THAT (B)(6) WAS THE PATIENT¿S PRIMARY CONTROLLER, IN USE DURING THE REPORTED EVENT. LOG FILE ANALYSIS REVEALED A DECREASE IN POWER CONSUMPTION AND ESTIMATED FLOWS ON (B)(6) 2025 AND NINETY-NINE (99) LOW FLOW ALARMS WERE LOGGED SINCE (B)(6) 2025. REVIEW OF THE ALARM LOG FILE REVEALED ONE (1) CRITICAL BATTERY ALARM INVOLVING (B)(6) WAS LOGGED ON (B)(6) 2025 AT 21:05:46, WHICH WAS DUE TO THE PATIENT ALLOWING THE BATTERY TO DEPLETE BELOW 10% RSOC. IN ADDITION, LOG FILES REVEALED THAT THE CONTROLLER HAD BEEN IN USE FOR MORE THAN TWO (2) YEARS. FAILURE ANALYSIS OF (B)(6) REVEALED THAT THE RETURNED CONTROLLER PASSED VISUAL INSPECTION AND FUNCTIONAL TESTING. INTERNAL VISUAL INSPECTION OF THE RETURNED CONTROLLER REVEALED NO ANOMALIES. LOG FILE ANALYSIS ASSOCIATED WITH (B)(6) REVEALED THAT THE CONTROLLER WAS NOT IN USE AND WAS LIKELY THE PATIENT¿S BACKUP CONTROLLER. FAILURE ANALYSIS OF THE CONTROLLER AC ADAPTERS, CONTROLLER DC ADAPTERS, BATTERY CHARGER, BATTERY CHARGER AC ADAPTERS AND BATTERIES REVEALED THAT THE UNITS PASSED VISUAL INSPECTION AND FUNCTIONAL TESTING. INTERNAL VISUAL INSPECTION OF THE DEVICES REVEALED DID NOT REVEAL ANY ANOMALIES. LOG FILE ANALYSIS REVEALED NO ANOMALIES INVOLVING AN ADAPTER OR THE BATTERIES WITHIN THE ANALYZED PERIOD. AS A RESULT, THE REPORTED EVENT WAS CONFIRMED. THE MOST LIKELY ROOT CAUSE OF THE DRIVELINE DAMAGE MAY BE ATTRIBUTED TO THE REPORTED CUTTING OF THE DRIVELINE BY THE PATIENT, AS DESCRIBED IN THE EVENT DETAILS. BASED ON HISTORICAL REVIEW OF SIMILAR EVENTS, THE MOST LIKELY ROOT CAUSE OF THE OBSERVED OUTER SHEATH DISCOLORATION MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING DESIGN ISSUES AND/OR EXPOSURE TO UV LIGHT. THE MOST LIKELY ROOT CAUSE OF THE OBSERVED DRIVELINE STRAIN RELIEF DAMAGE MAY BE ATTRIBUTED TO MULTIPLE FACTORS, INCLUDING BUT NOT LIMITED TO WEAR AND/OR HANDLING OF THE DRIVELINE. THE MOST LIKELY ROOT CAUSE OF THE OBSERVED CONTROLLER FAULT ALARM AND NO POWER ALARM NOT SOUNDING DURING TESTING CAN BE ATTRIBUTED TO A REDUCED CHARGE CAPACITY OF THE INTERNAL NIMH BATTERY. CAPA PR00492825 INVESTIGATED INTERNAL BATTERY ISSUES WITH CONTROLLER 2.0. EVEN THOUGH THIS CAPA IS CLOSED, (B)(6) FALLS WITHIN THE BOUNDS OF THIS CAPA. THE MOST LIKELY ROOT CAUSE OF THE CRITICAL BATTERY ALARM CAN BE ATTRIBUTED TO THE PATIENT ALLOWING BATTERIES TO D EPLETE BELOW 10%. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE LOW FLOW EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, AND/OR POOR VAD FILLING. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. INFORMATION RECEIVED FROM THE SITE INDICATED THAT THE PATIENT PASSED AWAY DUE TO SUICIDE. PER THE INSTRUCTIONS FOR USE, PSYCHIATRIC EPISODES AND DEATH ARE A KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ CONTROLLER 2.0 D4: MODEL #: 1420 / CATALOG #: 1420 / EXPIRATION DATE: 30-SEP-2023 / SERIAL OR LOT#: (B)(6) ,D9: YES, RETURN DATE: 23-JUL-2025,H3: YES, H4: MFG DATE: 16-SEP-2022, H5: NO, H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21998022&pc=LYJ | 1644487-2025-10114 | 2022-12-28 05:00 | Death | HOUSTON | 2025-05-09 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | Convulsion, Clonic | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2022 AND IT WAS SEIZURE-RELATED. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 CFR PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21999296&pc=GDW | 2955842-2025-20146 | 2025-04-02 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-05-09 04:00 | GDW | SUREFORM | Adverse Event Without Identified Device or Use Problem | Pancreatitis; Unspecified Tissue Injury; Insufficient Information | K173721 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI ASSISTED SIGMOID COLECTOMY FOR RECTOSIGMOID CANCER, THE PATIENT EXPIRED ON AN UNSPECIFIED DATE. THE SURGEON INFORMED THE INTUITIVE CLINICAL SALES MANAGER (CSM) THAT THE PATIENT DEVELOPED PANCREATITIS POST-OPERATIVELY, UNRELATED TO THE PROCEDURE, WHICH CAUSED PANCREATIC SECRETIONS TO COLLECT IN THE PELVIS AND CAUSE THE BOWEL ANASTOMOSIS TO BREAK DOWN AND EVENTUALLY LEAK. THE LEAK FROM THE BOWEL ANASTOMOSIS LED TO FURTHER UNSPECIFIED COMPLICATIONS. THE TIMELINE OF THE EVENTS WAS NOT PROVIDED AND THE SPECIFIC CAUSE OF DEATH WAS UNKNOWN. IT WAS REPORTED THAT NO MALFUNCTIONS OF THE DA VINCI SYSTEM OCCURRED DURING THE PROCEDURE. Manufacturer Narrative: REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ISSUE RELATED EVENT CODES OCCURRED. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 0 DEGREE, MONOPOLAR CURVED SCISSORS, LARGE CLIP APPLIERS (2), TIP-UP FENESTRATED GRASPER, FENESTRATED BIPOLAR FORCEPS, VESSEL SEALER EXTEND, SUREFORM 60 STAPLER WITH GREEN RELOAD. A SITE HISTORY REVIEW FOUND NO RELEVANT COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE FOUND NO NON-CONFORMANCES IDENTIFIED TO BE RELATED TO THE EVENT. REVIEW OF THE SUREFORM 60 STAPLER LOG FOUND THE INSTRUMENT WAS INSTALLED ON THE SYSTEM 1 TIME AND FIRED 1 GREEN RELOAD. CLAMPING WAS SUCCESSFUL, AND THE FIRING WAS COMPLETED WITH NO PAUSES FOR COMPRESSION. THERE WERE NO STAPLER RELATED ERRORS IN THE SYSTEM LOGS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING A SIGMOID COLECTOMY. THE PATIENT DEVELOPED PANCREATITIS POSTOPERATIVELY, AND THE SURGEON SPECULATED THAT THIS COMPLICATION CONTRIBUTED TO THE ANASTOMOTIC BREAKDOWN. INFORMATION WAS NOT PROVIDED REGARDING THE PATIENT¿S CLINICAL COURSE, HOW THE PANCREATITIS DEVELOPED, OR HOW THE ANASTOMOTIC BREAKDOWN WAS DISCOVERED. HOW THE PATIENT EVENTUALLY DIED WAS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21999297&pc=NAY | 2955842-2025-20136 | 2025-04-03 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-05-09 04:00 | NAY | DAVINCI XI | Insufficient Information | Necrosis; Unspecified Tissue Injury | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED CHOLECYSTECTOMY SURGICAL PROCEDURE, THE PATIENT EXPIRED POSTOPERATIVELY WITH REPORTS OF MESENTERY INJURY AND DEAD BOWEL. THE INTUITIVE CLINICAL SALES REPRESENTATIVE (CSR) WAS INFORMED 13 DAYS POST-PROCEDURE THAT THE PATIENT HAD EXPIRED ON AN UNSPECIFIED DATE. SITE OR STAFF INFORMED THE CSR THAT DURING THE INITIAL ROBOTIC PROCEDURE, THERE WAS DIFFICULTY MAKING THE INITIAL CANNULA PORT ENTRY. A LAPAROSCOPIC CAMERA WAS UTILIZED TO VISUALIZE IF THERE WAS ANY MESENTERY INJURY OR BOWEL INJURY FROM THE ATTEMPTS. AFTER INSPECTING THE BOWEL AND MESENTERY, IT WAS REPORTED THAT NO INJURY WAS NOTED BY THE SURGEON. THE PROCEDURE CONTINUED AND WAS COMPLETED ROBOTICALLY; THERE WERE NO REPORTS OF ANY INTRAOPERATIVE COMPLICATIONS RELATED TO THE DA VINCI SYSTEM, INSTRUMENT, OR ACCESSORY. NO ADDITIONAL INFORMATION WAS PROVIDED. ATTEMPTS TO OBTAIN INFORMATION FROM THE SURGEON AND SITE RISK MANAGEMENT WERE MADE; NO RESPONSE HAS BEEN RECEIVED. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ISSUE RELATED ERRORS OCCURRED. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, MEDIUM-LARGE CLIP APPLIER, AND FENESTRATED BIPOLAR FORCEPS, PERMANENT CAUTERY HOOK, PROGRASP FORCEPS, MEDIUM-LARGE CLIP APPLIER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING OR SUBSEQUENT TO THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING A CHOLECYSTECTOMY. THERE IS INFORMATION SUGGESTING THEY HAD DIFFICULTY DURING INITIAL PORT PLACEMENT. THIS WAS REPORTEDLY INVESTIGATED DURING THE ORIGINAL PROCEDURE AND NO INJURY WAS DETECTED. HOWEVER, A MESENTERIC INJURY AND DEAD BOWEL WERE LATER DISCOVERED. HOW THESE INJURIES WERE DISCOVERED AND HOW THIS MAY HAVE LED TO THE DEATH ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THESE EVENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21982789&pc=FRN | 2016493-2025-78709 | 2025-03-07 05:00 | Death | CAREFUSION SD | 2025-05-08 04:00 | FRN | ALARIS SYSTEM | Use of Device Problem; Inaccurate Delivery | Low Blood Pressure/ Hypotension | K133532 |  | 1 | Event Description: IT WAS INITIALLY REPORTED A "PROGRAMMING ERROR" OCCURRED INVOLVING CONTINUOUS INFUSION OF NOREPINEPHRINE. PER REPORT, THE INFUSION WAS "ERRONEOUSLY PROGRAMMED TO TITRATE BY CHANGING THE RATE FIELD OF THE PUMP WITH ML/HR. INSTEAD OF THE DOSE FIELD WITH MCG/MIN." THE CLINICIAN REPORTEDLY CONFUSED THE RATE WITH THE DOSE AS IT WAS THE FIRST OPTION (FIELD) ON THE SCREEN'S PROGRAMMING WORKFLOW. THE CUSTOMER REACHED OUT TO THE MANUFACTURER REQUESTING TO HAVE THE DOSE FIELD BE MOVED AS THE FIRST OPTION DURING PROGRAMMING WORKFLOW. BD LEARNED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE. IT WAS REPORTED THAT THE EVENT OCCURRED ON 07 MARCH 2025 AT 1430. PER REPORT, NOREPINEPHRINE 4MG/250ML WAS ORDERED AS A TITRATABLE INFUSION (0-30 MCG/MIN. INITIAL DOSE 3 MCG/MIN = 11.3ML/HR. INCREASE OR DECREASE RATE BY 2 MCG/MIN EVERY 5 MIN PRN TO MAINTAIN MAP 65 OR GREATER). HOWEVER, THE PUMP WAS PROGRAMMED TO START AT 3 ML/HR. INSTEAD OF 3 MCG/MIN, TITRATED TO MAX OF 30ML/HR. INSTEAD OF 30 MCG/MIN. THE PATIENT REPORTEDLY RECEIVED INCORRECT DOSE FROM 1430 TO 1627. THERE WAS NO ALLEGED DEVICE MALFUNCTION. SINCE THE PATIENT WAS RECEIVING LOWER DOSE THAN INTENDED DURING THAT TIMEFRAME, THE PATIENT WAS PERSISTENTLY HYPOTENSIVE REQUIRING ADDITIONAL VASOPRESSOR INITIATION TO INCREASE THE BLOOD PRESSURE. THE PATIENT REPORTEDLY DIED ON (B)(6) 2025 AT 0917, AND THE PRIMARY CAUSES OF DEATH WERE "END STAGE CARDIOMYOPATHY, DIABETIC KETOACIDOSIS, AND MEDICAL NONCOMPLIANCE" ACCORDING TO THE FACILITY'S MEDICATION SAFETY PROGRAM MANAGER (PHARM.D). NO AUTOPSY WAS PERFORMED.Event Description: IT WAS INITIALLY REPORTED A "PROGRAMMING ERROR" OCCURRED INVOLVING CONTINUOUS INFUSION OF NOREPINEPHRINE. PER REPORT, THE INFUSION WAS "ERRONEOUSLY PROGRAMMED TO TITRATE BY CHANGING THE RATE FIELD OF THE PUMP WITH ML/HR. INSTEAD OF THE DOSE FIELD WITH MCG/MIN." THE CLINICIAN REPORTEDLY CONFUSED THE RATE WITH THE DOSE AS IT WAS THE FIRST OPTION (FIELD) ON THE SCREEN'S PROGRAMMING WORKFLOW. THE CUSTOMER REACHED OUT TO THE MANUFACTURER REQUESTING TO HAVE THE DOSE FIELD BE MOVED AS THE FIRST OPTION DURING PROGRAMMING WORKFLOW. BD LEARNED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE. IT WAS REPORTED THAT THE EVENT OCCURRED ON (B)(6) 2025 AT 1430. PER REPORT, NOREPINEPHRINE 4MG/250ML WAS ORDERED AS A TITRATABLE INFUSION (0-30 MCG/MIN. INITIAL DOSE 3 MCG/MIN = 11.3ML/HR. INCREASE OR DECREASE RATE BY 2 MCG/MIN EVERY 5 MIN PRN TO MAINTAIN MAP 65 OR GREATER). HOWEVER, THE PUMP WAS PROGRAMMED TO START AT 3 ML/HR. INSTEAD OF 3 MCG/MIN, TITRATED TO MAX OF 30ML/HR. INSTEAD OF 30 MCG/MIN. THE PATIENT REPORTEDLY RECEIVED INCORRECT DOSE FROM 1430 TO 1627. THERE WAS NO ALLEGED DEVICE MALFUNCTION. SINCE THE PATIENT WAS RECEIVING LOWER DOSE THAN INTENDED DURING THAT TIMEFRAME, THE PATIENT WAS PERSISTENTLY HYPOTENSIVE REQUIRING ADDITIONAL VASOPRESSOR INITIATION TO INCREASE THE BLOOD PRESSURE. THE PATIENT REPORTEDLY DIED ON (B)(6) 2025 AT 0917, AND THE PRIMARY CAUSES OF DEATH WERE "END STAGE CARDIOMYOPATHY, DIABETIC KETOACIDOSIS, AND MEDICAL NONCOMPLIANCE" ACCORDING TO THE FACILITY'S MEDICATION SAFETY PROGRAM MANAGER ((B)(6)). NO AUTOPSY WAS PERFORMED. Manufacturer Narrative: A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. Manufacturer Narrative: OMIT : C20 - NO FINDINGS AVAILABLE, D15 - CAUSE NOT ESTABLISHED. ADDITIONAL INFORMATION: IMDRF ANNEX A, B, C, D, G CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY: THE REPORTED INCIDENT OF A PROGRAMMING ERROR WAS NOT REPLICATED THROUGH DEVICE TESTING BUT WAS CONFIRMED THROUGH REVIEW OF THE FACILITY RECORDS. THE SUSPECT PCU AND PUMP MODULE DEVICE LOGS WERE NOT RETURNED FOR INVESTIGATION; HOWEVER, THE CUSTOMER PROVIDED AN ¿ALL INFUSION DETAIL REPORT¿, SHOWING THE PROGRAMMING DETAILS FOR THE INFUSIONS PERFORMED ON (B)(6) 2025 BETWEEN 2:30 PM AND 4:27 PM. THE FACILITY ALSO PROVIDED THE DATA SET NAMED SJO V21241205. ALL INFUSION DETAIL REPORTS DO NOT CONTAIN KEYSTROKE PROGRAMMING AND ARE NOT VALIDATED FOR LOG ANALYSIS PURPOSES. IT IS NOT POSSIBLE TO CONFIRM THE DRUG ID WITHOUT THE PCU EVENT LOGS TO PERFORM THE KEYPRESS REPLICATION. ANALYSIS OF THE ALL INFUSION DETAIL REPORT SHOWS THE INITIAL INFUSION FOR THE DRUG ¿NOREPINEPHRINE¿ WAS PROGRAMMED AND STARTED ON (B)(6) 2025 AT 2:30 PM, AT A CONCENTRATION OF 4 MG / 250 ML AND A RATE OF 3.0 ML/HR, CALCULATING A DOSE OF 0.800 MCG/MIN (INSTEAD OF THE INTENDED 3.0 MCG/MIN). THE VOLUME TO BE INFUSED (VTBI) WAS RECORDED AT 30.0 ML. THE USER TITRATED THE RATE UPWARDS BETWEEN 2:42 PM AND 3:52 PM, STARTING BY INCREASING TO 5.0 ML/HR (FOR A DOSE OF 1.33 MCG/MIN) UNTIL IT REACHED 27.0 ML/HR (FOR A DOSE OF 7.20 MCG/MIN). IT WAS THEN CHANGED BACK DOWN TO 25.0 ML/HR (FOR A DOSE OF 6.67 MCG/MIN) AT 4:10 PM AND WAS LEFT RUNNING UNTIL THE INFUSION REACHED KEEP VEIN OPEN (KVO). AFTER THIS, THE INFUSION WAS RESTARTED WITH THE SAME PARAMETERS AND A VTBI OF 150.0 ML. THE DEVICE WAS KEPT IN USE AND ACTIVELY INFUSING NOREPINEPHRINE UNTIL AT LEAST (B)(6) 2025 AT 9:12 AM. THE FACILITY PROVIDED TWO PICTURES OF WHAT IS SUSPECTED TO BE THE SUSPECT PCU AND SUSPECT PUMP MODULE SHOWING THE ¿RATE¿ OPTION SELECTED, STATING THAT ¿THE RN CONFUSED THE RATE OPTION WITH THE DOSE AS IT WAS THE FIRST OPTION ON THE SCREEN¿. THE FACILITY PROPOSED THAT THE ¿DOSE¿ FIELD BE THE FIRST OPTION. THE PUMP MODULE WAS OBSERVED TO HAVE A THIRD-PARTY DOOR INSTALLED. NO OTHER ANOMALIES OR ISSUES CAN BE DISCERNED FROM THE PICTURES PROVIDED. NO DEVICES WERE RETURNED FOR LABORATORY TESTING; THEREFORE, IT WAS NOT POSSIBLE TO PERFORM AN INTERNAL, EXTERNAL INSPECTION OR LABORATORY TESTING. PROPER OPERATION OF THE BD ALARIS¿ SYSTEM REQUIRES THAT YOU ARE FAMILIAR WITH RELATED FEATURES, SETUP, PROGRAMMING, IV SETS, AND ACCESSORIES. READ ALL INSTRUCTIONS, INCLUDING THOSE FOR ALL ATTACHED MODULE(S) BEFORE USING THE BD ALARIS ¿ SYSTEM. ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED INCIDENT IS ATTRIBUTED TO USER PROGRAMMING ERROR, AS REPORTED BY THE FACILITY AND SUPPORTED BY THE CUSTOMER-PROVIDED ALL INFUSION DETAIL REPORT. NO OBVIOUS ANOMALIES WERE OBSERVED WITH THE SUSPECT DEVICE PICTURES THAT WOULD CONTRIBUTE TO THE REPORTED EVENT. PER 803.52(F)(11)(III), THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21982944&pc=FRN | 2016493-2025-78716 | 2025-03-07 05:00 | Death | CAREFUSION SD | 2025-05-08 04:00 | FRN | ALARIS SYSTEM | Use of Device Problem; Inaccurate Delivery | Low Blood Pressure/ Hypotension | K133532 |  | 1 | Event Description: IT WAS INITIALLY REPORTED A "PROGRAMMING ERROR" OCCURRED INVOLVING CONTINUOUS INFUSION OF NOREPINEPHRINE. PER REPORT, THE INFUSION WAS "ERRONEOUSLY PROGRAMMED TO TITRATE BY CHANGING THE RATE FIELD OF THE PUMP WITH ML/HR. INSTEAD OF THE DOSE FIELD WITH MCG/MIN." THE CLINICIAN REPORTEDLY CONFUSED THE RATE WITH THE DOSE AS IT WAS THE FIRST OPTION (FIELD) ON THE SCREEN'S PROGRAMMING WORKFLOW. THE CUSTOMER REACHED OUT TO THE MANUFACTURER REQUESTING TO HAVE THE DOSE FIELD BE MOVED AS THE FIRST OPTION DURING PROGRAMMING WORKFLOW. BD LEARNED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE. IT WAS REPORTED THAT THE EVENT OCCURRED ON (B)(6) 2025 AT 1430. PER REPORT, NOREPINEPHRINE 4MG/250ML WAS ORDERED AS A TITRATABLE INFUSION (0-30 MCG/MIN. INITIAL DOSE 3 MCG/MIN = 11.3ML/HR. INCREASE OR DECREASE RATE BY 2 MCG/MIN EVERY 5 MIN PRN TO MAINTAIN MAP 65 OR GREATER). HOWEVER, THE PUMP WAS PROGRAMMED TO START AT 3 ML/HR. INSTEAD OF 3 MCG/MIN, TITRATED TO MAX OF 30ML/HR. INSTEAD OF 30 MCG/MIN. THE PATIENT REPORTEDLY RECEIVED INCORRECT DOSE FROM 1430 TO 1627. THERE WAS NO ALLEGED DEVICE MALFUNCTION. SINCE THE PATIENT WAS RECEIVING LOWER DOSE THAN INTENDED DURING THAT TIMEFRAME, THE PATIENT WAS PERSISTENTLY HYPOTENSIVE REQUIRING ADDITIONAL VASOPRESSOR INITIATION TO INCREASE THE BLOOD PRESSURE. THE PATIENT REPORTEDLY DIED ON (B)(6) 2025 AT 0917, AND THE PRIMARY CAUSES OF DEATH WERE "END STAGE CARDIOMYOPATHY, DIABETIC KETOACIDOSIS, AND MEDICAL NONCOMPLIANCE" ACCORDING TO THE FACILITY'S MEDICATION SAFETY PROGRAM MANAGER ((B)(6)). NO AUTOPSY WAS PERFORMED.Event Description: IT WAS INITIALLY REPORTED A "PROGRAMMING ERROR" OCCURRED INVOLVING CONTINUOUS INFUSION OF NOREPINEPHRINE. PER REPORT, THE INFUSION WAS "ERRONEOUSLY PROGRAMMED TO TITRATE BY CHANGING THE RATE FIELD OF THE PUMP WITH ML/HR. INSTEAD OF THE DOSE FIELD WITH MCG/MIN." THE CLINICIAN REPORTEDLY CONFUSED THE RATE WITH THE DOSE AS IT WAS THE FIRST OPTION (FIELD) ON THE SCREEN'S PROGRAMMING WORKFLOW. THE CUSTOMER REACHED OUT TO THE MANUFACTURER REQUESTING TO HAVE THE DOSE FIELD BE MOVED AS THE FIRST OPTION DURING PROGRAMMING WORKFLOW. BD LEARNED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE. IT WAS REPORTED THAT THE EVENT OCCURRED ON (B)(6) 2025 AT 1430. PER REPORT, NOREPINEPHRINE 4MG/250ML WAS ORDERED AS A TITRATABLE INFUSION (0-30 MCG/MIN. INITIAL DOSE 3 MCG/MIN = 11.3ML/HR. INCREASE OR DECREASE RATE BY 2 MCG/MIN EVERY 5 MIN PRN TO MAINTAIN MAP 65 OR GREATER). HOWEVER, THE PUMP WAS PROGRAMMED TO START AT 3 ML/HR. INSTEAD OF 3 MCG/MIN, TITRATED TO MAX OF 30ML/HR. INSTEAD OF 30 MCG/MIN. THE PATIENT REPORTEDLY RECEIVED INCORRECT DOSE FROM 1430 TO 1627. THERE WAS NO ALLEGED DEVICE MALFUNCTION. SINCE THE PATIENT WAS RECEIVING LOWER DOSE THAN INTENDED DURING THAT TIMEFRAME, THE PATIENT WAS PERSISTENTLY HYPOTENSIVE REQUIRING ADDITIONAL VASOPRESSOR INITIATION TO INCREASE THE BLOOD PRESSURE. THE PATIENT REPORTEDLY DIED ON (B)(6) 2025 AT 0917, AND THE PRIMARY CAUSES OF DEATH WERE "END STAGE CARDIOMYOPATHY, DIABETIC KETOACIDOSIS, AND MEDICAL NONCOMPLIANCE" ACCORDING TO THE FACILITY'S MEDICATION SAFETY PROGRAM MANAGER (PHARM.D). NO AUTOPSY WAS PERFORMED. Manufacturer Narrative: OMIT : C20 - NO FINDINGS AVAILABLE, D15 - CAUSE NOT ESTABLISHED. ADDITIONAL INFORMATION: IMDRF ANNEX A, B, C, D, G CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY: THE REPORTED INCIDENT OF A PROGRAMMING ERROR WAS NOT REPLICATED THROUGH DEVICE TESTING BUT WAS CONFIRMED THROUGH REVIEW OF THE FACILITY RECORDS. THE SUSPECT PCU AND PUMP MODULE DEVICE LOGS WERE NOT RETURNED FOR INVESTIGATION; HOWEVER, THE CUSTOMER PROVIDED AN ¿ALL INFUSION DETAIL REPORT¿, SHOWING THE PROGRAMMING DETAILS FOR THE INFUSIONS PERFORMED ON (B)(6) 2025 BETWEEN 2:30 PM AND 4:27 PM. THE FACILITY ALSO PROVIDED THE DATA SET NAMED SJO V21241205. ALL INFUSION DETAIL REPORTS DO NOT CONTAIN KEYSTROKE PROGRAMMING AND ARE NOT VALIDATED FOR LOG ANALYSIS PURPOSES. IT IS NOT POSSIBLE TO CONFIRM THE DRUG ID WITHOUT THE PCU EVENT LOGS TO PERFORM THE KEYPRESS REPLICATION. ANALYSIS OF THE ALL INFUSION DETAIL REPORT SHOWS THE INITIAL INFUSION FOR THE DRUG ¿NOREPINEPHRINE¿ WAS PROGRAMMED AND STARTED ON (B)(6) 2025 AT 2:30 PM, AT A CONCENTRATION OF 4 MG / 250 ML AND A RATE OF 3.0 ML/HR, CALCULATING A DOSE OF 0.800 MCG/MIN (INSTEAD OF THE INTENDED 3.0 MCG/MIN). THE VOLUME TO BE INFUSED (VTBI) WAS RECORDED AT 30.0 ML. THE USER TITRATED THE RATE UPWARDS BETWEEN 2:42 PM AND 3:52 PM, STARTING BY INCREASING TO 5.0 ML/HR (FOR A DOSE OF 1.33 MCG/MIN) UNTIL IT REACHED 27.0 ML/HR (FOR A DOSE OF 7.20 MCG/MIN). IT WAS THEN CHANGED BACK DOWN TO 25.0 ML/HR (FOR A DOSE OF 6.67 MCG/MIN) AT 4:10 PM AND WAS LEFT RUNNING UNTIL THE INFUSION REACHED KEEP VEIN OPEN (KVO). AFTER THIS, THE INFUSION WAS RESTARTED WITH THE SAME PARAMETERS AND A VTBI OF 150.0 ML. THE DEVICE WAS KEPT IN USE AND ACTIVELY INFUSING NOREPINEPHRINE UNTIL AT LEAST (B)(6) 2025 AT 9:12 AM. THE FACILITY PROVIDED TWO PICTURES OF WHAT IS SUSPECTED TO BE THE SUSPECT PCU AND SUSPECT PUMP MODULE SHOWING THE ¿RATE¿ OPTION SELECTED, STATING THAT ¿THE RN CONFUSED THE RATE OPTION WITH THE DOSE AS IT WAS THE FIRST OPTION ON THE SCREEN¿. THE FACILITY PROPOSED THAT THE ¿DOSE¿ FIELD BE THE FIRST OPTION. THE PUMP MODULE WAS OBSERVED TO HAVE A THIRD-PARTY DOOR INSTALLED. NO OTHER ANOMALIES OR ISSUES CAN BE DISCERNED FROM THE PICTURES PROVIDED. NO DEVICES WERE RETURNED FOR LABORATORY TESTING; THEREFORE, IT WAS NOT POSSIBLE TO PERFORM AN INTERNAL, EXTERNAL INSPECTION OR LABORATORY TESTING. PROPER OPERATION OF THE BD ALARIS¿ SYSTEM REQUIRES THAT YOU ARE FAMILIAR WITH RELATED FEATURES, SETUP, PROGRAMMING, IV SETS, AND ACCESSORIES. READ ALL INSTRUCTIONS, INCLUDING THOSE FOR ALL ATTACHED MODULE(S) BEFORE USING THE BD ALARIS ¿ SYSTEM. ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED INCIDENT IS ATTRIBUTED TO USER PROGRAMMING ERROR, AS REPORTED BY THE FACILITY AND SUPPORTED BY THE CUSTOMER-PROVIDED ALL INFUSION DETAIL REPORT. NO OBVIOUS ANOMALIES WERE OBSERVED WITH THE SUSPECT DEVICE PICTURES THAT WOULD CONTRIBUTE TO THE REPORTED EVENT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. Manufacturer Narrative: A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21984290&pc=MKJ | 1220908-2025-01544 | 2025-04-14 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-05-08 04:00 | MKJ | AED PRO | Failure to Charge | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A PATIENT (AGE:70'S & GENDER: MALE) THE DEVICE FAILED TO CHARGE. THE CLINICIAN REPLACED THE DEVICE TO ATTEMPT TO CONTINUE TO TREAT THE PATIENT. THE COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS RECEIVED THE PRODUCT AND WILL BE PROVIDING A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS COMPLETED. Manufacturer Narrative: ZOLL MEDICAL CORPORATION EVALUATED THE DEVICE AND THE CUSTOMER'S REPORT WAS NOT REPLICATED OR CONFIRMED. THE DEVICE FUNCTIONED CORRECTLY AND AS DESIGNED. TWO CLINICAL DATA LOGS FROM THE REPORTED EVENT ON (B)(6) 2025, WERE REVIEWED. ALL THREE ECG ANALYSES CORRECTLY IDENTIFIED ASYSTOLE RHYTHMS AS NON-SHOCKABLE. THE PRESENCE OF TWO LOGS WAS DUE TO THE DEVICE BEING POWER CYCLED BY THE OPERATOR DURING THE EVENT. THE DEVICE PASSED ALL TESTING AND PROVIDED PROMPTS AS EXPECTED. THE DEVICE WILL BE RECERTIFIED AND RETURNED TO THE CUSTOMER. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21984740&pc=OZD | 1220648-2025-28204 | 2025-04-17 04:00 | Death | ABIOMED, INC. | 2025-05-08 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT EXPERIENCED A RUN OF VENTRICULAR TACHYCARDIA AND REQUIRED SHOCK A FEW TIMES. IT WAS FURTHER REPORTED THAT THE PATIENT DEVELOPED A GASTROINTESTINAL BLEED. BIVAL WAS STOPPED. THE PATIENT WAS SCOPED BY GI MD. AN ULCER WAS FOUND AND WAS CAUTERIZED/CLIPPED.Event Description: ADDITIONAL INFORMATION NOTED THE PATIENT EXPIRED ON SUPPORT. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: CORRECTED INFORMATION PROVIDED IN B2, B5, H1 AND H6. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21984997&pc=OZD | 1220648-2025-28206 | 2025-04-17 04:00 | Death | ABIOMED, INC. | 2025-05-08 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Renal Failure; Hemorrhagic Stroke | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED VENTRICULAR TACHYCARDIA (VT), CEREBRAL HEMORRHAGE AND ACUTE KIDNEY INJURY WITH A DEMISE OUTCOME. THE RIGHT AXILLARY CUT DOWN WAS PERFORMED. GRAFT ANASTOMOSED, AND INTRODUCER PLACED, SECURED WITH GRAFT LOCKS. DR SHARP ACCESSED LV ACCESSED WITH .035 GUIDEWIRE AND PIGTAIL CATHETER. THE GUIDEWIRE WAS EXCHANGED FOR AN .018 GUIDEWIRE. THE PIGTAIL CATHETER WAS REMOVED AND THE IMPELLA WAS PRIMED, BACKLOADED OVER WIRE AND ADVANCED INTO THE LEFT VENTRICULAR WITHOUT DIFFICULTY. THE GUIDEWIRE WAS REMOVED OUT AND THE IMPELLA WAS SLOWLY RAMPED UP. THE INCISION WAS CLOSED AND IMPELLA SECURED WITH 3-POINT FIXATION. THROUGHOUT THE PROCEDURE, THE PATIENT WAS IN AND OUT OF VENTRICULAR TACHYCARDIA. THE PATIENT WAS SENT TO THE UNIT ON SUPPORT, AND AFTER ARRIVING ON THE UNIT THE PATIENT WAS IN VENTRICULAR TACHYCARDIA. MILRINONE WAS STOPPED, LIDOCAINE WAS ADDED, AND AN ELECTROPHYSIOLOGIST WAS CONSULTED. THE PATIENT CONTINUED ON SUPPORT, AND APPROXIMATELY FIVE DAYS LATER, THE PATIENT WENT FOR A COMPUTED TOMOGRAPHY SCAN FOR A LEFT VENTRICULAR ASSIST DEVICE WORKUP, AND A "LARGE" BRAIN BLEED WAS NOTICED. ANTICOAGULATION WAS CURRENTLY BEING HELD. NEUROLOGICAL STATUS WAS MONITORED AND SUPPORT CONTINUED. THE NEXT DAY, IT WAS NOTED THE PATIENT REMAINED CRITICALLY ILL, WORSENING OXYGEN AND ACIDOSIS. TWO DAYS THEREAFTER, THE PATIENT HAD AN ACUTE KIDNEY INJURY (AKI) AND WAS PUT ON DIALYSIS. IT IS UNKNOWN IF THE PATIENT HAD A KIDNEY INJURY PRIOR TO IMPELLA SUPPORT AND/OR IF THE IMPELLA CONTRIBUTED TO THE AKI. THE PATIENT WAS CONSULTED FOR PALLIATIVE CARE. THE PATIENT'S CONDITION AND NEUROLOGICAL STATUS REMAINED UNCHANGED, WAS TRANSITIONED TO COMFORT CARE AND EXPIRED TWO DAYS LATER. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS: "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY" (INCLUDING VENTRICULAR PERFORATION). SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: "TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21985143&pc=OZD | 1220648-2025-28207 | 2025-04-19 04:00 | Death | ABIOMED, INC. | 2025-05-08 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Ischemia; Septic Shock; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT WAS CATHED AND MITRAL VALVE CORONARY ARTERY DISEASE WAS FOUND. IT WAS DETERMINED THAT THE PATIENT NEEDED CARDIAC SURGERY BUT THE IMPELLA WAS NOT PLACED DUE TO PERIPHERAL VASCULAR DISEASE. THE PATIENT WAS TAKEN FOR POSSIBLE IMPELLA 5.5 PLACEMENT FOR CARDIAC PROCEDURE. HOWEVER, THE PATIENT STARTED DECOMPENSATING SO THE PATIENT WAS SENT TO LAB FOR IMPELLA CP PLACEMENT WITH VASCULAR ASSISTANCE. THE IMPELLA CP WAS PLACED IN USUAL FASHION. LIMB ISCHEMIA WAS FOUND. THE PHYSICIAN KNEW THERE WOULD BE ISCHEMIA SO HE HAD VASCULAR DO THE CASE WITH HIM FROM THE START AND THE LEG WAS BYPASSED. THE CARDIAC PROCEDURE WAS NOT DONE DUE TO THE PATIENT WAS STABILIZED WITH THE IMPELLA CP AND WANTED KIDNEY FUNCTION TO IMPROVE. DURING SUPPORT, IT WAS NOTED THAT THE BYPASS SHEATH CLOTTED OFF AND THE RETROGRADE SHEATH WAS FOUND TO BE KINKED. A NEW REPERCUSSION SHEATH WAS PLACED AND FLOW WAS RESTORED. THE PATIENT ENDED UP HAVING THE INTERVENTION OF TWO MAIN VESSELS WITH NO ISSUE. HOWEVER, THE PATIENT SEEMED TO BE IN FLORID SEPTIC SHOCK AND NONRESPONSIVE TO VASOPRESSORS. THE PATIENT DID GROW BACTERIA ON THE BRONCHEALVEOLAR LAVAGE (BAL). ANTIBIOTICS WERE BEING GIVEN. DESPITE AGGRESSIVE INTERVENTION, THE PATIENT COULD NOT MAINTAIN BLOOD PRESSURE. THE PATIENT CODED AND A DECISION WAS MADE TO STOP THERAPY. THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." IMPELLA CP WITH SMARTASSIST SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA, THROMBUS, AND INFECTION/SEPSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21985368&pc=OZD | 1220648-2025-28210 | 2025-04-20 04:00 | Death | ABIOMED, INC. | 2025-05-08 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhagic Stroke | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT SUFFERED A HEMORRHAGIC STROKE DURING IMPELLA CP SUPPORT. THE PATIENT'S EYES WERE FOUND TO BE DILATED AND A COMPUTED TOMOGRAPHY SCAN REVEALED A MASSIVE BLEED AND INFARCTION. THE DECISION WAS MADE BY THE FAMILY TO WITHDRAW CARE. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21985418&pc=OZD | 1220648-2025-28211 | 2025-04-16 04:00 | Death | ABIOMED, INC. | 2025-05-08 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Application Program Problem | Cardiovascular Insufficiency; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT DURING USE THE AUTOMATED IMPELLA CONSOLE TURNED OFF FOR 35 SECONDS AND REBOOTED. THE PROCEDURE OUTCOME WAS SUCCESSFUL WITH ANOTHER DEVICE.Event Description: THE PATIENT EXPIRED WHILE ON SUPPORT. PER THE MEDICAL SAFETY OFFICER THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE AUTOMATED IMPELLA CONTROLLER (AIC) AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME GIVEN THE SHUT DOWN AND LOSS OF PATIENT SUPPORT. Manufacturer Narrative: THE INVESTIGATION OF AUTOMATED IMPELLA CONTROLLER (AIC) - SHUTDOWN OR RESTART ISSUE HAS BEEN COMPLETED. BASED ON THE DATA AND DEVICE ANALYSIS THE UNEXPECTED REBOOT WAS CAUSED BY THE CALCULATOR SOFTWARE SUBPROCESS CRASHING. THE CAUSE OF THE CALCULATOR CRASHING COULD NOT BE DETERMINED SINCE THE FAILURE MODE COULD NOT BE REPRODUCED. Manufacturer Narrative: DEVICE RETURNED 29-APR-2025, INVESTIGATION STILL PENDING. BASED ON THE MEDICAL SAFETY OFFICER'S (MSO) REVIEW OF THIS CASE THE PATIENT'S DEMISE IS ASSOCIATED WITH THE AUTOMATED IMPELLA CONTROLLER (AIC) ASSOCIATED WITH MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28211 HAS BEEN REVISED TO DEATH. THE FOLLOWING FIELDS HAVE BEEN REVISED. B1 REVISED TO REFLECT BOTH ADVERSE EVENT AND PRODUCT PROBLEM. B2 DEATH AND DATE OF DEATH ADDED ON RECOMMENDATION OF MSO. B5 REVISED WITH MSO STATEMENT. H1 REVISED TO DEATH. H6 HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE UPDATED BASED ON RECOMMENDATION OF MSO. THE IMPELLA PUMP WAS REVISED TO SERIOUS INJURY AND UPDATED ACCORDINGLY UNDER MANUFACTURER DEVICE REPORT NUMBER 11220648-2025-28159. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21985648&pc=BZD | 2518422-2025-044650 | 2023-05-11 04:00 | Death | RESPIRONICS, INC. | 2025-05-08 04:00 | BZD | DREAMSTATION CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21985725&pc=OZD | 1220648-2025-28215 | 2025-04-17 04:00 | Death | ABIOMED, INC. | 2025-05-08 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT HAD AN IMPELLA 5.5 PUMP IMPLANTED AS A BRIDGE TO TRANSPLANT EXHIBITED ARRHYTHMIC STORM. THE STORM PUT THE TRANSPLANT ON HOLD. THE PATIENT EXPIRED. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21985860&pc=DQA | 21985860 | 2025-03-25 04:00 | Death | MEDTRONIC | 2025-05-08 04:00 | DQA | MEDTRONIC CAPNO 20P | Device Alarm System | Insufficient Information |  |  | 1 | Event Description: A MEDICAL EVENT OCCURRED, AND TEAM MEMBERS DID NOT HEAR THE EQUIPMENT ALARM. NO EQUIPMENT ISSUES WERE IDENTIFIED UPON THE CLINICAL ENGINEERING EVALUATION, THE EQUIPMENT FUNCTIONED AS DESIGNED. AN INVESTIGATION REVEALED. THE ALARM SETTINGS WERE SET TO A LOW DECIBEL LEVEL. THE EQUIPMENT ALARM SETTINGS PERMIT AN INDEFINITE LOWING OF THE DECIBEL LEVEL OF THE AUDIBLE ALARM SOUND RATHER THAN LOWING THE AUDIBLE ALARM SOUND FOR A SHORT INTERVAL. THE EQUIPMENT ALSO DOES NOT RESET UPON TURNING THE EQUIPMENT OFF/ON BUT INSTEAD RETAINS THE LAST ALARM SETTING WHEN TURNED OFF/ON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21986054&pc=DSQ | 2916596-2025-03134 | 2025-04-16 04:00 | Death | THORATEC CORPORATION | 2025-05-08 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE PATIENT'S OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. DUE TO HOSPITAL POLICY, NO FURTHER INFORMATION COULD BE PROVIDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE, REV. C IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21986060&pc=QFG | MW5170108 |  | Death | TANDEM DIABETES CARE, INC. | 2025-05-08 04:00 | QFG | TANDEM INSULIN PUMP | Defective Device | Insufficient Information |  |  | 1 | Event Description: IN ACCORDANCE WITH 21 CFR 803.22 (B)(2), WE ARE NOTIFYING YOU THAT ON (B)(6) 2025 THE PATIENT¿S WIFE REPORTED THE PATIENT DIED DUE TO A FAULTY INSULIN PUMP. NO FURTHER INFORMATION WAS PROVIDED. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). REF REPORT: MW5170109. Event Description: ADDITIONAL INFORMATION RECEIVED ON 5/14/2025 FOR REPORT MW5170108 TO UPDATE PROCODE: QFG. REF REPORT: MW5170109.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21986067&pc=LZG | 3004464228-2025-20017 | 2025-04-03 04:00 | Death | INSULET CORPORATION | 2025-05-08 04:00 | LZG | OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | Insufficient Information | Insufficient Information | K211575 |  | 1 | Event Description: ON 14APRIL2025 A HEALTHCARE PROFESSIONAL HAS CONTACTED INSULET VIA EMAIL TO REPORT THAT A PATIENT HAS DIED IN A MOTORCYCLE ACCIDENT ON (B)(6) 2025. THE PATIENT (MALE) WAS USING DASH AT THE TIME OF PASSING. THE INVESTIGATION IS ONGOING AND INSULET IS ATTEMPTING TO RECOVER ADDITIONAL DETAILS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: FOLLOWING THE REPORT OF A FATAL MOTORCYCLE ACCIDENT INVOLVING THE PATIENT ON (B)(6) 2025, BETWEEN APPROXIMATELY 13:38 AND 13:51, AN INVESTIGATION WAS CONDUCTED AT THE PATIENT¿S RESIDENCE AT THE REQUEST OF THE FAMILY. THE PATIENT'S PERSONAL DIABETES MANAGER (PDM), SERIAL NUMBER (B)(6), LCD SCREEN WAS FOUND CRACKED. THE CAUSE AND TIMING OF THE DAMAGE COULD NOT BE DETERMINED. DESPITE THE DAMAGE, THE TOUCHSCREEN REMAINED FULLY FUNCTIONAL. ANALYSIS OF ANDROID LOG FILES FROM (B)(6) 2025 INDICATED THE SYSTEM DID NOT ENCOUNTER ANY ISSUES DURING INSULIN DELIVERY. ALL BOLUSES WERE SUCCESSFULLY DELIVERED IN FULL, AND ALL BASAL INSULIN WAS DELIVERED AS PROGRAMMED. THE PDM HISTORY SHOWS THAT THE USER PERFORMED FIVE BLOOD GLUCOSE MEASUREMENTS ON (B)(6) 2025, WITH AN AVERAGE READING OF 12.9 MMOL/L (242.4 MG/DL), WHICH ALIGNS WITH THE ANDROID LOG DATA. THE POD USED ON (B)(6) 2025 (LOT: PD1K06072441, SERIAL: (B)(6)) WAS VISUALLY INSPECTED. THE ADHESIVE PAD HAD BEEN REMOVED, THE PIEZO KILL SWITCH WAS ENGAGED, AND THE RESERVOIR WAS NEARLY FULL. NO EXTERNAL DAMAGE WAS OBSERVED. A FULL INVESTIGATION WAS NOT POSSIBLE DUE TO FIELD CONDITIONS. BASED ON THE INVESTIGATION OF THE PATIENT'S PERSONAL DIABETES MANAGER, ANDROID LOG FILES FROM (B)(6) 2025, AND THE POD USED ON (B)(6) 2025 THERE IS NO INDICATION THAT THE DEVICES CAUSED OR CONTRIBUTED TO THE REPORTED INCIDENT. Manufacturer Narrative: THE DEVICE HAS NOT BEEN RETURNED/RECEIVED TO DATE. IF THE DEVICE IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED WITH THE INVESTIGATION RESULTS. WE ARE UNABLE TO DETERMINE IF ANY PRODUCT CONDITION COULD HAVE CONTRIBUTED TO THE REPORTED EVENT. NO LOT RELEASE RECORDS WERE REVIEWED, AS THE PRODUCT LOT NUMBER WAS NOT PROVIDED. THE INVESTIGATION IS ONGOING AND INSULET IS ATTEMPTING TO RECOVER ADDITIONAL DETAILS, INCLUDING PRODUCT NAME, LOT NUMBER AND SERIAL NUMBER. IF ADDITIONAL INFORMATION IS RECEIVED, INSULET WILL SUBMIT A SUPPLEMENTAL REPORT AND CONDUCT ALL POSSIBLE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21986087&pc=QFG | MW5170109 |  | Death | TANDEM DIABETES CARE, INC. | 2025-05-08 04:00 | QFG | TANDEM INSULIN PUMP | Defective Device | Insufficient Information |  |  | 1 | Event Description: IN ACCORDANCE WITH 21 CFR 803.22 (B)(2), WE ARE NOTIFYING YOU THAT ON (B)(6) 2025 THE PATIENT¿S WIFE REPORTED THE PATIENT DIED DUE TO A FAULTY INSULIN PUMP. NO FURTHER INFORMATION WAS PROVIDED. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). REF REPORT: MW5170108. Event Description: ADDITIONAL INFORMATION RECEIVED ON 5/14/25 FOR REPORT MW5170109 TO UPDATE PROCODE QFG. REF REPORT: MW5170108.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21986447&pc=PYX | 1220648-2025-28222 | 2025-04-16 04:00 | Death | ABIOMED, INC. | 2025-05-08 04:00 | PYX | IMPELLA RP FLEX | Restricted Flow rate; Increase in Pressure | Cardiac Arrest | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN POSTCARDIOTOMY CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA RP FLEX DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). TWO DAYS PRIOR, THE PATIENT UNDERWENT CORONARY ARTERY BYPASS GRAFT AND 3-VALVE REPLACEMENT SURGERIES, AND SHORTLY AFTER ARRIVING TO THE CARDIOVASCULAR INTENSIVE CARE UNIT (CVICU), THE PATIENT WENT INTO VENTRICULAR FIBRILLATION AND WAS CODED BRIEFLY. AN EMERGENT STERNOTOMY WAS PERFORMED AT BEDSIDE AND CHEST WAS LEFT OPEN FOR TWO DAYS. THE PATIENT WAS TAKEN TO THE OPERATING ROOM FOR SURGICAL CHEST CLOSURE, AND AFTER CLOSING THE CHEST, THE PATIENT'S RIGHT VENTRICLE EXHIBITED DYSFUNCTION ON ECHOCARDIOGRAM WITH CENTRAL VENOUS PRESSURE OF 20 AND HYPOTENSION REQUIRING INOTROPIC SUPPORT. AT THIS TIME THE DECISION WAS MADE TO IMPLANT AN IMPELLA RP FLEX DEVICE WHICH WAS SUCCESSFULLY COMPLETED. APPROXIMATELY TWO DAYS AFTER START OF THE IMPELLA MCS, IT WAS NOTED FLOWS SLOWLY DROPPED OVER COURSE OF THERAPY FROM 3.0 LITERS PER MINUTE TO 2.4 LITERS PER MINUTE. DURING THIS TIME MOTOR CURRENT TRENDED DOWNWARD FROM 380¿S TO 240¿S. PULMONARY ARTERY PLACEMENT SIGNAL GRADUALLY INCREASED INTO 70¿S MEAN. AT THIS TIME, THE PHYSICIAN DECIDED TO INITIATE TISSUE PLASMINOGEN ACTIVATOR (TPA) PROTOCOL FOR PURGE SOLUTION TO PREVENT PUMP FAILURE. 2 MILLIGRAMS OF TPA IN 50 MILLILITERS WAS GIVEN THROUGH PURGE SYSTEM AT WHICH POINT FLOW RESTORATION OCCURRED WITH MOTOR CURRENT RETURNING TO UPPER 300¿S AND THE PULMONARY ARTERY PLACEMENT SIGNAL DECREASING TO MEAN OF 40¿S. THEREAFTER, IT WAS NOTED THE IMPELLA RP FLEX RAN WITHOUT ISSUE SINCE SAID RESOLUTION. THE PATIENT'S CONDITION CONTINUED TO WORSEN DESPITE AGGRESSIVE THERAPY. A DISCUSSION WAS ULTIMATELY MADE WITH THE FAMILY REGARDING PROGNOSIS, AND THE DECISION WAS MADE TO WITHDRAWAL CARE. THE IMPELLA WAS WEANED. THE PATIENT WAS EXTUBATED AND ALL DRIPS WERE TURNED OFF AND THE PATIENT PASSED AWAY SHORTLY THEREAFTER. Manufacturer Narrative: HIGH PURGE PRESSURE: CLINICAL DETAILS REPORT THAT WHILE THE PUMP WAS EXPERIENCING LOW PUMP FLOWS ON (B)(6) 2025, THE TEAM DECIDED TO RUN TPA THROUGH THE PURGE, AND IT RESOLVED THE COMPLICATIONS. DATA LOGS WERE REVIEWED, AND AN INCREASE IN PURGE PRESSURE WAS NOTED DURING THE LOW PUMP FLOWS. THE FIRST STEP UP IN PURGE PRESSURE WAS SMALL, BUT CORRELATED WITH THE SUDDEN DROP IN MC. SOON AFTER, THERE WAS A LARGER STEP UP IN THE PURGE PRESSURE TO ~600 MMHG, NOT HIGH ENOUGH TO TRIGGER AN ALARM. THERE WERE SEVERAL BAG CHANGES DURING THIS PERIOD, BUT NONE LONGER THAN 2 MINUTES. DATA LOGS ALSO SHOW SODIUM BICARB BEING TAKEN OUT OF THE PURGE, PRESUMABLY WHEN TPA WAS ADDED, AND THEN PURGE PRESSURE TRENDING BACK DOWN TO ~400 MMHG. THOUGH TPA BEING ADDED TO THE PURGE CORRELATES WITH THE RESOLUTION OF HIGH PURGE PRESSURE, THE BEHAVIOR OF THE MOTOR CURRENT DURING THE COMPLICATION DOES NOT CORRELATE WITH BIOMATERIAL INGESTION OR BUILD UP. ADDITIONALLY, THE NATURE OF THE RISE IN PURGE PRESSURE DOES NOT ALIGN WITH THE TYPICAL INCREASE DUE TO BIOMATERIAL. PRODUCT WAS NOT RETURNED FOR INVESTIGATION. SINCE DATA LOGS DID NOT REVEAL SUFFICIENT EVIDENCE FOR A CAUSE OF THE RISE IN PURGE PRESSURE AND PRODUCT WAS NOT RETURNED FOR INVESTIGATION, THE ROOT CAUSE OF THE HIGH PURGE PRESSURE COULD NOT BE DETERMINED. LOW PUMP FLOW: CLINICAL DETAILS REPORT THAT ON (B)(6) 2025 PUMP FLOWS HAD SLOWLY TRENDED DOWN FROM 3.0 TO 2.4 L/MIN. DURING THIS TIME, MOTOR CURRENT HAD ALSO TRENDED DOWN, AND PAP TRENDED UP. IT IS REPORTED THAT AFTER TPA WAS ADDED TO THE PURGE, ALL THE PREVIOUSLY MENTIONED PUMP METRICS RETURNED TO NORMAL. DATA LOGS WERE REVIEWED AND THE REPORTED TRENDS WERE CONFIRMED LATE ON (B)(6) 2025. THERE WAS A SUDDEN DROP IN MC AND STEP UP IN THE DP PS WHILE RUNNING AT P-7. THE MC CONTINUED TO TREND DOWN, PUMP FLOWS WENT BELOW SPECIFICATION, AND SHORTLY AFTER THERE WAS A LARGE STEP UP IN PURGE PRESSURE TO 600 MMHG. ALL PUMP METRICS DID RETURN TO NORMAL DURING THE TIME THAT THE PURGE DATA SHOWS SODIUM BICARBONATE WAS REMOVED FROM THE PURGE, PRESUMABLY WHEN TPA WAS ADDED. THOUGH THE ISSUE RESOLVED WHILE TPA WAS BEING RUN IN THE PURGE, THE DROP IN MC AT THE SAME TIME IS NOT INDICATIVE OF BIOMATERIAL. IT WAS NOTED THAT ROUGHLY 17 HOURS PRIOR TO THIS COMPLICATION, THERE WAS A MOTOR CURRENT SPIKE TO 1600 MA THAT WAS INDICATIVE OF A BIOMATERIAL INGESTION EVENT. HOWEVER, MC AND PURGE PRESSURE RESOLVED PRIOR TO THE REPORTED COMPLICATION ON (B)(6) 2025, SO THE TWO EVENTS CANNOT BE DIRECTLY CORRELATED. FINALLY, CLINICAL DETAILS REPORTED THAT THE PATIENT WAS ON CRRT WITH A LOT OF VOLUME BEING PULLED DURING THE TIME OF THE COMPLICATION, WHICH HAS THE POTENTIAL TO PLAY A ROLE IN PUMP FLOWS. PRODUCT WAS NOT RETURNED FOR INVESTIGATION. SINCE PRODUCT WAS NOT RETURNED FOR INVESTIGATION AND DATA LOGS DID NOT HAVE ANY CLEAR PATTERNS INDICATIVE OF A CAUSE, THE ROOT CAUSE OF THE LOW PUMP FLOWS COULD NOT BE DETERMINED. Manufacturer Narrative: ALTHOUGH THE DECISION WAS MADE TO DISCONTINUE IMPELLASUPPORT, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE (INCREASE IN PURGE PRESSURE WITH SUBSEQUENT DECREASED PUMP FLOW) AS AN ASSOCIATED FACTOR TO THE OVERALL OUTCOME. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21986737&pc=NVZ | 2017865-2025-52144 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-08 04:00 | NVZ | ENDURITY¿ CORE SR, JPD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P880086 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21986830&pc=BZD | 2518422-2025-105099 | 2025-04-28 04:00 | Death | RESPIRONICS, INC. | 2025-05-08 04:00 | BZD | DS2ADV AUTO CPAP W/HUMID CELL/BT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction | K200480 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DREAM STATION 2 DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAD PASSED AWAY DUE TO A HEART ATTACK. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21986842&pc=DSQ | 3007042319-2025-01120 | 2025-04-07 04:00 | Death | HEARTWARE, INC. | 2025-05-08 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Dyspnea; Respiratory Failure; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF CARDIOMYOPATHY EXPERIENCED RIGHT VENTRICULAR (RV) FAILURE, SHORTNESS OF BREATH, WORSENING HEART FAILURE AND MULTI-ORGAN SYSTEM FAILURE. AN ECHOCARDIOGRAM SHOWED AN EJECTION FRACTION (EF) OF 19% AND LEFT VENTRICULAR END-DIASTOLIC DIMENSION (LVEDD) OF 6.3 CM, INDICATING REDUCED RV SYSTOLIC FUNCTION. THE PATIENT WAS INTUBATED, HEMODIALYSIS WAS INITIATED, AND INFUSIONS OF DOBUTAMINE AND LASIX WERE ADMINISTERED ALONG WITH A HEPARIN DRIP. THE PATIENT WAS ADMITTED FOR RIGHT HEART FAILURE WITH MULTI-ORGAN SYSTEM FAILURE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. THE VAD REMAINS IN USE.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS EXTUBATED SEVERAL WEEKS LATER, AND HOSPICE CARE WAS INITIATED TWO WEEKS AFTER EXTUBATION PER PATIENT/FAMILY WISHES. THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) HW34736 WAS NOT RETURNED FOR EVALUATION AS IT REMAINS IN USE. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, RIGHT VENTRICULAR FAILURE, WORSENING HEART FAILURE, AND MULTI-ORGAN FAILURE ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21986871&pc=DSQ | 2916596-2025-02792 | 2025-04-18 04:00 | Death | THORATEC CORPORATION | 2025-05-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem |  | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED A SUDDEN DROP IN FLOW AND LOW FLOW ALARM. PATIENT WAS BEING MONITORED IN THE INTENSIVE CARE UNIT (ICU) WITH CATECHOLAMINE THERAPY. CIRCULATION WAS STABLE. VENTRICULAR ASSIST DEVICE (VAD) COORDINATOR CONFIRMED THE FLOW DROP. ON ECHOCARDIOGRAM (ECHO), THE PATIENT HAD A GOOD EJECTION FRACTION (EF) AND SUFFICIENT FLOW INTO BOTH THE VAD AND THE OUTFLOW GRAFT. THE RPM WAS REDUCED TO TEST WHETHER THE PATIENT'S EF WAS SUFFICIENT TO TURN OFF THE VAD SYSTEM. THE MEAN ARTERIAL PRESSURE (MAP) DECREASED, SO THE VAD WAS SET TO THE INITIAL RMP. THE PATIENT PASSED AWAY ON (B)(6) 2025. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME.Event Description: IT WAS REPORTED THAT THE CLINIC WAS UNABLE TO DETERMINE A CAUSE OF THE DROP IN FLOW, THERE WAS NO TECHNICAL DEFECT WITH THE CONTROLLER, ALTHOUGH IT IS NOT IN USE AFTER EXCHANGE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILE CONFIRMED THE REPORTED LOW FLOW ALARMS. A SPECIFIC CAUSE COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CAPTURED EVENTS FROM (B)(6) 2025 THROUGH (B)(6) 2024. CONTINUOUS LOW FLOW ALARMS WERE ACTIVE THROUGHOUT THE DURATION OF THE FILE. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE SET SPEED. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. ADDITIONALY, IT ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. THIS SECTION ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 5 OF THE HEARTMATE 3 LVAS PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING," AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿ OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21986900&pc=BYI | 2183816-2025-00024 | 2025-02-12 05:00 | Death | BAXTER HEALTHCARE CORPORATION | 2025-05-08 04:00 | BYI | THE VEST APX SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Dyspnea | NA |  | 1 | Event Description: THE REPORTER STATED THE PATIENT EXPIRED THE DAY AFTER USING THE VEST APX SYSTEM. IT WAS REPORTED THE PATIENT PERFORMED VEST THERAPY AND THE THERAPY DISLODGED THE MUCUS INTO THE PATIENT¿S BRONCHIAL TUBES WHICH THE REPORTER STATED CAUSED THE PATIENT TO HAVE PROBLEMS WITH BREATHING. IT WAS REPORTED THE PATIENT¿S DIFFICULTY BREATHING CONTINUED THROUGHOUT THE NEXT DAY. EMERGENCY SERVICES REPORTEDLY WERE CALLED; THE PATIENT REPORTEDLY WAS TRANSPORTED TO THE HOSPITAL AND THE REPORTER STATED THAT LATER THAT SAME DAY THE PATIENT DIED. NO FURTHER MEDICAL TREATMENT OR INTERVENTIONS WERE REPORTED. THERE WAS NO REPORT OF A DEVICE MALFUNCTION. NO FURTHER INFORMATION WAS AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: H11: THE DEVICE WAS RECEIVED FOR EVALUATION AND THE DEVICE PASSED ALL FUNCTIONAL TESTING. THE EVENT HISTORY LOG REVIEW SHOWED NO FUNCTION-RELATED ERROR CODES IN MEMORY. THERE WAS ONE INSTANCE OF ERROR 104 (PASSWORD INVALID) IN MEMORY. THE DEVICE WAS FOUND TO HAVE PERFORMED PER PRODUCT SPECIFICATIONS. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21986907&pc=OZD | 1220648-2025-28226 | 2025-04-17 04:00 | Death | ABIOMED, INC. | 2025-05-08 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT ON IMPELLA CP SUPPORT DEVELOPED BLEEDING FROM THE RIGHT GROIN SITE. REPERFUSION SHEATH, IMPELLA CP, AND VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION CANNULA WAS IN THE RIGHT GROIN. SIX UNITS OF BLOOD PRODUCTS WERE ADMINISTERED. THE ANGLE WAS ADJUSTED AND PRESSURE DRESSING APPLIED. IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RETURNED, AND THE INVESTIGATION HAS BEEN COMPLETED. THE ROOT CAUSE OF ACCESS SITE BLEEDING WAS MOST LIKELY ANTICOAGULATION MANAGEMENT SINCE THE ACTIVATED CLOTTING TIME WAS OUT OF RANGE WHICH IS NOT RECOMMENDED PER INSTRUCTIONS FOR USE. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21987266&pc=OZD | 1220648-2025-28228 | 2025-04-18 04:00 | Death | ABIOMED, INC | 2025-05-08 04:00 | OZD | IMPELLA 5.5 | Malposition of Device | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT HAD AN IMPELLA 5.5 AND ECMO SUPPORT ONGOING. AFTER FOURTEEN DAYS OF SUPPORT THE PUMP BECAME MALPOSITIONED AND WAS ADJUSTED. THE PATIENT HAD CARE WITHDRAWN AND EXPIRED. Manufacturer Narrative: A4 AND A5 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21987357&pc=OZP | 2032227-2025-170872 | 2025-04-03 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-08 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Myocardial Infarction; Renal Failure | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025 AT HOME. THE CAUSE OF DEATH WAS HYPERGLYCEMIA, HYPOGLYCEMIA AND KIDNEY PROBLEM. THE INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1884, MMT-332A, MMT-397A, AND MMT-7040A. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING. CUSTOMER HAS BEEN USING THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT AND THE CUSTOMER WAS USING SENSORS AND A TRANSMITTER. THERE WAS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-397A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21987363&pc=OZD | 1220648-2025-28230 | 2025-04-08 04:00 | Death | ABIOMED, INC. | 2025-05-08 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN POSTCARDIOTOMY CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). THE PATIENT HAD ELECTIVE MITRAL VALVE REPAIR SIX DAYS PRIOR AND STAYED IN THE HOSPITAL LOOKING WELL UNTIL OVERNIGHT. THE STARTED TO HAVE DECREASED URINE OUTPUT AND WAS TRANSFERRED TO CARDIOVASCULAR INTENSIVE CARE UNIT. THE PATIENT REQUIRED INCREASING CHEMICAL SUPPORT AND HAD BRIEF A CODE IN THE UNIT. THE DECISION MADE TO PLACE THE IMPELLA CP DEVICE, WHICH WAS SUCCESSFULLY COMPLETED. WHILE IN THE CATHETERIZATION LAB THE RIGHT VENTRICLE WAS ASSESSED, AND A DECISION WAS MADE TO PLACE BIVENTRICULAR SUPPORT (IMPELLA RP FLEX). SOON AFTER THE IMPLANT THE PHYSICIAN HAD CONCERNS REGARDING PERFUSION TO THE RIGHT LEG (IMPELLA CP LEG). LIMB ISCHEMIA WAS INDICATED AND THE IMPELLA CP DEVICE WAS EXPLANTED AND REPLACED WITH AN IMPELLA 5.5 DEVICE FOR CONTINUED BIPELLA MCS. NOW ON IMPELLA 5.5 AND IMPELLA RP FLEX, THE FOLLOWING DAY THE STAFF WAS STILL UNABLE TO GET A DOPPLER PULSE ON THE LOWER EXTREMITIES. THE PATIENT WAS CONTINUED ON BIPELLA SUPPORT AND CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) WAS STARTED. THE NEXT DAY AN AMIODARONE DRIP WAS STARTED DUE TO TACHYCARDIA. A NEUROLOGICAL ASSESSMENT WAS COMPLETED THE DAY AFTER AND THERE WAS NO MOVEMENT NOTED IN THE EXTREMITIES. TWO DAYS LATER, THE PATIENT EXPIRED. CARE WAS WITHDRAWN. Manufacturer Narrative: PRODUCT-SPECIFIC INFORMATION FOR THE IMPELLA CP DEVICE IS UNAVAILABLE AT THE TIME OF THIS MDR WRITING; THEREFORE, RESPECTIVE FIELDS REFLECT "UNKNOWN" OR LEFT BLANK ACCORDINGLY. MEDICAL SAFETY REVIEWED THE EVENT AND NOTED IT APPEARS ARTERIAL FLOW WAS BEING OCCLUDED ASSOCIATED WITH LIMB ISCHEMIA ON THE IMPELLA CP., WHICH APPEARED TO BE AN UNRESOLVED ISSUE (NO INFORMATION OF RESOLUTION ACHIEVED) AND IS THE LEADING ISSUE ASSOCIATED WITH THE PATIENT'S DEMISE OUTCOME. NEURO ASSESSMENT NOTED NO MOVEMENT IN EXTREMITIES (NO STROKE DIAGNOSIS), AND IT APPEARS THE ISSUE IS MOST LIKELY DUE TO THE LIMB ISCHEMIA. ARRHYTHMIC EVENT INVOLVING THE IMPELLA RP FLEX AND IMPELLA 5.5 OCCURRED AND APPEARED TO HAVE BEEN RESOLVED WITH AMIODARONE (THIS IS A SERIOUS INJURY FOR THE IMPELLA 5.5 AND IMPELLA RP FLEX). THE IMPELLA CP IS ONE OF THREE DEVICES ASSOCIATED WITH THE EVENT. NOTE THE FOLLOWING: REFER TO INITIAL MEDICAL DEVICE REPORT FOR THE IMPELLA RP FLEX-SERIAL NUMBER: (B)(6) WITH DATE OF EVENT 10-APR-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). REFER TO INITIAL MEDICAL DEVICE REPORT FOR THE IMPELLA 5.5-SERIAL NUMBER: (B)(6) WITH DATE OF EVENT 10-APR-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21987660&pc=QFG | 3013756811-2025-87167 | 2025-04-11 04:00 | Death | TANDEM DIABETES CARE | 2025-05-08 04:00 | QFG | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Hypoglycemia | K201214 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. REPORTEDLY, PARAMEDICS WERE CALLED AND TRANSPORTED THE CUSTOMER TO THE INTENSIVE CARE UNIT (ICU). THE CUSTOMERS WAS DIAGNOSED BRAIN DEAD IN THE ICU. THE SUSPECTED CAUSE OF DEATH WAS SEVERE HYPOGLYCEMIA. A REVIEW OF THE PUMP DATA LOGS IS EXPECTED; HOWEVER, THE PUMP DATA LOG REVIEW HAS NOT YET BEEN COMPLETED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21988171&pc=OZD | 1220648-2025-28236 | 2025-03-14 04:00 | Death | ABIOMED, INC. | 2025-05-08 04:00 | OZD | IMPELLA 5.5 | Positioning Failure; Restricted Flow rate | Atrial Fibrillation; Hemolysis; Unspecified Infection | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT ON IMPELLA 5.5 NEEDING MECHANICAL CIRCULATORY SUPPORT. THE PUMP HAD SUCTION ALARMS OCCUR DURING SUPPORT. FURTHERMORE, THE PATIENT EXPERIENCED ATRIAL FIBRILLATION AND WAS CARDIOVERTED. THE PATIENT ALSO HAD HEMOLYSIS DURING SUPPORT, THE PUMP WAS REPOSITIONED AWAY FROM THE MITRAL VALVE. MOREOVER, THE PATIENT HAD AN INFECTION ON SUPPORT WITH A LOW-GRADE FEVER. THE SWAN WAS DISCONTINUED AND THE RELATIONSHIP BETWEEN THE INFECTION AND THE IMPELLA IS UNKNOWN. LATER WE RECEIVED A NOTIFICATION VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED WORSENING CARDIOGENIC SHOCK WITH A "DEFINITE" RELATIONSHIP TO THE IMPELLA DEVICE USED: ¿(B)(6) 2025- LDH/PFH UPTRENDING, STARTING INTEGRIIN FOLLOWING C-SHOCK CALL (B)(6) 2025-INTERMITTENT HYPOTENSION/SUCTION ALARMS NOTED. LDH 1187-SLOWLY TRENDING DOWN, REMAINS ON HEPARIN/CANGRELOR (B)(6) 2025-R THORACENTESIS, 900 ML DARK, RED FLUID REMOVED, INTEGRELIN ON HOLD (B)(6) 2025-PRESENTED TO MRB, DECLINED FOR OHT/VAD (B)(6) 2025-PT DNR (B)(6) 2025-PT MADE CMO, ADMITTED TO HOSPICE. IMPELLA 5.5 TURNED OFF @ 1011.¿ THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21988405&pc=DSQ | 2916596-2025-02794 | 2016-01-01 05:00 | Death | THORATEC CORPORATION | 2025-05-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE, "THE IMPACT OF GENETIC POLYMORPHISM ON COMPLICATION DEVELOPMENT IN HEART FAILURE PATIENT", THAT THE HEARTMATE 3 (HM3) LEFT VENTRICULAR ASSIST DEVICE (LVAD) MAY BE ASSOCIATED WITH COMPLICATIONS SUCH AS: DEATH, THROMBOSIS, BLEEDING, INFECTIONS, STROKES, AND MYOCARDIAL INFARCTIONS. THE RETROSPECTIVE STUDY AIMED TO CHARACTERIZE THE INFLUENCE OF THE GENETIC POLYMORPHISMS IN THE UGT1A6 GENE ON COMPLICATIONS IN HEART FAILURE (HF) PATIENTS UNDER ANTIPLATELET THERAPY WITH IMPLANTED LVADS. THERE WERE 100 HEART FAILURE PATIENTS RECRUITED FOR THE STUDY, WHICH WERE IMPLANTED WITH LVADS BETWEEN 2011 AND 2016. PATIENTS WERE DIAGNOSED WITH ISCHEMIC CARDIOMYOPATHY (ICM), DILATED CARDIOMYOPATHY (DCM), HYPERTROPHIC CARDIOMYOPATHY (HCM), AND VALVULAR HEART DISEASE (VHD) PRIOR TO IMPLANTATION. THREE TYPES OF LVADS WERE IMPLANTED IN THE PATIENTS FROM THIS STUDY: 18 PATIENTS WITH THE HEARTWARE HVAD, 34 PATIENTS WITH THE HEARTMATE II (HM II), AND 46 PATIENTS WITH THE HEARTMATE 3 (HM3). 2 PATIENTS WERE EXCLUDED FROM THE STUDY DUE TO BEING UNDER 18 YEARS OF AGE. VENOUS BLOOD SAMPLES COLLECTED IN STERILE VACUTAINERS WITH K2EDTA WERE USED TO PERFORM GENETIC ANALYSIS. 41 OF THE HM3 PATIENTS DID NOT EXPERIENCE ANY COMPLICATIONS, WHEREAS 5 OF THE HM3 PATIENTS DID EXPERIENCE COMPLICATIONS (12 HM II PATIENTS AND 7 HVAD PATIENTS EXPERIENCED COMPLICATIONS). OF THE 71 PATIENTS WHO SURVIVED, 13 EXPERIENCED COMPLICATIONS. OF THE 27 PATIENTS WHO PASSED AWAY, 11 EXPERIENCED COMPLICATIONS. THERE WERE 13 PATIENTS (54.2%, P=0.0001) THAT EXPERIENCED THROMBOSIS COMPLICATIONS, 14 PATIENTS (58.3%, P=0.0001) THAT EXPERIENCED BLEEDING COMPLICATIONS, 15 PATIENTS (62.5%, P=0.015) THAT EXPERIENCED INFECTIONS, 8 PATIENTS THAT EXPERIENCED HEMORRHAGIC STROKES, 12 PATIENTS THAT EXPERIENCED ISCHEMIC STROKES, AND 44 PATIENTS THAT EXPERIENCED MYOCARDIAL INFARCTIONS. THE MEAN DURATION OF LVAD SUPPORT WAS 29.6 ± 17.3 MONTHS IN THE STUDY PERIOD (2011¿2016). 71 (72.4%) PATIENTS REACHED THE OUTCOME MEASUREMENT TIME, WHEREAS 27 (27.6%) PATIENTS DID NOT REACH THE OUTCOME MEASUREMENT. 10 (10.2%) PATIENTS WHO REACHED THE OUTCOME MEASUREMENT TIME HAD HEART TRANSPLANTATION (BTT). PATIENTS WERE PRESCRIBED WARFARIN AND ASPIRIN ANTITHROMBOTIC THERAPY FOR THE LONG TERM AFTER DEVICE IMPLANTATION. THE DOSE OF WARFARIN (2.99 ± 1.15 MG/DAY) WAS CORRECTED TO MAINTAIN THE TARGET INTERNATIONAL NORMALIZED RATIO RANGE (INR 2.25¿3.25). THE DAILY DOSE OF ASPIRIN (100 MG/DAY) WAS PRESCRIBED FOR HF PATIENTS ACCORDING TO THEIR MEDICAL INDICATION AND RANGED MONTHLY IN DYNAMICS (BEFORE 14 DAYS, AS WELL AS AFTER 1, 3, 6, 12, AND 18 MONTHS), UNLESS THERE WAS A CONTRAINDICATION, SUCH AS ULCERATIVE GASTRITIS OR LIVER CIRRHOSIS. ON THE 6TH MONTH, 49 PATIENTS (66.2%) WERE NOT RECEIVING ASPIRIN TREATMENT IN THE GROUP OF PATIENTS WITHOUT COMPLICATIONS, WHEREAS 15 PATIENTS (62.5%) WERE RECEIVING ASPIRIN TREATMENT IN THE GROUP OF PATIENTS WITH COMPLICATIONS (P = 0.02). IT WAS DETERMINED THAT POLYMORPHISMS RS9934438 AND RS9923231 IN THE VKORC1 GENE, RS5918 IN THEITGB3 GENE, AND RS2070959 IN THEUGT1A6 GENE WERE SIGNIFICANTLY ASSOCIATED WITH COMPLICATIONS OF HF PATIENTS ACCORDING TO THE LOGISTIC REGRESSION ANALYSIS. POLYMORPHISM RS2070959 IN THE UGT1A6 GENE IS SIGNIFICANTLY ASSOCIATED WITH HF PATIENTS WHO WERE ON ASPIRIN TREATMENT IN THE 12TH MONTH AND WITHOUT ASPIRIN IN THE 18TH MONTH OF THERAPY (P <0.05). THE GG GENOTYPE IN THE RECESSIVE AND CODOMINANT MODEL OF RS2070959 IN THE UGT1A6 GENE SHOWED A SIGNIFICANT ASSOCIATION WITH ASPIRIN TREATMENT IN THE 12TH MONTH [OR (95% CI): 5.10 (1.31¿19.87), P = 0.018 AND (OR (95% CI): 4.50 (1.05¿19.25), P = 0.054)]. THE AG GENOTYPE IN AN OVER-DOMINANT MODEL OF RS2070959 IN THE UGT1A6 GENE SHOWED A SIGNIFICANT ASSOCIATION WITHOUT ASPIRIN TREATMENT IN THE 18TH MONTH [OR (95% CI): 0.33 (0.12¿0.95), P = 0.032)]. IN SUMMARY, THIS STUDY REPORTS THAT GENOTYPING FOR POLYMORPHISMS OF ITGB3, UGT1A6, AND VKORC1 GENES COULD HELP TO PREVENT AND PREDICT COMPLICATION (THROMBOSIS/BLEEDING) DEVELOPMENT IN HF PATIENTS WITH IMPLANTED LVADS, WHICH WILL IMPROVE THE QUALITY OF LIFE, AND REDUCE THE MORBIDITY AND MORTALITY RATE. THEY STUDY CONCLUDED THAT LVAD COMPLICATIONS DEVELOP NOT ONLY DUE TO THE INCORRECT DOSE OF ANTITHROMBOTIC DRUGS AND THE INFLUENCE OF THE DEVICE¿S NON-PHYSIOLOGICAL SHEAR STRESS (NPSS), BUT ALSO DUE TO THE INHERITED INDIVIDUAL GENETIC CHARACTERISTICS OF HF PATIENTS. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: THE DATE OF THE EVENT HAS BEEN ENTERED AS 01JAN2016 AS PATIENTS WERE IMPLANTED BETWEEN 2011 AND 2016. NATIONAL LABORATORY ASTANA, NAZARBAYEV UNIVERSITY, ASTANA 010000, KAZAKHSTAN; EURASIAN SOCIETY OF PERSONALIZE MEDICINE, ASTANA 010000, KAZAKHSTAN; DEPARTMENT OF MEDICINE, SEMEY MEDICAL UNIVERSITY, PAVLODAR BRANCH, PAVLODAR 140000, KAZAKHSTAN; DEPARTMENT OF MEDICINE, SEMEY MEDICAL UNIVERSITY, SEMEY 071400, KAZAKHSTAN; SERGIEVSKY CENTER, TAUB INSTITUTE, COLUMBIA UNIVERSITY MEDICAL CENTER, NEW YORK, NY 10032, USA; NATIONAL RESEARCH CARDIAC SURGERY CENTER, ASTANA 010000, KAZAKHSTAN; DEPARTMENT OF GENERAL BIOLOGY AND GENOMICS, L. N. GUMILYOV EURASIAN NATIONAL UNIVERSITY ZHALBINOVA, M.R.; RAKHIMOVA, S.E.; KOZHAMKULOV, U.A.; AKILZHANOV, K.R.; SHAIMARDANOV, N.K.; AKILZHANOVA, G.A.; LEE, J.H.; PYA, Y.V.; BEKBOSSYNOVA, M.S.; AKILZHANOVA, A.R. THE IMPACT OF GENETIC POLYMORPHISM ON COMPLICATION DEVELOPMENT IN HEART FAILURE PATIENTS. J. CLIN. MED. 2025, 14, 35. HTTPS://DOI.ORG/10.3390/JCM14010035 NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21988482&pc=KNT | 3010757606-2025-00206 | 2025-04-11 04:00 | Death | ABBVIE - MEDICAL DEVICE CENTER | 2025-05-08 04:00 | KNT | DUODOPA_DUOPA | Adverse Event Without Identified Device or Use Problem | Pneumonia; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K142793 |  | 1 | Event Description: ON AN (B)(6) 2025, A PATIENT IN CROATIA UNDERWENT A PROCEDURE FOR THE PLACEMENT OF PERCUTANEOUS ENDOSCOPIC GASTROSTOMY (PEG) TUBE. IT WAS REPORTED THAT THE PATIENT EXPERIENCED ABDOMINAL PAIN IMMEDIATELY AFTER DISCHARGE FROM PEG TUBE PLACEMENT PROCEDURE. ON (B)(6) 2025, THE PATIENT WAS HOSPITALIZED FOR A PARALYTIC ILEUS. ON (B)(6) 2025, THE PATIENT DEVELOPED PNEUMONIA AND WAS PLACED ON OXYGEN THERAPY. DUE TO UNKNOWN COMPLICATIONS AND PROGRESSIVE DETERIORATION, THE PATIENT EXPIRED ON (B)(6) 2025, DURING HOSPITALIZATION. THE CAUSE OF DEATH WAS NOT PROVIDED AND THE DEVICE WAS NOT RETURNED. AT THE TIME OF THIS REPORT, IT IS UNKNOWN IF AN AUTOPSY HAS BEEN PERFORMED. THE REPORTER DID NOT HAVE ANY ADDITIONAL INFORMATION. ABBVIE MEDICAL SAFETY HAS DETERMINED THAT A CAUSAL RELATIONSHIP BETWEEN THE PEG PLACEMENT PROCEDURE AND THE EVENTS OF PARALYTIC ILEUS, PNEUMONIA, AND DEATH CANNOT BE RULED OUT. Manufacturer Narrative: H6 CODE OF E2402 WAS CHOSEN TO CAPTURE THE EVENT OF PARALYTIC ILEUS. CATALOG NUMBER IN D4 IS THE INTERNATIONAL LIST NUMBER WHICH IS SIMILAR TO US LIST NUMBER OF 062910. THE DEVICE INVOLVED IN THE EVENT WAS NOT RETURNED; THEREFORE, A RETURN SAMPLE EVALUATION IS UNABLE TO BE PERFORMED. ABDOMINAL PAIN, PNEUMONIA, PARALYTIC ILEUS, AND DEATH ARE KNOWN COMPLICATIONS OF A PEG TUBE PLACEMENT. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21988985&pc=DSQ | 2916596-2025-02740 | 2021-12-01 05:00 | Death | THORATEC CORPORATION | 2025-05-08 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED ¿THE IMPACT OF PREOPERATIVE ATRIAL FIBRILLATION ON SURVIVAL AND OUTCOMES IN HEARTMATE II AND HEARTMATE 3 PATIENTS¿ IDENTIFYING THAT HEARTMATE 3 (HM3) MAY BE RELATED TO ATRIAL INFECTION, SEPSIS, RIGHT HEART FAILURE, ISCHEMIC AND HEMORRHAGIC STROKE, PUMP THROMBOSIS, TRANSIENT ISCHEMIC ATTACK (TIA), HEMOTHORAX, KIDNEY DISEASE, GASTROINTESTINAL (GI) BLEEDING AND DEATH. THIS WAS A RETROSPECTIVE STUDY WERE IT WAS REVIEWED THE MEDICAL RECORDS OF ADULT PATIENTS WHO UNDERWENT HEARTMATE II (HMII) OR HM3 IMPLANTATION BETWEEN JANUARY 2012 AND DECEMBER 2021. THE AIM OF THIS STUDY WAS TO EVALUATE WHETHER PREOPERATIVE ATRIAL FIBRILLATION (AF) AFFECTS SURVIVAL AND ADVERSE EVENTS DURING LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY. THE INCIDENCE OF POSTOPERATIVE PNEUMONIA AND SEPSIS DID NOT DIFFER AMONG THE GROUPS. PATIENTS WITH PREOPERATIVE PAROXYSMAL ATRIAL FIBRILLATION (PAF) OR PERSISTENT ATRIAL FIBRILLATION (PEAF) HAD A GREATER INCIDENCE OF POSTOPERATIVE RIGHT HEART FAILURE (RHF) THAN DID PATIENTS WITH SINUS RHYTHM (SR) (SR: 7.5% VS. PAF: 13.2% VS. PEAF: 27.8%, P = 0.037). A TOTAL OF 31.1% OF PATIENTS WITH SR, 55.3% OF PATIENTS WITH PAF AND 72.2% OF PATIENTS WITH PEAF HAD POSTOPERATIVE KD (P = 0.001). THE INCIDENCE OF POSTOPERATIVE ISCHEMIC STROKE (IS) WAS SIGNIFICANTLY GREATER IN BOTH THE PAF AND PEAF GROUPS THAN IN THE SR GROUP (18.4% VS. 27.7% VS. 6.6%, P = 0.012). THE FREQUENCY OF POSTOPERATIVE HEMORRHAGIC STROKE (HS) DID NOT DIFFER AMONG THE THREE GROUPS. THE INCIDENCE OF PUMP THROMBOSIS ALSO DID NOT DIFFER AMONG THE THREE GROUPS. PATIENTS WITH PRE-OPERATIVE PAF OR PEAF HAD A GREATER INCIDENCE OF RETHORACOTOMY THAN DID THOSE IN THE SR GROUP (31.6% VS. 27.8% VS. 12.3%, P = 0.018). PATIENTS WITH PREOPERATIVE PEAF WHO UNDERWENT LVAD IMPLANTATION ALSO HAD AN INCREASED RISK OF MORTALITY (HR: 18.68 (8.04¿43.40, P < 0.001)). IT WAS ALSO CALCULATED THE AVERAGE TREATMENT EFFECT (ATE) FOR MORTALITY AFTER LVAD IMPLANTATION, AND FOUND THAT COMPARED WITH THOSE WITH SR, PATIENTS WITH PEAF HAD A 54% INCREASED RISK OF MORTALITY (P < 0.001). Manufacturer Narrative: SECTION B: DATE OF DEATH ESTIMATED AS EVENT DATE. SECTION D4: MODEL AND CATALOG NUMBERS CORRECTED. SECTION E1: CUSTOMER SITE WAS (B)(6) HOSPITAL AT (B)(6) GERMANY. SECTION G2: REPORT SOURCES CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS THE DATE OF DATA COLLECTION (01DEC2021). SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. AL-KHUSEIN ET AL. BMC CARDIOVASCULAR DISORDERS (2025). 25:236 HTTPS://DOI.ORG/10.1186/S12872-025-04680-1 FACULTY OF MEDICINE, CLINIC FOR CARDIOLOGY, ANGIOLOGY, AND INTENSIVE CARE MEDICINE, RWTH AACHEN UNIVERSITY, UNIVERSITY HOSPITAL AACHEN, AACHEN, GERMANY NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21989049&pc=OZD | 1220648-2025-28239 | 2025-04-14 04:00 | Death | ABIOMED, INC | 2025-05-08 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA 5.5 EXPERIENCED CARDIAC ARRHYTHMIA THAT WAS SUCCESSFULLY HANDLED. ULTIMATELY, THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21989238&pc=NPT | 2015691-2025-03661 | 2025-02-28 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-05-08 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES JAPANESE AFFILIATE, THE PATIENT UNDERWENT A TRANS-AXILLARY OR TRANS-SUBCLAVIAN TAVR PROCEDURE WITH A 29 MM SAPIEN 3 ULTRA RESILIA VALVE. AFTER THE PROCEDURE, THE PATIENT DEVELOPED A CARDIOEMBOLIC STROKE, ALTHOUGH THE EXACT OCCURRENCE DATE IS UNKNOWN. SUBSEQUENTLY, LOW BLOOD PRESSURE MADE IT DIFFICULT TO CONTINUE DIALYSIS. DURING THE COURSE OF TREATMENT, THE PATIENT EXPERIENCED SEPSIS AND INTESTINAL ISCHEMIA AND EXPIRED ON POSTOPERATIVE DAY 41. ACCORDING TO THE MEDICAL OPINION, THE DEATH WAS PROCEDURE-RELATED, NOT DEVICE-RELATED. THE CAUSE OF DEATH WAS REPORTED AS ''INFECTION.''Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES JAPANESE AFFILIATE, THE PATIENT UNDERWENT A TRANS-AXILLARY OR TRANS-SUBCLAVIAN TAVR PROCEDURE WITH A 29 MM SAPIEN 3 ULTRA RESILIA VALVE. ON POST-OPERATIVE DAY 25, THE PATIENT EXPERIENCED A CARDIOGENIC CEREBRAL EMBOLISM AND MECHANICAL CIRCULATORY SUPPORT WAS INITIATED. THIS SUPPORT IS SUSPECTED TO BE THE CAUSE OF THE CARDIOGENIC EMBOLISM. LOW BLOOD PRESSURE MADE IT DIFFICULT TO CONTINUE DIALYSIS. DURING THE COURSE OF TREATMENT, THE PATIENT EXPERIENCED SEPSIS AND INTESTINAL ISCHEMIA AND EXPIRED ON POSTOPERATIVE DAY 41. ACCORDING TO THE MEDICAL OPINION, THE DEATH WAS PROCEDURE-RELATED, NOT DEVICE-RELATED. THE CAUSE OF DEATH WAS REPORTED AS ''INFECTION.''Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES JAPANESE AFFILIATE, THE PATIENT UNDERWENT A TRANS-AXILLARY OR TRANS-SUBCLAVIAN TAVR PROCEDURE WITH A 29 MM SAPIEN 3 ULTRA RESILIA VALVE. AFTER THE PROCEDURE, THE PATIENT DEVELOPED A CARDIOEMBOLIC STROKE, ALTHOUGH THE EXACT OCCURRENCE DATE IS UNKNOWN. SUBSEQUENTLY, LOW BLOOD PRESSURE MADE IT DIFFICULT TO CONTINUE DIALYSIS. DURING THE COURSE OF TREATMENT, THE PATIENT EXPERIENCED SEPSIS AND INTESTINAL ISCHEMIA AND EXPIRED ON POSTOPERATIVE DAY 41. ACCORDING TO THE MEDICAL OPINION, THE DEATH WAS PROCEDURE-RELATED, NOT DEVICE-RELATED. THE CAUSE OF DEATH WAS REPORTED AS ''INFECTION.'' Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRAPROCEDURAL EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING< 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT FACTORS, SUCH AS DIABETES, MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, AND ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: ADDITION TO B.5. CORRECTION TO B.3. Manufacturer Narrative: CORRECTION MADE TO B.3. ADDITION MADE TO B.5.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21989294&pc=DSQ | 2916596-2025-02713 | 2025-04-13 04:00 | Death | THORATEC CORPORATION | 2025-05-08 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI ORGAN FAILURE. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21989353&pc=QGM | 3027664504-2025-00006 | 2025-04-07 04:00 | Death | HISTOSONICS, INC. | 2025-05-08 04:00 | QGM | EDISON SYSTEM | Adverse Event Without Identified Device or Use Problem | Adult Respiratory Distress Syndrome; Cardiac Arrest; Respiratory Acidosis; Thrombosis/Thrombus | K233466 |  | 1 | Event Description: AS PART OF THE BOOMBOX POST-MARKET CLINICAL TRIAL, PATIENT ID (B)(6), DIED DUE TO WHAT THE TREATING PHYSICIAN BELIEVES TO BE A THROMBOEMBOLIC EVENT POSSIBLY RELATED TO THE HISTOTRIPSY DEVICE AND HISTOTRIPSY PROCEDURE PERFORMED ON (B)(6) 2025. AN 87 YEAR OLD MALE PATIENT WITH A HISTORY OF SIGMOID COLON CANCER, WHICH HAD BEEN RESECTED 4 YEARS PRIOR, RECEIVED HISTOTRIPSY TREATMENT ON (B)(6) 2025 TO 2 LESIONS IN SEGMENTS V (X=25 MM, Y=30 MM, Z = 30 MM, PTV = 29.3 CC) AND VI (X=18 MM, Y=22 MM, Z = 14.5 MM, PTV = 11.5 CC) FOR A TOTAL PLANNED TREATMENT VOLUME (PTV) 40.8 CC. PER THE TREATING PHYSICIAN, NO ANTI-COAGULATION THERAPY WAS PROVIDED PRIOR TO THE HISTOTRIPSY PROCEDURE. THE PATIENT HAD PREVIOUSLY UNDERGONE SBRT IN (B)(6) 2023, WHICH POTENTIALLY OVERLAPPED WITH THE HISTOTRIPSY PTVS. DURING THE INTRAOPERATIVE COURSE, TWO SYSTEM ERRORS OCCURRED WHICH REQUIRED A SYSTEM RESET RESULTING IN A CUMULATIVE DELAY OF APPROXIMATELY 30 MINUTES. THE TOTAL PROCEDURE TIME UNDER ANESTHESIA WAS APPROXIMATELY 3.5 HOURS. PER THE TREATING PHYSICIAN, THE PATIENT'S IMMEDIATE POST PROCEDURE RECOVERY WAS UNEVENTFUL. HE EXHIBITED NORMAL NEUROLOGICAL FUNCTION, WAS ALERT, AMBULATING, AND ABLE TO TOLERATE ORAL INTAKE. HE WAS TRANSFERRED TO THE POST-ANESTHESIA CARE UNIT (PACU) IN STABLE CONDITION WITH PLANS FOR POSTOPERATIVE IMAGING, INCLUDING A CT SCAN. WHILE BEING TRANSPORTED FOR IMAGING, THE PATIENT SUFFERED A CARDIAC ARREST. RESUSCITATION EFFORTS WERE EXTENSIVE, INVOLVING 30 MINUTES OF CARDIOPULMONARY RESUSCITATION (CPR) AND THE ADMINISTRATION OF THROMBOLYTIC THERAPY. RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED; HOWEVER, THE PATIENT SUBSEQUENTLY DEVELOPED ACUTE RESPIRATORY DISTRESS SYNDROME (ARDS). THE PATIENT THEN EXPERIENCED REFRACTORY METABOLIC ACIDOSIS. DESPITE AGGRESSIVE CRITICAL CARE INTERVENTIONS, THE PATIENT SHOWED NO SIGNS OF RECOVERING, AND FOLLOWING DISCUSSIONS WITH THE FAMILY, LIFE-SUSTAINING TREATMENTS WERE WITHDRAWN. THE PATIENT WAS PRONOUNCED DECEASED ON (B)(6) 2025. A TRANSTHORACIC ECHOCARDIOGRAM PERFORMED DURING THE RESUSCITATION PERIOD DEMONSTRATED PRESERVED CARDIAC FUNCTION. NO POSTOPERATIVE IMAGING WAS OBTAINED. IT WAS NOTED THAT THE PATIENT HAD NO HISTORY OF ARRHYTHMIA. FOR TRIAL PURPOSES, THE EVENT WAS CLASSIFIED BY THE TREATING PHYSICIAN AS POSSIBLY RELATED TO THE DEVICE AND POSSIBLY RELATED TO THE PROCEDURE. PER THE TREATING PHYSICIAN, THE MOST PROBABLE CONTRIBUTING FACTORS WERE PERIOPERATIVE IMMOBILITY AND THE ABSENCE OF HEPARIN PROPHYLAXIS. GIVEN THE PATIENT'S CLINICAL COURSE AND KNOWN RISKS, THIS EVENT WAS CONSIDERED ANTICIPATED. THIS MDR IS BEING SUBMITTED OUT OF AN ABUNDANCE OF CAUTION DUE TO THE PATIENT'S DEATH POST-PROCEDURE ALTHOUGH NO DEFINITIVE CAUSE OF DEATH WAS DETERMINED. NO OTHER DEVICE MALFUNCTIONS OR OTHER NOTEWORTHY EVENTS OCCURRED DURING THE CASE. Manufacturer Narrative: NO OTHER DEVICE MALFUNCTIONS OR OTHER NOTEWORTHY EVENTS OCCURRED DURING THE CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21989633&pc=LWS | 2649622-2025-12580 | 2025-04-20 04:00 | Death | MPRI | 2025-05-08 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing; High Capture Threshold | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) FOR SENSING INTEGRITY COUNTER (SIC) AND HIGH RATE NON-SUSTAINED EPISODES. THE LEAD DISPLAYED HIGH THRESHOLDS AND INTERMITTENT OVERSENSING DURING AN EPISODE OF VENTRICULAR FIBRILLATION (VF). THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANTED: (B)(6) 2016. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21990941&pc=NAY | 2955842-2025-19893 | 2025-04-10 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-05-08 04:00 | NAY | DAVINCI XI | Insufficient Information | Hemorrhage/Bleeding; Insufficient Information | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PARAESOPHAGEAL HIATAL HERNIA REPAIR, BLEEDING OCCURRED. DURING CONVERSION TO AN OPEN PROCEDURE THE PATIENT CODED AND ULTIMATELY EXPIRED. THE SURGEON STATED THAT THE DA VINCI PRODUCTS USED DURING THIS EVENT HAD NOTHING TO DO WITH THE BLEEDING AND PATIENT DEATH. THE SURGEON STATED THAT THE BLEEDING AND SUBSEQUENT DEATH WERE ENTIRELY CAUSED BY THE PATIENT'S CONDITION AND THE STATUS OF HER BODY AT THE TIME OF THE SURGERY. THE SURGEON STATED THAT HE COULD NOT PROVIDE ANY ADDITIONAL PROCEDURE OR PATIENT SPECIFIC INFORMATION. Manufacturer Narrative: A REVIEW OF THE DA VINCI SYSTEM LOGS FOUND THAT NO ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, VESSEL SEALER EXTEND, FENESTRATED BIPOLAR FORCEPS, TIP-UP FENESTRATED GRASPER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR ANY OF THE INSTRUMENTS USED DURING THIS OR SUBSEQUENT PROCEDURES. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AFTER THE SURGEON ENCOUNTERED BLEEDING DURING THE CASE. ALTHOUGH THE SURGEON STATED THE PATIENT BASELINE CONDITION MAY HAVE CONTRIBUTED, HOW THE BLEEDING OCCURRED AND WHAT SPECIFICALLY WAS BLEEDING WAS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21975229&pc=MVK | 3008642652-2025-04571 | 2025-04-23 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-07 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 21:28:37 ON (B)(6) 2025. AT 03:04:57 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 130 BPM. AT 03:05:35, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 150 BPM. POST SHOCK RHYTHM WAS ASYSTOLE TRANSITIONING TO AN IDIOVENTRICULAR RHYTHM FROM 10-70 BPM. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE ELECTRODE BELT WAS DISCONNECTED AT 03:46:13 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21975741&pc=NVZ | 3008973940-2025-03131 | 2025-04-11 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-05-07 04:00 | NVZ | ADAPTA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. DURING IMPLANT THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) EXHIBITED HIGH THRESHOLDS IN VARIOUS POSITIONS. THE LEADLESS IPG WAS NOT USED AND REPLACED WITH A SINGLE CHAMBER TRANSVENOUS PACING SYSTEM. FOLLOWING THIS THE PATIENT DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21975763&pc=DTB | 3008973940-2025-03133 | 2025-04-11 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-05-07 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. DURING IMPLANT THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) EXHIBITED HIGH THRESHOLDS IN VARIOUS POSITIONS. THE LEADLESS IPG WAS NOT USED AND REPLACED WITH A SINGLE CHAMBER TRANSVENOUS PACING SYSTEM. FOLLOWING THIS THE PATIENT DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21975849&pc=DTB | 21975849 | 2025-03-28 04:00 | Death | MEDTRONIC, INC. | 2025-05-07 04:00 | DTB | CAPSUREFIX NOVUS MRI¿ SURESCAN® | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade |  |  | 1 | Event Description: TAMPONADE AND ARREST WITH MEDTRONIC HARDWARE IN THE POSTOP PHASE. IT APPEARS ALL OF THESE PATIENTS HAD LEAD #5076 IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21975984&pc=MUE | MW5170004 | 2025-04-10 04:00 | Death | HOLOGIC, INC. | 2025-05-07 04:00 | MUE | GE, SIEMENS AND HOLOGIC MAMMOGRAPHIC TOMOSYNTHESIS SYSTEM | Adverse Event Without Identified Device or Use Problem | Abdominal Pain; Hemorrhage/Bleeding; Intracranial Hemorrhage; Low Blood Pressure/ Hypotension; Hypoxia; Shock; Brain Injury; Fungal Infection; Cancer; Cerebral Edema; Hematemesis; Melena |  |  | 1 | Event Description: PATIENT HAS A HISTORY OF STAGE IV METASTATIC RCC. PATIENT RESTARTED CHEMO (LENVATINIB AND EVEROLIMUS) THE WEEK OF (B)(6) 2025. PATIENT STARTED HAVING ABDOMINAL PAIN, SO TOOK LAST DOSE ON (B)(6) 2025. PATIENT THOUGHT SHE STARTED SEEING BLOOD IN STOOL ON (B)(6) 2025, BUT ON (B)(6) 2025 PROGRESSED TO LARGER QUANTITIES. PATIENT ARRIVED AT (B)(6) EMERGENCY DEPARTMENT. PRE-TRANSFUSION HGB WAS 6.2 AND WAS GIVEN LU PRBC. PATIENT WAS TRANSFERRED TO (B)(6) (MEDICAL INTENSIVE CARE UNIT) FOR MANAGEMENT OF UNDIFFERENTIATED SHOCK. UPON TRANSFER TO (B)(6) SHE WAS HDS AND DID NOT REQUIRE FURTHER VASOPRESSORS. STARTED ON BROAD SPECTRUM ANTIBIOTICS. STABILIZED AND TRANSFERRED TO GOLD SERVICE. INTERVENTIONAL GI CONSULTED FOR HER GIB AS HAS HX OF AMPULLARY MASS S/P DUODENAL AND BILIARY STENTS IN (B)(6) AS WELL. AT THAT TIME, AS PT WAS STABLE, RECOMMENDED CLOSE MONITORING. HOWEVER, HAD RECURRENT MELENA, SO UNDERWENT EGD AND C-SCOPE ON (B)(6) FOR RECURRENT MELENA WHICH SHOWED FRIABLE MUCOSA OF MALIGNANT STRICTURE IN DUODENUM, OLD BLOOD/CLOTS IN R>L COLON BUT NO ACTIVE BLEEDING NOTED. FOUND TO HAVE CANDIDEMIA (BLOOD EX 4/1) SO PLACED ON MICAFUNGIN. TTE GROSSLY NORMAL & W/OUT VEGETATION. PORT REMOVED BY IR FOR LINE HOLIDAY ON (B)(6). PALLIATIVE FOLLOWING FOR PAIN MANAGEMENT. RETURNED TO MICU (B)(6) FOR RECURRENT SHOCK, NEW 02 REQUIREMENT & ONE EPISODE OF HEMATEMESIS. INITIALLY IMPROVED W/ IV FLUIDS, PRBC. FOUND TO HAVE NEW BACTEREMIA W/ E. FAECALIS, STARTED ON ZOSYN. PT & FAMILY DISCUSSING POSSIBLE TRANSITION TO HOME HOSPICE AT THAT TIME. HOWEVER, ON (B)(6), PATIENT BECAME HEMODYNAMICALLY UNSTABLE (3 PRESSORS) AFTER MULTIPLE EPISODES OF MELENA, LA UP TO 6. GI RE-ENGAGED. RAD ONE ALSO CONSULTED FOR PALLIATIVE XRT FOR BLEEDING DUODENAL MASS. PT UNDERWENT EGD ON (B)(6) WHICH SHOWED DIFFUSE OOZING OF BLOOD FROM FRIABLE MALIGNANT TISSUE TREATED W/ HEMOSTATIC POWDER. WHILE IN PROCEDURE, PT BECAME HYPOXIC AND HYPOTENSIVE. SHE WAS INTUBATED AND TRANSFERRED BACK TO MICU. AROUND 3PM ON (B)(6) 2025, PT'S PUPILS FOUND TO HAVE ACUTELY FIXED/DILATED PUPILS W/ ABSENT BRAINSTEM REFLEXES. STROKE CODE CALLED. CT BRAIN W/ DIFFUSE ANOXIC INJURY, CEREBRAL EDEMA, PNEUMOCEPHALUS AND POSSIBLE ICH (INTRACEREBRAL HEMORRHAGE). AFTER LEARNING THIS INFORMATION, FAMILY MADE DECISION TO CHANGE CODE STATUS TO DNR WITH PLAN TO TRANSITION TO COMFORT-FOCUSED CARE AFTER OTHER FAMILY MEMBERS HAS BEEN ABLE TO VISIT. ON (B)(6) 2025, PT WAS TRANSITIONED TO COMFORT-FOCUSED CARE AND PALLIATIVELY EXTUBATED. SHE PASSED AWAY AT 1:28PM SURROUNDED BY FAMILY. AUTOPSY WAS DECLINED. Event Description: ADDITIONAL INFORMATION RECEIVED FROM REPORTER ON 5/8/2025 FOR REPORT MW5170004. TOMOSYNTHESIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21975986&pc=OZD | 1220648-2025-28155 | 2025-04-16 04:00 | Death | ABIOMED, INC. | 2025-05-07 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT AN IMPELLA CP WAS IMPLANTED IN A PATIENT WITH ST-ELEVATED MYOCARDIAL INFARCTION AND A PLACED STENT. DURING SUPPORT, A HEMATOMA DEVELOPED AT THE IMPELLA ACCESS SITE. PRESSURE WAS HELD FOR 30 MINS FOLLOWED BY FEMO STOP PLACEMENT. AGGRASTAT WAS GIVEN AND HEPARIN WAS WITHHELD. IN ADDITION, THE PATIENT HAD DETERIORATED AND WAS REQUIRING INCREASED OXYGEN NEEDS AND PRESSER SUPPORT. CONTINUOUS RENAL REPLACEMENT THERAPY WAS STARTED IN RESPONSE TO ANURIA. EXTRACORPOREAL MEMBRANE OXYGENATION CANNULATION WAS PROVIDED DUE TO THE PATIENT'S OXYGENATION STATUS. HOWEVER, THE ARTERIAL CANNULA WAS NOT SECURED AND SOMEHOW FELL OUT RESULTING IN VERY SIGNIFICANT BLOOD LOSS. THE ARTERY WAS RE-CANNULATED AND THE VESSEL WAS REPAIRED AROUND THE CANNULA, AND MULTIPLE BLOOD PRODUCTS WERE ADMINISTERED. ULTIMATELY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION ACCESS SITE BLEEDING AND VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER, ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21976005&pc=DTZ | MW5170005 | 2025-04-12 04:00 | Death | UNK | 2025-05-07 04:00 | DTZ | EXTRACORPOREAL MEMBRANE OXYGENATION | Adverse Event Without Identified Device or Use Problem | Ischemia; Renal Failure; Respiratory Failure |  |  | 1 | Event Description: A 68 Y/O MALE ON PIRFENIDONE FOR IPF. UPON ROUTINE CHART REVIEW, PROVIDER NOTED PT ADMITTED (B)(6) 2025 FOR EXPEDITED LUNG TRANSPLANT EVALUATION, DEVELOPED DIC FROM ECMO COMPLICATIONS OF LIMB ISCHEMIA AND KIDNEY FAILURE, DECLARED INELIGIBLE FOR TRANSPLANT. PT AND FAMILY ELECTED DNR (DO NOT RESUSCITATE) AND COMFORT CARE. PT EXPIRED (B)(6) 2025 ON COMFORT MEASURES WITH FAMILY AT BEDSIDE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21976056&pc=MOZ | MW5170006 | 2025-04-13 04:00 | Death | SANOFI GENZYME CORP/GENZYME CORPORATION | 2025-05-07 04:00 | MOZ | SYNVISC PFS 8MG/ML (3X2ML) | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: PT DECEASED. CALLER'S NAME: (B)(6). CALLER'S RELATIONSHIP TO THE CUSTOMER: DAUGHTER. CUSTOMER'S DECEASED DATE: (B)(6) 2025. REASON OF DEATH UNK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21976066&pc=CBK | 2518422-2025-105292 | 2025-05-01 04:00 | Death | RESPIRONICS, INC. | 2025-05-07 04:00 | CBK | TRILOGY EVO, JAPAN | Adverse Event Without Identified Device or Use Problem | Unspecified Respiratory Problem | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A DEVICE WAS COLLECTED BECAUSE THE PATIENT PASSED AWAY. THE DEVICE EXTERIOR HAD TURNED BLACK AND TOBACCO SMELLS WERE OBSERVED. A REQUEST HAS BEEN MADE FOR REPAIR OF THE DEVICE. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED. THE REQUEST IS FOR THE REPAIR OF THE DEVICE NOT RELATED TO THE PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION ALLEGING A DEVICE WAS COLLECTED BECAUSE THE PATIENT PASSED AWAY. THE DEVICE EXTERIOR HAD TURNED BLACK AND TOBACCO SMELLS WERE OBSERVED. A REQUEST HAS BEEN MADE FOR REPAIR OF THE DEVICE. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED. THE REQUEST IS FOR THE REPAIR OF THE DEVICE NOT RELATED TO THE PATIENT DEATH. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND A LIFE-RELATED SMELL (TOBACCO SMELL) WAS CONFIRMED TO BE FOUND ON THE MUFFLER ASSEMBLY. THE MUFFLER ASSEMBLY WAS REPLACED AND THE DEVICE WAS CLEANED. DEVICE PASSED TESTING AND CONFIRMED TO OPERATE PROPERLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21976106&pc=OZD | 1220648-2025-28156 | 2025-04-16 04:00 | Death | ABIOMED, INC. | 2025-05-07 04:00 | OZD | IMPELLA CP | Malposition of Device | Arrhythmia; Hemolysis; Hematuria | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP HAD HEMATURIA AND HEMOLYZED LABS, AND THE IMPELLA WAS REPOSITIONED. THE PATIENT ALSO HAD SUSTAINED VENTRICULAR TACHYCARDIA (SVT) AND WAS SUCCESSFULLY CARDIOVERTED. FURTHERMORE, THERE WERE CRITICAL IN-AORTA ALARMS AND IMPELLA DISCONNECTED ALARMS. CPR WAS THEN IN PROGRESS FOR PEA ARREST. THE PUMP WAS THEN WEANED AND EXPLANTED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER, ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE, ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS, ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21976116&pc=DYB | 21976116 | 2025-02-10 05:00 | Death | MEDTRONIC, INC. | 2025-05-07 04:00 | DYB | MICRA INTRODUCER | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IV ANCEF 2GM AND IV EPINEPHRINE 1MG (4 DOSES) WERE ADMINISTERED. MANUFACTURER RESPONSE FOR MICRA INTRODUCER SHEATH WITH HYDROPHILIC COATING, MICRA INTRODUCER SHEATH WITH HYDROPHILIC COATING (PER SITE REPORTER). UNKNOWN AT THIS TIME.Event Description: IV ANCEF 2GM AND IV EPINEPHRINE 1MG (4 DOSES) WERE ADMINISTERED. MANUFACTURER RESPONSE FOR MICRA INTRODUCER SHEATH WITH HYDROPHILIC COATING, MICRA INTRODUCER SHEATH WITH HYDROPHILIC COATING (PER SITE REPORTER). UNKNOWN AT THIS TIME.Event Description: IV ANCEF 2GM AND IV EPINEPHRINE 1MG (4 DOSES) WERE ADMINISTERED. MANUFACTURER RESPONSE FOR MICRA INTRODUCER SHEATH WITH HYDROPHILIC COATING, MICRA INTRODUCER SHEATH WITH HYDROPHILIC COATING (PER SITE REPORTER): UNKNOWN AT THIS TIME. Manufacturer Narrative: THE FOLLOWING ELEMENTS HAVE BLANK DATA: [PATIENT'S AGE AT TIME OF EVENT:] FOR PATIENT IDENTIFIER: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21976268&pc=OZD | 1220648-2025-28160 | 2025-04-15 04:00 | Death | ABIOMED, INC. | 2025-05-07 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT ON IMPELLA 5.5 SUPPORT HAD RUNS OF VENTRICULAR TACHYCARDIA REQUIRING TWO SHOCKS. ON DAY EIGHT OF 5.5 SUPPORT, THE PATIENT WENT BRADYCARDIC REQUIRING ATROPINE ADMINISTRATION. IMPELLA REMAINED ON P-4 WITH NO ALARMS. ADDITIONALLY, IMPELLA KEPT HAVING SUCTION EVENTS WHICH REQUIRED LOWERING THE P-LEVEL. THE PATIENT WAS MADE A "DO NOT RESUSCITATE" AND EXPIRED ON IMPELLA SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/ AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. LOW PUMP FLOW: CLINICAL STATES THAT MULTIPLE TIMES DURING THE CASE, WHEN THE PUMP WAS ATTEMPTED TO GO UP IN P-LEVEL, SUCTION ALARMS WERE OBSERVED AND THE PUMP HAD TO BE REDUCED TO THE LOWER P-LEVEL. CLINICAL ALSO STATES THAT THE PATIENT HAD PULMONARY EDEMA AND PATIENT WAS GOING THROUGH DIALYSIS. PUMP WAS NOT RETURNED FOR INVESTIGATION. DATA LOGS WERE RETURNED AND SUCTION ALARMS WERE PRESENT AND OBSERVED CORRESPONDING TO THE P-LEVEL CHANGE ATTEMPT. PLACEMENT SIGNAL WAS SEEN TRENDING DOWNWARDS. NO MOTOR CURRENT SPIKES OR PURGE ISSUES WERE OBSERVED. THEREFORE, BASED ON THE DATA LOG REVIEW AND CLINICAL INFORMATION PROVIDED, THE ROOT CAUSE OF THE LOW FLOW ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED. VT/ VF: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED B2 REQUIRED INTERVENTION SELECTION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28160. H6 HEALTH EFFECT - IMPACT CODE 4647 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28160. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21976536&pc=OZD | 1220648-2025-28165 | 2025-04-17 04:00 | Death | ABIOMED, INC. | 2025-05-07 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT UPON PLACING THE IMPELLA 5.5 INTO THE GRAFT THE DOCTOR PERFORMED AN ARTERIOTOMY TO HELP PLACE THE IMPELLA ACROSS THE AORTIC VALVE. ONCE ACROSS THE VALVE AND WHAT WAS VISUALIZED ON TRANSESOPHAGEAL ECHOCARDIOGRAM SEEM TO BE DEEP AND THE PHYSICIAN WAS INSTRUCTED BY THE ANESTHESIOLOGIST TO ROTATE THE IMPELLA 5.5 AND NOT TO ADVANCE THE IMPELLA 5.5 BUT TO ROTATE THE IMPELLA 5.5 TO ORIENT THE INLET TO THE APEX OF THE LEFT VENTRICLE (LV). THE PHYSICIAN ELECTED TO ADVANCE WHILE ROTATING THE DEVICE THROUGH THE ARTERIOTOMY SITE AND IN THE CHEST THROUGH THE GRAFT RESULTING IN THE PERFORATION OF THE LV. THE PHYSICIAN THEN RETRACTED THE IMPELLA 5.5 AND FIXED THE PERFORATION. UPON REATTEMPTING TO PLACE THE IMPELLA 5.5 ACROSS FOR THE SECOND TIME THE IMPELLA 5.5 AGAIN APPEARED TO BE IN IDEAL DEPTH AND THE DOCTOR WAS INSTRUCTED TO ROTATE THE IMPELLA 5.5 AND THE PHYSICIAN AGAIN ADVANCED AND ROTATED USING THE SAME EXACT TECHNIQUE AS BEFORE AND PERFORATED THE LV FOR A SECOND TIME. THE DOCTOR AGAIN RETRACTED THE IMPELLA AND FIXED THE PERFORATION. ON THE THIRD ATTEMPT, THE PHYSICIAN SEATED THE IMPELLA 5.5 SHALLOW IN THE LV, AND ELECTED NOT TO ROTATE THE IMPELLA 5.5. THE PATIENT WAS THEN STARTED ON IMPELLA 5.5 SUPPORT. BYPASS WAS WEANED, GRAFT WAS TRIMMED AND PEEL-AWAY SHEATH WAS REMOVED. THE IMPELLA 5.5 WAS SECURED TO THE UPPER CHEST. HOWEVER, THE PATIENT'S CHEST REMAINED OPEN. CARE WAS ULTIMATELY WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF VENTRICLE PERFORATION HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. ACCORDING TO THE CLINICAL NOTES, THE LEFT VENTRICLE WAS FOUND PERFORATED TWO TIMES WHILE USING THE TECHNIQUE TO ADVANCE AND ROTATE THE PUMP. ON THE THIRD ATTEMPT, THE DOCTOR OPTED TO LEAVE THE PUMP IN A SHALLOW POSITION AND PROCEEDED WITH THE CASE. THE PUMP WAS UTILIZED FOR 9 DAYS FOLLOWING THE LEFT VENTRICLE (LV) REPAIR AND THIRD PLACEMENT. THE CAUSE OF THE VENTRICLE PERFORATION ISSUE WAS MOST LIKELY USE RELATED, BASED ON GIVEN CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21976771&pc=OZD | 1220648-2025-28167 | 2025-04-17 04:00 | Death | ABIOMED, INC. | 2025-05-07 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT BEFORE AN IMPELLA CP WAS IMPLANTED, A 12 FRENCH SHEATH WAS FIRST PLACED TO PERFORM VALVULOPLASTY. DURING THIS TIME, THE PATIENT BECAME HEMODYNAMICALLY CRITICAL. THE VALVULOPLASTY WAS PERFORMED SUCCESSFULLY, FOLLOWED BY THE 14 FRENCH PEEL-AWAY SHEATH, AND THE IMPELLA CP WAS IMPLANTED. AFTER THE IMPELLA CP WAS PLACED, IT WAS DISCOVERED THAT THE PATIENT WAS BLEEDING FROM THE FEMORAL ARTERY. A COVERED STENT WAS INSERTED INTO THE FEMORAL ARTERY AND THE IMPELLA CP WAS REMOVED. THE PATIENT EXPIRED IN THE PROCEDURAL AREA. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21976830&pc=QZI | 3012520654-2025-00050 | 2025-04-17 04:00 | Death | MEDTRONIC, INC. | 2025-05-07 04:00 | QZI | AFFERA | Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock; Multiple Organ Failure | P240013 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A COMPLETED CARDIAC ABLATION PROCEDURE, THE PATIENT EXPERIENCED SEVERE HEMODYNAMIC FAILURE AND SEVERE ACIDOSIS. THE PATIENT WAS DIAGNOSED WITH CARDIOGENIC SHOCK AND METABOLIC ACIDOSIS. THE PATIENT WAS NOT ELIGIBLE FOR LEFT VENTRICULAR ASSIST DEVICE (LVAD) AND THE PATIENT REFUSED THERAPEUTIC OBSTINATION. TESTING INCLUDED BLOOD TESTS REVEALING METABOLIC ACIDOSIS AND ECHOCARDIOGRAM RESULTS INDICATING SEVERE BIVENTRICULAR FAILURE. X-RAYS CONDUCTED SHOWED SUSPICION OF ATELECTASIS. MEDICAL INTERVENTIONS INCLUDED THE ADMINISTRATION OF BLOOD PRESSURE SUPPORT, A DIURETIC, ANTIBIOTICS, A NON-DEPOLARIZING NEUROMUSCULAR-BLOCKING DRUG AND A SUBSTANCE BLOCKING THE TRANSMISSION OF NERVE IMPULSES TO MUSCLES, CAUSING PARALYSIS. ADDITIONAL ACTIONS TAKEN INCLUDED THE IMPLANTATION OF LEADS AND THE USE OF AN EXTERNAL STIMULATION SYSTEM TO IMPROVE CARDIAC CONTRACTION AND HEMODYNAMIC FAILURE, AS WELL AS THE PLACEMENT OF A HEMOFILTRATION CATHETER. THE PATIENT'S HOSPITALIZATION WAS EXTENDED. DESPITE A SLOW INITIAL IMPROVEMENT, THE PATIENT'S CONDITION WORSENED, RESULTING IN MULTIVISCERAL FAILURE AND ULTIMATELY NON-SUDDEN CARDIAC DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE CARDIAC ABLATION PROCEDURE WAS FOR VENTRICULAR TACHYCARDIA (VT). Manufacturer Narrative: UPDATED B5 EVENT DESCRIPTION UPDATED E. INITIAL REPORTER MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21977033&pc=NPT | 2015691-2025-03615 | 2025-04-14 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-07 04:00 | NPT | COMMANDER DELIVERY SYSTEM | Burst Container or Vessel; Difficult to Remove; Detachment of Device or Device Component; Insufficient Information | Insufficient Information | P140031 |  | 1 | Event Description: AS REPORTED, IT WAS A CASE OF AN IMPLANT OF A 23MM SAPIEN 3 ULTRA VALVE, IN AORTIC POSITION BY TRANSFEMORAL APPROACH. DURING PROCEDURE, THE BALLOON OF THE COMMANDER DELIVERY SYSTEM BURST DURING INFLATION. THE VALVE WAS FULLY DEPLOYED WHEN THE BALLOON BURST. THERE WAS NO RESIDUAL LEAK NOTED BY TTE AND ANGIOGRAPHY, WITH A GRADIENT MEASURED AT 5MMHG. HOWEVER, THERE WERE DIFFICULTIES ENCOUNTERED DURING THE WITHDRAWAL OF THE DELIVERY SYSTEM THROUGH THE ESHEATH+ AND REMAINED STUCK IN THE FEMORAL-ILIAC VESSEL. A CUT-DOWN WAS PERFORMED. THE REMOVAL OF THE DELIVERY SYSTEM WAS VERY COMPLICATED, AND THE DISTAL RADIOPAQUE MARKER OF THE ESHEATH+ EMBOLIZED. THIS EVENT RESULTED IN A PROLONGED HOSPITALIZATION OF THE PATIENT WITH SEVERE POST OPERATIVE COMPLICATIONS, AND UNFORTUNATELY, THE PATIENT EXPIRED ONE DAY POST-PROCEDURE. Manufacturer Narrative: THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. PLEASE REFERENCE RELATED MANUFACTURER REPORT NO: 2015691-2025-03616. A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR CORRECTION AND ADDITIONAL INFORMATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: CORRECTED H.3 DEVICE EVALUATED BY MANUFACTURER AND H.6 INVESTIGATION CONCLUSIONS AND INVESTIGATION FINDINGS. ADDED NEW INFORMATION TO D.9 DATE RETURNED TO MANUFACTURER AND H.6 DEVICE CODES AND TYPE OF INVESTIGATION. THE DEVICE WAS RETURNED TO EDWARDS LIFESCIENCES FOR EVALUATION AND THE FOLLOWING WAS OBSERVED. BALLOON MATERIAL BUNCHED TOWARDS NOSE TIP. RADIAL AND LONGITUDINAL BURST. NO APPARENT BALLOON MATERIAL MISSING. NOSE TIP MATERIAL MISSING. IMAGERY WAS NOT PROVIDED FOR REVIEW. A REVIEW OF THE RISK MANAGEMENT DOCUMENTATION WAS PERFORMED, AND NO EVIDENCE OF PRODUCT NON-CONFORMANCES OR LABELING/IFU INADEQUACIES WERE IDENTIFIED IN THE EVALUATION. A REVIEW OF THE DHR AND LOT HISTORY DID NOT PROVIDE ANY INDICATION THAT A MANUFACTURING NON-CONFORMANCE WOULD HAVE CONTRIBUTED TO THE COMPLAINT EVENT. A REVIEW OF THE IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. FURTHERMORE, NO ABNORMALITIES WERE NOTED DURING DEVICE UNPACKING OR PREPARATION. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, BALLOON BURST EVENTS DURING VALVE DEPLOYMENT HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO CALCIFICATION IN THE LANDING ZONE OR OVER-INFLATION OF THE BALLOON. THE COMPLAINTS WERE CONFIRMED BASED ON EVALUATION OF RETURNED DEVICE. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION) LIKELY CONTRIBUTED TO THE EVENT AS PROVIDED INFORMATION INDICATED PRESENCE OF CALCIFICATION AT THE LANDING ZONE. THE PRESENCE OF CALCIFICATION CAN CREATE A CHALLENGING ANATOMY FOR BALLOON INFLATION. WHILE THE BALLOONS ARE SUFFICIENTLY DESIGNED AND TESTED TO ENSURE THE BURST PRESSURE IS AT OR ABOVE THE RATED BURST PRESSURE, CALCIFIED NODULES CAN COMPROMISE THE STRUCTURE OF THE BALLOON WALL VIA FOLLOWING MECHANISMS SUCH AS PUNCTURE, LOCAL OVERSTRETCHING, OPEN CELL IMPINGEMENT, OR STRESS CONCENTRATION. AVAILABLE INFORMATION ALSO SUGGESTS THAT PROCEDURAL FACTORS (WITHDRAWAL OF AN ALTERED BALLOON PROFILE AND DEVICE MANIPULATION) LIKELY CONTRIBUTED TO THE EVENT, AS EVALUATION OF THE RETURNED DEVICE SHOWED THE BALLOON MATERIAL LINER BUNCHED TOWARDS NOSE TIP. AS THE BALLOON WAS BURST, THE ALTERED BALLOON PROFILE COULD HAVE MADE IT MORE SUSCEPTIBLE TO CATCH ON PATIENT VASCULATURE AND/OR SHEATH'S DISTAL END, WHICH WOULD HAVE THEN LED TO THE EXPERIENCED RETRIEVAL DIFFICULTY. IN THIS CASE, A CUT DOWN WAS PERFORMED IN ORDER TO WITHDRAW THE DELIVERY SYSTEM WITH AN ALTERED BALLOON PROFILE (BURST BALLOON). IT IS POSSIBLE THAT THE BURST BALLOON COULD BE REMOVED, BY PULLING THE NOSE TIP USING A CLAMP DURING THE SURGICAL CUT DOWN PROCEDURE, AND IF ADDITIONAL PULL FORCE IS APPLIED IT CAN FURTHER DAMAGE THE NOSE TIP, AND POTENTIALLY RESULT IN THE OBSERVED SEPARATION. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE THERE IS NO PRODUCT NON-CONFORMANCES OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: REFERENCE NO. (B)(4). THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21977034&pc=MKJ | 1220908-2025-01731 | 2025-04-16 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-05-07 04:00 | MKJ | AED PLUS | Tear, Rip or Hole in Device Packaging; Unintended Electrical Shock | Cardiac Arrest; No Clinical Signs, Symptoms or Conditions | P160015 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A 38-YEAR-OLD FEMALE PATIENT, THE ELECTRODE PADS WERE DAMAGED DURING REMOVAL FROM THE PACKAGING, THE PADS WERE APPLIED TO THE PATIENT, AND THE CLINICIAN SUSTAINED AN ELECTRICAL SHOCK FROM THE DEVICE WHILE PERFORMING CPR. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. COMPLAINANT INDICATE THAT THERE WAS NO FURTHER DAMAGE AFTER A SHORT STAY IN THE HOSPITAL. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. THE CUSTOMER'S REPORT OF ELECTRODE DAMAGED DURING REMOVAL WAS CONFIRMED BY A PICTURE PROVIDED BY THE CUSTOMER. THE IMAGE SHOWED THE LEFT ANTERIOR ELECTRODE WITH EXPOSED COMPONENTS. HOWEVER, THERE IS NO UNDERSTANDING IF THE PACKAGING WAS OPENED IN ACCORDANCE WITH THE PACKAGING ILLUSTRATIONS AS WELL AS HOW THE PADS WERE REMOVED FROM THE STYRENE. THE ELECTRODES WERE NOT RETURNED AND HAVE BEEN CONFISCATED BY LOCAL AUTHORITIES PREVENTING A LOT NUMBER VERIFICATION. NO RETAIN SAMPLE DEVICE HISTORY REVIEW COULD BE PREFORMED. THE ZOLL AED PLUS WAS CONFIRMED TO FUNCTION AS INTENDED, PROVIDING AUDIBLE AND VISUAL PROMPTS DURING SHOCK DELIVERY. LOG REVIEW SHOWED THAT THREE SHOCKS WERE DELIVERED AND ALL ADVISORY PROMPTS WERE ISSUED PROPERLY. THE DEVICE PASSED TESTING SUCCESSFULLY. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: THIS SUPPLEMENTAL MEDWATCH REPORT IS CORRECTING INFORMATION SUBMITTED ON THE PREVIOUS MEDWATCH REPORTS. PLEASE REFERENCE SECTION E1 (COUNTRY).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21977696&pc=OZD | 1220648-2025-28176 | 2025-04-18 04:00 | Death | ABIOMED, INC. | 2025-05-07 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT AN IMPELLA CP PUMP WAS PLACED TO SUPPORT A PATIENT ADMITTED IN WITH CARDIOGENIC SHOCK AND ORGAN FAILURE. THE IMPELLA CP ACCESSED LEG WAS ISCHEMIC. THE ISCHEMIA WAS NOT TREATED BY FEMOROFEMORAL BYPASS OR PREMATURE EXPLANT AS THE PATIENT COULD NOT SURVIVE OFF THE PUMP. FURTHER DETAILS STATED THAT THE PATIENT EXPIRED ON SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21977969&pc=DSQ | 2916596-2025-03002 | 2025-02-16 05:00 | Death | THORATEC CORPORATION | 2025-05-07 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Fall; Hematoma; Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WHO EXPERIENCED A GROUND LEVEL FALL (GLF) WAS FOUND TO HAVE A SUBARACHNOID HEMORRHAGE (SAH)/SUBDURAL HEMATOMA (SDH). THE PATIENT WAS TRANSFERRED TO THE HOSPITAL ON (B)(6) 2025 FOR HIGHER LEVEL OF CARE (HLOC) AND EVALUATION BY NEUROSURGERY (NSGY). THE PATIENT UNDERWENT AM EMERGENT LEFT CRANIOTOMY ON (B)(6) 2025 FOR SDH EVACUATION DUE TO 2/2 WORSENING NEUROLOGICAL STATUS. THE PATIENT REMAINED IN THE INTENSIVE CARE UNIT (ICU) WHERE THEIR STAY WAS COMPLICATED BY ONGOING NEUROLOGICAL DEFICITS. THE PATIENT WAS EMERGENTLY INTUBATED FOR HYPERCARBIA AND WORSENING SOMNOLENCE. THEY HAD A BILATERAL MIDDLE MENINGEAL ARTERY (MMA) ON (B)(6) 2025 AND WENT TO INTERVENTIONAL RADIOLOGY (IR) FOR A CEREBRAL ANGIOGRAM AND INTERCRANIAL SPASMOLYSIS ON (B)(6) 2025. THE PATIENT WAS DO NOT RESUSCITATE - CONTINUE OTHER TREATMENT (DNR COT) AND WAS LIBERATED FROM THE VENTILATOR ON (B)(6) 2025 AT THE URGING OF THE FAMILIES WISH TO NOT BE REINTUBATED. THE PATIENT DECOMPENSATED FURTHER DUE TO INABILITY TO VENTILATE PROPERLY. PER FAMILY REQUEST PATIENT WAS TRANSITIONED TO COMFORT CARE. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE EVENT WAS NOT DEVICE RELATED. NO AUTOPSY WAS PERFORMED Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, DOCUMENT D ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS INCLUDING BLEEDING, OTHER NEUROLOGICAL EVENTS (NOT STROKE-RELATED), AND DEATH THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. IT ALSO LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21977992&pc=NVZ | 9614453-2025-01827 | 2025-04-11 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-05-07 04:00 | NVZ | ASTRA S DR MRI SURESCAN | Appropriate Term/Code Not Available | Pulmonary Embolism; Cardiac Arrest; Chest Pain; Dyspnea; High Blood Pressure/ Hypertension; Low Blood Pressure/ Hypotension; Myocardial Infarction; Tachycardia; Ventricular Fibrillation; Loss of consciousness; Diaphoresis; Low Oxygen Saturation; Asystole | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT POST OPERATIVELY AFTER THE IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM, THE PATIENT WAS PROFOUNDLY SWEATING AND LOST CONSCIOUSNESS. PATIENT EXPERIENCED CHEST TIGHTNESS, SHORTNESS OF BREATH AND TACHYCARDIA. OXYGEN SATURATION WAS ABOUT SEVENTY PERCENT AND BLOOD PRESSURE WAS HIGH. OXYGEN MASK WAS USED TO STIMULATE BREATHING. MEDICATION WAS ADMINISTERED. DESPITE THESE EFFORTS, OXYGEN SATURATION REMAINED LOW. BLOOD PRESSURE DROPPED, DOPAMINE WAS GIVEN TO INCREASE BLOOD PRESSURE AND INFUSE FLUIDS. LATER, RESPIRATORY AND CARDIAC ARREST OCCURRED. CARDIOPULMONARY RESUSCITATION WAS ATTEMPTED TO RESCUE THE PATIENT. ENDOTRACHEAL INTUBATION WAS PERFORMED. ADRENALINE WAS INTERMITTENTLY USED TO STRENGTHEN THE HEART, AND DOPAMINE WAS CONTINUED TO INCREASE BLOOD PRESSURE AND ALKALINIZE FLUIDS. THE BLOOD PRESSURE AND HEARTBEAT DID NOT RECOVER. EVENTUALLY, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS SUSPECTED TO BE MYOCARDIAL INFARCTION, CARDIAC TAMPONADE AND PERICARDIAL EFFUSION DUE TO LEADS PERFORATION, PULMONARY EMBOLISM AND ACUTE CORONARY SYNDROME POSSIBLY RELATED TO THE ARTERIAL OCCLUSION.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY DUE TO ACUTE MYOCARDIAL INFARCTION. IT WAS ALSO REPORTED THAT THE PATIENT'S DEATH WAS NOT RELATED TO THE IPG SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21978007&pc=DSQ | 2916596-2025-02704 | 2025-04-11 04:00 | Death | THORATEC CORPORATION | 2025-05-07 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hypoxia | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO HYPOXIA. THE PATIENT WAS A DELIRIUM TREMENS (DT) WHO OPTED TO DECOMMISSION THE PUMP AS THEY HAD REACHED THE END OF THEIR LIFE. THE DEVICE OPERATED AS EXPECTED AND THE DEATH WAS NOT DEVICE RELATED. THE DEVICE WAS NOT RETURNED FOR EVALUATION.Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO HYPOXIA. THE PATIENT WAS A DESTINATION THERAPY (DT) PATIENT WHO OPTED TO DECOMMISSION THE PUMP AS THEY HAD REACHED THE END OF THEIR LIFE. THE DEVICE OPERATED AS EXPECTED AND THE DEATH WAS NOT DEVICE RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT THE DEVICE WILL NOT RETURN FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING PSYCHIATRIC EPISODE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21978026&pc=NPT | 2025587-2025-03492 | 2025-02-03 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-07 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Device Damaged by Another Device; Patient Device Interaction Problem | Cardiac Arrest; Cardiogenic Shock | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE INVOLVING THE EVOLUT R 29 MM, THE VALVE DISLODGED AFTER DEPLOYMENT. SUBSEQUENTLY, A SECOND VALVE WAS IMPLANTED. THE PATIENT EXPERIENCED A SHOCK AND SUBSEQUENTLY DIED. ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE IMPLANT, NO PRE-IMPLANT BALLOON DILATATION WAS PERFORMED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS AORTIC DISSECTION. THE PATIENT'S SHOCK EXPERIENCED WAS CARDIOGENIC SHOCK DUE TO CARDIAC ARREST. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE PATIENT'S NATIVE AORTIC VALVE WAS HEAVILY CALCIFIED WHICH ADDED STRESS TO THE FIRST EVOLUT VALVE AND RESULTED IN DISLODGEMENT TO THE ASCENDING AORTA. DURING ADVANCEMENT OF THE DELIVERY CATHETER SYSTEM (DCS) WITH THE LOADED SECOND VALVE, THE FIRST VALVED MOVED MULTIPLE TIMES AND RESULTED IN AN ASCENDING AORTIC DISSECTION WITH SUBSEQUENT CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS INITIATED AND PERFORMED FOR MORE THAN ONE HOUR. AFTERWARD, THE PATIENT WAS PREPARED FOR OPEN HEART SURGERY TO REPAIR THE AORTIC DISSECTION. THE PATIENT WAS PLACED ON CARDIOPULMONARY BYPASS; HOWEVER, THE PATIENT'S HEART DID NOT RESPOND TO THE BYPASS PROCEDURE. SUBSEQUENTLY, THE PATIENT'S DEATH WAS PRONOUNCED AND A TIME OF DEATH WAS DECLARED. PER THE PHYSICIAN, THE MEDTRONIC VALVES AND THE SECOND MEDTRONIC DCS CANNOT BE EXCLUDED AS CONTRIBUTING FACTORS TO THE AORTIC DISSECTION AND SUBSEQUENT DEATH. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: MEDTRONIC PRODUCT BRAND NAME: DELIV SYS ENVEOR-L ENVEOR; PRODUCT ID: ENVEOR-L; (LOT: UNKNOWN); MANUFACTURED DATE: UNKNOWN; USE BY DATE: UNKNOWN; UDI: UNKNOWN CONTINUATION OF D10: OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME: EVOLUT R 29 MM; MEDTRONIC PRODUCT ID: EVOLUTR-29 (LOT: J168578); PRODUCT TYPE: 0195-HEART VALVES; FULL UDI#: (B)(4); MANUFACTURED DATE: 2024-02-13; USE BY DATE: 2025-02-12; IMPLANT DATE: (B)(6) 2025. H6. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. THIS EVENT WAS REVIEWED BY MEDTRONIC¿S MEDICAL SAFETY TEAM. THE EXCERPT OF THE MEDICAL SAFETY SUBJECT MATTER EXPERT (SME) REVIEW REPORT FOLLOWS: BASED ON THE SEQUENCE OF EVENTS AND MEDICAL EXPERTISE, THE EVENT OF VALVE DISLODGEMENT (TO THE ASCENDING AORTA), REQUIRING PLACEMENT OF A 2ND EVOLUT R VALVE, IS LIKELY RELATED TO THE DEVICE THAT COULD NOT SEAT WELL IN THE ANNULUS LIKELY DUE TO SIGNIFICANT AORTIC VALVE CALCIFICATION AS A CONTRIBUTING FACTOR. THE EVENT OF AORTIC DISSECTION, LEADING TO CARDIOGENIC SHOCK AND INITIATION OF LIFE-SAVING MEASURES, IS LIKELY RELATED TO THE 2ND DCS CROSSING THROUGH THE FIRST EVOLUT R VALVE CAUSING THE VALVE TO "MOVE" AND LIKELY RESULTING IN THE AORTIC DISSECTION PER SITE REPORTING. THE EVENT OF DEATH IS LIKELY RELATED TO COMPLICATIONS FROM THE AORTIC DISSECTION. PRODUCT ANALYSIS (DEVICE LISTED IN SECTION D): THE VALVE REMAINED IMPLANTED AND NO PROCEDURAL IMAGES OF THE IMPLANTATION OR THE IMPLANTED VALVE WERE SUBMITTED FOR MEDTRONIC REVIEW. THE DEVICE HISTORY RECORD OF THE SUSPECT VALVE WAS REVIEWED AND SHOWED THIS PRODUCT MET ALL MANUFACTURING SPECIFICATIONS FOR PRODUCT RELEASED FOR DISTRIBUTION. NO ISSUES WERE IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. CARDIOGENIC SHOCK AND CARDIAC ARREST ARE KNOWN POTENTIAL ADVERSE EFFECT PER THE DEVICE INSTRUCTIONS FOR USE (IFU). UNFORTUNATELY, A RELATIONSHIP OF THE REPORTED CARDIOGENIC SHOCK TO THE DEVICE OR PROCEDURE COULD NOT BE DETERMINED WITH THE LIMITED INFORMATION AVAILABLE. PATIENT DEATH IS A POTENTIAL RISK ASSOCIATED WITH THE IMPLANTATION OF A BIOPROSTHESIS VALVE PER THE DEVICE IFU. A PROCEDURE- OR VALVE-RELATED DEATH IS AN INHERENT RISK WHEN THE PATIENT CONDITION IS SUCH THAT A TRANSCATHETER AORTIC VALVE IS NEEDED TO SUSTAIN CARDIAC FUNCTION AND IT CAN OCCUR DESPITE AN IDEAL IMPLANT PROCEDURE OR DEVICE FUNCTIONALITY. IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS AORTIC DISSECTION. PER THE MEDICAL SAFETY ASSESSMENT, THE EVENT OF DEATH IS LIKELY RELATED TO COMPLICATIONS FROM THE AORTIC DISSECTION. HOWEVER, A CONCLUSIVE ASSESSMENT OF THE RELATIONSHIP BETWEEN THE EVENT AND THE DEVICE COULD NOT BE REACHED. NO OTHER TECHNICAL ASSESSMENTS WERE PERFORMED. NO NEW OR UNEXPECTED DEVICE OR THERAPY ISSUE WAS IDENTIFIED; THE EVENT IS FORESEEN, WITHIN EXPECTED SEVERITY, DOCUMENTED IN RISK MANAGEMENT, AND INCLUDED IN MONITORING/TRENDING. THERE IS NO INFORMATION TO SUGGEST A DEVICE QUALITY DEFICIENCY THAT MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. TRENDS FOR THESE EVENT TYPES ARE ASSESSED, AND NO FURTHER ACTION IS RECOMMENDED AT THIS TIME. MEDTRONIC WILL CONTINUE TO MONITOR FIELD PERFORMANCE OR SIMILAR EVENTS SHOULD THEY OCCUR. PRODUCT ANALYSIS (DEVICE LISTED IN H11 AS SYSTEM REPORTED): THE SUBJECT DELIVERY CATHETER SYSTEM (DCS) WAS NOT RETURNED TO MEDTRONIC AND PROCEDURAL IMAGES WERE NOT SUBMITTED TO MEDTRONIC FOR REVIEW; THEREFORE, ANALYSIS COULD NOT BE PERFORMED. A DEVICE HISTORY RECORD REVIEW WAS NOT REQUIRED AS THE EVENT DESCRIPTION DID NOT INDICATE A POTENTIAL MANUFACTURING ISSUE. HOWEVER, MANUFACTURING CONTROLS ARE IN PLACE TO ENSURE THAT THE DEVICE MET SPECIFICATION REQUIREMENTS PRIOR TO RELEASE FROM THE MANUFACTURING FACILITY. PRODUCT LABELING MATERIAL AND RISK MANAGEMENT FILES WERE REVIEWED. THE PRIMARY HARM IS DOCUMENTED IN THE PRODUCT LABELING AND THE PRIMARY HARM AND SEVERITY ARE SUFFICIENTLY DOCUMENTED IN THE RISK MANAGEMENT FILE. THERE IS NO INFORMATION TO SUGGEST A DEVICE QUALITY DEFICIENCY OR A FAILURE TO MEET SPECIFICATIONS THAT MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT, BUT A CONCLUSIVE ROOT CAUSE CANNOT BE DETERMINED AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21978034&pc=DTB | 2649622-2025-12364 | 2025-03-28 04:00 | Death | MPRI | 2025-05-07 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Cardiac Arrest; Cardiac Tamponade; Pericardial Effusion | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT POST OPERATIVELY THE PATIENT EXPERIENCED PERICARDIAL EFFUSION, TAMPONADE AND CARDIAC ARREST. EVENTUALLY, THE PATIENT PASSED AWAY. Manufacturer Narrative: CORRECTED: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21978036&pc=DTB | 3008973940-2025-03140 | 2025-04-11 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-05-07 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Pulmonary Embolism; Cardiac Arrest; Chest Pain; Dyspnea; High Blood Pressure/ Hypertension; Low Blood Pressure/ Hypotension; Myocardial Infarction; Perforation; Tachycardia; Ventricular Fibrillation; Cardiac Tamponade; Loss of consciousness; Diaphoresis; Low Oxygen Saturation; Pericardial Effusion; Asystole | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT POST OPERATIVELY AFTER THE IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM, THE PATIENT WAS PROFOUNDLY SWEATING AND LOST CONSCIOUSNESS. PATIENT EXPERIENCED CHEST TIGHTNESS, SHORTNESS OF BREATH AND TACHYCARDIA. OXYGEN SATURATION WAS ABOUT SEVENTY PERCENT AND BLOOD PRESSURE WAS HIGH. OXYGEN MASK WAS USED TO STIMULATE BREATHING. MEDICATION WAS ADMINISTERED. DESPITE THESE EFFORTS, OXYGEN SATURATION REMAINED LOW. BLOOD PRESSURE DROPPED, DOPAMINE WAS GIVEN TO INCREASE BLOOD PRESSURE AND INFUSE FLUIDS. LATER, RESPIRATORY AND CARDIAC ARREST OCCURRED. CARDIOPULMONARY RESUSCITATION WAS ATTEMPTED TO RESCUE THE PATIENT. ENDOTRACHEAL INTUBATION WAS PERFORMED. ADRENALINE WAS INTERMITTENTLY USED TO STRENGTHEN THE HEART, AND DOPAMINE WAS CONTINUED TO INCREASE BLOOD PRESSURE AND ALKALINIZE FLUIDS. THE BLOOD PRESSURE AND HEARTBEAT DID NOT RECOVER. EVENTUALLY, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS SUSPECTED TO BE MYOCARDIAL INFARCTION, CARDIAC TAMPONADE AND PERICARDIAL EFFUSION DUE TO LEADS PERFORATION, PULMONARY EMBOLISM AND ACUTE CORONARY SYNDROME POSSIBLY RELATED TO THE ARTERIAL OCCLUSION.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY DUE TO ACUTE MYOCARDIAL INFARCTION. IT WAS ALSO REPORTED THAT THE PATIENT'S DEATH WAS NOT RELATED TO THE IPG SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21978211&pc=DSQ | 3007042319-2025-01111 | 2025-04-26 04:00 | Death | HEARTWARE, INC. | 2025-05-07 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Intracranial Hemorrhage; Viral Infection | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A HISTORY OF CARDIOMYOPATHY EXPERIENCED A HEMORRHAGIC STROKE FOLLOWED BY A COVID INFECTION, LEADING TO PLACEMENT ON COMFORT CARE. THE PATIENT SUBSEQUENTLY EXPIRED WHILE ON COMFORT CARE. THE EVENT INVOLVED A VENTRICULAR ASSIST DEVICE, SPECIFICALLY THE (B)(6) VENTRICULAR ASSIST DEVICE (VAD), WHICH HAD NO CURRENT PUMP ISSUES NOTED. THE VAD REMAINS IMPLANTED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION SUMMARY. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6). THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED WHILE ON COMFORT CARE AFTER EXPERIENCING A HEMORRHAGIC STROKE FOLLOWED BY A COVID INFECTION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, NEUROLOGICAL DYSFUNCTION, INFECTION, AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED.### MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21978212&pc=DTB | 3008973940-2025-03141 | 2025-04-11 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-05-07 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Pulmonary Embolism; Cardiac Arrest; Chest Pain; Dyspnea; High Blood Pressure/ Hypertension; Low Blood Pressure/ Hypotension; Myocardial Infarction; Perforation; Tachycardia; Ventricular Fibrillation; Cardiac Tamponade; Loss of consciousness; Diaphoresis; Low Oxygen Saturation; Pericardial Effusion; Asystole | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT POST OPERATIVELY AFTER THE IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM, THE PATIENT WAS PROFOUNDLY SWEATING AND LOST CONSCIOUSNESS. PATIENT EXPERIENCED CHEST TIGHTNESS, SHORTNESS OF BREATH AND TACHYCARDIA. OXYGEN SATURATION WAS ABOUT SEVENTY PERCENT AND BLOOD PRESSURE WAS HIGH. OXYGEN MASK WAS USED TO STIMULATE BREATHING. MEDICATION WAS ADMINISTERED. DESPITE THESE EFFORTS, OXYGEN SATURATION REMAINED LOW. BLOOD PRESSURE DROPPED, DOPAMINE WAS GIVEN TO INCREASE BLOOD PRESSURE AND INFUSE FLUIDS. LATER, RESPIRATORY AND CARDIAC ARREST OCCURRED. CARDIOPULMONARY RESUSCITATION WAS ATTEMPTED TO RESCUE THE PATIENT. ENDOTRACHEAL INTUBATION WAS PERFORMED. ADRENALINE WAS INTERMITTENTLY USED TO STRENGTHEN THE HEART, AND DOPAMINE WAS CONTINUED TO INCREASE BLOOD PRESSURE AND ALKALINIZE FLUIDS. THE BLOOD PRESSURE AND HEARTBEAT DID NOT RECOVER. EVENTUALLY, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS SUSPECTED TO BE MYOCARDIAL INFARCTION, CARDIAC TAMPONADE AND PERICARDIAL EFFUSION DUE TO LEADS PERFORATION, PULMONARY EMBOLISM AND ACUTE CORONARY SYNDROME POSSIBLY RELATED TO THE ARTERIAL OCCLUSION.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY DUE TO ACUTE MYOCARDIAL INFARCTION. IT WAS ALSO REPORTED THAT THE PATIENT'S DEATH WAS NOT RELATED TO THE IPG SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21978274&pc=OZD | 1220648-2025-28180 | 2025-04-18 04:00 | Death | ABIOMED, INC. | 2025-05-07 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, A CLOT WAS INGESTED. THE IMPELLA CP DROPPED FLOWS FROM 3.5 TO 1.2 AND THEN TO 0.8. A CLOT WAS INGESTED BY THE IMPELLA CP PUMP. THE IMPELLA CP WAS REMOVED AND REPLACED WITH A NEW PUMP. THE PATIENT EXPIRED ON SUPPORT OF THE REPLACEMENT PUMP. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF THROMBUS HAS BEEN COMPLETED. CLINICAL STATED REPORTED CLOT DURING CASE. PUMP WAS RETURNED AND NO DAMAGES OR ABNORMALITIES WERE FOUND. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE THROMBUS WAS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA PUMP WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION, SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21978283&pc=FGE | 3005099803-2025-01885 | 2025-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-07 04:00 | FGE | EXTRACTOR PRO RX | Adverse Event Without Identified Device or Use Problem | Perforation; Peritonitis | K102082 |  | 1 | Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT AN EXTRACTOR PRO RX RETRIEVAL BALLOON WAS USED IN THE BILE DUCT DURING AN ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY (ERCP) PROCEDURE PERFORMED IN (B)(6) 2025. THE EXACT PROCEDURE DATE WAS NOT REPORTED. DURING THE PROCEDURE, A PERFORATION OCCURRED NEAR THE DUODENUM. IT WAS REPORTED THAT THE PHYSICIAN SUSPECTS THAT THE EXTRACTOR PRO RX RETRIEVAL BALLOON MAY BE A CONTRIBUTING FACTOR IN CAUSING THE PERFORATION. THE PROCEDURE WAS COMPLETED USING THE SAME ORIGINAL DEVICE. THE PATIENT PASSED AWAY IN (B)(6) 2025 DUE TO PERITONITIS. THE EXACT PATIENT DEATH DATE WAS NOT REPORTED. Manufacturer Narrative: BLOCK B2: APPROXIMATED BASED ON THE REPORTED PATIENT DEATH DATE OF (B)(6) 2025. BLOCK B3: APPROXIMATED BASED ON THE REPORTED EVENT DATE OF JANUARY 2025. BLOCK D4, H4: THE COMPLAINANT WAS UNABLE TO PROVIDE THE LOT NUMBER. THEREFORE, THE MANUFACTURE AND EXPIRATION DATES ARE UNKNOWN. BLOCK H6: IMDRF PATIENT CODE E2114 IS BEING USED TO CAPTURE THE REPORTABLE EVENT OF PERFORATION. IMDRF PATIENT CODE E1024 IS BEING USED TO CAPTURE THE REPORTABLE EVENT OF PERITONITIS. IMDRF IMPACT CODE F02 IS BEING USED TO CAPTURE THE REPORTABLE EVENT OF PATIENT DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21978428&pc=DSQ | 2916596-2025-02769 | 2025-04-16 04:00 | Death | THORATEC CORPORATION | 2025-05-07 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Tachycardia; Ventricular Fibrillation; Brain Injury; Hypovolemia; Loss of consciousness; Syncope/Fainting | P160054 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT WAS SCENE IN CLINIC ON (B)(6) 2025 AND WAS DOING WELL AND LEFT THE CLINIC. EMERGENCY MEDICAL SERVICES WERE CALLED DUE TO THE PATIENT BEING UNRESPONSIVE AROUND 20:30. THE PATIENT WAS FOUND TO BE IN VENTRICULAR TACHYCARDIA (VT) WHICH EVENTUALLY CONVERTED INTO VENTRICULAR FIBRILLATION (VF). THE PATIENT WAS UNCONSCIOUS WHEN THEY ARRIVED AT THE EMERGENCY DEPARTMENT (ED). THE ED DID NOT RECOGNIZE THE PATIENT WAS IN VF UPON ARRIVAL AND THUS DID NOT PERFORM ELECTRICAL CARDIOVERSION. THE PATIENT WAS SHOCKED AFTER 2 HOURS AT THE ED. THE PATIENT WAS NOTED TO HAVE NEVER REGAINED CONSCIOUSNESS. THE ARRHYTHMIA WAS NOTED TO HAVE BEEN SUSTAINED. THE PATIENT WAS NOTED TO NOT HAVE A HISTORY OF ARRHYTHMIA, AND THE PATIENT HAD NEVER HAD ANY PREMATURE VENTRICULAR CONTRACTIONS IN THE PAST, NOR HAD THE PATIENT EVER REPORTED ANY PALPITATIONS OR HEART IRREGULARITIES. THE PATIENT WAS TRANSFERRED TO A VENTRICULAR ASSIST DEVICE (VAD) CENTER AND AT THAT TIME THE PATIENT HAD A NORMAL SINUS RHYTHM AND ATRIAL FIBRILLATION. THE PATIENT PASSED AWAY AFTER THE FAMILY WITHDREW CARE DUE TO ANOXIC BRAIN INJURY IN THE SETTING OF PROLONGED DOWNTIME. THE CENTER WAS UNSURE IF THE EVENT WAS DEVICE OR THERAPY RELATED AS THEY NOTED THERE MAY HAVE BEEN SUCTION INITIALLY AS THE PATIENT WAS HYPOVOLEMIC. LOG FILE REVIEW REVEALED LOW FLOW EVENTS ON (B)(6) 2025. THERE WERE NO OTHER UNUSUAL EVENTS RECORDED IN THE LOG FILE EVENT HISTORY. THE LOW FLOW ALARMS WERE DUE TO ARRHYTHMIA LEADING TO RIGHT VENTRICLE (RV) FAILURE. THE LOW FLOW ALARMS RESOLVED, WHEN THE PATIENT WAS TRANSFERRED, THE VAD SPEED WAS DECREASED, AND THE PATIENT WAS PLACED ON ANTIARRHYTHMICS AS WELL AS INOTROPES DUE TO THE RV FAILING THE MECHANICAL CIRCULATORY SUPPORT EQUIPMENT APPEARED TO BE OPERATING AS INTENDED. ADDITIONAL INFORMATION REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED, AND THAT THE PUMP WOULD NOT BE EXPLANTED.Event Description: IT WAS REPORTED THAT THE DEVICE WAS OPERATING AS EXPECTED.Event Description: IT WAS REPORTED THAT EMERGENCY MEDICAL SERVICES WERE CALLED DUE TO THE PATIENT BEING UNRESPONSIVE AROUND 20:30 ON (B)(6) 2025. THE PATIENT WAS FOUND TO BE IN VENTRICULAR FIBRILLATION (VF). LOG FILE REVIEW REVEALED LOW FLOW EVENTS ON (B)(6) 2025. THERE WERE NO OTHER UNUSUAL EVENTS RECORDED IN THE LOG FILE EVENT HISTORY. THE PATIENT WAS UNCONSCIOUS WHEN THEY ARRIVED AT THE EMERGENCY DEPARTMENT (ED). THE ED DID NOT RECOGNIZE THE PATIENT WAS IN VF UPON ARRIVAL AND THUS DID NOT PERFORM ELECTRICAL CARDIOVERSION. THE PATIENT WAS SHOCKED AFTER 2 HOURS AT THE ED. THE SITE ASSUMED THAT THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA (VT) AND THEN EVENTUALLY CONVERTED TO VF. Manufacturer Narrative: ADDITIONAL CODES: HEALTH EFFECT - CLINICAL CODES: 4446 - HEART FAILURE/CONGESTIVE HEART FAILURE. 4868 - HEMODYNAMIC INSTABILITY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION B5 CORRECTION, SECTION E3 CORRECTION, SECTION H6 CORRECTION: HEALTH EFFECT - CLINICAL CODE 2418 - LOSS OF CONSCIOUSNESS ADDED. SECTION H6 CORRECTION: HEALTH EFFECT - IMPACT CODE 4607 - HOSPITALIZATION OR PROLONGED HOSPITALIZATION ADDED. SECTION H6 CORRECTION: HEALTH EFFECT - IMPACT CODE 4641 - UNEXPECTED MEDICAL INTERVENTION ADDED. ADDITIONAL CODES: HEALTH EFFECT - CLINICAL CODES: 4446 - HEART FAILURE/CONGESTIVE HEART FAILURE, 4868 - HEMODYNAMIC INSTABILITY, MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED THE LOW FLOW ALARMS. ALTHOUGH A SPECIFIC CAUSE FOR THESE ALARMS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION, THE ACCOUNT REPORTED THAT THE ALARMS WERE CAUSED BY PROLONGED ARRHYTHMIA LEADING TO RIGHT VENTRICULAR FAILURE. THE SUBMITTED SYSTEM CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2025. ON (B)(6) 2025, THE FILE CAPTURED TRANSIENT LOW FLOW HAZARD ALARMS. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE FIXED SPEED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU AND PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA, RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU PROVIDES AN EXPLANATION OF PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 4 ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AND HYPOVOLEMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6, UNDER ¿RIGHT HEART FAILURE¿, ALSO STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, OUTLINE SYSTEM CONTROLLER ALARMS, INCLUDING THE LOW FLOW HAZARD ALARM, AND PROVIDE INFORMATION REGARDING HOW TO RESPOND TO AND TROUBLESHOOT THE ALARMS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: ADDITIONAL CODES: HEALTH EFFECT - CLINICAL CODES: 4446 - HEART FAILURE/CONGESTIVE HEART FAILURE, 4868 - HEMODYNAMIC INSTABILITY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21978456&pc=OZD | 1220648-2025-28182 | 2025-04-18 04:00 | Death | ABIOMED, INC. | 2025-05-07 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP PUMP FOR SUPPORT DURING A HIGH-RISK PROCEDURE AND CARDIOGENIC SHOCK. THE PATIENT WAS SUFFERING FROM PULSELESS ELECTRICAL ACTIVITY (PEA) AND HAD CARDIAC INTERVENTION. DURING THE CARDIAC PROCEDURE AND WHILE ON SUPPORT, THERE WAS LOW FLOWS AND PS ISSUES BUT, THE PUMP WAS NOT REPLACED. AFTER ALL MEASURES AND UNDER FOUR HOURS OF SUPPORT, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE PATIENT WAS NOT AN EXTRACORPOREAL MEMBRANE OXYGENATION OR IMPELLA RP CANDIDATE. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW HAS BEEN COMPLETED. THE ROOT CAUSE OF THE LOW PUMP FLOW COULD NOT BE DEFINITIVELY DETERMINED AS NO PRODUCT OR LOGS WERE RETURNED FOR EVALUATION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21978476&pc=DSQ | 3007042319-2025-01112 | 2025-04-28 04:00 | Death | HEARTWARE, INC. | 2025-05-07 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Contamination; Connection Problem; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT WAS FOUND DECEASED IN THE PATIENT¿S APARTMENT ON (B)(6) 2025. NO FURTHER INFORMATION IS KNOWN.Event Description: IT WAS FURTHER REPORTED THAT A CONTROLLER WAS RETURNED TO THE MANUFACTURER. MANUFACTURER'S ANALYSIS SUBSEQUENTLY REVEALED AN ENVIRON MENT/USAGE/FRACTURE PROBLEM IDENTIFIED. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. CONTROLLER (B)(6) WAS RETURNED FOR EVALUATION. A REVIEW OF THE VAD MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE. A REVIEW OF THE CONTROLLER DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. VARIOUS ANALYSES WERE CONDUCTED AND REVIEWED IN ORDER TO EVALUATE THE PERFORMANCE OF THE DEVICES IN RELATION TO THE REPORTED EVENT. FAILURE ANALYSIS OF (B)(6) REVEALED THAT THE CONTROLLER PASSED FUNCTIONAL TESTING. VISUAL INSPECTION OF THE RETURNED CONTROLLER REVEALED CONTAMINATION WITHIN BOTH POWER PORTS AND A MISSING SERIAL PORT CAP. THESE ARE ADDITIONAL OBSERVATIONS UNRELATED TO THE REPORTED EVENT, LIKELY DUE TO THE HANDLING OF THE DEVICE. AN INTERNAL INSPECTION OF THE RETURNED DEVICES DID NOT REVEAL ANY ABNORMALITIES. LOG FILE ANALYSIS ASSOCIATED WITH (B)(6) REVEALED SIX (6) CRITICAL BATTERY ALARMS INVOLVING (B)(6) LOGGED ON (B)(6) 2025 STARTING AT 03:43:41 AND TWO (2) CONTROLLER FAULT ALARMS WERE LOGGED ON (B)(6) 2025 AT 04:23:59 AND 04:24:08, WHICH WERE DUE TO THE PATIENT ALLOWING THE BATTERIES TO DEPLETE BELOW 10% RSOC. ANALYSIS OF THE LOG FILES REVEALED THAT AT THE TIME OF THE FIRST CRITICAL BATTERY ALARM NO POWER SOURCE WAS CONNECTED TO POWER PORT ONE (1) AND (B)(6) WAS CONNECTED TO POWER PORT TWO (2) WITH 9% RSOC. THE BATTERY (B)(6) THEN CONTINUED TO DEPLETE. A CONTROLLER FAULT ALARM WILL OCCUR IF THE BATTERY IS APPROACHING 0% RSOC. LOG FILE ANALYSIS ALSO REVEALED THAT THE CONTROLLER LOST POWER AT 04:20:55. THE BATTERY HAD BEEN ALLOWED TO DEPLETE COMPLETELY, RESULTING IN A LOSS OF POWER TO THE CONTROLLER. CONTROLLER (B)(6) IS LOADED WITH THE ALTERNATE SOFTWARE FOR THE VAD START ALGORITHM. THE MOST LIKELY ROOT CAUSE OF THE CRITICAL BATTERY ALARMS, CONTROLLER FAULT ALARMS, AND LOSS OF POWER TO THE CONTROLLER CAN BE ATTRIBUTED TO THE BATTERY DEPLETING TO 0% RSOC. CAPA PR00551638 INVESTIGATED CONTROLLER LOSSES OF POWER. EVEN THOUGH THIS CAPA IS NOW CLOSED, (B)(6) FALLS IN SCOPE OF THIS CAPA. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH IS UNKNOWN. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. ADDITIONAL PRODUCT: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ CONTROLLER D4: SERIAL#: (B)(6) D9: YES, RETURNED ON 09-JUL-2025 H3: YES H5: NO H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21978647&pc=DSQ | 2916596-2025-02772 | 2025-04-10 04:00 | Death | THORATEC CORPORATION | 2025-05-07 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Encephalopathy; Low Blood Pressure/ Hypotension; Hypoxia; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT HEARTMATE 3 (HM3) IMPLANTATION DUE TO VALVULAR HEART FAILURE ON (B)(6) 2025. THIS WAS COMBINED WITH A SEVERE MITRAL STENOSIS (MS) REPAIR. HYPOTENSION AND METABOLIC ACIDOSIS WERE OBSERVED, WHICH OCCURRED DURING WEANING FROM CARDIOPULMONARY BYPASS (CPB), WHICH NECESSITATED THE USE OF A RIGHT VENTRICULAR ASSIST DEVICE (RVAD). BLOOD PRESSURE WAS STILL UNSTABLE. LEVOPHED (NOREPINEPHRINE), DOPAMINE, BOSMIN (EPINEPHRINE), AND DOBUTAMINE LINES WERE SUPPORTED. VASOPLEGIA WAS SUSPECTED. POSTOPERATIVE PARAMETERS FOR THE RVAD WERE 3050 RPM WITH FLOW RATE 3.3 LPM AND 5400 RPM WITH FLOW RATE 4.7¿4.8 LPM FOR THE LVAD. THE PATIENT'S CONSCIOUSNESS WAS UNCLEAR AFTER SURGERY. AN ARRANGED BRAIN COMPUTED TOMOGRAPHY (CT) REVEALED HYPOXIC ENCEPHALOPATHY (HE). THE FAMILY DECIDED TO SIGN DO-NOT-RESUSCITATE (DNR) AND REMOVE ALL THE LIFE SUPPORT SYSTEMS. HEALTHCARE PROVIDERS (HCPS) DETERMINED THAT THE EVENT WAS NOT RELATED TO THE DEVICES, AND THE DEVICES OPERATED AS EXPECTED.Event Description: THE CAUSE OF DEATH WAS MULTIPLE ORGAN FAILURE. THE MULTIPLE ORGAN FAILURE WAS THOUGHT TO NOT BE RELATED TO THE DEVICE. NO PRODUCTS WOULD BE RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION), NEUROLOGICAL EVENTS (NOT STROKE-RELATED), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE¿) STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THIS SECTION (UNDER ¿CAUTION!¿) ALSO EXPLAINS THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP. SECTION 6 ALSO LISTS NEUROLOGICAL DYSFUNCTION AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21980383&pc=QZI | 2124215-2025-29461 | 2025-04-15 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-07 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Pericardial Effusion |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED A PERICARDIAL EFFUSION AND CARDIAC ARREST. A PULSED FIELD ABLATION (PFA) PROCEDURE TO TREAT AN ATRIAL FIBRILLATION WITH A FARAWAVE CATHETER WAS COMPLETED WITHOUT ISSUES. DURING CASE CLOSURE, A PERICARDIAL EFFUSION WAS FOUND, AND A CARDIAC ARREST OCCURRED. A PERICARDIOCENTESIS WAS PERFORMED TO SOLVE THE EFFUSION, AND CARDIOPULMONARY RESUSCITATION WAS GIVEN UNTIL THE PATIENT WAS STABLE, AND THEN WAS TAKEN TO THE INTENSIVE CARE UNIT (ICU), IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY. A DECISION WAS MADE TO REMOVE THE PATIENT FROM THE VENTILATOR. PHYSICIAN COMMENTED THAT COMBINATION OF THE TAMPONADE AND UNDER-VENTILATION RESULTED IT ACIDOSIS. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. INFORMATION WAS NOT AVAILABLE BECAUSE DEVICE WAS DISCARDED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21981021&pc=OUT | 2029214-2025-01141 | 2025-04-21 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-05-07 04:00 | OUT | PIPELINE FLEX W/SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P100018 |  | 1 | Event Description: MEDTRONIC RECEIVED A REPORT THAT AFTER A PIPELINE PLACEMENT THE PED, DURING PTA WITH A BALLOON, OVER-INFLATION OF THE BALLOON RESULTED IN BLEEDING FROM THE MCA. SUBSEQUENTLY, THE MCA WAS EMBOLIZED WITH COILS TO ACHIEVE HEMOSTASIS. A PATIENT DEATH WAS REPORTED (THIS IS NOT A HEALTH HAZARD CAUSED BY PIPELINE). THE REPORTED DEVICE AND ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTION FOR USE (IFU). THE PATIENT WAS BEING TREATED FOR AN UNRUPTURED LEFT C7 ICA ANEURYSM WITH A SACCULAR MORPHOLOGY. SIZE OF ANEURYSM: MAXIMUM DIAMETER WAS AROUND 8 MM AND NECK DIAMETER WAS AROUND 5 MM. BLOOD VESSEL DIAMETER IN THE LANDING ZONE: DISTAL SIDE 3.0 MM AND PROXIMAL SIDE 4.5 MM. THE PATIENT¿S VESSEL TORTUOSITY WAS ORDINARY. DAPT WAS PERFORMED (PRU LEVEL WAS UNKNOWN). POSTOPERATIVE FINDINGS RELATED TO BLOOD FLOW IMAGING: NO PARTICULAR PROBLEMS AFTER PIPELINE PLACEMENT. NO PRODUCT DEFECT/MALFUNCTION WAS NOTED. ON THE DAY OF THE PROCEDURE, EXIT WAS MADE AFTER PLACING THE PED. SUBSEQUENTLY, A REPORT WAS RECEIVED FROM THE DISTRIBUTOR INDICATING THAT THE MCA WAS EMBOLIZED WITH COILS, COMPLETING THE PROCEDURE. SINCE THE PIPELINE (AND ITS PLACEMENT) WAS NOT THE DIRECT CAUSE OF THIS INCIDENT. IT WAS NOTED THAT THE CAUSE OF DEATH WAS PROBABLY DUE TO BLEEDING FROM THE MCA BUT NOT CONFIRMED IT, SO THE DIRECT CAUSE IS UNKNOWN. A DEPLOYMENT DEFECT WAS NOT OBSERVED. THIS IS A PTA INTENDED TO BE CRIMPED TO THE PARENT BLOOD VESSEL. NO RESISTANCES OR ABNORMALITIES DURING DELIVERY OF THE PIPELINE WERE NOTED.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT PASSED AWAY AT (B)(6) HOSPITAL, LOCATED IN (B)(6). THE CONCOMITANT BALLOON USED IN PAT WAS THE 7-7 OF THE SHOURYU HR.Event Description: ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT A BALLOON (NON-MEDTRONIC) USED WITH THE PIPELINE WAS DAMAGED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21981041&pc=KRA | 9612164-2025-02265 | 2025-04-21 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-05-07 04:00 | KRA | PHENOM CATHETER | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K180959 |  | 1 | Event Description: MEDTRONIC RECEIVED A REPORT THAT AFTER A PIPELINE PLACEMENT THE PED, DURING PTA WITH A BALLOON, OVER-INFLATION OF THE BALLOON RESULTED IN BLEEDING FROM THE MCA. SUBSEQUENTLY, THE MCA WAS EMBOLIZED WITH COILS TO ACHIEVE HEMOSTASIS. A PATIENT DEATH WAS REPORTED (THIS IS NOT A HEALTH HAZARD CAUSED BY PIPELINE). THE REPORTED DEVICE AND ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTION FOR USE (IFU). THE PATIENT WAS BEING TREATED FOR AN UNRUPTURED LEFT C7 ICA ANEURYSM WITH A SACCULAR MORPHOLOGY. SIZE OF ANEURYSM: MAXIMUM DIAMETER WAS AROUND 8 MM AND NECK DIAMETER WAS AROUND 5 MM. BLOOD VESSEL DIAMETER IN THE LANDING ZONE: DISTAL SIDE 3.0 MM AND PROXIMAL SIDE 4.5 MM. THE PATIENT¿S VESSEL TORTUOSITY WAS ORDINARY. DAPT WAS PERFORMED (PRU LEVEL WAS UNKNOWN). POSTOPERATIVE FINDINGS RELATED TO BLOOD FLOW IMAGING: NO PARTICULAR PROBLEMS AFTER PIPELINE PLACEMENT. NO PRODUCT DEFECT/MALFUNCTION WAS NOTED. ON THE DAY OF THE PROCEDURE, EXIT WAS MADE AFTER PLACING THE PED. SUBSEQUENTLY, A REPORT WAS RECEIVED FROM THE DISTRIBUTOR INDICATING THAT THE MCA WAS EMBOLIZED WITH COILS, COMPLETING THE PROCEDURE. SINCE THE PIPELINE (AND ITS PLACEMENT) WAS NOT THE DIRECT CAUSE OF THIS INCIDENT. IT WAS NOTED THAT THE CAUSE OF DEATH WAS PROBABLY DUE TO BLEEDING FROM THE MCA BUT NOT CONFIRMED IT, SO THE DIRECT CAUSE IS UNKNOWN. A DEPLOYMENT DEFECT WAS NOT OBSERVED. THIS IS A PTA INTENDED TO BE CRIMPED TO THE PARENT BLOOD VESSEL. NO RESISTANCES OR ABNORMALITIES DURING DELIVERY OF THE PIPELINE WERE NOTED.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT PASSED AWAY AT (B)(6) HOSPITAL, LOCATED IN (B)(6). THE CONCOMITANT BALLOON USED IN PAT WAS THE 7-7 OF THE SHOURYU HR.Event Description: ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT A BALLOON (NON-MEDTRONIC) USED WITH THE PIPELINE WAS DAMAGED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21981427&pc=LWR | 3004478276-2025-00133 | 2024-06-16 04:00 | Death | CORCYM CANADA CORP. | 2025-05-07 04:00 | LWR | PERCEVAL SUTURELESS AORTIC HEART VALVE | Insufficient Information | Insufficient Information | P150011 |  | 1 | Event Description: MANUFACTURER WAS INFORMED THAT A PERCEVAL VALVE PVS21 WAS IMPLANTED ON (B)(6) 2019. REPORTEDLY, PATIENT PASSED AWAY ON (B)(6) 2024 DUE TO UNKNOWN REASON. PER MEDICAL JUDGMENT RECEIVED, HEPATOCELLULAR CARCINOMA, CHRONIC DIC, AND POOR NUTRITIONAL STATUS COULD BE POSSIBLE REASONS, BUT THE EXACT CAUSE CANNOT BE DETERMINED. MANUFACTURER WAS INFORMED THAT NO FURTHER INFORMATION WILL BE RECEIVED. Manufacturer Narrative: MANUFACTURER IS RETRIEVING AND REVIEWING THE PRODUCTION RECORD. A FOLLOW UP REPORT WILL BE PROVIDED UPON COMPLETION OF THIS REVIEW. Manufacturer Narrative: UPDATED FIELDS: B4, G3, G, H2, H6. THE MANUFACTURING AND MATERIAL RECORDS FOR THE DEVICE INVOLVED IN THIS EVENT, AS THEY PERTAIN TO THE REPORTED ISSUE, WERE RETRIEVED AND REVIEWED BY THE MANUFACTURER¿S QUALITY ENGINEERING. THE RESULTS CONFIRMED THAT THIS PROSTHESIS SATISFIED ALL REQUIRED MATERIAL, VISUAL, AND PERFORMANCE STANDARDS AT THE TIME OF MANUFACTURE AND RELEASE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEFINITIVE ROOT CAUSE OF THE EVENT CANNOT BE ESTABLISHED AT THIS TIME. HOWEVER, FROM THE DOCUMENT REVIEW PERFORMED, NO MANUFACTURING DEFICIENCIES WERE NOTED. PER MEDICAL JUDGMENT RECEIVED, HEPATOCELLULAR CARCINOMA, CHRONIC DIC, AND POOR NUTRITIONAL STATUS COULD BE POSSIBLE REASONS, BUT THE EXACT CAUSE CANNOT BE DETERMINED AT THIS TIME. SHOULD FURTHER INFORMATION BE RECEIVED IN THE FUTURE, A FOLLOW UP REPORT WILL BE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21981429&pc=NVN | 2017865-2025-52177 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-07 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE PATIENT DEATH. AS RECEIVED, A COMPLETE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING WAS NORMAL. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21981533&pc=NVN | 2017865-2025-52180 | 2025-03-17 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-07 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE FOR ANALYSIS. ELECTRICAL TESTING THAT COULD BE MEASURED DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21981538&pc=LWP | 2017865-2025-52179 | 2025-03-17 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-07 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. THE REASON FOR RETURN WAS DEATH UNKNOWN. DEVICE IMAGE WAS SAVED SUCCESSFULLY. ONCE THE DEVICE WAS RECEIVED, IT DISPLAYED NORMAL DEVICE CHARACTERISTICS UPON THE FIRST STAGE OF ANALYSIS. FURTHERMORE, THE DEVICE IMAGE WAS SAVED AND ANALYZED, NO DISCREPANCIES WERE NOTED. AFTER ALL ELECTRICAL TESTS PERFORMED, INCLUDING THERMAL AND MECHANICAL, THE DEVICE CONTINUES EXHIBITING NORMAL DEVICE CHARACTERISTICS. LONGEVITY ASSESSMENT WAS PERFORMED AND DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21981542&pc=NVN | 2017865-2025-52181 | 2025-03-17 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-07 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. VISUAL INSPECTION OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21981875&pc=NVY | 2017865-2025-52189 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-07 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Capturing Problem | Low Blood Pressure/ Hypotension; Pericardial Effusion | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT A HIGH PACING THRESHOLD WAS OBSERVED ON THE RIGHT VENTRICULAR (RV) LEAD. THE PATIENT PRESENTED FOR A PROCEDURE TO EXPLANT THE RV LEAD. DURING THE PROCEDURE, A DECREASE IN BLOOD PRESSURE AND PERICARDIAL EFFUSION WAS OBSERVED. THE LOW BLOOD PRESSURE WAS MANAGED BY ADDRESSING ANESTHESIA BUT PERICARDIOCENTESIS AND TRANSFUSION WERE ALSO REQUIRED TO REGAIN STABILITY. THE RV LEAD WAS EXPLANTED. THE SYSTEM WAS REMOVED WITH THE LEFT VENTRICULAR LEAD CAPPED. THE PATIENT WAS TO RETURN AT A LATER DATE TO IMPLANT A NEW SYSTEM. THE PATIENT WAS STABLE.Event Description: NEW INFORMATION RECEIVED NOTES THE PATIENT WAS FOUND TO BE UNSTABLE, IMPLANTING NEW LEADS WAS NOT POSSIBLE DUE TO THE PATIENT'S CONDITION. THE PATIENT SUBSEQUENTLY PASSED AWAY ON (B)(6) 2025, SEVERAL DAYS AFTER THE INITIAL PROCEDURE. THE PATIENT HAD NO ACTIVE ABBOTT PRODUCT AT THE TIME OF DEATH. THERE WAS NO KNOWN COMPLAINT BY THE PHYSICIAN. FURTHER INFORMATION WAS REQUESTED BUT WAS NOT AVAILABLE. Manufacturer Narrative: CORRECTION: SECTION A: PATIENT IDENTIFIERS: A1, A3 AND DATE OF BIRTH WERE UPDATED. Manufacturer Narrative: THE REPORTED EVENTS WERE HIGH PACING THRESHOLD AND CARDIAC PERFORATION. AS RECEIVED, A PARTIAL LEAD (ONLY DISTAL PORTION) WAS RETURNED IN ONE PIECE WITH THE HELIX NOTED STRETCHED. THE REPORTED EVENT OF HIGH PACING THRESHOLD WAS NOT CONFIRMED. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. THE HELIX MECHANISM/ EXTENSION LENGTH TEST AND TIP STIFFNESS TEST NOT PERFORMED DUE TO CONDITION OF THE PARTIAL LEAD, AS RECEIVED. VISUAL AND X-RAY EXAMINATIONS OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. Manufacturer Narrative: THE REPORTED EVENTS WERE HIGH PACING THRESHOLD, PERICARDIAL EFFUSION AND PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY DISTAL PORTION) WAS RETURNED IN ONE PIECE WITH THE HELIX NOTED STRETCHED. THE REPORTED EVENT OF HIGH PACING THRESHOLD WAS NOT CONFIRMED. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. THE HELIX MECHANISM/ EXTENSION LENGTH TEST AND TIP STIFFNESS TEST NOT PERFORMED DUE TO CONDITION OF THE PARTIAL LEAD, AS RECEIVED. VISUAL AND X-RAY EXAMINATIONS OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21981960&pc=LJT | 1317056-2025-00124 | 2022-04-03 04:00 | Death | ANGIODYNAMICS, INC | 2025-05-07 04:00 | LJT | BIOFLO VORTEX | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Endocarditis; Muscle Weakness; Pericardial Effusion | K190559 |  | 1 | Event Description: ON (B)(6) 2021, THE DECEDENT, (B)(6), UNDERWENT PLACEMENT OF THE DEFENDANTS' SMARTPORT POWER INJECTABLE PORT PRODUCT, I.E., SMARTPORT PLASTIC PORT: MODEL/CAT NUMBER: H787CT80LPPDVI0; LOT #: 5708017, FOR HIS CHEMOTHERAPY TREATMENT. UPON INFORMATION AND BELIEF, THE AFOREMENTIONED DEVICE'S DESCRIPTION IS AS FOLLOWS: DETACHED 8F POLY CATH, VI INTRODUCER, & FILLED SH OR SMART PORT PLASTIC, 8F SINGLE PLASTIC PORT FILLED NON-VALVED PG, OR 8FRENCH ANGIODYNAMICS SMARTPORT POWER INJECTABLE PORT. UPON INFORMATION AND BELIEF, ON (B)(6) 2021, THE DEVICE AT ISSUE WAS IMPLANTED BY DR. (B)(6), MD, AT (B)(6) HOSPITAL IN (B)(6), WISCONSIN. UPON INFORMATION AND BELIEF, ON OR ABOUT (B)(6) 2022, THE DECEDENT PRESENTED TO THE EMERGENCY DEPARTMENT AT (B)(6) HOSPITAL IN (B)(6), WISCONSIN, IN AN AMBULANCE, WITH COMPLAINTS OF WEAKNESS, AND WAS THEN ADMITTED WITH PROGRESSIVE WEAKNESS AND OTHER ACCOMPANYING SYMPTOMS. UPON INFORMATION AND BELIEF, ON OR ABOUT (B)(6) 2022, THE CULTURE OF THE DECEDENT'S PORT-A-CATH WAS ORDERED. UPON INFORMATION AND BELIEF, THE DECEDENT'S BLOOD CULTURE AND THE BLOOD CULTURE AEROBIC PANEL CAME BACK AS "ABNORMAL". UPON INFORMATION AND BELIEF, THE DECEDENT'S AFOREMENTIONED "ABNORMAL" BLOOD CULTURES' RESULTS SHOWED THE PRESENCE OF STAPHYLOCOCCUS AND STAPHYLOCOCCUS AUREUS. UPON INFORMATION AND BELIEF, ON (B)(6) 2022, THE CT CHEST WAS ORDERED BY THE INFECTIOUS DISEASE DEPARTMENT. UPON INFORMATION AND BELIEF, ON (B)(6) 2022, THE AFOREMENTIONED CT CHEST PERFORMED ON THE DECEDENT SHOWED THE FOLLOWING RELEVANT FINDINGS: "INFECTED PORT/ENDOCARDITIS"; "MSSA BACTEREMIA PRELIMINARY". "VEGETATION ON PORT", "PERICARDIAL EFFUSION [OR ITS POTENTIAL COMPLICATION PERICARDIAL] TAMPONADE", ETC. UPON INFORMATION AND BELIEF, ON OR ABOUT (B)(6) 2022, DR. (B)(6), MD, AN INFECTIOUS DISEASES PHYSICIAN, NOTED THAT THE DECEDENT'S BLOOD CULTURE CAME BACK AS STAPHYLOCOCCUS AUREUS AND THAT THE DECEDENT HAD VEGETATION ON [HIS] PORT. DR. (B)(6) ALSO NOTED THAT THE DECEDENT'S BLOOD CULTURE CAME BACK AS MSSA [MSSA IS A COMMON MEDICAL ABBREVIATION FOR METHICILLIN-SUSCEPTIBLE STAPHYLOCOCCUS AUREUS]. DR. (B)(6) DID ALSO UNDERLINE THAT THERE WERE CLEARLY SEPTIC EMBOLI [FROM STAPHYLOCOCCUS]. UPON INFORMATION AND BELIEF, ON (B)(6) 2022, THE DECEDENT'S EC ECHO COMPLETE W/DOPPLER (I.E., DOPPLER ECHOCARDIOGRAPHY), ORDERED BY (B)(6), PA-C, FOUND LIKELY VEGETATION ON THE DECEDENT'S RIGHT ATRIAL CATHETER. ON (B)(6) 2022, DR. (B)(6), MD, WHO, UPON INFORMATION AND BELIEF, REVIEWED & ANALYZED THE RESULTS OF THE AFOREMENTIONED DOPPLER ECHOCARDIOGRAPHY COULD NOT "RULE OUT SMALL VEGETATION ON THE AORTIC VALVE" OF THE DECEDENT. UPON INFORMATION AND BELIEF, ON OR ABOUT (B)(6) 2022, DR. (B)(6), M.D., FCCM UNDERLINED THAT THE DECEDENT'S ILLNESS AT ISSUE WAS DUE TO "MSSA ENDOCARDITIS ASSOCIATED WITH THE PORT INFECTION, PULMONARY SEPTIC EMBOLI, POSSIBLY PLEURAL AND PERICARDIAL INFECTION." UPON INFORMATION AND BELIEF, ON OR ABOUT (B)(6) 2025, THE DECEDENT'S PHYSICIAN, DR. (B)(6), M.D., FCCM, WHILE DOCUMENTING THE DECEDENT'S PORT-RELATED DIAGNOSES, NOTED THAT THE DECEDENT WOULD REQUIRE CARDIAC SURGERY AND PROLONGED ANTIBIOTIC SUPPORT, HOWEVER DR. (B)(6) ALSO OPINED THAT THE DECEDENT WOULD NOT HAVE SURVIVED THE REQUIRED CARDIAC SURGERY. ON (B)(6) 2022, (B)(6) DIED AT (B)(6) HOSPITAL, IN (B)(6), WISCONSIN. Manufacturer Narrative: THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF PATIENT INFECTION AND THROMBOSIS (AND SUBSEQUENT EXPIRATION) CANNOT BE CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THIS SERIOUS ADVERSE EVENT (SAE). NO PORT DEVICE WAS RETURNED FOR EVALUATION SINCE THERE WAS NO REPORT OF PORT DEVICE MALFUNCTION, DAMAGE OR PERFORMANCE ISSUE DURING USE. A DEVICE HISTORY RECORD (DHR) REVIEW OF THE INDICATED PACKAGING LOT REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. STERILIZATION LOAD RELEASE RECORDS WERE REVIEWED FOR THE PACKAGING LOT IN QUESTION; NO ISSUES WERE OBSERVED. THE PORT WAS IMPLANTED INTO THE PATIENT MORE THAN 3 MONTHS PRIOR TO THE PATIENT'S INFECTION. THERE WAS NO REPORT OF PORT DEVICE MALFUNCTION OR PERFORMANCE ISSUE DURING USE. IF THE MANUFACTURING/PACKAGING/STERILIZATION OF THE PORT DEVICE IN QUESTION WAS TO CONTRIBUTE TO AN INFECTION IN THE PATIENT IT WOULD HAVE LIKELY OCCURRED SHORTLY AFTER THE IMPLANT DATE AND NOT MORE THAN 3 MONTHS LATER. DEVICE DIRECTIONS FOR USE (DFU) CAUTIONS THAT EACH ACCESS OF THE PORT SHOULD BE PERFORMED USING ASEPTIC TECHNIQUE, FOLLOWING THE INSTITUTIONS UNIVERSAL PRECAUTIONS. THERE IS NO INDICATION FROM THE REPORTED COMPLAINT THAT THE MANUFACTURING/PACKAGING/STERILIZATION OF THE PORT DEVICE COULD HAVE CONTRIBUTED TO THE PATIENT'S INFECTION. INFECTION AND THROMBOSIS IS CAUTIONED IN THE DEVICE DFU AS A POTENTIAL COMPLICATION OF PORT SYSTEM USE. LABELING REVIEW: THE DIRECTION FOR USE (DFU) PROVIDED WITH THE PORT DEVICE CONTAINS THE FOLLOWING STATEMENTS: CONTRAINDICATIONS: PRESENCE OF INFECTION, BACTEREMIA, OR SEPTICEMIA. WARNINGS: THE DEVICE IS TO BE IMPLANTED, USED, MAINTAINED AND REMOVED IN ACCORDANCE WITH CURRENT EVIDENCE-BASED GUIDANCE FOR INFECTION PREVENTION FROM AGENCIES/ORGANIZATIONS SUCH AS CENTERS OF DISEASE CONTROL (CDC) OR THE WORLD HEALTH ORGANIZATION. IT IS RECOMMENDED THAT INSTITUTIONAL POLICIES AND PROCEDURES ARE UPDATED PERIODICALLY TO REFLECT CURRENT EVIDENCE-BASED GUIDELINES FOR INFECTION PREVENTION. CONTAMINATION OF THE DEVICE MAY LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. PRECAUTIONS: CAREFULLY READ AND FOLLOW ALL INSTRUCTIONS PRIOR TO USE AFTER IMPLANTATION OR USAGE OF THE PORT, THE SYSTEM SHOULD BE FLUSHED TO CLEAR INFUSATES AND/OR BLOOD COMPONENTS IN ORDER TO MAINTAIN PATENCY. IF MORE THAN ONE DRUG IS TO BE ADMINISTERED, BETWEEN THE INDIVIDUAL DRUG APPLICATIONS, FLUSH THE SYSTEM WITH 5 TO 10 ML NORMAL SALINE FOR INJECTION TO PREVENT DRUG INTERACTION. POTENTIAL COMPLICATIONS: BACTEREMIA. CATHETER THROMBOSIS. DEATH. INFLAMMATION. INFECTION. THROMBOEMBOLISM. THROMBOPHLEBITIS. TUNNEL INFECTION. VASCULAR THROMBOSIS. USE AND MAINTENANCE: AFTER IMPLANTATION OR USAGE OF THE PORT, THE SYSTEM SHOULD BE FLUSHED AND LOCKED TO CLEAR INFUSATES AND/OR BLOOD COMPONENTS IN ORDER TO MAINTAIN PATENCY. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT WILL CONTINUE TO BE MONITORED FOR TRENDS. REFERENCE (B)(4). Manufacturer Narrative: THE DEVICE IS NOT AVAILABLE TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. REFERENCE (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21965607&pc=FRN | 2016493-2025-78261 | 2025-01-11 05:00 | Death | CAREFUSION SD | 2025-05-06 04:00 | FRN | ALARIS SYSTEM | Contamination; Corroded; Fracture; Excess Flow or Over-Infusion | Insufficient Information | K133532 |  | 1 | Event Description: IT WAS INITIALLY REPORTED THE DEVICE HAD AN EVENT WHERE THE PUMP DELIVERED AN INCORRECT DOSE. THE DEVICE WAS INVESTIGATED BY THE FACILITY'S BIOMEDICAL ENGINEERING AND PER REPORT, IT "PASSED PREVENTATIVE MAINTENANCE AND RETURNED BACK TO SERVICE". THERE WAS PATIENT INVOLVEMENT HOWEVER OUTCOME IS UNKNOWN. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE EVENT INVOLVED A "COMFORT CARE PATIENT" ON A CONTINUOUS MORPHINE DRIP INFUSING AT 5MG/HR. FROM A 100ML (CONCENTRATION: 1MG/1ML) BAG. THE INFUSION WAS INITIATED AT 12:14. AT 14:09 THE NURSE RESPONDED TO THE ROOM AND FOUND THAT THE BAG WAS EMPTY. THE BAG AND TUBING WERE INSPECTED BY CUSTOMER; NO LEAKS WERE NOTED PER REPORT. IT WAS REPORTED THE PATIENT LATER PASSED AWAY. THE PATIENT'S FAMILY REQUESTED NO FURTHER INTERVENTIONS. THE EVENT OCCURRED ON 11 JANUARY 2025. PER CUSTOMER, THE EVENT WAS NOT IMMEDIATELY REPORTED TO BD AS THEY "DID NOT HAVE REASON TO BELIEVE THERE WAS AN EQUIPMENT MALFUNCTION." THE CUSTOMER, AT THE TIME, BELIEVED THAT THE ISSUE WAS DUE TO "DRUG DIVERSION". HOWEVER, THE SAME DEVICE HAS BEEN INVOLVED IN ANOTHER (CAPTURED UNDER SEPARATE RECORD MFR REPORT # 2016493-2025-77596) AND THIS PROMPTED THE CUSTOMER TO REPORT THE EVENT TO BD.Event Description: IT WAS INITIALLY REPORTED THE DEVICE HAD AN EVENT WHERE THE PUMP DELIVERED AN INCORRECT DOSE. THE DEVICE WAS INVESTIGATED BY THE FACILITY'S BIOMEDICAL ENGINEERING AND PER REPORT, IT "PASSED PREVENTATIVE MAINTENANCE AND RETURNED BACK TO SERVICE". THERE WAS PATIENT INVOLVEMENT HOWEVER OUTCOME IS UNKNOWN. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE EVENT INVOLVED A "COMFORT CARE PATIENT" ON A CONTINUOUS MORPHINE DRIP INFUSING AT 5MG/HR. FROM A 100ML (CONCENTRATION: 1MG/1ML) BAG. THE INFUSION WAS INITIATED AT 12:14. AT 14:09 THE NURSE RESPONDED TO THE ROOM AND FOUND THAT THE BAG WAS EMPTY. THE BAG AND TUBING WERE INSPECTED BY CUSTOMER; NO LEAKS WERE NOTED PER REPORT. IT WAS REPORTED THE PATIENT LATER PASSED AWAY. THE PATIENT'S FAMILY REQUESTED NO FURTHER INTERVENTIONS. THE EVENT OCCURRED ON (B)(6) 2025. PER CUSTOMER, THE EVENT WAS NOT IMMEDIATELY REPORTED TO BD AS THEY "DID NOT HAVE REASON TO BELIEVE THERE WAS AN EQUIPMENT MALFUNCTION." THE CUSTOMER, AT THE TIME, BELIEVED THAT THE ISSUE WAS DUE TO "DRUG DIVERSION". HOWEVER, THE SAME DEVICE HAS BEEN INVOLVED IN ANOTHER (CAPTURED UNDER SEPARATE COMPLAINT RECORD (B)(4)) AND THIS PROMPTED THE CUSTOMER TO REPORT THE EVENT TO BD. Manufacturer Narrative: OMIT : B17 - DEVICE NOT RETURNED, C20 - NO FINDINGS AVAILABLE. CORRECTION: DESCRIBE EVENT OR PROBLEM. ADDITIONAL INFORMATION: DEVICE AVAILABLE FOR EVAL?, RETURNED TO MANUFACTURER ON, DEVICE EVAL BY MANUFACTURER?, IMDRF ANNEX A, B, C, D, G CODES, REMEDIAL ACTION REQUIRED, REMEDIAL ACTION # AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY THE REPORTED OVER INFUSION OF MORPHINE WAS NOT CONFIRMED THROUGH REVIEW OF THE DEVICE LOGS; HOWEVER, OVER INFUSION WAS REPLICATED DURING DEVICE INSPECTION AND TESTING. INSPECTION OF THE SUSPECT PUMP MODULE (SN (B)(6)) FOUND THE UPPER LEFT BEZEL POST TO BE SEPARATED AND ALL THREE (3) OF THE RIGHT BEZEL POSTS TO BE BROKEN AND SEPARATED COMPLETELY. UNREGULATED FLOW WAS OBSERVED DURING TESTING. IT WAS NOTED THAT THE BEZEL ASSEMBLY DATE CODE WAS NOTED AS 0114 (JANUARY 2014) AND IS RECALL AFFECTED FOR FR-110 PLASTIC. BD RELEASED RECALL NOTIFICATION MMS-21-4100 IN JUNE OF 2021 THAT INFORMED FACILITIES ABOUT PUMP MODULES THAT WERE MANUFACTURED BETWEEN APRIL OF 2011 TO JUNE OF 2017 WITH THE PLASTIC MATERIAL FR-110 HAVING A POTENTIAL TO SEPARATE AT THE BEZEL POSTS. THE SEPARATION OF ONE OR MORE POSTS MAY RESULT IN FREE FLOW, OVER INFUSION, UNDER INFUSION AND INTERRUPTION OF INFUSION. REVIEW OF THE PCU EVENT LOG DID NOT SHOW ANY RECORDED EVENTS ON 11JAN2025 OR PROGRAMMING MATCHING THE REPORTED EVENTS AROUND SAID DATE. THE ONLY RECORDED INFUSION OF MORPHINE (DRUGID 190) WAS PROGRAMMED ON A DIFFERENT PUMP MODULE (SN (B)(6)) ON THE DATE (B)(6) 2025. THE MORPHINE INFUSION THAT WAS PROGRAMMED ON A DIFFERENT PUMP MODULE (SN (B)(6)) WAS PROGRAMMED AT A CONCENTRATION OF 100 MG / 100 ML AND STARTED ON (B)(6) 2025 AT 3:24 PM. THE RATE WAS PROGRAMMED AT 2 ML/HR, CALCULATING THE DOSE AT 2 MG/HR AND THE VOLUME TO BE INFUSED (VTBI) WAS 100 ML. ROOT CAUSE THE REPORTED OVER INFUSION OF MORPHINE WAS NOT CONFIRMED SINCE THE PUMP MODULE INVOLVED (SN (B)(6)) WAS NOT SENT BACK FOR INVESTIGATION. THE PUMP MODULE THAT WAS SENT BACK FOR INVESTIGATION (SN (B)(6)) WAS FOUND TO HAVE BROKEN/SEPARATED BEZEL BOSS POSTS WHICH COULD LEAD TO FLOW RATE INACCURACIES (FREE FLOW, OVER INFUSION, UNDER INFUSION AND INTERRUPTION OF INFUSION). NOTE THAT THIS REPORT LISTS IMDRF ANNEX A040502, A1801, G02017, G04058, G04067, G04037, C0601, D01 CODES NOT ASSOCIATED WITH THE REPORTED EVENT BUT IDENTIFIED AS REPORTABLE MALFUNCTIONS OBSERVED ON THE DEVICE DURING INVESTIGATION ARE UNRELATED TO THE REPORTED ISSUE. THESE OTHER FAILURES PRESUMPTIVELY DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED ISSUE. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. Manufacturer Narrative: THE ACTUAL DATE OF DEATH IS UNKNOWN. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21965715&pc=QZI | 3002648230-2025-00392 | 2025-04-17 04:00 | Death | MEDTRONIC CRYOCATH LP | 2025-05-06 04:00 | QZI | PULSESELECT | Material Twisted/Bent; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Cardiac Tamponade; Pericardial Effusion | P230017 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A CARDIAC ABLATION PROCEDURE PATIENT PASSED AWAY. THE PHYSICIAN SUCCESSFULLY ABLATED THE LEFT ATRIAL POSTERIOR WALL, LEFT ATRIAL ANTERIOR WALL, AND MITRAL ISTHMUS. HOWEVER, UPON DELIVERING THE FIFTH APPLICATION NEAR THE LEFT ATRIAL APPENDAGE, THE PATIENT'S BLOOD PRESSURE RAPIDLY DROPPED AND WAS LOST. PATIENT THEN WENT INTO ASYSTOLE. CARDIOPULMONARY RESUSCITATION WAS IMMEDIATELY INITIATED. A PERICARDIOCENTESIS WAS PERFORMED TO DRAIN FLUID. DEFIBRILLATION WAS ALSO ATTEMPTED, BUT THE PATIENT'S RHYTHM COULD NOT BE RESTORED. LIFE-SAVING MEASURES CONTINUED FOR APPROXIMATELY SIXTY MINUTES BEFORE THE PATIENT WAS PRONOUNCED DECEASED.Event Description: IT WAS LATER REPORTED THAT THE PATIENT WAS IN PULSELESS ELECTRICAL ACTIVITY WHEN THE BLOOD PRESSURE DROPPED. PRIOR TO THE PROCEDURE, INTRACARDIAC ECHOCARDIOGRAPHY (ICE) DID NOT REVEAL ANY PERICARDIAL EFFUSION. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PSCC100 LOOP CATHETER WITH LOT NUMBER: 0012446760 WAS RETURNED AND ANALYZED. NO ANOMALIES WERE IDENTIFIED DURING EXTERNAL VISUAL INSPECTION OF THE SHAFT SEGMENT. VISUAL INSPECTION OF THE SPIRAL ARRAY SEGMENT REVEALED THE GUIDEWIRE LUMEN WAS KINKED AT 0.442 INCHES FROM THE DISTAL TIP. ON THE HANDLE SEGMENT, THE CAPTURE DEVICE ADAPTER WAS OBSERVED TO BE DETACHED (REMOVED AND NOT RETURNED WITH CATHETER). THE PRINTED CIRCUIT BOARD ASSEMBLY (PCBA) CHIP WAS REPROGRAMMED FOR FUNCTIONAL TESTING. THE CATHETER PASSED PERFORMANCE FUNCTIONAL TESTING WITH THE GENERATOR; THERAPY DELIVERIES WERE COMPLETED WITH NO SYSTEM ERRORS GENERATED. NO PERFORMANCE ISSUES WERE IDENTIFIED. DEFLECTION TESTING (ROTATING THE STEERING KNOB CLOCKWISE AND COUNTER-CLOCKWISE) WAS PERFORMED, AND THE DISTAL CATHETER TIP RESPONDED WITH APPROXIMATELY 170° ROTATION IN BOTH DIRECTIONS. NO ANOMALIES WERE OBSERVED. THE SLIDE CONTROL FUNCTION OPERATED AS EXPECTED WITHOUT ANY ISSUES. UPON FULL ADVANCEMENT OF THE SLIDE CONTROL TO EXTEND THE GUIDEWIRE LUMEN, NO KINKS OR OTHER ANOMALIES WERE OBSERVED ALONG THE EXTENDED PORTION. ELECTRICAL CONTINUITY TESTING DID NOT REVEAL ANY ANOMALIES. IN CONCLUSION, THE CATHETER DID NOT PASS THE RETURNED PRODUCT INSPECTION DUE TO A DETACHED ADAPTER TO THE CAPTURE DEVICE AND A GUIDEWIRE LUMEN KINK IN THE SHAFT REGION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21965718&pc=DRA | 3002648230-2025-00393 | 2025-04-17 04:00 | Death | MEDTRONIC CRYOCATH LP | 2025-05-06 04:00 | DRA | FLEXCATH CONTOUR¿ STEERABLE SHEATH | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Cardiac Tamponade; Pericardial Effusion | K232321 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A CARDIAC ABLATION PROCEDURE PATIENT PASSED AWAY. THE PHYSICIAN SUCCESSFULLY ABLATED THE LEFT ATRIAL POSTERIOR WALL, LEFT ATRIAL ANTERIOR WALL, AND MITRAL ISTHMUS. HOWEVER, UPON DELIVERING THE FIFTH APPLICATION NEAR THE LEFT ATRIAL APPENDAGE, THE PATIENT'S BLOOD PRESSURE RAPIDLY DROPPED AND WAS LOST. PATIENT THEN WENT INTO ASYSTOLE. CARDIOPULMONARY RESUSCITATION WAS IMMEDIATELY INITIATED. A PERICARDIOCENTESIS WAS PERFORMED TO DRAIN FLUID. DEFIBRILLATION WAS ALSO ATTEMPTED, BUT THE PATIENT'S RHYTHM COULD NOT BE RESTORED. LIFE-SAVING MEASURES CONTINUED FOR APPROXIMATELY SIXTY MINUTES BEFORE THE PATIENT WAS PRONOUNCED DECEASED.Event Description: IT WAS LATER REPORTED THAT THE PATIENT WAS IN PULSELESS ELECTRICAL ACTIVITY WHEN THE BLOOD PRESSURE DROPPED. PRIOR TO THE PROCEDURE, INTRACARDIAC ECHOCARDIOGRAPHY (ICE) DID NOT REVEAL ANY PERICARDIAL EFFUSION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21966113&pc=MHX | 1218950-2025-000283 | 2025-04-08 04:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-05-06 04:00 | MHX | INTELLISPACE EVENT MGMT PLATFORM | No Audible Alarm | Cardiac Arrest | K102974 |  | 1 | Event Description: IT WAS REPORTED THE DEVICE DID NOT GENERATE AN AUDIBLE ALARM ON THE EMERGIN PAGER AFTER THE PATIENT WENT INTO CARDIAC ARREST. IT WAS INDICATED THE VISUAL ALARM MESSAGE WAS GENERATED BUT NO AUDIBLE BEEP ON THE PAGING DEVICE. BIOMED TESTED THE ALARM CHAIN, REVEALING ALARMS OCCURRED VISUALLY AND AUDIBLY AT THE CENTRAL STATION AND ON THE EMERGIN PAGING DEVICE. IT WAS REPORTED THE NURSES BEGAN CARDIAC MASSAGE. IMMEDIATELY FOLLOWING THE ALARM AT THE CENTRAL STATION, THE USER DISCOVERED THE PATIENT IN CARDIAC ARREST, FROM WHICH THE PATIENT PASSED AWAY. THE LOGS WERE REVIEWED, REVEALING THE ALARM MESSAGE WAS DISTRIBUTED TO THE THREE PAGERS, WITH THE ALARM MESSAGE BEING FOUND ON THE PAGER SCREENS. THE CUSTOMER WAS NOT SURE WHETHER THE DEVICES PROVIDED AN AUDIBLE ALERT. Manufacturer Narrative: THE LOGS HAVE BEEN REVIEWED BY A PRODUCT SUPPORT ENGINEER (PSE). IT HAS BEEN ESTABLISHED THAT MULTIPLE ALERTS WERE SENT OUT OF THE IEM SERVER DURING THE DATE AND TIME OF THE EVENT. THE IEM SERVER PROCESSED THE ALERTS AND SENT THEM TO ASCOM FOR DELIVERY TO THE BEEPERS. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. REPORTING INSTITUTION PHONE NUMBER: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21966498&pc=MVK | 3008642652-2025-04510 | 2025-02-12 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-06 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION. THE DEVICE WAS STARTED UP AT 05:08:16 ON (B)(6) 2025. AT 06:20:43, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 150 BPM WITH CPR/MOTION ARTIFACT. THE LIFEVEST PROPERLY DETECTED VT. CPR/MOTION ARTIFACT AND VT RATES VARYING/UNDER THRESHOLD PREVENTED THE LIFEVEST DELIVERING A TREATMENT. AT 06:23:06, THE PATIENT RECEIVED THE NON-LIFEVEST DEFIBRILLATION. RHYTHM AT THE TIME OF THE NON-LIFEVEST DEFIBRILLATION WAS VT @ 140 BPM WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT DEGRADING TO VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH CPR/MOTION ARTIFACT FOR 18 SECONDS TRANSITIONING TO AF FROM 65-15. THE PATIENT RECEIVED THE NON-LIFEVEST DEFIBRILLATION 24 SECONDS INTO THE DETECTION SEQUENCE. THE DEVICE WAS SHUT DOWN AT 06:44:35 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21966963&pc=OZD | 1220648-2025-28098 | 2025-04-14 04:00 | Death | ABIOMED, INC. | 2025-05-06 04:00 | OZD | IMPELLA 5.5 | Positioning Failure; Deformation Due to Compressive Stress | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT AN IMPELLA 5.5 WAS IMPLANTED IN A PATIENT FOUND UNRESPONSIVE IN THEIR HOME AND WHO HAD RECEIVED CARDIOPULMONARY RESUSCITATION FOR TWO HOURS. UPON INSERTION IT WAS FOUND THAT THE IMPELLA GUIDEWIRE WAS KINKED. THE GUIDEWIRE WAS EXCHANGED AND THE IMPELLA WAS BACKLOADED IN STANDARD FASHION. THE IMPELLA WAS NOTED TO BE DEEP AND WAS REPOSITIONED. THE PATIENT'S FAMILY ULTIMATELY WITHDREW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES AND POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE FAILURE MODE (FM): PRODUCT DAMAGE (CANNULA KINK) CHANGED TO FM: DELIVERY ISSUE SINCE THE GUIDEWIRE KINK PRIOR TO INSERTION PER SALESFORCE REPORTED INFORMATION. DELIVERY ISSUE: THE CAUSE OF THE DELIVERY ISSUE CANNOT BE DETERMINED SINCE COMPLAINT DEVICE NOT RETURNED FOR ANALYSIS AND INSUFFICIENT CLINICAL DATA PROVIDED. POSITIONING ISSUES: THE CAUSE OF THE POSITIONING ISSUES CANNOT BE DETERMINED SINCE INSUFFICIENT CLINICAL DATA PROVIDED WHY POSITION ISSUE OCCURRED. INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. IMPELLA CP WITH SMARTASSIST SYSTEM WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER: SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA CP WITH SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER: SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21967296&pc=OZD | 1220648-2025-28104 | 2025-02-11 05:00 | Death | ABIOMED, INC. | 2025-05-06 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Sepsis; Tachycardia | P140003 |  | 1 | Event Description: IT WAS REPORTED VIA ABIOMED'S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY THAT A PATIENT ENDURED RECURRING NON-SUSTAINED VENTRICULAR TACHYCARDIA (VT). THE VT WAS BELIEVED TO HAVE A POSSIBLE RELATIONSHIP TO THE IMPELLA DEVICE. AS SUPPORT CONTINUED, THE PATIENT WAS NON-COMPLIANT, AND THE DECISION WAS ULTIMATELY MADE TO WITHDRAW CARE. THE PATIENT WAS TRANSFERRED TO PALLIATIVE CARE AND PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME.Event Description: UPON LATER NOTIFICATION VIA LOQI, IT WAS DOCUMENTED THAT ON (B)(6) 2025, THE PATIENT ENDURED SEPSIS BELIEVED TO HAVE A POSSIBLE RELATIONSHIP TO THE IMPELLA DEVICE USED. A BLOOD CULTURE CAME BACK POSITIVE FOR ENTEROBACTER CLOACAE. ANTIBIOTICS WERE GIVEN TO TREAT THE CONDITION Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: B5 UPDATED WITH ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. H6 HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE UPDATED TO REFLECT THE ADDITIONAL ADVERSE EVENT THAT WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21967381&pc=OZD | 1220648-2025-28105 | 2025-04-14 04:00 | Death | ABIOMED, INC. | 2025-05-06 04:00 | OZD | IMPELLA CP | Positioning Failure |  | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP. DURING THE PEEL AWAY SHEATH REMOVAL, THE PUMP WAS PULLED OUT OF LEFT VENTRICLE. AFTER SEVERAL ATTEMPTS, THE DOCTOR WAS ABLE TO RECROSS INTO THE LEFT VENTRICLE. THE PUMP WAS SUTURED IN A FORWARD FASHION. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE CAUSE OF THE POSITIONING ISSUES MOST LIKELY WAS USE RELATED DUE TO IMPROPER TECHNIQUE SINCE DURING PEEL AWAY SHEATH REMOVAL PUMP PULLED OUT OF LV AND WAS ABLE TO RECROSS IT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21967666&pc=QZI | 3002648230-2025-00401 | 2025-04-11 04:00 | Death | MEDTRONIC CRYOCATH LP | 2025-05-06 04:00 | QZI | PULSESELECT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P230017 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. IT WAS REPORTED THAT POST A PULSE FIELD ABLATION PROCEDURE, THE PATIENT EXPERIENCED APPARENT STROKE-LIKE SYMPTOMS AFTER WAKING FROM ANESTHESIA IN THE HOLDING AREA. THE PATIENT WAS SUBSEQUENTLY TAKEN TO INTERVENTIONAL RADIOLOGY (IR) FOR REVASCULARIZATION AND ADMITTED TO THE INTENSIVE CARE UNIT. DESPITE THE INTERVENTION, THE PATIENT CONTINUED TO SUFFER FROM GLOBAL DYSPHASIA AND STROKE-LIKE SYMPTOMS. THE PATIENT THEN UNDERWENT HEMORRHAGIC CONVERSION AND EXPERIENCED A SEIZURE. THREE DAYS LATER DO NOT RESUSCITATE (DNR) ORDERS WERE PUT IN PLACE FOR THE PATIENT AND THE PATIENT THEN PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21967677&pc=FKX | 1416980-2025-02699 | 2025-04-08 04:00 | Death | VANTIVE HEALTHCARE | 2025-05-06 04:00 | FKX | HOMECHOICE APD SYSTEM | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K102936 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. IT WAS NOT REPORTED IF THE PATIENT WAS HOSPITALIZED PRIOR TO DEATH. IT WAS REPORTED THAT THE PATIENT WAS CONNECTED DURING THE TIME OF THE EVENT. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. IT WAS NOT REPORTED IF PD THERAPY WAS ONGOING PRIOR TO DEATH OR AT THE TIME OF DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 AND H11. H11: THE DEVICE WAS NOT RECEIVED FOR EVALUATION; THEREFORE A DEVICE ANALYSIS COULD NOT BE COMPLETED. A SERVICE HISTORY REVIEW REVEALED NO INDICATION THAT THE PARTS REPLACED DURING SERVICING CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE REPORTED EVENT WAS NOT VERIFIED. THE CAUSE OF THE CONDITION COULD NOT BE DETERMINED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21967681&pc=DRA | 3002648230-2025-00402 | 2025-04-11 04:00 | Death | MEDTRONIC CRYOCATH LP | 2025-05-06 04:00 | DRA | FLEXCATH CONTOUR¿ STEERABLE SHEATH | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | K232321 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. IT WAS REPORTED THAT POST A PULSE FIELD ABLATION PROCEDURE, THE PATIENT EXPERIENCED APPARENT STROKE-LIKE SYMPTOMS AFTER WAKING FROM ANESTHESIA IN THE HOLDING AREA. THE PATIENT WAS SUBSEQUENTLY TAKEN TO INTERVENTIONAL RADIOLOGY (IR) FOR REVASCULARIZATION AND ADMITTED TO THE INTENSIVE CARE UNIT. DESPITE THE INTERVENTION, THE PATIENT CONTINUED TO SUFFER FROM GLOBAL DYSPHASIA AND STROKE-LIKE SYMPTOMS. THE PATIENT THEN UNDERWENT HEMORRHAGIC CONVERSION AND EXPERIENCED A SEIZURE. THREE DAYS LATER DO NOT RESUSCITATE (DNR) ORDERS WERE PUT IN PLACE FOR THE PATIENT AND THE PATIENT THEN PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21967990&pc=DTZ | 8010762-2025-0000200 | 2024-11-16 05:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-05-06 04:00 | DTZ | HLS SET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN SPAIN. MAQUET CARDIOPULMONARY GMBH HAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICE ON (B)(6) 2025: THE VIGILANCE UNIT OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT, INVOLVING ¿PLS SET¿, UDI-DI (B)(4) MANUFACTURED BY GETINGE WHICH OCCURRED ON (B)(6) 2024, IN HOSPITAL (B)(6), SPAIN. THE FOLLOWING INFORMATION WAS PROVIDED: INCIDENT DESCRIPTION: DEATH." THE GETINGE SALES AND SERVICE UNIT (SSU) CONFIRMED ON (B)(6) 2025, THAT THE AFFECTED PRODUCT IN QUESTION WAS AN HLS SET NOT A PLS SET. BASED ON THE INFORMATION AVAILABLE THAT THE PATIENT EXPIRED, A REPORT IS REQUIRED. COMPLAINT ID #(B)(4). Manufacturer Narrative: MAQUET CARDIOPULMONARY GMBH HAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICE ON (B)(6) 2025: "THE VIGILANCE UNIT OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT, INVOLVING ¿PLS SET¿, UDI-DI (B)(4) MANUFACTURED BY GETINGE WHICH OCCURRED ON (B)(6) 2024, IN HOSPITAL (B)(6), SPAIN. THE FOLLOWING INFORMATION WAS PROVIDED: INCIDENT DESCRIPTION: DEATH." BASED ON THE INFORMATION AVAILABLE THAT THE PATIENT EXPIRED, A REPORT IS REQUIRED. HOWEVER, THE SSU CONFIRMED ON (B)(6) 2025, THAT THE AFFECTED PRODUCT IN QUESTION WAS AN HLS SET NOT A PLS SET. AFTER CONSULTATION WITH THE CUSTOMER, THE SSU CONFIRMED ON (B)(6) 2025, THAT THE INVOLVED HLS SET WAS FUNCTIONING AS INTENDED AND HAD NO MALFUNCTION. THE PATIENT PASSED AWAY DUE TO HIS DISORDERS. THE AFFECTED HLS SET WAS DISCARDED BY THE CUSTOMER. BASED ON THE AVAILABLE INFORMATION THE REPORTED EVENT "DEATH" COULD BE CONFIRMED BUT WAS NOT RELATED TO A DEVICE MALFUNCTION. THIS COMPLAINT WAS FOUND IN THE DATABASE OF CUSTOMER COMPLAINTS AS A SINGLE EVENT (TIMEFRAME FROM (B)(6) 2024 TILL (B)(6) 2025). THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. NOTE: THIS EVENT OCCURRED ON THE SPANISH MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21968013&pc=PNJ | 9612164-2025-02230 | 2024-08-27 04:00 | Death | MEDTRONIC IRELAND | 2025-05-06 04:00 | PNJ | MICRA¿ AV2 | Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR WEAKNESS SHORTLY AFTER THE IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. LATER, IT WAS NOTED THAT THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS NOT IDENTIFIED. Manufacturer Narrative: B3: DATE IS APPROXIMATE. YEAR IS CONFIRMED VALID. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21968075&pc=DSQ | 2916596-2025-02734 | 2025-04-10 04:00 | Death | THORATEC CORPORATION | 2025-05-06 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED WAY DUE TO A CEREBROVASCULAR ACCIDENT (CVA). THE OUTCOME WAS SAID TO NOT HAVE BEEN DEVICE OR THERAPY RELATED. THE DEVICE WAS ALSO NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21968079&pc=QZI | 2124215-2025-28922 | 2025-04-10 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-06 04:00 | QZI | FARAWAVE NAV | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hypersensitivity/Allergic reaction; Low Blood Pressure/ Hypotension; Renal Failure; Ventricular Fibrillation |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A HISTORY OF ISCHEMIC CARDIOMYOPATHY WITH HEART FAILURE. THE ATRIAL FIBRILLATION ABLATION WAS SUCCESSFUL, AND THE PATIENT WAS IN NORMAL SINUS RHYTHM (NSR). THE PHYSICIAN REMOVED THE CATHETERS, CLOSED THE GROIN INCISIONS, AND ADMINISTERED PROTAMINE. SUBSEQUENTLY, THE PATIENTS BLOOD PRESSURE (BP) DROPPED. EPINEPHRINE WAS GIVEN, AND THE BP STABILIZED. AN ECHOCARDIOGRAM WAS PERFORMED, REVEALING NO EFFUSION. HOWEVER, THE PATIENTS BP DROPPED AGAIN, AND HE LOST HIS PULSE. A CODE WAS CALLED, AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. EACH TIME THE PATIENTS BP DROPPED, HE WENT INTO VENTRICULAR FIBRILLATION (VF), AND THIS CYCLE CONTINUED. ADDITIONALLY, A LEFT AND RIGHT CARDIAC CATHETERIZATION WAS PERFORMED, CLEARING THE LEFT ANTERIOR DESCENDING ARTERY (LAD), BUT IT WAS NOT BELIEVED TO BE THE ISSUE AS THE PATIENT NEVER EXHIBITED ST ELEVATION. THE TEAM SUSPECTED A SEVERE ALLERGIC REACTION TO PROTAMINE. THE PATIENT WAS PLACED ON LIFE SUPPORT WITH LEFT AND RIGHT IMPELLA DEVICES. DUE TO LACK OF PERFUSION DURING THE CODING EPISODES, THE PATIENT DEVELOPED KIDNEY FAILURE. THE STAFF NOTED THAT THE PATIENT WAS NEVER ANOXIC, AND THE CERTIFIED REGISTERED NURSE ANESTHETIST (CRNA) OBSERVED THAT EVEN WHEN THE BP WAS DROPPING, THE PATIENT WAS FIGHTING THE VENTILATOR, ATTEMPTING TO BREATHE INDEPENDENTLY. THEREFORE, THEY BELIEVED THAT NEUROLOGICALLY, THE PATIENT WOULD BE OKAY IF HIS HEART CONDITION STABILIZED. THE FARAWAVE NAV CATHETER IS NOT EXPECTED TO RETURN FOR ANALYSIS IT WAS DISPOSED, THEREFORE THE LOT NUMBER IS NOT AVAILABLE.Event Description: IT WAS REPORTED THAT THE PATIENT HAD A HISTORY OF ISCHEMIC CARDIOMYOPATHY WITH HEART FAILURE. THE ATRIAL FIBRILLATION ABLATION WAS SUCCESSFUL, AND THE PATIENT WAS IN NORMAL SINUS RHYTHM (NSR). THE PHYSICIAN REMOVED THE CATHETERS, CLOSED THE GROIN INCISIONS, AND ADMINISTERED PROTAMINE. SUBSEQUENTLY, THE PATIENTS BLOOD PRESSURE (BP) DROPPED. EPINEPHRINE WAS GIVEN, AND THE BP STABILIZED. AN ECHOCARDIOGRAM WAS PERFORMED, REVEALING NO EFFUSION. HOWEVER, THE PATIENTS BP DROPPED AGAIN, AND HE LOST HIS PULSE. A CODE WAS CALLED, AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. EACH TIME THE PATIENTS BP DROPPED, HE WENT INTO VENTRICULAR FIBRILLATION (VF), AND THIS CYCLE CONTINUED. ADDITIONALLY, A LEFT AND RIGHT CARDIAC CATHETERIZATION WAS PERFORMED, CLEARING THE LEFT ANTERIOR DESCENDING ARTERY (LAD), BUT IT WAS NOT BELIEVED TO BE THE ISSUE AS THE PATIENT NEVER EXHIBITED ST ELEVATION. THE TEAM SUSPECTED A SEVERE ALLERGIC REACTION TO PROTAMINE. THE PATIENT WAS PLACED ON LIFE SUPPORT WITH LEFT AND RIGHT IMPELLA DEVICES. DUE TO LACK OF PERFUSION DURING THE CODING EPISODES, THE PATIENT DEVELOPED KIDNEY FAILURE. THE STAFF NOTED THAT THE PATIENT WAS NEVER ANOXIC, AND THE CERTIFIED REGISTERED NURSE ANESTHETIST (CRNA) OBSERVED THAT EVEN WHEN THE BP WAS DROPPING, THE PATIENT WAS FIGHTING THE VENTILATOR, ATTEMPTING TO BREATHE INDEPENDENTLY. THEREFORE, THEY BELIEVED THAT NEUROLOGICALLY, THE PATIENT WOULD BE OKAY IF HIS HEART CONDITION STABILIZED. THE FARAWAVE NAV CATHETER IS NOT EXPECTED TO RETURN FOR ANALYSIS IT WAS DISPOSED, THEREFORE THE LOT NUMBER IS NOT AVAILABLE. IT WAS FURTHER REPORTED THAT THE PATIENT WAS NOT STABILIZED AS THE PATIENT PASSED AWAY. THE RENAL FAILURE WAS CONFIRMED BY THE PHYSICIAN. NO DIALYSIS WAS PERFORMED AS THE PATIENT HAD PASSED AWAY. THE PATIENTS PRE-PROCEDURE EF WAS AROUND 20 TO 25 PERCENT. NO INFORMATION AVAILABLE ON PATIENT AGE AND GENDER. THERE WAS NO INFORMATION ON HOW MANY ABLATION LESIONS BUT ABLATION WAS PERFORMED ON THE CTI, SVC, PVI AND POSTERIOR WALL. THIS WAS NOT A REPEAT ABLATION PROCEDURE.Event Description: IT WAS REPORTED THAT THE PATIENT HAD A HISTORY OF ISCHEMIC CARDIOMYOPATHY WITH HEART FAILURE. THE ATRIAL FIBRILLATION ABLATION WAS SUCCESSFUL, AND THE PATIENT WAS IN NORMAL SINUS RHYTHM (NSR). THE PHYSICIAN REMOVED THE CATHETERS, CLOSED THE GROIN INCISIONS, AND ADMINISTERED PROTAMINE. SUBSEQUENTLY, THE PATIENTS BLOOD PRESSURE (BP) DROPPED. EPINEPHRINE WAS GIVEN, AND THE BP STABILIZED. AN ECHOCARDIOGRAM WAS PERFORMED, REVEALING NO EFFUSION. HOWEVER, THE PATIENTS BP DROPPED AGAIN, AND HE LOST HIS PULSE. A CODE WAS CALLED, AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. EACH TIME THE PATIENTS BP DROPPED, HE WENT INTO VENTRICULAR FIBRILLATION (VF), AND THIS CYCLE CONTINUED. ADDITIONALLY, A LEFT AND RIGHT CARDIAC CATHETERIZATION WAS PERFORMED, CLEARING THE LEFT ANTERIOR DESCENDING ARTERY (LAD), BUT IT WAS NOT BELIEVED TO BE THE ISSUE AS THE PATIENT NEVER EXHIBITED ST ELEVATION. THE TEAM SUSPECTED A SEVERE ALLERGIC REACTION TO PROTAMINE. THE PATIENT WAS PLACED ON LIFE SUPPORT WITH LEFT AND RIGHT IMPELLA DEVICES. DUE TO LACK OF PERFUSION DURING THE CODING EPISODES, THE PATIENT DEVELOPED KIDNEY FAILURE. THE STAFF NOTED THAT THE PATIENT WAS NEVER ANOXIC, AND THE CERTIFIED REGISTERED NURSE ANESTHETIST (CRNA) OBSERVED THAT EVEN WHEN THE BP WAS DROPPING, THE PATIENT WAS FIGHTING THE VENTILATOR, ATTEMPTING TO BREATHE INDEPENDENTLY. THEREFORE, THEY BELIEVED THAT NEUROLOGICALLY, THE PATIENT WOULD BE OKAY IF HIS HEART CONDITION STABILIZED. THE FARAWAVE NAV CATHETER IS NOT EXPECTED TO RETURN FOR ANALYSIS IT WAS DISPOSED, THEREFORE THE LOT NUMBER IS NOT AVAILABLE. IT WAS FURTHER REPORTED THAT THE PATIENT WAS NOT STABILIZED AS THE PATIENT PASSED AWAY. THE RENAL FAILURE WAS CONFIRMED BY THE PHYSICIAN. NO DIALYSIS WAS PERFORMED AS THE PATIENT HAD PASSED AWAY. THE PATIENTS PRE-PROCEDURE EF WAS AROUND 20 TO 25 PERCENT. NO INFORMATION AVAILABLE ON PATIENT AGE AND GENDER. THERE WAS NO INFORMATION ON HOW MANY ABLATION LESIONS BUT ABLATION WAS PERFORMED ON THE CTI, SVC, PVI AND POSTERIOR WALL. THIS WAS NOT A REPEAT ABLATION PROCEDURE. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. IT WAS NOT AVAILABLE BECAUSE THE DEVICE WAS DISCARDED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. Manufacturer Narrative: CORRECTION TO SECTION H6 DEVICE CODES, CODE A2304 WAS ADDED AND H6 EVALUATION CONCLUSION CODES. ADDITIONAL INFORMATION WAS PROVIDED IN SECTION B2 OUTCOMES ATTRIB TO ADV EVENT, SECTION B5 DESCRIBE EVENT OR PROBLEM, AND SECTION H6 IMPACT CODES DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. IT WAS NOT AVAILABLE BECAUSE THE DEVICE WAS DISCARDED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS PROVIDED IN SECTION B2 OUTCOMES ATTRIB TO ADV EVENT, SECTION B5 DESCRIBE EVENT OR PROBLEM, AND SECTION H6 IMPACT CODES. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. IT WAS NOT AVAILABLE BECAUSE THE DEVICE WAS DISCARDED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21968875&pc=NIQ | 2124215-2025-28972 | 2025-04-18 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-06 04:00 | NIQ | PROMUS ELITE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Chest Pain |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH CORONARY ARTERY DISEASE AND UNDERWENT PERCUTANEOUS CORONARY INTERVENTION. THE 90% TARGET LESION WAS LOCATED IN THE MILDLY TORTUOUS AND MILDLY CALCIFIED RIGHT CORONARY ARTERY. AFTER VESSEL PREPARATION, A 38 X 2.50 PROMUS ELITE MR DRUG-ELUTING STENT WAS ADVANCED AND DEPLOYED. FOLLOWING THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO A CARDIAC CARE UNIT. THE NEXT MORNING, THE PATIENT EXPERIENCED CHEST PAIN. CARDIOPULMONARY RESUSCITATION WAS PERFORMED IMMEDIATELY BUT THE PATIENT DIED. THE OFFICIAL CAUSE OF DEATH WAS CARDIAC ARREST. AN AUTOPSY WAS NOT PERFORMED.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH CORONARY ARTERY DISEASE AND UNDERWENT PERCUTANEOUS CORONARY INTERVENTION. THE 90% TARGET LESION WAS LOCATED IN THE MILDLY TORTUOUS AND MILDLY CALCIFIED RIGHT CORONARY ARTERY. AFTER VESSEL PREPARATION, A 38 X 2.50 PROMUS ELITE MR DRUG-ELUTING STENT WAS ADVANCED AND DEPLOYED. FOLLOWING THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO A CARDIAC CARE UNIT. THE NEXT MORNING, THE PATIENT EXPERIENCED CHEST PAIN. CARDIOPULMONARY RESUSCITATION WAS PERFORMED IMMEDIATELY BUT THE PATIENT DIED. THE OFFICIAL CAUSE OF DEATH WAS CARDIAC ARREST. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: H6 PATIENT CODES: CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21968925&pc=OZD | 1220648-2025-28120 | 2025-04-15 04:00 | Death | ABIOMED, INC. | 2025-05-06 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Tachycardia; Ectopic Heartbeat | P140003 |  | 1 | Event Description: A PATIENT DROVE HIMSELF INTO URGENT CARE WITH SHORTNESS OF BREATH. HE WAS TRANSFERRED TO THE HOSPITAL WHERE THEY INTUBATED HIM. THEY DID AN ELECTROCARDIOGRAM AND SAW A ST-SEGMENT ELEVATED MYOCARDIAL INFARCTION. THE PATIENT WAS TREATED WITH PORTNEUF AND WAS BROUGHT TO THE CATHETERIZATION LAB. IT WAS FOUND THAT THE PATIENT HAD TOTALED RIGHT CORONARY ARTERY. THEY FIXED THAT AND PRESSURES DROPPED. THEY ENDED UP CODING HIM AND GOT RETURN OF SPONTANEOUS CIRCULATION. AN IMPELLA CP WAS SUBSEQUENTLY PLACED FOR MECHANICAL CIRCULATORY SUPPORT. UPON POSITIONING THE PUMP, IT WAS NOTED THAT THE PATIENT EXPERIENCED MORE ECTOPY. THE PUMP WAS SUBSEQUENTLY LEFT DEEPER TO ALLEVIATE THE CONDITION. THE FOLLOWING DAY, THE PATIENT WAS NOTED TO BE EXPERIENCING HEMOLYSIS. A UNIT OF BLOOD WAS GIVEN TO HELP STABILIZE THE PATIENT. THE PATIENT'S SYMPTOMS WERE NOTED TO BE CLEARING. FIVE DAYS LATER, THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA AND WAS SHOCKED FOUR TIMES. THE PATIENT EXPIRED LATER THAT DAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER. ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21969332&pc=DSQ | 2916596-2025-02692 | 2025-04-11 04:00 | Death | THORATEC CORPORATION | 2025-05-06 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE MULTI-ORGAN FAILURE (MOF). IT WAS REPORTED THAT THE PATIENT DEVELOPED MOF, BUT NO FURTHER INFORMATION WAS SHARED. THE DEVICE WORKED AS EXPECTED. THE OUTCOME WAS NOT CONSIDERED AS DEVICE THERAPY RELATED. AUTOPSY WAS NOT PERFORMED, AND THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION E1: REPORTER CONTACT INFORMATION WAS NOT AVAILABLE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS DEATH AS AN ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RIGHT HEART FAILURE, RESPIRATORY FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21969334&pc=CAW | 2518422-2025-105200 | 2025-04-23 04:00 | Death | RESPIRONICS, INC. | 2025-05-06 04:00 | CAW | EVERFLO | Adverse Event Without Identified Device or Use Problem | Hypoxia; Respiratory Failure | UNKNOWN |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL DUE TO PNEUMONITIS AND CHRONIC RESPIRATORY FAILURE WITH HYPOXIA ON (B)(6) 2023. AFTER BEING TREATED, THE PATIENT WAS DISCHARGED WITH AN EVERFLO OXYGEN CONCENTRATOR ON (B)(6) 2023. HOWEVER, THE CITY CUT THE POWER TO THE PATIENT'S APARTMENT, AND THE PATIENT ULTIMATELY DIED ON (B)(6) 2023. THE CLAIM AGAINST PHILIPS ALLEGES NEGLIGENT DESIGN, CITING A LACK OF A BACKUP BATTERY. THERE ARE NO CLAIMS OF DEVICE MALFUNCTION. IT IS IMPORTANT TO NOTE, ACCORDING TO THE EVERFLO USER MANUAL, THAT THE INTENDED USE OF THE EVERFLO AND EVERFLO Q OXYGEN CONCENTRATORS IS TO PROVIDE SUPPLEMENTAL OXYGEN TO INDIVIDUALS REQUIRING OXYGEN THERAPY. THESE DEVICES ARE NOT DESIGNED TO BE LIFE-SUPPORTING OR LIFE-SUSTAINING. ADDITIONALLY, THE MANUAL STATES THAT IF A HEALTHCARE PROFESSIONAL DETERMINES THAT AN INTERRUPTION IN THE OXYGEN SUPPLY COULD LEAD TO SERIOUS CONSEQUENCES FOR THE USER, AN ALTERNATE SOURCE OF OXYGEN SHOULD BE READILY AVAILABLE. IT ALSO ADVISES THAT THE POWER QUALITY SHOULD MEET TYPICAL HOME OR HOSPITAL STANDARDS, AND IF CONTINUOUS OPERATION DURING POWER INTERRUPTIONS IS NECESSARY, IT IS RECOMMENDED TO USE AN UNINTERRUPTIBLE POWER SUPPLY OR A BATTERY. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: THE DEVICE UDI, SERIAL NUMBER, AND 510K WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21969568&pc=DQX | 2126666-2025-00022 | 2025-04-10 04:00 | Death | LAKE REGION MEDICAL | 2025-05-06 04:00 | DQX | SAFARI2¿ | Malposition of Device | Stroke/CVA; Cardiac Perforation | K151244 |  | 1 | Event Description: EVENT DESCRIPTION: ON (B)(6) 2025, DURING TAVI CASE AFTER THE FINAL RELEASE OF THE ACURATE PRIME VALVE 27MM, THE VALVE WAS IMPLANTED AT HIGH POSITION (PROBABLY BECAUSE THERE WAS A RAPH BETWEEN RIGHT AND NON-CORONARY CUSP. HOWEVER, THE CT SCAN ANALYSIS WAS SUBOPTIMAL AND THIS IS NOT 100% CONFIRMED. BEFORE CASE, THE PHYSICIANS WERE ALSO AWARE OF THE RAPHE). THE PHYSICIANS ESTIMATED THAT THE HEMODYNAMIC RESULT NEEDED TO BE IMPROVED, SO THEY DECIDED TO DO POST-DILATATION (IGNORING MY WARNINGS OF THE RISK TO DO POST-DILATATION IN A HIGH POSITION) AND THE VALVE MOVED AORTIC. THEN, AFTER CHECKING THAT THE CORONARY FLOW WAS GOOD, THEY DECIDED TO DO VALVE IN VALVE. AFTER THE VALVE IN VALVE, THE HEMODYNAMICS WAS GOOD. ON (B)(6), THE PATIENT HAD A STROKE AND DISSECTION OF AORTA. ON (B)(6), THE PHYSICIANS ANNOUNCED DEATH. PATIENT PRESENT AT TIME OF EVENT? YES. PATIENT COMPLICATIONS: DEATH. IN THE PHYSICIAN'S OPINION, DID THE DEVICE OR PROCEDURE CAUSE OR CONTRIBUTE TO THE PATIENT COMPLICATION? NO. MEDICAL OR SURGICAL INTERVENTIONS: VALVE IN VALVE. PATIENT OUTCOME: DEATH. 23APR2025: THE FOLLOWING INFORMATION RECEIVED VIA A RELATED ACURATE PRIME COMPLAINT INSTIGATED THE NEED FOR A SAFARI2 COMPLAINT: DID THE PHYSICIAN ALLEGE ANY DIFFICULTIES WERE EXPERIENCED WITH THE GUIDEWIRE DURING THE PROCEDURE? DURING THE SECOND VALVE (VIV), THE SAFARI WIRE'S LOOP WAS NOT INTACT. ADDITIONAL INFORMATION: EVENT SUMMARY: IT WAS FURTHER REPORTED THE SEVERELY CALCIFIED NATIVE AORTIC ANNULUS MEASURED 25.2MM AND A SAFARI2 GUIDEWIRE AND 14F ISLEEVE INTRODUCER SHEATH WERE USED DURING THE TAVR PROCEDURE. PRE-DILATATION BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED ON THE FIRST ACURATE PRIME VALVE WITH A 23MM NON-BOSTON SCIENTIFIC BALLOON CATHETER. AFTER IMPLANTATION OF THE FIRST ACURATE PRIME VALVE, MODERATE AORTIC REGURGITATION WAS IDENTIFIED WHICH NECESSITATED POST-DILATATION BAV. WHEN THE FIRST ACURATE PRIME VALVE MOVED AORTICALLY IT LOST CONTACT WITH THE NATIVE AORTIC ANNULUS. DURING THE VALVE-IN-VALVE PROCEDURE THE LOOP OF THE SAFARI2 GUIDEWIRE WAS NOT INTACT. THE AORTIC DISSECTION WAS IDENTIFIED AS THE SUSPECTED SOURCE OF THE STROKE. 24APR2025: PLEASE REQUEST INFORMATION TO HELP OUR TEAM UNDERSTAND WHAT IS MEANT BY "DURING THE SECOND VALVE (VIV), THE SAFARI WIRE'S LOOP WAS NOT INTACT". WAS THERE DAMAGE TO THE FORMED CURVE DISTAL REGION OF THE WIRE AND IF SO, PLEASE SPECIFY THE TYPE OF DAMAGE AND ASK FOR IMAGES SHOWING ANY DAMAGE. FINAL RELEASE OF THE ACURATE PRIME VALVE WAS PERFORMED AND THE VALVE FRAME WAS SHAPED LIKE A CONE, COMPRESSED AT LOWER CROWNS AND IN A HIGH POSITION. THE SAFARI2 GUIDEWIRE WAS INVERTED AT FINAL VALVE RELEASE. AFTER CONFIRMING CORONARIES WERE PATENT A VALVE-IN-VALVE WAS PREPARED. THE SECOND PRIME VALVE WAS PREPARED AND THE SAFARI2 NOT IN AN OPTIMAL POSITION AT RELEASE. 25APR2025: QUESTION: LISTING AND MODEL OF OTHER DEVICES USED DURING THE PROCEDURE, INCLUDE THE MODEL, BRAND NAME, SIZES, ETC. ANSWER: ACURATE PRIME VALVE, ISLEEVE INTRODUCER SHEATH, ACURATE PRIME DELIVERY SYSTEM, 23MM OSYPKA VACS III. QUESTION: DID THE END USER MONITOR THE WIRE POSITION THROUGHOUT THE PROCEDURE FOR PROPER PLACEMENT OF THE CURVE AND DISTAL TIP? ANSWER: YES. QUESTION: WAS THE CURVE OF THE SAFARI2 GUIDEWIRE CONSTRAINED WITHIN A CATHETER DURING INSERTION INTO AND WITHDRAWAL FROM THE BODY OR TREATMENT SITE? ANSWER: NO. QUESTION: DID THE END USER ADVANCE THE SAFARI2 GUIDEWIRE THROUGH THE CATHETER UNDER FLUOROSCOPIC GUIDANCE? ANSWER: UNKNOWN. QUESTION: WAS THE WIRE POSITION MAINTAINED WHILE TRACKING OR ADVANCING A CATHETER OR DEVICE OVER THE WIRE? ANSWER: YES. QUESTION: IS THERE ANY ADDITIONAL IMAGING (E.G., INTRAOPERATIVE FLUOROSCOPY, ECG) THAT WAS PERFORMED DURING THE PROCEDURE THAT MIGHT PROVIDE MORE INSIGHTS INTO THE GUIDEWIRE POSITIONING AND VALVE DEPLOYMENT PROCESS? ANSWER: ALL AVAILABLE MEDIA HAS BEEN SUBMITTED. QUESTION: IS THIS A NEW USER OR EXPERIENCED USER? ANSWER: EXPERIENCED. QUESTION: WAS THE AORTIC DISSECTION IDENTIFIED BEFORE OR AFTER THE STROKE OCCURRED? ANSWER: AFTER, STROKE WAS SUSPECTED AND MRI CONFIRMED DISSECTION. QUESTION: WERE THERE ANY KNOWN ANATOMICAL OR PATHOLOGICAL FINDINGS PRE-PROCEDURE THAT MIGHT HAVE PREDISPOSED THE PATIENT TO DISSECTION? ANSWER: FUNCTIONALLY BICUSPID ANATOMY. QUESTION: WAS THE PRESENCE OF A RAPHE BETWEEN THE RIGHT AND NON-CORONARY CUSP CONFIRMED POST EVENT OR IS IT STILL PRESUMED? ANSWER: CONFIRMED VIA MEDIA. QUESTION: IS THE WIRE AVAILABLE TO BE RETURNED AND IF NOT, PLEASE PROVIDE A REASON FOR NO RETURN? ANSWER: DEVICE WAS DISCARDED AT HEALTHCARE FACILITY. ALTHOUGH THERE IS NO KNOWN RELATIONSHIP BETWEEN THE SAFARI2 DEVICE AND ADVERSE EVENT, PERFORATION AND DEATH ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE SAFARI2 DEVICE AND ARE LISTED IN THE INSTRUCTIONS FOR USE (IFU). THEREFORE, THIS INCIDENT IS BEING FILED AS A DEATH AS A GOOD FAITH MEASURE.Event Description: ON (B)(6) 2025: PATIENT WITH SUSPECTED BICUSPID ANATOMY - RAPHE BETWEEN RCC/NCC PHYSICIANS WERE CONSERVATIVE WITH PRE-DIL DUE TO CALCIUM CHUNK (23MM BALLOON/25.2MMANNULUS) - PRE-DIL SUCCESSFUL. HAD SOME DIFFICULTIES WITH 3 CUSP VIEW AND DURING INITIAL POSITIONING VALVE WAS SITTING HIGH AND GW WAS MOVING FROM INITIAL GOOD POSITION. AFTER KNOB1B PARALLAX IDENTIFIED BUT DECISION TO PROCEED WITH DEPLOYMENT VALVE WAS IN A HIGH POSITION AFTER FINAL VALVE RELEASE - DECISION FOR POST-DIL DUE TO UNDEREXPANSION AND BALLOON INTERACTION LED TO FURTHER VALVE MOVEMENT 2ND PRIME VALVE WAS PREPARED FOR VINV - RELEASE OF 2ND VALVE IN HIGH POSITION AFTER THE VIV THE RESULT WAS BETTER AND ECHO SHOWED NO PARAVALVULAR LEAK, NO REGURGITATION, NO AORTIC GRADIENT. 3HRS POST PROCEDURE A STROKE WAS SUSPECTED AND DISSECTION OD AORTA IDENTIFIED - 1 DAY LATER PT DIED. CAUSE OF STROKE AND DISSECTION UNKNOWN. ON 29MAY2025: MEDIA REVIEW RECEIVED: 3 MENSIO REPORT (MEASUREMENTS ARE CORRECT), PRE-PROCEDURAL CT SCAN AND FLUOROSCOPIC PROCEDURAL VIDEOS. THE INITIAL FINDINGS INDICATE A BICUSPID AORTIC VALVE (SIEVERS 1, POTENTIALLY TRICUSPID) WITH AN ICD-AORTIC ANNULUS EXHIBITING A FLARE-TYPE ANATOMY AND MILD CALCIFICATION (CALCIUM VOLUME OF 688MM³). THE ANGLE IS MEASURED AT 51 DEGREES, AND THE DIAMETER OF THE ASCENDING AORTA IS 37.3MM. ANNULAR DIMENSIONS ARE PROVIDED, WITH DIAMETERS OF 22.3 X 29.0MM (MEAN 25.6MM), A PERIMETER OF 79.6MM, AND AN AREA OF 478MM². THE LVOT DIMENSIONS ARE 20.3 X 28.3MM. FLUOROSCOPIC VIDEOS REVEAL INITIAL AORTOGRAPHY SHOWING A MILDLY CALCIFIED AORTIC VALVE WITH MODERATE TO SEVERE AORTIC REGURGITATION. AN ATTEMPT AT A DOUBLE COPLANAR VIEW WAS SUBOPTIMAL (RAO 6 / CRA 9). BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED WITH A PIGTAIL CATHETER MAINTAINED IN THE NON-CORONARY SINUS. THE SAFARI WIRE WAS POSITIONED INCORRECTLY, CONTRARY TO RECOMMENDATIONS. DEPLOYMENT OF THE PROSTHESIS BEGAN WITH IT POSITIONED HIGHER THAN RECOMMENDED BY THE MANUFACTURER. THE SAFARI WIRE REMAINED INCORRECTLY POSITIONED (FACING DOWN). THE VALVE FRAME APPEARED CONSTRICTED, WITH THE LOWER CROWN HYPOEXPANDED AND MIGRATING AORTIC (SERIES 12, FRAMES 21-22/51). SUBSEQUENTLY, THE ACURATE PRIME PROSTHESIS WAS NOTED TO BE SUPERFICIALLY POSITIONED RELATIVE TO THE ANNULAR PLANE. ANOTHER BAV WAS PERFORMED, INDICATING POSSIBLE AORTIC MOVEMENT OF THE PROSTHESIS (SERIES 15 - FRAMES 88-92/136 AND FRAMES 128-133). THE SEQUENCE SHOWED THE PROSTHESIS EMBOLIZED TO THE SINOTUBULAR JUNCTION, WITHOUT CONTACT WITH THE ANNULAR PLANE. SEVERE AORTIC REGURGITATION WAS NOTED. A SECOND ACURATE PRIME PROSTHESIS WAS INTRODUCED, AND DEPLOYMENT BEGAN IN AN APPROPRIATE POSITION. HOWEVER, THE FINAL VIDEO SHOWING THE LAST STAGES OF RELEASE WAS MISSING. IT WAS NOTED THAT UPON FINAL RELEASE, THE SECOND PROSTHESIS WAS POSITIONED HIGHER THAN RECOMMENDED. THE PROCEDURE CONCLUDED WITH THE SECOND PROSTHESIS'S LOWER CROWNS NEAR THE ANNULAR PLANE, WITH ADEQUATE BILATERAL CORONARY PERFUSION AND NO EVIDENCE OF DISSECTION OR CONTRAST RETENTION. Manufacturer Narrative: PRODUCT WAS DISCARDED AT THE HEALTHCARE FACILITY; THEREFORE, NO VISUAL OR PHYSICAL ANALYSIS OF THE PRODUCT COULD BE PERFORMED. A REVIEW OF THE DEVICE HISTORY RECORDS OF THE SPECIMEN DEVICE DOES NOT PRESENT ANY INDICATION OF MANUFACTURING DEFECT OR ANOMALY THAT COULD HAVE IMPACTED ON THE EVENT AS REPORTED. DURING MANUFACTURING AND PACKAGING OF THIS PRODUCT THE OPERATORS ARE INSTRUCTED TO 100% VISUALLY INSPECT FOR ANY OBVIOUS DEFECT PRIOR TO SHIPMENT. THE HISTORY RECORDS INDICATE THIS PRODUCT WAS FINAL INSPECTION TESTED AT LAKE REGION MEDICAL AND WAS DETERMINED TO BE ACCEPTABLE. AS DESCRIBED IN THE DEVICE INSTRUCTIONS FOR USE (IFU) WARNINGS SECTION: MONITOR THE WIRE POSITION THROUGHOUT THE PROCEDURE FOR PROPER PLACEMENT OF CURVE AND DISTAL TIP. WHEN ADVANCING OR REMOVING THE GUIDEWIRE, ALWAYS USE FLUOROSCOPIC GUIDANCE WITH RADIOGRAPHIC EQUIPMENT THAT PROVIDES HIGH-RESOLUTION IMAGES. NEVER POSITION THE GUIDEWIRE BLINDLY, AS THIS MAY RESULT IN MISPLACEMENT, DISSECTION OR PERFORATION. NEVER ADVANCE THE GUIDEWIRE AGAINST THE RESISTANCE WITHOUT FIRST DETERMINING THE REASON FOR RESISTANCE UNDER FLUOROSCOPY. EXCESSIVE FORCE AGAINST RESISTANCE MAY RESULT IN DAMAGE TO THE CATHETER OR VESSEL/ORGAN. CARE SHOULD BE TAKEN WHEN ADVANCING THE GUIDEWIRE AFTER DEVICE DEPLOYMENT. THE WIRE SHOULD ONLY BE INTRODUCED INTO OR WITHDRAWN FROM THE VENTRICLE THROUGH A CATHETER ALREADY POSITIONED INTO THE VENTRICLE. THE CURVE OF THE SAFARI2TM GUIDEWIRE SHOULD BE CONSTRAINED WITHIN A CATHETER DURING INSERTION INTO OR WITHDRAWAL FROM THE BODY OR TREATMENT SITE. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE: 1. ENSURE A CATHETER IS APPROPRIATELY PLACED IN THE VENTRICLE OR OTHER INTENDED TREATMENT AREA. 2. CAREFULLY REMOVE THE SAFARI2 GUIDEWIRE, FROM THE HOOP BY GRASPING THE PROXIMAL END OF THE J-STRAIGHTENER SEPARATING THE STRAIGHTENER FROM THE HOOP. 3. AFTER INSPECTION OF THE SAFARI2 GUIDEWIRE, ADVANCE THE J-STRAIGHTENER OVER THE DISTAL SECTION OF THE SAFARI2 WIRE TO STRAIGHTEN THE CURVE. 4. INSERT THE SAFARI2 GUIDEWIRE INTO THE HUB OF THE CATHETER WITH THE AID OF THE J- STRAIGHTENER. 5. CAREFULLY ADVANCE THE DISTAL TIP OF THE SAFARI2 GUIDEWIRE INTO THE LUMEN OF THE CATHETER. 6. REMOVE THE SAFARI2 GUIDEWIRE J- STRAIGHTENER BY WITHDRAWING IT OVER THE LENGTH OF THE SAFARI2 GUIDEWIRE. 7. ADVANCED THE SAFARI2 GUIDEWIRE THROUGH THE CATHETER UNDER FLUOROSCOPIC GUIDANCE. 8. CONFIRM SAFARI2 GUIDEWIRE CURVE POSITION TO ASSURE SAFARI2 GUIDEWIRE PLACEMENT AND STABILITY. 9. MAINTAIN WIRE POSITION WHILE TRACKING OR ADVANCING A CATHETER OR DEVICE OVER THE WIRE. 10. TO REMOVE THE GUIDEWIRE FROM THE TREATMENT AREA: A. A CATHETER MUST BE ADVANCED INTO THE TREATMENT AREA OVER THE SAFARI2 GUIDEWIRE PRIOR TO WITHDRAWING THE WIRE. B. THE SAFARI2 GUIDEWIRE IS PULLED BACK COMPLETELY INTO THE CATHETER. C. THE SAFARI2 GUIDEWIRE MAY THEN BE WITHDRAWN FROM THE CATHETER. AT THIS TIME, IT IS NOT POSSIBLE TO ASSIGN A DEFINITIVE ROOT CAUSE FOR THE EVENT AS REPORTED. BASED ON THE INFORMATION PROVIDED, IT APPEARS THAT PROCEDURAL AND/OR CLINICAL FACTORS IMPACTED ON THE EVENT AS REPORTED. SHOULD ADDITIONAL INFORMATION BE RECEIVED, A FOLLOW UP MEDWATCH REPORT WILL BE SUBMITTED. ALTHOUGH ANNEX G INSTRUCTIONS STATE THAT WITH STAND-ALONE DEVICES ANNEX G TERM SHOULD NOT BE USED, CODE 4755 WAS SELECTED DUE TO COMPONENT CODE (G) SHOWING UP ON THE MISSING DATA REPORT. Manufacturer Narrative: THIS MEDWATCH REPORT IS AN UPDATE TO THE PREVIOUS REPORT TO INCLUDE THE ADDITIONAL INFORMATION B(5) AND CODES IN SECTION H6 (C), AND (D).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21970048&pc=DTZ | 3013876692-2025-00037 | 2024-11-16 05:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-05-06 04:00 | DTZ | HLS SET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN SPAIN. MAQUET CARDIOPULMONARY GMBH HAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICE ON 2025-04-01: THE VIGILANCE UNIT OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT, INVOLVING ¿PLS SET¿, UDI-DI (B)(4) MANUFACTURED BY GETINGE WHICH OCCURRED ON 2024-11-16, IN HOSPITAL (B)(6). THE FOLLOWING INFORMATION WAS PROVIDED: INCIDENT DESCRIPTION: DEATH." THE GETINGE SALES AND SERVICE UNIT (SSU) CONFIRMED ON 2025-04-20, THAT THE AFFECTED PRODUCT IN QUESTION WAS AN HLS SET NOT A PLS SET. BASED ON THE INFORMATION AVAILABLE THAT THE PATIENT EXPIRED, A REPORT IS REQUIRED. COMPLAINT ID#: (B)(4). Manufacturer Narrative: MAQUET CARDIOPULMONARY GMBH HAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICE ON 2025-04-01: "THE VIGILANCE UNIT OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT, INVOLVING ¿PLS SET¿, UDI-DI (B)(4) MANUFACTURED BY GETINGE WHICH OCCURRED ON 2024-11-16, IN HOSPITAL (B)(6). THE FOLLOWING INFORMATION WAS PROVIDED: INCIDENT DESCRIPTION: DEATH." BASED ON THE INFORMATION AVAILABLE THAT THE PATIENT EXPIRED, A REPORT IS REQUIRED. HOWEVER, THE SSU CONFIRMED ON 2025-04-20, THAT THE AFFECTED PRODUCT IN QUESTION WAS AN HLS SET NOT A PLS SET. AFTER CONSULTATION WITH THE CUSTOMER, THE SSU CONFIRMED ON 2025-04-24, THAT THE INVOLVED HLS SET WAS FUNCTIONING AS INTENDED AND HAD NO MALFUNCTION. THE PATIENT PASSED AWAY DUE TO HIS DISORDERS. THE AFFECTED HLS SET WAS DISCARDED BY THE CUSTOMER. BASED ON THE AVAILABLE INFORMATION THE REPORTED EVENT "DEATH" COULD BE CONFIRMED BUT WAS NOT RELATED TO A DEVICE MALFUNCTION. THIS COMPLAINT WAS FOUND IN THE DATABASE OF CUSTOMER COMPLAINTS AS A SINGLE EVENT (TIMEFRAME FROM 2024-03-31 TILL 2025-03-31). THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. NOTE: THIS EVENT OCCURRED ON THE SPANISH MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER: 701069073.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21970161&pc=PYX | 1220648-2025-28131 | 2025-04-07 04:00 | Death | ABIOMED, INC | 2025-05-06 04:00 | PYX | IMPELLA RP FLEX | Adverse Event Without Identified Device or Use Problem | Hemolysis; Hemorrhage/Bleeding; Renal Failure | P170011 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT ON BIPELLA SUPPORT EXPERIENCED HEMOLYSIS. BLEEDING AND RENAL FAILURE WERE ALSO NOTED. THE SOURCE OF BLEEDING WAS UNKNOWN. THE PATIENT WAS SCHEDULED FOR PUMP REPOSITIONING AND RENAL THERAPY. THE FAMILY DECIDED TO WITHDRAW CARE, AND THE PATIENT EXPIRED. PER REVIEW BY ABIOMED'S MEDICAL SAFETY OFFICER, THE IMPELLA RP FLEX DEVICE EXPERIENCED POSITIONING ISSUES THAT WERE MOST LIKELY ASSOCIATED WITH THE PATIENT OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS, POSITIONING ISSUES, RENAL FAILURE, AND VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. HEMOLYSIS - REPORTED HEMOLYSIS ON (B)(6) 2025 WITH UNKNOWN DETAILS ON TROUBLESHOOTING OR ANY RESOLUTION OF THE HEMOLYSIS. DATA LOGS SHOW MOTOR CURRENT SPIKES WITH SPEED DEVIATIONS AND STEP UP IN PLACEMENT SIGNAL ON (B)(6) 2025. THE MOTOR CURRENT (MC) REMAINS SLIGHTLY ELEVATED AND HAS FLUCTUATIONS WHILE AT A CONSTANT P-LEVEL. THE LAST MC SPIKE OCCURS ON (B)(6) 2025. THE ROOT CAUSE OF THE HEMOLYSIS WAS MOST LIKELY BIOMATERIAL BASED ON MC SPIKES AND SPEED DEVIATIONS OBSERVED IN THE DATA LOGS PRIOR TO THE CLINICAL REPORT OF HEMOLYSIS. POSITIONING ISSUE - REPORTED PLANNING ON REPOSITIONING ON (B)(6) 2025 BUT UNKNOWN IF REPOSITIONING OCCURRED OR ANY DETAILS ON HOW/WHY. THE ROOT CAUSE OF THE POSITIONING ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL INFORMATION WAS PROVIDED. VASCULAR DAMAGE - REPORTED UNSPECIFIED BLEEDING ON (B)(6) 2025. THE ROOT CAUSE OF THE VASCULAR DAMAGE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS PROVIDED. RENAL FAILURE - REPORTED THE PATIENT WAS ANURIC ON (B)(6) 2025. THE CAUSE OF THE RENAL FAILURE WAS NOT DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. B1 PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28131. B2 SELECTION OF REQUIRED INTERVENTION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28131. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER. ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL. VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21970193&pc=NPT | 2025587-2025-03462 | 2025-04-03 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-06 04:00 | NPT | COREVALVE TRANSCATHETER AORTIC VALVE | Fluid/Blood Leak; Perivalvular Leak; Insufficient Information; Patient Device Interaction Problem | Myocardial Infarction; Heart Failure/Congestive Heart Failure; Myocardial Hypertrophy | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 9 YEARS AND 7 MONTHS FOLLOWING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC AORTIC VALVE THE PATIENT DIED AT HOME. THE DEATH WAS CLASSIFIED AS CARDIOVASCULAR. THE CAUSE OF DEATH WAS LISTED AS CONGESTIVE HEART FAILURE EXACERBATION. MYOCARDIAL INFARCTION AND CORONARY ARTERY DISEASE ALSO WAS REPORTED. THE PHYSICIAN REPORTED THE EVENT WAS UNRELATED TO THE MEDTRONIC PRODUCTS OR IMPLANT PROCEDURE. AN AUTOPSY WAS NOT PERFORMED.Event Description: IT WAS REPORTED THAT APPROXIMATELY 9 YEARS AND 7 MONTHS FOLLOWING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC AORTIC VALVE THE PATIENT DIED AT HOME. THE DEATH WAS CLASSIFIED AS CARDIOVASCULAR. THE CAUSE OF DEATH WAS LISTED AS CONGESTIVE HEART FAILURE EXACERBATION. MYOCARDIAL INFARCTION AND CORONARY ARTERY DISEASE ALSO WAS REPORTED. THE PHYSICIAN REPORTED THE EVENT WAS UNRELATED TO THE MEDTRONIC PRODUCTS OR IMPLANT PROCEDURE. AN AUTOPSY WAS NOT PERFORMED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PHYSICIAN THAT FOLLOWING THE IMPLANT OF THE VALVE, THE FOLLOW-UP ECHOCARDIOGRAMS REVEALED STABLE TRACE TO MILD RESIDUAL REGURGITATION. MYOCARDIAL PERFUSION IMAGING (MPI) WAS PERFORMED APPROXIMATELY THREE AND A HALF YEARS FOLLOWING IMPLANT OF THE VALVE AND SHOWED NONTRANS MURAL INFEROLATERAL SCARE WITH NO DEFINITE REVERSIBLE ISCHEMIA. DYSLIPIDEMIA AND HYPERTENSION WERE WELL CONTROLLED, AND NO MEDICATION CHANGES WERE REQUIRED. APPROXIMATELY TWO MONTHS LATER A FOLLOW-UP VISIT, THE PATIENT DENIED ANY CHEST PAIN, PAROXYSMAL NOCTURNAL DYSPNEA (PND) OR ORTHOPNEA. THE ECHOCARDIOGRAM PERFORMED AT FOLLOW-UP SHOWED A NORMAL LEFT VENTRICULAR SIZE AND SYSTOLIC FUNCTION. THE EJECTION HYPERDYNAMIC WAS AROUND 65-70%. MODERATE SEPTAL HYPERTROPHY AND MILD POSTERIOR HYPE1TROPHY WERE NOTED. NORMAL RIGHT VENTRICULAR SIZE AND SYSTOLIC FUNCTION WERE OBSERVED WITH A MILDLY ENLARGED LEFT ATRIUM. NORMAL SIZED RIGHT ATRIUM WAS IDENTIFIED WITH A PROPERLY FUNCTIONING TRANSCATHETER AORTIC VALVE (TAV) WITH A MEAN GRADIENT OF 12 MILLIMETERS OF MERCURY (MM HG) AND MILD PARAVALVULAR LEAKAGE. THE CAUSE OF DEATH WAS REPORTED AS CHF AND MI. Manufacturer Narrative: UPDATED DATA: B5 B7 D4 - UDI H6 - ANNEX E, ANNEX A, ANNEX G CORRECTION: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT CHANGED THE EVENT DESCRIPTION AND DEVICE RELATEDNESS OF THIS PREVIOUSLY REPORTED EVENT. THERE IS NO EVIDENCE TO SUGGEST THE DEVICE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY. B2 OUTCOME ATTRIBUTED TO ADVERSE EVENT - REMOVED DEATH AND LIFE THREATENING MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21970587&pc=NPT | 2015691-2025-03589 | 2025-04-15 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-06 04:00 | NPT | SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Bradycardia; Cardiac Arrest; Rupture; Pericardial Effusion | P140031 |  | 1 | Event Description: AS REPORTED BY AFFILIATES FROM ITALY, AFTER DEPLOYMENT OF A 29 MM SAPIEN 3 VALVE IN THE AORTIC POSITION BY TRANSFEMORAL APPROACH, THE PATIENT EXPERIENCED BRADYCARDIA UNTIL CARDIAC ARREST OCCURRED. A PERICARDIOCENTESIS WAS PERFORMED AS PERICARDIAL EFFUSION WAS SEEN. THE PATIENT WAS TAKEN TO CARDIAC SURGERY, WHERE EXTRACORPOREAL CIRCULATION (ECC) WAS INITIATED. A LACERATION WAS OBSERVED, LIKELY CAUSED BY A GUIDEWIRE. AN AORTIC VALVE REPLACEMENT WAS PERFORMED, BUT UNFORTUNATELY, THE PATIENT DIED IMMEDIATELY AFTER THE SURGERY.Event Description: AS REPORTED BY AFFILIATES FROM ITALY, AFTER DEPLOYMENT OF A 29MM SAPIEN 3 VALVE IN THE AORTIC POSITION BY TRANSFEMORAL APPROACH, THE PATIENT EXPERIENCED BRADYCARDIA UNTIL CARDIAC ARREST OCCURRED. A PERICARDIOCENTESIS WAS PERFORMED AS PERICARDIAL EFFUSION WAS SEEN. PERICARDIAL EFFUSION WAS CAUSED BY A SINUS RUPTURE. THE PATIENT WAS TAKEN TO CARDIAC SURGERY, WHERE EXTRACORPOREAL CIRCULATION (ECC) WAS INITIATED. A LACERATION WAS OBSERVED BETWEEN THE RIGHT CORONARY CUSP AND NON-CORONARY CUSP, LIKELY CAUSED BY A GUIDEWIRE. AN AORTIC VALVE REPLACEMENT WAS PERFORMED, BUT UNFORTUNATELY, THE PATIENT DIED IMMEDIATELY AFTER THE SURGERY. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR CORRECTION AND ADDITIONAL INFORMATION BASED ON NEW INFORMATION RECEIVED FROM THE SITE. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: B5, ADDITIONAL CODE ADDED TO H6 HEALTH EFFECT - CLINICAL CODE, AND CORRECTED H11. PER THE INSTRUCTIONS FOR USE (IFU), ARRHYTHMIAS AND CONDUCTION SYSTEM DEFECTS WHICH MAY OR MAY NOT REQUIRE A PERMANENT PACEMAKER IMPLANTATION ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH BALLOON VALVULOPLASTY, THE USE OF LOCAL AND/OR GENERAL ANESTHESIA, BIOPROSTHETIC HEART VALVES, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. SICK SINUS SYNDROME (SSS), ALSO CALLED SINUS NODE DYSFUNCTION (SND), IS AN UMBRELLA TERM FOR A GROUP OF ABNORMAL HEART RHYTHMS (ARRHYTHMIAS) PRESUMABLY CAUSED BY A MALFUNCTION OF THE SINUS NODE, THE HEART'S PRIMARY PACEMAKER. BRADYCARDIA-TACHYCARDIA SYNDROME IS A VARIANT OF SICK SINUS SYNDROME IN WHICH SLOW ARRHYTHMIAS AND FAST ARRHYTHMIAS ALTERNATE. IT IS OFTEN ASSOCIATED WITH ISCHEMIC HEART DISEASE AND VALVULAR LESIONS. SICK SINUS SYNDROME IS MORE COMMON IN ELDERLY ADULTS, WHERE THE CAUSE IS OFTEN A NON-SPECIFIC, SCAR-LIKE DEGENERATION OF THE CARDIAC CONDUCTION SYSTEM. CORONARY ARTERY DISEASE, HIGH BLOOD PRESSURE, AND AORTIC AND MITRAL VALVE DISEASES MAY BE ASSOCIATED WITH SICK SINUS SYNDROME, ALTHOUGH THIS ASSOCIATION MAY ONLY BE INCIDENTAL. ACQUIRED SND MAY OCCUR AFTER DAMAGE TO THE SN ARTERY DURING CARDIAC SURGERY OR MAY BE DUE TO OCCLUSION, SUCH AS AFTER MYOCARDIAL INFARCTION. ANOTHER SURGICAL CAUSE OF SND INCLUDES SN TISSUE DAMAGE DURING CANNULATION OF THE SUPERIOR VENA CAVA FOR CARDIOPULMONARY BYPASS OR EXTRACORPOREAL MEMBRANE OXYGENATION. ISCHEMIC CARDIAC ARREST MAY ALSO CAUSE SND. THE NATURAL HISTORY OF SND MAY BE HIGHLY VARIABLE, ALTHOUGH IT TENDS TO BE PROGRESSIVE. THE ONLY EFFECTIVE TREATMENT FOR PATIENTS WITH CHRONIC SYMPTOMATIC SND IS PACEMAKER THERAPY. SLOW HEART RATES OR BRADYCARDIA HAVE MANY POTENTIAL CAUSES INCLUDING DISTURBANCES IN AUTOMATICITY AND CONDUCTION, MEDICATIONS INCLUDING GENERAL ANESTHESIA, ELECTROLYTE IMBALANCES, ADVANCED AGE, AND CARDIAC DISEASES. IF THIS CANNOT BE MANAGED WITH MEDICATION, PERMANENT PACEMAKER IMPLANTATION MAY BE NECESSARY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION THAT A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. WITH THE LIMITED INFORMATION AVAILABLE, THE EXACT CAUSE OF THE BRADYCARDIA IS UNKNOWN BUT MAY BE RELATED TO THE MECHANISMS ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PROCEDURAL FACTORS (LACERATION POSSIBLY CAUSED BY GUIDEWIRE) MAY HAVE CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), ARRHYTHMIAS AND CONDUCTION SYSTEM DEFECTS WHICH MAY OR MAY NOT REQUIRE A PERMANENT PACEMAKER IMPLANTATION ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH BALLOON VALVULOPLASTY, THE USE OF LOCAL AND/OR GENERAL ANESTHESIA, BIOPROSTHETIC HEART VALVES, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. SICK SINUS SYNDROME (SSS), ALSO CALLED SINUS NODE DYSFUNCTION (SND), IS AN UMBRELLA TERM FOR A GROUP OF ABNORMAL HEART RHYTHMS (ARRHYTHMIAS) PRESUMABLY CAUSED BY A MALFUNCTION OF THE SINUS NODE, THE HEART'S PRIMARY PACEMAKER. BRADYCARDIA-TACHYCARDIA SYNDROME IS A VARIANT OF SICK SINUS SYNDROME IN WHICH SLOW ARRHYTHMIAS AND FAST ARRHYTHMIAS ALTERNATE. IT IS OFTEN ASSOCIATED WITH ISCHEMIC HEART DISEASE AND VALVULAR LESIONS. SICK SINUS SYNDROME IS MORE COMMON IN ELDERLY ADULTS, WHERE THE CAUSE IS OFTEN A NON-SPECIFIC, SCAR-LIKE DEGENERATION OF THE CARDIAC CONDUCTION SYSTEM. CORONARY ARTERY DISEASE, HIGH BLOOD PRESSURE, AND AORTIC AND MITRAL VALVE DISEASES MAY BE ASSOCIATED WITH SICK SINUS SYNDROME, ALTHOUGH THIS ASSOCIATION MAY ONLY BE INCIDENTAL. ACQUIRED SND MAY OCCUR AFTER DAMAGE TO THE SN ARTERY DURING CARDIAC SURGERY OR MAY BE DUE TO OCCLUSION, SUCH AS AFTER MYOCARDIAL INFARCTION. ANOTHER SURGICAL CAUSE OF SND INCLUDES SN TISSUE DAMAGE DURING CANNULATION OF THE SUPERIOR VENA CAVA FOR CARDIOPULMONARY BYPASS OR EXTRACORPOREAL MEMBRANE OXYGENATION. ISCHEMIC CARDIAC ARREST MAY ALSO CAUSE SND. THE NATURAL HISTORY OF SND MAY BE HIGHLY VARIABLE, ALTHOUGH IT TENDS TO BE PROGRESSIVE. THE ONLY EFFECTIVE TREATMENT FOR PATIENTS WITH CHRONIC SYMPTOMATIC SND IS PACEMAKER THERAPY. SLOW HEART RATES OR BRADYCARDIA HAVE MANY POTENTIAL CAUSES INCLUDING DISTURBANCES IN AUTOMATICITY AND CONDUCTION, MEDICATIONS INCLUDING GENERAL ANESTHESIA, ELECTROLYTE IMBALANCES, ADVANCED AGE, AND CARDIAC DISEASES. IF THIS CANNOT BE MANAGED WITH MEDICATION, PERMANENT PACEMAKER IMPLANTATION MAY BE NECESSARY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION THAT A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. WITH THE LIMITED INFORMATION AVAILABLE, THE EXACT CAUSE OF THE BRADYCARDIA IS UNKNOWN BUT MAY BE RELATED TO THE MECHANISMS ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT TRANSCATHETER HEART VALVE OVERSIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA, AND/OR PRESENCE OF BULKY CALCIFICATION AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN VALVE (ALL MODELS) RELIES ON VALVE CALCIUM TO SECURELY ANCHOR TO THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO CARDIOVASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. PERICARDIAL EFFUSION IS AN ABNORMAL ACCUMULATION OF FLUID IN THE PERICARDIAL CAVITY. BECAUSE OF THE LIMITED AMOUNT OF SPACE IN THE PERICARDIAL CAVITY, FLUID ACCUMULATION WILL LEAD TO AN INCREASED INTRA PERICARDIAL PRESSURE WHICH CAN NEGATIVELY AFFECT HEART FUNCTION. WHEN THERE IS A PERICARDIAL EFFUSION WITH ENOUGH PRESSURE TO ADVERSELY AFFECT HEART FUNCTION, THIS IS CALLED CARDIAC TAMPONADE. IT CAN BE CAUSED BY A VARIETY OF LOCAL AND SYSTEMIC DISORDERS, OR IT MAY BE IDIOPATHIC. PERICARDIAL EFFUSIONS CAN BE ACUTE OR CHRONIC, AND THE TIME COURSE OF DEVELOPMENT HAS A SIGNIFICANT IMPACT ON THE PATIENT'S SYMPTOMS. IN TVR PATIENTS, IT MAY BE CAUSED BY GUIDEWIRE PERFORATIONS OR ANNULAR RUPTURES. IN TRANSAPICAL PATIENTS, POSTOPERATIVE PERICARDIAL EFFUSIONS ARE COMMON, MOST WILL RESOLVE SPONTANEOUSLY, AND SOME MAY REQUIRE INTERVENTION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION THAT A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PROCEDURAL FACTORS (LACERATION POSSIBLY CAUSED BY GUIDEWIRE) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21970725&pc=PAW | 1649833-2025-00012 |  | Death | ON-X LIFE TECHNOLOGIES, INC. | 2025-05-06 04:00 | PAW | CHORDX UNKNOWN | Insufficient Information | Mitral Valve Insufficiency/ Regurgitation | K174060 |  | 1 | Event Description: ACCORDING TO PUBLICATION "MINIMALLY INVASIVE MITRAL VALVE REPAIR WITH ARTIFICIAL CHORDAE. INSIGHTS FROM A 6 YEAR SINGLE CENTER STUDY¿, THIS PUBLICATION REPORTS ON A RETROSPECTIVE COHORT STUDY THAT WAS CONDUCTED ON CONSECUTIVE PATIENTS WHO UNDERWENT MIMVR USING ARTIFICIAL CHORDAE AT A SINGLE CENTER IN VIETNAM BETWEEN 2016 AND 2022. 90 PATIENTS WERE TREATED, 41 SIMPLE AND 49 INTERMEDIATE TO COMPLEX REPAIRS. NO CASES REQUIRED MR REPLACEMENT AND ALL OUTCOMES WERE FAVORABLE. THE MEDIAN FOLLOW UP TIME WAS 30.3 MONTHS AND THERE WERE TWO CARDIAC DEATHS, ONE IN EACH GROUP. THE KAPLAN¿MEIER SURVIVAL ESTIMATES FOR GROUPS 1 AND 2 AT 12 AND 24 MONTHS WERE 97% VS. 100% AND 97% VS. 96%, RESPECTIVELY, AND ESTIMATES FOR FREEDOM FROM RECURRENT MR WERE 97% VS. 92% AND 97% VS. 88%, RESPECTIVELY. THERE WAS A NEGATIVE ASSOCIATION BETWEEN THE VOLUME OF OPERATIONS AND THE DURATION OF OPERATION AND AORTIC CROSS CLAMP TIME, LEADING TO SHORTER DURATIONS. CONCLUSION. BASED ON OUR SINGLE-CENTER EXPERIENCE, MIMVR USING ARTIFICIAL CHORDAE VIA RIGHT MINI-THORACOTOMY CA. BE SAFELY AND ELECTIVELY PERFORMED IN RESOURCE-LIMITED COUNTRIES FOR PATIENTS WITH MR. THIS APPROACH HAS BEEN SHOWN TO BE APPLICABLE FOR A RANGE OF MR COMPLEXITIES, FROM SIMPLE TO INTERMEDIATE-TO-COMPLEX MV REPAIRS, AND HAS DEMONSTRATED PROMISING RESULTS IN TERMS OF MIDTERM FREEDOM FROM MR RECURRENCE. THIS INVESTIGATION IS RELEGATED TO CHORD-X UNKNOWN CONFIGURATION FOR DEATH AND MITRAL VALVE REGURGITATION.. NO PRODUCT IDENTIFYING INFORMATION PROVIDED. Manufacturer Narrative: ACCORDING TO PUBLICATION "MINIMALLY INVASIVE MITRAL VALVE REPAIR WITH ARTIFICIAL CHORDAE. INSIGHTS FROM A 6 YEAR SINGLE CENTER STUDY¿, THIS PUBLICATION REPORTS ON A RETROSPECTIVE COHORT STUDY THAT WAS CONDUCTED ON CONSECUTIVE PATIENTS WHO UNDERWENT MIMVR USING ARTIFICIAL CHORDAE AT A SINGLE CENTER IN VIETNAM BETWEEN 2016 AND 2022. 90 PATIENTS WERE TREATED, 41 SIMPLE AND 49 INTERMEDIATE TO COMPLEX REPAIRS. NO CASES REQUIRED MR REPLACEMENT AND ALL OUTCOMES WERE FAVORABLE. THE MEDIAN FOLLOW UP TIME WAS 30.3 MONTHS AND THERE WERE TWO CARDIAC DEATHS, ONE IN EACH GROUP. THE KAPLAN¿MEIER SURVIVAL ESTIMATES FOR GROUPS 1 AND 2 AT 12 AND 24 MONTHS WERE 97% VS. 100% AND 97% VS. 96%, RESPECTIVELY, AND ESTIMATES FOR FREEDOM FROM RECURRENT MR WERE 97% VS. 92% AND 97% VS. 88%, RESPECTIVELY. THERE WAS A NEGATIVE ASSOCIATION BETWEEN THE VOLUME OF OPERATIONS AND THE DURATION OF OPERATION AND AORTIC CROSS CLAMP TIME, LEADING TO SHORTER DURATIONS. CONCLUSION. BASED ON OUR SINGLE-CENTER EXPERIENCE, MIMVR USING ARTI5CIAL CHORDAE VIA RIGHT MINI-THORACOTOMY CAN BE SAFELY AND EFFECTIVELY PERFORMED IN RESOURCE-LIMITED COUNTRIES FOR PATIENTS WITH MR. THIS APPROACH HAS BEEN SHOWN TO BE APPLICABLE FOR A RANGE OF MR COMPLEXITIES, FROM SIMPLE TO INTERMEDIATE-TO-COMPLEX MV REPAIRS AND HAS DEMONSTRATED PROMISING RESULTS IN TERMS OF MIDTERM FREEDOM FROM MR RECURRENCE. THIS INVESTIGATION IS RELEGATED TO CHORD-X UNKNOWN CONFIGURATION FOR THE FOLLOWING: DEATH (2). MITRAL VALVE REGURGITATION (33). MULTIPLE ATTEMPTS FOR ADDITIONAL INFORMATION HAVE GONE UNMET. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. A REVIEW OF MANUFACTURING RECORDS COULD NOT BE PERFORMED AS A DEFINITIVE LOT/SERIAL NUMBER AND DATE OF SURGERY WERE NOT PROVIDED BY THE COMPLAINANT. A REVIEW OF THE AVAILABLE INFORMATION WAS PERFORMED. THIS INVESTIGATION REVIEWS CLINICAL OUTCOMES REPORTED IN THE 2023 PUBLICATION TITLED "MINIMALLY INVASIVE MITRAL VALVE REPAIR WITH ARTIFICIAL CHORDAE: INSIGHTS FROM A 6-YEAR SINGLE-CENTER STUDY" BY BUI ET AL. THE STUDY REPORTS A RETROSPECTIVE COHORT ANALYSIS OF 90 PATIENTS UNDERGOING MINIMALLY INVASIVE MITRAL VALVE REPAIR (MIMVR) AT A SINGLE CENTER IN VIETNAM BETWEEN APRIL 2016 AND APRIL 2022. ALL PROCEDURES UTILIZED ARTIFICIAL CHORDAE (CHORD-X LOOPS OR GORE CV-5 SUTURES) THROUGH A RIGHT MINITHORACOTOMY APPROACH. PATIENTS WERE DIVIDED INTO TWO GROUPS BASED ON THE ANYANWU COMPLEXITY SCORE: GROUP 1 INCLUDED SIMPLE REPAIRS (N=41), AND GROUP 2 INCLUDED INTERMEDIATE-TO-COMPLEX REPAIRS (N=49). THE MEAN AGE OF THE COHORT WAS 50.5 ± 12.9 YEARS, AND 77.8% WERE MALE. ALL CASES INVOLVED CARPENTIER TYPE II MITRAL REGURGITATION, PRIMARILY DUE TO FIBROELASTIC DEFICIENCY. BOTH CHORD-X AND GORE CV-5 SUTURES WERE USED FOR NEOCHORDAE IMPLANTATION. CHORD-X WAS USED IN 56% OF CASES, SINGLE GORE CV-5 IN 33%, AND BOTH TECHNIQUES COMBINED IN 11%. ALL PATIENTS ALSO UNDERWENT RING ANNULOPLASTY USING THE CARPENTIER PHYSIO II RING. THE STUDY REPORTS NO CONVERSIONS TO MITRAL VALVE REPLACEMENT AND NO REOPERATIONS RELATED TO THE MITRAL VALVE WERE OBSERVED DURING THE FOLLOW UP PERIOD. POSTOPERATIVE OUTCOMES INCLUDED TWO CARDIAC DEATHS (2.2%) DURING THE MEDIAN 30.3-MONTH FOLLOW-UP PERIOD, WITH ONE PATIENT READMITTED AFTER FOUR MONTHS DUE TO ACUTE HEART FAILURE AND THE OTHER PRESENTING WITH INFECTIVE ENDOCARDITIS TWO MONTHS POST-IMPLANT. ONE PATIENT (1.1%) EXPERIENCED A STROKE WITH SEQUELAE. REOPERATION FOR BLEEDING WAS REQUIRED IN THREE CASES (3.3%). MODERATE OR SEVERE RECURRENT MITRAL REGURGITATION OCCURRED IN 11.1% OF PATIENTS OVERALL, WITH A HIGHER INCIDENCE IN GROUP 2 (18.4%) COMPARED TO GROUP 1 (2.4%). SURVIVAL ESTIMATES AT 48 MONTHS WERE 97% FOR GROUP 1 AND 96% FOR GROUP 2 (P=0.850). KAPLAN-MEIER ESTIMATES FOR FREEDOM FROM RECURRENT MR AT 48 MONTHS WERE 97% FOR GROUP 1 AND 88% FOR GROUP 2 (P=0.260). THE AUTHORS CONCLUDED THAT MIMVR USING ARTIFICIAL CHORDAE (INCLUDING BOTH GORE CV-5 SUTURES AND CHORD-X LOOPS) IS SAFE AND EFFECTIVE ACROSS A RANGE OF MITRAL REGURGITATION COMPLEXITIES, EVEN IN RESOURCE-LIMITED SETTINGS. THIS CLINICAL OUTCOME ANALYSIS IS BASED ON A SINGLE-SURGEON EXPERIENCE USING ARTIFICIAL CHORDAE (GORE CV-5 AND CHORD-X) FOR MINIMALLY INVASIVE MITRAL VALVE REPAIR. THE STUDY DEMONSTRATED FAVORABLE MIDTERM SURVIVAL, LOW RATES OF REOPERATION, AND ACCEPTABLE RECURRENCE RATES OF MITRAL REGURGITATION. WHILE BOTH GORE CV-5 SUTURES AND CHORD-X LOOPS WERE EMPLOYED, THE PUBLICATION DOES NOT DISAGGREGATE OUTCOMES BY DEVICE TYPE. THEREFORE, IT IS NOT POSSIBLE TO ASSOCIATE SPECIFIC COMPLICATIONS OR EVENTS WITH EITHER PRODUCT INDIVIDUALLY. ALL REPORTED COMPLICATIONS, INCLUDING RECURRENT MR, PERIOPERATIVE STROKE, BLEEDING, AND MORTALITY, ARE CONSISTENT WITH THE KNOWN RISKS FOR MITRAL VALVE REPAIR PROCEDURES USING ARTIFICIAL CHORDAE AND ARE IDENTIFIED IN THE IFU. A COMPLAINT AND RECALL QUERY COULD NOT BE PERFORMED AS PRODUCT IDENTIFYING INFORMATION IS UNKNOWN. BASED ON THE AVAILABLE INFORMATION, A DEFINITIVE ROOT CAUSE FOR THIS EVENT CANNOT BE DETERMINED. THIS EVENT DOES NOT IDENTIFY ADDITIONAL HAZARDS OR MODIFY THE PROBABILITY AND SEVERITY OF EXISTING HAZARDS. THERE IS NO INDICATION THAT AN ERROR OR DEFICIENCY OCCURRED AT ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC AND THE IFU ADEQUATELY COMMUNICATES RISK. THIS COMPLAINT WAS REVIEWED FOR A CAPA EVALUATION AND A CAPA IS NOT WARRANTED AT THIS TIME. ARTIVION WILL CONTINUE TO MONITOR SIMILAR COMPLAINTS TO DETERMINE IF ADDITIONAL ACTIONS ARE WARRANTED; HOWEVER, AT THIS TIME NO FURTHER ACTIONS ARE NECESSARY. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC. Manufacturer Narrative: ACCORDING TO PUBLICATION "MINIMALLY INVASIVE MITRAL VALVE REPAIR WITH ARTIFICIAL CHORDAE. INSIGHTS FROM A 6 YEAR SINGLE CENTER STUDY¿, THIS PUBLICATION REPORTS ON A RETROSPECTIVE COHORT STUDY THAT WAS CONDUCTED ON CONSECUTIVE PATIENTS WHO UNDERWENT MIMVR USING ARTIFICIAL CHORDAE AT A SINGLE CENTER IN VIETNAM BETWEEN 2016 AND 2022. 90 PATIENTS WERE TREATED, 41 SIMPLE AND 49 INTERMEDIATE TO COMPLEX REPAIRS. NO CASES REQUIRED MR REPLACEMENT AND ALL OUTCOMES WERE FAVORABLE. THE MEDIAN FOLLOW UP TIME WAS 30.3 MONTHS AND THERE WERE TWO CARDIAC DEATHS, ONE IN EACH GROUP. THE KAPLAN¿MEIER SURVIVAL ESTIMATES FOR GROUPS 1 AND 2 AT 12 AND 24 MONTHS WERE 97% VS. 100% AND 97% VS. 96%, RESPECTIVELY, AND ESTIMATES FOR FREEDOM FROM RECURRENT MR WERE 97% VS. 92% AND 97% VS. 88%, RESPECTIVELY. THERE WAS A NEGATIVE ASSOCIATION BETWEEN THE VOLUME OF OPERATIONS AND THE DURATION OF OPERATION AND AORTIC CROSS CLAMP TIME, LEADING TO SHORTER DURATIONS. CONCLUSION. BASED ON OUR SINGLE-CENTER EXPERIENCE, MIMVR USING ARTI5CIAL CHORDAE VIA RIGHT MINI-THORACOTOMY CAN BE SAFELY AND E2ECTIVELY PERFORMED IN RESOURCE-LIMITED COUNTRIES FOR PATIENTS WITH MR. THIS APPROACH HAS BEEN SHOWN TO BE APPLICABLE FOR A RANGE OF MR COMPLEXITIES, FROM SIMPLE TO INTERMEDIATE-TO-COMPLEX MV REPAIRS, AND HAS DEMONSTRATED PROMISING RESULTS IN TERMS OF MIDTERM FREEDOM FROM MR RECURRENCE. THIS INVESTIGATION IS RELEGATED TO CHORD-X UNKNOWN CONFIGURATION FOR THE FOLLOWING: DEATH (2). MITRAL VALVE REGURGITATION (33). MULTIPLE ATTEMPTS FOR ADDITIONAL INFORMATION HAVE GONE UNMET. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. A REVIEW OF MANUFACTURING RECORDS COULD NOT BE PERFORMED AS A DEFINITIVE LOT/SERIAL NUMBER AND DATE OF SURGERY WERE NOT PROVIDED BY THE COMPLAINANT. A REVIEW OF THE AVAILABLE INFORMATION WAS PERFORMED. THIS INVESTIGATION REVIEWS CLINICAL OUTCOMES REPORTED IN THE 2023 PUBLICATION TITLED "MINIMALLY INVASIVE MITRAL VALVE REPAIR WITH ARTIFICIAL CHORDAE: INSIGHTS FROM A 6-YEAR SINGLE-CENTER STUDY" BY BUI ET AL. THE STUDY REPORTS A RETROSPECTIVE COHORT ANALYSIS OF 90 PATIENTS UNDERGOING MINIMALLY INVASIVE MITRAL VALVE REPAIR (MIMVR) AT A SINGLE CENTER IN VIETNAM BETWEEN APRIL 2016 AND APRIL 2022. ALL PROCEDURES UTILIZED ARTIFICIAL CHORDAE (CHORD-X LOOPS OR GORE CV-5 SUTURES) THROUGH A RIGHT MINITHORACOTOMY APPROACH. PATIENTS WERE DIVIDED INTO TWO GROUPS BASED ON THE ANYANWU COMPLEXITY SCORE: GROUP 1 INCLUDED SIMPLE REPAIRS (N=41), AND GROUP 2 INCLUDED INTERMEDIATE-TO-COMPLEX REPAIRS (N=49). THE MEAN AGE OF THE COHORT WAS 50.5 ± 12.9 YEARS, AND 77.8% WERE MALE. ALL CASES INVOLVED CARPENTIER TYPE II MITRAL REGURGITATION, PRIMARILY DUE TO FIBROELASTIC DEFICIENCY. BOTH CHORD-X AND GORE CV-5 SUTURES WERE USED FOR NEOCHORDAE IMPLANTATION. CHORD-X WAS USED IN 56% OF CASES, SINGLE GORE CV-5 IN 33%, AND BOTH TECHNIQUES COMBINED IN 11%. ALL PATIENTS ALSO UNDERWENT RING ANNULOPLASTY USING THE CARPENTIER PHYSIO II RING. THE STUDY REPORTS NO CONVERSIONS TO MITRAL VALVE REPLACEMENT AND NO REOPERATIONS RELATED TO THE MITRAL VALVE WERE OBSERVED DURING THE FOLLOW UP PERIOD. POSTOPERATIVE OUTCOMES INCLUDED TWO CARDIAC DEATHS (2.2%) DURING THE MEDIAN 30.3-MONTH FOLLOW-UP PERIOD, WITH ONE PATIENT READMITTED AFTER FOUR MONTHS DUE TO ACUTE HEART FAILURE AND THE OTHER PRESENTING WITH INFECTIVE ENDOCARDITIS TWO MONTHS POST-IMPLANT. ONE PATIENT (1.1%) EXPERIENCED A STROKE WITH SEQUELAE. REOPERATION FOR BLEEDING WAS REQUIRED IN THREE CASES (3.3%). MODERATE OR SEVERE RECURRENT MITRAL REGURGITATION OCCURRED IN 11.1% OF PATIENTS OVERALL, WITH A HIGHER INCIDENCE IN GROUP 2 (18.4%) COMPARED TO GROUP 1 (2.4%). SURVIVAL ESTIMATES AT 48 MONTHS WERE 97% FOR GROUP 1 AND 96% FOR GROUP 2 (P=0.850). KAPLAN-MEIER ESTIMATES FOR FREEDOM FROM RECURRENT MR AT 48 MONTHS WERE 97% FOR GROUP 1 AND 88% FOR GROUP 2 (P=0.260). THE AUTHORS CONCLUDED THAT MIMVR USING ARTIFICIAL CHORDAE (INCLUDING BOTH GORE CV-5 SUTURES AND CHORD-X LOOPS) IS SAFE AND EFFECTIVE ACROSS A RANGE OF MITRAL REGURGITATION COMPLEXITIES, EVEN IN RESOURCE-LIMITED SETTINGS. THIS CLINICAL OUTCOME ANALYSIS IS BASED ON A SINGLE-SURGEON EXPERIENCE USING ARTIFICIAL CHORDAE (GORE CV-5 AND CHORD-X) FOR MINIMALLY INVASIVE MITRAL VALVE REPAIR. THE STUDY DEMONSTRATED FAVORABLE MIDTERM SURVIVAL, LOW RATES OF REOPERATION, AND ACCEPTABLE RECURRENCE RATES OF MITRAL REGURGITATION. WHILE BOTH GORE CV-5 SUTURES AND CHORD-X LOOPS WERE EMPLOYED, THE PUBLICATION DOES NOT DISAGGREGATE OUTCOMES BY DEVICE TYPE. THEREFORE, IT IS NOT POSSIBLE TO ASSOCIATE SPECIFIC COMPLICATIONS OR EVENTS WITH EITHER PRODUCT INDIVIDUALLY. ALL REPORTED COMPLICATIONS, INCLUDING RECURRENT MR, PERIOPERATIVE STROKE, BLEEDING, AND MORTALITY, ARE CONSISTENT WITH THE KNOWN RISKS FOR MITRAL VALVE REPAIR PROCEDURES USING ARTIFICIAL CHORDAE AND ARE IDENTIFIED IN THE IFU. A COMPLAINT AND RECALL QUERY COULD NOT BE PERFORMED AS PRODUCT IDENTIFYING INFORMATION IS UNKNOWN. BASED ON THE AVAILABLE INFORMATION, A DEFINITIVE ROOT CAUSE FOR THIS EVENT CANNOT BE DETERMINED. THIS EVENT DOES NOT IDENTIFY ADDITIONAL HAZARDS OR MODIFY THE PROBABILITY AND SEVERITY OF EXISTING HAZARDS. THERE IS NO INDICATION THAT AN ERROR OR DEFICIENCY OCCURRED AT ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC AND THE IFU ADEQUATELY COMMUNICATES RISK. THIS COMPLAINT WAS REVIEWED FOR A CAPA EVALUATION AND A CAPA IS NOT WARRANTED AT THIS TIME. ARTIVION WILL CONTINUE TO MONITOR SIMILAR COMPLAINTS TO DETERMINE IF ADDITIONAL ACTIONS ARE WARRANTED; HOWEVER, AT THIS TIME NO FURTHER ACTIONS ARE NECESSARY. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC. Manufacturer Narrative: THIS INVESTIGATION IS CURRENTLY ONGOING. ANY ADDITIONAL INFORMATION WILL BE PROVIDED IN THE FOLLOW-UP REPORT. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21970822&pc=OZD | 1220648-2025-28137 | 2025-04-16 04:00 | Death | ABIOMED, INC. | 2025-05-06 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Malposition of Device | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED SUBOPTIMAL POSITIONING AND PUMP FLOW. THE PATIENT EXPIRED THE SAME DAY. AN ECHOCARDIOGRAM WAS COMPLETED AND SHOWED EJECTION FRACTION BETWEEN 15-20% WITH MODERATE MITRAL REGURGITATION AND SEVERE TRICUSPID REGURGITATION. LEFT AND RIGHT HEART CATHETERIZATIONS WERE ORDERED AND COMPLETED AND SHOWED MULTIVESSEL DISEASE INVOLVING ALL MAJOR VESSELS INCLUDING THE LEFT MAIN. END-DIASTOLIC PRESSURE WAS GREATER THAN 30. AT COMPLETION OF THE PROCEDURE, THE PATIENT WAS STABLE. ONCE THE PATIENT WAS MOVED TO A BED, THE PATIENT ARRESTED. THE PATIENT WAS, THEREFORE, MOVED BACK TO CATHETERIZATION LAB TABLE AND CODED. CODE INVOLVED CARDIOPULMONARY RESUSCITATION, INTUBATION, EMERGENCY MEDICATIONS, AND DEFIBRILLATION. ONCE RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED, THE PHYSICIAN OPTED TO PLACE AN IMPELLA CP FOR MCS. THE IMPELLA CP DEVICE WAS SUCCESSFULLY COMPLETED, AND THE PATIENT WAS TRANSFERRED TO THE CARDIOVASCULAR INTENSIVE CARE UNIT ON SUPPORT. FOR CONTINUED EVALUATION. THE FOLLOWING DAY THE PATIENT REMAINED INTUBATED AND WAS NOTED TO BE FOLLOWING THE NURSES' COMMANDS. NEPHROLOGY WAS CONSULTED FOR CONTINUOUS RENAL REPLACEMENT (CRRT) THERAPY. THE NEXT MORNING THE PATIENT'S GROIN SITE BEGAN TO BLEED AND THERE WERE WAVEFORM CHANGES. MANUAL PRESSURE WAS APPLIED TO THE GROIN SITE. REGARDING THE WAVEFORM CHANGES, AN ECHOCARDIOGRAM FOR PLACEMENT ASSESSMENT WAS COMPLETED AND CONFIRMED THE IMPELLA CP DEVICE WAS "VERY" SHALLOW. REPOSITIONING AT BEDSIDE WAS COMPLETED. THE DEVICE WAS INITIALLY ADVANCED TO APPROXIMATELY 4.6 CENTIMETERS AND THEN PULLED BACK TO 3.3 CENTIMETERS. IT WAS NOTED THAT THE PIGTAIL OF THE IMPELLA APPEARED TO BE BEING PULLED BY VENTRICLE CORD OR PAPILLARY BED. THE PHYSICIAN WAS UNABLE TO REPOSITION WITHOUT RISK OF THE DEVICE BEING PULLED THROUGH THE AORTIC VALVE. THE DEVICE WAS ONLY ABLE TO ACHIEVE A PERFORMANCE LEVEL OF P-5 WITHOUT SUCTION AND GETTING APPROXIMATELY 2.8 LITERS OF FLOW. THE PATIENT CONTINUED TO DETERIORATE. THERE WERE SEVERAL DISCUSSIONS MADE REGARDING THE INITIATION OF CRRT AND ESCALATION OF IMPELLA THERAPY. IT WAS NOTED THAT THE IMPELLA POSITION AND FLOW WERE TO BE SUBOPTIMAL. HOWEVER DUE TO THE FAMILY'S DECISION, NO MEASURES WERE TAKEN TO RECTIFY. THE FAMILY MADE THE DECISION TO PLACE THE PATIENT ON DO NOT RESUSCITATE STATUS. THE PATIENT WOULD EXPIRE SHORTLY THEREAFTER. Manufacturer Narrative: ACCESS SITE BLEEDING - MINOR: CLINICAL STATES THAT MILD BLEEDING WAS OBSERVED AT THE ACCESS SITE AFTER THE LINEN OF THE PATIENTS BED WAS CHANGED WITH WAS EASILY CONTROLLED BY MANUAL PRESSURE. PUMP WAS NOT RETURNED FOR INVESTIGATION. THEREFORE, BASED OF THE CLINICAL INFORMATION PROVIDED, THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY PATIENT MOVEMENT SINCE THE BLEEDING STARTED RIGHT AFTER THE LINEN CHANGE. POSITIONING ISSUE: CLINICAL STATES THAT AFTER THE BED LINEN CHANGE, THE PATIENT HAD ACCESS SITE OOZE AND THE ECHO FOUND PUMP TO BE ADVANCED FURTHER INTO THE VENTRICLE INITIALLY AND PULLED BACK AGAIN DUE TO PIGTAIL APPEARING TO BE PULLED BY VENTRICLE CORD OR PAP BED. PUMP WAS NOT RETURNED FOR INVESTIGATION. THEREFORE, AS PER THE CLINICAL INFORMATION PROVIDED, THE ROOT CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY PATIENT MOVEMENT RELATED TO THE MOVEMENT OF PATIENT DURING THE LINEN CHANGE. SATURATED FLOW: FM CHANGED FROM LOW PUMP FLOW TO SATURATED FLOW PER REVIEW OF THE DATA LOGS. CLINICAL STATES THAT THE PUMP COULD ONLY ACHIEVE P5 WITHOUT ANY SUCTION ALARMS AFTER THE POSITIONING ISSUE. PUMP WAS NOT RETURNED FOR INVESTIGATION. DATA LOGS CONFIRM SUCTION ALARMS WHEN P-LEVEL IS TURNED ABOVE P5. PLACEMENT SIGNAL IS SEEN TRENDING DOWNWARD. ON (B)(6) 2025, AFTER 12PM THE PUMP WAS AUTO ADJUSTED TO P5 WITH SATURATED FLOWS. SIMILAR SATURATED FLOWS AT P4 AND P3 AS WELL TILL PUMP EXPLANT. NO MOTOR CURRENT SPIKE OR PURGE ISSUES. PATIENT HAD POSITIONING ISSUES EARLIER ON (B)(6) 2025 AND PUMP WAS REPOSITIONED WITH RISK INVOLVED BUT THE OUTCOME OF REPOSITIONING IS UNKNOWN. THEREFORE IT IS NOT CONCLUSIVE THE TRUE REASON FOR SUCTION ALARMS. BASED ON THE CLINICAL INFORMATION PROVIDED AND DATA LOGS REVIEW, THE ROOT CAUSE OF THE SATURATED FLOW WAS NOT DETERMINED SINCE PUMP WAS NOT RETURNED AND CLINICAL DATA AND DATA LOG REVIEW WERE INCONCLUSIVE. Manufacturer Narrative: THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE SUBOPTIMAL POSITION AND FLOW AS AN ASSOCIATED FACTOR TO THE OVERALL OUTCOME. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21970988&pc=NPT | 2015691-2025-03591 | 2025-04-15 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-06 04:00 | NPT | COMMANDER DELIVERY SYSTEM | Improper or Incorrect Procedure or Method | Cardiac Perforation | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES FIELD CLINICAL SPECIALIST, A 23MM SAPIEN 3 ULTRA VALVE WAS TO BE IMPLANTED IN THE AORTIC POSITION VIA TRANSFEMORAL APPROACH USING A COMMANDER DELIVERY SYSTEM AND ESHEATH+. AFTER THE DELIVERY SYSTEM BALLOON ALIGNMENT WHILE ADVANCING THE VALVE AROUND THE ARCH, THE WIRE CAME OUT OF THE VENTRICLE. THE COMMANDER DELIVERY SYSTEM WAS CAREFULLY ADVANCED OVER THE WIRE INTO THE ASCENDING AORTA. THE PHYSICIAN WAS ABLE TO CROSS THE VALVE AND THE DELIVERY SYSTEM SLOWLY ADVANCED OVER THE WIRE UNTIL THE TIP OF THE DELIVERY SYSTEM WAS IN THE VENTRICLE. THE PLAN WAS TO REMOVE THE WIRE AND PLACE A STIFF WIRE IN THE VENTRICLE. ONCE THE WIRE WAS REMOVED, THE DELIVERY SYSTEM WAS ADVANCED WITHOUT THE WIRE AND PUNCTURED THE VENTRICLE. THE PATIENT WAS CONVERTED TO OPEN SURGERY. HOWEVER, DESPITE ALL EFFORTS, THE PATIENT EXPIRED. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR CORRECTION AND ADDITIONAL INFORMATION BASED ON THE ENGINEER EVALUATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: CODES ADDED TO H6 TYPE OF INVESTIGATION, CORRECTED H6 INVESTIGATION FINDINGS, AND CORRECTED H6 INVESTIGATION CONCLUSIONS. THE EVENT REPORTED IS ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PRE-PROCEDURAL IMAGERY WAS PROVIDED, AND THE FOLLOWING WAS OBSERVED: TORTUOSITY, CALCIFICATION AND PRE-EXISTING STENT/GRAFT IS PRESENT IN THE PATIENT ANATOMY. THE INSTRUCTIONS FOR USE (IFU), AND TRAINING MANUALS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THE REPORTED PUNCTURE OF THE VENTRICLE WITH THE DELIVERY SYSTEM RESULTING IN DEATH WAS UNABLE TO BE CONFIRMED AS NO RELEVANT DEVICE OR IMAGERY WAS PROVIDED FOR EVALUATION. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR VENTRICULAR PERFORATION DURING A TVR PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS EVENT. INVESTIGATION RESULTS SUGGEST THAT IT IS POSSIBLE THAT THE INADVERTENT ADVANCEMENT OF THE DELIVERY SYSTEM DURING WIRE EXCHANGE MAY HAVE RESULTED IN THE TIP OF THE DELIVERY SYSTEM CONTACTING THE VENTRICLE RESULTING IN THE PERFORATION REPORTED. AVAILABLE INFORMATION SUGGESTS THAT USE ERROR (ADVANCING DELIVERY SYSTEM WITHOUT GUIDEWIRE) MAY HAVE CONTRIBUTED TO THE REPORTED DAMAGE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21971231&pc=IYN | 3019216-2025-000188 | 2025-04-16 04:00 | Death | PHILIPS ULTRASOUND, INC | 2025-05-06 04:00 | IYN | EPIQ CVXI | Application Program Freezes, Becomes Nonfunctional | Cardiac Arrest | K132304 |  | 1 | Event Description: IT WAS REPORTED THAT THE EPIQ CVXI ULTRASOUND SYSTEM BECAME UNRESPONSIVE AND LOCKED UP DURING AN EXAMINATION WHERE A PATIENT WAS CODING. THE CLINICIAN HAD TO BRING IN ANOTHER SYSTEM IN AN ATTEMPT TO COMPLETE THE EXAMINATION. UNFORTUNATELY, THE PATIENT DID NOT RECOVER FROM THE CODE AND ULTIMATELY EXPIRED; HOWEVER, THE CUSTOMER HAS CLEARLY STATED THE ULTRASOUND DID NOT CAUSE THE PATIENT INCIDENT. PHILIPS HAS STARTED AN INVESTIGATION, AND UPON COMPLETION, A FOLLOW UP REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21971395&pc=DSQ | 2916596-2025-02916 | 2025-03-16 04:00 | Death | ABBOTT MEDICAL-CARDIOMEMS | 2025-05-06 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Neuropathy; Coma; Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED NEUROLOGICAL DYSFUNCTION WHILE IN HOSPITAL. THE EVENT WAS CLASSIFIED AS "NEUROPATHY-LESS THAN 24 HOURS". THE PATIENT PRESENTED WITH A STROKE AND WAS IN A COMA. A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED AND SHOWED AN EMBOLISM ON THE BRAIN STEM. THE PATIENT WAS NOT ON BLOOD THINNER TREATMENT AT THE TIME OF THE EVENT. THE PATIENT THEN PASSED AWAY. IT WAS NOTED THAT THE NEUROLOGICAL DISORDERS CONTRIBUTED TO THE DEATH. THE MAJOR CAUSE OF THE PATIENT DEATH WAS CENTRAL NERVOUS SYSTEM (CNS) DEATH. THE CAUSE OF THE NEUROPATHY WAS COMPLEXITY OF MEDICAL MANAGEMENT. THE DEVICE WAS OPERATING PROPERLY AT THE TIME OF DEATH. THE DEATH WAS DUE TO POSTOPERATIVE COMPLICATIONS RESULTING FROM THE COMPLEXITY OF THE PATIENT'S CONDITION AND WAS NOT CONSIDERED RELATED TO VENTRICULAR ASSIST DEVICE (VAD). Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21971443&pc=DSQ | 3007042319-2025-01104 | 2025-05-01 04:00 | Death | HEARTWARE, INC. | 2025-05-06 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Battery Problem; Adverse Event Without Identified Device or Use Problem; Power Problem; Pumping Problem | Pneumonia; Thrombosis/Thrombus | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT HAD BEEN HOSPITALIZED DUE TO PNEUMONIA, ELEVATED C-REACTIVE PROTEIN (CRP) LEVELS AND A CONFIRMED INCREASE IN LACTIC ACID DEHYDROGENASE (LDH). THERE IS CONCERN FOR SUSPECTED PUMP THROMBOSIS. IT WAS ADVISED TO TURN OFF THE LAVARE CYCLE UNTIL CLINICAL AND TECHNICAL PARAMETERS RETURN TO NORMAL. OF NOTE, WHILE IN HOSPITAL THE CONTROLLER HAD A LOSS OF POWER AND EXHIBITED A CONTROLLER FAULT ALARM WHICH HAD BEEN PREVIOUSLY PERMANENTLY SILENCED AND WAS PERMANENTLY SILENCED AGAIN. THE VAD AND THE CONTROLLER REMAIN IN USE.Event Description: IT WAS FURTHER REPORTED THAT LOG FILES SHOW THAT THE VAD EXHIBITED LOW FLOW ALARMS, BUT THE PATIENT IS CURRENTLY STABLE. HOWEVER, DUE TO THE PRESENCE OF CONFIRMED GERMS, SUSPECTED TO BE UNRELATED TO THE EXISTING PNEUMONIA AND POTENTIALLY TARGETING TISSUE OR THE HEART, AN INVESTIGATION INTO POSSIBLE PATHOGENS HAS BEEN INITIATED. A LABORATORY UPDATE FROM (B)(6) NOTED LDH 541, CRP 95 (REFERENCE 145), LEUKOCYTES 13.2, AND HAPTOGLOBIN 3.14. ADDITIONALLY, A COMPUTERIZED TOMOGRAPHY (CT) SCAN HAS BEEN SCHEDULED TO ASSESS ANY POTENTIAL RISKS TO BOTH THE PATIENT AND THE VAD. THE SITE IS EAGER TO RECEIVE FEEDBACK ON THE HVAD PUMP¿S PERFORMANCE, PARTICULARLY TO DETERMINE WHETHER ANY VAD RELATED ISSUES CAN BE IDENTIFIED OR RULED OUT. AT THIS TIME, IT IS UNCLEAR WHY THE VAD TEAM HAS NOT PROCEEDED WITH A CONTROLLER EXCHANGE, DESPITE A CONTROLLER FAULT BEING REPORTED CONCURRENTLY WITH THE PATIENT¿S HOSPITALIZATION FOR PNEUMONIA. ENCOURAGINGLY, TODAY¿S LAB RESULTS SHOW A DECREASE IN LDH (FROM 700 TO 500) AND DECLINING CRP LEVELS, INDICATING THAT THE ANTIBIOTIC TREATMENT MAY BE TAKING EFFECT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS PLACED IN PALLIATIVE CARE AND EXPIRED. THE PASSING WAS NOT RELATED TO ANY PUMP-RELATED ISSUES. THE VAD TEAM SUPPORTED THE PALLIATIVE CARE TEAM THROUGHOUT, AND EVERYTHING TOOK PLACE IN A PEACEFUL ENVIRONMENT. NO FURTHER DETAILS ABOUT THE CIRCUMSTANCES ARE AVAILABLE. Manufacturer Narrative: CORRECTION: A REVISED SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. REVISED PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) AND THE CONTROLLER (B)(6) WERE NOT RETURNED FOR EVALUATION AS THE VAD REMAINS IN THE PATIENT AND THE CONTROLLER'S LOCATION IS UNKNOWN. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. REVIEW OF THE CONTROLLER LOG FILES REVEALED INTERMITTENT DECREASES IN POWER CONSUMPTION AND ESTIMATED FLOWS THROUGHOUT THE ANALYZED PERIOD, AND ONE (1) LOW FLOW ALARM LOGGED ON (B)(6) 2025 AT 12:15:08. LOG FILE ANALYSIS REVEALED TWO (2) CONTROLLER FAULT ALARMS LOGGED ON (B)(6) 2025 AT 22:10:46 AND ON (B)(6) 2025 AT 11:45:00, INDICATING AN ISSUE WITH THE INTERNAL BATTERY. OF NOTE, THE ALARM WAS SILENCED AT 11:45:01. IN ADDITION, LOG FILES REVEALED THAT THE CONTROLLER HAD BEEN IN USE FOR MORE THAN TWO (2) YEARS. TWO (2) CONTROLLER POWER-UP EVENTS, WITH AN ASSOCIATED MOTOR START EVENTS, WERE LOGGED ON (B)(6) 2025 AT 21:35:05 AND AT 21:37:23, INDICATING A LOSS OF POWER TO THE CONTROLLER. THE DATA POINT RECORDED PRIOR TO THE FIRST LOSS OF POWER REVEALED THAT (B)(6) WAS CONNECTED TO POWER PORT ONE (1) WITH 70% RSOC AND THAT (B)(6) WAS CONNECTED TO POWER PORT TWO (2) WITH 25% RSOC. THE DATA POINT RECORDED AFTER THE FIRST LOSS OF POWER REVEALED THAT NO POWER SOURCE WAS CONNECTED TO POWER PORT ONE (1) AND (B)(6) WAS CONNECTED TO POWER PORT TWO (2). THE DATA POINT RECORDED PRIOR TO THE SECOND LOSS OF POWER REVEALED THAT NO POWER SOURCE WAS CONNECTED TO POWER PORT ONE (1) AND THAT (B)(6) WAS CONNECTED TO POWER PORT TWO (2) WITH 24% RSOC. THE DATA POINT RECORDED AFTER THE SECOND LOSS OF POWER REVEALED THAT NO POWER SOURCE WAS CONNECTED TO POWER PORT ONE (1) AND (B)(6) WAS CONNECTED TO POWER PORT TWO (2). THE CONTROLLER WAS WITHOUT POWER FOR 43 SECONDS AND 17 SECONDS, RESPECTIVELY. NO ANOMALIES WERE RECORDED LEADING UP TO THE LOSSES OF POWER. AS A RESULT, THE REPORTED EVENTS WERE CONFIRMED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED LOW FLOW EVENT. BASED ON THE RISK DOCU- MENTATION, POSSIBLE CAUSES OF THE REPORTED LOW FLOW EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. THE MOST LIKELY ROOT CAUSE OF THE LOSS OF POWER TO THE CONTROLLER CAN BE ATTRIBUTED TO A DISCONNECTION OF BOTH POWER SOURCES AND/OR TO AN INTERMITTENT DISCONNECTION ON ONE OR BOTH POWER SOURCES. CAPA: PR00551638 INVESTIGATED CONTROLLER LOSSES OF POWER. EVEN THOUGH THIS CAPA IS NOW CLOSED, (B)(6) FALLS IN SCOPE OF THIS CAPA. APPLICABLE RISK DOCUMENTATION, EXPERIENCE WITH EVENTS OF SIMILAR CIRCUMSTANCES, AND THE AVAILABLE INFORMATION WERE CONSIDERED; A POSSIBLE ROOT CAUSE OF THE REPORTED CONTROLLER FAULT ALARM EVENT MAY BE ATTRIBUTED, BUT NOT LIMITED, TO A REDUCED CHARGE CAPACITY OF THE INTERNAL BATTERY AND/OR A FAULTY INTERNAL BATTERY CHARGER INTEGRATED CIRCUIT. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT WAS IN PALLIATIVE CARE AND SUBSEQUENTLY EXPIRED. PER THE INSTRUCTIONS FOR USE, DEVICE THROMBUS, RESPIRATORY DYSFUNCTION AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ CONTROLLER D4: SERIAL#: (B)(6), H3: YES, H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) AND THE CONTROLLER (B)(6) WERE NOT RETURNED FOR EVALUATION AS THE VAD REMAINS IN THE PATIENT AND THE CONTROLLER REMAINS IN USE. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. REVIEW OF THE CONTROLLER LOG FILES REVEALED INTERMITTENT DECREASES IN POWER CONSUMPTION AND ESTIMATED FLOWS THROUGHOUT THE ANALYZED PERIOD, AND ONE (1) LOW FLOW ALARM LOGGED ON (B)(6) 2025 AT 12:15:08. LOG FILE ANALYSIS REVEALED TWO (2) CONTROLLER FAULT ALARMS LOGGED ON (B)(6) 2025 AT 22:10:46 AND ON (B)(6) 2025 AT 11:45:00, INDICATING AN ISSUE WITH THE INTERNAL BATTERY. OF NOTE, THE ALARM WAS SILENCED AT 11:45:01. IN ADDITION, LOG FILES REVEALED THAT THE CONTROLLER HAD BEEN IN USE FOR MORE THAN TWO (2) YEARS. TWO (2) CONTROLLER POWER-UP EVENTS, WITH AN ASSOCIATED MOTOR START EVENTS, WERE LOGGED ON (B)(6) 2025 AT 21:35:05 AND AT 21:37:23, INDICATING A LOSS OF POWER TO THE CONTROLLER. THE DATA POINT RECORDED PRIOR TO THE FIRST LOSS OF POWER REVEALED THAT (B)(6) WAS CONNECTED TO POWER PORT ONE (1) WITH 70% RSOC AND THAT (B)(6) WAS CONNECTED TO POWER PORT TWO (2) WITH 25% RSOC. THE DATA POINT RECORDED AFTER THE FIRST LOSS OF POWER REVEALED THAT NO POWER SOURCE WAS CONNECTED TO POWER PORT ONE (1) AND (B)(6) WAS CONNECTED TO POWER PORT TWO (2). THE DATA POINT RECORDED PRIOR TO THE SECOND LOSS OF POWER REVEALED THAT NO POWER SOURCE WAS CONNECTED TO POWER PORT ONE (1) AND THAT (B)(6) WAS CONNECTED TO POWER PORT TWO (2) WITH 24% RSOC. THE DATA POINT RECORDED AFTER THE SECOND LOSS OF POWER REVEALED THAT NO POWER SOURCE WAS CONNECTED TO POWER PORT ONE (1) AND (B)(6) WAS CONNECTED TO POWER PORT TWO (2). THE CONTROLLER WAS WITHOUT POWER FOR 43 SECONDS AND 17 SECONDS, RESPECTIVELY. NO ANOMALIES WERE RECORDED LEADING UP TO THE LOSSES OF POWER. AS A RESULT, THE REPORTED EVENTS WERE CONFIRMED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED LOW FLOW EVENT. BASED ON THE RISK DOCU MENTATION, POSSIBLE CAUSES OF THE REPORTED LOW FLOW EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. THE MOST LIKELY ROOT CAUSE OF THE LOSS OF POWER TO THE CONTROLLER CAN BE ATTRIBUTED TO A DISCONNECTION OF BOTH POWER SOURCES AND/OR TO AN INTERMITTENT DISCONNECTION ON ONE OR BOTH POWER SOURCES. CAPA: PR00551638 INVESTIGATED CONTROLLER LOSSES OF POWER. EVEN THOUGH THIS CAPA IS NOW CLOSED, (B)(6) FALLS IN SCOPE OF THIS CAPA. APPLICABLE RISK DOCUMENTATION, EXPERIENCE WITH EVENTS OF SIMILAR CIRCUMSTANCES, AND THE AVAILABLE INFORMATION WERE CONSIDERED; A POSSIBLE ROOT CAUSE OF THE REPORTED CONTROLLER FAULT ALARM EVENT MAY BE ATTRIBUTED, BUT NOT LIMITED, TO A REDUCED CHARGE CAPACITY OF THE INTERNAL BATTERY AND/OR A FAULTY INTERNAL BATTERY CHARGER INTEGRATED CIRCUIT. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT WAS IN PALLIATIVE CARE AND SUBSEQUENTLY EXPIRED. PER THE INSTRUCTIONS FOR USE, DEVICE THROMBUS, RESPIRATORY DYSFUNCTION AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. D1: HEARTWARE VENTRICULAR ASSIST SYSTEM CONTROLLER, D4: SERIAL#: (B)(6), H3: YES, H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED AS ADDITIONAL INFORMATION HAS BEING RECEIVED FOR THIS EVENT AND SECTIONS HAVE BEEN UPDATED. ADDITIONAL PRODUCT: SERIAL# (B)(6) H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED AS ADDITIONAL INFORMATION HAS BEING RECEIVED FOR THIS EVENT AND SECTIONS HAVE BEEN UPDATED. ADDITIONAL PRODUCT: SERIAL# (B)(6), H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM CONTROLLER 2.0 D4: MODEL #: 1420 / CATALOG #: 1420 / EXPIRATION DATE: 31-MAR-2021 / SERIAL OR LOT#: (B)(6) D9: NO H3: NO H4: MFG DATE: 12-MAR-2019 H5: NO H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21971899&pc=IYN | 3019216-2025-000189 | 2025-04-16 04:00 | Death | PHILIPS ULTRASOUND, INC | 2025-05-06 04:00 | IYN | EPIQ CVXI | Application Program Freezes, Becomes Nonfunctional | Cardiac Arrest | K132304 |  | 1 | Event Description: IT WAS REPORTED THAT THE EPIQ CVXI ULTRASOUND SYSTEM BECAME UNRESPONSIVE AND LOCKED UP DURING AN EXAMINATION WHERE A PATIENT WAS CODING. THE CLINICIAN HAD TO BRING IN ANOTHER SYSTEM IN AN ATTEMPT TO COMPLETE THE EXAMINATION. UNFORTUNATELY, THE PATIENT DID NOT RECOVER FROM THE CODE AND ULTIMATELY EXPIRED; HOWEVER, THE CUSTOMER HAS CLEARLY STATED THE ULTRASOUND DID NOT CAUSE THE PATIENT INCIDENT. PHILIPS HAS STARTED AN INVESTIGATION, AND UPON COMPLETION, A FOLLOW UP REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21972325&pc=NIQ | 2124215-2025-27611 | 2025-04-11 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-06 04:00 | NIQ | SYNERGY? XD | Failure to Deflate | Chest Pain | P150003 |  | 1 | Event Description: IT WAS REPORTED THAT DEFLATION FAILURE OCCURRED, RESULTING IN PATIENT DEATH. THE TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY (LAD). A 3.00 X 48MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT AND PLACED IN THE PROXIMAL LAD. HOWEVER, THE DEVICE WOULD NOT DEFLATE. RE-INFLATION ATTEMPTS WERE MADE FOLLOWED BY DEFLATION ATTEMPTS, AND APPROXIMATELY 30 SECONDS TO ONE MINUTE OF WAIT TIME BEFORE ATTEMPTING TO REMOVE THE DEVICE. HOWEVER, PATIENT EXPERIENCED CHEST PAIN AND ALL EQUIPMENT WAS REMOVED. THE PATIENT WAS STABLE POST PROCEDURE WITH A NON-BOSTON SCIENTIFIC HEART PUMP IMPLANTED. UNFORTUNATELY, PATIENT DIED. IT IS UNKNOWN IF THE DEVICE OR PROCEDURE WERE A CAUSE OR CONTRIBUTION TO THE PATIENT DEATH.Event Description: IT WAS REPORTED THAT DEFLATION FAILURE OCCURRED, RESULTING IN PATIENT DEATH. THE TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY (LAD). A 3.00 X 48MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT AND PLACED IN THE PROXIMAL LAD. HOWEVER, THE DEVICE WOULD NOT DEFLATE. RE-INFLATION ATTEMPTS WERE MADE FOLLOWED BY DEFLATION ATTEMPTS, AND APPROXIMATELY 30 SECONDS TO ONE MINUTE OF WAIT TIME BEFORE ATTEMPTING TO REMOVE THE DEVICE. HOWEVER, PATIENT EXPERIENCED CHEST PAIN AND ALL EQUIPMENT WAS REMOVED. THE PATIENT WAS STABLE POST PROCEDURE WITH A NON-BOSTON SCIENTIFIC HEART PUMP IMPLANTED. UNFORTUNATELY, PATIENT DIED. IT IS UNKNOWN IF THE DEVICE OR PROCEDURE WERE A CAUSE OR CONTRIBUTION TO THE PATIENT DEATH. IT WAS FURTHER REPORTED THAT THE BALLOON DEFLATION WAS APPROXIMATELY 20-30 SECONDS AND IN EMDR TW# (B)(4) THAT THE PATIENT SUBSEQUENTLY DIED OF UNKNOWN CAUSE SOMETIME DURING THE NEXT TWO DAYS. Manufacturer Narrative: B2 DATE OF DEATH: EITHER (B)(6) 2025 AS IT WAS REPORTED THE PATIENT 'DIED OVER THE WEEKEND'.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21973292&pc=NVZ | 2017865-2025-52053 | 2025-04-16 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-06 04:00 | NVZ | ELLIPSE DR | Failure to Deliver Shock/Stimulation; Failure to Convert Rhythm; Incorrect Interpretation of Signal | Arrhythmia | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED AN ARRHYTHMIA. DEVICE INTERROGATION REVEALED MULTIPLE ISSUES WITH THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), INCLUDING FAILURE TO DELIVER A SHOCK, INCORRECT SIGNAL INTERPRETATION, AND FAILURE TO CONVERT THE RHYTHM. THE FAILURE TO DELIVER A SHOCK WAS ASSOCIATED WITH THE RIGHT VENTRICULAR (RV) LEAD. AN ADVERSE EVENT WITHOUT DEVICE MALFUNCTION WAS NOTED ON THE ATRIAL (RA) LEAD. THE PHYSICIAN DECIDED TO EXPLANT THE ICD ALONG WITH THE RV AND RA LEADS. THE PATIENT SUBSEQUENTLY PASSED AWAY.Event Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED AN ARRHYTHMIA. DEVICE INTERROGATION REVEALED MULTIPLE ISSUES WITH THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), INCLUDING FAILURE TO DELIVER A SHOCK, INCORRECT SIGNAL INTERPRETATION, AND FAILURE TO CONVERT THE RHYTHM. THE FAILURE TO DELIVER A SHOCK WAS ASSOCIATED WITH THE RIGHT VENTRICULAR (RV) LEAD. AN ADVERSE EVENT WITHOUT DEVICE MALFUNCTION WAS NOTED ON THE ATRIAL (RA) LEAD. THE PHYSICIAN DECIDED TO EXPLANT THE ICD ALONG WITH THE RV AND RA LEADS. THE PATIENT SUBSEQUENTLY PASSED AWAY, AND THERE IS NO ALLEGATION FROM THE PHYSICIAN OR ANY OTHER HEALTHCARE PROFESSIONAL THAT THE PATIENT'S DEATH WAS CAUSED OR CONTRIBUTED BY ANY OF THE PATIENT'S ABBOTT PRODUCTS OR ANY PROCEDURE INVOLVING THE ABBOTT PRODUCTS. Event Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED AN ARRHYTHMIA. DEVICE INTERROGATION REVEALED MULTIPLE ISSUES WITH THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), INCLUDING FAILURE TO DELIVER A SHOCK, INCORRECT SIGNAL INTERPRETATION, AND FAILURE TO CONVERT THE RHYTHM. THE FAILURE TO DELIVER A SHOCK WAS ASSOCIATED WITH THE RIGHT VENTRICULAR (RV) LEAD. AN ADVERSE EVENT WITHOUT DEVICE MALFUNCTION WAS NOTED ON THE ATRIAL (RA) LEAD. NO INFORMATION IF THE PHYSICIAN DECIDED TO EXPLANT THE ICD ALONG WITH THE RV AND RA LEADS. THE PATIENT SUBSEQUENTLY PASSED AWAY, AND THERE IS NO ALLEGATION FROM THE PHYSICIAN OR ANY OTHER HEALTHCARE PROFESSIONAL THAT THE PATIENT¿S DEATH WAS CAUSED OR CONTRIBUTED BY ANY OF THE PATIENT¿S ABBOTT PRODUCTS OR ANY PROCEDURE INVOLVING THE ABBOTT PRODUCTS Manufacturer Narrative: CORRECTION: B5 SHOULD HAVE NOT STATED THAT THE PRODUCTS WERE EXPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21973294&pc=NVN | 2017865-2025-52057 | 2025-04-16 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-06 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED AN ARRHYTHMIA. DEVICE INTERROGATION REVEALED MULTIPLE ISSUES WITH THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), INCLUDING FAILURE TO DELIVER A SHOCK, INCORRECT SIGNAL INTERPRETATION, AND FAILURE TO CONVERT THE RHYTHM. THE FAILURE TO DELIVER A SHOCK WAS ASSOCIATED WITH THE RIGHT VENTRICULAR (RV) LEAD. AN ADVERSE EVENT WITHOUT DEVICE MALFUNCTION WAS NOTED ON THE ATRIAL (RA) LEAD. THE PHYSICIAN DECIDED TO EXPLANT THE ICD ALONG WITH THE RV AND RA LEADS. THE PATIENT SUBSEQUENTLY PASSED AWAY.Event Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED AN ARRHYTHMIA. DEVICE INTERROGATION REVEALED MULTIPLE ISSUES WITH THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), INCLUDING FAILURE TO DELIVER A SHOCK, INCORRECT SIGNAL INTERPRETATION, AND FAILURE TO CONVERT THE RHYTHM. THE FAILURE TO DELIVER A SHOCK WAS ASSOCIATED WITH THE RIGHT VENTRICULAR (RV) LEAD. AN ADVERSE EVENT WITHOUT DEVICE MALFUNCTION WAS NOTED ON THE ATRIAL (RA) LEAD. THE PHYSICIAN DECIDED TO EXPLANT THE ICD ALONG WITH THE RV AND RA LEADS. THE PATIENT SUBSEQUENTLY PASSED AWAY, AND THERE IS NO ALLEGATION FROM THE PHYSICIAN OR ANY OTHER HEALTHCARE PROFESSIONAL THAT THE PATIENT¿S DEATH WAS CAUSED OR CONTRIBUTED BY ANY OF THE PATIENT¿S ABBOTT PRODUCTS OR ANY PROCEDURE INVOLVING THE ABBOTT PRODUCTS.Event Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED AN ARRHYTHMIA. DEVICE INTERROGATION REVEALED MULTIPLE ISSUES WITH THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), INCLUDING FAILURE TO DELIVER A SHOCK, INCORRECT SIGNAL INTERPRETATION, AND FAILURE TO CONVERT THE RHYTHM. THE FAILURE TO DELIVER A SHOCK WAS ASSOCIATED WITH THE RIGHT VENTRICULAR (RV) LEAD. AN ADVERSE EVENT WITHOUT DEVICE MALFUNCTION WAS NOTED ON THE ATRIAL (RA) LEAD. NO INFORMATION IF THE PHYSICIAN DECIDED TO EXPLANT THE ICD ALONG WITH THE RV AND RA LEADS. THE PATIENT SUBSEQUENTLY PASSED AWAY, AND THERE IS NO ALLEGATION FROM THE PHYSICIAN OR ANY OTHER HEALTHCARE PROFESSIONAL THAT THE PATIENT¿S DEATH WAS CAUSED OR CONTRIBUTED BY ANY OF THE PATIENT¿S ABBOTT PRODUCTS OR ANY PROCEDURE INVOLVING THE ABBOTT PRODUCTS Manufacturer Narrative: CORRECTION: B5 SHOULD HAVE NOT STATED THAT THE PRODUCTS WERE EXPLANTED. Manufacturer Narrative: CORRECTION: B5 STATEMENT WAS NOT UPDATED IN ERROR AND SHOULD HAVE BEEN "THERE IS NO ALLEGATION FROM THE PHYSICIAN OR ANY OTHER HEALTHCARE PROFESSIONAL THAT THE PATIENT¿S DEATH WAS CAUSED OR CONTRIBUTED BY ANY OF THE PATIENT¿S ABBOTT PRODUCTS OR ANY PROCEDURE INVOLVING THE ABBOTT PRODUCTS".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21973297&pc=NVY | 2017865-2025-52054 | 2025-04-16 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-06 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Failure to Deliver Shock/Stimulation | Arrhythmia | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED AN ARRHYTHMIA. DEVICE INTERROGATION REVEALED MULTIPLE ISSUES WITH THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), INCLUDING FAILURE TO DELIVER A SHOCK, INCORRECT SIGNAL INTERPRETATION, AND FAILURE TO CONVERT THE RHYTHM. THE FAILURE TO DELIVER A SHOCK WAS ASSOCIATED WITH THE RIGHT VENTRICULAR (RV) LEAD. AN ADVERSE EVENT WITHOUT DEVICE MALFUNCTION WAS NOTED ON THE ATRIAL (RA) LEAD. THE PHYSICIAN DECIDED TO EXPLANT THE ICD ALONG WITH THE RV AND RA LEADS. THE PATIENT SUBSEQUENTLY PASSED AWAY.Event Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED AN ARRHYTHMIA. DEVICE INTERROGATION REVEALED MULTIPLE ISSUES WITH THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), INCLUDING FAILURE TO DELIVER A SHOCK, INCORRECT SIGNAL INTERPRETATION, AND FAILURE TO CONVERT THE RHYTHM. THE FAILURE TO DELIVER A SHOCK WAS ASSOCIATED WITH THE RIGHT VENTRICULAR (RV) LEAD. AN ADVERSE EVENT WITHOUT DEVICE MALFUNCTION WAS NOTED ON THE ATRIAL (RA) LEAD. THE PHYSICIAN DECIDED TO EXPLANT THE ICD ALONG WITH THE RV AND RA LEADS. THE PATIENT SUBSEQUENTLY PASSED AWAY, AND THERE IS NO ALLEGATION FROM THE PHYSICIAN OR ANY OTHER HEALTHCARE PROFESSIONAL THAT THE PATIENT'S DEATH WAS CAUSED OR CONTRIBUTED BY ANY OF THE PATIENT'S ABBOTT PRODUCTS OR ANY PROCEDURE INVOLVING THE ABBOTT PRODUCTSEvent Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED AN ARRHYTHMIA. DEVICE INTERROGATION REVEALED MULTIPLE ISSUES WITH THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), INCLUDING FAILURE TO DELIVER A SHOCK, INCORRECT SIGNAL INTERPRETATION, AND FAILURE TO CONVERT THE RHYTHM. THE FAILURE TO DELIVER A SHOCK WAS ASSOCIATED WITH THE RIGHT VENTRICULAR (RV) LEAD. AN ADVERSE EVENT WITHOUT DEVICE MALFUNCTION WAS NOTED ON THE ATRIAL (RA) LEAD. NO INFORMATION IF THE PHYSICIAN DECIDED TO EXPLANT THE ICD ALONG WITH THE RV AND RA LEADS. THE PATIENT SUBSEQUENTLY PASSED AWAY, AND THERE IS NO ALLEGATION FROM THE PHYSICIAN OR ANY OTHER HEALTHCARE PROFESSIONAL THAT THE PATIENT¿S DEATH WAS CAUSED OR CONTRIBUTED BY ANY OF THE PATIENT¿S ABBOTT PRODUCTS OR ANY PROCEDURE INVOLVING THE ABBOTT PRODUCTS Manufacturer Narrative: CORRECTION: B5 SHOULD HAVE NOT STATED THAT THE PRODUCTS WERE EXPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21958498&pc=OZP | 2032227-2025-167191 | 2025-03-30 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-05 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hypoglycemia; Brain Injury | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. THE PRIMARY CAUSE OF DEATH WAS HYPOGLYCEMIA, CARDIAC ARREST AND BRAIN INJURY. THE CUSTOMER WAS HOSPITALIZED ON (B)(6) 2025 WITH A BLOOD GLUCOSE VALUE OF 17 MG/DL. THE PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1884, MMT-431AG, MMT-342G, MMT-7040A. TROUBLESHOOTING WAS PERFORMED AND IT WAS NOTICED THAT THE CUSTOMER WAS USING THE SENSOR AND IT WAS WORN DURING THE EVENT. MMT-1884 WAS REQUESTED AND CALLER RESPONSE WAS THE DEVICE WILL BE RETURNED. NO PRODUCT RETURN IS REQUIRED FOR MMT-431AG. NO PRODUCT RETURN IS REQUIRED FOR MMT-342G. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21958817&pc=OZP | 2032227-2025-167398 | 2019-03-21 04:00 | Death | MEDTRONIC MINIMED | 2025-05-05 04:00 | OZP | UNKNOWN NGP | Adverse Event Without Identified Device or Use Problem | Hyperglycemia |  |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY AND THE CUSTOMER ALSO HAD EXPERIENCED HYPERGLYCEMIA. THE CUSTOMER DIED DUE TO THE TUBING FELL OFF AND DID NOT RECEIVE INSULIN. THE EVENT INVOLVED PRODUCT(S) UNKNOWN INSULIN PUMP. TROUBLESHOOTING WAS NOT PERFORMED. IT WAS UNKNOWN IF THE CUSTOMER WAS USING THE INSULIN PUMP WITHIN 48 HOURS AND IF THE AUTO MODE/SMARTGUARD FEATURE WAS ACTIVE AT THE TIME OF HIGH BLOOD GLUCOSE EVENT. THE PATIENT WAS WEARING THE PUMP AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR UNKNOWN INSULIN PUMP. Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21959333&pc=OZD | 1220648-2025-28059 | 2025-04-07 04:00 | Death | ABIOMED, INC. | 2025-05-05 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Thrombocytopenia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 AND WAS POSITIVE FOR HEPARIN INDUCED THROMBOCYTOPENIA. THREE DAYS LATER, SOME SUCTION ALARMS WERE NOTED AND ADDRESSED. CARE WAS WITHDRAWN TWO AFTER THE ALARMS AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW AND THROMBOCYTOPENIA HAS BEEN COMPLETED. THE CAUSE OF THE LOW PUMP FLOW ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED. THE CAUSE OF THE THROMBOCYTOPENIA ISSUE WAS NOT DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21959396&pc=NPT | 2015691-2025-03522 | 2025-04-03 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-05 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140031 |  | 1 | Event Description: PER REPORT RECEIVED FROM GERMANY, IT WAS AN IMPLANT OF A 26 MM SAPIEN 3 ULTRA VALVE IN AORTIC POSITION BY TRANSFEMORAL APPROACH. AFTER THE COMPLETION OF THE PROCEDURE WITHOUT SIGNS OF COMPLICATION, THE PATIENT'S CONDITION WORSENED AND A PERICARDIAL EFFUSION WAS DETECTED BY TTE PRESUMABLY DUE TO A PERICARDIAL PUNCTURE OF THE LEFT VENTRICLE BY THE GUIDEWIRE. A PERICARDIOCENTESIS WAS PERFORMED AND PATIENT CONDITION WAS STABLE. CT PERFORMED FOR ROOT CAUSE DETERMINATION SHOWED NO SIGNS OF ANNULAR RUPTURE OR AORTIC DISSECTION. LATER THAT EVENING, THE PATIENT PASSED AWAY. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR VENTRICULAR PERFORATION DURING A TVR PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PROCEDURAL FACTORS (PERICARDIAL PUNCTURE WITH THE GUIDEWIRE) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21959434&pc=OZD | 1220648-2025-28060 | 2025-04-11 04:00 | Death | ABIOMED, INC. | 2025-05-05 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT FOLLOWING EXPLANT OF AN IMPELLA CP THE PATIENT'S BLOOD PRESSURE SUDDENLY DROPPED AND CARDIOPULMONARY RESUSCITATION WAS PERFORMED FOR OVER AN HOUR. THE SUSPECTED CAUSE FOR THE DROP IN BLOOD PRESSURE WAS A BLEED IN THE ILIAC ARTERY. THE BLEED COULD NOT BE TREATED AND THE PATIENT PASSED AWAY. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21959512&pc=MNS | 2518422-2025-105153 | 2025-04-22 04:00 | Death | RESPIRONICS, INC. | 2025-05-05 04:00 | MNS | DREAMSTATION BIPAP AUTOSV, DS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K090539 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION BIPAP AUTOSV DEVICE HAS PASSED AWAY AND THE WIFE WOULD LIKE TO RETURN IT. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RECEIVED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21959605&pc=OZD | 1220648-2025-28062 | 2025-04-12 04:00 | Death | ABIOMED, INC. | 2025-05-05 04:00 | OZD | IMPELLA CP | Positioning Failure | Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE 62-YEAR-OLD MALE PATIENT ON IMPELLA CP SUPPORT DEVELOPED GI BLEED AND GI WAS CONSULTED. THE PLAN WAS TO SCOPE THE PATIENT FOR SOURCE OF BLEEDING. TWO UNITS OF BLOOD WERE TRANSFUSED. IT WAS FURTHER REPORTED THAT SUCTION EVENTS WERE NOTED OVERNIGHT AND IMPELLA WAS PULLED BACK WITH ECHO GUIDANCE IN THE MORNING. THERE WAS NO-GOOD LONG-TERM OPTIONS FOR THE PATIENT AS SUBSTANCE ABUSE AND NONCOMPLIANCE ISSUES COMPLICATE ADVANCED THERAPY STRATEGY. THE PATIENT¿S FAMILY MADE DECISION TO STOP SUPPORT, AND DONATE ORGANS. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21959780&pc=BZD | 2518422-2025-105154 | 2025-03-17 04:00 | Death | RESPIRONICS, INC. | 2025-05-05 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY AND THE SON WOULD LIKE TO HAVE THE FATHER'S INFORMATION CLEARED AND HIS ENTERED SO HE COULD USE THE DEVICE. THE RESPIRATORY THERAPIST DIRECTED THE SON TO THE DME FOR THOSE CHANGES. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION AS THE SON WOULD LIKE TO USE THE DEVICE. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21960529&pc=OZD | 1220648-2025-28073 | 2025-04-11 04:00 | Death | ABIOMED, INC. | 2025-05-05 04:00 | OZD | IMPELLA CP | Failure to Prime | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE PATIENT PRESENTED WITH COMPLAINTS OF CHEST PAIN APPROXIMATELY 1 HOUR PRIOR TO ARRIVAL THAT WAS DESCRIBED AS HEAVINESS TO HIS CHEST. HE THEN STATED HE HAD PAIN DOWN HIS LEFT ARM AND UP HIS LEFT SIDE OF HIS NECK. PATIENT DENIED ANY PREVIOUS CORONARY ARTERY DISEASE. UPON ARRIVAL TO THE ER, EKG WAS OBTAINED AND REVEALED POSSIBLE ANTEROSEPTAL INFARCT. THERE WAS CONCERN ABOUT THORACIC AORTIC ANEURYSM THAT WAS BEING MONITORED IN WAGONER SO HE WAS SENT FOR A CT WHICH CAME BACK NEGATIVE. PATIENT WAS TRANSFERRED TO THE CATH LAB AS ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION AND PERCUTANEOUS CORONARY INTERVENTION (PCI) WAS PERFORMED. DURING THE PCI, THE PHYSICIAN PERFORMED THROMBECTOMY, BALLOONED AND STENTED THE LEFT ANTERIOR DESCENDING AND THE LEFT CIRCUMFLEX ARTERY. ONCE STENTS WERE IN, ARTERIES KEPT CLOSING OFF AND THERE WAS NO FLOW TO THE ARTERIES. PATIENT BECOME HEMODYNAMICALLY UNSTABLE SO THE DECISION WAS MADE TO PLACE IMPELLA CP. IMPELLA CP ACCESS WAS GAINED AND THE DEVICE WAS THEN ADVANCED OVER THE WIRE. THEY THEN ENTERED THE BODY WITH THE DEVICE BUT IT WAS AFTER THE DEVICE WAS INSERTED THAT STAFF REALIZED THE CP HAD NOT BEEN PRIMED. THEY HAD NOT SPIKED THE PURGE SOLUTION AND THE DECISION WAS MADE TO PULL THE DEVICE FROM THE BODY. STAFF WAS UNDER THE IMPRESSION THAT ONCE THE DEVICE HAS ENTERED THE BODY, IT COULD NOT BE RE-INSERTED SO THEY OPTED TO OPEN A NEW IMPELLA CP. THE PATIENT EXPIRED ON SUPPORT OF THE REPLACEMENT PUMP. Manufacturer Narrative: THE INVESTIGATION OF PRIMING ISSUE HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED. DATA LOG REVIEW SHOWS THAT THE PUMP WAS CONNECTED AND THE CONSOLE PROMPTED CASE START. LOGS ALSO SHOW AN ELEVATED PLACEMENT SIGNAL BEFORE PRIMING IS STARTED, WHICH MATCHES THE CLINICAL DETAILS THAT THEY DID NOT PRIME THE PUMP BEFORE INSERTING INSIDE THE PATIENT. THE ROOT CAUSE OF THE PRIMING ISSUE WAS MOST LIKELY USE-RELATED SINCE CLINICAL DETAILS NOTE THAT THEY DID NOT PRIME OR SPIKE THE PURGE SOLUTION PRIOR TO INSERTING THE PUMP INTO THE PATIENT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21960998&pc=DSQ | 2916596-2025-02715 | 2025-04-09 04:00 | Death | THORATEC CORPORATION | 2025-05-05 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. THE CAUSE OF DEATH WAS HEART FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED, THE DEVICE OPERATED AS EXPECTED, AND WOULD NOT RETURN FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV D IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21961397&pc=OZD | 1220648-2025-28079 | 2025-04-06 04:00 | Death | ABIOMED, INC. | 2025-05-05 04:00 | OZD | IMPELLA CP | Failure to Advance; Patient-Device Incompatibility | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD AN IMPELLA CP PUMP SUPPORT WAS BEING INITIATED FOR A PATIENT ADMITTED IN WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK. THE TEAM WAS UNABLE TO ADVANCE THE PUMP AND MET MUCH RESISTANCE. NATIVE ANATOMY PRECLUDED PLACEMENT OF THE PUMP AND THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA ARREST. CATHETERIZATION LAB TEAM ATTEMPTED RESUSCITATION EFFORTS BUT WERE UNABLE TO OBTAIN RETURN OF SPONTANEOUS CIRCULATION. THE PATIENT EXPIRED AFTERWARDS. ABIOMED'S MEDICAL SAFETY OFFICER REVIEWED THE DEATH AND HAS STATED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. ABIOMED'S MEDICAL SAFETY OFFICER REVIEWED THE DEATH, AND HAS STATED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR TO THE PATIENTS OUTCOME. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: CAUTIONS. ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION SINCE THE IMPELLA MET RESISTANCE AT AORTIC ARCH WHICH MADE IT UNABLE TO ADVANCE. B2 DATE OF DEATH WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28079.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21961407&pc=NIK | 2182208-2025-02187 | 2024-10-11 04:00 | Death | MEDTRONIC, INC. | 2025-05-05 04:00 | NIK | MEDTRONIC CRT-D | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ASSOCIATION EXISTS BETWEEN RIGHT VENTRICULAR (RV) DYSFUNCTION AND ANTIARRHYTHMIC THERAPIES OR ALL-CAUSE MORTALITY IN REMOTELY MONITORED PATIENTS WITH HEART FAILURE WITH CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATORS (CRT-D). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE. THERE WERE PATIENTS WHO EXPERIENCED MYOCARDIAL INFARCTIONS, STROKES, AND INAPPROPRIATE THERAPY WHICH INCLUDED ANTI-TACHYCARDIA PACING (ATP) AND SHOCKS. CAUSES OF INAPPROPRIATE THERAPY INCLUDED DETECTION OF SUPRA VENTRICULAR TACHYCARDIA (SVT), ATRIAL FIBRILLATION (AF), ATRIAL FLUTTER, AND T-WAVE OVERSENSING (TWOS). THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/63 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: RIGHT VENTRICULAR SYSTOLIC PRESSURE PREDICTS OUTCOMES IN PATIENTS WITH CARDIAC RESYNCHRONIZATION THERAPY DEFIBRI LLATORS. POLISH HEART JOURNAL. 2024. 82 (12): 1239-1246. DOI: 10.33963/V.PHJ.102981. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21961438&pc=DXY | 2182208-2025-02188 | 2024-10-11 04:00 | Death | MEDTRONIC, INC. | 2025-05-05 04:00 | DXY | MEDTRONIC LEAD | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ASSOCIATION EXISTS BETWEEN RIGHT VENTRICULAR (RV) DYSFUNCTION AND ANTIARRHYTHMIC THERAPIES OR ALL-CAUSE MORTALITY IN REMOTELY MONITORED PATIENTS WITH HEART FAILURE WITH CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATORS (CRT-D). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE. THERE WERE PATIENTS WHO EXPERIENCED MYOCARDIAL INFARCTIONS, STROKES, AND INAPPROPRIATE THERAPY WHICH INCLUDED ANTI-TACHYCARDIA PACING (ATP) AND SHOCKS. CAUSES OF INAPPROPRIATE THERAPY INCLUDED DETECTION OF SUPRA VENTRICULAR TACHYCARDIA (SVT), ATRIAL FIBRILLATION (AF), ATRIAL FLUTTER, AND T-WAVE OVERSENSING (TWOS). THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/63 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: RIGHT VENTRICULAR SYSTOLIC PRESSURE PREDICTS OUTCOMES IN PATIENTS WITH CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATORS. POLISH HEART JOURNAL. 2024. 82 (12): 1239-1246. DOI: 10.33963/V.PHJ.102981. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21961833&pc=MNS | 2518422-2025-044544 | 2021-09-02 04:00 | Death | RESPIRONICS, INC. | 2025-05-05 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION ST30. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRM FOAM PARTICLES IN THE DEVICE EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21961839&pc=BZD | 2518422-2025-044545 | 2022-09-14 04:00 | Death | RESPIRONICS, INC. | 2025-05-05 04:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRM FOAM PARTICLES IN THE DEVICE EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21961851&pc=DSQ | 2916596-2025-02746 | 2025-04-13 04:00 | Death | THORATEC CORPORATION | 2025-05-05 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY SECONDARY TO SEPSIS AND RIGHT VENTRICULAR (RV) DYSFUNCTION. THE PUMP WAS NOT EXPLANTED AND WOULD NOT RETURN FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENTS; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CUSTOMER. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING SEPSIS, RIGHT HEART FAILURE, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, "PATIENT CARE AND MANAGEMENT" (UNDER "RIGHT HEART FAILURE"), DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. ALTHOUGH ONLY SEPSIS WAS REPORTED BY THE ACCOUNT, THE IFU ALSO LISTS INFECTION AS AN ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21961936&pc=DSQ | 2916596-2025-02824 | 2025-04-13 04:00 | Death | THORATEC CORPORATION | 2025-05-05 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE FUNCTIONED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21962257&pc=DSQ | 3007042319-2025-01098 | 2025-02-23 05:00 | Death | HEARTWARE, INC. | 2025-05-05 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DIED OF CONGESTIVE HEART FAILURE (CHF) WHILE IN HOSPICE CARE. THE PATIENT HAD A VENTRICULAR ASSIST DEVICE (VAD) SYSTEM IMPLANTED. THE VAD WAS DECOMMISSIONED AFTER THE PATIENT'S DEATH. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) HW41588 WAS NOT RETURNED FOR EVALUATION AS IT REMAINED IMPLANTED AND DECOMMISSIONED. THE COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT DIED OF CONGESTIVE HEART FAILURE (CHF) WHILE IN HOSPICE CARE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, WORSENING HEART FAILURE AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 1458Q LEAD IMPLANTED (B)(6)2014, CD3365-40Q ICD IMPLANTED (B)(6) 2015, 7120Q LEAD IMPLANTED (B)(6) 2015, 5019 LEAD ADAPTER IMPLANTED (B)(6) 2016, 6721M-50 ICD IMPLANTED (B)(6) 2016, INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21962365&pc=NPT | 2025587-2025-03403 | 2025-05-01 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-05 04:00 | NPT | EVOLUT FX VALVE | Fluid/Blood Leak; Material Invagination; Device Dislodged or Dislocated; Insufficient Information; Patient Device Interaction Problem; Central Regurgitation | Cardiac Arrest; Foreign Body In Patient; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANT PROCEDURE USING THE EVOLUTFX-26, THE PATIENT WAS UNSTABLE AND REQUIRED ANESTHESIA SUPPORT. A MEDTRONIC GUIDEWIRE WAS INSERTED INTO THE LEFT VENTRICLE WHILE THE VALVE WAS SUCCESSFULLY LOADED INTO THE DELIVERY CATHETER SYSTEM (DCS) AND CONFIRMED A GOOD LOAD VIA FLUOROSCOPIC INSPECTION. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED AND WAS MODERATELY TOLERATED BY THE PATIENT. A SECOND BAV WAS PERFORMED, AND THE VALVE AND DCS WERE INSERTED INTO THE PATIENT; HOWEVER, THE PIGTAIL DISLODGED FROM THE NON-CORONARY SINUS, AND THE PATIENT BEGAN TO DECOMPENSATE WITH LOW OUTPUT. THE VALVE WAS QUICKLY IMPLANTED, IMPROVING CARDIAC OUTPUT, BUT AN INFOLD IN THE FRAME WAS OBSERVED. REGURGITATION WAS DETECTED AT THE INFOLDING UNDER ECHOCARDIOGRAPHY. A POST-IMPLANT BAV WAS PERFORMED USING A 20MM BALLOON AT AN AVERAGE INFLATION PRESSURE OF 35. DURING THE BAV, THE PATIENT WENT INTO CARDIAC ARREST. RESUSCITATION MEASURES WERE UNSUCCESSFUL, AND THE LEFT CORONARY ARTERY WAS NOT VISIBLE. CANNULATION WAS ATTEMPTED. WITH ANOTHER INJECTION, IT WAS NOTED THAT THE VALVE HAD RISEN, MAKING IT MORE DIFFICULT TO CANNULATE. ATTEMPTS TO RAISE THE VALVE WITH AN INFLATED BALLOON TO CLEAR THE CORONARY OSTIA WERE SUCCESSFUL; HOWEVER, THE PATIENT DIED AFTER MULTIPLE RESUSCITATION ATTEMPTS.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANT PROCEDURE USING THE EVOLUTFX-26, THE PATIENT WAS UNSTABLE AND REQUIRED ANESTHESIA SUPPORT. A MEDTRONIC GUIDEWIRE WAS INSERTED INTO THE LEFT VENTRICLE WHILE THE VALVE WAS SUCCESSFULLY LOADED INTO THE DELIVERY CATHETER SYSTEM (DCS) AND CONFIRMED A GOOD LOAD VIA FLUOROSCOPIC INSPECTION. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED AND WAS MODERATELY TOLERATED BY THE PATIENT. A SECOND BAV WAS PERFORMED, AND THE VALVE AND DCS WERE INSERTED INTO THE PATIENT; HOWEVER, THE PIGTAIL DISLODGED FROM THE NON-CORONARY SINUS, AND THE PATIENT BEGAN TO DECOMPENSATE WITH LOW OUTPUT. THE VALVE WAS QUICKLY IMPLANTED, IMPROVING CARDIAC OUTPUT, BUT AN INFOLD IN THE FRAME WAS OBSERVED. REGURGITATION WAS DETECTED AT THE INFOLDING UNDER ECHOCARDIOGRAPHY. A POST-IMPLANT BAV WAS PERFORMED USING A 20MM BALLOON AT AN AVERAGE INFLATION PRESSURE OF 35. DURING THE BAV, THE PATIENT WENT INTO CARDIAC ARREST. RESUSCITATION MEASURES WERE UNSUCCESSFUL, AND THE LEFT CORONARY ARTERY WAS NOT VISIBLE. CANNULATION WAS ATTEMPTED. WITH ANOTHER INJECTION, IT WAS NOTED THAT THE VALVE HAD RISEN, MAKING IT MORE DIFFICULT TO CANNULATE. ATTEMPTS TO RAISE THE VALVE WITH AN INFLATED BALLOON TO CLEAR THE CORONARY OSTIA WERE SUCCESSFUL; HOWEVER, THE PATIENT DIED AFTER MULTIPLE RESUSCITATION ATTEMPTS. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE VALVE WAS INITIALLY IMPLANTED AT A DEPTH OF 3MM; HOWEVER, FOLLOWING DISLODGEMENT, THE DEPTH WAS 0MM. PER THE PHYSICIAN, THE PRESENCE OF SEVERE AORTIC CALCIFICATION CONTRIBUTED TO THE INFOLD NOTED WITHIN THE VALVE FRAME. THE AORTIC REGURGITATION WAS CLARIFIED TO BE SEVERE CENTRAL AORTIC REGURGITATION. THE CAUSE OF DEATH WAS CORONARY OBSTRUCTION AND SEVERE INSUFFICIENCY. PER THE PHYSICIAN, THE MEDTRONIC GUIDEWIRE, VALVE, AND DCS CAN BE EXCLUDED AS CONTRIBUTING FACTORS TO THE PATIENT'S DEATH. Manufacturer Narrative: IMAGE REVIEW: PRE-IMPLANT COMPUTED TOMOGRAPHY (CT) IMAGING WAS PROVIDED WITH SUBOPTIMAL IMAGE QUALITY. THE AORTIC VALVE WAS TRILEAFLET WITH LEAFLETS RIGHT CORONARY CUSP (RCC) AND LEFT CORONARY CUSP (LCC) APPEARING MILDLY CALCIFIED, AND THE NON-CORONARY CUSP (NCC) APPEARING MODERATELY CALCIFIED. THE ANNULUS PERIMETER WAS 72.1 MM WITH A PERIMETER DERIVED DIAMETER OF 22.9 MM SUGGESTING A 26 MM VALVE. THE SOV MINIMUM DIAMETER (RCC) WAS LESS THAN THE RECOMMENDED DIAMETER OF 27 MM, HOWEVER THE MEAN DIAMETER WAS 27.5 MM WHICH ALIGNS WITH THE SIZING MATRIX. THE LCC SINUS HEIGHT MEASURED LESS THAN THE RECOMMENDED HEIGHT OF 15 MM, HOWEVER THE MEAN HEIGHT WAS 15.4 MM WHICH ALIGNS WITH THE SIZING MATRIX. PROXIMAL CALCIFICATION SEEN IN THE RIGHT CORONARY ARTERY (RCA) AND THE TAKE-OFF MEASURED 8.5 MM. CALCIFICATION SEEN IN SINOTUBULAR JUNCTION (STJ). AORTIC ROOT ANGULATION WAS 55°. DIFFERENCE IN SOV DIAMETERS AND SINUS HEIGHTS COMPARED TO ORIGINAL CASE SUMMARY REPORT MAY BE DUE TO CT IMAGE QUALITY AND INDIVIDUAL ADJUSTMENTS MADE TO IMAGE GAINS.THE CASE REPORT PROVIDED INDICATED MEASUREMENTS WERE ROUGH ESTIMATES. THE ANNULUS PERIMETER REPORTED WAS 74.3 MM WITH A PERIMETER DERIVED DIAMETER OF 23.7 MM SUGGESTING A 29 MM VALVE. THE SOV MINIMUM DIAMETER (RCC) WAS LESS THAN THE RECOMMENDED DIAMETER OF 29 MM, HOWEVER THE MEAN DIAMETER WAS 29.1 MM WHICH ALIGNS WITH THE SIZING MATRIX. SINUS HEIGHTS MEASURED OVER THE RECOMMENDED HEIGHT OF =15 MM. CALCIFICATION SEEN IN STJ AND AORTIC ROOT ANGULATION WAS 60°. INTRAPROCEDURAL FLUOROSCOPIC IMAGES WERE PROVIDED FOR REVIEW. FLUOROSCOPIC VALVE LOAD INSPECTION WAS PERFORMED AND CONFIRMED A GOOD LOAD. A PRE-IMPLANT BALLOON DILATION WAS PERFORMED PRIOR TO THE VALVE IMPLANT. THE VALVE WAS DEPLOYED JUST PRIOR TO THE POINT OF NO RECAPTURE IN THE CUSP OVERLAP VIEW AT A DEPTH OF APPROXIMATELY 2-3 MM ON THE NCC AND AN INFOLD WAS EVIDENT. THE INFOLD IS CHARACTERIZED BY AN INWARD FOLD OR CREASE IN THE VALVE, EXTENDING FROM THE INFLOW. INFOLDING COULD OCCUR DUE A MISLOADED VALVE, ANATOMICAL CHARACTERISTICS SUCH AS CALCIUM, AND RECAPTURES. IF AN INFOLD IS IDENTIFIED, THE VALVE SHOULD BE RECAPTURED, REMOVED FROM THE BODY, AND NEW STERILE COMPONENTS MUST BE USED. HOWEVER, THE INFOLDED VALVE WAS RELEASED. CONSEQUENTLY, A POST IMPLANT DILATATION WAS PERFORMED WITHOUT COMPLETE RESOLUTION OF THE INFOLD. CARDIOPULMONARY RESUSCITATION (CPR) EFFORTS WERE INITIATED AND SUBSEQUENT ANGIOGRAM RE VEALED ABSENCE OF FLOW TO THE LEFT CORONARY ARTERY AND DIMINISHED FLOW TO THE RCA. ATTEMPTS TO ACCESS THE LEFT CORONARY ARTERY FAILED AND CPR CONTINUED. THE FOLLOWING ANGIOGRAM REVEALED THAT THE VALVE HAD DISLODGED AORTIC AND THERE WAS COMPLETE ABSENCE OF FLOW TO BOTH CORONARY ARTERIES. A BALLOON WAS USED TO MOVE THE VALVE INTO A HIGHER POSITION IN THE AORTA; HOWEVER, CORONARY FLOW WAS NEVER RESTORED AND IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY DIED. UPDATED: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED H.6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. THE SYSTEM REPORTED DEVICE: BRAND NAME: EVOLUT FX DCS; PRODUCT ID: D-EVOLUTFX-2329 (LOT: 0012483145); PRODUCT TYPE: 0195-HEART VALVES; MANUFACTURED DATE: 2024-10-09; USE BEFORE DATE: 2026-10-09; FULL UDI #: (B)(4); IMPLANT DATE: NON-IMPLANTABLE DEVICE H6. PATIENT CODES AND DEVICE CODES WERE ADDED. EVAL CODE METHOD FOR THE SYSTEM REPORTED DCS WAS CHANGED. ADDITIONAL CODES. THE ANNEX G CODE WAS UPDATED. CORRECTED DATA: B6. LABORATORY DATA - MEDICAL ACRONYM WAS WRITTEN OUT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME EVOLUT FX DCS; PRODUCT ID D-EVOLUTFX-2329 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: L-EVOLUTFX-2329; PRODUCT LOT NUMBER: UNKNOWN; PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID: 20MM BALLOON VALVULOPLASTY CATHETER (NON-MEDTRONIC DEVICE); LOT #: UNKNOWN. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21962427&pc=DSQ | 3007042319-2025-01099 | 2024-01-01 05:00 | Death | HEARTWARE, INC. | 2025-05-05 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CLINICAL AND CARDIAC PARAMETERS ASSOCIATED WITH AORTIC VALVE OPENING IN PATIENTS WITH A LEFT VENTRICULAR ASSIST DEVICE (VAD). PATIENTS WERE CLASSIFIED INTO THOSE WHO SHOWED OPENING OF THE AORTIC VALVE (AV) INTERMITTENTLY, EVERY 2¿3 BEATS, OR IN EVERY CYCLE (AVO) AND THOSE WHOSE AV WAS CONTINUOUSLY CLOSED (NAVO) AFTER SIX MONTHS OF IMPLANTATION. THE AUTHORS DESCRIBED PATIENT DEATHS IN BOTH GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO EXPERIENCED WORSENING (MILD TO MODERATE OR MODERATE TO SEVERE) MITRAL, TRICUSPID, AND AORTIC REGURGITATION AND THREE PATIENTS WHO REQUIRED THE IMPLANTATION OF A RIGHT VAD. THE STATUS OF THE DEVICES IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/57 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: POTENTIAL BENEFITS OF AORTIC VALVE OPENING IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES. ARTIFICIAL ORGANS. 2025; 49:441¿450. DOI: 10.1111/AOR.14891. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21963569&pc=NIQ | 2124215-2025-28713 | 2025-04-11 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-05 04:00 | NIQ | SYNERGY XD | Failure to Deflate | Chest Pain | P150003 |  | 1 | Event Description: IT WAS REPORTED THAT DEFLATION FAILURE OCCURRED AND THE PATIENT LATER DIED. THE PATIENT PRESENTED WITH ST-ELEVATION MYOCARDIAL INFARCTION. THE TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY (LAD). A 3.00 X 48 MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT. HOWEVER, POST DEPLOYMENT, THE BALLOON WOULD NOT DEFLATE AFTER 30-60 SECONDS. RE-INFLATION ATTEMPTS FOLLOWED BY DE-INFLATION ATTEMPTS WERE MADE. THE PATIENT EXPERIENCED CHEST PAIN AND THE DEVICE WAS REMOVED ALONG WITH THE GUIDE CATHETER AND GUIDEWIRE. THE PATIENT WAS STABLE POST PROCEDURE WITH A NON-BOSTON SCIENTIFIC HEART PUMP IMPLANTED. THE PATIENT SUBSEQUENTLY DIED OF UNKNOWN CAUSE SOMETIME DURING THE NEXT TWO DAYS.Event Description: IT WAS REPORTED THAT DEFLATION FAILURE OCCURRED AND THE PATIENT LATER DIED. THE PATIENT PRESENTED WITH ST-ELEVATION MYOCARDIAL INFARCTION. THE TARGET LESION WAS LOCATED IN THE PROXIMAL LEFT ANTERIOR DESCENDING ARTERY (LAD). A 3.00 X 48MM SYNERGY XD DRUG-ELUTING STENT WAS ADVANCED FOR TREATMENT. HOWEVER, POST DEPLOYMENT, THE BALLOON WOULD NOT DEFLATE AFTER 30-60 SECONDS. RE-INFLATION ATTEMPTS FOLLOWED BY DE-INFLATION ATTEMPTS WERE MADE. THE PATIENT EXPERIENCED CHEST PAIN AND THE DEVICE WAS REMOVED ALONG WITH THE GUIDE CATHETER AND GUIDEWIRE. THE PATIENT WAS STABLE POST PROCEDURE WITH A NON-BOSTON SCIENTIFIC HEART PUMP IMPLANTED. THE PATIENT SUBSEQUENTLY DIED OF UNKNOWN CAUSE SOMETIME DURING THE NEXT TWO DAYS. IT WAS FURTHER REPORTED THAT THIS EVENT WAS ALREADY REPORTED IN EMDR 2124215-2025-27611. Manufacturer Narrative: B2 DATE OF DEATH: EITHER (B)(6) OR (B)(6) 2025 AS IT WAS REPORTED THE PATIENT 'DIED OVER THE WEEKEND'. Manufacturer Narrative: B2 DATE OF DEATH: EITHER (B)(6) 2025 AS IT WAS REPORTED THE PATIENT 'DIED OVER THE WEEKEND'.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21963591&pc=BZD | 2518422-2025-105197 | 2025-02-28 05:00 | Death | RESPIRONICS, INC. | 2025-05-05 04:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT OF THE DREAMSTATION AUTO BIPAP DEVICE PASSED AWAY IN 2021 WHILE ON HOSPICE CARE. THE PATIENT WAS NOT USING THE DEVICE TOWARD THE END OF HER LIFE, AND THE DAUGHTER CONTACTED PHILIPS TO RECYCLE THE DEVICE. A SHIPPING LABEL WAS CREATED, AND THE DEVICE WAS RETURNED TO PHILIPS. THE DEVICE WAS EVALUATED, AND NO VISIBLE FOAM DEGRADATION OR PROBLEMS WERE FOUND. THERE IS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21963592&pc=BZD | 2518422-2025-105201 | 2025-04-25 04:00 | Death | RESPIRONICS, INC. | 2025-05-05 04:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT PASSED AWAY ONE MONTH AGO AND THE REP RECRT DREAMSTATION AUTO CPAP DEVICE WAS THROWN AWAY AND THE PATIENT "WOULD NOT HAVE BEEN USING THIS DEVICE". A REPLACEMENT FOR THE DEVICE WAS DECLINED. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FOLLOW UP REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21963604&pc=NPT | 2025587-2025-03415 | 2025-04-28 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-05 04:00 | NPT | EVOLUT FX PLUS VALVE | Gradient Increase; Device Damaged by Another Device; Device Dislodged or Dislocated; Activation Failure; Patient Device Interaction Problem | Foreign Body In Patient; Vascular Dissection; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANTATION OF THE EVOLUT VALVE, MODERATE CALCIUM WAS PRESENT IN THE ANNULUS. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WITH A 20MM NON-MEDTRONIC BALLOON WAS PERFORMED. THE VALVE WAS POSITIONED IN THE ANNULUS USING CUSP OVERLAP TECHNIQUE AT 3MM DEPTH. A GRADIENT OF 14MMHG WAS OBSERVED POST-IMPLANTATION. THE IMPLANTING PHYSICIAN NOTED THAT THE PIGTAIL CATHETER IN THE LEFT VENTRICLE WASN'T DEEP ENOUGH AND THAT MIGHT BE CONTRIBUTING TO THE GRADIENT. THE PIGTAIL WAS POSITIONED JUST BELOW THE SKIRT OF THE VALVE. THE PHYSICIAN ATTEMPTED TO REPOSITION THE PIGTAIL FURTHER INTO THE CHAMBER OF THE LEFT VENTRICLE A COUPLE OF TIMES BUT THE PIGTAIL DID NOT MOVE FORWARD. ONE MORE PUSH RESULTED IN THE VALVE DISLODGING UP INTO THE AORTIC ARCH. A SECOND EVOLUT WAS LOADED TO POSITION IN THE ANNULUS, AND THE NATIVE AORTIC VALVE WAS RECROSSED WITH THE NON-MEDTRONIC GUIDEWIRE. THE SECOND VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE ADVANCED INTO THE DESCENDING AORTA; HOWEVER, THE SYSTEM WAS NOT ABLE TO CROSS AROUND THE ARCH DUE TO THE FIRST DISLODGED VALVE. WITH CONTINUED ATTEMPTS TO ADVANCE THE SYSTEM, THE INITIALLY IMPLANTED VALVE MOVED WITH THE SECOND DCS. THE FIRST VALVE WAS THEN DEEPER THAN THE ANNULUS, AND A SNARE WAS USED TO HOLD THE FIRST VALVE IN PLACE. THE VALVE WAS CAPTURED WITH THE SNARE AND PULLED BACK TO SIT JUST ABOVE THE CUSPS. RESISTANCE WAS NOTED DURING AN ATTEMPT TO PUT A PIGTAIL CATHETER ALONGSIDE THE FIRST VALVE TO VISUALIZE THE POSITION OF THE FIRST VALVE, WHICH APPEARED HIGHER THAN ANTICIPATED. IT WAS NOTED THAT WHEN THE VALVE WAS MOVED, A DISSECTION IN THE NON-CORONARY CUSP OCCURRED. THE SECOND VALVE WAS ADVANCED INSIDE THE FIRST VALVE AND DEPLOYED. THE OUTFLOW OF THE SECOND VALVE WAS COMPROMISED BY THE WAIST OF THE SECOND VALVE. A 22MM NON-MEDTRONIC BALLOON WAS USED TO POST DILATE, BUT THERE WAS NO CHANGE TO THE APPEARANCE OF THE SECOND VALVE. AN AORTOGRAM WAS PERFORMED AND CONFIRMED THAT A DISSECTION HAD OCCURRED ASSOCIATED WITH THE ROOT OF THE AORTA. THE PATIENT'S HEMODYNAMICS WERE TRYING TO COMPENSATE. AN ECHOCARDIOGRAM WAS PERFORMED AND A PERICARDIAL EFFUSION WAS OBSERVED. THE TEAM CONTINUED TO SUPPORT THE PATIENT'S RECOVERY FOR 50 MINUTES INCLUDING ATTEMPTING CARDIOPULMONARY RESUSCITATION, OBSERVING VARIOUS GASES, ECHOCARDIOGRAPHY, AND HEMODYNAMICS; HOWEVER, WERE UNSUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DECEASED.Event Description: IT WAS REPORTED THAT DURING THE IMPLANTATION OF THE EVOLUT VALVE, MODERATE CALCIUM WAS PRESENT IN THE ANNULUS. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WITH A 20MM NON-MEDTRONIC (OSYPKA) BALLOON WAS PERFORMED. THE VALVE WAS POSITIONED IN THE ANNULUS USING CUSP OVERLAP TECHNIQUE AT 3MM DEPTH. A GRADIENT OF 14MMHG WAS OBSERVED POST-IMPLANTATION. THE IMPLANTING PHYSICIAN NOTED THAT THE PIGTAIL CATHETER IN THE LEFT VENTRICLE WASN'T DEEP ENOUGH AND THAT MIGHT BE CONTRIBUTING TO THE GRADIENT. THE PIGTAIL WAS POSITIONED JUST BELOW THE SKIRT OF THE VALVE. THE PHYSICIAN ATTEMPTED TO REPOSITION THE PIGTAIL FURTHER INTO THE CHAMBER OF THE LEFT VENTRICLE A COUPLE OF TIMES BUT THE PIGTAIL DID NOT MOVE FORWARD. ONE MORE PUSH RESULTED IN THE VALVE DISLODGING UP INTO THE AORTIC ARCH. A SECOND EVOLUT WAS LOADED TO POSITION IN THE ANNULUS, AND THE NATIVE AORTIC VALVE WAS RECROSSED WITH THE NON-MEDTRONIC GUIDEWIRE. THE SECOND VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE ADVANCED INTO THE DESCENDING AORTA; HOWEVER, THE SYSTEM WAS NOT ABLE TO CROSS AROUND THE ARCH DUE TO THE FIRST DISLODGED VALVE. WITH CONTINUED ATTEMPTS TO ADVANCE THE SYSTEM, THE INITIALLY IMPLANTED VALVE MOVED WITH THE SECOND DCS. THE FIRST VALVE WAS THEN DEEPER THAN THE ANNULUS, AND A SNARE WAS USED TO HOLD THE FIRST VALVE IN PLACE. THE VALVE WAS CAPTURED WITH THE SNARE AND PULLED BACK TO SIT JUST ABOVE THE CUSPS. RESISTANCE WAS NOTED DURING AN ATTEMPT TO PUT A PIGTAIL CATHETER ALONGSIDE THE FIRST VALVE TO VISUALIZE THE POSITION OF THE FIRST VALVE, WHICH APPEARED HIGHER THAN ANTICIPATED. IT WAS NOTED THAT WHEN THE VALVE WAS MOVED, A DISSECTION IN THE NON-CORONARY CUSP OCCURRED. THE SECOND VALVE WAS ADVANCED INSIDE THE FIRST VALVE AND DEPLOYED. THE OUTFLOW OF THE SECOND VALVE WAS COMPROMISED BY THE WAIST OF THE SECOND VALVE. A 22MM NON-MEDTRONIC (TRUE) BALLOON WAS USED TO POST DILATE, BUT THERE WAS NO CHANGE TO THE APPEARANCE OF THE SECOND VALVE. AN AORTOGRAM WAS PERFORMED AND CONFIRMED THAT A DISSECTION HAD OCCURRED ASSOCIATED WITH THE ROOT OF THE AORTA. THE PATIENT'S HEMODYNAMICS WERE TRYING TO COMPENSATE. AN ECHOCARDIOGRAM WAS PERFORMED AND A PERICARDIAL EFFUSION WAS OBSERVED. THE TEAM CONTINUED TO SUPPORT THE PATIENT'S RECOVERY FOR 50 MINUTES INCLUDING ATTEMPTING CARDIOPULMONARY RESUSCITATION, OBSERVING VARIOUS GASES, ECHOCARDIOGRAPHY, AND HEMODYNAMICS; HOWEVER, WERE UNSUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DECEASED. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE FIRST VALVE COMPROMISED THE OUTFLOW PORTION OF THE SECOND VALVE FRAME AND PREVENTED FULL EXPANSION. THE PATIENT'S CAUSE OF DEATH WAS AORTIC DISSECTION AND RIGHT CORONARY OCCLUSION. PER THE PHYSICIAN, NEITHER THE VALVE NOR THE DCS CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: D-EVOLUTFX-2329 (0012710774); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID: EVFXPLUS-26 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: ON (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. H6. H10. ADDITIONAL CODES. CORRECTED DATA: D. SUSPECT MEDICAL DEVICE FULL UDI #. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21963691&pc=NPT | 2025587-2025-03416 | 2025-04-28 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-05 04:00 | NPT | EVOLUT FX DCS | Device Damaged by Another Device; Difficult to Advance; Activation Failure; Patient Device Interaction Problem | Vascular Dissection; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANTATION OF THE EVOLUT VALVE, MODERATE CALCIUM WAS PRESENT IN THE ANNULUS. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WITH A 20MM NON-MEDTRONIC BALLOON WAS PERFORMED. THE VALVE WAS POSITIONED IN THE ANNULUS USING CUSP OVERLAP TECHNIQUE AT 3MM DEPTH. A GRADIENT OF 14MMHG WAS OBSERVED POST-IMPLANTATION. THE IMPLANTING PHYSICIAN NOTED THAT THE PIGTAIL CATHETER IN THE LEFT VENTRICLE WASN'T DEEP ENOUGH AND THAT MIGHT BE CONTRIBUTING TO THE GRADIENT. THE PIGTAIL WAS POSITIONED JUST BELOW THE SKIRT OF THE VALVE. THE PHYSICIAN ATTEMPTED TO REPOSITION THE PIGTAIL FURTHER INTO THE CHAMBER OF THE LEFT VENTRICLE A COUPLE OF TIMES BUT THE PIGTAIL DID NOT MOVE FORWARD. ONE MORE PUSH RESULTED IN THE VALVE DISLODGING UP INTO THE AORTIC ARCH. A SECOND EVOLUT WAS LOADED TO POSITION IN THE ANNULUS, AND THE NATIVE AORTIC VALVE WAS RECROSSED WITH THE NON-MEDTRONIC GUIDEWIRE. THE SECOND VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE ADVANCED INTO THE DESCENDING AORTA; HOWEVER, THE SYSTEM WAS NOT ABLE TO CROSS AROUND THE ARCH DUE TO THE FIRST DISLODGED VALVE. WITH CONTINUED ATTEMPTS TO ADVANCE THE SYSTEM, THE INITIALLY IMPLANTED VALVE MOVED WITH THE SECOND DCS. THE FIRST VALVE WAS THEN DEEPER THAN THE ANNULUS, AND A SNARE WAS USED TO HOLD THE FIRST VALVE IN PLACE. THE VALVE WAS CAPTURED WITH THE SNARE AND PULLED BACK TO SIT JUST ABOVE THE CUSPS. RESISTANCE WAS NOTED DURING AN ATTEMPT TO PUT A PIGTAIL CATHETER ALONGSIDE THE FIRST VALVE TO VISUALIZE THE POSITION OF THE FIRST VALVE, WHICH APPEARED HIGHER THAN ANTICIPATED. IT WAS NOTED THAT WHEN THE VALVE WAS MOVED, A DISSECTION IN THE NON-CORONARY CUSP OCCURRED. THE SECOND VALVE WAS ADVANCED INSIDE THE FIRST VALVE AND DEPLOYED. THE OUTFLOW OF THE SECOND VALVE WAS COMPROMISED BY THE WAIST OF THE SECOND VALVE. A 22MM NON-MEDTRONIC BALLOON WAS USED TO POST DILATE, BUT THERE WAS NO CHANGE TO THE APPEARANCE OF THE SECOND VALVE. AN AORTOGRAM WAS PERFORMED AND CONFIRMED THAT A DISSECTION HAD OCCURRED ASSOCIATED WITH THE ROOT OF THE AORTA. THE PATIENT'S HEMODYNAMICS WERE TRYING TO COMPENSATE. AN ECHOCARDIOGRAM WAS PERFORMED AND A PERICARDIAL EFFUSION WAS OBSERVED. THE TEAM CONTINUED TO SUPPORT THE PATIENT'S RECOVERY FOR 50 MINUTES INCLUDING ATTEMPTING CARDIOPULMONARY RESUSCITATION, OBSERVING VARIOUS GASES, ECHOCARDIOGRAPHY, AND HEMODYNAMICS; HOWEVER, WERE UNSUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DECEASED.Event Description: IT WAS REPORTED THAT DURING THE IMPLANTATION OF THE EVOLUT VALVE, MODERATE CALCIUM WAS PRESENT IN THE ANNULUS. A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WITH A 20MM NON-MEDTRONIC (OSYPKA) BALLOON WAS PERFORMED. THE VALVE WAS POSITIONED IN THE ANNULUS USING CUSP OVERLAP TECHNIQUE AT 3MM DEPTH. A GRADIENT OF 14MMHG WAS OBSERVED POST-IMPLANTATION. THE IMPLANTING PHYSICIAN NOTED THAT THE PIGTAIL CATHETER IN THE LEFT VENTRICLE WASN'T DEEP ENOUGH AND THAT MIGHT BE CONTRIBUTING TO THE GRADIENT. THE PIGTAIL WAS POSITIONED JUST BELOW THE SKIRT OF THE VALVE. THE PHYSICIAN ATTEMPTED TO REPOSITION THE PIGTAIL FURTHER INTO THE CHAMBER OF THE LEFT VENTRICLE A COUPLE OF TIMES BUT THE PIGTAIL DID NOT MOVE FORWARD. ONE MORE PUSH RESULTED IN THE VALVE DISLODGING UP INTO THE AORTIC ARCH. A SECOND EVOLUT WAS LOADED TO POSITION IN THE ANNULUS, AND THE NATIVE AORTIC VALVE WAS RECROSSED WITH THE NON-MEDTRONIC GUIDEWIRE. THE SECOND VALVE AND DELIVERY CATHETER SYSTEM (DCS) WERE ADVANCED INTO THE DESCENDING AORTA; HOWEVER, THE SYSTEM WAS NOT ABLE TO CROSS AROUND THE ARCH DUE TO THE FIRST DISLODGED VALVE. WITH CONTINUED ATTEMPTS TO ADVANCE THE SYSTEM, THE INITIALLY IMPLANTED VALVE MOVED WITH THE SECOND DCS. THE FIRST VALVE WAS THEN DEEPER THAN THE ANNULUS, AND A SNARE WAS USED TO HOLD THE FIRST VALVE IN PLACE. THE VALVE WAS CAPTURED WITH THE SNARE AND PULLED BACK TO SIT JUST ABOVE THE CUSPS. RESISTANCE WAS NOTED DURING AN ATTEMPT TO PUT A PIGTAIL CATHETER ALONGSIDE THE FIRST VALVE TO VISUALIZE THE POSITION OF THE FIRST VALVE, WHICH APPEARED HIGHER THAN ANTICIPATED. IT WAS NOTED THAT WHEN THE VALVE WAS MOVED, A DISSECTION IN THE NON-CORONARY CUSP OCCURRED. THE SECOND VALVE WAS ADVANCED INSIDE THE FIRST VALVE AND DEPLOYED. THE OUTFLOW OF THE SECOND VALVE WAS COMPROMISED BY THE WAIST OF THE SECOND VALVE. A 22MM NON-MEDTRONIC (TRUE) BALLOON WAS USED TO POST DILATE, BUT THERE WAS NO CHANGE TO THE APPEARANCE OF THE SECOND VALVE. AN AORTOGRAM WAS PERFORMED AND CONFIRMED THAT A DISSECTION HAD OCCURRED ASSOCIATED WITH THE ROOT OF THE AORTA. THE PATIENT'S HEMODYNAMICS WERE TRYING TO COMPENSATE. AN ECHOCARDIOGRAM WAS PERFORMED AND A PERICARDIAL EFFUSION WAS OBSERVED. THE TEAM CONTINUED TO SUPPORT THE PATIENT'S RECOVERY FOR 50 MINUTES INCLUDING ATTEMPTING CARDIOPULMONARY RESUSCITATION, OBSERVING VARIOUS GASES, ECHOCARDIOGRAPHY, AND HEMODYNAMICS; HOWEVER, WERE UNSUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DECEASED. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT THE FIRST VALVE COMPROMISED THE OUTFLOW PORTION OF THE SECOND VALVE FRAME AND PREVENTED FULL EXPANSION. THE PATIENT'S CAUSE OF DEATH WAS AORTIC DISSECTION AND RIGHT CORONARY OCCLUSION. PER THE PHYSICIAN, NEITHER THE VALVE NOR THE DCS CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: PRODUCT ID: EVFXPLUS-26, SERIAL/LOT #: (B)(6), PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID EVFXPLUS-26 (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21963707&pc=BZD | 2518422-2025-105203 | 2022-05-15 04:00 | Death | RESPIRONICS, INC. | 2025-05-05 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | K113068 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT OF THE REMSTAR AUTO A-FLEX DEVICE PASSED AWAY FROM RESPIRATORY FAILURE. THERE WAS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FOLLOW UP REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21963820&pc=ONU | 2124215-2025-27133 | 2024-11-18 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-05 04:00 | ONU | RANGER | Adverse Event Without Identified Device or Use Problem | Ischemia |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. ON (B)(6) 2024, THE SUBJECT UNDERWENT TREATMENT WITH THE RANGER DRUG COATED BALLOON AS A PART OF THE CLINICAL TRIAL. THE TARGET LESION WAS IN THE RIGHT DISTAL SUPERFICIAL FEMORAL ARTERY WITH 4.9 MM PROXIMAL REFERENCE VESSEL DIAMETER AND 3.9 MM DISTAL REFERENCE VESSEL DIAMETER WITH LESION LENGTH 85 MM WITH 100% STENOSIS AND WAS CLASSIFIED AS TRANSATLANTIC INTER-SOCIETY CONSENSUS (TASC II C) LESION. PRIOR TO THE TARGET LESION TREATMENT WITH THE STUDY DEVICE PRE-DILATION WAS PERFORMED USING 3 MM X 150 MM AND 5 MM X 200 MM STERLING PTA BALLOONS. TREATMENT OF TARGET LESION WAS PERFORMED BY DILATION OF 6 MM X 150 MM RANGER DRUG-COATED BALLOON, STUDY DEVICE. FOLLOWING TREATMENT, 6 MM X 80 MM PULSAR-18 T3 DRUG ELUTING STENT WAS PLACED, AND THE FINAL RESIDUAL STENOSIS WAS NOTED TO BE 0%. ON (B)(6) 2024, THE SUBJECT WAS DISCHARGED FROM HOSPITAL ON ASPIRIN. ON (B)(6) 2024, THE SUBJECT WAS PRESENTED TO THE EMERGENCY DEPARTMENT (ED) WITH POOR PAIN CONTROL DUE TO PERIPHERAL VASCULAR DISEASE (PVD) AND ASSOCIATED NEUROPATHIC PAIN. THE SUBJECT RECEIVED PAIN MANAGEMENT WITH OXYNORM 5 MG WITHOUT EFFECT. IN ADDITION, SUBJECT WAS ALSO NOTED WITH INTERMITTENT DELIRIUM AND DECREASED APPETITE, POOR ORAL INTAKE. ON (B)(6) 2024, THE SUBJECT WAS UNSTEADY AND CANNOT BEAR WEIGHT ON THE RIGHT FOOT DUE TO PAIN. REGULAR PARACETAMOL WAS ADMINISTERED, AND A HEAT PACK WAS APPLIED TO THE RIGHT LEG FOR PAIN RELIEF. SUBJECT ALSO UNDERWENT CHEMICAL SYMPATHECTOMY ALONG WITH INCREASED TAPENTADOL 50 MG BD, BUPRENORPHINE 100-200 MG, KETAMINE LOZENGES 25MG. SUBJECT WAS ALSO NOTED WITH LOWER RESPIRATORY TRACT INFECTION (LRTI) AGAINST THE BACKGROUND OF ATELECTASIS WHICH WAS TREATED USING AMOXICILLIN AND DOXYCYCLINE. SUBJECT WAS OFTEN NOTED WITH LOW SP O2 OF 66-82%. HOWEVER, DURING DEEP INSPIRATIONS SP O2 IMPROVED TO 92% BUT THEN REDUCED LESS THAN 90% WITHOUT VERBAL STIMULATION. SUBJECT WAS VITAL SIGNS SUCH AS PH OF 7.33; P O2 OF 26; P CO2 OF 59; LACTATE OF 2.6. SUBJECT EXPERIENCED TYPE 2 RESPIRATORY DEPRESSION CONSISTENT WITH OPIATE TOXICITY. ON (B)(6) 2024, A REGIONAL SCIATIC BLOCK WAS INSERTED, AND THE DOSE OF ANESTHETIC WAS TITRATED UP RESULTING IN GOOD RELIEF FROM PAIN. SHE WOULD BENEFIT FROM CLOSE OVERNIGHT MONITORING AND SUPERVISION TO PREVENT FURTHER EVALUATION AND INCIDENTS WHILE SHE IS AWAITING HER PROCEDURE. ON (B)(6) 2024, IT WAS DECIDED AGAINST AMPUTATION OF THE LEG OR FURTHER SURGICAL OPTIONS DUE TO EXPLAINED POOR CANDIDACY FOR REHABILITATION. PRIORITIZED SYMPTOM RELIEF AND PAIN MANAGEMENT OVERACTIVE MEDICAL OR SURGICAL TREATMENT. INTRATHECAL CATHETER INSERTION FOR LONG-TERM PAIN RELIEF AND TRANSFER FOR PALLIATIVE CARE. ON (B)(6) 2024, THE SUBJECT WAS DISCHARGED FROM HOSPITAL ON ASPIRIN. ON (B)(6) 2024,146 DAYS POST INDEX PROCEDURE, THE SUBJECT DIED. THE ETIOLOGY OF CHRONIC LIMB THREATING ISCHEMIA THAT LED TO DEATH WAS STILL TO BE CONFIRMED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED OR NOT. THE MEDICAL PROFESSIONAL CONFIRMED IT IS UNLIKELY THAT THE DEVICE IS THE CAUSE OF THE ISCHEMIA SINCE THE SUBJECT IS A WELL-KNOWN VASCULOPATHY, BUT IT COULD NOT BE RULED OUT INDEFINITELY.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. ON (B)(6) 2024, THE SUBJECT UNDERWENT TREATMENT WITH THE RANGER DRUG COATED BALLOON AS A PART OF THE CLINICAL TRIAL. THE TARGET LESION WAS IN THE RIGHT DISTAL SUPERFICIAL FEMORAL ARTERY WITH 4.9 MM PROXIMAL REFERENCE VESSEL DIAMETER AND 3.9 MM DISTAL REFERENCE VESSEL DIAMETER WITH LESION LENGTH 85 MM WITH 100% STENOSIS AND WAS CLASSIFIED AS TRANSATLANTIC INTER-SOCIETY CONSENSUS (TASC II C) LESION. PRIOR TO THE TARGET LESION TREATMENT WITH THE STUDY DEVICE PRE-DILATION WAS PERFORMED USING 3 MM X 150 MM AND 5 MM X 200 MM STERLING PTA BALLOONS. TREATMENT OF TARGET LESION WAS PERFORMED BY DILATION OF 6 MM X 150 MM RANGER DRUG-COATED BALLOON, STUDY DEVICE. FOLLOWING TREATMENT, 6 MM X 80 MM PULSAR-18 T3 DRUG ELUTING STENT WAS PLACED, AND THE FINAL RESIDUAL STENOSIS WAS NOTED TO BE 0%. ON (B)(6) 2024, THE SUBJECT WAS DISCHARGED FROM HOSPITAL ON ASPIRIN. ON (B)(6) 2024, THE SUBJECT WAS PRESENTED TO THE EMERGENCY DEPARTMENT (ED) WITH POOR PAIN CONTROL DUE TO PERIPHERAL VASCULAR DISEASE (PVD) AND ASSOCIATED NEUROPATHIC PAIN. THE SUBJECT RECEIVED PAIN MANAGEMENT WITH OXYNORM 5 MG WITHOUT EFFECT. IN ADDITION, SUBJECT WAS ALSO NOTED WITH INTERMITTENT DELIRIUM AND DECREASED APPETITE, POOR ORAL INTAKE. ON (B)(6) 2024, THE SUBJECT WAS UNSTEADY AND CANNOT BEAR WEIGHT ON THE RIGHT FOOT DUE TO PAIN. REGULAR PARACETAMOL WAS ADMINISTERED, AND A HEAT PACK WAS APPLIED TO THE RIGHT LEG FOR PAIN RELIEF. SUBJECT ALSO UNDERWENT CHEMICAL SYMPATHECTOMY ALONG WITH INCREASED TAPENTADOL 50 MG BD, BUPRENORPHINE 100-200 MG, KETAMINE LOZENGES 25MG. SUBJECT WAS ALSO NOTED WITH LOWER RESPIRATORY TRACT INFECTION (LRTI) AGAINST THE BACKGROUND OF ATELECTASIS WHICH WAS TREATED USING AMOXICILLIN AND DOXYCYCLINE. SUBJECT WAS OFTEN NOTED WITH LOW SP O2 OF 66-82%. HOWEVER, DURING DEEP INSPIRATIONS SP O2 IMPROVED TO 92% BUT THEN REDUCED LESS THAN 90% WITHOUT VERBAL STIMULATION. SUBJECT WAS VITAL SIGNS SUCH AS PH OF 7.33; P O2 OF 26; P CO2 OF 59; LACTATE OF 2.6. SUBJECT EXPERIENCED TYPE 2 RESPIRATORY DEPRESSION CONSISTENT WITH OPIATE TOXICITY. ON (B)(6) 2024, A REGIONAL SCIATIC BLOCK WAS INSERTED, AND THE DOSE OF ANESTHETIC WAS TITRATED UP RESULTING IN GOOD RELIEF FROM PAIN. SHE WOULD BENEFIT FROM CLOSE OVERNIGHT MONITORING AND SUPERVISION TO PREVENT FURTHER EVALUATION AND INCIDENTS WHILE SHE IS AWAITING HER PROCEDURE. ON (B)(6) 2024, IT WAS DECIDED AGAINST AMPUTATION OF THE LEG OR FURTHER SURGICAL OPTIONS DUE TO EXPLAINED POOR CANDIDACY FOR REHABILITATION. PRIORITIZED SYMPTOM RELIEF AND PAIN MANAGEMENT OVERACTIVE MEDICAL OR SURGICAL TREATMENT. INTRATHECAL CATHETER INSERTION FOR LONG-TERM PAIN RELIEF AND TRANSFER FOR PALLIATIVE CARE. ON (B)(6) 2024, THE SUBJECT WAS DISCHARGED FROM HOSPITAL ON ASPIRIN. ON (B)(6) 2024,146 DAYS POST INDEX PROCEDURE, THE SUBJECT DIED. THE ETIOLOGY OF CHRONIC LIMB THREATING ISCHEMIA THAT LED TO DEATH AND TO REPLY IN PI OPINION ABOUT THE PRIMARY, SECONDARY CAUSES THAT LED TO DEATH WAS STILL TO BE CONFIRMED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED OR NOT. IT WAS CONFIRMED THAT, IN THE OPINION OF THE PI, IT IS UNLIKELY THAT THE DEVICE IS THE CAUSE OF THE ISCHEMIA SINCE THE SUBJECT IS A WELL-KNOWN VASCULOPATHY, BUT IT CANNOT BE RULED OUT INDEFINITELY. IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY AT AN OUTSIDE INSTITUTION AND NO FURTHER INFORMATION WAS AVAILABLE TO DETERMINE THE EXACT CAUSE OF DEATH. HENCE, UPON QUERYING PER PI OPINION THE EVENT TERM WAS UPDATED TO UNKNOWN CAUSE OF DEATH.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. ON (B)(6) 2024, THE SUBJECT UNDERWENT TREATMENT WITH THE RANGER DRUG COATED BALLOON AS A PART OF THE CLINICAL TRIAL. THE TARGET LESION WAS IN THE RIGHT DISTAL SUPERFICIAL FEMORAL ARTERY WITH 4.9 MM PROXIMAL REFERENCE VESSEL DIAMETER AND 3.9 MM DISTAL REFERENCE VESSEL DIAMETER WITH LESION LENGTH 85 MM WITH 100% STENOSIS AND WAS CLASSIFIED AS TRANSATLANTIC INTER-SOCIETY CONSENSUS (TASC II C) LESION. PRIOR TO THE TARGET LESION TREATMENT WITH THE STUDY DEVICE PRE-DILATION WAS PERFORMED USING 3 MM X 150 MM AND 5 MM X 200 MM STERLING PTA BALLOONS. THE TREATMENT OF TARGET LESION WAS PERFORMED BY DILATION OF 6 MM X 150 MM RANGER DRUG-COATED BALLOON, STUDY DEVICE. FOLLOWING TREATMENT, 6 MM X 80 MM PULSAR-18 T3 DRUG ELUTING STENT WAS PLACED, AND THE FINAL RESIDUAL STENOSIS WAS NOTED TO BE 0%. IN ADDITION, DURING INDEX PROCEDURE, PERONEAL ARTERY WAS TREATED BY BALLOON ANGIOPLASTY USING 3MM BALLOON. ON (B)(6) 2024, THE SUBJECT WAS DISCHARGED FROM HOSPITAL ON ASPIRIN AND CLOPIDOGREL. ON (B)(6) 2024, THE SUBJECT PRESENTED TO THE EMERGENCY DEPARTMENT (ED) WITH POOR PAIN CONTROL SECONDARY TO PERIPHERAL VASCULAR DISEASE (PVD) AND ASSOCIATED NEUROPATHIC PAIN. THE SUBJECT RECEIVED PAIN MANAGEMENT BY OXYNORM 5 MG WITHOUT EFFECT. IN ADDITION, THE SUBJECT WAS ON HIGH FLOW NASAL PRONG (HFNP) WITH O2 SATURATION OF 90% AND ALSO NOTED HAVE INTERMITTENT DELIRIUM, DECREASED APPETITE AND POOR ORAL INTAKE. THE SUBJECT WAS ALERT AND ORIENTED WITH STABLE BLOOD PRESSURE AND HEART RATE. THE SPO2 WAS 92% ON 2 L NP AND 97% ON 4L NP. ON (B)(6) 2024, THE SUBJECT WAS UNSTEADY AND CANNOT BEAR WEIGHT ON THE RIGHT FOOT DUE TO PAIN. REGULAR PARACETAMOL WAS ADMINISTERED, AND A HEAT PACK WAS APPLIED TO THE RIGHT LEG FOR PAIN RELIEF. THE SUBJECT PREVIOUSLY UNDERWENT CHEMICAL SYMPATHECTOMY ((B)(6) 2024) ALONG WITH INCREASED TAPENTADOL 50 MG BD, BUPRENORPHINE 100-200 MG, KETAMINE LOZENGES 25MG. THE SUBJECT EXPERIENCED COUGH AND FEVER WAS ALSO NOTED WITH LOWER RESPIRATORY TRACT INFECTION (LRTI) AGAINST THE BACKGROUND OF ATELECTASIS WHICH WAS TREATED USING AMOXICILLIN AND DOXYCYCLINE. THE SUBJECT WAS OFTEN NOTED WITH LOW SP O2 OF 66-82%. HOWEVER, DURING DEEP INSPIRATIONS SP O2 IMPROVED TO 92% BUT THEN REDUCED LESS THAN 90% WITHOUT VERBAL STIMULATION. THE SUBJECT LABORATORY FINDINGS REVEALED PH OF 7.33; P O2PO2 OF 26; P CO2 OF 59; LACTATE OF 2.6. SUBJECT, SODIUM 137, POTASSIUM 5, UREA 10.5, CREATININE 105 (UNITS AND REFERENCE RANGES NOT PROVIDED). THE SUBJECT EXPERIENCED TYPE 2 RESPIRATORY DEPRESSION CONSISTENT WITH OPIATE TOXICITY. NO DELIRIUM WAS NOTED FOR PAST 24 HOURS AFTER STOPPING GABAPENTIN AND DUE TO RENAL IMPAIRMENT PREGABALIN WAS STOPPED. UPON DISCUSSION WITH FAMILY MEMBERS, THE SUBJECT GOAL OF CARE WAS CATEGORY C INDICATING A PALLIATIVE APPROACH DUE TO NON-CURABLE CONDITION. ON (B)(6) 2024, THE SUBJECT WAS SEEN BY PAIN MANAGEMENT AND WAS TO CONTINUE KETAMINE LOZENGES PRN, LIDOCAINE PATCH, BUPRENORPHINE PATCH ALONG WITH THERMAL LIST. THE VASCULAR TEAM WAS CONSULTED WITH CONCERN FOR SUPERIMPOSED INFECTION IN THE RIGHT LOWER LIMB. ON (B)(6) 2024, A REGIONAL SCIATIC BLOCK WAS INSERTED, AND THE DOSE OF ANAESTHETIC WAS TITRATED UP RESULTING IN GOOD RELIEF FROM PAIN. IT WAS SUGGESTED THAT THE SUBJECT WOULD BENEFIT FROM CLOSE OVERNIGHT MONITORING AND SUPERVISION TO PREVENT FURTHER EVALUATION AND INCIDENTS WHILE SHE IS AWAITING HER PROCEDURE. ON (B)(6) 2024, IT WAS DECIDED AGAINST AMPUTATION OF THE LEG OR FURTHER SURGICAL OPTIONS DUE TO EXPLAINED POOR CANDIDACY FOR REHABILITATION. THE CLINICAL APPROACH PRIORITIZED SYMPTOM RELIEF AND PAIN MANAGEMENT OVERACTIVE MEDICAL OR SURGICAL TREATMENT. THE SUBJECT WAS TRANSFERRED TO ANOTHER HOSPITAL FOR ELECTIVE ADMISSION FOR INTRATHECAL CATHETER INSERTION FOR LONG-TERM PAIN RELIEF. ON (B)(6) 2024, THE SUBJECT WAS DISCHARGED FROM HOSPITAL ON ASPIRIN TO HOSPICE FOR ONGOING PALLIATIVE CARE. ON (B)(6) 2024,146 DAYS POST INDEX PROCEDURE, THE SUBJECT DIED DUE TO UNKNOWN CAUSE. THE ETIOLOGY OF CHRONIC LIMB THREATING ISCHEMIA THAT LED TO DEATH AND TO REPLY IN PI OPINION ABOUT THE PRIMARY, SECONDARY CAUSES THAT LED TO DEATH WAS STILL TO BE CONFIRMED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED OR NOT. IT WAS CONFIRMED THAT, IN THE OPINION OF THE PI, IT IS UNLIKELY THAT THE DEVICE IS THE CAUSE OF THE ISCHEMIA SINCE THE SUBJECT IS A WELL-KNOWN VASCULOPATHY, BUT IT CANNOT BE RULED OUT INDEFINITELY. Manufacturer Narrative: A2 - AGE AT TIME OF EVENT: (B)(6) YEARS OLD AT THE TIME OF STUDY ENROLLMENT. Manufacturer Narrative: A2 - AGE AT TIME OF EVENT: (B)(6) YEARS OLD AT THE TIME OF STUDY ENROLLMENT. Manufacturer Narrative: A2 - AGE AT TIME OF EVENT: (B)(6) AT THE TIME OF STUDY ENROLLMENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21964853&pc=NKE | 2124215-2025-28793 | 2025-03-14 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-05 04:00 | NKE | VALITUDE? X4 CRT-P | Defective Device | No Clinical Signs, Symptoms or Conditions | P030005 |  | 1 | Event Description: IT WAS REPORTED THAT THIS CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) RECORDED A YELLOW ALERT IN THE REMOTE MONITORING SYSTEM FOR AN ARRHYTHMIA ON (B)(6) 2025. THE DATA WAS TRANSMITTED ON (B)(6) 2025, BUT THE ALERT WAS RECEIVED BY THE CLINIC TWO DAYS LATER. BASED ON THE ELECTROGRAM (EGM) IN THE VT EPISODE, IT APPEARED THE PATIENT HAD EXPIRED. TECHNICAL SERVICES REVIEWED THE EPISODE AND AGREED WITH THE HEALTH CARE PROFESSIONAL (HCP) AND RECOMMENDED A WELLNESS CHECK. THE PATIENT WAS FOUND DECEASED IN THEIR HOME ON (B)(6) 2025. LATITUDE CUSTOMER SUPPORT WAS CONTACTED AND DETERMINED THE REASON THE ALERT WAS DELAYED IS THAT THE DATA APPEARS TO HAVE BEEN TRAPPED IN THE COMMUNICATOR UNTIL IT PERFORMED A POWER CYCLE. THE DEVICE WAS SUBSEQUENTLY RECEIVED FOR ANALYSIS. A NOTE ON THE DEATH CERTIFICATE INDICATED THE CAUSE OF DEATH WAS ASYSTOLE DUE TO "PACEMAKER FAILURE" OR MYOCARDIAL INFARCTION, CAUSED BY CORONARY ARTERY DISEASE AND HEART FAILURE. THE MANNER OF DEATH WAS LISTED AS NATURAL. AN ALLEGATION AGAINST THE IMPLANTED DEVICE WAS NOT ABLE TO BE RULED OUT AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE SUBMITTED AFTER THE DEVICE ANALYSIS IS COMPLETED.Event Description: IT WAS REPORTED THAT THIS CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) RECORDED A YELLOW ALERT IN THE REMOTE MONITORING SYSTEM FOR AN ARRHYTHMIA ON (B)(6) 2025. THE DATA WAS TRANSMITTED ON (B)(6) 2025, BUT THE ALERT WAS RECEIVED BY THE CLINIC TWO DAYS LATER. BASED ON THE ELECTROGRAM (EGM) IN THE VT EPISODE, IT APPEARED THE PATIENT HAD EXPIRED. TECHNICAL SERVICES REVIEWED THE EPISODE AND AGREED WITH THE HEALTH CARE PROFESSIONAL (HCP) AND RECOMMENDED A WELLNESS CHECK. THE PATIENT WAS FOUND DECEASED IN THEIR HOME ON (B)(6) 2025. LATITUDE CUSTOMER SUPPORT WAS CONTACTED AND DETERMINED THE REASON THE ALERT WAS DELAYED IS THAT THE DATA APPEARS TO HAVE BEEN TRAPPED IN THE COMMUNICATOR UNTIL IT PERFORMED A POWER CYCLE. THE DEVICE WAS SUBSEQUENTLY RECEIVED FOR ANALYSIS. A NOTE ON THE DEATH CERTIFICATE INDICATED THE CAUSE OF DEATH WAS ASYSTOLE DUE TO "PACEMAKER FAILURE" OR MYOCARDIAL INFARCTION, CAUSED BY CORONARY ARTERY DISEASE AND HEART FAILURE. THE MANNER OF DEATH WAS LISTED AS NATURAL. AN ALLEGATION AGAINST THE IMPLANTED DEVICE WAS NOT ABLE TO BE RULED OUT AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE SUBMITTED AFTER THE DEVICE ANALYSIS IS COMPLETED. Manufacturer Narrative: THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS. THE RETURNED DEVICE WAS THOROUGHLY INSPECTED AND ANALYZED UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY. VISUAL EXAMINATION OF THE DEVICE HEADER AND CASE NOTED NO ANOMALIES. THE DEVICE WAS THEN EXPOSED TO SIMULATED HEART LOAD CONDITIONS, AND THE PACING AND SENSING FUNCTIONS WERE TESTED. IMPEDANCE TESTING WAS COMPLETED AND ALL MEASUREMENTS WERE WITHIN NORMAL LIMITS. THE DEVICE OPERATED APPROPRIATELY WITH NO INTERRUPTIONS IN THERAPY OUTPUT AT THE RETURNED PROGRAMMED SETTINGS. A SERIES OF ELECTRICAL TESTS WAS ALSO PERFORMED, AND AGAIN, NORMAL DEVICE FUNCTION WAS OBSERVED. Manufacturer Narrative: THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21964928&pc=KOC | 0008030665-2025-01024 | 2025-03-26 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-05-05 04:00 | KOC | COMBISET SMARTECH BLOODLINE (NO HEPARIN) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Exsanguination | K201207 |  | 1 | Event Description: IT WAS REPORTED TO FRESENIUS THAT A BLOOD LEAK OCCURRED DURING A HEMODIALYSIS (HD) PATIENT¿S TREATMENT, WHICH SUBSEQUENTLY RESULTED IN THEIR DEATH. ADDITIONAL DETAILS WERE PROVIDED UPON FOLLOW-UP AND REVIEW OF THE PATIENT¿S TREATMENT SHEET. THE PATIENT¿S PRE-TREATMENT VITAL SIGNS AT 10:33AM WERE CONFIRMED AS BLOOD PRESSURE (BP) 115/80, PULSE (P) 73, RESPIRATIONS (R) 16, TEMPERATURE (T) 96.6°. THE TREATMENT WAS INITIATED UTILIZING A CENTRAL VENOUS CATHETER (CVC) IN THE CHEST, ALONG WITH FRESENIUS COMBISET SMARTECH BLOODLINES, 2008T HD MACHINE, AND PATIENT CONNECTOR CLIP (NON-STERILE). THE CVC WAS REPORTEDLY WIPED WITH ALCOHOL PRIOR TO CONNECTING THE BLOODLINES. THE TREATMENT SHEET INDICATED THAT THE CVC CATHETER DRESSING WAS GAUZE WITH A SOLUTION OF 2% CHLORHEXIDINE GLUCONATE AND 70% ISOPROPYL ALCOHOL. THE ARTERIAL PRESSURE SETTING ON THE 2008T MACHINE WAS 160MMHG AND THE VENOUS PRESSURE WAS 100 ASYMMETRICAL. THE PATIENT DID NOT REPORT ANY CONCERNS PRIOR TO THE START OF TREATMENT. THE TREATMENT SHEET DOCUMENTED THE ULTRAFILTRATION RATE (UFR) 760, FLUID REMOVED 38, FLUID GIVEN 300, THE ACCESS AND CONNECTOR CLIP WERE CHECKED, THE PATIENT WAS RESTING COMFORTABLY, THE PATIENT WAS ALERT, THE ACCESS WAS VISIBLE, AND TREATMENT WAS INITIATED WITHOUT PROBLEMS. THE GREEN AMP LIGHT WAS ON. THE FOLLOWING VITAL SIGNS AND NOTES WERE RECORDED DURING THE TREATMENT: 11:00AM BP 144/54, P 64, ACCESS AND CONNECTOR CLIP CHECKED; 11:02AM BP 114/64, P 64, ACCESS AND CONNECTOR CLIP CHECKED, PATIENT WAS ALERT WITH NO COMPLAINTS, THE GREEN AMP LIGHT WAS ON; 11:31AM BP 154/71, P 66, ARTERIAL PRESSURE 182, VENOUS PRESSURE 125, GREEN AMP LIGHT ON, ACCESS VISIBLE, PATIENT DENIES COMPLAINTS, ACCESS AND CONNECTOR CLIP CHECKED, UFR 760, FLUID REMOVED 723. AT 11:45AM THE MACHINE ALARMED WITH AN UNSPECIFIED MACHINE ALARM (INITIALLY REPORTED AS AN ARTERIAL PRESSURE ALARM). THE PATIENT¿S BP WAS DOCUMENTED AS BP 87/31, P 125. THE TREATMENT SHEET DOCUMENTED THAT THE ACCESS WAS CHECKED, AND THE GREEN AMP LIGHT WAS ON. AT THIS TIME, IT WAS NOTED THERE WAS A BLOOD LEAK. THE LEAK WAS REPORTEDLY COMING FROM THE CONNECTION OF THE CATHETER LUMEN TO THE COMBISET VENOUS BLOOD LINE. THE CONNECTION WAS LOOSE, BUT NOT COMPLETELY DISCONNECTED. THE PATIENT CONNECTOR CLIP WAS REPORTEDLY STILL CONNECTED TO THE BLOOD LINE AND CATHETER. WHEN THE BLOOD LEAK WAS DETECTED, THE PATIENT WENT UNRESPONSIVE. A PULSE WAS NOT PALPABLE. TREATMENT WAS DISCONTINUED AND THE PATIENT¿S BLOOD FROM THE EXTRACORPOREAL CIRCUIT WAS IMMEDIATELY RETURNED WITH NORMAL SALINE (NS). CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. ANOTHER STAFF MEMBER ACTIVATED 911. THE PATIENT WAS ADMINISTERED OXYGEN VIA AMBU MASK. 3L OF NS WERE ADMINISTERED WHILE WAITING FOR EMERGENCY MEDICAL SERVICES (EMS). NO PATIENT VITALS WERE DOCUMENTED AFTER BECOMING UNRESPONSIVE. EMS ARRIVED AND IMMEDIATELY TRANSPORTED THE PATIENT TO THE HOSPITAL. AT THE HOSPITAL, THE PATIENT RECEIVED CPR, MULTIPLE MEDICATIONS (UNKNOWN), AND WAS INTUBATED. THE PATIENT EXPIRED AT 4:00PM. THE CAUSE OF DEATH DOCUMENTED IN THE MEDICAL RECORD DISCHARGE IS CARDIAC ARREST, ANEMIA, AND END STAGE RENAL DISEASE (ESRD). NO ESRD DEATH NOTIFICATION FORM CMS-2746-U2 WAS AVAILABLE. THE MEDICAL RECORD DOCUMENTED THE PATIENT¿S ESTIMATED BLOOD LOSS (EBL) AT 2L. IT WAS DOCUMENTED THAT PRIOR TO THE START OF THE PATIENT¿S TREATMENT, THE CONNECTION AT THE COMBISET VENOUS BLOODLINE AND CATHETER WAS SECURE. THE PATIENT CONNECTOR CLIP WAS REPORTEDLY ON AT THE START OF TREATMENT AND AT THE TIME OF THE BLOOD LEAK, HOWEVER; IT COULD NOT BE CONFIRMED IF THE CLIP REMAINED ON THE BLOODLINE POST EVENT. THE PATIENT DID NOT COMPLAIN OF ANY SYMPTOMS AT ANY TIME PRIOR TO THE ADVERSE EVENT. THERE WERE NO ADDITIONAL ALARMS OR TREATMENT ISSUES DOCUMENTED PRIOR TO THE UNKNOWN MACHINE ALARM. THE PATIENT¿S ACCESS SITE WAS REPORTED TO BE VISIBLE DURING THE TREATMENT. NO FURTHER INFORMATION WAS PROVIDED. THE SAMPLE WAS REPORTED TO BE AVAILABLE FOR MANUFACTURER EVALUATION.Event Description: IT WAS REPORTED TO FRESENIUS THAT A BLOOD LEAK OCCURRED DURING A HEMODIALYSIS (HD) PATIENT¿S TREATMENT, WHICH SUBSEQUENTLY RESULTED IN THEIR DEATH. ADDITIONAL DETAILS WERE PROVIDED UPON FOLLOW-UP AND REVIEW OF THE PATIENT¿S TREATMENT SHEET. THE PATIENT¿S PRE-TREATMENT VITAL SIGNS AT 10:33AM WERE CONFIRMED AS BLOOD PRESSURE (BP) 115/80, PULSE (P) 73, RESPIRATIONS (R) 16, TEMPERATURE (T) 96.6°. THE TREATMENT WAS INITIATED UTILIZING A CENTRAL VENOUS CATHETER (CVC) IN THE CHEST, ALONG WITH FRESENIUS COMBISET SMARTECH BLOODLINES, 2008T HD MACHINE, AND PATIENT CONNECTOR CLIP (NON-STERILE). THE CVC WAS REPORTEDLY WIPED WITH ALCOHOL PRIOR TO CONNECTING THE BLOODLINES. THE TREATMENT SHEET INDICATED THAT THE CVC CATHETER DRESSING WAS GAUZE WITH A SOLUTION OF 2% CHLORHEXIDINE GLUCONATE AND 70% ISOPROPYL ALCOHOL. THE ARTERIAL PRESSURE SETTING ON THE 2008T MACHINE WAS 160MMHG AND THE VENOUS PRESSURE WAS 100 ASYMMETRICAL. THE PATIENT DID NOT REPORT ANY CONCERNS PRIOR TO THE START OF TREATMENT. THE TREATMENT SHEET DOCUMENTED THE ULTRAFILTRATION RATE (UFR) 760, FLUID REMOVED 38, FLUID GIVEN 300, THE ACCESS AND CONNECTOR CLIP WERE CHECKED, THE PATIENT WAS RESTING COMFORTABLY, THE PATIENT WAS ALERT, THE ACCESS WAS VISIBLE, AND TREATMENT WAS INITIATED WITHOUT PROBLEMS. THE GREEN AMP LIGHT WAS ON. THE FOLLOWING VITAL SIGNS AND NOTES WERE RECORDED DURING THE TREATMENT: 11:00AM BP 144/54, P 64, ACCESS AND CONNECTOR CLIP CHECKED; 11:02AM BP 114/64, P 64, ACCESS AND CONNECTOR CLIP CHECKED, PATIENT WAS ALERT WITH NO COMPLAINTS, THE GREEN AMP LIGHT WAS ON; 11:31AM BP 154/71, P 66, ARTERIAL PRESSURE 182, VENOUS PRESSURE 125, GREEN AMP LIGHT ON, ACCESS VISIBLE, PATIENT DENIES COMPLAINTS, ACCESS AND CONNECTOR CLIP CHECKED, UFR 760, FLUID REMOVED 723. AT 11:45AM THE MACHINE ALARMED WITH AN UNSPECIFIED MACHINE ALARM (INITIALLY REPORTED AS AN ARTERIAL PRESSURE ALARM). THE PATIENT¿S BP WAS DOCUMENTED AS BP 87/31, P 125. THE TREATMENT SHEET DOCUMENTED THAT THE ACCESS WAS CHECKED, AND THE GREEN AMP LIGHT WAS ON. AT THIS TIME, IT WAS NOTED THERE WAS A BLOOD LEAK. THE LEAK WAS REPORTEDLY COMING FROM THE CONNECTION OF THE CATHETER LUMEN TO THE COMBISET VENOUS BLOOD LINE. THE CONNECTION WAS LOOSE, BUT NOT COMPLETELY DISCONNECTED. THE PATIENT CONNECTOR CLIP WAS REPORTEDLY STILL CONNECTED TO THE BLOOD LINE AND CATHETER. WHEN THE BLOOD LEAK WAS DETECTED, THE PATIENT WENT UNRESPONSIVE. A PULSE WAS NOT PALPABLE. TREATMENT WAS DISCONTINUED AND THE PATIENT¿S BLOOD FROM THE EXTRACORPOREAL CIRCUIT WAS IMMEDIATELY RETURNED WITH NORMAL SALINE (NS). CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. ANOTHER STAFF MEMBER ACTIVATED 911. THE PATIENT WAS ADMINISTERED OXYGEN VIA AMBU MASK. 3L OF NS WERE ADMINISTERED WHILE WAITING FOR EMERGENCY MEDICAL SERVICES (EMS). NO PATIENT VITALS WERE DOCUMENTED AFTER BECOMING UNRESPONSIVE. EMS ARRIVED AND IMMEDIATELY TRANSPORTED THE PATIENT TO THE HOSPITAL. AT THE HOSPITAL, THE PATIENT RECEIVED CPR, MULTIPLE MEDICATIONS (UNKNOWN), AND WAS INTUBATED. THE PATIENT EXPIRED AT 4:00PM. THE CAUSE OF DEATH DOCUMENTED IN THE MEDICAL RECORD DISCHARGE IS CARDIAC ARREST, ANEMIA, AND END STAGE RENAL DISEASE (ESRD). NO ESRD DEATH NOTIFICATION FORM CMS-2746-U2 WAS AVAILABLE. THE MEDICAL RECORD DOCUMENTED THE PATIENT¿S ESTIMATED BLOOD LOSS (EBL) AT 2L. IT WAS DOCUMENTED THAT PRIOR TO THE START OF THE PATIENT¿S TREATMENT, THE CONNECTION AT THE COMBISET VENOUS BLOODLINE AND CATHETER WAS SECURE. THE PATIENT CONNECTOR CLIP WAS REPORTEDLY ON AT THE START OF TREATMENT AND AT THE TIME OF THE BLOOD LEAK, HOWEVER; IT COULD NOT BE CONFIRMED IF THE CLIP REMAINED ON THE BLOODLINE POST EVENT. THE PATIENT DID NOT COMPLAIN OF ANY SYMPTOMS AT ANY TIME PRIOR TO THE ADVERSE EVENT. THERE WERE NO ADDITIONAL ALARMS OR TREATMENT ISSUES DOCUMENTED PRIOR TO THE UNKNOWN MACHINE ALARM. THE PATIENT¿S ACCESS SITE WAS REPORTED TO BE VISIBLE DURING THE TREATMENT. NO FURTHER INFORMATION WAS PROVIDED. THE SAMPLE WAS REPORTED TO BE AVAILABLE FOR MANUFACTURER EVALUATION. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN HD THERAPY UTILIZING THE COMBISET SMARTECH BLOODLINES AND THE PATIENT EVENT OF BLOOD LOSS WITH SUBSEQUENT DEATH. ALTHOUGH IT WAS REPORTED THAT A BLOOD LEAK OCCURRED FROM THE CONNECTION OF THE CATHETER LUMEN AND THE COMBISET VENOUS BLOOD LINE, THERE IS NO DOCUMENTATION OF ANY DEFECTS OR DAMAGE NOTED AT THE CONNECTION SITE PRIOR TO THE START OF TREATMENT. ADDITIONALLY, THERE WERE NO OTHER REPORTED ISSUES LEADING UP TO THE EVENT THAT WOULD CAUSE THE REPORTED LOOSE CONNECTION. IT WAS REPORTED THAT THE CONNECTION WAS SECURE PRIOR TO THE START OF TREATMENT. IN ADDITION, THERE ARE SEVERAL DOCUMENTED CHECKS THROUGHOUT THE TREATMENT THAT INDICATE THE PATIENT ACCESS, AND THE PATIENT CONNECTOR CLIP WERE CHECKED. THE COMBISET SMARTECH BLOODLINE IFU STATES TO ¿ENSURE ALL CONNECTIONS ARE SECURE BEFORE USE AND MONITOR FOR LEAKS REGULARLY DURING PATIENT USE. BLOOD LEAKS CAN RESULT IF CONNECTIONS ARE NOT SECURE.¿ ALTHOUGH IT WAS REPORTED THAT THE PATIENT CONNECTOR CLIP WAS PLACED PRIOR TO THE INITIATION OF TREATMENT AND DOCUMENTED AS STILL BEING CONNECTED AT THE TIME OF THE EVENT, IT COULD NOT BE CONFIRMED IF THE CLIP WAS STILL ATTACHED POST EVENT. THE COMBISET SMARTECH BLOODLINES IFU ADDITIONALLY STATES ¿USE OF CATHETERS, AVF NEEDLE SETS OR OTHER DEVICES NOT COMPATIBLE WITH THIS SET MAY RESULT IN BLOOD/AIR LEAKAGE OR ACCIDENTAL DISCONNECTION, AIR EMBOLUS, SIGNIFICANT BLOOD LOSS, PATIENT INJURY OR DEATH.¿ IT WAS ALSO DOCUMENTED SEVERAL TIMES THAT THE PATIENT¿S ACCESS SITE WAS UNCOVERED AND VISIBLE DURING TREATMENT. Manufacturer Narrative: ADDITIONAL INFORMATION: D4, D9, H3, H4 PLANT INVESTIGATION: THE COMPLAINT SAMPLE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. THE DEVICE WAS NOT RETURNED IN ITS ORIGINAL PACKAGING. THE RETURNED SAMPLE INCLUDED BOTH AN ARTERIAL AND A VENOUS LINE. THE PATIENT CONNECTOR CLIP WAS NOT ATTACHED OR INCLUDED. THE SAMPLE UNDERWENT FUNCTIONAL TESTING AND WAS DISINFECTED ACCORDING TO PLANT PROCEDURES. NO LEAKS, KINKS, OR OTHER FAILURES WERE OBSERVED DURING DISINFECTION OF THE SAMPLE. NO IRREGULARITIES WERE FOUND DURING VISUAL INSPECTION OF THE RETURNED DEVICE. THE ARTERIAL AND VENOUS PATIENT END-CONNECTORS WERE CLOSELY INSPECTED FOR DAMAGE (INCLUDING COLLAR AND TAPER DAMAGE), AND NO DEFECTS WERE FOUND. SINCE THE PATIENT CONNECTOR CLIP WAS NOT RETURNED WITH THE SAMPLE, THE DEVICE WAS TESTED WITHOUT IT. THERE WERE NO DISCONNECTIONS OR LEAKS DURING TESTING. NO AIR BUBBLES WERE NOTED IN THE CIRCUIT, AND THERE WERE NO NOTED LEVEL VARIATIONS OF THE VENOUS OR ARTERIAL CHAMBERS. THERE WERE NO ALARMS EITHER. THE DIMENSIONS OF THE MALE CONICAL FITTINGS ON THE ARTERIAL AND VENOUS PATIENT CONNECTORS WERE INSPECTED AND FOUND TO BE ACCEPTABLE. THE DEVICE WORKED AS INTENDED, WITHOUT ANY PROBLEMS OR FAILURES. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE COMBISET WORKED AS INTENDED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21957263&pc=OZP | 2032227-2025-166291 | 2025-04-04 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-04 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025 AT HOME. THE CAUSE OF DEATH WAS COMPLICATIONS DUE TO DIABETES. THE INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-396A, MMT-1884, MMT-7040A. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING. THE CUSTOMER HAS BEEN USING THE INSULIN PUMP SYSTEM WITHIN 48HOURS OF THE REPORTED EVENT, AND THE CUSTOMER WAS USING SENSORS. THERE IS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-396A. PRODUCT RETURN REQUESTED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21957299&pc=BZD | 2518422-2025-105145 | 2025-05-01 04:00 | Death | RESPIRONICS, INC. | 2025-05-04 04:00 | BZD | REMSTAR AUTO A-FLEX W/HUMID, SD CARD | Adverse Event Without Identified Device or Use Problem | Cancer | K091319 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REMSTAR AUTO DEVICE DEVELOPED CANCER AND HAS PASSED AWAY. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION THAT A PATIENT WITH A REMSTAR AUTO DEVICE DEVELOPED CANCER AND HAS PASSED AWAY. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER AS PART OF THE UNO RECALL PROCESS. DURING THE EVALUATION OF THE DEVICE, THE SERVICE TECHNICIAN OBSERVED BLOWER FOAM DEGRADATION. THE DEVICE WAS SCRAPPED BY THE SERVICE CENTER AFTER THE EVALUATION WAS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21957302&pc=BZD | 2518422-2025-105141 | 2024-02-27 05:00 | Death | RESPIRONICS, INC. | 2025-05-04 04:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Sepsis |  |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH AN UNKNOWN PHILIPS DEVICE DEVELOPED SEPSIS AND PASSED AWAY. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21957321&pc=BZD | 2518422-2025-105144 | 2024-08-17 04:00 | Death | RESPIRONICS, INC. | 2025-05-04 04:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION DEVICE DEVELOPED MELANOMA AND HAS PASSED AWAY. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21957610&pc=NAY | 2955842-2025-18306 | 2025-03-24 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-05-04 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Myocardial Infarction | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI-ASSISTED VENTRAL HERNIA ¿ COMPLEX TRANSVERSUS ABDOMINUS RELEASE (TAR) REPAIR, THE PATIENT WAS KEPT OVERNIGHT FOR OBSERVATION AND LABORATORY RESULTS. THE SURGEON INFORMED THE INTUITIVE SURGICAL CLINICAL SALES REPRESENTATIVE (CSR) THAT THE PATIENT WAS DOING WELL POST-OPERATIVELY; HOWEVER, EXPIRED OVER NIGHT DUE TO UNKNOWN CAUSES. THE SURGEON REPORTED THAT THE CAUSE OF DEATH WAS LIKELY A HEART ATTACK OR A PULMONARY EMBOLISM. THE SURGEON SPECIFIED THAT THE PATIENT OUTCOME WAS NOT CAUSED BY THE DA VINCI-ASSISTED PROCEDURE AND THERE WERE NO DA VINCI ISSUES DURING THE PROCEDURE. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION WERE REQUESTED; NO RESPONSE WAS RECEIVED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, MARYLAND BIPOLAR FORCEPS, FENESTRATED BIPOLAR FORCEPS, MEGA SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING OR SUBSEQUENT TO THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED UNEXPECTEDLY AFTER AN UNEVENTFUL HERNIA REPAIR. NO ALLEGATION AGAINST ANY INTUITIVE SURGICAL PRODUCT HAS BEEN MADE AND THE CASE HAD NO ISSUES. THE SURGEON STATED THE PATIENT MAY HAVE DIED OF A MYOCARDIAL INFARCTION OR PULMONARY EMBOLISM BUT NO SPECIFIC INFORMATION ALLUDING TO THESE CAUSES IS PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION B2 DATE OF DEATH IS ESTIMATED - SURGEON REPORTED THE PATIENT EXPIRED DURING THE NIGHT, TIME NOT PROVIDED. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOUND THAT NO ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21948300&pc=OZP | 2032227-2025-165222 | 2025-03-03 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-02 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY AT HOME ON (B)(6) 2025. THE DATE THE CUSTOMER WAS NOT ADMITTED TO THE HOSPITAL. THE CAUSE OF DEATH WAS HEART DIABETES, KIDNEY DISEASE, AND SLEEP APNEA. THE CUSTOMER¿S BLOOD GLUCOSE WAS UNKNOWN. THE EVENT INVOLVED PRODUCT(S) MMT-1884, UNK_SENSOR, UNK_RESERVOIR, UNOMEDICAL. TROUBLESHOOTING WAS NOT PERFORMED. IT WAS UNKNOWN THAT THE CUSTOMER WAS WEARING THE INSULIN PUMP AT THE TIME OF DEATH. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884. NO PRODUCT RETURN IS REQUIRED FOR UNK_SENSOR. NO PRODUCT RETURN IS REQUIRED FOR UNK_RESERVOIR. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21948389&pc=IZQ | 9612420-2025-00101 | 2025-02-20 05:00 | Death | ULRICH GMBH & CO. KG | 2025-05-02 04:00 | IZQ | ULRICH MEDICAL | Adverse Event Without Identified Device or Use Problem | Anaphylactoid | K171392 |  | 1 | Event Description: ACCORDING TO REPORTING PERSON: "WE HAVE BEEN INFORMED TODAY OF SERIOUS ALLERGIC INCIDENT RESULTING IN THE DEATH OF 1 PATIENT AT THE (B)(6) HOSPITAL. THIS INCIDENT HAS OCCURRED IN FEBRUARY, A PATIENT WAS INJECTED WITH IOMERON USING THE ULRICH INJECTOR. THE PATIENT TUBING XD 2040 ANNOUNCED TO BE INVOLVED COMES FROM BATCH E000095." Manufacturer Narrative: THE MENTIONED ARTICLE XD 2040 IS NOT SOLD IN THE USA. HOWEVER, XD 2042 IS FUNCTIONALLY IDENTICAL AND IS SOLD IN THE USA. BASED ON THE INFORMATION AVAILABLE, THE DEVICE CAN BE EXCLUDED AS A SOURCE OF ERROR. IT IS APPROVED BY THE AUTHORITIES FOR USE WITH THE SAID CONTRAST MEDIUM. ALL DEVICES ON SITE HAVE SUCCESSFULLY PASSED A TECHNICAL SAFETY TEST, THERE IS NO INDICATION OF A ENDOGENOUS CAUSE OF THE INJECTION SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21948403&pc=IZQ | 9612420-2025-00102 | 2025-04-04 04:00 | Death | ULRICH GMBH & CO. KG | 2025-05-02 04:00 | IZQ | UL_INJECT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K171392 |  | 1 | Event Description: ACCORDING TO USER INFORMATION: "ON (B)(6) 2025 AT 11:45 A.M., MS. (B)(6) REPORTED THE DEATH OF A PATIENT DUE TO AIR IN THE HEART IN RELATION TO (B)(6). A CT THORAX ABDOMEN EXAMINATION WITH ULTRAVIST 300 IN 500 ML GLASS BOTTLES AT A FLOW RATE OF 3ML/S WITH 21ML NACL / 100ML KM / 30ML NACL. THE CONTRAST AGENT WAS NOT PREHEATED. A 320 CM PATIENT TUBE WAS USED. ACCESS WAS VIA A PORT SYSTEM THAT IS ALSO USED FOR CHEMOTHERAPY. MS. (B)(6) WAS UNABLE TO STATE THE AMOUNT OF AIR. THE INJECTOR DID NOT SHOW ANY ERROR MESSAGES."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21948481&pc=OZO | 2032227-2025-165288 | 2025-04-02 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-02 04:00 | OZO | 530G INSULIN PUMP MMT-551NAS | Appropriate Term/Code Not Available | Insufficient Information | P120010 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED CUSTOMER EXPRESSED DISSATISFACTION ON PUMP AND CONTINUOUS GLUCOSE MONITORING SYSTEM (UNKNOWN ALLEGATION). THE CUSTOMER MENTIONED INTRAUTERINE FETAL. THE CAUSE OF DEATH WAS UNKNOWN. THE INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS MMT-7703A, MMT-551AS. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING. CUSTOMER HAS BEEN USING THE INSULIN PUMP SYSTEM WITHIN 48HRS OF REPORTED EVENT. THERE IS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-7703A. NO PRODUCT RETURN IS REQUIRED FOR MMT-551AS. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. THIS IS 1 OF 2 MEDWATCH REPORTS REGARDING THIS EVENT. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21948493&pc=MDS | 2032227-2025-165295 | 2025-04-02 04:00 | Death | MEDTRONIC MINIMED | 2025-05-02 04:00 | MDS | MINILINK TRANSMITTER MMT-7703NA | Appropriate Term/Code Not Available | Insufficient Information | P150019 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED CUSTOMER EXPRESSED DISSATISFACTION ON PUMP AND CONTINUOUS GLUCOSE MONITORING SYSTEM (UNKNOWN ALLEGATION). THE CUSTOMER MENTIONED INTRAUTERINE FETAL. THE CAUSE OF DEATH WAS UNKNOWN. THE INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS MMT-7703A, MMT-551AS. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING. CUSTOMER HAS BEEN USING THE INSULIN PUMP SYSTEM WITHIN 48HRS OF REPORTED EVENT. THERE IS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-7703A. NO PRODUCT RETURN IS REQUIRED FOR MMT-551AS. Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. THIS IS 2 OF 2 MEDWATCH REPORTS REGARDING THIS EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21948889&pc=OZD | 1220648-2025-28021 | 2025-04-11 04:00 | Death | ABIOMED, INC. | 2025-05-02 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Renal Failure | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED BLEEDING WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA DEVICE USED: ¿PATIENT ARRIVED AT FACILITY INTUBATED AND WAS FOUND TO BE IN BIVENTRICULAR FAILURE. WORSENING SEPTIC SHOCK AND ACIDOSIS PROMPTED INITIATION OF V-A ECMO AND IMPELLA PLACEMENT. PATIENT CODED ON TABLE BEFORE INITIATION OF ECMO OR IMPELLA. PATIENT WAS GIVEN 2 UNITS OF RBC¿S ON (B)(6) 2025 POST IMPELLA AND ECMO PLACEMENT AS INDICATED BY HER UNSTABLE ANGINA, CARDIOGENIC SHOCK, OR A H&H BELOW THE 24 AND 8 THRESHOLDS. THE PATIENT¿S HEMOGLOBIN DROPPED TO 6.9 G/DL BUT IMPROVED TO 9.0 G/DL AFTER THE ADMINISTRATION OF RBC¿S NOT REQUIRING ANY ADDITIONAL BLOOD PRODUCTS AT THAT TIME.¿ THE PATIENT RECOVERED WITH SEQUELAE. TWO DAYS POST IMPELLA CP IMPLANT AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) PLACEMENT INDICATED ACTIVE BLEEDING AND 750 CUBIC CENTIMETER¿S OF BLOOD LOSS WERE REPORTED. THE PATIENT'S HEMOGLOBIN DROPPED TO 6.5 G/DL, PROBABLY DUE TO BLOOD LOSS. AFTER ADMINISTRATION OF RED BLOOD CELL¿S, THE PATIENT'S HEMOGLOBIN WAS FOUND TO BE 10.2 G/DL AND HER BLOOD VOLUME WAS STABLE ENOUGH TO TRANSPORT THE PATIENT TO AN OUTLYING FACILITY. LOQI REPORTED RENAL FAILURE. THE RELATIONSHIP IS NOTED AS "POSSIBLY" RELATED. LOQI STATES: PATIENT DEVELOPED VERY SEVERE METABOLIC ACIDOSIS AND WITH THE CREATININE RISING, THE DECISION WAS MADE TO PLACE THE PATIENT ON CRRT. WITH LOW POTASSIUM THE PATIENT STARTED ON CRRT WITH PHOSPHATE CONTAINING SOLUTIONS. POTASSIUM, CO2, AND MAGNESIUM WERE ALL MAINTAINING WELL ON CRRT. ADDITIONALLY, THE PATIENT PROGNOSIS WAS POOR, AND CARE WAS WITHDRAWN. THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21949204&pc=OZD | 1220648-2025-28025 | 2023-05-15 04:00 | Death | ABIOMED, INC. | 2025-05-02 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Pericardial Effusion | P140003 |  | 1 | Event Description: IT WAS REPORTED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY THAT AN IMPELLA CP PUMP WAS IMPLANTED TO SUPPORT A PATIENT ADMITTED IN WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK. THE PATIENT ENDURED CARDIAC TAMPONADE WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA DEVICE USED. AFTER IMPELLA PLACEMENT, THE PATIENT LATER HAD A PERICARDIOCENTESIS DUE TO SMALL-MODERATE PERICARDIAL EFFUSION WITH CARDIAC TAMPONADE. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT AFTER THE PROCEDURE. AFTER FIVE HOURS OF SUPPORT, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF VENTRICLE PERFORATION HAS BEEN COMPLETED. BASED ON INSUFFICIENT CLINICAL DETAILS, PRODUCT NOT BEING RETURNED, AND UNCLEAR EVIDENCE FROM DATA LOGS WHETHER OR NOT THE POSITION IN VENTRICLE ALARMS OCCURRED AT THE TIME OF THE COMPLICATION, THE ROOT CAUSE OF THE VENTRICLE PERFORATION COULD NOT BE DETERMINED. B3 DATE OF EVENT WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21949309&pc=PYX | 1220648-2025-28027 | 2025-04-12 04:00 | Death | ABIOMED, INC. | 2025-05-02 04:00 | PYX | IMPELLA RP FLEX | Restricted Flow rate | Hemorrhage/Bleeding; Cardiovascular Insufficiency | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH AN IMPELLA RP FLEX REQUIRED CARDIAC PACING. DURING SUPPORT, THE PATIENT BEGAN HAVING INCREASED SUCTION ALARMS AND THE CENTRAL VENOUS PRESSURE (CVP) WAS MUCH LOWER THAN IT HAD BEEN THE PREVIOUS DAY. BLOOD LOSS WITH LINE PLACEMENT AND CONTINUOUS RENAL REPLACEMENT INITIATION COULD HAVE PLAYED A ROLE. GIVEN THE PATIENTS HEMODYNAMICS, THEY ELECTED TO CHANGE IMPELLA FLOW TO P6. HOWEVER, THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. THAT THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME.Event Description: DURING IMPELLA RP FLEX SUPPORT, THE PATIENT EXPERIENCED A SUBSTANTIAL BLEED DURING A LINE PLACEMENT FOR DIALYSIS. THE PATIENT'S SYSTEMIC HEPARIN WAS DISCONTINUED AND THE PATIENT WAS TRANSFUSED WITH A UNIT OF PACKED RED BLOOD CELLS. THE FOLLOWING DAY, THERE WERE REPORTED SUCTION ALARMS AND THE P-LEVEL WAS LOWERED FROM 8 TO 6 AND THE PATIENT WAS TRANSFUSED WITH AN ADDITIONAL UNIT OF PACKED RED BLOOD CELLS AND 2 UNITS OF FRESH FROZEN PLASMA. Manufacturer Narrative: THE IMPELLA PUMP WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW HAS BEEN COMPLETED. BASED ON THE INVESTIGATION, LOW PUMP FLOW WAS REMOVED AS A FAILURE MODE BECAUSE THE LOGS DO NOT SHOW ANY ABNORMAL CHANGE IN MOTOR CURRENT AND LOW PUMP FLOW OCCURRENCE FOR MATCHING P-LEVEL OUTSIDE WHEN FLOW CALCULATION WAS DISABLED AS A RESULT OF THE PLACEMENT SIGNAL ISSUE. INSUFFICIENT PRODUCT WAS RETURNED FOR EVALUATION. THE RETURNED PUMP WAS SEVERED AT THE PUMP PLUG SO ONLY THE CANNULA WITH CATHETER WAS INSPECTED. BIOMATERIAL WAS OBSERVED ON THE IMPELLER. Manufacturer Narrative: B5 REVISED TO INCLUDE ADDITIONAL INFORMATION THAT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28027. H6 REVISED TO REFLECT THE ADDITIONAL INFORMATION THAT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28027.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21949370&pc=MIH | 9612164-2025-02194 | 2025-04-08 04:00 | Death | MEDTRONIC IRELAND | 2025-05-02 04:00 | MIH | VALIANT CAPTIVIA - FF | Material Deformation; Positioning Problem | Disseminated Intravascular Coagulation (DIC); Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Myocardial Infarction; Perforation of Vessels | P100040 |  | 1 | Event Description: VALIANT CAPTIVIA STENT GRAFTS WERE INTENDED TO BE IMPLANTED IN THE ENDOVASCULAR TREATMENT OF A PAU. THE PAU WAS SUSPECTED TO BE INFECTED AND THE PATIENT WAS TREATED WITH ANTIBIOTICS FOR 2 WEEKS. SURGERY PROCEEDED ONCE BLOOD RESULTS HAD NORMALIZED. IT WAS REPORTED THAT DURING THE INDEX PROCEDURE, THE VAMF2222C100TU WAS ATTEMPTED TO BE INSERTED VIA THE RIGHT ACCESS BUT THE DEVICE WOULD NOT ADVANCE. SHOCKWAVE THERAPY WAS THEN USED AND THE DEVICE INSERTION WAS ATTEMPTED AGAIN. THE DEVICE WOULD NOT ADVANCE PAST THE RIGHT COMMON ILIAC ARTERY. TWO OVERLAPPING NON-MDT 8MM STENTS WERE PLACED TO SPAN THE RIGHT COMMON ILIAC ARTERY. DESPITE FURTHER ATTEMPTS, INCLUDING SERIAL DILATION USING SEVERAL DILATORS ,THE LARGEST DILATOR USED WAS A 18 FR SENTRANT SHEATH WHICH HAD A OUTER DIAMETER COMPARABLE WITH THE DEVICE. ANOTHER ATTEMPT TO INSERT THE DEVICE WAS MADE BUT THIS WAS UNSUCCESSFUL. THE PHYSICIAN THEN IDENTIFIED A SMALL GAP BETWEEN THE NOSE CONE AND THE GRAFT COVER, PROMPTING A SWITCH TO THE VAMF2222C150TU WHICH WAS ALSO INSERTED INTO THE PATIENT BUT WOULD NOT ADVANCE. THE PHYSICIAN THEN ATTEMPTED TO FLUSH THE RIGHT ACCESS SHEATH AND NO BLOOD RETURNED. A DIAGNOSTIC ARTERIOGRAM REVEALED EXTRAVASATION DUE TO A PERFORATION IN RIGHT ILIAC ARTERY. A RELIANT BALLOON WAS USED TO PREVENT FURTHER BLOOD LOSS. AN ADDITIONAL STENT WAS THEN PLACED AT THE JUNCTION OF THE RIGHT COMMON ILIAC AND EXTERNAL ILIAC AT THE SITE OF EXTRAVASATION. THE PROXIMAL PORTION OF THIS STENT OVERLAPPED WITH THE PREVIOUSLY DEPLOYED NON MDT STENTS. AT THIS POINT THE PATIENTS BLOOD PRESSURE WAS DROPPING. THE PHYSICIAN DEFLATED THE RELIANT AND ADVANCED IT UP TO THE DISTAL AORTA JUST ABOVE THE BIFURCATION AND REINFLATED. THEN AN OPEN REPAIR OF THE RIGHT ILIAC ARTERY WAS PERFORMED . THE PATIENT LEFT THE OPERATING ROOM IN STABLE CONDITION TO ICU. THE FOLLOWING DAY THE PATIENT DEVELOPED DISSEMINATED INTRAVASCULAR COAGULATION (DIC) AND ELEVATED TROPONINS, INDICATING A LIKELY MYOCARDIAL INFARCTION. THE PATIENT EXPIRED ON THE (B)(6) 2025. IT WAS DETERMINED THAT NEITHER THE SENTRANT SHEATH OR THE RELIANT BALLOON CONTRIBUTED TO THE ADVERSE EVENTS OR THE PATIENT'S DEATH. IT WAS CONFIRMED THE DIFFICULTY ADVANCING THE STENT GRAFTS WAS DUE TO PATIENT ANATOMY. Manufacturer Narrative: FILM ANALYSIS SUMMARY: THE REPORTED POSITIONING DIFFICULTIES AND DEFORMATION IN VIVO COULD NOT BE CONFIRMED ON THE PRE-IMPLANT IMAGES PROVIDED; THEREFORE, THE CAUSE OF THE EVENTS COULD NOT BE DETERMINED. LACK OF FULL PRE-IMPLANT CT¿S DID NOT ALLOW FOR A THOROUGH ASSESSMENT OF THE PRE-IMPLANT ANATOMY. PROCEDURAL ANGIOGRAMS SHOWING ATTEMPTS TO POSITION THE VALIANT CAPTIVIA STENT GRAFTS WERE NOT RECEIVED FOR A THOROUGH ANALYSIS OF THE EVENT. THE REPORTED VESSEL PERFORATION, BLOOD LOSS, DIC , MYOCARDIAL INFRACTION , AND PATIENT DEATH COULD NOT BE ASSESSED. IT IS POSSIBLE PATIENT ANATOMY AS REPORTED MAY HAVE BEEN A FACTOR IN THE POSITIONING DIFFICULTIES BUT THIS COULD NOT CONFIRMED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21949374&pc=OZD | 1220648-2025-28028 | 2025-04-04 04:00 | Death | ABIOMED, INC. | 2025-05-02 04:00 | OZD | IMPELLA 5.5 | Positioning Failure; Restricted Flow rate | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT EXPERIENCED PERSISTENT BLEEDING AND SUCTION ISSUES DURING IMPELLA 5.5 SUPPORT. ROUGHLY TWO DAYS INTO SUPPORT, A GRAFT BLEED WAS NOTED, PROMPTING REPEATED WASHOUTS AND EXPLORATIONS TO MANAGE THE CONDITION. HEPARIN WAS EVENTUALLY PAUSED IN RESPONSE TO THE ONGOING BLEEDS. THEY PATIENT ALSO LATER BECAME SYMPTOMATIC FROM A POSSIBLE GASTROINTESTINAL BLEED. TRANSFUSIONS WERE PERFORMED THROUGHOUT SUPPORT IN RESPONSE TO THE BLEEDS. COINCIDING WITH THE BLEEDS, THE PUMP EXPERIENCED PERSISTENT SUCTION EVENTS. IT WAS NOTED THAT FLOWS HIGHER THAN P-4 WOULD RESULT IN IMMEDIATE SUCTION OR ECTOPY. REPOSITIONING WAS ATTEMPTED, BUT WAS ULTIMATELY UNSUCCESSFUL AND THE DEVICE REMAINED IN SUBOPTIMAL POSITIONING FOR SUPPORT. THE PATIENT PASSED AWAY AFTER 27 DAYS OF SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL, POSITIONING ISSUES, AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. DATA LOGS WERE REVIEWED AND PERSISTENT SUCTION ALARMS WERE PRESENT THROUGHOUT CASE. CLINICAL DETAILS NOTE THAT PATIENT WAS HAVING BLEEDING ISSUES AND REQUIRED MULTIPLE BLOOD PRODUCTS THROUGHOUT SUPPORT. PUMP WAS UNABLE TO FLOW HIGHER THAN P-4 WITHOUT SUCTION ALARMS. THE ROOT CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY PATIENT CONDITION RELATED AS PATIENT REQUIRED MULTIPLE SUCTION ALARMS THROUGHOUT SUPPORT AND PATIENT NEEDED MANY UNITS OF BLOOD AND REQUIRED VOLUME RESUSCITATION. THE ROOT CAUSE OF THE VASCULAR DAMAGE - PERIPHERAL VESSEL COULD NOT BE DEFINITIVELY DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED OF BLEEDING ON SUPPORT. THE ROOT CAUSE OF THE POSITIONING ISSUES COULD NOT BE DEFINITIVELY DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED AND NO POSITIONING ISSUES WERE OBSERVED IN DATA LOGS. B2 REQUIRED INTERVENTION SELECTION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28028. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: WARNINGS & CAUTIONS: CAUTIONS: "PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA." CAUTION: "PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE." SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. "TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION." "TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY." "TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21949394&pc=MIH | 9612164-2025-02195 | 2025-04-08 04:00 | Death | MEDTRONIC IRELAND | 2025-05-02 04:00 | MIH | VALIANT CAPTIVIA - FF | Positioning Problem | Disseminated Intravascular Coagulation (DIC); Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Myocardial Infarction; Perforation of Vessels | P100040 |  | 1 | Event Description: VALIANT CAPTIVIA STENT GRAFTS WERE INTENDED TO BE IMPLANTED IN THE ENDOVASCULAR TREATMENT OF A PAU . THE PAU WAS SUSPECTED TO BE INFECTED AND THE PATIENT WAS TREATED WITH ANTIBIOTICS FOR 2 WEEKS. SURGERY PROCEEDED ONCE BLOOD RESULTS HAD NORMALIZED. IT WAS REPORTED THAT DURING THE INDEX PROCEDURE, THE VAMF2222C100TU WAS ATTEMPTED TO BE INSERTED VIA THE RIGHT ACCESS BUT THE DEVICE WOULD NOT ADVANCE. SHOCKWAVE THERAPY WAS THEN USED AND THE DEVICE INSERTION WAS ATTEMPTED AGAIN. THE DEVICE WOULD NOT ADVANCE PAST THE RIGHT COMMON ILIAC ARTERY. TWO OVERLAPPING NON-MDT 8MM STENTS WERE PLACED TO SPAN THE RIGHT COMMON ILIAC ARTERY. DESPITE FURTHER ATTEMPTS, INCLUDING SERIAL DILATION USING SEVERAL DILATORS ,THE LARGEST DILATOR USED WAS A 18 FR SENTRANT SHEATH WHICH HAD A OUTER DIAMETER COMPARABLE WITH THE DEVICE. ANOTHER ATTEMPT TO INSERT THE DEVICE WAS MADE BUT THIS WAS UNSUCCESSFUL. THE PHYSICIAN THEN IDENTIFIED A SMALL GAP BETWEEN THE NOSE CONE AND THE GRAFT COVER, PROMPTING A SWITCH TO THE VAMF2222C150TU WHICH WAS ALSO INSERTED INTO THE PATIENT BUT WOULD NOT ADVANCE. THE PHYSICIAN THEN ATTEMPTED TO FLUSH THE RIGHT ACCESS SHEATH AND NO BLOOD RETURNED. A DIAGNOSTIC ARTERIOGRAM REVEALED EXTRAVASATION DUE TO A PERFORATION IN RIGHT ILIAC ARTERY. A RELIANT BALLOON WAS USED TO PREVENT FURTHER BLOOD LOSS. AN ADDITIONAL STENT WAS THEN PLACED AT THE JUNCTION OF THE RIGHT COMMON ILIAC AND EXTERNAL ILIAC AT THE SITE OF EXTRAV ASATION. THE PROXIMAL PORTION OF THIS STENT OVERLAPPED WITH THE PREVIOUSLY DEPLOYED NON-MDT STENTS. AT THIS POINT THE PATIENTS BLOOD PRESSURE WAS DROPPING. THE PHYSICIAN DEFLATED THE RELIANT AND ADVANCED IT UP TO THE DISTAL AORTA JUST ABOVE THE BIFURCATION AND REINFLATED. THEN AN OPEN REPAIR OF THE RIGHT ILIAC ARTERY WAS PERFORMED. THE PATIENT LEFT THE OPERATING ROOM IN STABLE CONDITION TO ICU. THE FOLLOWING DAY THE PATIENT DEVELOPED DISSEMINATED INTRAVASCULAR COAGULATION (DIC) AND ELEVATED TROPONINS, INDICATING A LIKELY MYOCARDIAL INFARCTION. THE PATIENT EXPIRED THE NEXT DAY. IT WAS DETERMINED THAT NEITHER THE SENTRANT SHEATH OR THE RELIANT BALLOON CONTRIBUTED TO THE ADVERSE EVENTS OR THE PATIENT'S DEATH. IT WAS CONFIRMED THE DIFFICULTY ADVANCING THE STENT GRAFTS WAS DUE TO PATIENT ANATOMY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: FILM ANALYSIS SUMMARY: THE REPORTED POSITIONING DIFFICULTIES AND DEFORMATION IN VIVO COULD NOT BE CONFIRMED ON THE PRE-IMPLANT IMAGES PROVIDED; THEREFORE, THE CAUSE OF THE EVENTS COULD NOT BE DETERMINED. LACK OF FULL PRE-IMPLANT CT¿S DID NOT ALLOW FOR A THOROUGH ASSESSMENT OF THE PRE-IMPLANT ANATOMY. PROCEDURAL ANGIOGRAMS SHOWING ATTEMPTS TO POSITION THE VALIANT CAPTIVIA STENT GRAFTS WERE NOT RECEIVED FOR A THOROUGH ANALYSIS OF THE EVENT. THE REPORTED VESSEL PERFORATION, BLOOD LOSS, DIC , MYOCARDIAL INFRACTION , AND PATIENT DEATH COULD NOT BE ASSESSED. IT IS POSSIBLE PATIENT ANATOMY AS REPORTED MAY HAVE BEEN A FACTOR IN THE POSITIONING DIFFICULTIES BUT THIS COULD NOT CONFIRMED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21949417&pc=NIM | 2183870-2025-00231 | 2025-04-15 04:00 | Death | COVIDIEN | 2025-05-02 04:00 | NIM | PROTEGE RX CAROTID STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P060001 |  | 1 | Event Description: DURING THE INDEX PROCEDURE A PROTEGE RX WAS USED TO TREAT THE LEFT COMMON ILIAC. PATIENT SUFFERED STROKE DURING SURGERY OF THE INDEX LESION. ADMISSION OF PATIENT TO INTENSIVE CARE. IMPLEMENTATION OF COMFORT CARE. PATIENT DIED. THE INVESTIGATOR ASSESSED THE EVENT AS NOT RELATED TO THE INDEX DEVICE, POSSIBLY RELATED TO INDEX PROCEDURE. THE SPONSOR ASSESSED THE EVENT AS NOT RELATED TO THE INDEX DEVICE, RELATED TO INDEX PROCEDURE. Manufacturer Narrative: UPDATE RECEIVED 26 MAY, 04, 05 JUN 2025: ADVERSE EVENT TERM UPDATED TO - ISCHEMIC STROKE. STROKE CONFIRMED BY SCANNER POST-OPERATIVELY. PATIENT ADMITTED TO NEUROVASCULAR UNIT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21949527&pc=NVZ | 3008973940-2025-03080 | 2025-04-03 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-05-02 04:00 | NVZ | AZURE XT SR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT'S DAUGHTER THAT THE PATIENT HAD PASSED AWAY TEN DAYS POST IMPLANT OF AN IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. IT WAS ALSO NOTED THAT THE SYSTEM HAD BEEN TURNED OFF PRIOR TO DISCHARGE POST IMPLANT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21949536&pc=DTB | 2649622-2025-11953 | 2025-04-03 04:00 | Death | MPRI | 2025-05-02 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENTS DAUGHTER THAT THE PATIENT HAD PASSED AWAY TEN DAYS POST IMPLANT OF AN IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. IT WAS ALSO NOTED THAT THE SYSTEM HAD BEEN TURNED OFF PRIOR TO DISCHARGE POST IMPLANT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21949889&pc=OZD | 1220648-2025-28034 | 2025-04-12 04:00 | Death | ABIOMED, INC. | 2025-05-02 04:00 | OZD | IMPELLA CP | Defective Device | Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT FOR CARDIOGENIC SHOCK. DURING THE PLACEMENT WITH A 14 FRENCH SHEATH, THE SHEATH KINKED. THE TEAM REPLACED THE SHEATH AND PLACED AN IMPELLA CP BUT DURING THE FIRST HOUR OF SUPPORT, THE PATIENT SUFFERED VENTRICULAR TACHYCARDIA AND HAD OPTICAL SIGNAL ISSUES. THE PATIENT AND FAMILY MADE THE DECISION TO NOT ESCALATE CARE AND THE IMPELLA CP AND LUCAS DEVICE WERE EXPLANTED AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF INTRODUCER DAMAGE ¿ MECHANICAL AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. DURING IMPLANT PROCEDURE, ATTEMPTED TO EXCHANGE SHEATH FOR LONG IMPELLA SHEATH IN RFA HOWEVER IMPELLA SHEATH KINKED AND UNABLE TO PASS A WIRE. SHORT SHEATH THEN USED SUCCESSFULLY. THE PUMP AND INTRODUCER WERE NOT RETURNED. THE CAUSE OF THE ISSUE WAS A INTRODUCER SHEATH KINK. THE CAUSE OF THE SHEATH KINK WAS UNKNOWN. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21950332&pc=OZD | 1220648-2025-28039 | 2025-04-04 04:00 | Death | ABIOMED, INC. | 2025-05-02 04:00 | OZD | IMPELLA 5.5 | Positioning Failure | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP AND WAS BEING ESCALTED TO AN IMPELLA 5.5 AND DURING THE IMPLANT, THE PUMP WAS BEING REPOSITIONED AND THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF). THE PATIENT WAS DEFIBRILLATED. ADDITIONALLY, THE PATIENT HAD RECURRENT PERSISTENT VENTRICULAR TACHYCARDIA. IMPELLA 5.5 WAS REPOSITIONED TWICE. THE IMPELLA WAS TOUCHING WHAT APPEARS TO BE THE SEPTUM PER TEAM. THE PUMP WAS REPOSITIONED AGAIN FACING TOWARDS MITRAL VALVE. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THE PATIENT RECEIVED SUBOPTIMAL HEMODYNAMIC SUPPORT AND CONTRIBUTING TO ARRHYTHMIC EVENTS. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA 5.5 AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. ABIOMED'S MEDICAL SAFETY OFFICER REVIEWED THE PATIENTS OUTCOME AND THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA 5.5 AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THE IMPELLA CP FOR THE SAME PATIENT IS REPORTED ON MANUFACTURE REPORT NUMBER 1220648-2025-28037.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21950637&pc=DSQ | 3007042319-2025-01092 | 2025-04-17 04:00 | Death | HEARTWARE, INC. | 2025-05-02 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Shock; Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED FROM COMPLICATIONS AFTER THE VENTRICULAR ASSIST DEVICE (VAD) WAS EXPLANTED FOR A HEART TRANSPLANT. THE CAUSE OF DEATH WAS NOT SPECIFIED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT DIED DUE TO VASOPLEGIC SHOCK. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION SUMMARY. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT DIED DUE TO VASOPLEGIC SHOCK. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21950697&pc=DSQ | 2916596-2025-02953 | 2025-04-14 04:00 | Death | ABBOTT MEDICAL-CARDIOMEMS | 2025-05-02 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO AN UNKNOWN ETIOLOGY. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), (REV. C) IS CURRENTLY AVAILABLE. SECTION 1, ENTITLED ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21950722&pc=DSQ | 2916596-2025-02995 | 2025-04-26 04:00 | Death | THORATEC CORPORATION | 2025-05-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI ORGAN FAILURE. THE DEATH WAS A NATURAL PROGRESSION. THE DEVICE OPERATED AS EXPECTED AND THE DEATH WAS NOT DEVICE RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED MULTI ORGAN FAILURE AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21950902&pc=NPT | 2025587-2025-03317 | 2024-06-27 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-02 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Intracranial Hemorrhage | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS A CEREBRAL HEMORRHAGE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21951016&pc=CBK | 2518422-2025-105105 | 2025-04-19 04:00 | Death | RESPIRONICS, INC. | 2025-05-02 04:00 | CBK | TRILOGY EVO, JAPAN | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT FAILED TO PRESS THE START BUTTON BEFORE GOING TO SLEEP. THE PATIENT WAS FOUND IN THE MORNING IN CARDIOPULMONARY ARREST AND WAS ADMITTED TO THE HOSPITAL. THE DEVICE WAS RETRIEVED FROM THE HOSPITAL. NO EVALUATION OF THE DEVICE HAS BEEN PERFORMED YET. THE CLINICAL ASSESSMENT WAS PERFORMED, AND THE FOLLOWING RATIONALE WAS PROVIDED. THE FAMILY REPORTED THAT PATIENT TESTS HAVE BEEN COMPLETED, AND IT WAS DETERMINED THERE IS NO BRAIN ACTIVITY, SO A DECISION WAS MADE NOT TO PURSUE LIFE-PROLONGING TREATMENT. THEREFORE, DEATH IS PRESUMED BASED ON "4/2025-NO BRAIN ACTIVITY AND THEY DECIDED NOT TO PURSUE LIFESAVING INTERVENTIONS". BASED ON INFORMATION PROVIDED AND AVAILABLE, THIS EVENT IS LIKELY A FORESEEABLE UNINTENTIONAL USE ERROR. THEREFORE, DEVICE CAUSE AND/OR CONTRIBUTION TO THE REPORTED EVENT OF DEATH CAN BE CONFIRMED BASED ON THE EVIDENCE PRESENT. AT THIS TIME, WE ARE UNABLE TO CONFIRM IF ANY MALFUNCTION OF THE DEVICE OCCURRED. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED THIS ISSUE AS A PRODUCT PROBLEM AND SERIOUS INJURY. THE PATIENT FAILED TO START THE DEVICE PRIOR TO GOING TO SLEEP AND CARDIOPULMONARY ARRESTED. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE NOTIFICATION WAS CORRECTED TO REFLECT A SERIOUS INJURY ONLY. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED A PATIENT FAILED TO PRESS THE START BUTTON BEFORE GOING TO SLEEP. THE PATIENT WAS FOUND IN THE MORNING IN CARDIOPULMONARY ARREST AND WAS ADMITTED TO THE HOSPITAL. THE DEVICE WAS RETRIEVED FROM THE HOSPITAL. THE DEVICE EVALUATION FOUND NO EVIDENCE OF PRODUCT MALFUNCTION IN WHICH THE ABILITY OF THE DEVICE TO DELIVER THE PRESCRIBED THERAPY TO THE PUBLISHED SPECIFICATIONS WOULD BE IMPACTED. IN ADDITION, THERE IS NO EVIDENCE THAT THE DEVICE MALFUNCTIONED IN A MANNER THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY IF IT WERE TO RECUR. THE DEVICE WAS TESTED AND PASSED ALL TESTING. Manufacturer Narrative: THIS ISSUE WAS PREVIOUSLY REPORTED AS A SERIOUS INJURY. A CORRECTION IS BEING SENT TO INDICATE A DEATH. THE MANUFACTURER PREVIOUSLY REPORTED A PATIENT FAILED TO PRESS THE START BUTTON BEFORE GOING TO SLEEP. THE PATIENT WAS FOUND IN THE MORNING IN CARDIOPULMONARY ARREST AND WAS ADMITTED TO THE HOSPITAL. THE PATIENT PASSED AWAY WHILE IN THE HOSPITAL. THE DEVICE WAS EVALUATED, AND NO MALFUNCTION OF THE DEVICE WAS OBSERVED. THERE WAS NO EVIDENCE THAT THE DEVICE MALFUNCTIONED IN A MANNER THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY IF IT WERE TO RECUR. THE DEVICE WAS TESTED AND PASSED ALL TESTING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21952217&pc=QZI | 3012520654-2025-00045 | 2025-03-11 04:00 | Death | MEDTRONIC, INC. | 2025-05-02 04:00 | QZI | AFFERA | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemoptysis; Unspecified Tissue Injury | P240013 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN CARDIAC ABLATION PROCEDURE USING THE AFFERA SYSTEM, THE PATIENT INCURRED A LUNG INJURY, AND IT WAS SURMISED THAT IT WAS MOST LIKELY CAUSED BY THE GUIDEWIRE AFTER THE TRANSSEPTAL PUNCTURE. IT WAS FURTHER REPORTED THAT TOWARDS THE END OF THE PROCEDURE THE PATIENT HAD A VIOLENT COUGH AND SEVERE HEMOPTYSIS OCCURRED (COUGHING UP BLOOD). THE PATIENT DEVELOPED A LUNG HEMATOMA, THE HOSPITALIZATION WAS EXTENDED, AND DESPITE AGGRESSIVE MEASURE TO RESUSCITATE THE PATIENT, THE FAMILY ELECTED TO WITHDRAW SUPPORT AFTER FEW DAYS, AND THE PATIENT DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THREE IMAGES AND THREE VIDEOS WERE RETURNED FOR ANALYSIS. ALL THE THREE IMAGES SHOW THE MAPPING IMAGES OF THE HEART WITH LESIONS MARKED. THE VIDEO FILES SHOW THE SAME WITH NO NOTIFICATIONS TRIGGERED DURING THE PROCEDURE. IN CONCLUSION, THE REPORTED DEATH CANNOT BE CONFIRMED THROUGH ANALYSIS OF THE RETURNED CLINICAL DATE. THE PHYSICAL PRODUCT WAS NOT RETURNED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21952271&pc=EOQ | 3021325287-2025-00024 | 2025-04-09 04:00 | Death | NOAH MEDICAL CORP. | 2025-05-02 04:00 | EOQ | GALAXY SYSTEM | Adverse Event Without Identified Device or Use Problem | Bradycardia; Low Oxygen Saturation | K223144 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEVELOPED BRADYCARDIA AND OXYGEN DESATURATION DURING A GALAXY-ASSISTED BIOPSY PROCEDURE FOR A RIGHT UPPER LOBE LESION. THE INJURY OCCURRED DURING THE CASE AND THE ROBOT WAS SAFELY UNDOCKED TO ADDRESS THE PATIENT. MEDICATIONS WERE ADMINISTERED TO RAISE THE PATIENT'S VITALS (NOT PROVIDED BY THE HOSPITAL STAFF) AND THE PATIENT WAS ADMITTED FOR OBSERVATION. TWO WEEKS LATER, NOAH MEDICAL WAS INFORMED THAT THE PATIENT PASSED AWAY 2 DAYS AFTER THE PROCEDURE. THERE WERE NO MALFUNCTIONS REPORTED. THE PHYSICIAN DOES NOT BELIEVE THE GALAXY SYSTEM CONTRIBUTED TO THE BRADYCARDIA AND OXYGEN DESATURATION, ATTRIBUTING IT INSTEAD TO THE BREATH HOLD DURING THE TOOL INSERTION IN THE LESION TILT. IT IS IMPORTANT TO NOTE THAT THE PATIENT HAD PRE-EXISTING VENTRICULAR ISSUES. NOAH MEDICAL WAS INFORMED THAT THE PATIENT EXPIRED 2 DAYS AFTER THE BRONCHOSCOPY PROCEDURE. THE PHYSICIAN ATTRIBUTED THE DEATH TO MULTIORGAN FAILURE. ADDITIONALLY, THE PATIENT HAD RIGHT VENTRICULAR DYSFUNCTION, ELEVATED POTASSIUM LEVELS AND MULTIPLE MYOCARDIAL INFARCTIONS THE NIGHT OF PASSING. ADDITIONAL INFORMATION WAS REQUESTED FROM THE HOSPITAL, A SUPPLEMENTAL WILL BE SUBMITTED IF THE INFORMATION IS ACQUIRED BY NOAH MEDICAL. Manufacturer Narrative: SYSTEM AND BRONCHOSCOPE MANUFACTURING RECORDS WERE REVIEWED AND FOUND NO ISSUES ASSOCIATED WITH THIS CASE. FAILURE ANALYSIS OF THE SCOPE COULD NOT BE PERFORMED AS THE SCOPE WAS DISCARDED. VIDEO LOGS WERE REVIEWED AND NO USE ERRORS WERE IDENTIFIED. AT TIMESTAMP 22:21, A SUDDEN SCOPE DISCONNECTION AND DISENGAGEMENT OF THE CART BRAKES WERE NOTED, WHICH IS LIKELY DUE TO THE NEED TO ADDRESS THE PATIENT'S CONDITION. THE PHYSICIAN DOES NOT BELIEVE THE GALAXY SYSTEM CONTRIBUTED TO THE BRADYCARDIA AND OXYGEN DESATURATION, ATTRIBUTING IT INSTEAD TO THE BREATH HOLD DURING THE TOOL INSERTION IN THE LESION TILT. IT IS IMPORTANT TO NOTE THAT THE PATIENT HAD PRE-EXISTING VENTRICULAR ISSUES. WHILE PERFORMING BREATH HOLDS DURING BRONCHOSCOPY IS NOT INHERENTLY HIGH-RISK, IT MAY PRESENT POTENTIAL RISKS, SUCH AS DECREASED OXYGEN LEVELS, PARTICULARLY IN PATIENTS WITH UNDERLYING CARDIAC CONDITIONS, AS SEEN IN THIS CASE. 10 DAYS AFTER THE EVENT, NOAH MEDICAL WAS INFORMED THAT THE PATIENT EXPIRED 2 DAYS AFTER THE BRONCHOSCOPY PROCEDURE. THE PHYSICIAN ATTRIBUTED THE DEATH TO MULTIORGAN FAILURE. ADDITIONALLY, THE PATIENT HAD RIGHT VENTRICULAR DYSFUNCTION, ELEVATED POTASSIUM LEVELS AND MULTIPLE MYOCARDIAL INFARCTIONS THE NIGHT OF PASSING. THIS CASE IS BEING REPORTED DUE TO THE FOLLOWING CONCLUSIONS: THE PATIENT DEVELOPED BRADYCARDIA AND OXYGEN DESATURATION DURING A GALAXY-ASSISTED BIOPSY PROCEDURE THAT REQUIRED MEDICAL INTERVENTION TO STABILIZE THE VITAL SIGNS AND THE PATIENT WAS ADMITTED. 2 DAYS LATER, THE PATIENT PASSED AWAY. THE PHYSICIAN DID NOT ATTRIBUTE THE BRADYCARDIA AND OXYGEN DESATURATION TO THE GALAXY SYSTEM AND STATES IT IS DUE TO THE BREATH HOLD AND THE PHYSICIAN ATTRIBUTED THE DEATH TO MULTIORGAN FAILURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21952546&pc=OZD | 1220648-2025-28050 | 2025-04-04 04:00 | Death | ABIOMED, INC. | 2025-05-02 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Stroke/CVA; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND THE NEXT DAY WAS ESCALATED TO THE IMPELLA 5.5 DEVICE DUE TO LIMB ISCHEMIA COMPLICATION RELATED TO THE IMPELLA CP. THE IMPELLA 5.5 DEVICE WAS IMPLANTED WITH SOME COMPLICATIONS AND DAYS FOLLOWING THE PATIENT WOULD EXPERIENCE A STROKE AND SUBSEQUENTLY EXPIRE. THE IMPELLA 5.5 (PUMP 2) WAS PLACED AND DURING THE TYING ON OF THE AXILLARY GRAFT, THE PATIENT WENT INTO ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE AND WAS SUCCESSFULLY CARDIOVERTED AFTER THREE SHOCKS. THE SURGEON MADE THE DECISION TO GO ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) FIRST AND THEN INSERT THE IMPELLA 5.5. ACCESS IN LEFT FEMORAL VEIN TO CANNULATE FOR ECMO WAS MADE. SECOND GRAFT ON AXILLARY USED FOR ARTERIAL CANNULATION OF ECMO AND ECMO WAS STARTED. THE IMPELLA CP DEVICE SUPPORT LEVEL WAS LOWERED TO P-2 AND PULLED BACK INTO AORTA. THE PLAN FOR VASCULAR CLOSURE OF RIGHT FEMORAL ARTERY (RFA) WAS AT END OF THE CASE. THE PATIENT WAS NOTED TO HAVE A SIGNIFICANT AMOUNT OF BLEEDING FROM AXILLARY ACCESS SITE AND WAS GIVEN VOLUME AND BLOOD PRODUCTS THROUGHOUT. THE PHYSICIAN PLACED A 23FF AXILLARY SHEATH IN GRAFT WITH GRAFT LOCKS X2. LEFT VENTRICLE WAS ACCESSED WITH AL1 CATHETER AND CROSSED WITH 0.035 STRAIGHT WIRE. THE GUIDEWIRE WAS EXCHANGED FOR 0.018 ABIOMED WIRE AND IMPELLA 5.5 BACKLOADED. THE IMPELLA 5.5 DEVICE WAS STARTED AT P-4 WITH FLOWS OF APPROXIMATELY 2 LPM. ECMO FLOWS WERE 3.4 LPM. THE IMPELLA 5.5 WAS REPOSITIONED USING FLUOROSCOPY AND TRANSESOPHAGEAL CARDIOGRAM GUIDANCE DUE TO FLIPPING UP. THE IMPELLA MEASURED 5.6CM, AND THE PHYSICIAN WAS GOOD WITH THE POSITION. THE GRAFT WAS TRIMMED DOWN AND BLUE SUTURE HUB PLACED AND SECURED. THE RIGHT AXILLARY SITE WAS CLOSED. LARGE DRESSING APPLIED OVER GRAFT THAT WAS ON THE OUTSIDE OF THE BODY. GRAFT WAS SECURED AND INCLUDED IMPELLA AND ECMO CATHETERS TO CHEST. VASCULAR THEN REMOVED THE IMPELLA CP FROM THE RFA AND SITE WAS SURGICALLY CLOSED. AT END OF PROCEDURE, THE PATIENT WAS TRANSFERRED TO THE UNIT ON ECMO AND IMPELLA MCS. APPROXIMATELY THREE DAYS LATER, HOWEVER, THE PATIENT EXPERIENCED A STROKE AND EXPIRED THE NEXT DAY. CARE WAS WITHDRAWN. Manufacturer Narrative: MEDICAL SAFETY REVIEW OF THE EVENT NOTED THE IMPELLA CP IS A SERIOUS INJURY GIVEN THE LIMB ISCHEMIA. THE PATIENT'S DEMISE IS ASSOCIATED WITH THE IMPELLA 5.5 DEVICE. WITH IMPELLA 5.5 RELATED COMPLICATIONS, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THIS IS ONE OF TWO DEVICES ASSOCIATED WITH THE EVENT. REFER TO INITIAL MEDICAL DEVICE REPORT FOR IMPELLA CP SERIAL NUMBER: (B)(6) WITH DATE OF EVENT 04-APR-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMARTASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: POTENTIAL ADVERSE EVENTS: ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21952807&pc=KNT | 3011270181-2025-00013 | 2024-12-10 05:00 | Death | AVANOS MEDICAL INC. | 2025-05-02 04:00 | KNT | CORTRAK 2 ENTERAL ACCESS SYSTEM | Malposition of Device | Cardiac Arrest; Fistula; Pleural Effusion | K240965 |  | 1 | Event Description: IT WAS REPORTED VIA FDA USER FACILITY REPORT MDR NUMBER MW5168383 THE FOLLOWING INFORMATION: PT ADMITTED FOR A FALL AT HOME RESULTING IN A FRACTURE OF THE DISTAL CLAVICLE W MINIMAL DISPLACEMENT. NGT (NASOGASTRIC TUBE)/ CORTRACK WAS ORDERED THAT RESULTED IN 3-MINUTE CARDIAC ARREST. INJURY TO THE RIGHT MAINSTEM AND THROUGH THE LUNG. ABDOMINAL X-RAY SHOWED FEEDING TUBE WAS LOOPED OVER RIGHT HEMITHORAX CONCERNING FOR TRAUMATIC BRONCHOPLEURAL FISTULA AND POSSIBLE EXTENDING INTO THE PLEURAL SPACE. TYPE OF EVENT: DEATH, HEALTH EFFECT: CARDIAC ARREST, FISTULA, PLEURAL EFFUSION. Manufacturer Narrative: THE ACTUAL COMPLAINT PRODUCT WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD IS NOT POSSIBLE AS NO LOT NUMBER WAS PROVIDED; THEREFORE, THE UDI-PI IS UNAVAILABLE. THE LOT NUMBER WAS NOT PROVIDED BY THE CUSTOMER. ALL INFORMATION REASONABLY KNOWN AS OF 01-MAY-2025 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AVANOS MEDICAL, INC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AVANOS MEDICAL, INC. AVANOS MEDICAL, INC. HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AVANOS MEDICAL, INC. COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN AVANOS MEDICAL, INC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21952856&pc=LWS | 2124215-2025-28225 | 2025-04-08 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-02 04:00 | LWS | RESONATE EL ICD VR | Premature Discharge of Battery; Incorrect, Inadequate or Imprecise Result or Readings; Low impedance; Delayed Charge Time | Insufficient Information | P960040/S385 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND ASSOCIATED RIGHT VENTRICULAR (RV) LEAD WAS FOUND DECEASED. A RED ALERT WAS ISSUED IN THE REMOTE MONITORING SYSTEM SHOWING A CODE 1004, INDICATING A SHOCK INTO A SHORTED LEAD HAD OCCURRED (SHOCK IMPEDANCE WAS LOW, OUT-OF-RANGE AT LESS THAN 12.5 OHMS). ALSO, CODE 1007 WAS OBSERVED, INDICATING A CHARGE TIME OUT DUE TO NOT BEING ABLE TO COMPLETE A CHARGE IN OVER 45 SECONDS. FOLLOWING THESE CODES, THE DEVICE TRIGGERED BATTERY CAPACITY DEPLETED STATUS. THE DATA SHOWED THE PATIENT HAD RECEIVED MULTIPLE SHOCKS ON THE DATE OF DEATH. TECHNICAL SERVICES DISCUSSED LIMITED INFORMATION WOULD BE AVAILABLE WITH A PROGRAMMER INTERROGATION AND THE DEVICE SHOULD BE RETURNED FOR LABORATORY ANALYSIS. THE LOCAL SALES REPRESENTATIVE CONFIRMED THE DEVICE AND POSSIBLY THE RV LEAD WOULD BE RETURNED. BOTH PRODUCTS WERE SUBSEQUENTLY RECEIVED AND ARE NOW UNDERGOING LABORATORY ANALYSIS.Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND ASSOCIATED RIGHT VENTRICULAR (RV) LEAD WAS FOUND DECEASED. IT WAS NOTED THAT A RED ALERT WAS ISSUED VIA THE LATITUDE REMOTE MONITORING SYSTEM DUE TO A CODE 1004, INDICATING A SHOCK INTO A SHORTED CONDITION. A CODE 1007 WAS ALSO ISSUED, INDICATING A CHARGE TIME OUT DUE TO THE HIGH VOLTAGE CAPACITORS NOT CHARGING WITHIN THE EXPECTED TIMEFRAME (45 SECONDS). FOLLOWING THESE CODES, THE DEVICE TRIGGERED THE BATTERY CAPACITY DEPLETED STATUS. REVIEW OF DEVICE DATA NOTED THAT THE DEVICE ATTEMPTED TO DELIVER MULTIPLE SHOCKS ON (B)(6) 2025, WHICH WAS REPORTED TO BE THE DATE OF THE PATIENT'S DEATH. BOSTON SCIENTIFIC'S TECHNICAL SERVICES (TS) DISCUSSED THAT LIMITED INFORMATION WOULD BE AVAILABLE WITH A PROGRAMMER INTERROGATION AND RECOMMENDED THAT THE DEVICE BE RETURNED FOR LABORATORY ANALYSIS. THE DEVICE AND LEAD WERE EXPLANTED POST-MORTEM AND RETURNED FOR ANALYSIS. BOTH PRODUCTS WERE SUBSEQUENTLY RECEIVED AND ARE NOW UNDERGOING LABORATORY ANALYSIS. Manufacturer Narrative: THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS. UPON RECEIPT AT BOSTON SCIENTIFIC'S POST MARKET QUALITY ASSURANCE LABORATORY, A THOROUGH EVALUATION OF THE DEVICE WAS PERFORMED. REVIEW OF DEVICE MEMORY CONFIRMED THAT A CODE 1004 OCCURRED ON (B)(6) 2025, INDICATING A SHOCK INTO A SHORTED CONDITION. THIS WAS FOLLOWED BY EIGHT (8) CODE 1007 ALERTS ON THE SAME DATE DUE TO THE CHARGE TIME EXCEEDING THE EXPECTED TIMEFRAME. THE DEVICE TRIGGERED THE BATTERY CAPACITY DEPLETED STATUS DUE TO THE CHARGE TIME FAULTS (RATHER THAN DUE TO BATTERY DEPLETION). HIGH-POWERED VISUAL INSPECTION OF THE DEVICE CASE DID NOT IDENTIFY ANY ARC MARKS, WHICH CAN OCCUR WHEN A SHOCK IS DIRECTED BACK TO THE DEVICE CASE DUE TO A LEAD ISSUE (I.E., INSULATION BREACH). AN X-RAY OF THE DEVICE REVEALED THAT THE INTERNAL HIGH VOLTAGE FUSE WAS DAMAGED (I.E., NO LONGER INTACT). THE DAMAGE TO THE FUSE MOST LIKELY OCCURRED DURING DELIVERY OF THE SHOCK INTO A SHORTED CONDITION THAT RESULTED IN THE CODE 1004. THE DAMAGE TO THE HIGH VOLTAGE FUSE PREVENTED ANY SUBSEQUENT CHARGING OF THE HIGH VOLTAGE CAPACITORS WHICH TRIGGERED THE BATTERY CAPACITY DEPLETED STATUS AND PREVENTED FURTHER DELIVERY OF SHOCK THERAPY. THE DEVICE CASE WAS OPENED FOR FURTHER INSPECTION. NO FOREIGN MATERIAL WAS NOTED IN THE DEVICE ASSEMBLY. HIGH POWERED VISUAL INSPECTION NOTED THE PRESENCE OF ELECTRICAL OVERSTRESS DAMAGE ON THE HIGH VOLTAGE CAPACITOR CONNECTIONS AND IN THE AREA OF THE HIGH VOLTAGE FUSE. THE FUSE WAS NOTED TO HAVE BUBBLING AND WAS CRACKED OPEN, CONSISTENT WITH ELECTRICAL OVERSTRESS DAMAGE. DUE TO THE LACK OF AN ARC MARK ON THE DEVICE CASE AND THERE BEING NO ANOMALIES ON THE ASSOCIATED DEFIBRILLATION LEAD, THE INTERNAL DEVICE DAMAGE WAS DETERMINED TO BE CONSISTENT WITH A SHORT THAT OCCURRED WITHIN THE DEVICE DURING SHOCK DELIVERY. HOWEVER, DESPITE EXHAUSTIVE LABORATORY ANALYSIS, THE SPECIFIC ROOT CAUSE OF THE SHORT CONDITION COULD NOT BE ESTABLISHED DUE TO THE INTERNAL DAMAGE INCURRED BY THE ELECTRICAL SHORT. Manufacturer Narrative: THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21952932&pc=NPT | 2015691-2025-03496 | 2025-03-18 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-02 04:00 | NPT | EDWARDS LIFESCIENCES SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Thrombosis/Thrombus | P140031 |  | 1 | Event Description: IT WAS REPORTED BY THE JAPANESE TAVI REGISTRY THAT THE PATIENT HAD A TRANSFEMORAL TAVR WITH A 20MM SAPIEN 3 ULTRA RESILA VALVE, ON POSTOPERATIVE DAY (POD) 50 A VALVE THROMBOSIS WAS OBSERVED, AND ON POD 59 THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL FOR CEREBRAL INFARCTION AND EXPIRED. ACCORDING TO THE REPORT FROM THE TAVI FACILITY, THE CLOSE TIMING OF THE ADVERSE EVENTS SUGGESTS A POSSIBLE ASSOCIATION THAT CANNOT BE EXCLUDED. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRAPROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ALTHOUGH THE EXACT CAUSE AND SOURCE OF THE STROKE CANNOT BE DETERMINED, IT IS POSSIBLE PATIENT FACTORS NOT PROVIDED MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. PER THE INSTRUCTIONS FOR USE (IFU), VALVE THROMBOSIS IS A POTENTIAL RISK ASSOCIATED WITH THE USE OF THE TRANSCATHETER HEART VALVES (THV). USE OF THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS) IS CONTRAINDICATED IN PATIENTS WHO CANNOT TOLERATE AN ANTICOAGULATION/ANTIPLATELET REGIMEN. IN ADDITION, IT WARNS THAT VALVE RECIPIENTS SHOULD BE MAINTAINED ON ANTICOAGULANT/ANTIPLATELET THERAPY, EXCEPT WHEN CONTRAINDICATED, AS DETERMINED BY THEIR PHYSICIAN. THROMBOSIS IS A RARE AND WELL-RECOGNIZED COMPLICATION ASSOCIATED WITH THE USE OF BIOPROSTHETIC HEART VALVES. VALVE THROMBOSIS IS THE FORMATION OF SIGNIFICANT BLOOD CLOTS FORMING ON THE VALVE. POTENTIAL PATIENT RISK FACTORS SUCH AS ATRIAL FIBRILLATION, SYSTEMIC DISEASE (E.G., SYSTEMIC LUPUS ERYTHEMATOSUS, INFLAMMATION, AND DAMAGE TO VARIOUS BODY TISSUES, INCLUDING JOINTS, SKIN, KIDNEYS, HEART, LUNGS, BLOOD VESSELS, AND BRAIN), THE VALVULAR DISEASE ITSELF, AND REDUCED CARDIAC EJECTION FRACTION CAN CONTRIBUTE TO INCREASED RISK OF THROMBUS/THROMBOSIS. SUB-OPTIMAL ANTICOAGULATION DURING THE PROCEDURE AND UNDERLYING PATIENT CONDITIONS CAN ALSO RESULT IN INCREASED THROMBOGENICITY. THESE PATIENTS ARE ANTICOAGULATED FOR THE PROCEDURE AND INTERVENTIONAL BEST PRACTICES MANDATE METICULOUS PREPARATION, ASPIRATION, AND FLUSHING OF THE DEVICES TO PREVENT AND/OR REMOVE CLOT(S). SHORT-TERM ANTICOAGULATION THERAPY MAY ALSO BE NECESSARY AFTER VALVE REPAIR, ANTICOAGULATION IS PRESCRIBED PER INSTITUTIONAL GUIDELINES. INTRA-PROCEDURAL INTRA-CARDIAC THROMBUS AND POST-PROCEDURE OR LATE VALVE THROMBOSIS ARE COMPLEX PROCESSES TRIGGERED BY THE INTERACTION BETWEEN THE HOST AND THE DEVICE WHICH ARE HIGHLY VARIABLE AMONG PATIENTS. IT IS THE NATURAL TENDENCY OF THE BODY TO FORM A CLOT ON FOREIGN OBJECTS IN THE VASCULAR SPACE AND A DEFINITIVE ROOT CAUSE CANNOT ALWAYS BE CONFIRMED. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR TO CONSIDER ALL PROCEDURAL AND ANATOMICAL FACTORS. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE DEVICE SYSTEM. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY IN CONJUNCTION WITH ECHOCARDIOGRAPHY FOR OPTIMAL VISUALIZATION DURING POSITIONING AND DEPLOYMENT AND CAUTIONS INCLUDE THE MAINTENANCE OF THE PATIENT'S ANTICOAGULATION (ACT STATUS GREATER THAN 250 SECONDS) DURING THE PROCEDURE IN ORDER TO PREVENT THROMBOSIS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THESE ADVERSE EVENTS. ALTHOUGH THE EXACT CAUSE AND SOURCE OF THE THROMBOSIS AND STROKE CANNOT BE DETERMINED, IT IS POSSIBLE PATIENT FACTORS NOT PROVIDED MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21953133&pc=NVY | 2124215-2025-28242 | 2025-04-08 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-02 04:00 | NVY | RELIANCE 4-FRONT | Unexpected Therapeutic Results; Low impedance | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND ASSOCIATED RIGHT VENTRICULAR (RV) LEAD WAS FOUND DECEASED. A RED ALERT WAS ISSUED IN THE REMOTE MONITORING SYSTEM SHOWING A CODE 1004, INDICATING A SHOCK INTO A SHORTED LEAD HAD OCCURRED (SHOCK IMPEDANCE WAS LOW, OUT-OF-RANGE AT LESS THAN 12.5 OHMS). ALSO, CODE 1007 WAS OBSERVED, INDICATING A CHARGE TIME OUT DUE TO NOT BEING ABLE TO COMPLETE A CHARGE IN OVER 45 SECONDS. FOLLOWING THESE CODES, THE DEVICE TRIGGERED BATTERY CAPACITY DEPLETED STATUS. THE DATA SHOWED THE PATIENT HAD RECEIVED MULTIPLE SHOCKS ON THE DATE OF DEATH. TECHNICAL SERVICES DISCUSSED LIMITED INFORMATION WOULD BE AVAILABLE WITH A PROGRAMMER INTERROGATION AND THE DEVICE SHOULD BE RETURNED FOR LABORATORY ANALYSIS. THE LOCAL SALES REPRESENTATIVE CONFIRMED THE DEVICE AND POSSIBLY THE RV LEAD WOULD BE RETURNED. BOTH PRODUCTS WERE SUBSEQUENTLY RECEIVED AND ARE NOW UNDERGOING LABORATORY ANALYSIS.Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND ASSOCIATED RIGHT VENTRICULAR (RV) LEAD WAS FOUND DECEASED. IT WAS NOTED THAT A RED ALERT WAS ISSUED VIA THE LATITUDE REMOTE MONITORING SYSTEM DUE TO A CODE 1004, INDICATING A SHOCK INTO A SHORTED CONDITION. A CODE 1007 WAS ALSO ISSUED, INDICATING A CHARGE TIME OUT DUE TO THE HIGH VOLTAGE CAPACITORS NOT CHARGING WITHIN THE EXPECTED TIMEFRAME (45 SECONDS). FOLLOWING THESE CODES, THE DEVICE TRIGGERED THE BATTERY CAPACITY DEPLETED STATUS. REVIEW OF DEVICE DATA NOTED THAT THE DEVICE ATTEMPTED TO DELIVER MULTIPLE SHOCKS ON (B)(6) 2025, WHICH WAS REPORTED TO BE THE DATE OF THE PATIENT'S DEATH. BOSTON SCIENTIFIC'S TECHNICAL SERVICES (TS) DISCUSSED THAT LIMITED INFORMATION WOULD BE AVAILABLE WITH A PROGRAMMER INTERROGATION AND RECOMMENDED THAT THE DEVICE BE RETURNED FOR LABORATORY ANALYSIS. THE DEVICE AND LEAD WERE EXPLANTED POST-MORTEM AND RETURNED FOR ANALYSIS. Manufacturer Narrative: E1: PATIENT CITY: (B)(6). PATIENT COUNTRY: (B)(6). H11: SINCE NO LOT NUMBER IS AVAILABLE, A DETAILED INVESTIGATION, TESTS ON REFERENCE SAMPLES OR BATCH REVIEW CANNOT BE CONDUCTED. THEREFORE, THIS COMPLAINT WILL BE CLOSED, THIS ISSUE WILL BE MONITORED THROUGH PMS PRODUCT TRENDS AND MALFUNCTION. IF ANY TRENDS PICKED UP, THIS WILL FLAG ON THE TRIPS AND ALERTS ACCORDING TO THE OMQR PROCEDURE. Manufacturer Narrative: THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS. A THOROUGH EVALUATION OF THE RV DEFIBRILLATION LEAD WAS PERFORMED. THE COMPLETE LEAD WAS RETURNED. SURFACE CUTS WERE NOTED ON THE INSULATION AND SUTURE SLEEVE, CONSISTENT WITH HAVING BEEN INDUCED DURING THE EXPLANT PROCEDURE (THESE FINDINGS WERE COSMETIC IN NATURE AND DID NOT IMPACT THE LEAD INTEGRITY). SETSCREW MARKS WERE CONFIRMED TO BE IN THE APPROPRIATE LOCATION ON THE TERMINAL PIN. TESTING WAS COMPLETED TO ASSESS LEAD ELECTRICAL PERFORMANCE AND INSULATION INTEGRITY. MEASUREMENTS THROUGHOUT THESE TESTS WERE WITHIN NORMAL LIMITS, THEREBY CONFIRMING THAT THE LEAD WAS ELECTRICALLY CONTINUOUS (I.E., NOT FRACTURED) AND THE OUTER AND INNER INSULATION WERE INTACT, INDICATING NO ELECTRICAL SHORTS BETWEEN CONDUCTORS. AN X-RAY OF THE LEAD WAS PERFORMED, WHICH REVEALED NO ANOMALIES. AS A RESULT, LABORATORY TESTING WAS UNABLE TO IDENTIFY ANY LEAD CHARACTERISTICS THAT WOULD HAVE CAUSED OR CONTRIBUTED TO THE SHORT CIRCUIT CONDITION AND INTERNAL DEVICE DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21953560&pc=OZD | 1220648-2025-28054 | 2025-04-05 04:00 | Death | ABIOMED, INC. | 2025-05-02 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED A IMPELLA CP FOR THE SUPPORT OF A PATIENT ADMITTED IN WITH ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK. THE DAY AFTER IMPLANT THE TEAM OBSERVED LEFT VENTRICLE THROMBUS AND 2 DAYS LATER A THROMBUS IN THE ATRIUM. THE PUMP REMAINED ON DESPITE THE THROMBUS. AFTER 6 DAYS OF SUPPORT THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION, SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21953757&pc=LYJ | 1644487-2025-10084 | 2024-07-25 04:00 | Death | HOUSTON | 2025-05-02 04:00 | LYJ | PULSE GEN MODEL 106 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. NO FURTHER DETAILS WERE PROVIDED. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 CFR PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR "MALFUNCTIONS". THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21953927&pc=BZD | 2518422-2025-105124 | 2022-02-24 05:00 | Death | RESPIRONICS, INC. | 2025-05-02 04:00 | BZD | DREAMSTATION CPAP PRO | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION A USER OF A DREAMSTATION CPAP PRO PASSED AWAY THREE YEARS AGO. THE PATIENT'S SISTER STATES THEY ARE RECEIVING LETTERS REGARDING THE DEVICE AND THE PATIENT WANTED THE LETTERS TO STOP BEFORE PASSING AWAY. THE PATIENT'S SISTER COULD NOT RECALL WHAT HAPPENED WITH THE RECALLED DEVICE OR THE REPLACEMENT DEVICE. IT WAS NOT RECEIVED BY THE MANUFACTURER. THE CAUSE OF DEATH WAS NOT PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21954148&pc=DSQ | 2916596-2025-02580 | 2025-04-09 04:00 | Death | THORATEC CORPORATION | 2025-05-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD RESPIRATORY ILLNESS AND PASSED AWAY ON (B)(6) 2025. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. NO FURTHER INFORMATION WAS AVAILABLE. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED AND FUNCTIONED AS INTENDED. Manufacturer Narrative: D4 - UDI: CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED RESPIRATORY ILLNESS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS INVESTIGATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENTS; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A: SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. SECTION D4: DEVICE EXPIRATION DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21954149&pc=LYJ | 1644487-2025-10086 | 2025-04-17 04:00 | Death | HOUSTON | 2025-05-02 04:00 | LYJ | GENERATOR MODEL 1000-D | Adverse Event Without Identified Device or Use Problem | Status Epilepticus | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY DUE TO STATUS EPILEPTICUS. THE PHYSICIAN DID NOT BELIEVE THAT THIS DEATH WAS RELATED TO VNS. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE.Event Description: A DEATH FOLLOW-UP FORM WAS SENT TO THE PATENT'S MANAGING PHYSICIAN. THE COMPLETED FORM WAS FILLED OUT AND REVEALED THE FOLLOWING SURROUNDING THE EVENT: THE PHYSICIAN RE-ITERATES THAT THE PATIENT PASSED AWAY FROM STATUS EPILEPTICUS. THE PATIENT'S RESPONSE TO VNS THERAPY WAS A SEIZURE REDUCTION/SEIZURE FREE. BOTH THE GENERATOR AND LEAD WERE EXPLANTED AFTER DEATH AND HAVE NOT BEEN RECEIVED BY THE MANUFACTURER TO DATE. THE RELATIONSHIP OF DEATH TO THE VNS WAS DETERMINED TO BE NOT RELATED. WHEN THE PATIENT WAS FOUND DEAD, THEY WERE NOTED TO BE LAYING ON THE FLOOR WITH THEIR HEAD [ILLEGIBLE]. THE PATIENT DID HAVE A HISTORY OF NOCTURNAL SEIZURES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21954336&pc=NIQ | 2024168-2025-04610 | 2024-01-01 05:00 | Death | ABBOTT VASCULAR | 2025-05-02 04:00 | NIQ | XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IT WAS REPORTED IN A RESEARCH ARTICLE THAT THE EXPLORE (EVALUATING XIENCE AND LEFT VENTRICULAR FUNCTION IN PCI ON OCCLUSIONS AFTER STEMI) TRIAL WAS THE FIRST AND ONLY RANDOMIZED TRIAL INVESTIGATING CHRONIC TOTAL OCCLUSION (CTO) PERCUTANEOUS CORONARY INTERVENTION (PCI) EARLY AFTER PRIMARY PCI FOR ST-SEGMENT¿ ELEVATION MYOCARDIAL INFARCTION, COMPARED WITH MEDICAL THERAPY FOR THE CTO. IN THIS STUDY, FOLLOW UP WAS PERFORMED AT 10 YEARS TO INVESTIGATE LONG-TERM SAFETY AND CLINICAL IMPACT OF CTO PCI AFTER ST-SEGMENT¿ ELEVATION MYOCARDIAL INFARCTION, COMPARED WITH NO-CTO PCI. IN EXPLORE, 302 PATIENTS POST-ST- SEGMENT¿ ELEVATION MYOCARDIAL INFARCTION WITH CONCURRENT CTO WERE RANDOMIZED TO CTO PCI WITHIN 1 WEEK OR NO-CTO PCI. EXTENDED CLINICAL FOLLOW-UP WAS PERFORMED FOR THE PRIMARY END POINT OF MAJOR ADVERSE CARDIAC EVENTS, CONSISTING OF CARDIOVASCULAR DEATH, CORONARY ARTERY BYPASS GRAFTING, OR MYOCARDIAL INFARCTION. SECONDARY END POINTS INCLUDED ALL-CAUSE DEATH, ANGINA, AND DYSPNEA. MEDIAN FOLLOW-UP WAS 10 YEARS. ADDITIONAL INFORMATION CAN BE FOUND IN THE ARTICLE "TEN-YEAR OUTCOME OF RECANALIZATION OR MEDICAL THERAPY FOR CONCOMITANT CHRONIC TOTAL OCCLUSION AFTER MYOCARDIAL INFARCTION". Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER A SEPARATE MEDWATCH REPORT. ATTACHMENT ARTICLE TITLED, "TEN-YEAR OUTCOME OF RECANALIZATION OR MEDICAL THERAPY FOR CONCOMITANT CHRONIC TOTAL OCCLUSION AFTER MYOCARDIAL INFARCTION". B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. C4: TREATMENT/THERAPY START DATE IS ESTIMATED. D6: DATE OF IMPLANT IS ESTIMATED. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT OF DEATH AND THE RELATIONSHIP TO PRODUCT IF ANY, CANNOT BE DETERMINED. THE INFORMATION RECEIVED IN THE COMPLAINT WAS OBTAINED THROUGH A PROACTIVE LITERATURE SEARCH. SPECIFIC MODEL AND LOT NUMBERS WERE NOT AVAILABLE AND CAUSALITY BETWEEN THE DOCUMENTED DEVICE USAGE AND THE PATIENT EFFECTS COULD NOT BE ESTABLISHED. FACTORS THAT MAY CONTRIBUTE TO THE PERFORMANCE OUTCOMES LISTED IN THE ARTICLE INCLUDE, BUT ARE NOT LIMITED TO, DEVICE TO PATIENT INTERACTION, PATIENT ANATOMICAL CONDITIONS, DEVICE TO USER INTERACTION AND MANUFACTURING OR MATERIAL ISSUES. DUE TO THE LIMITED INFORMATION AVAILABLE, THE EXACT CAUSE CANNOT BE DETERMINED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21954549&pc=BZD | 2518422-2025-105132 | 2025-05-01 04:00 | Death | RESPIRONICS, INC. | 2025-05-02 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION VIA EMAIL TO THE NHS (NATIONAL HEALTH SERVICE) (SBS.APINVOICING@NHS.NET) ON 06/30/2025, 07/03/2025, 07/09/2025. NO RESPONSE WAS RECEIVED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21955177&pc=BZD | 2518422-2025-105134 | 2025-02-26 05:00 | Death | RESPIRONICS, INC. | 2025-05-02 04:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Insufficient Information | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING USER OF UNKNOWN DEVICE PASSED AWAY. THE PATIENT'S WIFE REPORTED THEY NEED TO RETURN THE PATIENT'S REPLACEMENT DEVICE SINCE THE PATIENT HAD PASSED AWAY BEFORE HE RECEIVED THE REPLACEMENT. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THIS REPORT WILL CAPTURE THE DEATH OF THE PATIENT ON THE DEVICE THAT WAS RECALLED SINCE THE PATIENT DID NOT USE THE REPLACEMENT DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING USER OF UNKNOWN DEVICE PASSED AWAY. THE PATIENT'S WIFE REPORTED THEY NEED TO RETURN THE PATIENT'S REPLACEMENT DEVICE SINCE THE PATIENT HAD PASSED AWAY BEFORE HE RECEIVED THE REPLACEMENT. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THIS REPORT WILL CAPTURE THE DEATH OF THE PATIENT ON THE DEVICE THAT WAS RECALLED SINCE THE PATIENT DID NOT USE THE REPLACEMENT DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DEVICE WAS IN GOOD CONDITION AND PASSED EVALUATION. THE CUSTOMER COMPLAINT WAS NOT ADDRESSED. THE DEVICE WAS SCRAPPED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21955188&pc=MNS | 2518422-2025-105135 | 2025-05-01 04:00 | Death | RESPIRONICS, INC. | 2025-05-02 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION VIA EMAIL TO THE NHS (NATIONAL HEALTH SERVICE) (SBS.APINVOICING@NHS.NET) ON 06/30/2025, 07/03/2025, 07/09/2025. NO RESPONSE WAS RECEIVED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21955399&pc=QEW | 3005168196-2025-00152 | 2025-01-12 05:00 | Death | PENUMBRA, INC. | 2025-05-02 04:00 | QEW | INDIGO SYSTEM FLASH ASPIRATION CATHETER | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism | K240030 |  | 1 | Event Description: THE PATIENT WAS UNDERGOING A THROMBECTOMY PROCEDURE IN THE LEFT ILIAC AND FEMORAL VEIN FOR A DEEP VEIN THROMBOSIS (DVT) USING AN INDIGO SYSTEM LIGHTNING FLASH ASPIRATION TUBING (FLASH TUBING), AN INDIGO SYSTEM FLASH ASPIRATION CATHETER (FLASH CATHETER), A BALLOON (10X40), GUIDEWIRES, A MICROSHEATH, AND SHEATHS (6FR AND 16FR) IN A PRONE POSITION VIA THE LEFT LOWER EXTREMITY. DURING THE PROCEDURE, AN INITIAL VENOGRAM WAS TAKEN WHICH SHOWED EXTENSIVE CLOT IN THE FEMORAL, COMMON FEMORAL, AND ILIAC VENOUS SYSTEM. THE 6FR SHEATH WAS UPSIZED TO A 16FR SHEATH. SEVERAL PASSES WERE MADE IN THE TARGET LOCATION USING THE FLASH CATHETER. A VENOGRAM SHOWED SOME RESIDUAL NARROWING OF THE COMMON FEMORAL AND FEMORAL VEIN SO AN ANGIOPLASTY WAS PERFORMED WITH A BALLOON CATHETER. A POST-INTERVENTION VENOGRAM SHOWED NO RESIDUAL STENOSIS. THE VEIN WAS WIDELY PATENT. THE FLASH CATHETER AND SHEATH WERE THEN REMOVED FROM THE PATIENT. THE PROCEDURE WAS COMPLETED AND THE PATIENT WAS PLACED IN A SUPINE POSITION ON THE STRETCHER. IT WAS REPORTED THAT THE PATIENT WAS GETTING READY TO BE EXTUBATED AND BEGAN TO EXPERIENCE OXYGEN DESATURATION AND TACHYCARDIA, WHICH BECAME PROGRESSIVELY WORSE. THE PATIENT THEN BECAME HYPOTENSIVE AND BRADYCARDIC. A CODE WAS CALLED, MULTIPLE ROUNDS OF MEDICATIONS WERE GIVEN, AND 3 ROUNDS OF CARDIOPULMONARY RESUSCITATION (CPR) WERE PERFORMED. TWO UNITS OF PACKED RED BLOOD CELLS (PRBCS) WERE TRANSFUSED. THE PATIENT WAS ABLE TO REGAIN HER BLOOD PRESSURE. THE PHYSICIAN SUSPECTED THAT THE PATIENT HAD A PULMONARY EMBOLISM, SO AN INTERVENTIONAL RADIOLOGIST WAS CONSULTED. ONCE THE PATIENT WAS STABLE, SHE WAS MOVED BACK ONTO THE TABLE. THE PATIENT UNDERWENT A THROMBECTOMY PROCEDURE IN BILATERAL PULMONARY ARTERIES FOR PULMONARY EMBOLISM USING NON-PENUMBRA CATHETERS, A NON-PENUMBRA THROMBECTOMY SYSTEM, AND A SHEATH (6FR) VIA THE RIGHT GROIN. POST-INTERVENTION ANGIOGRAM DEMONSTRATED SCATTERED BILATERAL SEGMENTAL AND SUBSEGMENTAL EMBOLI. THE CATHETER AND SHEATH WERE REMOVED, AND HEMOSTASIS WAS ACHIEVED. THE PROCEDURE WAS COMPLETED AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). IT WAS REPORTED THAT THE PATIENT'S CONDITION CONTINUED TO DETERIORATE WITH MAXIMUM PRESSOR SUPPORT, AND THE PATIENT WAS SHOCKED TWICE WITH NO IMPROVEMENT IN STATUS. THE FAMILY DECIDED TO WITHDRAW CARE, AND THE PATIENT EXPIRED ON (B)(6) 2025. THE CEC CHAIR ADJUDICATED SHOCK/HEMODYNAMIC INSTABILITY DUE TO PULMONARY EMBOLISM TO BE A SERIOUS ADVERSE EVENT WITH A PROBABLE RELATIONSHIP TO THE INDEX PROCEDURE AND A POSSIBLE RELATIONSHIP TO THE INDIGO ASPIRATION SYSTEM. Manufacturer Narrative: THE PRODUCT WAS NOT RETURNED FOR EVALUATION. THEREFORE, A PHYSICAL DEVICE INVESTIGATION WAS UNABLE TO BE PERFORMED. FROM THE INFORMATION PROVIDED, THERE IS NO INDICATION THAT THERE WAS ANY DEVICE MALFUNCTION, NONCONFORMANCE, OR MISUSE THAT CONTRIBUTED TO THE REPORTED EVENT. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21955569&pc=BZD | 2518422-2025-105136 | 2024-03-29 04:00 | Death | RESPIRONICS, INC. | 2025-05-02 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K113068 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE USER OF A REMSTAR PRO C-FLEX+ PASSED AWAY. THE CAUSE OF THE DEATH WAS NOT PROVIDED. THE DEVICE WAS RETURNED TO THE MANUFACTURER IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. DURING THE EVALUATION OF THE DEVICE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. IF ADDITIONAL INFORMATION IS AVAILABLE A FOLLOW-UP REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21955785&pc=QLO | 3014015528-2025-00018 | 2024-04-07 04:00 | Death | EXTHERA MEDICAL CORPORATION | 2025-05-02 04:00 | QLO | SERAPH 100 | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: PER THE REPORTER, THE PATIENT STARTED "BEING FILTERED" IN (B)(6) 2024 IN (B)(6). WHEN SHE ARRIVED IN (B)(6) SHE WAS "VERY SICK" AND NEEDED TO BE STABILIZED AND MONITORED OVERNIGHT IN THE (B)(6) CLINIC PRIOR TO TREATING. THE PATIENT DIED (B)(6) 2024. THE DEVICE WAS BEING USED OUTSIDE OF THE UNITED STATES IN (B)(6). THE PATIENT WAS BEING TREATED BY DR. (B)(6) AND DR. (B)(6). EXTHERA HAS REACHED OUT TO BOTH OF THESE PHYSICIANS AS WELL AS THE FACILITY IN (B)(6). THE FACILITY HAS STATED IT WAS NOT AWARE OF ANY ADVERSE EVENTS. THE PHYSICIANS HAVE NOT RESPONDED. Manufacturer Narrative: THE MANUFACTURER FOLLOWED UP WITH THE FACILITY IN (B)(6) WHERE THE CLAIMED ADVERSE EVENT TOOK PLACE AND WITH THE TREATING PHYSICIAN. THE FACILITY HAS STATED THAT IT WAS NOT AWARE OF ANY ADVERSE EVENTS. THE PHYSICIAN HAS NOT RESPONDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21940598&pc=MIH | 2017233-2025-06125 | 2025-03-27 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-05-01 04:00 | MIH | GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS | Insufficient Information | Ischemia | P210032 |  | 1 | Event Description: IT WAS REPORTED TO GORE THAT ON (B)(6) 2025, THIS 84-YEAR-OLD MALE PATIENT UNDERWENT ENDOVASCULAR TREATMENT AS AN EMERGENT PROCEDURE FOR A COMMON ILIAC ARTERY ANEURYSM WITH UNSPECIFIED GORE® EXCLUDER® AAA ENDOPROSTHESES. VESSEL ACCESS WAS PERFORMED VIA THE PERCUTANEOUS ROUTE. THE GORE® DEVICES WERE SUCCESSFULLY NAVIGATED TO THE INTENDED LOCATION AND SUCCESSFULLY DEPLOYED. APART FROM ONE ADDITIONAL BUT UNSPECIFIED PROCEDURE NO OTHER INFORMATION WAS REPORTED. ON (B)(6) 2025, THE PATIENT REPORTEDLY HAD A REINTERVENTION FOR AN UNKNOWN REASON. DURING THE REINTERVENTION, IT APPEARS THAT THE GORE® EXCLUDER® TRUNK-IPSILATERAL LEG COMPONENT IMPLANTED PREVIOUSLY WAS PRESUMABLY EXTENDED WITH A GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS AORTIC EXTENDER. IT WAS ALSO REPORTED THAT ON THE SAME DAY, THE PATIENT WAS DIAGNOSED WITH BOWEL ISCHEMIA. THE EVENT WAS REPORTED TO HAVE BEEN FATAL AND THE PATIENT WAS NOTED AS DECEASED AS OF (B)(6) 2025. Manufacturer Narrative: H6, INVESTIGATION FINDINGS, CODE C19 - THE REVIEW OF THE MANUFACTURING RECORDS VERIFIED THAT THE LOT INVOLVED IN THIS EVENT MET ALL PRE-RELEASE SPECIFICATIONS. THE DEVICE REMAINS IMPLANTED. THEREFORE, AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED. GORE HAS REACHED OUT TO THE PHYSICIAN FOR ADDITIONAL INFORMATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21941063&pc=PNJ | 9612164-2025-02159 | 2025-04-04 04:00 | Death | MEDTRONIC IRELAND | 2025-05-01 04:00 | PNJ | MICRA AV | Migration or Expulsion of Device; Appropriate Term/Code Not Available | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED BY A PATIENT REPRESENTATIVE THAT A PATIENT EXPRESSED FEELING LIKE "SOMETHING MOVED, SOMETHING IS WRONG WITH THE PACER" BEFORE PASSING AWAY. THE DEVICE INVOLVED WAS A LEADLESS IMPLANTABLE PULSE GENERATOR. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21941117&pc=PNJ | 9612164-2025-02161 | 2025-04-11 04:00 | Death | MEDTRONIC IRELAND | 2025-05-01 04:00 | PNJ | MICRA | Adverse Event Without Identified Device or Use Problem; High Capture Threshold | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING IMPLANT THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) EXHIBITED HIGH THRESHOLDS IN VARIOUS POSITIONS. THE LEADLESS IPG WAS NOT USED AND REPLACED WITH A TRANSVENOUS SYSTEM. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. DURING IMPLANT THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) EXHIBITED HIGH THRESHOLDS IN VARIOUS POSITIONS. THE LEADLESS IPG WAS NOT USED AND REPLACED WITH A SINGLE CHAMBER TRANSVENOUS PACING SYSTEM. FOLLOWING THIS THE PATIENT DIED. Manufacturer Narrative: CORRECTION: B1, B2, B5, H6, H11. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC1VR01-DELSYS] (SERIAL: (B)(6)). D9: NO MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21941246&pc=DYB | 3007284313-2025-03933 | 2025-04-15 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-05-01 04:00 | DYB | GORE® DRYSEAL FLEX INTRODUCER SHEATH | Patient Device Interaction Problem | Perforation of Vessels | K160254 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT OF AN ABDOMINAL AORTIC ANEURYSM. IT WAS PLANNED TO IMPLANT NON-GORE STENTGRAFTS (AFX, MAIN AND CUFF) AND A GORE® EXCLUDER® AAA ENDOPROSTHESIS (AORTA EXTENDER ENDOPROSTHESIS) TO RELINE THE JUNCTION BETWEEN THE NON-GORE STENTGRAFTS. A NON-GORE STENTGRAFT (AFX, MAIN DEVICE) WAS ATTEMPTED TO DELIVER FOLLOWING IFU, BUT IT WAS CAUGHT INSIDE THE ANEURYSM AND COULD NOT BE ADVANCED. THE SHAFT OF THE MAIN DEVICE WAS DEFORMED. THEREFORE, THIS DEVICE WAS CHANGED TO A SECOND NON-GORE STENTGRAFT (AFX, MAIN DEVICE). THE SECOND MAIN DEVICE WAS ATTEMPTED TO DELIVER, BUT IT WAS DIFFICULT. THEN, THE SECOND MAIN DEVICE WAS TRIED TO DELIVER USING A 20FR GORE® DRYSEAL FLEX INTRODUCER SHEATH. HOWEVER, THE MAIN DEVICE WAS NOT ABLE TO BE DELIVERED OVER THE ANEURYSM. DURING THE DELIVERY, THE BLOOD PRESSURE DECREASED DUE TO AN ANEURYSM RUPTURE. A BALLOON OCCLUSION WAS PERFORMED AND A RESCUE WAS ATTEMPTED BY OPEN SURGERY, BUT IT WAS NOT SUCCESSFUL AND THE PATIENT DIED.Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT OF AN ABDOMINAL AORTIC ANEURYSM. IT WAS PLANNED TO IMPLANT NON-GORE STENTGRAFTS (AFX, MAIN AND CUFF) AND A GORE® EXCLUDER® AAA ENDOPROSTHESIS (AORTA EXTENDER ENDOPROSTHESIS) TO RELINE THE JUNCTION BETWEEN THE NON-GORE STENTGRAFTS. A 20FR GORE DRYSEAL FLEX INTRODUCER SHEATH WAS USED VIA LEFT ACCESS TO DELIVER THE MAIN DEVICE ACCORDING TO PROCEDURE, BUT IT COULD NOT BE ADVANCED OVER THE ANEURYSM. AS THE LEADING TIP OF THE MAIN DEVICE WAS FOUND TO BE BENT, IT WAS REPLACED WITH ANOTHER MAIN DEVICE (SECOND MAIN DEVICE). AFTER INSERTING A NON-GORE SHEATH (AFX SHEATH) INTO THE 20FR GORE DRYSEAL FLEX INTRODUCER SHEATH, DELIVERY OF THE SECOND MAIN DEVICE WAS ATTEMPTED, BUT IT WAS DIFFICULT AND COULD NOT BE ADVANCED TO THE INTENDED POSITION. UPON REMOVAL OF THE SECOND MAIN DEVICE ALONG WITH THE NON-GORE SHEATH, DAMAGE TO THE SHEATH AND BENDING OF THE TIP OF THE MAIN DEVICE WERE CONFIRMED. AS LEFT ACCESS WAS DEEMED DIFFICULT, THE APPROACH WAS SWITCHED TO RIGHT ACCESS. WHILE PREPARING TO ATTEMPT DELIVERY AGAIN, THE PATIENT¿S BLOOD PRESSURE DROPPED. A CONTRAST-ENHANCED CT REVEALED RUPTURE OF THE ABDOMINAL AORTIC ANEURYSM. EVAR WAS ABANDONED, AND OPEN SURGICAL GRAFT REPLACEMENT WAS PLANNED; HOWEVER, DUE TO UNSTABLE VITAL SIGNS, OPEN SURGERY WAS DEEMED UNFEASIBLE AND THE PROCEDURE WAS TERMINATED. THE PATIENT SUBSEQUENTLY DIED. THE CAUSE OF DEATH IS CONSIDERED TO BE HEMORRHAGIC SHOCK DUE TO RUPTURE OF THE ABDOMINAL AORTIC ANEURYSM. Manufacturer Narrative: W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: EMDR SECTION H6, C21 AND D16 ARE REMOVED AND C19, D12 AND D15 ARE ADDED TO REFLECT RESULTS OF INVESTIGATION. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: ACCORDING TO THE GORE® DRYSEAL FLEX INTRODUCER SHEATH INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO: VASCULAR TRAUMA (I.E., DISSECTION, RUPTURE, PERFORATION, TEAR, ETC.), DEATH. Manufacturer Narrative: UPDATED A2, B5, B7.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21941407&pc=LWS | 2649622-2025-11751 | 2025-04-11 04:00 | Death | MPRI | 2025-05-01 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT PASSED AWAY TWO WEEKS AFTER THE RIGHT VENTRICULAR (RV) LEAD WAS IMPLANTED. THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21941500&pc=LPB | 2124215-2025-27832 | 2025-04-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-01 04:00 | LPB | BLAZER II TEMPERATURE ABLATION CATHETER | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Coma |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. DURING AN ABLATION PROCEDURE TO TREAT A PAROXYSMAL ATRIAL FIBRILLATION (PAF) AND A RIGHT ATRIAL FLUTTER (AFL), A FARAWAVE CATHETER AND A BLAZER CATHETER WERE SELECTED FOR USE. THE PROCEDURE WAS PERFORMED CONVENTIONALLY ON TUESDAY APRIL 1ST; THE PATIENT WAS UNDER GENERAL ANESTHESIA AND ACT WAS GREATER THAN 300 SECONDS. THE PHYSICIAN ABLATED ALL 4 VEINS, NO PULMONARY VEIN ISOLATION+. NO COMPLICATIONS OCCURRED DURING THE PROCEDURE. HOWEVER, POST PROCEDURE, THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT ON THURSDAY, APRIL 3RD. HE WAS IN A COMA ON FRIDAY, APRIL 4TH, AND DIED ON SATURDAY, APRIL 5TH FROM A CEREBRAL HEMORRHAGE (DIAGNOSED WITH AN MRI). THE PATIENT WAS YOUNG (59 YEARS OLD), HAD NO COMORBID FACTORS, AND DIED 4 DAYS AFTER THE OPERATION. NO AUTOPSY WAS PERFORMED AND NO DIRECT CAUSE TO THE PROCEDURE NOR DEVICE CAN BE ESTABLISHED. DEVICES ARE NOT EXPECTED TO BE RETURNED AS THEY WERE DISPOSED. MOREOVER, IT WAS INFORMED THAT THE PATIENT WAS NOT ON DOAC OR COUMADIN BEFORE THE PROCEDURE. ANTICOAGULATION WAS INTERRUPTED 12 HOURS BEFORE THE PROCEDURE. THE PATIENT DID NOT HAVE A HISTORY OF HYPERTENSION. ACCORDING TO THE DOCTOR, THE BRAIN MRI DID NOT SHOW A STROKE, BUT ONLY A CEREBRAL HEMORRHAGE; IT WAS A PURE HEMORRHAGIC STROKE. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE BECAUSE THE DEVICE WAS DISCARDED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21941502&pc=QZI | 2124215-2025-27831 | 2025-04-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-01 04:00 | QZI | FARAWAVE PULSED FIELD ABLATION CATHETER | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Coma |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. DURING AN ABLATION PROCEDURE TO TREAT A PAROXYSMAL ATRIAL FIBRILLATION (PAF) AND A RIGHT ATRIAL FLUTTER (AFL), A FARAWAVE CATHETER AND A BLAZER CATHETER WERE SELECTED FOR USE. THE PROCEDURE WAS PERFORMED CONVENTIONALLY ON TUESDAY APRIL 1ST; THE PATIENT WAS UNDER GENERAL ANESTHESIA AND ACT WAS GREATER THAN 300 SECONDS. THE PHYSICIAN ABLATED ALL 4 VEINS, NO PULMONARY VEIN ISOLATION+. NO COMPLICATIONS OCCURRED DURING THE PROCEDURE. HOWEVER, POST PROCEDURE, THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT ON THURSDAY, APRIL 3RD. HE WAS IN A COMA ON FRIDAY, APRIL 4TH, AND DIED ON SATURDAY, APRIL 5TH FROM A CEREBRAL HEMORRHAGE (DIAGNOSED WITH AN MRI). THE PATIENT WAS YOUNG (59 YEARS OLD), HAD NO COMORBID FACTORS, AND DIED 4 DAYS AFTER THE OPERATION. NO AUTOPSY WAS PERFORMED AND NO DIRECT CAUSE TO THE PROCEDURE NOR DEVICE CAN BE ESTABLISHED. DEVICES ARE NOT EXPECTED TO BE RETURNED AS THEY WERE DISPOSED. MOREOVER, IT WAS INFORMED THAT THE PATIENT WAS NOT ON DOAC OR COUMADIN BEFORE THE PROCEDURE. ANTICOAGULATION WAS INTERRUPTED 12 HOURS BEFORE THE PROCEDURE. THE PATIENT DID NOT HAVE A HISTORY OF HYPERTENSION. ACCORDING TO THE DOCTOR, THE BRAIN MRI DID NOT SHOW A STROKE, BUT ONLY A CEREBRAL HEMORRHAGE; IT WAS A PURE HEMORRHAGIC STROKE. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE BECAUSE THE DEVICE WAS DISCARDED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21941729&pc=MVK | 3008642652-2025-04414 | 2025-03-27 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-01 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FIVE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 10:37:50 ON (B)(6) 2025. AT 22:28:05, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/INTERMITTENT CARDIAC ACTIVITY. AT 22:28:40, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. AT 22:33:51, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 22:34:43, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH TACTILE ARTIFACT. AT 22:35:11, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH TACTILE ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 22:35:38, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 22:36:06, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH TACTILE ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH CPR/ELECTRODE LEAD FALL OFF. THE ELECTRODE BELT WAS DISCONNECTED AT 22:45:16 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21941732&pc=MVK | 3008642652-2025-04417 | 2025-04-19 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-01 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 22:54:38 ON (B)(6) 2025. AT 11:50:16 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. AT 11:51:16, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 12:21:36 ON 4/19/2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21941739&pc=MVK | 3008642652-2025-04415 | 2024-10-05 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-01 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 08:36:38 ON (B)(6) 2024. AT 23:46:42, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS IDIOVENTRICULAR RHYTHM @ 20 BPM. AT 23:47:18, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS IDIOVENTRICULAR RHYTHM @ 30 BPM. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 20 BPM. THE DEVICE WAS SHUT DOWN AT 06:40:03 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21941808&pc=DYE | 2025587-2025-03276 |  | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-01 04:00 | DYE | MOSAIC MITRAL BIOPROSTHETIC HEART VALVE | Insufficient Information | Insufficient Information | P990064 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT AN UNKNOWN DURATION POST IMPLANT OF THIS 25MM MITRAL BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: B3: EVENT DATE IS NOT KNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21941963&pc=OZD | 1220648-2025-27993 | 2025-04-08 04:00 | Death | ABIOMED, INC. | 2025-05-01 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Stroke/CVA; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH CARDIOMYOPATHY WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND DURING SUPPORT, LOW FLOWS WERE ENCOUNTERED INTERMITTENTLY. REPOSITIONING OF THE PUMP WORKED WITH VARIED RESULTS AND BLOOD VOLUME WAS GIVEN TO HELP STABILIZE THE FLOW RATE. IT WAS ADDITIONALLY NOTED THAT BIOMATERIAL WAS OBSERVED AT THE TIME OF THE FLOW ISSUES. AS SUPPORT CONTINUED, THE PATIENT BECAME AGITATED, AND AN EVALUATION CONFIRMED THAT THE PATIENT HAD SUFFERED A STROKE. SUPPORT CONTINUED WITH THE IMPLANTED PUMP FOR FIVE MORE DAYS, AT WHICH POINT THE PATIENT'S PLAN OF CARE WAS CHANGED TO COMFORT MEASURES ONLY AND THE PATIENT EXPIRED. Manufacturer Narrative: GIVEN THE DETAILS OF THE EVENT, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMART ASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21942011&pc=NGV | 2124215-2025-27605 | 2025-04-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-01 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Chest Pain; Pericardial Effusion | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED. A 24MM WATCHMAN FLX PRO LAA WAS IMPLANTED ON (B)(6) 2025, AND THE PATIENT WAS ADMITTED OVERNIGHT DUE TO A LATE CASE TIME. THE PATIENT WAS DISCHARGED ON (B)(6) 2025, IN STABLE CONDITION. THE PATIENT RETURNED TO THE HOSPITAL (B)(6) 2025, WITH SEVERE CHEST PAIN. THE PATIENT HAD A COMPUTED TOMOGRAPHY (CT) SCAN THAT SHOWED NO ANEURYSM OR PERICARDIAL EFFUSION. TRANSTHORACIC ECHOCARDIOGRAM (TTE) SHOWED A TRACE PERICARDIAL EFFUSION ONLY WHICH WAS UNCHANGED. THE PATIENT WAS SEEN BY AN ELECTROPHYSIOLOGIST (EP) AND BELIEVED THE PATIENT'S SYMPTOMS WERE CONSISTENT WITH RESPIRATORY INSUFFICIENCY. THE PATIENT HAD ALSO BEEN DIAGNOSED WITH COMPLETE HEART BLOCK REQUIRING TEMPORARY VENOUS PACEMAKER, ACUTE KIDNEY INJURY SECONDARY TO ACUTE TUBULAR NECROSIS, AND HOSPITAL ACQUIRED BILATERAL PNEUMONIA. THE PATIENT'S FAMILY DECIDED TO WITHDRAW MEDICAL CARE ON (B)(6) 2025, AND THE PATIENT PASSED ON (B)(6) 2025. THE CAUSE OF DEATH WAS NOTED TO BE DUE TO ACUTE KIDNEY FAILURE. THE PHYSICIAN DID NOT PROVIDE ANY ASSESSMENT OF THE RELATIONSHIP OF THE DEATH TO THE WATCHMAN DEVICE(S) AND/OR IMPLANT PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21942115&pc=OZD | 1220648-2025-27994 | 2025-04-08 04:00 | Death | ABIOMED, INC. | 2025-05-01 04:00 | OZD | IMPELLA CP | Defective Device; Malposition of Device | Arrhythmia; Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS). HOWEVER, DURING INSERTION USING THE 14X13 PEEL-AWAY SHEATH, HALF OF THE ORANGE COLLAR OF THE SHEATH FELL OFF, AND THERE WAS OOZING FROM AROUND THE DIAPHRAGM ON THE HALF MISSING THE ORANGE PART. IMPELLA WAS ALREADY THROUGH THE SHEATH AT THIS TIME. THEREFORE, THE PHYSICIAN PROCEEDED WITH PLACEMENT AND SHEATH WAS EXCHANGED. THE BLEEDING RESOLVED WHEN 14X13 PEEL-AWAY SHEATH WAS REMOVED. THE PATIENT WAS SENT TO THE UNIT ON SUPPORT. DISCUSSION WAS MADE REGARDING TRANSFERRING OF THE PATIENT TO BE EVALUATED BY HEART FAILURE FOR VENTRICULAR ASSIST DEVICE/ORTHOTOPIC HEART TRANSPLANTATION. THE NEXT DAY AFTER START OF THE IMPELLA MCS, THE PATIENT WAS HAVING HEMOLYSIS. THE IMPELLA CP DEVICE WAS DEEPER THAN EXPECTED, AND THE PIGTAIL WAS INTERACTING WITH THE PAPILLARY MUSCLE CAUSING ECTOPIC RHYTHMS. MULTIPLE ATTEMPTS TO REPOSITION WERE MADE. THE IMPELLA CP PUMP LOST ITS POSITION, COMING OF VENTRICLE AFTER EACH REPOSITIONING. AFTER MULTIPLE REPOSITIONING ATTEMPTS, THE PHYSICIAN DECIDED TO NOT REPOSITION ANY FURTHER. SUPPORT CONTINUED. TWO DAYS LATER, IT WAS NOTED THE PATIENT WAS LISTED FOR HEART AND KIDNEY TRANSPLANT, CURRENTLY AT THAT TIME ON CONTINUOUS RENAL REPLACEMENT THERAPY. OF FURTHER NOTE WAS THE IMPELLA POSITION UNKNOWN ALARMS WERE OCCURRING DUE TO LOW PULSATILITY. SUPPORT WAS NOTED TO HAVE CONTINUED, AND APPROXIMATELY TWO DAYS LATER, THE PATIENT WOULD EXPIRE. THE PATIENT CHANGED TO DO NOT RESUSCITATE AND WITHDREW CARE. EVEN THOUGH THE PATIENT WITHDREW THE IMPELLA SUPPORT, WITH THE IMPELLA-RELATED HEMOLYSIS, ARRHYTHMIC EVENTS AND DEVICE MALPOSITIONINGS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR. Manufacturer Narrative: THE INVESTIGATION OF INTRODUCER DAMAGE ¿ MECHANICAL, HEMOLYSIS, POSITIONING ISSUES, AND VF/VT/AF/AT HAS BEEN COMPLETED. THE INTRODUCER WAS RETURNED AND INSPECTED. VISUAL INSPECTION REVEALED INSUFFICIENT ADHESIVE WAS USED ON THE ORANGE HUB WHICH ALLOWED THE HUB TO BECOME DETACHED DURING USE. THE CAUSE OF THE INTRODUCER DAMAGE - MECHANICAL WAS A DAMAGED VALVE DUE TO THE RETURNED ORANGE INTRODUCER HUB SHOWING INSUFFICIENT ADHESIVE TO REMAIN ATTACHED DURING USE. HEMOLYSIS: THE CAUSE OF THE HEMOLYSIS WAS NOT DETERMINED SINCE PUMP S/N (B)(6) WAS NOT RETURNED AND INSUFFICIENT CLINICAL INFORMATION PROVIDED. POSITIONING ISSUES: DURING IMPLANTATION PROCEDURE, IT WAS EVIDENT IMPELLA WAS BEHIND THE PAPILLARY MUSCLE AND MD REPOSITIONED. PUMP S/N (B)(6) WAS NOT RETURNED. DATA LOGS WERE NOT RECOVERED. THE CAUSE OF THE POSITIONING ISSUES WAS MOST LIKELY USE RELATED DUE TO THE PUMP BEING INSERTED TOO DEEP INTO THE PAP MUSCLES DURING IMPLANTATION PROCEDURE. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. H6 COMPONENT CODE 3038 WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-27994. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION/ANALYSIS IS UNDERWAY. UPON COMPLETION OF THE ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS. ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21942391&pc=OZD | 1220648-2025-27995 | 2025-04-08 04:00 | Death | ABIOMED, INC. | 2025-05-01 04:00 | OZD | IMPELLA CP | Malposition of Device | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT EXPERIENCED ACCESS SITE BLEEDING, DEVICE MALPOSITIONING, AND WOULD LATER EXPIRE. THE PATIENT PRESENTED WITH ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION WITH SUPRAPUBIC PAIN AND BURNING, PURPLE FINGERS AND TOES HYPOTENSIVE. IT WAS NOTED FOUR DAYS PRIOR THE PATIENT HAD A TRANSCAROTID ARTERY REVASCULARIZATION. THE DECISION WAS MADE TO IMPLANT AN IMPELLA CP DEVICE WHICH WAS COMPLETED. TEMPORARY PACER WAS PLACED LEFT FEMORAL VEIN. THE LEFT HEART CATHETERIZATION REVEALED MULTIVESSEL DISEASE AND SURGICAL CONSULT INITIATED. THE SWAN-GANZ CATHETER REPLACED THE TEMPORARY PACER. THE PEEL-AWAY SHEATH WAS REMOVED AND THE REPOSITIONING SHEATH ADVANCED ACCORDINGLY. THE PATIENT WAS TRANSFERRED TO THE UNIT ON SUPPORT. ONCE ON THE UNIT, PURSE SUTURES WERE PLACED FOR TRACK OOZING. ADDITIONALLY, ATTEMPTS TO REPOSITION WAS MADE BUT UNSUCCESSFUL UNDER ECHOCARDIOGRAM GUIDANCE. THE PUMP WAS SEEN LESS THAN OPTIMAL POSITION AT THE MITRAL. HOWEVER, THE PUMP WAS EVENTUALLY REPOSITIONED. THE PATIENT CONTINUED ON SUPPORT BUT WOULD SUBSEQUENTLY EXPIRE A LITTLE OVER SEVEN HOURS OF IMPELLA MCS. THE PATIENT MADE THE DECISION FOR NO FURTHER TREATMENT AND CARE WAS WITHDRAWN. Manufacturer Narrative: ALTHOUGH THE PATIENT DECIDED NOT TO CONTINUE WITH CARE, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE (RELATED COMPLICATIONS) AN ASSOCIATED FACTOR TO THE OVERALL OUTCOME. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21942428&pc=NVN | 2017865-2025-51671 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-01 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Syncope/Fainting | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PRESENTED IN EMERGENCY ROOM EXPERIENCING FEVER AND SYNCOPAL EPISODES. IT WAS FOUND THAT THE PATIENT HAD SEPSIS AND WENT INTO SEPTIC SHOCK. ECHOCARDIOGRAM WAS PERFORMED AND CONFIRMED VEGETATION ON MITRAL VALVE. THE PACEMAKER, RIGHT VENTRICULAR (RV) AND RIGHT ATRIAL (RA) LEADS WERE EXPLANTED AS A RESULT AND TEMPORARY RIGHT VENTRICULAR (RV) LEAD WAS PLACED. PATIENT CONDITION WAS STABLE.Event Description: NEW INFORMATION NOTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE AND INFECTIVE ENDOCARDITIS. IT WAS SUSPECTED THAT THE INFECTION CAME FROM THE TOE AMPUTATION AND THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO DEVICE OR DEVICE RELATED PROCEDURE. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21942429&pc=NVN | 2017865-2025-51672 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-01 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Syncope/Fainting | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PRESENTED IN EMERGENCY ROOM EXPERIENCING FEVER AND SYNCOPAL EPISODES. IT WAS FOUND THAT THE PATIENT HAD SEPSIS AND WENT INTO SEPTIC SHOCK. ECHOCARDIOGRAM WAS PERFORMED AND CONFIRMED VEGETATION ON MITRAL VALVE. THE PACEMAKER, RIGHT VENTRICULAR (RV) AND RIGHT ATRIAL (RA) LEADS WERE EXPLANTED AS A RESULT AND TEMPORARY RIGHT VENTRICULAR (RV) LEAD WAS PLACED. PATIENT CONDITION WAS STABLE.Event Description: NEW INFORMATION NOTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE AND INFECTIVE ENDOCARDITIS. IT WAS SUSPECTED THAT THE INFECTION CAME FROM THE TOE AMPUTATION AND THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO DEVICE OR DEVICE RELATED PROCEDURE. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21942443&pc=LWP | 2017865-2025-51665 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-01 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Syncope/Fainting | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PRESENTED IN EMERGENCY ROOM EXPERIENCING FEVER AND SYNCOPAL EPISODES. IT WAS FOUND THAT THE PATIENT HAD SEPSIS AND WENT INTO SEPTIC SHOCK. ECHOCARDIOGRAM WAS PERFORMED AND CONFIRMED VEGETATION ON MITRAL VALVE. THE PACEMAKER, RIGHT VENTRICULAR (RV) AND RIGHT ATRIAL (RA) LEADS WERE EXPLANTED AS A RESULT AND TEMPORARY RIGHT VENTRICULAR (RV) LEAD WAS PLACED. PATIENT CONDITION WAS STABLE.Event Description: NEW INFORMATION NOTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE AND INFECTIVE ENDOCARDITIS. IT WAS SUSPECTED THAT THE INFECTION CAME FROM THE TOE AMPUTATION AND THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO DEVICE OR DEVICE RELATED PROCEDURE. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21942640&pc=NPT | 2025587-2025-03285 | 2025-04-24 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-01 04:00 | NPT | EVOLUT FX DCS | Insufficient Information; Patient Device Interaction Problem | Perforation; Pericardial Effusion | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF A TRANSCATHETER HEART VALVE, A LEFT VENTRICLE PERFORATION OCCURRED. THE PERFORATION WAS REPORTEDLY DUE TO THE NON-MEDTRONIC GUIDEWIRE. THE VALVE WAS QUICKLY IMPLANTED; HOWEVER, THE PATIENT SUBSEQUENTLY DIED ON THE TABLE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE DELIVERY CATHETER SYSTEM DID NOT CAUSE OR CONTRIBUTE TO THE PERFORATION AND DEATH. THE CAUSE OF DEATH WAS AN EFFUSION FROM THE GUIDEWIRE. Manufacturer Narrative: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT CHANGED THE EVENT DESCRIPTION AND DEVICE RELATEDNESS OF THIS PREVIOUSLY REPORTED EVENT. THERE IS NO EVIDENCE TO SUGGEST THE DEVICE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY. UPDATED H.6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVFXPLUS-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21942663&pc=NHL | 3006630150-2025-02965 | 2025-04-07 04:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-05-01 04:00 | NHL | STIMULATOR, ELECTRICAL, IMPLANTED, FOR PARKINSONIAN SYMPTOMS | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Pain; Respiratory Arrest |  |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE COMPLETION OF A DEEP BRAIN STIMULATION (DBS) IMPLANT PROCEDURE, THE PATIENT INITIALLY EXHIBITED STABLE POSTOPERATIVE PROGRESS. HOWEVER, WHILE IN THE POST-ANESTHESIA CARE UNIT, SHE AWOKE IN PAIN AND WAS ADMINISTERED PAIN MEDICATION, WHICH RESULTED IN RESPIRATORY ARREST, NECESSITATING INTUBATION. SUBSEQUENTLY, THE ATTENDING PHYSICIAN IDENTIFIED A SEVERE HEMORRHAGE WITH SIGNIFICANT BLOOD LOSS AND CHARACTERIZED THE CONDITION AS CRITICAL. THE PHYSICIAN DID NOT ATTRIBUTE THE COMPLICATION TO THE IMPLANTED HARDWARE. THE PATIENT WAS THEREAFTER HOSPITALIZED AND PLACED ON LIFE SUPPORT BUT, UNFORTUNATELY, LATER SUCCUMBED TO HER CONDITION.Event Description: IT WAS REPORTED THAT FOLLOWING THE COMPLETION OF A DEEP BRAIN STIMULATION (DBS) IMPLANT PROCEDURE, THE PATIENT INITIALLY EXHIBITED STABLE POSTOPERATIVE PROGRESS. HOWEVER, WHILE IN THE POST-ANESTHESIA CARE UNIT, SHE AWOKE IN PAIN AND WAS ADMINISTERED PAIN MEDICATION, WHICH RESULTED IN RESPIRATORY ARREST, NECESSITATING INTUBATION. SUBSEQUENTLY, THE ATTENDING PHYSICIAN IDENTIFIED A SEVERE HEMORRHAGE WITH SIGNIFICANT BLOOD LOSS AND CHARACTERIZED THE CONDITION AS CRITICAL. THE PHYSICIAN DID NOT ATTRIBUTE THE COMPLICATION TO THE IMPLANTED HARDWARE. THE PATIENT WAS THEREAFTER HOSPITALIZED AND PLACED ON LIFE SUPPORT BUT, UNFORTUNATELY, LATER SUCCUMBED TO HER CONDITION. ADDITIONAL INFORMATION WAS RECEIVED THAT THE EVENT OCCURRED APPROXIMATELY 60 MINUTES AFTER THE PROCEDURE. THE PATIENT EXPERIENCED HEAD PAIN, AND THE CAUSE OF DEATH WAS DETERMINED TO BE HEMORRHAGE. Manufacturer Narrative: BLOCK B2: APPROXIMATED BASED ON THE DATE THE MANUFACTURER BECAME AWARE OF THE EVENT. ADDITIONAL SUSPECT MEDICAL DEVICE COMPONENT INVOLVED IN THE EVENT: PRODUCT FAMILY: DBS-LINEAR LEADS, UPN: M365DB2203450, MODEL: DB-2203-45, SERIAL: (B)(6), BATCH: (B)(6). GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED. Manufacturer Narrative: BLOCK B2: APPROXIMATED BASED ON THE DATE THE MANUFACTURER BECAME AWARE OF THE EVENT. ADDITIONAL SUSPECT MEDICAL DEVICE COMPONENT INVOLVED IN THE EVENT: PRODUCT FAMILY: DBS-LINEAR LEADS, UPN: M365DB2203450, MODEL: DB-2203-45, SERIAL: (B)(6), BATCH: 5002609. GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21942673&pc=MIH | 3011063223-2025-00059 | 2025-04-15 04:00 | Death | ENDOLOGIX | 2025-05-01 04:00 | MIH | AFX2 | Failure to Advance; Difficult or Delayed Activation | Hemorrhage/Bleeding; Shock; Ruptured Aneurysm | P040002 |  | 1 | Event Description: THE PATIENT WAS BEING TREATED FOR AN ABDOMINAL AORTIC ANEURYSM (AAA) WITH IMPLANT OF AN AFX2 BIFURCATED STENT GRAFT ON (B)(6) 2025. IT WAS REPORTED THAT (NON-ENDOLOGIX) 8 FR SHEATH WAS USED IN THE RIGHT SIDE AND A DRYSEAL 20 FR ON THE LEFT SIDE. A LUNDERQUIST (NON-ENDOLOGIX) EXTRA STIFF GUIDEWIRE WAS USED AND AN INDY (NON-ENDOLOGIX) OVER THE WIRE (OTW) VASCULAR RETRIEVER WAS USED AS A SNARE CATHETER. THE DELIVERY SYSTEM OF THE AFX2 BIFURCATED STENT GRAFT WAS INSERTED THROUGH THE LEFT ACCESS. THE CONTRALATERAL WIRE WAS CAUGHT BY THE SNARE CATHETER ON THE AORTIC BIFURCATION, THEN THE INNER CORE WAS ADVANCED TO THE POINT WHERE THE TIP CAME OUT FROM THE TOP OF THE DRYSEAL SHEATH. WHEN THE PHYSICIAN WAS REMOVING THE TRANSFER STOP AND TRYING TO ADVANCE THE INNER CORE FURTHER, THE GUIDEWIRE GOT WARPED ON THE SIDE OF THE GREATER CURVE OF THE ANEURYSM, CAUSING THE AFX2 BIFURCATED STENT GRAFT TO DROP DOWN TOWARD THE ANEURYSM AND TO BE STUCK. SUBSEQUENTLY, DEPLOYMENT WAS ATTEMPTED BY ADVANCING THE SHEATH PROXIMALLY IN THE ANEURYSM AND THEN PULLING THE SHEATH. THE INNER CORE WAS PULLED BACK TO MINIMIZE THE GAP BETWEEN THE TOP OF THE SHEATH AND THE TIP AND THE PHYSICIAN ADVANCED THE SHEATH. HOWEVER, IT STOPPED ADVANCING BECAUSE THE GUIDEWIRE GOT WARPED ON THE SIDE OF THE GREATER CURVE OF THE ANEURYSM. THE PHYSICIAN THEN ATTEMPTED TO ADVANCE THE SHEATH ONLY TO THE PROXIMAL NECK AND THEN INSERT THE AFX2 INTO THE SHEATH FOR DELIVERY. HOWEVER, THE SHEATH GOT PULLED DOWN DUE TO ADVANCING THE AFX2. AS A RESULT, THE AFX2 STENT GRAFT WAS EXPOSED AND DID NOT ADVANCE. IT WAS CONFIRMED THAT THE TIP OF THE AFX2 GOT BENT; THEREFORE, THE DECISION WAS MAD TO REMOVE THE AFX2 DEVICE FROM THE PATIENT. THE STIFF WIRE ALSO GOT BENT AND IT WAS REPLACED WITH A NEW ONE. AN AFX SHEATH WAS INSERTED INTO THE DRYSEAL 20 FR (SHEATH IN SHEATH) TO DELIVER A NEW AFX2 BIFURCATED STENT GRAFT. WHEN THE INNER CORE WAS ADVANCED UNTIL IT REACHED THE TRANSFER STOP PER THE INSTRUCTIONS, THE STIFF WIRE GOT WARPED IN THE ANEURYSM AND DEPLOYMENT WAS UNSUCCESSFUL. BOTH THE AFX2 DELIVERY SYSTEM AND THE AFX HEATH WERE REMOVED FROM THE PATIENT. IT WAS CONFIRMED THAT THE AFX SHEATH GOT DEFORMED INTO AN ACCORDION-LIKE SHAPE AND THE TIP WAS BENT. THE PHYSICIAN CONCLUDED THAT THE LEFT APPROACH WAS DIFFICULT AND THE RIGHT APPROACH WAS BEING CONSIDERED INSTEAD. HOWEVER, WHILE THE PHYSICIAN WAS PREPARING THE RIGHT APPROACH THE PATIENT'S BLOOD PRESSURE DECREASED. A CONTRAST ENHANCED COMPUTED TOMOGRAPHY (CT) CONFIRMED THE ANEURYSM WAS RUPTURED, THEREFORE, THE ENDOVASCULAR ANEURYSM REPAIR (EVAR) PROCEDURE WAS ABORTED. AN ARTIFICIAL VESSEL REPLACEMENT SURGERY WAS PLANNED INSTEAD; HOWEVER, THE PATIENT'S VITAL SIGNS WERE UNSTABLE AND THE PUPILS WERE DILATED. CONCLUDING THAT THE OPEN SURGERY WAS NOT POSSIBLE THE PHYSICIAN ENDED THE PROCEDURE. IT WAS LATER REPORTED THAT THE PATIENT EXPIRED AND THE CAUSE OF DEATH WAS NOTED TO BE HEMORRHAGIC SHOCK DUE TO ABDOMINAL AORTIC ANEURYSMS RUPTURE. THE PHYSICIAN NOTED THAT THE EVENT COULD OCCUR BECAUSE THE DEVICE WAS UNABLE TO FOLLOW THE ANGULATED PART OF THE ANEURYSM AND THAT CAUSED DIFFICULTY OF INSERTING THE DELIVERY SYSTEM. THE BLOOD VESSEL GOT DAMAGED DURING MANIPULATION IN THE ANEURYSM (SUCH AS ADVANCING THE SHEATH) AND THE ANEURYSM RUPTURED. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. AN EVALUATION OF THE MANUFACTURING RECORD WAS COMPLETED. A REVIEW OF THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION SHOWS THE DEVICE TO HAVE BEEN PROPERLY MANUFACTURED AND RELEASED IN ACCORDANCE WITH THE DEVICE MASTER RECORD. THE REVIEW CONFIRMS THERE WERE NO MANUFACTURING OR PROCESSING NON-CONFORMITIES IDENTIFIED THAT WOULD CONTRIBUTE TO THE REPORTED ADVERSE EVENT/INCIDENT. ENDOLOGIX RECEIVED TWO (2) AFX2 BIFURCATED ENDOGRAFT SYSTEMS, INCLUDING THE STENTS, AND ONE (1) AFX INTRODUCER SYSTEM. THE DEVICES WERE PROPERLY BOXED AND BAGGED. UPON INITIAL INSPECTION, TRACES OF BLOOD AND TISSUE RESIDUE WERE OBSERVED, PROMPTING DECONTAMINATION. VISUAL INSPECTION REVEALED THAT BOTH STENT GRAFTS REMAINED IN THEIR ORIGINAL, UNEXPANDED STATE. EXCESSIVE KINKING WAS OBSERVED ON BOTH AFX2 BIFURCATED ENDOGRAFT SYSTEMS, SPECIFICALLY PROXIMAL TO THE HANDLE AND THE OUTER SHEATH NEAR THE RADIOPAQUE TIP. ONE OF THE DEVICES EXHIBITED A SIGNIFICANT BEND AT THE CENTER OF THE INNER CORE. THE STENTS WERE DEPLOYED FOR FURTHER EVALUATION AND DEMONSTRATED OVERALL STRUCTURAL INTEGRITY WITH NO VISIBLE DAMAGE. THE REPORTED ISSUES OF KINKED COMPONENT AND UNABLE TO EPLOY WERE CONFIRMED. HOWEVER, DUE TO THE EXTENT OF THE DAMAGE, ENDOLOGIX COULD NOT REPRODUCE THE REPORTED ISSUE OF DIFFICULT TO ADVANCE. THE ROOT CAUSE OF THE OBSERVED DAMAGES REMAINS UNDETERMINED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT WAS COMPLETED. AN EXAMINATION OF MEDICAL RECORDS AND/OR MEDICAL IMAGING RECEIVED BY ENDOLOGIX SHOWS THAT THE INTRAOPERATIVELY UNABLE TO ADVANCE DELIVERY SYSTEM, KINKED COMPONENT, UNABLE TO DEPLOY MAIN BODY, ANEURYSM RUPTURE, HEMORRHAGIC SHOCK AND DEATH ARE UNCONFIRMED. THIS IS NOT CONSISTENT WITH THE REPORTED ADVERSE EVENT/INCIDENT. THE DEATH WAS DETERMINED TO BE INCONCLUSIVE. THE INFRARENAL ANGULATION WAS 69.7 DEGREES AND SHOULD BE LESS THAN OR EQUAL TO 60 DEGREES. IT IS UNCLEAR IF THE OFF-LABEL ANGULATION CONTRIBUTED TO THE REPORTED EVENT. DEVICE, USER PROCEDURE OR ANATOMY RELATEDNESS OF COMPLAINT COULD NOT BE DETERMINED. PROCEDURE RELATED HARMS COULD NOT BE DETERMINED. THE FINAL PATIENT STATUS WAS REPORTED AS DECEASED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS. CORRECTIONS: D9: DEVICE AVAILABLE FOR EVAL HAS BEEN UPDATED. D9: DATE RECEIVED HAS BEEN UPDATED. H3: DEVICE EVALUATED BY MFG HAS BEEN UPDATED. H3: DEVICE RET TO MFG FOR EVAL HAS BEEN UPDATED. G3: AWARENESS DATE HAS BEEN UPDATED. H6: INVESTIGATION FINDING CODES: REMOVE CODE 3233. H6: INVESTIGATION CONCLUSION CODES: REMOVE CODE 11. Manufacturer Narrative: THE DELIVERY SYSTEM INVOLVED IN THIS EVENT IS EXPECTED TO BE RETURNED FOR EVALUATION. PATIENT MEDICAL RECORDS AND IMAGING STUDIES WILL BE REQUESTED FOR FURTHER EVALUATION BY THE CLINICAL SPECIALIST. IF ADDITIONAL INFORMATION PERTINENT TO THE INCIDENT IS OBTAINED, A FOLLOW-UP REPORT WILL BE SUBMITTED. DEVICE ITERATION IS AFX2.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21942844&pc=OZD | 1220648-2025-27997 | 2025-04-09 04:00 | Death | ABIOMED, INC. | 2025-05-01 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT WHILE AN IMPELLA CP WAS IMPLANTED THE PATIENT HAD EPISODES OF VENTRICULAR TACHYCARDIA (VT) WITH DEFIBRILLATION. THE PUMP WAS REPOSITIONED AND SUPPORT CONTINUED. THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21943612&pc=OZD | MW5169721 | 2024-05-09 04:00 | Death | ABIOMED, INC. | 2025-05-01 04:00 | OZD | IMPELLA CP | Therapeutic or Diagnostic Output Failure | Failure to Anastomose; Hemorrhage/Bleeding; Cardiogenic Shock; Thrombocytopenia |  |  | 1 | Event Description: ON (B)(6) 2024 DR. (B)(6) PLANNED TO PLACE THE IMPELLA 5.5 THROUGH THE SUBCLAVIAN TO BRIDGE UNTIL CABG/MV REPAIR ON (B)(6) 2024. DURING THE IN INITIAL PROCEDURE THE IMPELLA 5.5 "WOULD NOT FIT" PER DR. (B)(6) SO THE DECIDED TO PLACE THE IMPELLA CP AND DO SURGERY THE FOLLOWING MORNING (B)(6) 2025. AT 3:32 (B)(6) 2024 IT WAS DISCOVERED THAT "THE IMPELLA CP HAD BASICALLY QUIT WORKING BY THE TIME THEY GOT MS. (B)(6) TO THE OPERATING ROOM" PER DR. (B)(6). ALSO, THERE WAS DIFFICULT EXPOSURE DUE TO THE LEAK AROUND THE IMPELLA DEVICE IN THE ASCENDING AORTA. THE IMPELLA WAS REMOVED AND REPLACED WITH AN IABP (INTRA-AORTIC BALLOON PUMP). MS. (B)(6) CONTINUED TO HAVE BLEEDING WHICH LED DR. (B)(6) TO RETURN HER TO THE OPERATING ROOM. THERE WAS EPICARDIAL BLEEDING ON THE ANTERIOR SURFACE OF THE HEART AND THERE WAS ALSO BLEEDING FROM THE PROXIMAL ANASTOMOSIS VEIN GRAFT TO THE AORTA. BLEEDING WAS CONTROLLED AND PATIENT WAS SENT BACK TO CICU. PATIENT EXPIRED ON (B)(6) 2025 AT 2:12 AM FROM CARDIOGENIC SHOCK. THROMBOCYTOPENIA MOST LIKELY RELATED TO IMPELLA.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21943714&pc=DSQ | 2916596-2025-02861 | 2025-04-15 04:00 | Death | THORATEC CORPORATION | 2025-05-01 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Dyspnea; High Blood Pressure/ Hypertension; Pneumonia; Pulmonary Edema | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT LOG FILES WERE SUBMITTED FOR EVALUATION FOR A RECENTLY RE-ADMITTED PATIENT IN THE INTENSIVE CARE UNIT (ICU) FROM THE CARDIAC FLOOR. THE PATIENT WAS IMPLANTED WITH HEARTMATE 3 ON (B)(6) 2025. THEY WENT TO THE FLOOR THE NEXT DAY AND THEN WAS READMITTED TO THE INTENSIVE CARE UNIT (ICU) FOR LOW FLOW ALARMS AND HIGH MEAN ARTERIAL PRESSURE (MAP) AFTER STOPPING MILRINONE. MILRINONE WAS RESTARTED. THE PATIENT STILL HAD SHORTNESS OF BREATH (SOB) WITH SUSPECTED PULMONARY EDEMA AND BREATHING VERY FAST. THE LVAD RPMS WERE ALSO BROUGHT BACK DOWN FROM 5400 TO 5300 RPM. IT WAS NOTED THE PATIENT WAS ON A DIURETIC. AN ECHOCARDIOGRAM WAS PLANNED, BUT THE SITE WANTED TO MAKE SURE THEY WERE NOT MISSING ANYTHING OTHER THAN HYPERTENSION AND RIGHT-SIDE ISSUES. THE EVENT LOG FILE CAPTURED LOW FLOW ALARM EVENTS ON 15APR2025. THERE WERE ALSO SEVERAL MOMENTARY LOW FLOW EVENTS RECORDED. SOME OF THESE EVENTS WERE VERY BRIEF AND DID NOT GENERATE AN ALARM. THERE DOES NOT APPEAR TO BE ANY TYPE OF EXTERNAL MECHANICAL CIRCULATORY SUPPORT (MCS) EQUIPMENT ISSUES CAUSING THESE EVENTS. THE LOW FLOW EVENTS MAY BE DUE TO A PATIENT RELATED ISSUE.Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FROM PNEUMONIA ON 11MAY2025. THE INFECTION WAS NOT RELATED TO DEVICE AND THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: CORRECTED DATA: SECTION A3A: PATIENT GENDER CORRECTED. SECTION D4: PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER CORRECTED. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21944272&pc=NPT | 2025587-2025-03297 | 2025-04-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-01 04:00 | NPT | EVOLUT FX PLUS VALVE | Device Dislodged or Dislocated; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Foreign Body In Patient; Unspecified Heart Problem | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING A TRANSCATHETER AORTIC VALVE (EVFXPLUS-26), SEVERAL COMPLICATIONS AROSE. THE LEFT CORONARY ARTERY EXHIBITED A 95% STENOSIS, WHICH WAS ADDRESSED WITH WIRING AND STENTING. THE VALVE WAS DEPLOYED UNDER PACING AT 120 BEATS PER MINUTE, BUT UPON RELEASE AT A DEPTH OF 4-5 MM, THE PATIENT EXPERIENCED HYPOTENSION AND A COMPLICATION WITH THE RIGHT VENTRICLE. AN ATTEMPT WAS MADE TO CANNULATE THE RIGHT CORONARY ARTERY (RCA) WHICH WAS UNSUCCESSFUL. ATTEMPTS TO ENGAGE A GUIDE OR WIRE DOWN THE RCA WERE ALSO UNSUCCESSFUL. LIMITED BLOOD FLOW WAS NOTED. A 20-MILLIMETER NON-MEDTRONIC BALLOON WAS INFLATED WITHIN THE VALVE FRAME IN AN ATTEMPT TO REPOSITION THE VALVE, BUT IT REMAINED IMMOBILE. THE VALVE WAS SNARED INTO THE ASCENDING AORTA; HOWEVER, BLOOD FLOW WAS STILL LIMITED ON THE RIGHT. CARDIOPULMONARY RESUSCITATION WAS PERFORMED FOR AN HOUR, AND THE PATIENT WAS SHOCKED MULTIPLE TIMES FOLLOWING ADVANCED CARDIOVASCULAR LIFE SUPPORT PROTOCOLS. DESPITE THESE EFFORTS, THE PATIENT WAS DECLARED DEAD. THE PROCEDURE AND THE MEDTRONIC DEVICE WERE NOTED TO HAVE CONTRIBUTED TO THE DEATH, WITH SMALL SINUSES OF VALSALVA IDENTIFIED AS A CONTRIBUTING ANATOMICAL FACTOR.Event Description: IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING A TRANSCATHETER AORTIC VALVE (EVFXPLUS-26), SEVERAL COMPLICATIONS AROSE. THE LEFT CORONARY ARTERY EXHIBITED A 95% STENOSIS, WHICH WAS ADDRESSED WITH WIRING AND STENTING. THE VALVE WAS DEPLOYED UNDER PACING AT 120 BEATS PER MINUTE, BUT UPON RELEASE AT A DEPTH OF 4-5 MM, THE PATIENT EXPERIENCED HYPOTENSION AND A COMPLICATION WITH THE RIGHT VENTRICLE. AN ATTEMPT WAS MADE TO CANNULATE THE RIGHT CORONARY ARTERY (RCA) WHICH WAS UNSUCCESSFUL. ATTEMPTS TO ENGAGE A GUIDE OR WIRE DOWN THE RCA WERE ALSO UNSUCCESSFUL. LIMITED BLOOD FLOW WAS NOTED. A 20-MILLIMETER NON-MEDTRONIC BALLOON WAS INFLATED WITHIN THE VALVE FRAME IN AN ATTEMPT TO REPOSITION THE VALVE, BUT IT REMAINED IMMOBILE. THE VALVE WAS SNARED INTO THE ASCENDING AORTA; HOWEVER, BLOOD FLOW WAS STILL LIMITED ON THE RIGHT. CARDIOPULMONARY RESUSCITATION WAS PERFORMED FOR AN HOUR, AND THE PATIENT WAS SHOCKED MULTIPLE TIMES FOLLOWING ADVANCED CARDIOVASCULAR LIFE SUPPORT PROTOCOLS. DESPITE THESE EFFORTS, THE PATIENT WAS DECLARED DEAD. THE PROCEDURE AND THE MEDTRONIC DEVICE WERE NOTED TO HAVE CONTRIBUTED TO THE DEATH, WITH SMALL SINUSES OF VALSALVA IDENTIFIED AS A CONTRIBUTING ANATOMICAL FACTOR. ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER. THE ONLY BALLOON USAGE WAS THE ATTEMPT TO MOVE THE VALVE OUT OF THE AORTIC ANNULUS EMERGENTLY. THE VALVE DID NOT MOVE WITH BALLOON; IT WAS PULLED UP WITH A SNARE INTENTIONALLY. THE IMPLANT DEPTH WAS 4 MM ON THE NON-CORONARY CUSP (NCC). A NON-MEDTRONIC (SAFARI) GUIDEWIRE WAS USED FOR THE PROCEDURE. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {D-EVOLUTFX-2329}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEART VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5, H6. CORRECTED H6 TO REMOVE A051201 CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21944509&pc=NGV | 2135147-2025-02293 | 2025-04-08 04:00 | Death | ABBOTT MEDICAL | 2025-05-01 04:00 | NGV | AMPLATZER AMULET | Patient-Device Incompatibility | Intracranial Hemorrhage; Perforation; Cardiac Tamponade; Pericardial Effusion | P200049 |  | 1 | Event Description: IT WAS REPORTED ON (B)(6) 2025 A 25MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS SELECTED FOR IMPLANT USING A 14F AMULET DELIVERY SHEATH TO TREAT A LEFT ATRIAL APPENDAGE (LAA). THE PATIENT HAD A PRE-EXISTING ATRIAL FIBRILLATION (AF). TRANSSEPTAL PUNCTURE WAS INFERIOR AND MIDDLE IN THE OVAL FOSSA. DURING THE PROCEDURE, 5000 UNITS OF HEPARIN WERE ADMINISTERED. WHILE DEPLOYING THE OCCLUDER IN BALL POSITION, A CIRCUMFERENTIAL PERICARDIAL EFFUSION WAS DETECTED IN THE LEFT ATRIUM. THE EFFUSION LED TO A CARDIAC TAMPONADE. A DRAINAGE WAS PERFORMED; HOWEVER, IT WAS NOTED THAT THERE WAS A HOLE INSIDE THE LEFT ATRIUM. THE OCCLUDER WAS NOT RELEASED FROM THE DELIVERY CABLE AND WAS REMOVED FROM THE PATIENT. A SURGICAL PROCEDURE WAS REQUIRED TO SUTURE THE HOLE CLOSED. THE LAA WAS ALSO SURGICALLY CLOSED. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). PER THE PHYSICIAN, THE PUNCTURE THAT LED TO THE PERICARDIAL EFFUSION AND SUBSEQUENT CARDIAC TAMPONADE WAS CAUSED BY THE END OF THE OCCLUDER. ON (B)(6) 2025 A CEREBRAL HEMATOMA WAS DISCOVERED. NO INTERVENTION WAS PERFORMED. THE PATIENT REMAINED IN INTENSIVE CARE INTUBATED VENTILATED. ON (B)(6) 2025 THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS A CEREBRAL HEMATOMA. THE PHYSICIAN BELIEVED THE DEATH WAS DUE TO COMPLICATIONS RELATED TO THE PROCEDURE AND THE PATIENT'S COMORBIDITIES. Manufacturer Narrative: AN EVENT OF PERICARDIAL EFFUSION, CARDIAC TAMPONADE, PERFORATION, CEREBRAL HEMATOMA, PROLONGED HOSPITALIZATION, AND DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO SIMILAR COMPLAINTS REPORTED FROM THIS LOT. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED THE CAUSE OF THE REPORTED PERICARDIAL EFFUSION AND CARDIAC TAMPONADE APPEARS TO BE RELATED TO PROCEDURAL COMPLICATION (PERFORATION). HOWEVER, THE CAUSE OF THE PERFORATION COULD NOT BE CONCLUSIVELY DETERMINED. THE CAUSE OF THE REPORTED HEMATOMA AND DEATH COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED SURGICAL INTERVENTION AND PROLONGED HOSPITALIZATION WERE RESULTS OF CASE-SPECIFIC CIRCUMSTANCES AS THE HOLE AND THE LEFT ATRIAL APPENDAGE WERE SURGICALLY CLOSED, AND THE PATIENT WAS TRANSFERRED TO THE ICU. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21944515&pc=MCX | 3004742232-2025-00108 | 2019-04-05 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-05-01 04:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P130005 |  | 1 | Event Description: A DIAMONDBACK 360 ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A LESION ON (B)(6) 2019. THE PATIENT EXPIRED WEEKS AFTER THE PROCEDURE DUE TO CARDIAC DEATH. THE CLINICAL EVENTS COMMITTEE REVIEWED THE ANGIOGRAPHIC IMAGES AND CONCLUDED STENT THROMBOSIS OCCURRED WHICH WAS POSSIBLY THE CAUSE OF DEATH AND THE OAD MAY HAVE CONTRIBUTED. THE SITE DID NOT REPORT ANY ADVERSE EFFECTS. NO FURTHER INFORMATION IS AVAILABLE. SUBJECT ID: ECL-178-004.Event Description: A DIAMONDBACK 360 ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A LESION ON 14 MARCH 2019. THE PATIENT EXPIRED WEEKS AFTER THE PROCEDURE DUE TO CARDIAC DEATH. THE CLINICAL EVENTS COMMITTEE REVIEWED THE ANGIOGRAPHIC IMAGES AND CONCLUDED STENT THROMBOSIS OCCURRED WHICH WAS POSSIBLY THE CAUSE OF DEATH AND THE OAD MAY HAVE CONTRIBUTED. THE SITE DID NOT REPORT ANY ADVERSE EFFECTS. NO FURTHER INFORMATION IS AVAILABLE. SUBJECT ID: (B)(6). Manufacturer Narrative: ADDITIONAL INFORMATION: D4, H4, H6, H10. THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED PATIENT THROMBOSIS, AND DEATH APPEAR TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE IT IS LIKELY THAT THE REPORTED PATIENT THROMBOSIS AND DEATH ARE A RESULT OF THE PATIENT¿S DISEASE SEVERITY COMBINED WITH USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21944640&pc=OZD | 1220648-2025-28004 | 2025-03-06 05:00 | Death | ABIOMED, INC. | 2025-05-01 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Tachycardia; Ventricular Fibrillation; Cardiogenic Shock | P140003 |  | 1 | Event Description: IT WAS REPORTED VIA A NOTIFICATION THAT WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED CARDIOGENIC SHOCK WITH A RELATEDNESS TO THE IMPELLA LISTED AS UNKNOWN (UNDETERMINED). IT WAS FURTHER REPORTED THAT THE PATIENT WAS EXHIBITING ARRHYTHMIA AND HYPOTENSION AND IMPELLA WAS IMPLANTED. AICD WAS PLACED, BUT VENTRICULAR TACHYCARDIA/VENTRICULAR FIBRILLATION WAS PERSISTENT, AND THE FAMILY DECIDED TO WITHDRAW CARE WHILE STILL ON IMPELLA SUPPORT. THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21944696&pc=MCX | 3004742232-2025-00110 | 2020-12-28 05:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-05-01 04:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Myocardial Infarction; Vascular Dissection | P130005 |  | 1 | Event Description: A DIAMONDBACK CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS SELECTED FOR TREATMENT OF A LESION IN THE PROXIMAL CIRCUMFLEX ARTERY VIA FEMORAL ACCESS. THE SEVERELY CALCIFIED LESION WAS 12MM IN LENGTH WITH A VESSEL DIAMETER OF 2.5MM. AFTER ONE TREATMENT ON LOW SPEED, A TYPE D DISSECTION OCCURRED. THE OAD WAS REMOVED, AND BALLOON ANGIOPLASTY AND STENT PLACEMENT WERE PERFORMED. THE VESSEL LOOKED DECENT FOLLOWING STENT DEPLOYMENT. THE PATIENT BECAME HYPOTENSIVE. DOPAMINE, EPINEPHRINE AND LEVOPHED WERE ADMINISTERED. AN IMPELLA HEART PUMP WAS INSERTED. THE HEART PUMP WAS THEN REMOVED. THE PATIENT ALSO EXPERIENCED ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE DURING STENT POSITIONING AND DEPLOYMENT. AMIODARONE WAS ADMINISTERED BUT DID NOT CONVERT THE ARRHYTHMIA. CARDIOVERSION WAS PERFORMED, AND THE CARDIOVERSION WAS SUCCESSFUL IN RETURNING THE PATIENT TO A NORMAL SINUS RHYTHM. THE OPINION OF THE PHYSICIAN WAS THAT THE DISSECTION LED TO THE HYPOTENSION AND ARRYTHMIA. THE PATIENT WAS IN CRITICAL CONDITION AND WAS ADMITTED TO THE INTENSIVE CARE UNIT. THE PATIENT REMAINED HEMODYNAMICALLY UNSTABLE AND EXPERIENCED WORSENING BRADYCARDIA AND PULSELESS ELECTRICAL ACTIVITY. A TEMPORARY PACEMAKER WAS PLACED. THE PATIENT CODED, AND SINUS RHYTHM WAS ABLE TO BE RESTORED. THE PATIENT WAS TRANSFERRED BACK TO THE INTENSIVE CARE UNIT AND PULSELESS ELECTRICAL ACTIVITY CONTINUED INTERMITTENTLY. A DO NOT RESUSCITATE WAS SIGNED BY THE PATIENT'S FAMILY, AND HEMODYNAMIC SUPPORT WAS DISCONTINUED. THE PATIENT EXPIRED ON (B)(6) 2020. THE CAUSES OF DEATH WERE REPORTED TO BE BRADYCARDIA AND PULSELESS ELECTRICAL ACTIVITY. ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL EVENTS COMMITTEE INDICATED THAT A MYOCARDIAL INFARCTION (MI) OCCURRED AND IS POSSIBLY RELATED TO ORBITAL ATHERECTOMY. MI WAS NOT REPORTED BY THE SITE. SUBJECT ID: (B)(6).Event Description: A DIAMONDBACK CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS SELECTED FOR TREATMENT OF A LESION IN THE PROXIMAL CIRCUMFLEX ARTERY VIA FEMORAL ACCESS. THE SEVERELY CALCIFIED LESION WAS 12MM IN LENGTH WITH A VESSEL DIAMETER OF 2.5MM. AFTER ONE TREATMENT ON LOW SPEED, A TYPE D DISSECTION OCCURRED. THE OAD WAS REMOVED, AND BALLOON ANGIOPLASTY AND STENT PLACEMENT WERE PERFORMED. THE VESSEL LOOKED DECENT FOLLOWING STENT DEPLOYMENT. THE PATIENT BECAME HYPOTENSIVE. DOPAMINE, EPINEPHRINE AND LEVOPHED WERE ADMINISTERED. AN IMPELLA HEART PUMP WAS INSERTED. THE HEART PUMP WAS THEN REMOVED. THE PATIENT ALSO EXPERIENCED ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE DURING STENT POSITIONING AND DEPLOYMENT. AMIODARONE WAS ADMINISTERED BUT DID NOT CONVERT THE ARRHYTHMIA. CARDIOVERSION WAS PERFORMED, AND THE CARDIOVERSION WAS SUCCESSFUL IN RETURNING THE PATIENT TO A NORMAL SINUS RHYTHM. THE OPINION OF THE PHYSICIAN WAS THAT THE DISSECTION LED TO THE HYPOTENSION AND ARRYTHMIA. THE PATIENT WAS IN CRITICAL CONDITION AND WAS ADMITTED TO THE INTENSIVE CARE UNIT. THE PATIENT REMAINED HEMODYNAMICALLY UNSTABLE AND EXPERIENCED WORSENING BRADYCARDIA AND PULSELESS ELECTRICAL ACTIVITY. A TEMPORARY PACEMAKER WAS PLACED. THE PATIENT CODED, AND SINUS RHYTHM WAS ABLE TO BE RESTORED. THE PATIENT WAS TRANSFERRED BACK TO THE INTENSIVE CARE UNIT AND PULSELESS ELECTRICAL ACTIVITY CONTINUED INTERMITTENTLY. A DO NOT RESUSCITATE WAS SIGNED BY THE PATIENT'S FAMILY, AND HEMODYNAMIC SUPPORT WAS DISCONTINUED. THE PATIENT EXPIRED ON (B)(6) 2020. THE CAUSES OF DEATH WERE REPORTED TO BE BRADYCARDIA AND PULSELESS ELECTRICAL ACTIVITY. ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL EVENTS COMMITTEE INDICATED THAT A MYOCARDIAL INFARCTION (MI) OCCURRED AND IS POSSIBLY RELATED TO ORBITAL ATHERECTOMY. MI WAS NOT REPORTED BY THE SITE. SUBJECT ID: (B)(6). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED RELATED TO PREVIOUSLY SUBMITTED MDR. CARDIOVASCULAR SYSTEMS INCORPORATED HAS TRANSITIONED TO A NEW COMPLAINT HANDLING SYSTEM AND NO LONGER HAS ACCESS TO SUBMIT FOLLOW-UP MDRS USING ESUBMITTER/WEBTRADER. THIS MDR IS BEING FILED FROM OUR NEW SYSTEM USING THE AS2 GATEWAY WITH REFERENCE TO THE PREVIOUSLY REPORTED MANUFACTURER REPORT NUMBER IN H10. MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED PATIENT MYOCARDIAL INFARCTION, VASCULAR DISSECTION, ARRHYTHMIA, AND DEATH APPEAR TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE IT IS POSSIBLE THAT THE REPORTED PATIENT MYOCARDIAL INFARCTION, VASCULAR DISSECTION, ARRHYTHMIA, AND DEATH ARE A RESULT OF THE PATIENT¿S DISEASE SEVERITY AND/OR USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21945084&pc=DSQ | 2916596-2025-02669 | 2025-04-06 04:00 | Death | THORATEC CORPORATION | 2025-05-01 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT WHILE THE PATIENT WAS ADMITTED FOR COVID-19, THE PATIENT PASSED AWAY DUE TO A BRAIN HEMORRHAGE. THE VENTRICULAR ASSIST DEVICE (VAD) COORDINATOR DID NOT MENTION ANY DEVICE RELATED ISSUES. IT WAS REPORTED THAT THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21945369&pc=OZD | 1220648-2025-28013 | 2025-04-10 04:00 | Death | ABIOMED, INC. | 2025-05-01 04:00 | OZD | IMPELLA CP | Defective Device; Material Split, Cut or Torn | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING THE IMPELLA CP PROCEDURE, AFTER POSITIONING THE GUIDEWIRE IN THE VENTRICLE THE WIRE COULDN¿T BE REMOVED. THE WIRE STOPPED IN THE PEEL AWAY SHEATH AND THEY BOTH WERE REMOVED TOGETHER. IT WAS NOTED THE TIP OF THE INTRODUCER WAS SPLIT. IT WAS ALSO REPORTED THAT THE PATIENT EXPIRED WHILE ON IMPELLA CP SUPPORT. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: DELIVERY ISSUES: THE FM "PRODUCT DAMAGE (GUIDEWIRE DAMAGE - PERIPHERAL VESSEL)" WAS UPDATED TO THE FM "DELIVERY ISSUES" SINCE THERE WERE NO FULL/PARTIAL DETACHMENTS OBSERVED ON THE RETURNED GUIDEWIRE. CLINICAL DETAILS STATE THAT THE 0.018" WIRE COULD NOT BE REMOVED AFTER DELIVERING THE PUMP, STOPPING IN THE PEEL-AWAY SHEATH. THE GUIDEWIRE WAS RETURNED AND MINOR BENDING WAS OBSERVED NEAR THE TIP AS WELL AS SOME UNRAVELING AT THE BEGINNING OF THE COILED SECTION ON THE WIRE. THE ROOT CAUSE OF THE DELIVERY ISSUES WAS NOT DETERMINED SINCE INSUFFICIENT CLINICAL DETAILS WERE PROVIDED. INTRODUCER DAMAGE - MECHANICAL: CLINICAL DETAILS STATE THAT THE 0.018" GUIDEWIRE COULD NOT BE REMOVED THROUGH THE SHEATH AND SHEATH TIP DAMAGE WAS OBSERVED UPON REMOVAL OF THE SHEATH. IT WAS NOTED THAT THERE WERE NO VESSEL CONDITIONS AND IT'S UNKNOWN IF EXCESSIVE FORCE WAS USED. THE 14FRX13CM SHEATH WAS RETURNED AND THERE WAS DAMAGE FOUND ON THE TIP AS WELL AS A SLIGHT KINK AT THE BASE OF THE SHEATH NEAR THE HUB. THE ROOT CAUSE OF THE INTRODUCER DAMAGE WAS MOST LIKELY SHEATH TIP SPLIT BUT THE CAUSE OF THE DAMAGE IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21945646&pc=NGV | 2124215-2025-27258 | 2025-04-23 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-01 04:00 | NGV | SYSTEM, APPENDAGE CLOSURE, LEFT ATRIAL | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED VIA A WATCHMAN NEWSLETTER CREATED BY AN IMPLANTING FACILITY THAT DEATHS OCCURRED. A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) QUALITY REPORT FOCUSING IN ON THE 45 DAY FOLLOW UP METRIC FOR PATIENT WHO HAD A WATCHMAN IMPLANTS WAS CREATED BY AN IMPLANTING FACILITY. THE REPORT PROVIDED INFORMATION FOR LAAC PROCEDURES WHERE THE PATIENT HAD ANY EVENT FROM PROCEDURE THROUGH THE 45 DAY INSIDE WINDOW FOLLOW UP PERIOD FROM THE FOURTH QUARTER (Q4) 2022 THROUGH THE THIRD QUARTER (Q3) 2024. THERE WERE SIX (6) REPORTED PATIENT DEATHS. NO FURTHER INFORMATION IS AVAILABLE AS SPECIFIC PATIENT DETAILS WERE NOT PROVIDED. Manufacturer Narrative: B3 DATE OF EVENT: AWARE DATE WAS USED AS EVENT DATE AS ACTUAL EVENT DATE WAS NOT KNOWN. D1: BRAND NAME IS BLANK BECAUSE THE DEVICE IS KNOWN TO BE A WATCHMAN LAA CLOSURE DEVICE, BUT THE EXACT BRAND NAME IS UNKNOWN. D4 UNIQUE IDENTIFIER (UDI) #: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21946395&pc=LWS | 2017865-2025-51506 | 2025-04-07 04:00 | Death | ABBOTT | 2025-05-01 04:00 | LWS | ELLIPSE VR ICD, US | Failure to Deliver Shock/Stimulation; Charging Problem | Arrhythmia; Syncope/Fainting | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A CARDIAC ARREST AND SYNCOPE. THE DEVICE FAILED TO DELIVER HIGH VOLTAGE THERAPY DURING EPISODES OF VENTRICULAR FIBRILLATION. LOW IMPEDANCE WAS OBSERVED ON THE RIGHT VENTRICULAR (RV) LEAD. AN OVER CURRENT DETECTION (OCD) ALERT WAS OBSERVED ON THE DEVICE. THE PATIENT WAS RESUSCITATED BY EXTERNAL SHOCKS BUT ULTIMATELY PASSED AWAY. Manufacturer Narrative: THE REPORTED EVENT OF NO HIGH VOLTAGE OUTPUT WAS NOT CONFIRMED. THE REPORTED EVENT OF OVER CURRENT DETECTION (OCD) WAS DETECTED IN THE DEVICE IMAGE¿S LOG. HOWEVER, THE CAUSE OF OCD WAS DETERMINED TO BE DUE TO EXTERNAL (NON-DEVICE RELATED) FACTORS.THE DEVICE WAS RECEIVED WITH NORMAL CHARGING AND NORMAL OUTPUTS. ELECTRICAL AND MECHANICAL TESTS PERFORMED INCLUDING LEADS IMPEDANCE, SENSING, PACING, CAPACITOR CHARGING, AND HIGH VOLTAGE (HV) OUTPUT VERIFICATION DID NOT IDENTIFY ANY FUNCTIONAL ISSUES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21946401&pc=NVY | 2017865-2025-51507 | 2025-04-07 04:00 | Death | ABBOTT | 2025-05-01 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Failure to Deliver Shock/Stimulation; Low impedance; Charging Problem | Arrhythmia; Syncope/Fainting | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A CARDIAC ARREST AND SYNCOPE. THE DEVICE FAILED TO DELIVER HIGH VOLTAGE THERAPY DURING EPISODES OF VENTRICULAR FIBRILLATION. LOW IMPEDANCE WAS OBSERVED ON THE RIGHT VENTRICULAR (RV) LEAD. AN OVER CURRENT DETECTION (OCD) ALERT WAS OBSERVED ON THE DEVICE. THE PATIENT WAS RESUSCITATED BY EXTERNAL SHOCKS BUT ULTIMATELY PASSED AWAY. Manufacturer Narrative: THE REPORTED EVENT OF FAILURE TO DELIVER SHOCKS DUE TO LOW HIGH VOLTAGE (HV) IMPEDANCE WAS NOT CONFIRMED. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION AND X-RAY EXAMINATION DID NOT FIND ANY ANOMALIES EXCEPT FOR THE DAMAGES FOUND CONSISTENT WITH PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21946898&pc=MNS | 2518422-2025-105068 | 2025-04-30 04:00 | Death | RESPIRONICS, INC. | 2025-05-01 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21946899&pc=MNS | 2518422-2025-105071 | 2025-04-30 04:00 | Death | RESPIRONICS, INC. | 2025-05-01 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22120413&pc=BZD | 2518422-2025-045896 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-05-31 04:00 | BZD | DREAMSTATION AUTO BIPAP | Degraded | Cancer; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING THROAT CANCER, RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY AND PATIENT HAS PASSED AWAY (DEATH). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22120418&pc=MNS | 2518422-2025-045899 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-31 04:00 | MNS | BIPAP AUTOSV ADV | Degraded | Heart Failure/Congestive Heart Failure; Insufficient Information; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K090539 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CONGESTIVE HEART FAILURE, FLUID RETENTION AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22120604&pc=BZD | 2518422-2025-045851 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-31 04:00 | BZD | DREAMSTATION CPAP PRO | Degraded | Cancer; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, CANCER, DEATH AND GLIOBLASTOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22112246&pc=PNJ | 9612164-2025-02664 | 2025-05-12 04:00 | Death | MEDTRONIC IRELAND | 2025-05-30 04:00 | PNJ | MICRA¿ VR2 | Deformation Due to Compressive Stress; Adverse Event Without Identified Device or Use Problem; Positioning Problem | Bacterial Infection; Respiratory Failure; Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IT WAS REPORTED THAT THERE WAS OUT OF PLANE DEFLECTION ON THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) DELIVERY SYSTEM DURING AN ATTEMPTED IMPLANT. IT WAS NOTED IT WAS ALSO KINKED. IT WAS NOTED THE PATIENT PASSED AWAY THREE DAYS POST THE ATTEMPTED IMPLANT.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IT WAS REPORTED THAT THERE WAS OUT OF PLANE DEFLECTION ON THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) DELIVERY SYSTEM DURING AN ATTEMPTED IMPLANT. IT WAS NOTED IT WAS ALSO KINKED. IT WAS NOTED THE PATIENT PASSED AWAY THREE DAYS POST THE ATTEMPTED IMPLANT. IT WAS REPORTED THAT THERE WAS A SUCCESSFUL IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) ON THE SAME DAY.Event Description: IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS AND HYPER-CARBIC RESPIRATORY FAILURE AND METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS BACTEREMIA Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2VR01-DELSYS] (SERIAL: [SERIAL (B)(6). D9: NO MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION B5 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22112542&pc=BZD | 2518422-2025-045411 | 2023-09-18 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR PRO SERIES C-FLEX + ASSY | Degraded | Cancer; Unspecified Hepatic or Biliary Problem | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER DISEASE/TOXICITY, CANCER, DEATH AND LIVER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22112724&pc=BZD | 2518422-2025-045454 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | BIPAP PRO SERIES ASSY | Degraded | Asthma; Headache; Nausea; Vomiting; Dizziness; Cancer; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA / VOMITING, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, CANCER, DEATH AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22112964&pc=MVK | 3008642652-2025-05319 | 2025-04-11 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-30 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 07:32:55 ON (B)(6) 2025. AT 01:50:00 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT FROM 150-170 BPM WITH CPR/ELECTRODE LEAD FALL OFF. THE DEVICE PROPERLY DETECTED VT. CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALL OFF PREVENTED THE DEVICE FROM TREATING THE PATIENT. THE ELECTRODE BELT WAS DISCONNECTED AT 02:44:51 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22112970&pc=MVK | 3008642652-2025-05321 | 2024-12-27 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-30 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2024 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 07:45:38 ON (B)(6) 2024. AT 15:32:28, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 150 BPM DEGRADING TO ASYSTOLE. THE RHYTHM THEN TRANSITIONS TO FINE VF. THE DEVICE PROPERLY DETECTED VT/VF. HEART RATE AT OR BELOW THE THRESHOLD FOR TREATMENT, FINE VF, RESPONSE BUTTON USE, VARYING AMPLITUDES, MOTION ARTIFACT, AND ELECTRODE LEAD FALL OFF ALL PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE DEVICE WAS SHUT DOWN AT 15:33:32 ON (B)(6) 2024. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22112971&pc=MVK | 3008642652-2025-05320 | 2025-03-08 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-30 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 13:11:13 ON (B)(6) 2025. AT 06:39:34 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. AT 06:41:37, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM SINUS BRADYCARDIA @ 30 BPM WITH CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. THE ELECTRODE BELT WAS DISCONNECTED AT 06:41:49 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22113048&pc=OZD | 1220648-2025-28829 | 2025-05-22 04:00 | Death | ABIOMED, INC. | 2025-05-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Air Embolism; Shock; Cardiac Tamponade; Cardiac Perforation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT DURING INSERTION OF AN IMPELLA CP SMARTASSIST IN A PATIENT, A LEFT VENTRICLE (LV) PERFORATION OCCURRED. DEVICE PLACEMENT WAS CHECKED AND FOUND THE DEVICE TO BE APPROXIMATELY SEVEN CENTIMETERS FROM THE INLET AREA TO THE AORTIC VALVE. IT WAS CONFIRMED THAT THE PIGTAIL WAS PRESSED AGAINST THE LEFT VENTRICULAR WALL. THE 0.018-INCH WIRE WAS PLACED FACING THE POSTERIOR WALL AND THREADED BETWEEN THE HEALTHY CHODARE TENDINEAE. THE LV PERFORATION RESULTED IN CARDIAC TAMPONADE AND SHOCK VITAL SIGNS. FREE WALL WAS REPAIR WAS PERFORMED IN THE OPERATING ROOM. CEREBRAL AIR EMBOLISM OCCURRED AFTERWARDS. THE PERICARDIUM WAS OPENED UNDER IMPELLA DRIVE, AND THE PIGTAIL WAS PROTRUDING, AND IT REPEATEDLY DISLOCATED. THERE WAS NO INSERTION OF A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). AFTER THE REPAIR SURGERY, A TTE WAS PERFORMED TO CHECK THE INSIDE OF THE HEART AND FOUND THAT IT WAS FILLED WITH AIR. AFTER CHECKING THE LIGHT REFLEX, THE PUPILS WERE DILATED AND IT WAS DETERMINED THAT THE PATIENT HAD AN AIR EMBOLISM. THE PATIENT'S OUTCOME IS SERIOUS.Event Description: THE PATIENT OUTCOME WAS LATER SHARED, AS DEATH. THE PATIENT EXPIRED FROM THE HYPOXEMIA DUE TO SHOCK CAUSED BY CARDIAC TAMPONADE AT THE TIME OF LEFT VENTRICULAR PERFORATION, OR AIR EMBOLISM DURING FREE WALL REPAIR SURGERY. THE TRUE CAUSE OF DEATH IS NOT KNOWN. Manufacturer Narrative: THE INVESTIGATION OF VENTRICLE PERFORATION AND EMBOLISM HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. A DATA LOG REVIEW WAS NOT REQUIRED FOR THIS CASE. CLINICAL DETAILS CONFIRM THAT THE PUMP WAS PLACED DEEP INTO THE VENTRICLE AND PERFORATED THE LV DURING DELIVERY. THERE WAS NO INDICATION OF A FILTER LEAK DURING THE CASE RUN. THE HEART WAS FOUND TO BE FILLED WITH AIR AFTER THE LEFT VENTRICLE (LV) REPAIR. THE CAUSE OF THE VENTRICLE ISSUE WAS MOST LIKELY USE RELATED SINCE PUMP WAS INSERTED DEEP INTO LV. THE CAUSE OF THE EMBOLISM ISSUE WAS MOST LIKELY USE RELATED SINCE AIR WAS FOUND AFTER LV WAS REPAIRED AFTER PERFORATION. B2 UPDATED TO REFLECT THE PATIENT OUTCOME DEATH. B5 UPDATED TO INCLUDE PATIENT OUTCOME. H1 TYPE OF REPORT UPDATED TO DEATH. H6 HEALTH EFFECT - IMPACT CODE UPDATED TO INCLUDE DEATH. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION. SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22113202&pc=BZD | 2518422-2025-045559 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR PLUS | Degraded | Cancer | K091319 |  | 1 | Event Description: (B)(4) - ALLEGATION OF PATIENT HARM. Manufacturer Narrative: IN THE PREVIOUS REPORT, THE MANUFACTURER CAPTURED INCORRECT INFORMATION IN GENERAL INFORMATION, BOX B, BOX G AND BOX H. THE FOLLOWING CORRECTION HAS BEEN CORRECTED IN THIS REPORT. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. GENERAL INFORMATION: 510K HAS BEEN CORRECTED. IN BOX B: DESCRIBE EVENT OR PROBLEM HAS BEEN CORRECTED. IN BOX G: 510K HAS BEEN CORRECTED. IN BOX H: TYPE OF REPORTED COMPLAINT HAS BEEN CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22113665&pc=BZD | 2518422-2025-045570 | 2024-01-22 05:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR PRO C-FLEX+ | Degraded | Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, RESPIRATORY ISSUES AND SARCOMA CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22113756&pc=OZD | 1220648-2025-28836 | 2025-05-08 04:00 | Death | ABIOMED, INC. | 2025-05-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hematoma | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THERE WAS A HEMATOMA AFTER THE REPOSITIONING SHEATH WAS INSERTED. THE DOCTOR FORCEFULLY INSERTED THE SHEATH AND THIS IS THOUGHT TO BE THE CAUSE OF THE HEMATOMA. MANUAL PRESSURE WAS ATTEMPTED TO STABILIZE THE HEMATOMA. THIS ATTEMPT WAS UNSUCCESSFUL, AND THE IMPELLA CP WAS REMOVED. AFTER WAITING FOR THE DOCTOR TO DECIDE HOW BEST TO APPROACH FOR ACHIEVING HEMOSTASIS, THEY GOT CONTRALATERAL ACCESS AND DEPLOYED A BALLOON AT THE SITE ACCOMPANIED BY MANUAL EXTERNAL PRESSURE AND SECUREMENT OF THEIR PRIOR PLACED PERCLOSE SUTURES. DUE TO THE DELAYED CONTRALATERAL ACCESS AND BALLOONING, THE PATIENT WENT INTO CARDIAC ARREST AND WAS UNABLE TO BE RESUSCITATED. THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY USE RELATED. THE DOCTOR FORCEFULLY INSERTED THE SHEATH AND THIS IS THOUGHT TO BE THE CAUSE OF THE HEMATOMA. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22113773&pc=BZD | 2518422-2025-045588 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR PLUS SERIES ASSY | Degraded | Cancer | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, ESOPHAGEAL CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22113774&pc=BZD | 2518422-2025-045618 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR PLUS SERIES ASSY | Degraded | Cancer | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING ESOPHAGEAL CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: IN BOX B: ADVERSE EVENT/PRODUCT PROBLEM CHANGED TO BOTH AND DEATH. DATE OF DEATH WAS UPDATED IN THIS REPORT FROM NOT APPLICABLE TO NO INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22113781&pc=NIM | 2183870-2025-00277 | 2025-04-01 04:00 | Death | COVIDIEN | 2025-05-30 04:00 | NIM | PROTEGE RX CAROTID STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Renal Failure; Septic Shock; Gastrointestinal Hemorrhage | P060001 |  | 1 | Event Description: ABSTRACT OBJECTIVE: THE AIM OF THIS STUDY WAS TO COMPARE SHORT- AND MID-TERM OUTCOMES AFTER CAROTID ARTERY STENTING (CAS) RELATED TO STENT CONFIGURATION. METHODS: THIS IS A RETROSPECTIVE STUDY OF ALL CAS INCLUDING TRANSCAROTID AND TRANSFEMORAL APPROACHES, PERFORMED AT OUR INSTITUTION FROM 2015 TO 2024. GROUPS WERE IDENTIFIED ACCORDING TO THE STENT USED AT THE INDEX PROCEDURE BY OPEN-CELL STENTS (OCS) AND CLOSED-CELL STENTS (CCS). SUBGROUP ANALYSIS BY SYMPTOMATIC STATUS WAS ALSO PERFORMED. STROKE, MYOCARDIAL INFARCTION (MI), AND DEATH WERE ANALYZED, INCLUDING THOSE OCCURRING WITHIN 30 DAYS OF PROCEDURE AND ON LAST FOLLOW-UP (14.5 6 15.1 MONTHS). RESTENOSIS AND NEED FOR REINTERVENTION WERE ALSO ASSESSED. RESULTS: TWO HUNDRED FORTY-SIX PATIENTS WERE INCLUDED IN OUR STUDY COHORT, WITH 128 OCS (ENROUTE STENT [(B)(6) MEDICAL], 76; PROTÉGÉ STENT [MEDTRONIC], 34; PRECISE STENT [CORDIS], 15; ACCULINK STENT [ABBOTT], 3), AND 118 CCS (WALLSTENT [BOSTON SCIENTIFIC], 118). BASELINE CHARACTERISTICS AND COMORBIDITIES OF THE OVERALL COHORT SHOWED DIFFERENCES IN BODY MASS INDEX (CCS: 28.57 6 6.19 KG/M2, OCS: 29.79 6 5.98 KG/M2; P ¼ .018) AND SYMPTOMATIC STATUS (CCS: 78; 66.1%, OCS: 34; 26.6%; P .001). IN SUBGROUP ANALYSIS WITHIN SYMPTOMATIC AND ASYMPTOMATIC PATIENTS, THESE DIFFERENCES DID NOT PERSIST. OF NOTE, IN SYMPTOMATIC PATIENTS, A HIGHER RATE OF DIABETICS RECEIVED CCS (33.3%) COMPARED WITH OCS (14.7%; P ¼ .043). COMPOSITE OUTCOMES OF STROKE, MI, AND DEATH WERE NO DIFFERENT AT 30 DAYS AND ON MID-TERM FOLLOW-UP (14.5 6 15.1 MONTHS). ONLY MI AT FOLLOW-UP IN THE OVERALL COHORT WAS SIGNIFICANTLY HIGHER IN THE OCS GROUP (4.7%) COMPARED WITH CCS (0%; P ¼ .03); HOWEVER, THIS DIFFERENCE DID NOT PERSIST IN SUBGROUP ANALYSIS AMONG SYMPTOMATIC AND ASYMPTOMATIC PATIENTS. RESTENOSIS (CCS: 16; 13.6%, OCS: 2; 1.6%; P .001) AND REINTERVENTIONS (CCS: 13; 11%, OCS: 0; 0%; P .001) WERE HIGHER IN THE CCS GROUP. FURTHERMORE, THIS DIFFERENCE PERSISTED FOR BOTH OUTCOMES IN SYMPTOMATIC AND ASYMPTOMATIC SUBGROUP ANALYSIS. MULTIVARIATE ANALYSIS SHOWED INCREASED RISK FOR RESTENOSIS IN CCS (ADJUSTED HAZARD RATIO, 10.28; 95% CONFIDENCE INTERVAL, 2.25-47.09; P ¼ .003). CONCLUSIONS: NO DIFFERENCE IN SHORT- AND MID-TERM OUTCOMES WAS PRESENT IN EITHER CCS OR OCS REGARDING STROKE, MI, OR DEATH. ON MID-TERM FOLLOW-UP, CCS PATIENTS HAD A HIGHER RATE OF RESTENOSIS AND REINTERVENTION, AND THIS DIFFERENCE PERSISTED IN BOTH SYMPTOMATIC AND ASYMPTOMATIC SUBGROUPS. Manufacturer Narrative: JOURNAL TITLE COMPARISON OF RESTENOSIS AND REINTERVENTION RATES IN CLOSED CELL STENTS COMPARED TO OPEN CELL STENTS AFTER CAROTID ARTERY STENTING 2025 SOCIETY FOR VASCULAR SURGERY. HTTPS://DOI.ORG/10.1016/J.JVS.2025.03.199 A2 AVERAGE AGE A3 MAJORITY GENDER B3 DATE OF PUBLICATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22114096&pc=QFG | 3013756811-2025-99862 | 2025-05-19 04:00 | Death | TANDEM DIABETES CARE | 2025-05-30 04:00 | QFG | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K203234 |  | 1 | Event Description: VOLUNTARY MEDWATCH RECEIVED ON 05/19/2025 REPORTED: IN ACCORDANCE WITH 21 CFR 803.22 (B)(2), WE ARE NOTIFYING YOU THAT ON (B)(6) 2025 THE PATIENTS WIFE REPORTED THE PATIENT DIED DUE TO A FAULTY INSULIN PUMP. NO FURTHER INFORMATION WAS PROVIDED. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). REF REPORT: MW5170109. IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY "DUE TO A FAULTY INSULIN PUMP." DUE TO THE ANONYMOUS NATURE OF THE REPORT, TANDEM CUSTOMER TECHNICAL SERVICE UNABLE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE ISSUE OR VERIFY AN EVENT DATE. NO ADDITIONAL PATIENT OR EVENT INFORMATION WAS AVAILABLE. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. RELATED MEDWATCH REPORT # MW5170108/MW5170109.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22114293&pc=OJX | 2649622-2025-14761 | 2024-12-14 05:00 | Death | MPRI | 2025-05-30 04:00 | OJX | ATTAIN PERFORMA MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Endocarditis; Hypoxia; Pneumonia; Sepsis; Septic Shock; Loss of consciousness; Respiratory Failure | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL WITH ACUTE HYPOXEMIC AND HYPERCARBIC RESPIRATORY FAILURE. THE PATIENT UNDERWENT EMERGENT INTUBATION AFTER BECOMING UNRESPONSIVE AND OXYGEN SATURATION DROPPED WHILE ON BILEVEL POSITIVE AIRWAY PRESSURE. TESTING REVEALED SEPSIS/SEPTIC SHOCK AND EXTENSIVE PNEUMONIA IN THE LEFT LOWER LUNG INVOLVING LINGULA WITH STREPTOCOCCUS PNEUMONIAE. AN ECHOCARDIOGRAM REVIEWED ¿SIGNIFICANT TRICUSPID VALVE VEGETATION,¿ WHICH WAS MOBILE AND IN CLOSE PROXIMITY TO THE LEADS. ENDOCARDITIS WAS THE CAUSE OF THE VEGETATION. A SEPTIC EMBOLISM WAS THOUGHT TO BE THE CAUSE OF THE PNEUMONIA. THE PATIENT WAS TRANSFERRED TO AN OUTSIDE CARDIAC SURGERY FACILITY FOR A SYSTEM EXTRACTION AND TRICUSPID VALVE REPLACEMENT, HOWEVER, EXPIRED 12 DAYS LATER. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22114304&pc=DTB | 2649622-2025-14762 | 2024-12-14 05:00 | Death | MPRI | 2025-05-30 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Endocarditis; Hypoxia; Pneumonia; Sepsis; Septic Shock; Loss of consciousness; Respiratory Failure | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL WITH ACUTE HYPOXEMIC AND HYPERCARBIC RESPIRATORY FAILURE. THE PATIENT UNDERWENT EMERGENT INTUBATION AFTER BECOMING UNRESPONSIVE AND OXYGEN SATURATION DROPPED WHILE ON BILEVEL POSITIVE AIRWAY PRESSURE. TESTING REVEALED SEPSIS/SEPTIC SHOCK AND EXTENSIVE PNEUMONIA IN THE LEFT LOWER LUNG INVOLVING LINGULA WITH STREPTOCOCCUS PNEUMONIAE. AN ECHOCARDIOGRAM REVIEWED ¿SIGNIFICANT TRICUSPID VALVE VEGETATION,¿ WHICH WAS MOBILE AND IN CLOSE PROXIMITY TO THE LEADS. ENDOCARDITIS WAS THE CAUSE OF THE VEGETATION. A SEPTIC EMBOLISM WAS THOUGHT TO BE THE CAUSE OF THE PNEUMONIA. THE PATIENT WAS TRANSFERRED TO AN OUTSIDE CARDIAC SURGERY FACILITY FOR A SYSTEM EXTRACTION AND TRICUSPID VALVE REPLACEMENT, HOWEVER, EXPIRED 12 DAYS LATER. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22114319&pc=DTB | 2649622-2025-14764 | 2024-12-14 05:00 | Death | MPRI | 2025-05-30 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Endocarditis; Hypoxia; Pneumonia; Sepsis; Septic Shock; Loss of consciousness; Respiratory Failure | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL WITH ACUTE HYPOXEMIC AND HYPERCARBIC RESPIRATORY FAILURE. THE PATIENT UNDERWENT EMERGENT INTUBATION AFTER BECOMING UNRESPONSIVE AND OXYGEN SATURATION DROPPED WHILE ON BILEVEL POSITIVE AIRWAY PRESSURE. TESTING REVEALED SEPSIS/SEPTIC SHOCK AND EXTENSIVE PNEUMONIA IN THE LEFT LOWER LUNG INVOLVING LINGULA WITH STREPTOCOCCUS PNEUMONIAE. AN ECHOCARDIOGRAM REVIEWED ¿SIGNIFICANT TRICUSPID VALVE VEGETATION,¿ WHICH WAS MOBILE AND IN CLOSE PROXIMITY TO THE LEADS. ENDOCARDITIS WAS THE CAUSE OF THE VEGETATION. A SEPTIC EMBOLISM WAS THOUGHT TO BE THE CAUSE OF THE PNEUMONIA. THE PATIENT WAS TRANSFERRED TO AN OUTSIDE CARDIAC SURGERY FACILITY FOR A SYSTEM EXTRACTION AND TRICUSPID VALVE REPLACEMENT, HOWEVER, EXPIRED 12 DAYS LATER. THE PATIENT IS A PARTICIPANT IN A CLINICAL STUDY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22114542&pc=FKX | 0008030665-2025-01219 | 2025-05-17 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-05-30 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K222318 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT REPORTED THE PATIENT PASSED AWAY. IT COULD NOT BE CONFIRMED IF THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF THE EVENT. ADDITIONAL DETAILS WERE OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT¿S SON REPORTED THAT THE PATIENT WAS FOUND UNRESPONSIVE ON (B)(6) 2025. THE SON WAS NOT AT THE PATIENT¿S HOME AT THE TIME OF THE EVENT. 911 WAS CALLED BY SOMEONE IN THE HOME. EMERGENCY MEDICAL SERVICES (EMS) RESPONDED TO THE HOME AND RESUSCITATIVE EFFORTS WERE ADMINISTERED, BUT NOT SUCCESSFUL. THE PATIENT¿S SON REPORTED THAT HE DID NOT ENTER THE ROOM WHERE THE PATIENT WAS FOUND. HE COULD NOT CONFIRM IF THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF THE EVENT. THE CAUSE OF DEATH IS UNKNOWN. THE DEATH WAS REPORTEDLY UNEXPECTED (NO HEALTH ISSUES REPORTED). NO ISSUES WERE REPORTED WITH THE LIBERTY SELECT CYCLER OR PD THERAPY PRIOR TO THE DEATH.Event Description: IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT REPORTED THE PATIENT PASSED AWAY. IT COULD NOT BE CONFIRMED IF THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF THE EVENT. ADDITIONAL DETAILS WERE OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT'S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT'S SON REPORTED THAT THE PATIENT WAS FOUND UNRESPONSIVE ON (B)(6) 2025. THE SON WAS NOT AT THE PATIENT¿S HOME AT THE TIME OF THE EVENT. 911 WAS CALLED BY SOMEONE IN THE HOME. EMERGENCY MEDICAL SERVICES (EMS) RESPONDED TO THE HOME AND RESUSCITATIVE EFFORTS WERE ADMINISTERED, BUT NOT SUCCESSFUL. THE PATIENT¿S SON REPORTED THAT HE DID NOT ENTER THE ROOM WHERE THE PATIENT WAS FOUND. HE COULD NOT CONFIRM IF THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF THE EVENT. THE CAUSE OF DEATH IS UNKNOWN. THE DEATH WAS REPORTEDLY UNEXPECTED (NO HEALTH ISSUES REPORTED). NO ISSUES WERE REPORTED WITH THE LIBERTY SELECT CYCLER OR PD THERAPY PRIOR TO THE DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED IN D9, G4, AND H3. PLANT INVESTIGATION: A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED SIGNS OF PHYSICAL DAMAGE: TOUCH SCREEN MISALIGNED. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE AN (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST PASSED. SYSTEM AIR LEAK TEST PASSED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: THERE IS A POSSIBLE TEMPORAL RELATIONSHIP BETWEEN THE PERITONEAL DIALYSIS (PD) AND USE OF THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH PRECEDED BY UNRESPONSIVENESS. HOWEVER, IT IS NOT CONFIRMED IF THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF THE ADVERSE EVENT. THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE PATIENT DEATH AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. ALTHOUGH NO CAUSE OF DEATH WAS CONFIRMED, PATIENTS ON DIALYSIS HAVE A SUBSTANTIALLY HIGHER MULTIVARIABLE-ADJUSTED MORTALITY THAN PATIENTS NOT ON DIALYSIS WHO HAVE CANCER, DIABETES, OR CARDIOVASCULAR DISEASE. A LARGE NUMBER OF RISK FACTORS THAT ARE UNRELATED TO THE DIALYSIS PROCEDURE HAVE BEEN ASSOCIATED WITH DECREASED SURVIVAL AMONG PATIENTS ON DIALYSIS. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22114620&pc=BZD | 2518422-2025-045669 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR PRO C-FLEX+ | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22114822&pc=OZD | 1220648-2025-28845 | 2025-05-07 04:00 | Death | ABIOMED, INC. | 2025-05-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Sepsis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT WAS TAKEN TO HAVE A SEPSIS WORKUP. THE LOCATION OF THE SEPSIS IS UNKNOWN AND THEREFORE IT IS UNKNOWN IF THE IMPELLA WAS RELATED TO THE SEPSIS. ADDITIONALLY, THE PATIENT HAD ARRHYTHMIAS. THE P LEVEL WAS SWITCHED TO P 8. SUPPORT WAS CONTINUED. CARE WAS WITHDRAWN TWO DAYS LATER AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RETURNED, AND THE INVESTIGATION HAS BEEN COMPLETED. AS THE NECESSARY CLINICAL INFORMATION WAS NOT PROVIDED AND THE DEVICE WAS NOT RETURNED, THE ROOT CAUSE OF THE ARRYTHMIA WAS NOT DETERMINED. EVIDENCE HAS NOT BEEN PROVIDED, THE ROOT CAUSE OF THE SEPSIS CANNOT BE DETERMINED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22114825&pc=BZD | 2518422-2025-045531 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR AUTO A-FLEX | Degraded | Headache; Dizziness; Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, LUNG DISEASE, DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22114829&pc=BZD | 2518422-2025-045741 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR PRO C-FLEX + | Degraded | Stroke/CVA; Dyspnea; Headache; Inflammation; Nausea; Pneumonia; Vomiting; Dizziness; Unspecified Respiratory Problem; Insufficient Information | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, DEATH, SARCOIDOSIS, STROKE, PNEUMONIA, DIFFICULTY BREATHING AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22114940&pc=BZD | 2518422-2025-045758 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR PRO C-FLEX + | Degraded | Inflammation; Cancer | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING INFLAMMATORY RESPONSE, CANCER, DEATH, LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22115131&pc=NIQ | 2124215-2025-35245 | 2025-04-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-30 04:00 | NIQ | SYNERGY XD | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest; Ventricular Fibrillation |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT WAS ADMITTED TO THE HOSPITAL WITH CHEST HEAVINESS AND PROBLEMS BREATHING. A SYNERGY XD MR OUS 3.50 X 48MM WAS SELECTED FOR TREATMENT OF THE TARGET LESION, WHICH WAS LOCATED IN THE PROXIMAL TO MID LEFT ANTERIOR DESCENDING ARTERY (LAD). A NON-BOSTON SCIENTIFIC STENT WAS PLACED IN THE LEFT CIRCUMFLEX ARTERY (LCX). DURING THE PROCEDURE THE PATIENT EXPERIENCED EPISODES OF VENTRICULAR FIBRILLATION AND VENTRICULAR TACHYCARDIA, BUT THE PROCEDURE WAS ABLE TO BE COMPLETED SUCCESSFULLY. AFTER THIS, THE PATIENT WAS MOVED TO A CORONARY CARE UNIT. AFTERWARDS, THE PATIENT EXPERIENCED SUDDEN CARDIAC ARREST. AN ELECTROCARDIOGRAM SHOWED VENTRICULAR FIBRILLATION, AND A DIRECT CURRENT SHOCK WAS GIVEN. CARDIOPULMONARY RESUSCITATION WAS PERFORMED FOR 45 MINUTES; HOWEVER, THE PATIENT HAD DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22115146&pc=MKJ | 1220908-2025-02041 | 2025-05-16 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-05-30 04:00 | MKJ | X SERIES | Unable to Obtain Readings | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A 77-YEAR-OLD MALE PATIENT, THE DEVICE WAS UNABLE TO OBTAIN AN ECG SIGNAL VIA ELECTRODE PADS. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER SET OF PADS TO CONTINUE TREATING THE PATIENT; HOWEVER, THE SAME ISSUE ENSUED. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: ZOLL MEDICAL CANADA EVALUATED THE DEVICE AND MULTIFUNCTION CABLE (MFC) AND THE CUSTOMER'S REPORT WAS NOT REPLICATED OR CONFIRMED. THE DEVICE WAS PUT THROUGH EXTENSIVE TESTING USING THE RETURNED MFC, A KNOWN GOOD ROC ADAPTOR, AND PADS WITHOUT DUPLICATING THE REPORT. THE MFC CONNECTOR WAS REPLACED AS A PRECAUTION. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. THE LOG SHOWED AT THE START OF THE EVENT, THE OPERATOR HAD ATTACHED CPR PADS TO THE DEVICE. THE DEVICE PROMPTED ATTACH PADS REPEATEDLY WITH NO ECG SIGNAL DISPLAYED, INDICATING THAT IT DID NOT DETECT A VALID PATIENT IMPEDANCE. THE LOG DID NOT RECORD ANY PATIENT IMPEDANCE, CPR WAVEFORM, OR FEEDBACK FROM CPR PADS DURING THIS TIME. THIS IS NORMAL BEHAVIOR IF THE DEVICE IS NOT ABLE TO DETECT A VALID PATIENT IMPEDANCE AT ANY POINT. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE. Manufacturer Narrative: JUSTIFICATION FOR NO UDI: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22115519&pc=BZD | 2518422-2025-106212 | 2025-05-06 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | DREAMSTATION | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION STATING THE PATIENT HAS PASSED AWAY DUE TO LUNG CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22115957&pc=BZD | 2518422-2025-045814 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR AUTO A-FLEX | Degraded | Cancer; Lymphoma | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND NON-HODGKIN'S LYMPHOMA (FOLLICULAR). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22115969&pc=PYX | 1220648-2025-28853 | 2025-05-07 04:00 | Death | ABIOMED, INC. | 2025-05-30 04:00 | PYX | IMPELLA RP FLEX | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Hypovolemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT ON IMPELLA RP FLEX SUPPORT WOULD NOT STOP BLEEDING FROM ALL SITES. LEADING TO HYPOVOLEMIA AND FINALLY RESULTING IN THE PASSING OF THE PATIENT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA RP FLEX WITH SMARTASSIST INSTRUCTIONS FOR USE & CLINICAL REFERENCE MANUAL SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP FLEX SMARTASSIST SYSTEM CATHETER. USE OF THE REPOSITIONING SHEATH AND THE 23 FR PEEL-AWAY INTRODUCER: ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE INTERNAL JUGULAR VEIN. CLOSE AND DRESS THE WOUND.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22116198&pc=BZD | 2518422-2025-045830 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR AUTO A-FLEX | Degraded | Cancer; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, UNSPECIFIED KIDNEY PROBLEM AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22116419&pc=OZD | 1220648-2025-28856 | 2025-04-30 04:00 | Death | ABIOMED, INC. | 2025-05-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemolysis; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT IMPLANTED WITH AN IMPELLA CP DEVELOPED HEMOLYSIS AND ECTOPY WHICH RESOLVED WITHOUT INTERVENTION ON THE FOLLOWING DAY. THE PATIENT ALSO EXHIBITED CONSUMPTIVE COAGULOPATHY DUE TO BLEEDING/DRAINAGE FROM A STERNOTOMY. MULTIPLE BLOOD PRODUCTS WERE GIVEN, AND A WASHOUT WAS PERFORMED. TWO DAYS LATER THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5 AND THE PATIENT EXPIRED. ABIOMED'S MEDICAL SAFETY OFFICER FOUND THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST SYSTEM LISTS VENTRICULAR ARRYTHMIA AS AN ADVERSE EVENT THAT OCCURRED WHILE IN USE WITH THE IMPELLA DEVICE. IT STATES THAT "ONLY A FRACTION OF THESE RATES WERE ATTRIBUTED TO THE IMPELLA DEVICE AND THE EVENTS RESOLVED WITHOUT RESIDUAL EFFECT IN MOST OF THE CASES, UNLESS THE EVENT OF DEATH OCCURRED. OVERALL, THE RESULTS DID NOT SHOW ANY EVIDENCE OF INCREASED MORBIDITY ASSOCIATED WITH THE IMPELLA SUPPORT IN CARDIOMYOPATHY, MYOCARDITIS, AND PERIPARTUM CARDIOMYOPATHY (PPCM) PATIENTS." Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT, VASCULAR DAMAGE - PERIPHERAL VESSEL, AND HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE CAUSE OF THE VASCULAR DAMAGE - PERIPHERAL VESSEL WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS. HEMOLYSIS: DATA LOGS WERE RECOVERED. DATA LOGS SHOW POSITIONING ISSUES THAT OCCUR ON THE DAY HEMOLYSIS WAS NOTED, WITH IMPELLA POSITION IN VENTRICLE AND SUCTION ALARMS. THE CAUSE OF THE HEMOLYSIS WAS MOST LIKELY IMPROPER POSITIONING OF THE DEVICE DUE TO DATA LOG REVIEW SHOWING PUMP OUT OF POSITION AND CLINICAL DETAILS STATING PUMP REQUIRED REPOSITIONING THAT BOTH OCCUR ON DAY HEMOLYSIS WAS NOTED WHICH RESOLVED WITHIN 24 HOURS. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED B2 REQUIRED INTERVENTION SELECTION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28856. H6 HEALTH EFFECT - IMPACT CODE 1802 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28856.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22116866&pc=OZD | 1220648-2025-28858 | 2025-05-04 04:00 | Death | ABIOMED, INC. | 2025-05-30 04:00 | OZD | IMPELLA CP | Failure to Advance | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE INSERTION OF AN IMPELLA CP IN A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS UNSUCCESSFUL. THERE WAS A DELAY IN STARTING SUPPORT. ANOTHER IMPELLA CP WOULD BE USED. HOWEVER, THE PATIENT WOULD SUBSEQUENTLY EXPIRE. AT THE START OF THE PROCEDURE, THE PATIENT RECEIVED CONTRAST TO REVIEW VESSELS. EJECTION FRACTION AT SUCH TIME WAS APPROXIMATELY 45%. THE PATIENT WENT INTO FLASH PULMONARY EDEMA REQUIRING EMERGENT INTUBATION AND POST INTUBATION. EJECTION FRACTION WAS LESS THAN 20%. THE INITIAL IMPELLA CP PLACEMENT WAS ATTEMPTED. HOWEVER, THERE WAS AN ISSUE DELIVERING THE IMPELLA CATHETER OVER A 0.018 GUIDEWIRE. IT COULD NOT BE DETERMINED IF THE ISSUE WAS WITH THE GUIDEWIRE OR THE IMPELLA CATHETER. CONSEQUENTLY, ANOTHER IMPELLA CP DEVICE WAS USED AND PLACED PRE PERCUTANEOUS CORONARY INTERVENTION (PCI) DUE TO THE SEVERITY OF GRAFTS AND REDUCED EJECTION FRACTION. THE PCI (DRUG-ELUTING STENTING) WAS COMPLETED, AND THE PATIENT'S MEAN ARTERIAL PRESSURE REMAINED GREATER THAN 65 THROUGHOUT CASE. THE DECISION WAS MADE TO LEAVE THE IMPELLA CP DEVICE IN PLACE OVERNIGHT AND THE PATIENT WAS TRANSFERRED TO THE UNIT ON MECHANICAL CIRCULATORY SUPPORT. HOWEVER, APPROXIMATELY FIVE HOURS AFTER THE START OF THE SUPPORT, THE PATIENT EXPIRED. NO FURTHER DETAILS WERE NOTED. Manufacturer Narrative: MEDICAL SAFETY REVIEWED THE EVENT AND NOTED LIMITED DETAILS SURROUNDING THE DEMISE OUTCOME, AND THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP PUMP 1 AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME. AN ISSUE WITH DELIVERING PUMP 1 OVER THE GUIDEWIRE; THE DEVICE WAS REPLACED AND PATIENT EXPIRED APPROXIMATELY FIVE HOURS AFTER THE START OF IMPELLA MECHANICAL CIRCULATORY SUPPORT ON THE REPLACEMENT DEVICE. BASED ON THE INFORMATION AT HAND, THE IMPELLA CP PUMP 2 APPEARS TO HAVE PROVIDED SUPPORT AS INTENDED. THERE IS NO INFORMATION THAT DIRECTLY OR INDIRECTLY INDICATES OTHERWISE. THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22116934&pc=NVN | 2017865-2025-75117 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-30 04:00 | NVN | TENDRIL MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FOR AN UNKNOWN REASON. THERE WAS NO ALLEGATION THAT THE DEATH WAS RELATED TO THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22116936&pc=NIK | 2017865-2025-75114 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-30 04:00 | NIK | GALLANT DR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FOR AN UNKNOWN REASON. THERE WAS NO ALLEGATION THAT THE DEATH WAS RELATED TO THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22116945&pc=NVY | 2017865-2025-75119 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-30 04:00 | NVY | OPTISURE ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FOR AN UNKNOWN REASON. THERE WAS NO ALLEGATION THAT THE DEATH WAS RELATED TO THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. VISUAL EXAMINATION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING OF THE LEAD DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22117166&pc=BZD | 2518422-2025-045533 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR PLUS | Degraded | Cancer | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22117367&pc=BZD | 2518422-2025-045864 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR PLUS | Degraded | Asthma; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22117551&pc=BZD | 2518422-2025-045867 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Bacterial Infection; Cancer; Lymphoma; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CANCER NON-HODGKINS LYMPHOMA, BACTERIAL LUNG INFECTION AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22117603&pc=DSQ | 2916596-2025-03305 | 2025-05-08 04:00 | Death | THORATEC CORPORATION | 2025-05-30 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Increased Pump Speed; Adverse Event Without Identified Device or Use Problem | Aneurysm; Hemorrhage/Bleeding; Pneumonia; Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE CLINIC STATING THE SYSTEM CONTROLLER WAS READING ABNORMAL NUMBERS AT ONE POINT IN THE PREVIOUS WEEK, 9000RPM, HIGH POWER AND HIGH PULSATILITY INDEX. THE PATIENT STATED THAT THIS HAS NOT HAPPENED SINCE. THE EVENT LOG FILE WAS FULL OF PI EVENTS ON (B)(6) 2025. THE DATA FROM THE PREVIOUS WEEK HAD ALREADY BEEN OVERWRITTEN BY RECENT PI EVENTS. THE CAUSE OF THE ABNORMAL READINGS WAS NOT DETERMINED, AND THE PATIENT HAD NO MORE OF THOSE EVENTS. THE PATIENT HAD NO SYMPTOMS AT THE TIME OF THE READINGS. THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO ACUTE RESPIRATORY FAILURE SECONDARY TO PNEUMONIA AND INTERMURAL AND PERIAORTIC HEMORRHAGE IN THE THORACIC ANEURYSM. THE DEATH WAS NOT THOUGHT TO BE RELATED TO THE DEVICE AND THE DEVICE OPERATED AS EXPECTED.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT WAS ADMITTED ON (B)(6) 2025, AND COMPUTED TOMOGRAPHY OF THE CHEST AND CHEST X RAY WERE PERFORMED. THE PATIENT WAS THEN PLACED ON INTRAVENOUS CEFEPIME, CEFTRIAXONE, AND CULTURES WERE NEGATIVE. THE PATIENT'S RESPIRATORY CONDITION CONTINUED TO DECLINE AND WAS PLACED ON COMFORT CARE ON (B)(6) 2025. THEIR CODE STATUS CHANGED TO NO CODE AND THEIR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR WAS DEACTIVATED. THE PATIENT PASSED AWAY ON (B)(6) 2025 AND THE DEVICE WAS TURNED OFF. NO AUTOPSY WAS PERFORMED. Manufacturer Narrative: SECTION B1: TYPE OF REPORT CORRECTED. SECTION H6: MEDICAL DEVICE PROBLEM CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILES COULD NOT CONFIRM THE REPORTED ELEVATIONS IN POWER AND PULSATILITY INDEX (PI), AS THE PUMP APPEARED TO BE FUNCTIONING AS INTENDED AT THE FIXED SPEED. NO UNUSUAL EVENTS OR ALARMS WERE NOTED. ADDITIONALLY, A SPECIFIC CAUSE FOR THE REPORTED EVENTS, AS WELL AS A DIRECT CORRELATION TO HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS INVESTIGATION. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE, INFECTION, BLEEDING, AND DEATH THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 1 ALSO PROVIDES AN EXPLANATION OF ALL PUMP PARAMETERS, INCLUDING PUMP POWER AND PULSATILITY INDEX. THIS SECTION EXPLAINS THAT PUMP POWER IS A DIRECT MEASUREMENT OF MOTOR VOLTAGE AND CURRENT. THEREFORE, CHANGES IN PUMP SPEED, FLOW, OR PHYSIOLOGICAL DEMAND CAN AFFECT PUMP POWER. THIS SECTION ALSO EXPLAINS THAT DEVICE FLOW AND POWER GENERALLY RETAIN A LINEAR RELATIONSHIP AT A GIVEN SPEED; HOWEVER, WHILE POWER IS DIRECTLY MEASURED BY THE SYSTEM CONTROLLER, THE REPORTED FLOW IS ESTIMATED BASED ON POWER. SECTION 1 ALSO NOTES THAT ANY INCREASE IN POWER NOT RELATED TO INCREASED FLOW CAUSES ERRONEOUSLY HIGH FLOW READINGS. SECTIONS 1 AND 4 OF THE IFU ALSO STATE THAT THE PI CALCULATION REPRESENTS CARDIAC PULSATILITY, WITH VALUES TYPICALLY RANGING FROM 1 TO 10. GENERALLY, THE MAGNITUDE OF THE PI VALUE IS RELATED TO THE AMOUNT OF ASSISTANCE PROVIDED BY THE PUMP. HIGHER VALUES INDICATE MORE VENTRICULAR FILLING AND HIGHER PULSATILITY (I.E. THE PUMP IS PROVIDING LESS SUPPORT TO THE LEFT VENTRICLE). SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. FURTHERMORE, SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. THIS SECTION ALSO EXPLAINS THAT PUMP PERFORMANCE IS SENSITIVE TO CHANGES IN SYSTEMIC VASCULAR RESISTANCE AND LEFT VENTRICULAR FILLING. SECTION 6 ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INR RANGE, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22117615&pc=LWS | 2649622-2025-14808 | 2025-05-10 04:00 | Death | MPRI | 2025-05-30 04:00 | LWS | SPRINT FIDELIS | Fracture; High impedance; Over-Sensing; Impedance Problem | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED WITH DEVICE TONING WHICH WAS A LEAD INTEGRITY ALERT (LIA) DUE TO UNDEFINED HIGH PACING IMPEDANCE AND SENSING INTEGRITY COUNTER (SIC) WITH ONE THOUSAND THREE HUNDRED AND TWELVE SHORT V-V INTERVALS IN ADDITION TO ONE HUNDRED FIFTY-TWO HIGH RATE NON-SUSTAINED (HRNS) EPISODES AND ONE HUNDRED FORTY-ONE NON-SUSTAINED VENTRICULAR TACHYCARDIA (NSVT) EPISODES. THE LEAD ALSO EXHIBITED OUT OF RANGE (OOR) SPIKED PACING IMPEDANCE. A FRACTURE WAS SUSPECTED IN THE LOW VOLTAGE PORTION AND REPROGRAMMING WAS PERFORMED. IT WAS FURTHER REPORTED THAT THE LEAD WAS EXPLANTED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT IMMEDIATELY AFTER THE RV LEAD EXTRACTION, A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) SHOWED THE RIGHT VENTRICLE WAS NOT CONTRACTING. THERE WERE NO SIGNS OF PERICARDIAL EFFUSION OR TAMPONADE. THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO), INTUBATED, A TEMPORARY MECHANICAL HEART PUMP WAS IMPLANTED, AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). THE NEXT DAY, LIFE SUPPORT MEASURES WERE DISCONTINUED AND THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. THE ANALYST NOTED A PROXIMAL SEGMENT OF THE LEAD 12 CM LONG WAS RETURNED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22117743&pc=BZD | 2518422-2025-045873 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | BIPAP PRO | Degraded | Asthma; Cardiac Arrest; Ischemia Stroke; Unspecified Respiratory Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), DEATH, HEART ATTACK, STROKE AND OTHER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22117745&pc=BZD | 2518422-2025-045866 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR PLUS | Degraded | Cardiac Arrest; Cancer; Unspecified Hepatic or Biliary Problem; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY, CANCER, DEATH, OTHER, LIVER AND LUNG CANCER AND HEART ATTACK. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22117799&pc=LGW | 3008514029-2025-00434 | 2025-05-15 04:00 | Death | NEVRO CORP. | 2025-05-30 04:00 | LGW | SENZA | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Sepsis; Insufficient Information | P130022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. AN HCP ASSESSMENT REGARDING THE CAUSE OF DEATH WAS NOT AVAILABLE. THERE WERE NO REPORTS OF DEVICE-RELATED ISSUES FROM THE PATIENT PRIOR TO THE PASSING. Manufacturer Narrative: A REVIEW OF THE DEVICE'S MANUFACTURING AND STERILIZATION RECORDS WAS COMPLETED, INCLUDING BATCH HISTORY, DEVICE HISTORY, AND STERILIZATION RECORDS; THERE WERE NO DEVIATIONS OR NONCONFORMITIES ASSOCIATED WITH THE REPORTED EVENT. THE DEVICE WAS NOT RETURNED FOR EVALUATION, WHICH LIMITED THE ABILITY TO CONDUCT ANY PHYSICAL, FUNCTIONAL, AND/OR ROOT CAUSE ANALYSIS. B2: THE EXACT DATE FOR THE DEATH IS UNKNOWN. NEVRO SUBMITS THIS REPORT IN COMPLIANCE WITH FDA¿S MEDICAL DEVICE REPORTING REGULATIONS UNDER 21 CFR PART 803. NEVRO HAS COMPLIED WITH REGULATORY INVESTIGATION REQUIREMENTS AND IS SUBMITTING ALL INFORMATION THAT IS REASONABLY KNOWN TO US AT THIS TIME. HOWEVER, WE MAY NOT HAVE BEEN ABLE TO CONFIRM THIS INFORMATION OR COMPLETE THE INVESTIGATION WITHIN THE TIMEFRAME FOR FILING THIS REPORT. WE MAY HAVE GIVEN NO RESPONSE OR AN INCOMPLETE RESPONSE TO CERTAIN QUESTIONS BECAUSE WE DO NOT CURRENTLY HAVE INFORMATION AVAILABLE TO PROVIDE A COMPLETE RESPONSE. IF WE LATER OBTAIN ANY REQUIRED INFORMATION THAT WAS NOT AVAILABLE AT THE TIME OF THIS INITIAL REPORT, WE WILL SUBMIT A SUPPLEMENTAL REPORT. THIS REPORT IS NOT AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT WAS DEFECTIVE, THAT IT MALFUNCTIONED, OR THAT IT CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THIS REPORT. WE MAY CONCLUDE THAT THE DEVICE HAD NO DEFECT, DID NOT MALFUNCTION, OR DID NOT CAUSE OR CONTRIBUTE TO A REPORTABLE EVENT. SOME OF THE ITEMS ON THIS FORM INCLUDE FORCED-CHOICE TERMS USED BY THE FDA FOR REPORTING PURPOSES THAT DO NOT NECESSARILY REFLECT NEVRO¿S CONCLUSIONS ABOUT THE CAUSES OR NATURE OF THE EVENT. THIS STATEMENT SHOULD BE INCLUDED WITH ANY FREEDOM OF INFORMATION ACT RESPONSE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22117804&pc=DSQ | 2916596-2025-03463 | 2021-01-01 05:00 | Death | THORATEC CORPORATION | 2025-05-30 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Aortic Valve Insufficiency/ Regurgitation; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED ¿TRANSCATHETER AORTIC VALVE REPLACEMENT FOR AORTIC REGURGITATION IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES: AN INSTITUTIONAL EXPERIENCE?¿ IDENTIFYING THAT HEARTMATE 3 (HM3) MAY BE RELATED TO AORTIC INSUFFICIENCY/REGURGITATION. THIS OBSERVATIONAL STUDY INCLUDED 348 LVAD PATIENTS IMPLANTED FROM JANUARY 2014 TO OCTOBER 2024. OF THOSE 348, 7(2%) PATIENTS RECEIVED A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) FOR SIGNIFICANT AORTIC REGURGITATION (AR) AND WERE INCLUDED 5 WITH HM3 (PATIENTS 2 TO 6). ALL PATIENTS WERE DISCHARGED ALIVE; HOWEVER, 2 PATIENTS DIED WITHIN 3 MONTHS AFTER THE PROCEDURE, AND 1 WAS READMITTED FOR ACUTE HEART FAILURE. THESE PATIENTS WERE NOT IDENTIFIED IN THE STUDY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A, D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. SECTION B3: DATE OF EVENT HAS BEEN ENTERED AS 01OCT2024, SINCE THE LVAD IMPLANT OCCURRED BETWEEN JANUARY 2014 TO OCTOBER 2024. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. AUTHOR INFORMATION: BASHIR H, ET AL. J. CLIN. MED. JOURNAL OF THE SOCIETY FOR CARDIOVASCULAR ANGIOGRAPHY & INTERVENTIONS, ARTICLE IN PRESS. HTTPS://DOI.ORG/10.1016/J.JSCAI.2025.103662. THE CHRIST HOSPITAL HEART AND VASCULAR INSTITUTE AND THE LINDNER CENTER FOR RESEARCH AND EDUCATION, CINCINNATI, OHIO, USA. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22117849&pc=MVK | 3008642652-2025-05349 | 2025-05-09 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-30 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 11:12:56 ON (B)(6) 2025. AT 06:12:31 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 06:13:42, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. THE ELECTRODE BELT WAS DISCONNECTED AT 06:17:15 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22117951&pc=PNJ | 2017865-2025-75138 | 2025-05-15 04:00 | Death | ABBOTT | 2025-05-30 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE, THE DELIVERY CATHETER AND DEVICE WERE DIFFICULT TO INSERT INTO THE RIGHT VENTRICLE DUE TO TORTUOUS HEART ANATOMY. AS A RESULT, THE RIGHT VENTRICULAR OUTFLOW TRACK WAS PERFORATED. THE PERFORATION RESULTED IN EFFUSION AND TAMPONADE. PERICARDIOCENTESIS WAS PERFORMED, BUT THE PATIENT'S CONDITION CONTINUED TO DETERIORATE. THE PATIENT WAS UNSUCCESSFULLY RESUSCITATED AND PASSED AWAY. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22117956&pc=PNJ | 2017865-2025-75137 | 2025-05-15 04:00 | Death | ABBOTT | 2025-05-30 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE, THE DELIVERY CATHETER AND DEVICE WERE DIFFICULT TO INSERT INTO THE RIGHT VENTRICLE DUE TO TORTUOUS HEART ANATOMY. AS A RESULT, THE RIGHT VENTRICULAR OUTFLOW TRACK WAS PERFORATED. THE PERFORATION RESULTED IN EFFUSION AND TAMPONADE. PERICARDIOCENTESIS WAS PERFORMED, BUT THE PATIENT'S CONDITION CONTINUED TO DETERIORATE. THE PATIENT WAS UNSUCCESSFULLY RESUSCITATED AND PASSED AWAY. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22118204&pc=LYJ | 1644487-2025-10187 | 2025-05-07 04:00 | Death | LIVANOVA USA, INC. | 2025-05-30 04:00 | LYJ | PULSE GEN MODEL 103 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22118477&pc=BZD | 2518422-2025-045870 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-05-30 04:00 | BZD | REMSTAR PRO C-FLEX+ | Degraded | Cancer; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, CANCER, LIVER (FATAL). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22120002&pc=FKX | 0008030665-2025-01235 | 2025-05-14 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-05-30 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K222318 |  | 1 | Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED TO FRESENIUS CUSTOMER SERVICE THAT A PERITONEAL DIALYSIS (PD) PATIENT EXPIRED WHILE CONNECTED TO THEIR LIBERTY SELECT CYCLER. UPON FOLLOW-UP WITH THE PATIENT¿S PDRN, IT WAS REVEALED THAT THE PATIENT EXPIRED IN THE EARLY MORNING ON (B)(6) 2025 DUE TO A CARDIAC ARREST ATTRIBUTED TO A SIGNIFICANT CARDIAC DISEASE HISTORY. IT WAS AFFIRMED THAT THE PATIENT WAS CONNECTED TO THEIR CYCLER AT THE TIME OF THE ARREST AND SUBSEQUENT DEATH. IT WAS EXPLAINED THAT THE PATIENT RECENTLY UNDERWENT AN EMERGENCY AORTIC VALVE REPLACEMENT (EXACT DATE NOT REPORTED) AND WAS DISCHARGED HOME WITH A HOLTER HEART MONITOR. THE OVERSEEING FACILITY NOTIFIED THE PATIENT¿S SPOUSE THAT THE PATIENT WAS PULSELESS, AND EMERGENCY MEDICAL SERVICES WERE ACTIVATED. THE PATIENT WAS NOT TRANSPORTED AND THEY WERE PRONOUNCED DECEASED IN THEIR HOME. IT WAS CONFIRMED THAT THE PATIENT¿S CARDIAC ARREST LEADING TO THEIR DEATH WAS UNRELATED TO PD THERAPY AND NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). THE PDRN STATED THE PATIENT'S CYCLER WOULD BE RETURNED TO FRESENIUS. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S CARDIAC ARREST, LEADING TO THEIR DEATH. THE CAUSE OF THIS PATIENT¿S CARDIAC ARREST AND DEATH CAN BE ATTRIBUTED TO A SIGNIFICANT HISTORY OF CARDIOVASCULAR DISEASE AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE END-STAGE RENAL DISEASE (ESRD) POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION, PARTICULARLY IN THE ENVIRONMENT OF CARDIOVASCULAR DISEASE AND RECENT INVASIVE CARDIAC SURGERY. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22120282&pc=OZD | 1220648-2025-28875 | 2025-05-04 04:00 | Death | ABIOMED, INC. | 2025-05-30 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT HAD VENTRICULAR TACHYCARDIA (VT). THE PATIENT WAS SHOCKED AND SUPPORT CONTINUED. THE VT OCCURRED AGAIN LATER IN SUPPORT. IMPELLA PLACEMENT WAS VERIFIED. ADDITIONALLY, THE PATIENT HAD A HEMATOMA AROUND THE AXILLARY POCKET SITE. ONE UNIT OF PACKED RED BLOOD CELLS WAS GIVEN TO THE PATIENT. FURTHERMORE, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22102780&pc=OZP | 2032227-2025-182776 | 2023-05-10 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-29 04:00 | OZP | 670G INSULIN PUMP MMT-1780KL | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Cardiac Arrest; Hyperglycemia; Hypoglycemia | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2023 AT HOME. THE CAUSE OF DEATH AS REPORTED WAS HEART ATTACK AND THE CUSTOMER HAD BEEN DIABETIC SINCE LONG. THE BLOOD GLUCOSE LEVEL AT THE TIME OF DEATH WAS 1000 MG/DL. THE CUSTOMER HAD WORN THE PUMP AT THE TIME OF PASSING. THE EVENT INVOLVED PRODUCTSMMT-1780KL. TROUBLESHOOTING WAS PERFORMED. IT WAS UNKNOWN WHETHER THE CUSTOMER USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT OR NOT. IT WAS UNKNOWN WHETHER THE AUTO MODE FEATURE WAS ACTIVE AT THE TIME OF THE EVENT OR NOT. NO PRODUCT RETURN IS REQUIRED FORMMT-1780KL.Event Description: UPDATED SUMMARY: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2023, AT HOME. THE CAUSE OF DEATH, AS REPORTED, WAS A HEART CONDITION AND DIABETES. THE BLOOD GLUCOSE LEVEL AT THE TIME OF DEATH WAS 1000 MG/DL. THE CUSTOMER HAD WORN THE PUMP AT THE TIME OF PASSING. THE LOCATION OF THE CUSTOMER¿S PASSING AT THE HOSPITAL. THE REASON FOR ADMISSION TO THE HOSPITAL WAS THE HEART SURGERY. THE EVENT INVOLVED PRODUCTS MMT-1780KL, UNOMEDICAL, AND MMT-332A. TROUBLESHOOTING WAS PERFORMED. THE REASON THE CUSTOMER WAS NOT WEARING A PUMP AT THE TIME OF THE DECEASED EVENT WAS SURGERY. THE MEDTRONIC INSULIN PUMP DELIVERY SYSTEM COMPONENTS WORN OR USED AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KL, UNOMEDICAL, AND MMT-332A. Manufacturer Narrative: ADDITIONAL INFORMATION HAS BEEN RECEIVED WHICH WAS NOT INCLUDED IN THE INITIAL REPORT. THE UPDATED INFORMATION HAS BEEN PROVIDED IN BELOW SECTIONS WITH THIS FOLLOW-UP REPORT. 1. SECTION B5 - UPDATED THE SUMMARY. 2. SECTION D10 - UPDATED CONCOMITANT PRODUCTS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22103335&pc=LPB | 2029046-2025-01737 | 2025-05-08 04:00 | Death | BIOSENSE WEBSTER INC | 2025-05-29 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade | P030031 |  | 1 | Event Description: A PATIENT WITH HISTORY OF HYPERTENSION UNDERWENT A LEFT ATRIAL TACHYCARDIA ABLATION PROCEDURE WITH A THERMOCOOL SMARTTOUCH SF AND THE PATIENT EXPERIENCED PERICARDIAL EFFUSION AND DEATH. DURING THE PROCEDURE, A MASSIVE PERICARDIAL EFFUSION WAS DIAGNOSED VIA INTRACARDIAC ECHOCARDIOGRAPHY (ICE) AFTER MAPPING THE LEFT ATRIUM. THE PATIENT'S PRESSURE DROPPED AND THE HEART WAS NOT PERFUSING AFTER A PERICARDIOCENTESIS WAS PERFORMED UNSUCCESSFULLY BY THE PHYSICIAN. PACING WAS PERFORMED IN THE LEFT VENTRICLE WITH THE DECANAV CATHETER. NO CODE WAS CALLED AND THE PATIENT EXPIRED. ADDITIONAL INFORMATION WAS LATER RECEIVED INDICATING THE PHYSICIAN'S OPINION ON THE CAUSE OF DEATH WAS PERICARDIAL DECOMPRESSION SYNDROME. IT WAS REPORTED THAT THERE WAS LOW IMPEDANCE DURING THE CAVOTRICUSPID ISHMUS (CTI) PORTION DUE TO THE PATIENT¿S WEIGHT. SECOND GROUNDING PAD WAS UNPLUGGED HALFWAY THROUGH THE LINE. NO EFFUSION NOTICED BEFORE TRANSSEPTAL PUNCTURE. A LARGE EFFUSION WAS NOTICED DURING THE MAPPING OF THE LEFT ATRIUM SINCE PRESSURE WAS DROPPING. THEY PREPARED FOR PERICARDIOCENTESIS (TAP) AND THE PHYSICIAN ABLATED THE EARLIEST TACHYCARDIC (TACH) LOCATION DURING PREPARATION. TACH BROKE AND TAP WAS PERFORMED. LARGE DRAINAGE WAS DONE BUT THE HEART WAS NOT MOVING, (STUNNED). THE PHYSICIAN THOUGHT THE PATIENT HAD PERICARDIAL DECOMPRESSION SYNDROME. THERE WAS NO EVIDENCE OF STEAM POP. THE EVENT WAS NOTICED DURING THE MAPPING PHASE. PROCEDURAL ACTIVATED CLOTTING TIME (ACT) NEVER EXCEEDED 400 IT WAS BETWEEN 350 AND 400. THE NUMBER OF ABLATIONS PERFORMED WAS 19 AND THEY WERE ALL PERFORMED OUTSIDE OF THE PULMONARY VEINS. THERE WERE NO NOTICEABLE MALFUNCTIONS WITH ANY BWI DEVICE, SECOND GROUNDING PAD WAS SIMPLY UNPLUGGED FROM NGEN GENERATOR. THE IMPEDANCE LOW BASED WAS BASED ON THE PATIENT¿S WEIGHT, BECAUSE THE PATIENT WAS FRAIL, SKINNY PATIENT. NOT MUCH FAT ON THE BODY. Manufacturer Narrative: A PATIENT WITH HISTORY OF HYPERTENSION UNDERWENT A LEFT ATRIAL TACHYCARDIA ABLATION PROCEDURE WITH A THERMOCOOL SMARTTOUCH SF AND THE PATIENT EXPERIENCED PERICARDIAL EFFUSION AND DEATH. DURING THE PROCEDURE, A MASSIVE PERICARDIAL EFFUSION WAS DIAGNOSED VIA INTRACARDIAC ECHOCARDIOGRAPHY (ICE) AFTER MAPPING THE LEFT ATRIUM. THE PATIENT'S PRESSURE DROPPED AND THE HEART WAS NOT PERFUSING AFTER A PERICARDIOCENTESIS WAS PERFORMED UNSUCCESSFULLY BY THE PHYSICIAN. PACING WAS PERFORMED IN THE LEFT VENTRICLE WITH THE DECANAV CATHETER. NO CODE WAS CALLED AND THE PATIENT EXPIRED. ADDITIONAL INFORMATION WAS LATER RECEIVED INDICATING THE PHYSICIAN'S OPINION ON THE CAUSE OF DEATH WAS PERICARDIAL DECOMPRESSION SYNDROME. DEVICE EVALUATION DETAILS: THE PRODUCT WAS RETURNED TO JOHNSON & JOHNSON MEDTECH (J&J MEDTECH) FOR EVALUATION. A VISUAL INSPECTION AND REVISION OF ALL FEATURES WERE PERFORMED FOLLOWING J&J MEDTECH PROCEDURES. VISUAL ANALYSIS REVEALED NO DAMAGE OR ANOMALIES ON THE DEVICE. THE DEVICE FEATURES WERE REVIEWED, AND NO ISSUES WERE OBSERVED DURING THE PRODUCT INVESTIGATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER, AND NO INTERNAL ACTIONS WERE IDENTIFIED. NO MALFUNCTION WAS OBSERVED DURING THE PRODUCT ANALYSIS. THE ROOT CAUSE OF THE ADVERSE EVENT REMAINS UNKNOWN. IN ADDITION, NO DEVICE MALFUNCTION WAS REPORTED, THERE MAY HAVE BEEN OTHER CIRCUMSTANCES OR ISSUES THAT OCCURRED DURING THE USE OF THE DEVICE THAT COULD NOT BE REPLICATED DURING THE ANALYSIS. THE INSTRUCTION FOR USE (IFU) STATES THAT CAREFUL CATHETER MANIPULATION MUST BE PERFORMED TO AVOID CARDIAC DAMAGE, PERFORATION, OR TAMPONADE. AS PART OF JOHNSON & JOHNSON MEDTECH¿S QUALITY PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REF. # (B)(4). Manufacturer Narrative: ON 23-MAY-2025, THE BWI PRODUCT ANALYSIS LAB RECEIVED THE COMPLAINT DEVICE FOR EVALUATION. THE PRODUCT ANALYSIS HAS BEGUN BUT IS NOT COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REF. # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22104245&pc=OZD | 1220648-2025-28807 | 2025-05-03 04:00 | Death | ABIOMED, INC. | 2025-05-29 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels; Cardiac Perforation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT HAD AN IMPELLA CP PLACED DUE TO MYOCARDIAL INFARCTION. DURING SUPPORT, THE FIRST CP WAS PUNCTURED, IT WAS SUCCESSFULLY WEANED AND REMOVED THROUGH THE PEEL AWAY SHEATH. PIGTAIL RE-INSERTED INTO THE LEFT VENTRICLE, AND THE SECOND CP WAS PRIMED AND BACKLOADED OVER THE 0.018 WIRE AND PLACED ACROSS HEMOSTATIC VALVE INTO PROPER POSITION UNDER FLUOROSCOPY WITHOUT ISSUE. THE PATIENT WAS MOVED OVER TO THE ICU BED AND SHORTLY AFTER BECAME NAUSEOUS AND DIAPHORETIC. PATIENT VOICED INCREASING NAUSEA AND SHORTNESS OF BREATH. RESPIRATORY TOOK BIPAP OFF AND PLACED ON NON REBREATHER. O2 SATURATIONS DECREASED AND PATIENT TRIED TO SIT UP IN BED HAVING INCREASED SHORTNESS OF BREATH. PATIENT THEN BECAME SEVERELY HYPOTENSIVE AND CODE WAS CALLED. SUCTION ALARMS PRESENT ON AIC (CONSOLE) AND CONTINUED WITHOUT RESOLUTION TO P1. PATIENT WAS INTUBATED AND TAKEN FOR A REPEAT ANGIOGRAM AND RIGHT HEART CATH. ECHO WAS PERFORMED IN CATH LAB AND REVEALED LARGE PERICARDIAL EFFUSION. THE PHYSICIAN THEN PERFORMED A PERICARDIOCENTESIS WHILE CPR WAS BEING PERFORMED. SUCTION ALARMS WERE STILL PERSISTENT AND P LEVEL WAS MAINTAINED ON P1. PATIENT EXPIRED IN CATH LAB WHILE ON IMPELLA SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE PATIENT'S OUTCOME WAS REVIEWED BY ABIOMED'S MEDICAL SAFETY OFFICER (MSO) AND IT WAS DETERMINED THAT THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THIS IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THE LEFT VENTRICULAR PERFORATION IS A SEVERE INJURY, CONSIDERED TO BE A LIFE-THREATENING CONDITION, AND THEREFORE, THE PATIENT DEMISE OUTCOME IS MORE LIKELY ASSOCIATED WITH THIS EVENT INVOLVING THE IMPELLA CP PUMP. *THIS REPORTABLE EVENT IS FOR THE SAME PATIENT IN MANUFACTURER MDR REPORT #1220648-2025-28876. HOWEVER, THIS MANUFACTURER REPORT CAPTURES REPORTABLE EVENTS AS IT RELATES TO THE REPLACEMENT IMPELLA CP, SERIAL NUMBER (B)(6). Manufacturer Narrative: THE INVESTIGATION OF VENTRICLE PERFORATION HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. VENTRICLE PERFORATION: DATA LOG REVIEW SHOWS THE PLACEMENT SIGNAL BEGIN TRENDING DOWNWARDS AFTER THE PUMP WAS TURNED DOWN FROM P-8. THERE WERE ALSO IMPELLA IN VENTRICLE AND AORTA ALARMS NEAR THE END OF THE CASE WHILE THE PUMP WAS AT P-1. IT IS UNCLEAR WHEN THE ALARMS OCCURRED IN RELATION TO THE CLINICAL TIMELINE. NO PRODUCT WAS RETURNED. THE ROOT CAUSE OF THE VENTRICLE PERFORATION WAS MOST LIKELY PATIENT MOVEMENT SINCE THE CLINICAL DETAILS STATE THAT THE PATIENT'S CONDITION BEGAN DETERIORATING AFTER TRANSFER TO ICU AND FURTHER WORSENED AFTER THE PATIENT SAT UP. B3 DATE OF EVENT WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28807. H6 HEALTH EFFECT - CLINICAL CODE 2135 WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28807. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22104425&pc=OZD | 1220648-2025-28810 | 2025-05-06 04:00 | Death | ABIOMED, INC. | 2025-05-29 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT HAD EPISODES OF VENTRICULAR TACHYCARDIA. THE PATIENT WAS SHOCKED OUT OF IT. SUPPORT CONTINUED UNTIL CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22104668&pc=OUT | 2029214-2025-01257 | 2025-05-15 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-05-29 04:00 | OUT | PIPELINE FLEX W/SHIELD TECHNOLOGY | Positioning Failure; Retraction Problem; Structural Problem; Activation, Positioning or Separation Problem; Physical Resistance/Sticking | Hemorrhage/Bleeding; Intracranial Hemorrhage; High Blood Pressure/ Hypertension; Rupture | P100018 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN INITIAL IMPLANT PROCEDURE USING THE PIPELINE EMBOLIZATION DEVICE 4.00MM X 30MM FOR MULTIPLE UNRUPTURED, SACCULAR ANEURYSMS LOCATED IN THE POSTERIOR COMMUNICATING SEGMENT OF THE LEFT INTERNAL CAROTID ARTERY (ICA), SEVERAL COMPLICATIONS OCCURRED. THE ANEURYSMS INCLUDED AN UPPER POSTERIOR COMMUNICATING ARTERY ANEURYSM MEASURING 7.1MM IN MAXIMUM DIAMETER WITH A 5.8MM NECK, AND ANOTHER ANEURYSM APPROXIMATELY 5MM IN SIZE BELOW. THE DISTAL LANDIN ZONE WAS 3.3MM AND THE PROXIMAL LANDING ZONE WAS 3.9MM. THE VESSEL EXHIBITED SEVERE TORTUOSITY. FOLLOWING FEMORAL ARTERY PUNCTURE, A 6F 90CM LONG SHEATH, 6F 115CM TETHYS GUIDE CATHETER, PHENOM 27 MICROCATHETER, AND A 0.014 INCH 200CM GUIDEWIRE WERE USED TO ACCESS THE TARGET SITE. THE DEVICES WERE PREPARED ACCORDING TO THE INSTRUCTIONS FOR USE (IFU). 3D ROTATIONAL ANGIOGRAPHY WAS PERFORMED TO GUIDE DEVICE SELECTION AND PLACEMENT. THE STENT WAS HYDRATED AND DELIVERED, WITH THE TIP OPENED AT THE M1 SEGMENT OF THE MIDDLE CEREBRAL ARTERY. DURING PULL-BACK POSITIONING, THE STENT COULD NOT BE ANCHORED DUE TO A SHORT ANCHORING DISTANCE, RESULTING IN STENT RETRACTION. WHEN ATTEMPTING TO RETRIEVE THE STENT FOR REPOSITIONING, EXCESSIVE RESISTANCE WAS ENCOUNTERED. SUBSEQUENTLY, A RAPID RISE IN BLOOD PRESSURE WAS OBSERVED, AND ANGIOGRAPHY REVEALED RUPTURE AND BLEEDING OF THE INTERNAL CAROTID ARTERY WITH SUBARACHNOID HEMORRHAGE. PUPILLARY DILATION WAS NOTED ON THE OPERATING TABLE. CRANIOTOMY WAS PERFORMED AS A RESCUE INTERVENTION, BUT THE RESCUE WAS UNSUCCESSFUL AND THE PATIENT DIED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE PROBLEM OCCURRED THAT DAY AND THE PATIENT DIED. THE PATIENT PASSED AWAY AT (B)(6), DOCTOR INFORMED THAT THE PATIENT DIED IN THE HOSPITAL. NO AUTOPSY RECORDS WILL BE AVAILABLE. THE PATIENT'S DEATH MAY BE RELATED TO THE PIPELINE SHIELD, AND IT MAY ALSO BE RELATED TO THE OPERATOR'S OPERATION. THERE WAS INTRACRANIAL ANEURYSM LEADING UP TO THE PATIENT'S DEATH. THE CAUSE OF THE RESISTANCE DURING RETRIEVAL MAY BE VASCULAR TORTUOSITY. Manufacturer Narrative: PRODUCT ANALYSIS: AS FOUND CONDITION: THE PIPELINE FLEX SHIELD WAS RETURNED INSIDE THE INNER POUCH, BIOHAZARD BAG, AND SHIPPING BOX. THE CATHETER WAS NOT RETURNED. VISUAL INSPECTION/DAMAGE LOCATION DETAILS: THE DISTAL AND PROXIMAL DPS RESTRAINTS WERE FOUND TO BE INTACT. THE DPS SLEEVES WERE FOUND INTACT WITH NO SIGNS OF DAMAGE. THE DISTAL HYPOTUBE AND PTFE SHRINK TUBING WERE FOUND TO BE INTACT WITH NO SIGNS OF ELONGATION. THE BRAID'S DISTAL, MIDDLE, AND PROXIMAL WERE FOUND OPEN, WITH NO DAMAGE. NO BENDS WERE FOUND WITH THE PUSHWIRE. NO DEFECTS WERE FOUND WITH THE TIP COIL, DISTAL MARKER, RE-SHEATHING MARKER, RE-SHEATHING PAD, OR THE PROXIMAL BUMPER. NO OTHER ANOMALIES WERE OBSERVED. TESTING/ANALYSIS: N/A MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22104769&pc=MHX | 8030229-2025-04952 | 2025-04-29 04:00 | Death | NIHON KOHDEN CORPORATION | 2025-05-29 04:00 | MHX | PU-621R | Protective Measures Problem | Arrhythmia | K102376 |  | 1 | Event Description: THE BIOMEDICAL ENGINEER (BME) REPORTED AN ALARM INCIDENT. THEY WERE UNSURE WHETHER THE ALARM HAD BEEN MISSED OR SILENCED AT THE CENTRAL NURSE'S STATION (CNS). IT WAS UNCLEAR WHEN THE PATIENT PASSED AWAY.Event Description: THE BIOMEDICAL ENGINEER (BME) REPORTED AN ALARM INCIDENT. THEY WERE UNSURE WHETHER THE ALARM HAD BEEN MISSED OR SILENCED AT THE CENTRAL NURSE'S STATION (CNS). IT WAS UNCLEAR WHEN THE PATIENT PASSED AWAY. Manufacturer Narrative: DETAILS OF ORIGINAL COMPLAINT: THE BIOMEDICAL ENGINEER (BME) REPORTED AN ALARM INCIDENT. THEY WERE UNSURE WHETHER THE ALARM HAD BEEN MISSED OR SILENCED AT THE CENTRAL NURSE'S STATION (CNS). IT WAS UNCLEAR WHEN THE PATIENT PASSED AWAY. THE BME WOULD LIKE TO KNOW THE FULL PICTURE OF WHAT HAD HAPPENED DURING THAT TIME. TECHNICAL SUPPORT (TS) REVIEWED THE LOGS FROM THIS CNS FOR THE TIME PERIOD BETWEEN 9:07 AM AND 12:35 PM PST, AND CONFIRMED THE CNS WAS ALARMING, BUT THE ALARMS WERE BEING SILENCED CONSTANTLY BY THE USER. TS AND OUR NK CLINICAL APPLICATION SPECIALIST (CAS) INFORMED THE BME THAT THE CNS DID ALARM CORRECTLY AND THAT THE ALARMS HAD BEEN SILENCED NUMEROUS TIMES OVER THE COURSE OF THAT PERIOD. THE CNS WAS SENT IN FOR EVALUATION. INVESTIGATION SUMMARY: THE ISSUE WAS NOT DUPLICATED DURING EVALUATION. THE ROOT CAUSE COULD NOT BE DETERMINED. THE REPORTED DEVICE HAS REACHED ITS END-OF-LIFE. NO FURTHER INVESTIGATION IS REQUIRED. THE FOLLOWING FIELD CONTAIN NO INFORMATION (NI), AS ATTEMPTS TO OBTAIN THE INFORMATION WERE MADE, BUT NOT PROVIDED: B2 (IN EVENT OF DEATH, DATE OF DEATH) ATTEMPT # 1: 04/30/2025 EMAILED THE BME FOR ALL INFORMATION IN THE NI LIST ABOVE: THE BME SENT THE ADVERSE EVENT FORM AND DESCRIPTION OF THE EVENT, WHICH CONTRADICTED THE INITIAL REPORT. THE BME SELECTED THERE WAS NO INJURY OR DEATH. ATTEMPT # 2: 05/28/2025 EMAILED THE BME FOR CLAIRIFICATION OF PATIENT EVENT: NO REPLY WAS RECEIVED. ATTEMPT # 3: 05/29/2025 EMAILED THE BME AND THE CUSTOMER FOR CLAIRIFICATION OF PATIENT EVENT: NO REPLY WAS RECEIVED. ADDITIONAL DEVICE INFORMATION: D10 CONCOMITANT MEDICAL DEVICE: THE FOLLOWING DEVICE WAS USED IN CONJUNCTION WITH THE CNS: ZM-520 TRANSMITTER, MODEL #: ZM-520PA, SERIAL #: (B)(6), DEVICE MANUFACTURER DATA: 07/18/2018 (JULY 18, 2018), UNIQUE IDENTIFIER (UDI) #: (B)(4). ADDITIONAL INFORMATION: B4 DATE OF THIS REPORT. G3 DATE RECEIVED BY MANUFACTURER. G6 TYPE OF REPORT. H2 IF FOLLOW-UP, WHAT TYPE? H3 DEVICE EVALUATED BY MANUFACTURER. H6 EVENT PROBLEM AND EVALUATION CODES. H11 ADDITIONAL MANUFACTURER NARRATIVE. Manufacturer Narrative: THE BIOMEDICAL ENGINEER (BME) REPORTED AN ALARM INCIDENT. THEY WERE UNSURE WHETHER THE ALARM HAD BEEN MISSED OR SILENCED AT THE CENTRAL NURSE'S STATION (CNS). IT WAS UNCLEAR WHEN THE PATIENT PASSED AWAY. THE BME WOULD LIKE TO KNOW THE FULL PICTURE OF WHAT HAD HAPPENED DURING THAT TIME. TECHNICAL SUPPORT (TS) REVIEWED THE LOGS FROM THIS CNS FOR THE TIME PERIOD BETWEEN 9:07 AM AND 12:35 PM PST, AND CONFIRMED THE CNS WAS ALARMING, BUT THE ALARMS WERE BEING SILENCED CONSTANTLY BY THE USER. TS AND OUR NK CLINICAL APPLICATION SPECIALIST (CAS) INFORMED THE BME THAT THE CNS DID ALARM CORRECTLY AND THAT THE ALARMS HAD BEEN SILENCED NUMEROUS TIMES OVER THE COURSE OF THAT PERIOD. THE CNS WAS SENT IN FOR EVALUATION. NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. THE FOLLOWING FIELD CONTAIN NO INFORMATION (NI), AS ATTEMPTS TO OBTAIN THE INFORMATION WERE MADE, BUT NOT PROVIDED: B2 (IN EVENT OF DEATH, DATE OF DEATH). ATTEMPT # 1: 04/30/2025 EMAILED THE BME FOR ALL INFORMATION IN THE NI LIST ABOVE: THE BME SENT THE ADVERSE EVENT FORM AND DESCRIPTION OF THE EVENT, WHICH CONTRADICTED THE INITIAL REPORT. THE BME SELECTED THERE WAS NO INJURY OR DEATH. ATTEMPT # 2: 05/28/2025 EMAILED THE BME FOR CLAIRIFICATION OF PATIENT EVENT: NO REPLY WAS RECEIVED. ATTEMPT # 3: 05/29/2025 EMAILED THE BME AND THE CUSTOMER FOR CLAIRIFICATION OF PATIENT EVENT: NO REPLY WAS RECEIVED. ADDITIONAL DEVICE INFORMATION: D10 CONCOMITANT MEDICAL DEVICE: THE FOLLOWING DEVICE WAS USED IN CONJUNCTION WITH THE CNS: ZM-520 TRANSMITTER, MODEL #: ZM-520PA, SERIAL #: (B)(6), DEVICE MANUFACTURER DATA: 07/18/2018 (JULY 18, 2018), UNIQUE IDENTIFIER (UDI) #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22104903&pc=OZD | 1220648-2025-28812 | 2025-05-06 04:00 | Death | ABIOMED, INC. | 2025-05-29 04:00 | OZD | IMPELLA CP | Failure to Advance; Patient-Device Incompatibility | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN EMERGENT PERCUTANEOUS CORONARY INTERVENTION AN IMPELLA CP WAS ATTEMPTED TO BE PLACED, BUT UNABLE TO CROSS HIGHLY CALCIFIED ATRIOVENTRICULAR PROBABLY DUE TO AORTIC STENOSIS. AN INTRA-AORTIC BALLOON PUMP WAS PLACED HOWEVER THE PATIENT EXPIRED IN PROCEDURE AREA. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. NO DAMAGES OR ABNORMALITIES WERE FOUND ON THE RETURNED PUMP. THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED SINCE THE PATIENT HAD A HIGHLY CALCIFIED ATRIOVENTRICULAR (AV). Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS. ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22105027&pc=DQA | MW5170850 |  | Death | MASIMO CORPORATION | 2025-05-29 04:00 | DQA | MASIMO RADICAL-7 PULSE OXIMETER | Unable to Obtain Readings; Incorrect, Inadequate or Imprecise Result or Readings; Communication or Transmission Problem; Protective Measures Problem; Application Program Freezes, Becomes Nonfunctional | Low Oxygen Saturation |  |  | 1 | Event Description: DEVICE-SETTING ERROR (2015 ¿ 2024) ¿ FROM (B)(6) 2015 TO (B)(6) 2024 (B)(6) CONDUCTED PEDIATRIC POLYSOMNOGRAPHY (PSG) WITH MASIMO RADICAL-7 PULSE-OXIMETERS SET TO A REDUCED SAMPLING RATE. INTERNAL TESTING SHOWED A PERSISTENT +3 ¿ 5 % SPO? BIAS. DIY INTERFACE CABLE ¿ (B)(6) NEURO-IT BUILT AN UNVALIDATED RS-232-TO-USB CABLE TO STREAM SPO? INTO THE NATUS PSG SYSTEM. THE CABLE EXPERIENCED >120 COMMUNICATION DROPOUTS; WHEN THE LINK FAILED, THE PSG TRACE FROZE ON THE LAST VALUE, HIDING DESATURATIONS FOR MINUTES. SENTINEL DEATH ((B)(6) 2022) ¿ A SIX-MONTH-OLD INFANT DIED DURING PSG WHEN A HIDDEN FREEZE MASKED A DROP IN OXYGEN SATURATION. THE (B)(6) CLOSED THE LAB FOR ONE MONTH. DETAILS OF THE FROZEN TRACE ARE DESCRIBED IN A (B)(6) GLOBE ARTICLE (01-MAY-2023). NO FDA MDR WAS FILED. SCOPE OF EXPOSURE ¿ ?10 000 PEDIATRIC PSG STUDIES ((B)(6) 2015 ¿ (B)(6) 2024) MAY UNDER-REPORT HYPOXEMIA. UNDER-DIAGNOSED OBSTRUCTIVE SLEEP-APNEA (OSA) AND CHRONIC NOCTURNAL HYPOXIA IN CHILDREN CAN CONTRIBUTE TO NEURO-COGNITIVE DELAY, CARDIOVASCULAR CHANGES, IMPAIRED GROWTH, AND BEHAVIORAL DISORDERS. STAFF CORROBORATION ¿ MORE THAN A DOZEN TECHNOLOGISTS AND PHYSICIANS DOCUMENTED SCREEN FREEZES IN STAFF MEETINGS (MINUTES ATTACHED). THE REPORTER ISSUED A FORMAL WARNING ON 18-OCT-2021, NO VALIDATION STUDY FOLLOWED. PARTIAL FIX / UNRESOLVED RISK ¿ ON (B)(6) 2024 THE REPORTER RECALIBRATED THE DEVICES AND REPLACED THE DIY CABLE WITH A MANUFACTURER-APPROVED PART, BUT LEADERSHIP HAS DECLINED TO REVIEW LEGACY STUDIES OR NOTIFY FAMILIES. EXTERNAL FILINGS ¿ JOINT COMMISSION PATIENT-SAFETY EVENT JC-24-07-12345 (22-JUL-2024) AND (B)(6) DPH COMPLAINT (26-MAY-2025) CITE THE SAME CONCERNS. REQUEST TO FDA/CDRH ¿ (1) DETERMINE WHETHER THE SAMPLING-RATE BIAS AND UNAPPROVED CABLE CONSTITUTES A REPORTABLE DEVICE MODIFICATION; (2) INVESTIGATE (B)(6)'S FAILURE TO FILE AN MDR AFTER THE 2022 DEATH; (3) ISSUE GUIDANCE TO PREVENT SIMILAR UNDER-DIAGNOSIS OF PEDIATRIC HYPOXEMIA IN PSG LABS. THE REPORTER CAN SUPPLY RAW PSG FILES, DROPOUT LOGS, AND STAFF STATEMENTS ON REQUEST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22105090&pc=OZD | 1220648-2025-28815 | 2025-05-07 04:00 | Death | ABIOMED, INC. | 2025-05-29 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT HAD A COMPUTED TOMOGRAPHY SCAN THAT RESULTED IN THE PRESENCE OF TRANSTENTORIAL HERNIATION, INDICATIONS OF HYPOXIA/ISCHEMIC ENCEPHALOPATHY AND EDEMA. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE PATIENT'S BRAIN INJURIES WERE RELATED TO THE IMPELLA. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22105330&pc=OZD | 1220648-2025-28816 | 2025-05-06 04:00 | Death | ABIOMED, INC. | 2025-05-29 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD VENTRICULAR TACHYCARDIA. THE PATIENT WAS SHOCKED MULTIPLE TIMES. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22105514&pc=KNT | MW5170864 | 2022-06-18 04:00 | Death | CONVATEC, INC. | 2025-05-29 04:00 | KNT | FLEXI-SEAL FECAL MANAGEMENT SYSTEM | Off-Label Use; Improper or Incorrect Procedure or Method | Diarrhea; Hemorrhage/Bleeding; Vomiting; Abdominal Distention |  |  | 1 | Event Description: THE REPORTER RESPONDED TO CALL TO MAKE COMPLAIN REGARDING DEVICE FLEXI-SEAL FECAL MANAGEMENT SYSTEM. THE REPORTER SAID HER HUSBAND WAS IN INTENSIVE CARE UNIT FOR TREATMENT OF HEART ATTACK. SHE ALSO MENTIONED HE WAS ON BLOOD THINNER AND ANTIBIOTICS WHICH WAS INFORMED TO THE HEALTH CARE SERVICE PROVIDERS IN THE INTENSIVE CARE UNIT. DURING THE TIME IN ICU, HE HAD DIARRHEA AND FLEXI-SEAL FECAL MANAGEMENT SYSTEM WAS INSERTED ON (B)(6) 2025. THE CALLER STATED, "THE DEVICE WAS USED ON MY HUSBAND WITHOUT CONSENT AND WITHOUT RECORDS OR DOCUMENTATIONS." HE EXPERIENCED VOMITING, ABDOMINAL DISTENTION AND RECTAL BLEEDING. THE TUBE WAS REMOVED ON (B)(6) 2025. HE PASSED AWAY ON (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22106272&pc=LOX | 2124215-2025-33530 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-29 04:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIOGENIC SHOCK THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22106480&pc=DTB | 2649622-2025-14612 | 2025-05-08 04:00 | Death | MPRI | 2025-05-29 04:00 | DTB | CAPSURE SENSE LEAD MRI SURESCAN | Failure to Capture; Mechanical Problem; Use of Device Problem; Unstable Capture Threshold | Cardiac Arrest; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Ventricular Fibrillation; Pericardial Effusion; Syncope/Fainting; Thrombosis/Thrombus; Asystole | P830061 |  | 1 | Event Description: THE PATIENT PRESENTED TO HOSPITAL WITH SYNCOPE. NORMAL DEVICE FUNCTION, NORMAL THRESHOLDS AND CLEAR CAPTURE WAS NOTED ON THE VENTRICULAR ELECTROGRAM (EGM). IT WAS REPORTED THAT LATER, THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED LOSS OF VENTRICULAR CAPTURE ON TELEMETRY. VARYING RV CAPTURE THRESHOLDS WERE NOTED ON INTERROGATION. THE RV LEAD WAS REPROGRAMMED TO ENSURE CAPTURE UNTIL A MORE PERMANENT FIX COULD BE ACHIEVED. THE PATIENT WAS TRANSFERRED TO ANOTHER FACILITY FOR IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) AS A REPLACEMENT FOR THE RV LEAD. THE LEADLESS IPG WAS DEPLOYED AND CAPTURE THRESHOLD WAS NOTED TO BE WITHIN NORMAL LIMITS. THREE TINES WERE NOTED TO BE ENGAGED DURING THE PULL AND HOLD TEST. THE TETHER OF THE DELIVERY SYSTEM WAS CUT AND THRESHOLD INCREASED. THE IPG COULD NOT BE RECAPTURED AND TETHER REMOVAL WAS COMPLETED. THE THRESHOLD DECREASED GRADUALLY BUT IT WAS DECIDED TO IMPLANT A SECOND LEADLESS IPG TO GET A BETTER THRESHOLD AS THE PATIENT WAS PACEMAKER DEPENDENT. BEFORE IMPLANT OF THE SECOND IPG, AN ATTEMPT WAS MADE TO SNARE THE FIRST IPG. IT WAS DIFFICULT TO REACH THE IPG. A GOOSENECK SNARE WAS ENGAGED AROUND THE MAIN BODY OF THE IPG AND IT WASN'T COAXIAL TO BE BROUGHT INTO THE SHEATH. A SECOND SNARE WAS ATTEMPTED TO SNARE THE RETRIEVAL BUTTON. HOWEVER, DURING THIS PROCESS THE DELIVERY SHEATH HAD BEEN PULLED BACK TO WHERE IT WAS NEARLY OUT OF THE BODY AND SHEATH POSITION WAS WATCHED SO IT WOULD NOT EXIT THE BODY WITHOUT THE SNARED IPG. HOWEVER, THE SHEATH WAS ACCIDENTALLY PULLED OUT AND THE IPG WAS LEFT INSIDE THE BODY. AT THIS POINT, THE PATIENT'S BLOOD PRESSURE WAS DOWN. A QUICK ECHO WAS PERFORMED WHICH SHOWED NO EFFUSION AND THE PRESSURE DROP WAS ATTRIBUTED TO BLOOD LOSS OUT OF THE INTRODUCER SHEATH DUE TO MULTIPLE CATHETERS AND SNARES INSIDE. AN ARTERIAL LINE WAS THEN PLACED AND THE PATIENT WAS INTUBATED AND VASCULAR SURGERY WAS PAGED FOR BACKUP. NEXT, JUGULAR ACCESS WAS OBTAINED TO ATTEMPT TO SNARE THE IPG, WHICH WAS IN THE PATIENT'S GROIN NEAR WHERE THE FEMORAL SHEATH HAD EXITED THE BODY. A WIRE WAS ADVANCED ALL THE WAY TO THE FEMORAL VEIN AND THE SHEATH WAS ADVANCED, HOWEVER THE WIRE NEVER REACHED THE LOCATION OF THE IPG. A VENOGRAM WAS PERFORMED WHICH REVEALED THE IPG WAS OUTSIDE OF THE VEIN. AT THIS POINT, IT WAS ESTABLISHED THAT IPG WOULD NEED TO BE REMOVED SURGICALLY. HOWEVER, SINCE THERE WAS ALREADY ACCESS IN THE JUGULAR VEIN, IT WAS DECIDED TO IMPLANT THE SECOND LEADLESS IPG VIA THE JUGULAR APPROACH. THE SECOND LEADLESS IPG WAS ADVANCED TO THE RV. AFTER DEPLOYMENT THERE WAS NO CAPTURE AT MAXIMUM OUTPUT. THE SECOND IPG WAS THEN RECAPTURED WITH SOME DIFFICULTY AND WAS REMOVED FROM THE BODY. AT THIS POINT, THE PRESSURE DROPPED. AN INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER WAS INSERTED VIA GROIN SHEATH WHICH SHOWED LARGE EFFUSION. EMERGENT PERICARDIOCENTESIS WAS PERFORMED WITH BLOOD BEING WITHDRAWN AND READMINISTERED VIA THE JUGULAR ACCESS. HOWEVER, THE EFFUSION WAS NOT ABLE TO BE DECREASED RAPIDLY ENOUGH TO MAKE A DIFFERENCE IN THE EFFUS ION. CODE BLUE WAS CALLED. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF) AND DEFIBRILLATION WAS ATTEMPTED UNSUCCESSFULLY. A SECOND EXTERNAL DEFIBRILLATOR WAS BROUGHT IN AND BOTH EXTERNAL DEFIBRILLATORS WERE USED TO DELIVER SIMULTANEOUS 360 JOULE SHOCKS. HOWEVER, VF WAS UNABLE TO BE TERMINATED. ON THE ICE CATHETER A LEFT VENTRICULAR (LV) THROMBUS WAS SEEN. CPR WAS STOPPED AND PATIENT WAS DECLARED DECEASED. CAUSE OF DEATH WAS PROVIDED AS: PATIENT DIED DUE TO INCESSANT VF SECONDARY TO EFFUSION THAT WAS UNABLE TO BE RESOLVED DURING LEADLESS IPG IMPLANT TO REPLACE THE RV LEAD. THE LEADLESS IPGS WERE ATTEMPTED/NOT USED. THE RV LEAD REMAINS IN THE PATIENT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22106503&pc=NPT | 2025587-2025-03945 | 2025-02-28 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-29 04:00 | NPT | MEDTRONIC TRANSCATHETER VALVE | Patient Device Interaction Problem | Atrial Fibrillation | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING A SYSTEMATIC REVIEW OF NEW-ONSET ATRIAL FIBRILLATION (NOAF) FOLLOWING TRANSCATHETER AORTIC VALVE I MPLANTATION (TAVI). THE AUTHORS ANALYZED 18 STUDIES PUBLISHED BETWEEN 2009 AND 2024. VARIOUS VALVE TYPES WERE USED IN THE POOLED PATIENT POPULATION, ENCOMPASSING MEDTRONIC (COREVALVE / EVOLUT R/ EVOLUT PRO / EVOLUT PRO+), NON-MEDTRONIC (SAPIEN FAMILY), AND UNSPECIFIED. ADVERSE OUTCOMES ASSOCIATED WITH NOAF AFTER TAVI INCLUDED: MORTALITY, STROKE, SYSTEMIC EMBOLISM/THROMBOEMBOLIC EVENT, BLEEDING, HEART FAILURE, AND HOSPITAL READMISSION. Manufacturer Narrative: CITATION: MONTENEGRO-PALACIOS, J.F.; VIDAL-CAÑAS, S.; MURILLO-BENÍTEZ, N.E.; QUINTANA-OSPINA, J.; CARDONA MURILLO, C.A.; LISCANO, Y. COMMON RISK FACTORS FOR ATRIAL FIBRILLATION AFTER TRANSCATHETER AORTIC VALVE IMPLANTATION: A SYSTEMATIC REVIEW FROM 2009 TO 2024. J. CARDIOVASC. DEV. DIS. 2025, 12, 90. HTTPS://DOI.ORG/10.3390/JCDD12030090 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. EARLIEST APPROVED MEDTRONIC TRANSCATHETER VALVE USED FOR PRODUCT CODE AND PMA#. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22106510&pc=PNJ | 9612164-2025-02658 | 2025-05-08 04:00 | Death | MEDTRONIC IRELAND | 2025-05-29 04:00 | PNJ | MICRA¿ AV2 | Mechanical Problem; Migration or Expulsion of Device; Use of Device Problem; Capturing Problem; Positioning Problem; High Capture Threshold | Cardiac Arrest; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Ventricular Fibrillation; Pericardial Effusion; Thrombosis/Thrombus; Asystole | P150033 |  | 1 | Event Description: THE PATIENT PRESENTED TO HOSPITAL WITH SYNCOPE. NORMAL DEVICE FUNCTION, NORMAL THRESHOLDS AND CLEAR CAPTURE WAS NOTED ON THE VENTRICULAR ELECTROGRAM (EGM). IT WAS REPORTED THAT LATER, THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED LOSS OF VENTRICULAR CAPTURE ON TELEMETRY. VARYING RV CAPTURE THRESHOLDS WERE NOTED ON INTERROGATION. THE RV LEAD WAS REPROGRAMMED TO ENSURE CAPTURE UNTIL A MORE PERMANENT FIX COULD BE ACHIEVED. THE PATIENT WAS TRANSFERRED TO ANOTHER FACILITY FOR IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) AS A REPLACEMENT FOR THE RV LEAD. THE LEADLESS IPG WAS DEPLOYED AND CAPTURE THRESHOLD WAS NOTED TO BE WITHIN NORMAL LIMITS. THREE TINES WERE NOTED TO BE ENGAGED DURING THE PULL AND HOLD TEST. THE TETHER OF THE DELIVERY SYSTEM WAS CUT AND THRESHOLD INCREASED. THE IPG COULD NOT BE RECAPTURED AND TETHER REMOVAL WAS COMPLETED. THE THRESHOLD DECREASED GRADUALLY BUT IT WAS DECIDED TO IMPLANT A SECOND LEADLESS IPG TO GET A BETTER THRESHOLD AS THE PATIENT WAS PACEMAKER DEPENDENT. BEFORE IMPLANT OF THE SECOND IPG, AN ATTEMPT WAS MADE TO SNARE THE FIRST IPG. IT WAS DIFFICULT TO REACH THE IPG. A GOOSENECK SNARE WAS ENGAGED AROUND THE MAIN BODY OF THE IPG AND IT WASN'T COAXIAL TO BE BROUGHT INTO THE SHEATH. A SECOND SNARE WAS ATTEMPTED TO SNARE THE RETRIEVAL BUTTON. HOWEVER, DURING THIS PROCESS THE DELIVERY SHEATH HAD BEEN PULLED BACK TO WHERE IT WAS NEARLY OUT OF THE BODY AND SHEATH POSITION WAS WATCHED SO IT WOULD NOT EXIT THE BODY WITHOUT THE SNARED IPG. HOWEVER, THE SHEATH WAS ACCIDENTALLY PULLED OUT AND THE IPG WAS LEFT INSIDE THE BODY. AT THIS POINT, THE PATIENT'S BLOOD PRESSURE WAS DOWN. A QUICK ECHO WAS PERFORMED WHICH SHOWED NO EFFUSION AND THE PRESSURE DROP WAS ATTRIBUTED TO BLOOD LOSS OUT OF THE INTRODUCER SHEATH DUE TO MULTIPLE CATHETERS AND SNARES INSIDE. AN ARTERIAL LINE WAS THEN PLACED AND THE PATIENT WAS INTUBATED AND VASCULAR SURGERY WAS PAGED FOR BACKUP. NEXT, JUGULAR ACCESS WAS OBTAINED TO ATTEMPT TO SNARE THE IPG, WHICH WAS IN THE PATIENT'S GROIN NEAR WHERE THE FEMORAL SHEATH HAD EXITED THE BODY. A WIRE WAS ADVANCED ALL THE WAY TO THE FEMORAL VEIN AND THE SHEATH WAS ADVANCED, HOWEVER THE WIRE NEVER REACHED THE LOCATION OF THE IPG. A VENOGRAM WAS PERFORMED WHICH REVEALED THE IPG WAS OUTSIDE OF THE VEIN. AT THIS POINT, IT WAS ESTABLISHED THAT IPG WOULD NEED TO BE REMOVED SURGICALLY. HOWEVER, SINCE THERE WAS ALREADY ACCESS IN THE JUGULAR VEIN, IT WAS DECIDED TO IMPLANT THE SECOND LEADLESS IPG VIA THE JUGULAR APPROACH. THE SECOND LEADLESS IPG WAS ADVANCED TO THE RV. AFTER DEPLOYMENT THERE WAS NO CAPTURE AT MAXIMUM OUTPUT. THE SECOND IPG WAS THEN RECAPTURED WITH SOME DIFFICULTY AND WAS REMOVED FROM THE BODY. AT THIS POINT, THE PRESSURE DROPPED. AN INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER WAS INSERTED VIA GROIN SHEATH WHICH SHOWED LARGE EFFUSION. EMERGENT PERICARDIOCENTESIS WAS PERFORMED WITH BLOOD BEING WITHDRAWN AND READMINISTERED VIA THE JUGULAR ACCESS. HOWEVER, THE EFFUSION WAS NOT ABLE TO BE DECREASED RAPIDLY ENOUGH TO MAKE A DIFFERENCE IN THE EFFUS ION. CODE BLUE WAS CALLED. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF) AND DEFIBRILLATION WAS ATTEMPTED UNSUCCESSFULLY. A SECOND EXTERNAL DEFIBRILLATOR WAS BROUGHT IN AND BOTH EXTERNAL DEFIBRILLATORS WERE USED TO DELIVER SIMULTANEOUS 360 JOULE SHOCKS. HOWEVER, VF WAS UNABLE TO BE TERMINATED. ON THE ICE CATHETER A LEFT VENTRICULAR (LV) THROMBUS WAS SEEN. CPR WAS STOPPED AND PATIENT WAS DECLARED DECEASED. CAUSE OF DEATH WAS PROVIDED AS: PATIENT DIED DUE TO INCESSANT VF SECONDARY TO EFFUSION THAT WAS UNABLE TO BE RESOLVED DURING LEADLESS IPG IMPLANT TO REPLACE THE RV LEAD. THE LEADLESS IPGS WERE ATTEMPTED/NOT USED. THE RV LEAD REMAINS IN THE PATIENT. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6)). D9: <(><<)>NO> MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22106546&pc=PNJ | 9612164-2025-02659 | 2025-05-08 04:00 | Death | MEDTRONIC IRELAND | 2025-05-29 04:00 | PNJ | MICRA¿ AV2 | Failure to Capture; Mechanical Problem; Use of Device Problem; Positioning Problem | Cardiac Arrest; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Ventricular Fibrillation; Pericardial Effusion; Thrombosis/Thrombus; Asystole | P150033 |  | 1 | Event Description: THE PATIENT PRESENTED TO HOSPITAL WITH SYNCOPE. NORMAL DEVICE FUNCTION, NORMAL THRESHOLDS AND CLEAR CAPTURE WAS NOTED ON THE VENTRICULAR ELECTROGRAM (EGM). IT WAS REPORTED THAT LATER, THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED LOSS OF VENTRICULAR CAPTURE ON TELEMETRY. VARYING RV CAPTURE THRESHOLDS WERE NOTED ON INTERROGATION. THE RV LEAD WAS REPROGRAMMED TO ENSURE CAPTURE UNTIL A MORE PERMANENT FIX COULD BE ACHIEVED. THE PATIENT WAS TRANSFERRED TO ANOTHER FACILITY FOR IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) AS A REPLACEMENT FOR THE RV LEAD. THE LEADLESS IPG WAS DEPLOYED AND CAPTURE THRESHOLD WAS NOTED TO BE WITHIN NORMAL LIMITS. THREE TINES WERE NOTED TO BE ENGAGED DURING THE PULL AND HOLD TEST. THE TETHER OF THE DELIVERY SYSTEM WAS CUT AND THRESHOLD INCREASED. THE IPG COULD NOT BE RECAPTURED AND TETHER REMOVAL WAS COMPLETED. THE THRESHOLD DECREASED GRADUALLY BUT IT WAS DECIDED TO IMPLANT A SECOND LEADLESS IPG TO GET A BETTER THRESHOLD AS THE PATIENT WAS PACEMAKER DEPENDENT. BEFORE IMPLANT OF THE SECOND IPG, AN ATTEMPT WAS MADE TO SNARE THE FIRST IPG. IT WAS DIFFICULT TO REACH THE IPG. A GOOSENECK SNARE WAS ENGAGED AROUND THE MAIN BODY OF THE IPG AND IT WASN'T COAXIAL TO BE BROUGHT INTO THE SHEATH. A SECOND SNARE WAS ATTEMPTED TO SNARE THE RETRIEVAL BUTTON. HOWEVER, DURING THIS PROCESS THE DELIVERY SHEATH HAD BEEN PULLED BACK TO WHERE IT WAS NEARLY OUT OF THE BODY AND SHEATH POSITION WAS WATCHED SO IT WOULD NOT EXIT THE BODY WITHOUT THE SNARED IPG. HOWEVER, THE SHEATH WAS ACCIDENTALLY PULLED OUT AND THE IPG WAS LEFT INSIDE THE BODY. AT THIS POINT, THE PATIENT'S BLOOD PRESSURE WAS DOWN. A QUICK ECHO WAS PERFORMED WHICH SHOWED NO EFFUSION AND THE PRESSURE DROP WAS ATTRIBUTED TO BLOOD LOSS OUT OF THE INTRODUCER SHEATH DUE TO MULTIPLE CATHETERS AND SNARES INSIDE. AN ARTERIAL LINE WAS THEN PLACED AND THE PATIENT WAS INTUBATED AND VASCULAR SURGERY WAS PAGED FOR BACKUP. NEXT, JUGULAR ACCESS WAS OBTAINED TO ATTEMPT TO SNARE THE IPG, WHICH WAS IN THE PATIENT'S GROIN NEAR WHERE THE FEMORAL SHEATH HAD EXITED THE BODY. A WIRE WAS ADVANCED ALL THE WAY TO THE FEMORAL VEIN AND THE SHEATH WAS ADVANCED, HOWEVER THE WIRE NEVER REACHED THE LOCATION OF THE IPG. A VENOGRAM WAS PERFORMED WHICH REVEALED THE IPG WAS OUTSIDE OF THE VEIN. AT THIS POINT, IT WAS ESTABLISHED THAT IPG WOULD NEED TO BE REMOVED SURGICALLY. HOWEVER, SINCE THERE WAS ALREADY ACCESS IN THE JUGULAR VEIN, IT WAS DECIDED TO IMPLANT THE SECOND LEADLESS IPG VIA THE JUGULAR APPROACH. THE SECOND LEADLESS IPG WAS ADVANCED TO THE RV. AFTER DEPLOYMENT THERE WAS NO CAPTURE AT MAXIMUM OUTPUT. THE SECOND IPG WAS THEN RECAPTURED WITH SOME DIFFICULTY AND WAS REMOVED FROM THE BODY. AT THIS POINT, THE PRESSURE DROPPED. AN INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER WAS INSERTED VIA GROIN SHEATH WHICH SHOWED LARGE EFFUSION. EMERGENT PERICARDIOCENTESIS WAS PERFORMED WITH BLOOD BEING WITHDRAWN AND READMINISTERED VIA THE JUGULAR ACCESS. HOWEVER, THE EFFUSION WAS NOT ABLE TO BE DECREASED RAPIDLY ENOUGH TO MAKE A DIFFERENCE IN THE EFFUSION. CODE BLUE WAS CALLED. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF) AND DEFIBRILLATION WAS ATTEMPTED UNSUCCESSFULLY. A SECOND EXTERNAL DEFIBRILLATOR WAS BROUGHT IN AND BOTH EXTERNAL DEFIBRILLATORS WERE USED TO DELIVER SIMULTANEOUS 360 JOULE SHOCKS. HOWEVER, VF WAS UNABLE TO BE TERMINATED. ON THE ICE CATHETER A LEFT VENTRICULAR (LV) THROMBUS WAS SEEN. CPR WAS STOPPED AND PATIENT WAS DECLARED DECEASED. CAUSE OF DEATH WAS PROVIDED AS: PATIENT DIED DUE TO INCESSANT VF SECONDARY TO EFFUSION THAT WAS UNABLE TO BE RESOLVED DURING LEADLESS IPG IMPLANT TO REPLACE THE RV LEAD. THE LEADLESS IPGS WERE ATTEMPTED/NOT USED. THE RV LEAD REMAINS IN THE PATIENT. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6)). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22106565&pc=DYB | 9612164-2025-02660 | 2025-05-08 04:00 | Death | MEDTRONIC IRELAND | 2025-05-29 04:00 | DYB | MICRA | Mechanical Problem; Use of Device Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Ventricular Fibrillation; Pericardial Effusion; Thrombosis/Thrombus; Asystole | K132030 |  | 1 | Event Description: THE PATIENT PRESENTED TO HOSPITAL WITH SYNCOPE. NORMAL DEVICE FUNCTION, NORMAL THRESHOLDS AND CLEAR CAPTURE WAS NOTED ON THE VENTRICULAR ELECTROGRAM (EGM). IT WAS REPORTED THAT LATER, THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED LOSS OF VENTRICULAR CAPTURE ON TELEMETRY. VARYING RV CAPTURE THRESHOLDS WERE NOTED ON INTERROGATION. THE RV LEAD WAS REPROGRAMMED TO ENSURE CAPTURE UNTIL A MORE PERMANENT FIX COULD BE ACHIEVED. THE PATIENT WAS TRANSFERRED TO ANOTHER FACILITY FOR IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) AS A REPLACEMENT FOR THE RV LEAD. THE LEADLESS IPG WAS DEPLOYED AND CAPTURE THRESHOLD WAS NOTED TO BE WITHIN NORMAL LIMITS. THREE TINES WERE NOTED TO BE ENGAGED DURING THE PULL AND HOLD TEST. THE TETHER OF THE DELIVERY SYSTEM WAS CUT AND THRESHOLD INCREASED. THE IPG COULD NOT BE RECAPTURED AND TETHER REMOVAL WAS COMPLETED. THE THRESHOLD DECREASED GRADUALLY BUT IT WAS DECIDED TO IMPLANT A SECOND LEADLESS IPG TO GET A BETTER THRESHOLD AS THE PATIENT WAS PACEMAKER DEPENDENT. BEFORE IMPLANT OF THE SECOND IPG, AN ATTEMPT WAS MADE TO SNARE THE FIRST IPG. IT WAS DIFFICULT TO REACH THE IPG. A GOOSENECK SNARE WAS ENGAGED AROUND THE MAIN BODY OF THE IPG AND IT WASN'T COAXIAL TO BE BROUGHT INTO THE SHEATH. A SECOND SNARE WAS ATTEMPTED TO SNARE THE RETRIEVAL BUTTON. HOWEVER, DURING THIS PROCESS THE DELIVERY SHEATH HAD BEEN PULLED BACK TO WHERE IT WAS NEARLY OUT OF THE BODY AND SHEATH POSITION WAS WATCHED SO IT WOULD NOT EXIT THE BODY WITHOUT THE SNARED IPG. HOWEVER, THE SHEATH WAS ACCIDENTALLY PULLED OUT AND THE IPG WAS LEFT INSIDE THE BODY. AT THIS POINT, THE PATIENT'S BLOOD PRESSURE WAS DOWN. A QUICK ECHO WAS PERFORMED WHICH SHOWED NO EFFUSION AND THE PRESSURE DROP WAS ATTRIBUTED TO BLOOD LOSS OUT OF THE INTRODUCER SHEATH DUE TO MULTIPLE CATHETERS AND SNARES INSIDE. AN ARTERIAL LINE WAS THEN PLACED AND THE PATIENT WAS INTUBATED AND VASCULAR SURGERY WAS PAGED FOR BACKUP. NEXT, JUGULAR ACCESS WAS OBTAINED TO ATTEMPT TO SNARE THE IPG, WHICH WAS IN THE PATIENT'S GROIN NEAR WHERE THE FEMORAL SHEATH HAD EXITED THE BODY. A WIRE WAS ADVANCED ALL THE WAY TO THE FEMORAL VEIN AND THE SHEATH WAS ADVANCED, HOWEVER THE WIRE NEVER REACHED THE LOCATION OF THE IPG. A VENOGRAM WAS PERFORMED WHICH REVEALED THE IPG WAS OUTSIDE OF THE VEIN. AT THIS POINT, IT WAS ESTABLISHED THAT IPG WOULD NEED TO BE REMOVED SURGICALLY. HOWEVER, SINCE THERE WAS ALREADY ACCESS IN THE JUGULAR VEIN, IT WAS DECIDED TO IMPLANT THE SECOND LEADLESS IPG VIA THE JUGULAR APPROACH. THE SECOND LEADLESS IPG WAS ADVANCED TO THE RV. AFTER DEPLOYMENT THERE WAS NO CAPTURE AT MAXIMUM OUTPUT. THE SECOND IPG WAS THEN RECAPTURED WITH SOME DIFFICULTY AND WAS REMOVED FROM THE BODY. AT THIS POINT, THE PRESSURE DROPPED. AN INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER WAS INSERTED VIA GROIN SHEATH WHICH SHOWED LARGE EFFUSION. EMERGENT PERICARDIOCENTESIS WAS PERFORMED WITH BLOOD BEING WITHDRAWN AND READMINISTERED VIA THE JUGULAR ACCESS. HOWEVER, THE EFFUSION WAS NOT ABLE TO BE DECREASED RAPIDLY ENOUGH TO MAKE A DIFFERENCE IN THE EFFUSION. CODE BLUE WAS CALLED. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF) AND DEFIBRILLATION WAS ATTEMPTED UNSUCCESSFULLY. A SECOND EXTERNAL DEFIBRILLATOR WAS BROUGHT IN AND BOTH EXTERNAL DEFIBRILLATORS WERE USED TO DELIVER SIMULTANEOUS 360 JOULE SHOCKS. HOWEVER, VF WAS UNABLE TO BE TERMINATED. ON THE ICE CATHETER A LEFT VENTRICULAR (LV) THROMBUS WAS SEEN. CPR WAS STOPPED AND PATIENT WAS DECLARED DECEASED. CAUSE OF DEATH WAS PROVIDED AS: PATIENT DIED DUE TO INCESSANT VF SECONDARY TO EFFUSION THAT WAS UNABLE TO BE RESOLVED DURING LEADLESS IPG IMPLANT TO REPLACE THE RV LEAD. THE LEADLESS IPGS WERE ATTEMPTED/NOT USED. THE RV LEAD REMAINS IN THE PATIENT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22106649&pc=DSQ | 2916596-2025-03225 | 2025-04-27 04:00 | Death | THORATEC CORPORATION | 2025-05-29 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Hemorrhage/Bleeding; Intracranial Hemorrhage; Hemorrhagic Stroke; Ischemia Stroke; Unspecified Nervous System Problem | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED WITH A CHANGE IN MENTAL STATUS, CEREBRAL HEMORRHAGE, AND BACTEREMIA. THE INFECTION WAS A DIFFUSE INFECTION. METHICILLIN-SUSCEPTIBLE STAPHYLOCOCCUS AUREUS (MSSA) GREW FROM BLOOD, URINE, AND SPUTUM. MOREOVER, THERE WERE MULTIPLE AREAS OF CEREBRAL INFARCTS. NO DEVICE RELATED REASON FOR A MECHANICAL FALL WERE NOTED. THE PATIENT'S FIRST INTERNATIONAL NORMALIZED RATIO (INR) WAS WITHIN RANGE. DIAGNOSTICS INCLUDED A COMPUTED TOMOGRAPHY (CT) SCAN, BLOOD CULTURES, TRANSTHORACIC ECHOCARDIOGRAM (TTE), AND TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). PATIENT STATUS HAD QUICKLY DETERIORATED, AND FAMILY DECIDED TO WITHDRAW CARE. THE VENTRICULAR ASSIST DEVICE (VAD) WAS DECOMMISSIONED. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION H6: HEALTH EFFECT - CLINICAL CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING STROKE, INFECTION, BLEEDING AND DEATH. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6 ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. FURTHERMORE, SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22106650&pc=NPW | 2015691-2025-04239 | 2025-04-26 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-29 04:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Patient Device Interaction Problem | Heart Block | P230013 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS LIFESCIENCES AFFILIATE IN NETHERLANDS, EDWARDS RECEIVED NOTIFICATION OF A 56MM EVOQUE PROCEDURE WHERE THE PATIENT WAS DISCHARGED ON POST OPERATIVE DAY (POD) #1. ON POD #2, THE PATIENT DEVELOPED A COMPLETE ATRIOVENTRICULAR (AV) BLOCK, PRESENTING AT HOME WITH A VENTRICULAR ESCAPE RHYTHM. RESUSCITATION EFFORTS WERE UNSUCCESSFUL, AND THE PATIENT PASSED AWAY. Manufacturer Narrative: THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED/ADDED: B4, G3, G6, H2, H6 AND H11. ADDITIONAL TYPE OF INVESTIGATION CODES THAT WERE UNABLE TO BE ADDED IN H6: LABELLING REVIEW. THE COMPLAINT FOR VALVE INTERACTS WITH CONDUCTION SYSTEM WAS CONFIRMED WITH OTHER EMPIRICAL EVIDENCE VIA INFORMATION REPORTED BY EDWARDS CLINICAL SPECIALIST. NO MANUFACTURING NON-CONFORMITIES WAS FOUND DURING DHR REVIEW. A DEFINITE ROOT CAUSE IS UNABLE TO BE DETERMINED. SINCE NO MANUFACTURING NON-CONFORMANCES WERE FOUND AND NO LABELING, TRAINING, OR IFU DEFICIENCIES WERE IDENTIFIED, NO PREVENTATIVE OR CORRECTIVE ACTIONS WERE REQUIRED. ACCORDING TO THE IFU, CONDUCTION DISTURBANCES-POTENTIALLY REQUIRING TREATMENT LIKE A PERMANENT PACEMAKER-ARE KNOWN RISKS. THESE EVENTS ARE OFTEN LINKED TO PRE-EXISTING CARDIOVASCULAR CONDITIONS AND MAY BE WORSENED BY ANESTHESIA OR INTRA-OPERATIVE MEDICATIONS. OTHER POSSIBLE CAUSES INCLUDE CORONARY OBSTRUCTION (E.G., AIR EMBOLISM, SPASM, OR EDEMA NEAR THE AV NODE). TRANSCATHETER PROCEDURES CAN ALSO TRIGGER CONDUCTION ISSUES DUE TO DIRECT INTERACTION WITH CARDIAC STRUCTURES, ESPECIALLY IN PREDISPOSED PATIENTS. THE VALVE'S DESIGN, WHICH APPLIES RADIAL FORCE AND ANCHOR INTERACTION TO THE SEPTUM, MAY CONTRIBUTE TO SUCH DISTURBANCES, THOUGH A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22106652&pc=KDI | 3003464075-2025-00073 | 2025-05-05 04:00 | Death | NXSTAGE MEDICAL, INC. | 2025-05-29 04:00 | KDI | NXSTAGE SYSTEM ONE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K230632 |  | 1 | Event Description: A REPORT WAS RECEIVED ON (B)(6) 2025 FROM THE HOME THERAPY NURSE (HTN) OF A 45-YEAR-OLD MALE PATIENT WITH A MEDICAL HISTORY INCLUDING MULTIPLE COMORBIDITIES AND END STAGE RENAL DISEASE, WHO STATED THE PATIENT WAS FOUND EXPIRED AND CONNECTED TO THE DEVICE ON (B)(6) 2025. ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2025 FROM THE HOME THERAPY NURSE (HTN) WHO STATED THAT THE SUSPECTED CAUSE OF DEATH WAS CARDIAC ARREST. Manufacturer Narrative: THERE WAS NO DEVICE MALFUNCTION. THE CYCLER WAS RECEIVED FOR EVALUATION AND SUCCESSFULLY PASSED TESTING. EVALUATION OF THE AVAILABLE LOG FILES WAS UNREMARKABLE WITH NO PRODUCT DEFICIENCY IDENTIFIED. Manufacturer Narrative: THE CYCLER WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4) .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22106993&pc=OZD | 1220648-2025-28821 | 2025-05-06 04:00 | Death | ABIOMED, INC. | 2025-05-29 04:00 | OZD | IMPELLA CP | Deformation Due to Compressive Stress; Positioning Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS UNDERGOING ESCALATION FROM AN IMPELLA CP TO AN IMPELLA 5.5. REMOVAL OF THE GUIDEWIRE WAS DIFFICULT AS IT WAS KINKED. OOZING WAS NOTED FROM THE ACCESS SITE SO A JACKSON-PRATT DRAIN WAS ADDED, IN ADDITION TO A GRAFT CUTDOWN WITH SUTURE. FIVE UNITS OF BLOOD PRODUCTS WERE TRANSFUSED.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF PRODUCT DAMAGE (GUIDEWIRE DAMAGE - GREAT VESSEL), POSITIONING ISSUES, AND ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. PRODUCT DAMAGE (GUIDEWIRE DAMAGE - GREAT VESSEL): UPON WIRE REMOVAL, WIRE BECAME KINKED DISTAL TO IMPELLA INLET AND WAS UNABLE TO BE REMOVED. CLINICAL DETAIL STATE THE GW UNRAVELED BUT DO NOT MENTION THE STATE OF THE VESSELS. THE WIRE WAS NOT RETURNED. THE CAUSE OF THE PRODUCT DAMAGE (GUIDEWIRE DAMAGE) WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAIL. POSITIONING ISSUES: AFTER THE GUIDEWIRE (GW) BECAME KINKED DISTAL TO IMPELLA INLET, IMPELLA 5.5 AND WIRE AND COME BACK INTO THE AORTA AND THE IMPELLA 5.5 CATHETER WAS WRAPPED BELOW THE IMPELLA OUTLET. CATHETER WAS MANIPULATED UP AND DOWN AND TORQUED TO RELIEVE THE MALPOSITIONED CATHETER. THE PHYSICIAN STATED PUMP WAS "FINE," NO KINKS, BENDS, OR ABNORMALITIES NOTED ON INSPECTION. THE PUMP AND GW WERE NOT RETURNED. THE CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY AN INTERACTION WITH ANOTHER DEVICE BEING THE KINKED GW THAT LODGED ONTO THE IMPELLA INLET CAGE AND PULLED THE PUMP INTO THE AORTA DURING GW REMOVAL ATTEMPTS. ACCESS SITE BLEEDING - MAJOR: AFTER PLACING THE PUMP SOME OOZING NOTED FROM IMPELLA POCKET AND JP DRAIN ADDED PRIOR TO CLOSING THE SITE. PATIENTS ACT WAS NOTED TO BE AT 308 AFTER TOTAL HEPARIN OF 28000U GIVEN. THE PUMP WAS NOT RETURNED. THE CAUSE OF THE ACCESS SITE BLEEDING - MAJOR WAS MOST LIKELY HIGH PATIENT ACTIVATED CLOTTING TIMES (ACT) VALUES OF 308 AFTER 28000U OF HEPARIN GIVEN. B2 DEATH AND DATE OF DEATH WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28821. B5 PATIENT OUTCOME WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28821. H1 TYPE OF REPORT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28821. H6 HEALTH EFFECT - IMPACT CODE 1802 AND MEDICAL DEVICE PROBLEM CODE 3009 WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28821.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22107167&pc=NPT | 2025587-2025-03949 | 2025-04-30 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-29 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem; Central Regurgitation | Stroke/CVA; Low Blood Pressure/ Hypotension; Paralysis; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER AORTIC VALVE, THE GUIDEWIRE WAS DIFFICULT TO CROSS THROUGH THE AORTIC VALVE AND A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED USING A 20 MILLIMETER (MM) NON-MEDTRONIC BALLOON DUE TO CALCIFICATION AND STANDARD FOR IMPLANTING PHYSICIAN. FOLLOWING THE PRE-IMPLANT BAV, THE PATIENT CODED DUE TO "WIDE OPEN" AORTIC REGURGITATION; THEREFORE, THE TRANSCATHETER VALVE WAS IMPLANTED AT A DEPTH OF APPROXIMATELY 6 MM ON THE LEFT CORONARY CUSP (LCC) AND 6 MM ON THE NON-CORONARY CUSP (NCC). FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT HAD A MEAN GRADIENT OF 4-5 MILLIMETERS OF MERCURY (MM HG) AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED TO GET THE PATIENT'S BLOOD PRESSURE UP. FOLLOWING CPR AND ANESTHESIA, THE PATIENT RECOVERED. FOLLOWING THE IMPLANT PROCEDURE, THE PATIENT WOKE UP WITH A STABLE BLOOD PRESSURE AND WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). ON THE SAME DAY AS THE IMPLANT PROCEDURE, THE PATIENT DEVELOPED ONE-SIDED PARALYSIS. A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED THAT REVEALED SEVERE CAROTID DISEASE AND A STROKE. SUBSEQUENTLY, THE PATIENT UNDERWENT NEUROLOGY PROTOCOL. IT WAS REPORTED THAT 7 DAYS FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT DIED DUE TO STROKE COMPLICATIONS.Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER AORTIC VALVE, THE GUIDEWIRE WAS DIFFICULT TO CROSS THROUGH THE AORTIC VALVE AND A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED USING A 20 MILLIMETER (MM) NON-MEDTRONIC (TRUE) BALLOON DUE TO CALCIFICATION AND STANDARD FOR IMPLANTING PHYSICIAN. FOLLOWING THE PRE-IMPLANT BAV, THE PATIENT CODED DUE TO "WIDE OPEN" AORTIC REGURGITATION; THEREFORE, THE TRANSCATHETER VALVE WAS IMPLANTED AT A DEPTH OF APPROXIMATELY 6 MM ON THE LEFT CORONARY CUSP (LCC) AND 6 MM ON THE NON-CORONARY CUSP (NCC). FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT HAD A MEAN GRADIENT OF 4-5 MILLIMETERS OF MERCURY (MM HG) AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED TO GET THE PATIENT'S BLOOD PRESSURE UP. FOLLOWING CPR AND ANESTHESIA, THE PATIENT RECOVERED. FOLLOWING THE IMPLANT PROCEDURE, THE PATIENT WOKE UP WITH A STABLE BLOOD PRESSURE AND WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). ON THE SAME DAY AS THE IMPLANT PROCEDURE, THE PATIENT DEVELOPED ONE-SIDED PARALYSIS. A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED THAT REVEALED SEVERE CAROTID DISEASE AND A STROKE. SUBSEQUENTLY, THE PATIENT UNDERWENT NEUROLOGY PROTOCOL. IT WAS REPORTED THAT 7 DAYS FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT DIED DUE TO STROKE COMPLICATIONS. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE GUIDEWIRE WAS A NON-MEDTRONIC PRODUCT (LUNDERQUIST). THE REGURGITATION WAS AGAIN DESCRIBED AS WIDE OPEN, "ENOUGH TO MAKE THIS PATIENT DECOMPENSATE QUICKLY." IT IS UNKNOWN WHETHER THE STROKE WAS ISCHEMIC OR HEMORRHAGIC. THE PHYSICIAN DOESN'T BELIEVE THE STROKE WAS RELATED TO THE VALVE, BUT NO WAY TO TRULY TELL. LIKEWISE, THE PHYSICIAN WAS NOT SURE IF THE STROKE WAS RELATED TO THE DELIVERY CATHETER SYSTEM. ACCORDING TO THE PHYSICIAN, THE STROKE COULD HAVE BEEN RELATED TO MULTIPLE FACTORS, SUCH AS THE BAV, CPR, AND THE PATIENT'S CALCIFIED ANATOMY AND CAROTID DISEASE. THE PHYSICIAN ALSO STATED THAT IT WAS POSSIBLE THAT THE PATIENT'S UNDERLYING MEDICAL HISTORY CAUSED OR CONTRIBUTED TO THE DEATH. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {D-EVOLUTFX-2329}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HEARTVALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} PRODUCT ID {NON-MEDTRONIC PRE-IMPLANT BAV}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {BALLOON AORTIC VALVULOPLASTY}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} PRODUCT ID {GUIDEWIRE};PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {GUIDEWIRE}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 (SECOND PARAGRAPH). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22107179&pc=DQX | 2025587-2025-03950 | 2025-04-30 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-29 04:00 | DQX | CONFIDA BRECKER CURVE GUIDEWIRE | Patient Device Interaction Problem | Stroke/CVA; Low Blood Pressure/ Hypotension; Paralysis; Aortic Valve Insufficiency/ Regurgitation | K132623 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER AORTIC VALVE, THE GUIDEWIRE WAS DIFFICULT TO CROSS THROUGH THE AORTIC VALVE AND A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED USING A 20 MILLIMETER (MM) NON-MEDTRONIC BALLOON DUE TO CALCIFICATION AND STANDARD FOR IMPLANTING PHYSICIAN. FOLLOWING THE PRE-IMPLANT BAV, THE PATIENT CODED DUE TO "WIDE OPEN" AORTIC REGURGITATION; THEREFORE, THE TRANSCATHETER VALVE WAS IMPLANTED AT A DEPTH OF APPROXIMATELY 6 MM ON THE LEFT CORONARY CUSP (LCC) AND 6 MM ON THE NON-CORONARY CUSP (NCC). FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT HAD A MEAN GRADIENT OF 4-5 MILLIMETERS OF MERCURY (MM HG) AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED TO GET THE PATIENT'S BLOOD PRESSURE UP. FOLLOWING CPR AND ANESTHESIA, THE PATIENT RECOVERED. FOLLOWING THE IMPLANT PROCEDURE, THE PATIENT WOKE UP WITH A STABLE BLOOD PRESSURE AND WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). ON THE SAME DAY AS THE IMPLANT PROCEDURE, THE PATIENT DEVELOPED ONE-SIDED PARALYSIS. A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED THAT REVEALED SEVERE CAROTID DISEASE AND A STROKE. SUBSEQUENTLY, THE PATIENT UNDERWENT NEUROLOGY PROTOCOL. IT WAS REPORTED THAT 7 DAYS FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT DIED DUE TO STROKE COMPLICATIONS.Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER AORTIC VALVE, THE GUIDEWIRE WAS DIFFICULT TO CROSS THROUGH THE AORTIC VALVE AND A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED USING A 20 MILLIMETER (MM) NON-MEDTRONIC (TRUE) BALLOON DUE TO CALCIFICATION AND STANDARD FOR IMPLANTING PHYSICIAN. FOLLOWING THE PRE-IMPLANT BAV, THE PATIENT CODED DUE TO "WIDE OPEN" AORTIC REGURGITATION; THEREFORE, THE TRANSCATHETER VALVE WAS IMPLANTED AT A DEPTH OF APPROXIMATELY 6 MM ON THE LEFT CORONARY CUSP (LCC) AND 6 MM ON THE NON-CORONARY CUSP (NCC). FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT HAD A MEAN GRADIENT OF 4-5 MILLIMETERS OF MERCURY (MM HG) AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED TO GET THE PATIENT'S BLOOD PRESSURE UP. FOLLOWING CPR AND ANESTHESIA, THE PATIENT RECOVERED. FOLLOWING THE IMPLANT PROCEDURE, THE PATIENT WOKE UP WITH A STABLE BLOOD PRESSURE AND WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). ON THE SAME DAY AS THE IMPLANT PROCEDURE, THE PATIENT DEVELOPED ONE-SIDED PARALYSIS. A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED THAT REVEALED SEVERE CAROTID DISEASE AND A STROKE. SUBSEQUENTLY, THE PATIENT UNDERWENT NEUROLOGY PROTOCOL. IT WAS REPORTED THAT 7 DAYS FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT DIED DUE TO STROKE COMPLICATIONS. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE GUIDEWIRE WAS A NON-MEDTRONIC PRODUCT (LUNDERQUIST). THE REGURGITATION WAS AGAIN DESCRIBED AS WIDE OPEN, "ENOUGH TO MAKE THIS PATIENT DECOMPENSATE QUICKLY." IT IS UNKNOWN WHETHER THE STROKE WAS ISCHEMIC OR HEMORRHAGIC. THE PHYSICIAN DOESN'T BELIEVE THE STROKE WAS RELATED TO THE VALVE, BUT NO WAY TO TRULY TELL. LIKEWISE, THE PHYSICIAN WAS NOT SURE IF THE STROKE WAS RELATED TO THE DELIVERY CATHETER SYSTEM. ACCORDING TO THE PHYSICIAN, THE STROKE COULD HAVE BEEN RELATED TO MULTIPLE FACTORS, SUCH AS THE BAV, CPR, AND THE PATIENT'S CALCIFIED ANATOMY AND CAROTID DISEASE. THE PHYSICIAN ALSO STATED THAT IT WAS POSSIBLE THAT THE PATIENT'S UNDERLYING MEDICAL HISTORY CAUSED OR CONTRIBUTED TO THE DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 (SECOND PARAGRAPH). THIS EVENT NO LONGER MEETS COMPLAINT OR REPORTABILITY CRITERIA FOR THE MEDTRONIC CONFIDA GUIDEWIRE DEVICE AS THE GUIDEWIRE USED DURING THIS CASE WAS CONFIRMED TO BE A NON-MEDTRONIC PRODUCT (LUNDERQUIST). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ID {D-EVOLUTFX-2329};; PRODUCT TYPE: {0195-HEARTVALVES}; PRODUCT ID {NON-MEDTRONIC PRE-IMPLANT BAV}; PRODUCT TYPE: {BALLOON AORTIC VALVULOPLASTY}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A} PRODUCT ID {EVFXPLUS-29};PRODUCT LOT/SERIAL NUMBER (B)(6); PRODUCT TYPE: {0195-HEARTVALVES}; IMPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22107215&pc=DXF | 2124215-2025-35058 | 2025-05-09 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-05-29 04:00 | DXF | VERSACROSS CONNECT LAAC ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Pleural Effusion; Cardiac Perforation | K150709 |  | 1 | Event Description: IT WAS REPORTED A PATIENT DEATH OCCURRED. DURING A LEFT ATRIAL APPENDAGE OCCLUSION (LAAO) PROCEDURE, A VERSACROSS CONNECT LAAC KIT WAS SELECTED FOR USE. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS BEING USED FOR IMAGING AND IT WAS CONFIRMED THAT THERE WAS NO EFFUSION WAS NOTED PRIOR TO PROCEDURE. THE PHYSICIAN PERFORMED TRANSSEPTAL PUNCTURE (TSP) WITH NO DIFFICULTY REPORTED, ONCE THE PATIENT STARTED TO DECLINE. THE PHYSICIAN THEN NOTICED THE PATIENT ACCUMULATED FLUID IN RIGHT LUNG CAVITY. THERE WAS A PERFORATION INTO CHEST CAVITY, THAT HAVE OCCURRED IN THE LEFT ATRIUM WHILE GOING TRANSSEPTAL. HENCE, A SURGERY WAS CALLED. THE PATIENT PASSED AWAY. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). IN THE PHYSICIAN'S OPINION, THE TRANSSEPTAL PORTION OF THE PROCEDURE CONTRIBUTE TO THE PATIENT COMPLICATION. THE WATCHMAN DEVICE WAS NOT IMPLANTED. ACT WAS THERAPEUTIC. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22107287&pc=DSQ | 2916596-2025-03173 | 2025-03-28 04:00 | Death | THORATEC CORPORATION | 2025-05-29 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Dyspnea; Fatigue; Nausea; Vomiting; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED ON (B)(6) 2025 AND STARTED DOBUTAMINE ON (B)(6) 2025. WITHDRAWAL OF CARE WAS SECONDARY TO RIGHT HEART FAILURE AND FAILURE TO THRIVE. THE PATIENT PASSED AWAY IN THE HOSPITAL ON (B)(6) 2025. AN AUTOPSY WAS NOT PERFORMED. THE LEFT VENTRICULAR ASSIST DEVICE FUNCTIONED AS INTENDED. PER THE SITE, THE OUTCOME WAS NOT DEVICE RELATED.Event Description: RIGHT HEART FAILURE EXISTED PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) TO SOME DEGREE. A DEVICE RELATED ISSUE WAS NOT CONSIDERED TO HAVE CAUSED OR CONTRIBUTED TO RIGHT HEART FAILURE. THE PATIENT HAD EXHIBITED SHORTNESS OF BREATH AND OCCASIONAL FOAMY EMESIS. AT ADMISSION, THE PATIENT WAS ON BROAD SPECTRUM ANTIBIOTICS AND DIURESIS.Event Description: IT WAS ADDITIONALLY REPORTED THAT RIGHT HEART FAILURE WAS DUE TO NORMAL PROGRESSION OF THE DISEASE. THE SOURCE OF INFECTION WAS METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) PNEUMONIA (PNA). SYMPTOMS INCLUDED COUGH AND SHORTNESS OF AIR (SOA). THE PATIENT HAD A CHRONIC ONGOING COUGH FOR GREATER THAN 6 MONTHS AND STATED THAT THEIR COUGH WORSENED IN THE LAST 24 HOURS. THEY ALSO REPORTED NAUSEA, VOMITING, AND EXTREME FATIGUE. THE EVENT WAS TREATED WITH BROAD SPECTRUM ANTIBIOTICS. PATIENT STATUS HAD DETERIORATED, AND SUPPORT WAS WITHDRAWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE OR PUMP POCKET INFECTION), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING INFECTION PREVENTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22107356&pc=OZD | 1220648-2025-28823 | 2025-05-17 04:00 | Death | ABIOMED, INC. | 2025-05-29 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT ON IMPELLA CP SUPPORT THAT A VENTRICULAR SEPTAL PERFORATION (VSP) WAS CONFIRMED A FEW DAYS AFTER PERCUTANEOUS CORONARY INTERVENTION. THE DOCTOR BELIEVED THAT THE VSP WAS NOT CAUSED BY IMPELLA, BUT THE PERFORATION OCCURRED DURING SUPPORT. LATER THE PATIENT EXPIRED FROM MULTIPLE ORGAN FAILURE. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22107384&pc=HGM | 9610816-2025-000443 | 2025-02-19 05:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-05-29 04:00 | HGM | AVALON ULTRASOUND TRANSDUCER | Erratic Results | Hypoxia in Utero | K140535 |  | 1 | Event Description: IT WAS REPORTED FOLLOWING A FETAL DEATH, THE USER WAS UNCERTAIN WHETHER FETAL AND MATERNAL PULSES WERE SWITCHED DURING DELIVERY. THE BABY WAS DELIVERED VIA C-SECTION BUT WAS STILLBORN; HOWEVER, THERE WAS A FETAL HEART RATE PRESENT ON THE TRACE. IT WAS THOUGHT BY THE USERS THAT THERE WERE HEART RATES APPEARING FOR BOTH MOTHER AND BABY DURING DELIVERY. IT WAS ALSO INDICATED THE CTG TRACING WAS TURNED OFF AND 50 MINUTES LATER THE C-SECTION WAS PERFORMED, REVEALING THE STILLBORN BABY. IT REMAINS UNCLEAR TO THE CUSTOMER IF THERE WAS A MISINTERPRETATION OF THE HEART RATE TRACE. ADDITIONAL INFORMATION WAS RECEIVED. REVIEW OF THE DATA PROVIDED WAS PERFORMED BY A CLINICAL SCIENTIST, REVEALING THAT MONITORING OF THE FETUS WAS STOPPED APPROXIMATELY 30 MINUTES PRIOR TO THE STILLBIRTH. THE FETUS WAS DELIVERED VIA C-SECTION DUE TO THE MOTHER HAVING A RUPTURED UTERUS, WHICH MAY HAVE LED TO FETAL HYPOXIA IN THAT 30-MINUTE WINDOW OF TIME. MONITORING WAS RESTARTED JUST PRIOR TO THE C-SECTION AND THE FETAL HEART RATE WAS CLEARLY INDICATED BY OTHER MEANS SUCH AS ULTRASOUND. BASED ON THIS INFORMATION, THE CAUSE OF THE FETAL DEMISE REMAINS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED FROM THE CUSTOMER. PER THE CUSTOMER, IT WAS DETERMINED THE DEATH WAS DUE TO UTERINE RUPTURE DURING LABOR. ALTHOUGH THE CAUSE OF DEATH WAS UTERINE RUPTURE DURING LABOR AND DELIVERY, IT REMAINS UNKNOWN WHETHER THE USERS MAY HAVE MISTAKEN MATERIAL HEART RATE (MHR) FOR FETAL HEART RATE (FHR) DURING THE EVENT, WHICH MAY HAVE PREVENTED THE USERS FROM SEEING THE CHANGE IN FETAL CONDITION. THE MOTHER AT 50 WEEKS PLUS 5 DAYS GESTATION AND EXPECTING HER FIFTH CHILD PRESENTED TO LABOR AND DELIVERY WITH CONTRACTIONS. LABOR PROGRESSION WAS SLOW WITH A LOT OF PAIN. THERE WAS CONTINUOUS EXTERNAL FETAL MONITORING WITH NO INTERNAL MONITOR DUE TO HER PRE-EXISTING HIV INFECTION. AT APPROXIMATELY 1630, THE CONTRACTIONS DECREASED AND THERE WAS CHANGE IN THE CTG, BUT NO DECELERATIONS WERE PRESENT. TWO HOURS LATER, IT WAS DETERMINED A C-SECTION WOULD BE PERFORMED DUE TO LACK OF LABOR PROGRESS. THROUGHOUT THIS TIME, CTG WAS REGISTERING WITH GOOD QUALITY. AT 1730, AN ULTRASOUND WAS PERFORMED, REVEALING NORMAL FETAL HEARTBEAT. BETWEEN 1950 AND 2240, A TOCO WITH A CORD WAS USED BUT THE MHR WAS NOT RECORDED SO THE STAFF CHANGED BACK TO A CORDLESS TOCO. DUE TO THE NORMAL CTG, A UTERINE RUPTURE WAS NOT SUSPECTED. WHEN THE BABY WAS DELIVERED, THERE WAS NO CARDIAC ACTIVITY AND RESUSCITATION WAS UNSUCCESSFUL. IT WAS THEN DETERMINED THE UTERUS HAD RUPTURED, WITH A LOT OF BLOOD PRESENT IN THE PATIENT'S ABDOMEN. THE CUSTOMER INDICATED THERE WAS NO MALFUNCTION OF THE DEVICE; THE QUESTION THEY HAD WAS WHETHER IT IS ACTUALLY POSSIBLE TO GET SUCH A GOOD REGISTRATION FOR SUCH A LONG TIME OF THE MOTHER´S PULSE AND MISTAKE IT FOR THE CHILD´S. A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) REVIEWED PATIENT STRIPS. THE PSE INDICATED THAT IT IS POSSIBLE TO GET GOOD REGISTRATION FOR AN EXTENDED PERIOD OF TIME. LOOKING AT THE STRIPS, THEY COULD ALSO SEE WHERE THE MOTHER AND BABY¿S SIGNALS MERGED (AROUND 15:00); THEY ALSO NOTED THAT THE BABY SEEMED TO BE ALIVE TO THE END OF THE STRIPS WE WERE PROVIDED (WHICH ENDED SOON AFTER 23:20). THIS WOULD HAVE PROVIDED A ¿COINCIDENCE ALARM¿, WHICH WOULD SHOW AS A ¿?¿ ON THE STRIPS. HOWEVER, THE PSE WAS UNFAMILIAR WITH A MILOU SYSTEM SO COULD THEY NOT SAY IF THAT WOULD SHOW THE SAME DETAIL. THIS INFORMATION WAS PROVIDED TO THE CUSTOMER. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS PATIENT RELATED. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. Manufacturer Narrative: IT WAS DETERMINED THAT THE AVALON BASE STATION WAS INCORRECTLY REPORTED. THE CORRECT DEVICE SHOULD HAVE BEEN THE WIRELESS FETAL TRANSDUCER (867246 AVALON US TRANSDUCER). CORRECTIONS MADE TO BOX D AND BOX G. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. GOOD FAITH EFFORTS ARE BEING MADE TO OBTAIN INFORMATION RELATING TO THE SERIAL NUMBER. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22107683&pc=DSQ | 2916596-2025-03052 | 2025-03-05 05:00 | Death | THORATEC CORPORATION | 2025-05-29 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Dyspnea; High Blood Pressure/ Hypertension; Heart Failure/Congestive Heart Failure; Swelling/ Edema | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT SYSTEM CONTROLLER INTERROGATION REVEALED THAT THE PATIENT HAD 20 LOW FLOW ALARMS ON THE MORNING OF (B)(6) 2025. LOG FILE REVIEW REVEALED THAT THE PATIENT HAD PERSISTENT UNSUSTAINED LOW FLOW ESTIMATES. THE ESTIMATED FLOW APPEARED TO BE INTERMITTENTLY FLUCTUATING BELOW THE 2.0 LPM ALARM THRESHOLD. THE LOW FLOW ESTIMATES WERE NOTED TO BE TOO BRIEF TO SHOW UP AS UP AS HAZARD ALARMS. IT WAS LATER REPORTED THAT THE PATIENT WAS RE-ADMITTED TO THE HOSPITAL ON (B)(6) 2025, AND DISCHARGED ON (B)(6) 2025 DUE TO RIGHT VENTRICLE FAILURE, WITH LOW FLOW ALARMS AND SHORTNESS OF BREATH. AT THIS TIME, MIDODRINE WAS RESUMED, A MILRINONE DRIP WAS INITIATED AND THE PATIENT WAS DISCHARGED WITH HOME MILRINONE. IT WAS ALSO REPORTED THAT THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU) ON (B)(6) FOR PROFOUND RIGHT VENTRICLE FAILURE DESPITE THE HOME MILRINONE. THE PATIENT WAS ALSO NOTED TO BE HYPERTENSIVE, WITH SUDDEN SHORTNESS OF BREATH AND MILD LOWER EXTREMITY EDEMA. THE PATIENT RECEIVED BILEVEL POSITIVE AIRWAY PRESSURE (BIPAP) TREATMENT WHILE IN THE ICU AND EVENTUALLY HAD AN IMPELLA RP FLEX PLACED. THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL FOR HEART/LUNG TRANSPLANT EVALUATION. THE PATIENT'S RIGHT HEART FAILURE WAS REPORTED TO EXIST PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT, AND THAT THE LVAD DID NOT CONTRIBUTE TO THE RIGHT HEART FAILURE. LOG FILES CAPTURED LOW FLOW ALARMS ON (B)(6) 2025, AS WELL AS 118 LOW FLOW EVENTS THAT WERE TOO BRIEF TO GENERATE AN ALARM. THE PATIENT'S AVERAGE FLOW APPEARED TO BE AROUND 2.0 LPM DURING THE SPAN OF TIME COVERED BY THE LOG FILES. THERE DID NOT APPEAR TO BE ANY TYPE OF EXTERNAL MECHANICAL CIRCULATORY SUPPORT EQUIPMENT ISSUES CAUSING THESE EVENTS. ADDITIONAL INFORMATION REPORTED THAT THE LOW FLOWS STOPPED ONCE THE PATIENT WAS ADMITTED AND THAT THE LOW FLOWS WERE NOT THOUGHT TO BE RELATED TO THE RIGHT VENTRICLE FAILURE AND MAY BE RELATED TO THE PATIENT'S HYPERTENSION. IT WAS LATER REPORTED THROUGH PATIENT OUTCOME THAT THE PATIENT PASSED AWAY DUE TO RIGHT VENTRICLE FAILURE. THE DEVICE WAS NOT EXPLANTED. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED, AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION B3: EVENT DATE UPDATED MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THE ALARMS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SUBMITTED SYSTEM CONTROLLER EVENT LOG FILES COLLECTIVELY CONTAINED DATA ON (B)(6) 2025, AND (B)(6) 2025 THROUGH (B)(6) 2025. THE MAJORITY OF THE FILES CONSISTED OF INTERMITTENT LOW FLOW FAULTS, MULTIPLE OF WHICH RESULTED IN LOW FLOW ALARMS ON (B)(6) 2025. PULSATILITY INDEX WAS ELEVATED AT THE TIME OF THE LOW FLOWS. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE DEVICE APPEARED TO OPERATE AS INTENDED AT THE SET SPEED. THE PATIENT EXPIRED ON (B)(6) 2025. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED, AND THE DEVICE OPERATED AS EXPECTED. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING HYPERTENSION AND DEATH. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR,¿ EXPLAINS THAT THE LOW FLOW HAZARD ALARM OCCURS WHEN PUMP FLOW IS LESS THAN 2.5 LPM. A 10-SECOND DELAY IS IMPOSED BETWEEN THE DETECTION OF THE LOW FLOW STATUS AND THE ACTIVATION OF THE ASSOCIATED AUDIO AND VISUAL INDICATORS ON THE SYSTEM CONTROLLER. CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING,¿ EXPLAINS ALL SYSTEM ALARMS AND THE RECOMMENDED ACTIONS ASSOCIATED WITH THEM. THE HEARTMATE 3 LVAS PATIENT HANDBOOK ALSO OUTLINES ALL SYSTEM CONTROLLER ALARMS AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. FOLLOWING SOFTWARE UPGRADES, THE HEARTMATE 3 LVAS POCKET GUIDES TO ALARMS FOR PATIENTS AND CLINICIANS EXPLAIN THAT THE LOW FLOW HAZARD ALARM WILL BE TRIGGERED WHEN THE ESTIMATED PUMP FLOW IS LESS THAN 2.0 LPM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22107827&pc=DSQ | 2916596-2025-03278 | 2025-05-05 04:00 | Death | THORATEC CORPORATION | 2025-05-29 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WOULD NOT RETURN. THE DEVICE OPERATED AS EXPECTED AND THE DEATH WAS NOT CONSIDERED DEVICE RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WILL NOT COMMUNICATE ANY FURTHER PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22107900&pc=KDI | 0008030665-2025-01217 | 2025-01-01 05:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-05-29 04:00 | KDI | 2008K2 HEMODIALYSIS SYS. OLC/DIASAFE PLS | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K153449 |  | 1 | Event Description: IT WAS REPORTED THAT A HOME HEMODIALYSIS (HD) PATIENT WAS IN TREATMENT WHEN THEY REPORTED NOT FEELING WELL (UNSPECIFIED). THE PATIENT WANTED TO END THE TREATMENT. IT WAS REPORTED THAT THE NURSE WANTED THE PATIENT TO CONTINUE WITH THE TREATMENT. THE PATIENT THEN WENT INTO CARDIAC ARREST. THE PATIENT WAS DISCONNECTED FROM THE MACHINE AND PASSED AWAY. IT IS BELIEVED THAT EMERGENCY MEDICAL SERVICES (EMS) ARRIVED AT THE PATIENT¿S HOME AND ATTEMPTED TO REVIVE THE PATIENT. ACCORDING TO THE PATIENT¿S SPOUSE, THE PATIENT HAD CARDIAC AND OTHER HEALTH ISSUES. NO ALARMS OR ISSUES WERE REPORTED DURING THE DIALYSIS TREATMENT. THE DATE OF THE EVENT WAS NOT PROVIDED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION, AND THUS FAR NO FURTHER DETAILS HAVE BEEN PROVIDED. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN HOME HEMODIALYSIS UTILIZING THE 2008K2 HEMODIALYSIS MACHINE AND THE PATIENT EVENT OF CARDIAC ARREST AND DEATH. HOWEVER, THERE IS NO DOCUMENTATION TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND USE OF THE 2008K2 MACHINE. ADDITIONALLY, THERE IS NO ALLEGATION OF A MACHINE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE PATIENT WAS REPORTEDLY NOT FEELING WELL DURING TREATMENT PRIOR TO GOING INTO CARDIAC ARREST. THE PATIENT¿S SPOUSE DID REPORT THAT THE PATIENT HAD CARDIAC AND OTHER HEALTH ISSUES. CARDIAC ARREST ACCOUNTS FOR A QUARTER OF DEATHS AMONG PATIENTS ON DIALYSIS. THE PROGRESSION OF CARDIOVASCULAR DISEASES IN DIALYSIS PATIENTS IS AFFECTED NOT ONLY BY TRADITIONAL RISK FACTORS BUT ALSO BY OTHER FACTORS ASSOCIATED WITH UREMIA. BASED ON THE LIMITED INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE 2008K2 HEMODIALYSIS MACHINE CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S CARDIAC ARREST AND DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22108245&pc=DSQ | 2916596-2025-03265 | 2025-05-09 04:00 | Death | THORATEC CORPORATION | 2025-05-29 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Disseminated Intravascular Coagulation (DIC); Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH PATIENT OUTCOME THAT THE PATIENT PASSED AWAY DUE TO SEVERE VASOPLEGIA AND DISSEMINATED INTRAVASCULAR COAGULATION (DIC). THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. PATIENT WAS SUPPORTED BY LEFT VENTRICULAR ASSIST DEVICE (LVAD) PLUS A TEMPORARY RIGHT VENTRICULAR ASSIST DEVICE (RVAD) AND EXPERIENCED SEVERE VASOPLEGIA, THE PATIENT WAS NOT RESPONSIVE TO MEDICAL TREATMENT, LEADING TO UNCONTROLLED DIC AND THEN DEATH. THE PATIENTS RVAD WAS IMPLANTED DUE LEFT VENTRICLE DYSFUNCTION. THE PATIENT WAS NOTED TO HAVE LOW LEFT VENTRICLE EJECTION FRACTION BEFORE LVAD IMPLANT. ACCORDING THE CLINICIANS, THE PATIENTS LVAD DID NOT CONTRIBUTE TO THE RIGHT HEART FAILURE. THE LVAD WAS NOTED TO HAVE FUNCTIONED AS EXPECTED. Manufacturer Narrative: D6A - DATE OF IMPLANT: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE PROVIDED INFORMATION INDICATED THE PATIENT EXPIRED DUE TO SEVERE VASOPLEGIA AND DISSEMINATED INTRAVASCULAR COAGULATION (DIC). THE PATIENT WAS SUPPORTED BY THE LVAS, AND PER PROVIDED INFORMATION, A TEMPORARY RIGHT VENTRICULAR ASSIST DEVICE (RVAD) WAS IMPLANTED IN THE ANATOMICAL LEFT VENTRICLE TO SUPPORT PULMONARY CIRCULATION DUE TO SEVERE DISFUNCTION. THE PATIENT EXPERIENCED SEVERE VASOPLEGIA AND WAS NOT RESPONSIVE TO MEDICAL TREATMENT, LEADING TO UNCONTROLLED DIC AND DEATH. THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH AND RIGHT HEART FAILURE, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 "PATIENT CARE AND MANAGEMENT" (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6 (UNDER CAUTION!) STATES: "RIGHT HEART FAILURE CAN OCCUR FOLLOWING IMPLANTATION OF THE PUMP. RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH ELEVATED PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP." NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A3: PATIENT GENDER WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22108469&pc=OWB | 3003768277-2025-004895 | 2025-05-01 04:00 | Death | PHILIPS MEDICAL SYSTEMS NEDERLAND B.V. | 2025-05-29 04:00 | OWB | AZURION | Difficult or Delayed Positioning | Insufficient Information; No Clinical Signs, Symptoms or Conditions | K200917 |  | 1 | Event Description: IT WAS REPORTED TO PHILIPS THAT INITIAL CARDIOPULMONARY RESUSCITATION (CPR) WAS DIFFICULT TO PERFORM DUE TO THE TABLE HEIGHT AND THAT, A STEP STOOL WAS USED. IT WAS NOTED THAT THE PATIENT PASSED AWAY ON THE TABLE. AN INVESTIGATION INTO THIS COMPLAINT HAS BEEN INITIATED BY PHILIPS. Manufacturer Narrative: PHILIPS HAS CONCLUDED ITS INVESTIGATION INTO THE REPORTED EVENT. BASED ON THE AVAILABLE DATA, THE SYSTEM OPERATED WITHIN ITS INTENDED DESIGN SPECIFICATIONS. THE CUSTOMER INDICATED THAT THE TABLE HEIGHT COULD NOT BE ADJUSTED AS LOW AS DESIRED. THIS SYSTEM'S TABLE INCLUDES A SWIVEL OPTION WITH A VERTICAL RANGE OF 87 CM TO 112 CM. THE SWIVEL FUNCTION ALSO SUPPORTS LONGITUDINAL MOVEMENT OF THE TABLE BASE TOWARD THE STAND, ENABLING FULL-BODY COVERAGE DURING EXAMINATIONS FOR FLOOR MOUNTED SYSTEMS. THE TABLE HEIGHT IS PART OF THE TABLE¿S DESIGN TO SUPPORT SWIVEL OPERATION AND DOES NOT REPRESENT A MALFUNCTION. SYSTEM LOG FILES DID NOT SHOW ERRORS OR ANOMALIES RELATED TO TABLE HEIGHT. FROM A CLINICAL PERSPECTIVE, THE CUSTOMER CONFIRMED CPR (CARDIOPULMONARY RESUSCITATION) WAS STARTED WITHOUT DELAY USING A STEP STOOL WHEN THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE. RETURN OF SPONTANEOUS CIRCULATION (ROSC) WAS ACHIEVED, SUGGESTING EFFECTIVE INITIAL CPR. THE PATIENT'S CONDITION AND MULTIPLE COMORBIDITIES¿INCLUDING CIRRHOSIS, ACUTE PULMONARY EMBOLISM, AND STROKE¿WERE LIKELY CONTRIBUTING FACTORS TO SUBSEQUENT DETERIORATION AND DEATH. IN SUMMARY, THE INVESTIGATION DID NOT IDENTIFY ANY DEVICE MALFUNCTION OR DEVIATION FROM SPECIFICATIONS. THE EVENT APPEARS ASSOCIATED WITH THE CLINICAL COMPLEXITY OF THE CASE. NO CORRECTIVE OR PREVENTIVE ACTIONS ARE RECOMMENDED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22108579&pc=DSQ | 2916596-2025-03244 | 2025-05-05 04:00 | Death | THORATEC CORPORATION | 2025-05-29 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY SECONDARY TO ACUTE RESPIRATORY FAILURE. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE PATIENT'S DEATH WAS RELATED TO HEALTH CONDITION. THE PUMP WAS WORKING AS INTENDED DURING TIME OF SUPPORT. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS RESPIRATORY FAILURE AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22108600&pc=DSQ | 2916596-2025-03299 | 2025-03-17 04:00 | Death | THORATEC CORPORATION | 2025-05-29 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Septic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO SEPTIC SHOCK. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE RELATED. THERE WAS NO ALARM. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22108787&pc=LNH | 1000513161-2025-00019 | 2025-05-27 04:00 | Death | FUJIFILM CORPORATION | 2025-05-29 04:00 | LNH | ECHELON OVAL MRI SYSTEM | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Insufficient Information; No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: ON MAY 27TH, 2025, FUJIFILM HEALTHCARE AMERICAS CORPORATION WAS INFORMED OF AN EVENT INVOLVING SYS/MR/ECHELON OVAL 32CH. PATIENT HAS DIED WHILE BEING SCANNED FOR ABDOMEN ON ECHELON OVAL SYSTEM. THE DETAILED INVESTIGATION IS CURRENTLY ONGOING. DUE TO A NOTIFICATION OF A PATIENT DEATH, FUJIFILM HEALTHCARE AMERICAS CORPORATION IS REPORTING THIS EVENT.Event Description: REF COMP #(B)(4). UPDATE: FUJIFILM HEALTHCARE AMERICAS CORPORATION CONFIRMED THAT THE PATIENT PASSED AWAY AFTER BEING PLACED ON THE TABLE. THIS INCIDENT WAS NOT A RESULT OF THE ECHELON OVAL MRI SYSTEM MALFUNCTION, SINCE THE PROCEDURE HAD NOT YET STARTED PRIOR TO THE PATIENT'S PASSING. ON 07/08/2025, THE FDA REQUESTED ADDITIONAL INFORMATION REGARDING MANUFACTURER MDR 3018423337-2025-00004, AND FUJIFILM HEALTHCARE AMERICAS CORPORATION RESPONDED ON 07/28/2025 BY PROVIDING AN INCIDENT REPORT AND CLARIFYING THAT THE DEVICE DID NOT CAUSE THE PATIENT'S DEATH. INITIAL IMPORTER MDR 1000513161-2025-00019 AND FDA CORRESPONDENCE ARE ATTACHED IN THIS SUPPLEMENTAL REPORT (S1). Manufacturer Narrative: REF COMP: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22109139&pc=MNS | 2518422-2025-045400 | 2022-03-16 04:00 | Death | RESPIRONICS, INC. | 2025-05-29 04:00 | MNS | BIPAP S/T | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT DIED IN THE HOSPITAL. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22109190&pc=NVN | 2017865-2025-65006 | 2025-05-13 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-29 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED THE DAY AFTER PACEMAKER IMPLANT PROCEDURE. PULMONARY EMBOLISM WAS SUSPECTED BY THE PHYSICIAN AS A POSSIBLE CAUSE, BUT UNABLE TO BE CONFIRMED. IT WAS UNKNOWN WHETHER THE PACEMAKER SYSTEM WAS RELATED TO THE PATIENT DEATH. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. THE STERILIZATION RECORDS WERE REVIEWED AND NO EVIDENCE OF ABNORMAL STERILIZATION CYCLE WAS FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22109198&pc=NVN | 2017865-2025-65004 | 2025-05-13 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-29 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED THE DAY AFTER PACEMAKER IMPLANT PROCEDURE. PULMONARY EMBOLISM WAS SUSPECTED BY THE PHYSICIAN AS A POSSIBLE CAUSE, BUT UNABLE TO BE CONFIRMED. IT WAS UNKNOWN WHETHER THE PACEMAKER SYSTEM WAS RELATED TO THE PATIENT DEATH. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. THE STERILIZATION RECORDS WERE REVIEWED AND NO EVIDENCE OF ABNORMAL STERILIZATION CYCLE WAS FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22109216&pc=LWP | 2017865-2025-65002 | 2025-05-13 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-29 04:00 | LWP | ENDURITY MRI | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED THE DAY AFTER PACEMAKER IMPLANT PROCEDURE. PULMONARY EMBOLISM WAS SUSPECTED BY THE PHYSICIAN AS A POSSIBLE CAUSE, BUT UNABLE TO BE CONFIRMED. IT WAS UNKNOWN WHETHER THE PACEMAKER SYSTEM WAS RELATED TO THE PATIENT DEATH. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. THE STERILIZATION RECORDS WERE REVIEWED AND NO EVIDENCE OF ABNORMAL STERILIZATION CYCLE WAS FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22109646&pc=DSQ | 2916596-2025-03392 | 2023-12-01 05:00 | Death | THORATEC CORPORATION | 2025-05-29 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿PUMP EXCHANGE FOR MAJOR DEVICE RELATED INFECTION IN PATIENTS RECEIVING A DURABLE LEFT VENTRICULAR ASSIST DEVICE¿ IDENTIFYING THAT HEARTMATE 3 (HM3) MAY BE RELATED TO DEVICE INFECTION AND DEATH. THIS RETROSPECTIVE STUDY EVALUATED 152 PATIENTS IMPLANTED BETWEEN JAN2007 AND DEC2023. THE STUDY DISCUSSED HOW THE DURABLE LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY HAS BECOME MORE COMMON FOR PATIENTS WITH ADVANCED HEART FAILURE DUE TO TECHNOLOGICAL ADVANCEMENTS AND IMPROVED OUTCOMES. DESPITE IMPROVEMENTS IN SURVIVAL AND QUALITY OF LIFE, DEVICE-RELATED INFECTIONS (DRIS) REMAIN A SIGNIFICANT ISSUE, AFFECTING AROUND 41% OF PATIENTS WITHIN A YEAR OF IMPLANTATION. THE STUDY WAS CONDUCTED AT A SINGLE CENTER USING DATA FROM AN INSTITUTIONAL DURABLE LVAD REGISTRY, INVOLVING 152 PATIENTS DIAGNOSED WITH DRI AFTER DURABLE LVAD IMPLANTATION BETWEEN 2007 AND 2022. OF THESE, 32% UNDERWENT PUMP EXCHANGE (PE) AND 14% UNDERWENT TRANSPLANT. THE STUDY DEFINED DRI ACCORDING TO SPECIFIC CRITERIA AND COLLECTED BASELINE CHARACTERISTICS FOR THE PE GROUP. MAJOR OUTCOMES INCLUDED PATIENT SURVIVAL, FREEDOM FROM SUBSEQUENT DRIS, AND INCIDENCE OF NEW OR RECURRENT DRI WITHIN A YEAR POST-PE. SUCCESS WAS DEFINED AS FREEDOM FROM DRI INFECTION AT ONE YEAR POST-EXCHANGE. OUTCOMES WERE ANALYZED USING THE KAPLAN-MEIER METHOD. THE STUDY SHOWED THE OPERATIVE MORTALITY RATE WAS 4.1%. THE OVERALL PATIENT SURVIVAL RATE DURING DURABLE LVAD SUPPORT WAS 71.4% AT TWO YEARS FOLLOWING PE. FREEDOM FROM DRI AT ONE YEAR POST-PE WAS 74.7%. THE MEDIAN TIME TO SUBSEQUENT DRI WAS 5.8 MONTHS, WITH 33% EXPERIENCING DRI AFTER 12 MONTHS. SUCCESS IN AVOIDING SUBSEQUENT DRI WAS HIGHER (96%) WHEN PE WAS PERFORMED WITHIN FOUR MONTHS OF DRI DIAGNOSIS, COMPARED TO 56% WHEN PE WAS DELAYED BEYOND FOUR MONTHS. MAJOR DRIVELINE AND PUMP POCKET INFECTIONS REPRESENTED 31/49 (63%) AND 18/49 (37%) OF DRIS, RESPECTIVELY, WITH STAPHYLOCOCCUS AUREUS BEING THE MOST COMMON ORGANISM (METHICILLIN-SENSITIVE, N = 17; METHICILLIN-RESISTANT, N = 5). OF THE GRAM-NEGATIVE INFECTIONS, SERRATIA MARCESCENS (N = 7) WAS THE MOST COMMON ORGANISM FOLLOWED BY PSEUDOMONAS (N = 5) SPECIES AND PROTEUS MIRABILIS (N = 2). PE IS A THERAPEUTIC OPTION FOR DRIS FOLLOWING PRIMARY DURABLE LVAD IMPLANTATION WHEN INFECTIONS ARE UNRESPONSIVE TO MEDICAL THERAPY. THE RISK OF RECURRENCE POST-EXCHANGE IS MINIMIZED BY SHORTENING THE TIME INTERVAL BETWEEN THE INITIAL DIAGNOSIS OF DRI AND PE. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3)(6: DATE OF EVENT IS APPROXIMATE AS THE DATA WERE COLLECTED BETWEEN JANUARY 2007 AND (B)(6) 2023. DATE OF DEATH WAS NOT PROVIDED AND IS ENTERED AS THE DATE OF EVENT. AUTHOR INFORMATION: DEPARTMENT OF BIOMEDICAL ENGINEERING, UNIVERSITY OF MICHIGAN RACKHAM GRADUATE SCHOOL, ANN ARBOR, MICHIGAN; AND THE DEPARTMENT OF CARDIAC SURGERY, MICHIGAN MEDICINE, ANN ARBOR, MICHIGAN. PERSAUD, S. R., PIENTA, M., LIKOSKY, D. S., & PAGANI, F. D. (2025). PUMP EXCHANGE FOR MAJOR DEVICE RELATED INFECTION IN PATIENTS RECEIVING A DURABLE LEFT VENTRICULAR ASSIST DEVICE. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION, 44(4), 686¿689. HTTPS://DOI.ORG/10.1016/J.HEALUN.2024.11.032. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) DEVICES AND THE REPORTED OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 DEVICE SERIAL NUMBERS AND OTHER SPECIFIC CASE/PATIENT INFORMATION ARE NOT AVAILABLE. THE HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22110401&pc=BZD | 2518422-2025-106188 | 2025-05-13 04:00 | Death | RESPIRONICS, INC. | 2025-05-29 04:00 | BZD | DS2ADV AUTO CPAP W/HUMID+HT CELL/BT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED THAT A DS2ADV AUTO CPAP W/HUMID+HT CELL/BT DEVICE; PATIENT HAS PASSED AWAY. THE PATIENT'S DAUGHTER CALLED IN REGARDING A NOTIFICATION INSTRUCTING THEM TO RETURN AN AFFECTED DEVICE. SHE STATED THAT THE PATIENT HAD PASSED AWAY AND EXPRESSED HER DESIRE TO RETURN THE DEVICE. THE DAUGHTER REQUESTED THAT A SHIPPING LABEL BE SENT TO HER EMAIL ADDRESS, WHICH WAS DOCUMENTED. WHILE NO ASSERTIONS HAVE BEEN MADE ABOUT THE DEVICE CAUSING THE PATIENT'S DEATH, FURTHER DETAILS ARE NECESSARY. A REPORT WILL BE SUBMITTED TO ALL RELEVANT AGENCIES, AND IF MORE INFORMATION BECOMES AVAILABLE, A FOLLOW-UP REPORT WILL BE FILED. THE THIRD-PARTY SERVICE CENTER VISUALLY INSPECTED AND EVALUATED THE DEVICE DURING THE DEVICE
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22110403&pc=BZD | 2518422-2025-106189 | 2025-05-19 04:00 | Death | RESPIRONICS, INC. | 2025-05-29 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K113068 |  | 1 | Event Description: THE MANUFACTURER HAS BEEN INFORMED THAT A REMSTAR AUTOA-FLEX W/HUM, SYSONE60SRS, DOM DEVICE; PATIENT HAS PASSED AWAY (DEATH). THE CALLER STATED THAT HIS BROTHER HAD PASSED AWAY, AND HE IS NOW THE CARETAKER OF HIS BROTHER'S ESTATE. HE INDICATED THAT HE HAD RECEIVED MONETARY COMPENSATION FOR A RECALLED DEVICE BUT WAS UNABLE TO CASH THE CHECK DUE TO LEGALITIES SURROUNDING HIS BROTHER'S ESTATE PAPERWORK. THE CALLER REPORTED THAT THE CHECK HAS SINCE EXPIRED, AND HE IS NOW UNABLE TO CASH IT. HE WOULD LIKE TO KNOW IF ANOTHER CHECK CAN BE SENT. WHILE NO ASSERTIONS HAVE BEEN MADE ABOUT THE DEVICE CAUSING THE PATIENT'S DEATH, FURTHER DETAILS ARE NECESSARY. A REPORT WILL BE SUBMITTED TO ALL RELEVANT AGENCIES, AND IF MORE INFORMATION BECOMES AVAILABLE, A FOLLOW-UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22110787&pc=PNJ | 2017865-2025-65033 | 2025-05-09 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-29 04:00 | PNJ | LEADLESS PACEMAKER | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR A SCHEDULED PROCEDURE FOR A RIGHT VENTRICLE LEADLESS PACEMAKER (VLP). DURING THE PROCEDURE, HIGH THRESHOLDS WERE OBSERVED ON THE VLP. THE DEVICE WAS FIXATED IN FOUR DIFFERENT LOCATIONS AND THREE OUT OF THE FOUR IMPLANT LOCATIONS HAD HIGH CAPTURE THRESHOLDS. MAPPING WAS PERFORMED AND THE PATIENTS BLOOD PRESSURE DROPPED. SURGICAL STAFF WAS CONSULTED; HOWEVER, THE PATIENT EXPIRED DUE TO COMPLICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22110793&pc=PNJ | 2017865-2025-65032 | 2025-05-09 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-29 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Capturing Problem | Low Blood Pressure/ Hypotension | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR A SCHEDULED PROCEDURE FOR A RIGHT VENTRICLE LEADLESS PACEMAKER (VLP). DURING THE PROCEDURE, HIGH THRESHOLDS WERE OBSERVED ON THE VLP. THE DEVICE WAS FIXATED IN FOUR DIFFERENT LOCATIONS AND THREE OUT OF THE FOUR IMPLANT LOCATIONS HAD HIGH CAPTURE THRESHOLDS. MAPPING WAS PERFORMED AND THE PATIENTS BLOOD PRESSURE DROPPED. SURGICAL STAFF WAS CONSULTED HOWEVER, THE PATIENT EXPIRED DUE TO COMPLICATIONS.Event Description: NEW INFORMATION NOTED THAT OPEN CHEST SURGERY WAS PERFORMED BUT WAS UNSUCCESSFUL. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22110873&pc=NPT | 2015691-2025-04266 | 2025-05-06 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-29 04:00 | NPT | COMMANDER DELIVERY SYSTEM | Difficult to Remove; Device Slipped; Failure to Advance; Difficult to Advance | Hemorrhage/Bleeding; Foreign Body In Patient; No Clinical Signs, Symptoms or Conditions | P140031 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN THE UNITED KINGDOM. AS REPORTED, THIS WAS A CASE OF AN IMPLANT OF A 29MM SAPIEN 3 ULTRA RESILIA IN AORTIC POSITION BY RIGHT TRANSFEMORAL APPROACH. SIGNIFICANT CALCIFICATION AND TORTUOSITY IN THE FEMORAL AND ABDOMINAL AORTA WERE UNDERESTIMATED. DURING ADVANCEMENT OF THE DELIVERY SYSTEM, IT BECAME STUCK IN THE CALCIUM IN THE DESCENDING AORTA ONCE IT EXITED THE ESHEATH. AN ATTEMPT WAS MADE TO WITHDRAW THE SYSTEM BUT, EVERYTHING MOVED BACK EXCEPT THE VALVE REMAINED STUCK IN THE CALCIUM OF THE DESCENDING AORTA. AFTER SEVERAL MANEUVERS, THE VALVE WAS SUCCESSFULLY RETRIEVED FROM THAT POSITION. HOWEVER, SINCE IT COULD NOT BE REINTRODUCED INTO THE ESHEATH, THE DECISION WAS MADE TO DEPLOY THE VALVE IN THE ILIAC ARTERY, WHICH ALLOWED THE REMOVAL OF THE DELIVERY SYSTEM THROUGH THE ESHEATH. COVERED STENTS WERE PLACED TO ANCHOR THE VALVE PLACED IN THE NON-TARGET LOCATION, AND ADEQUATE BLOOD FLOW WAS OBSERVED. SOME HOURS AFTER THE PROCEDURE, THE PATIENT BECAME UNWELL DUE TO BLEEDING, LIKELY CAUSED BY THE MANEUVERS IN THE DESCENDING AORTA FOR DISLODGING THE VALVE. UNFORTUNATELY, THE PATIENT PASSED AWAY. AS PER THE REPORT, THE SYSTEM WAS PROPERLY LOCKED AND THE STEP OF ALIGNING THE VALVE WAS NOT REACHED. Manufacturer Narrative: THE INVESTIGATION IS IN PROGRESS. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION. Manufacturer Narrative: THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE IFU, CURRENT RISK MITIGATION'S INCLUDE DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THE PROVIDED IMAGERY WAS EVALUATED AND REVEALED THE FOLLOWING: THE COMMANDER WITH CRIMPED THV INSIDE PATIENTS ACCESS VESSEL. THE CRIMPED VALVE WAS DISPLACED TOWARDS COMMANDER TIP. FOR THE COMPLAINT OF UNABLE TO TRACK SYSTEM THROUGH ANATOMY, THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, DIFFICULTY OR INABILITY IN TRACKING THE DELIVERY SYSTEM TO THE LANDING ZONE, INCLUDING TRAVERSING THE AORTIC ARCH AND CROSSING THE NATIVE ANNULUS/BIOPROSTHESIS HAS NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCE'S. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PATIENT FACTORS, POOR GUIDEWIRE MANAGEMENT, DAMAGED GUIDEWIRE, AND/OR EXCESSIVE MANIPULATION OF THE FLEX WHEEL. THE COMPLAINT UNABLE TO TRACK SYSTEM THROUGH ANATOMY WAS CONFIRMED BASED ON EVALUATION OF PROVIDED IMAGERY. THE AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION, TORTUOSITY) LIKELY CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT ''SIGNIFICANT CALCIFICATION AND TORTUOSITY IN THE FEMORAL AND ABDOMINAL AORTA WERE UNDERESTIMATED'' AND ''DURING ADVANCEMENT OF THE DELIVERY SYSTEM IT BECAME STUCK IN THE CALCIUM''. PATIENT FACTORS SUCH AS CALCIFICATION AND TORTUOSITY MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERES WITH PASSAGE OF THE DELIVERY SYSTEM AND CAUSES DIFFICULTY DURING TRACKING/CROSSING. FOR THE COMPLAINTS OF WITHDRAW DIFFICULTY AND VALVE MOVES ON BALLOON DURING WITHDRAWAL, THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING DIFFICULTY/INABILITY WITHDRAWING CATHETER WITH CRIMPED VALVE THROUGH PATIENT VASCULATURE OR THROUGH THE SHEATH HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCE'S. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PATIENT FACTORS, NON-COAXIAL WITHDRAWAL, ALTERED CRIMP PROFILE, DAMAGED SHEATH, NOT CENTERING THE THV ON THE FLEX TIP, AND/OR WITHDRAWING THE THV THROUGH THE SHEATH HUB. THE COMPLAINTS WERE CONFIRMED BASED ON THE EVALUATION OF PROVIDED IMAGERY. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION, TORTUOSITY) LIKELY CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT THE PATIENT HAD SIGNIFICANT CALCIFIED AND TORTUOUS ACCESS VESSELS AND ABDOMINAL AORTA. PATIENT FACTORS SUCH AS CALCIFICATION AND TORTUOSITY MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERES WITH PASSAGE OF THE DELIVERY SYSTEM AND CAUSES DIFFICULTY DURING WITHDRAWAL. PROPER WITHDRAWAL TECHNIQUE IS IMPORTANT, AS CALCIFICATION, TORTUOSITY AND NON-COAXIAL WITHDRAWAL OF THE DELIVERY SYSTEM WITH CRIMPED THV MAY CAUSE THE SYSTEM OR THV TO INTERACT WITH VASCULATURE OR THE SHEATH, RESULTING IN DIFFICULTY WITHDRAWING THE SYSTEM AND VALVE MOVEMENT ON THE BALLOON AS SEEN DURING IMAGERY EVALUATION. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22094203&pc=MIH | 2017233-2025-06243 | 2025-05-09 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-05-28 04:00 | MIH | GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Vascular Dissection | P210032 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE ON MAY 15, 2025 FROM THE IMEDIDATA DATABASE: ON (B)(6) 2025, A PATIENT UNDERWENT AN EMERGENT ENDOVASCULAR TREATMENT FOR A THORACOABDOMINAL TYPE B DISSECTION (DISTAL TO LEFT SUBCLAVIAN ARTERY). THE PATIENT WAS TREATED WITH A GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS. THE GORE® DEVICE WAS SUCCESSFULLY NAVIGATED TO THE INTENDED LOCATION AND SUCCESSFULLY DEPLOYED, WITH THE PROXIMAL LANDING ZONE IN ZONE 2 AND DISTAL LANDING ZONE IN ZONE 4. THE SITE WAS ACCESSED PERCUTANEOUSLY. ON (B)(6) 2025, IT WAS NOTED THE PATIENT HAD A RETROGRADE TYPE A DISSECTION. THIS NEW AORTIC DISSECTION WAS ADJACENT TO THE PROXIMAL END OF THE DEVICE. TREATMENT WAS NOTED TO BE SUBXIPHOID GENERATION OF PERICARDIUM. THE PATIENT PASSED AWAY ON THE SAME DAY. Manufacturer Narrative: THE PATIENT REFERENCED IN THIS EVENT FILE IS ENROLLED IN THE TGR 23-02TB TOGETHER AORTIC REGISTRY AND INFORMATION REGARDING THIS EVENT WAS GATHERED FROM THE IMEDIDATA RAVE CLINICAL STUDY DATABASE (CSD). AS PART OF OUR ROUTINE QUALITY PROCEDURES, EACH BATCH OF DEVICES UNDERGOES COMPREHENSIVE QUALITY CONTROL TESTING AND INSPECTIONS PRIOR TO RELEASE FOR DISTRIBUTION. GORE REVIEWED THE MANUFACTURING RECORDS ASSOCIATED WITH THE REPORTED LOT NUMBER AND VERIFIED THAT ALL RELEASE CRITERIA HAD BEEN MET. THE DEVICE REMAINS IMPLANTED AND THEREFORE NO PRODUCT EVALUATION HAS BEEN PERFORMED. THE AVAILABLE INFORMATION REPORTED IN THE COMPLAINT DOES NOT REASONABLY SUGGEST A POTENTIAL MALFUNCTION OR PRODUCT PACKAGING AND/OR LABELING ISSUE HAS OCCURRED. BASED ON THE INCIDENT DESCRIPTION AND THE SUBSEQUENT INVESTIGATION, NO FURTHER INFORMATION WAS PROVIDED TO GORE, WE ARE UNABLE TO DETERMINE THE CAUSE OF THIS INCIDENT AND ASSIGN A ROOT CAUSE. PER THE GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), POSSIBLE ADVERSE EVENTS AND COMPLICATIONS THAT MAY OCCUR WITH THE USE OF GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS INCLUDE, BUT ARE NOT LIMITED TO: DEATH, DISSECTION, PERFORATION, OR RUPTURE OF THE AORTIC VESSEL AND/OR SURROUNDING VASCULATURE. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22094382&pc=MIH | 2017233-2025-06244 | 2025-05-22 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-05-28 04:00 | MIH | GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM | Break; Migration; Activation Problem | Stroke/CVA; Obstruction/Occlusion | P040043 |  | 1 | Event Description: IT WAS REPORTED TO GORE THAT ON (B)(6) 2025, A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (CTAG) WAS USED TO TREAT A PATIENT WITH A THORACIC AORTIC ANEURYSM. THE CASE PLANNING INCLUDED A BRANCHED CUSTOM-MADE COOK DEVICE IN THE AORTIC ARCH (BRACHIOCEPHALIC TRUNK, THE LEFT COMMON CAROTID ARTERY, AND THE LEFT SUBCLAVIAN ARTERY), AS WELL AS A SECOND COOK STENT GRAFT FOR THE EXTENSION IN THE THORACIC AORTA. THIS SECOND COOK STENT GRAFT WAS NOT ABLE TO REACH THE BRANCHED MODULE AND COULD NOT CROSS THE AORTIC ISTHMUS. THEREFORE, THE PHYSICIAN DECIDED TO USE THE CTAG DEVICE INSTEAD. IN A 22FR INTRODUCER SHEATH, THE CTAG DEVICE WAS ABLE TO REACH THE FIRST STENT FROM COOK BRANCHED DEVICE. HOWEVER, AN UNKNOWN BREAKAGE OF THE DELIVERY SYSTEM OCCURRED DURING THE DEPLOYMENT. THE PHYSICIAN WAS ABLE TO DEPLOY THE STENT FULLY; IT WAS NOT CONNECTED TO THE COOK'S MODULE AND THE CTAG GRAFT SLID DOWN IN THE ABDOMINAL AORTA AND COVERED THE DIGESTIVE ARTERIES (CELIAC ARTERY AND SUPERIOR MESENTERIC ARTERY) AND PART OF THE RENAL ARTERIES. IT WAS DECIDED TO OPEN THE PATIENT WITH A LAPAROTOMY ACCESS, REMOVE THE CTAG DEVICE, AND ANOTHER CTAG DEVICE WAS USED TO TREAT THE PATIENT.Event Description: IT WAS REPORTED TO GORE THAT ON (B)(6) 2025, A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (CTAG) WAS USED TO TREAT A PATIENT WITH A THORACIC AORTIC ANEURYSM. THE CASE PLANNING INCLUDED A BRANCHED CUSTOM-MADE COOK DEVICE IN THE AORTIC ARCH (BRACHIOCEPHALIC TRUNK, THE LEFT COMMON CAROTID ARTERY, AND THE LEFT SUBCLAVIAN ARTERY), AS WELL AS A SECOND COOK STENT GRAFT FOR THE EXTENSION IN THE THORACIC AORTA. THIS SECOND COOK STENT GRAFT WAS NOT ABLE TO REACH THE BRANCHED MODULE AND COULD NOT CROSS THE AORTIC ISTHMUS. HENCE, THIS COOK STENT GRAFT WAS NOT IMPLANTED, AND INSTEAD, THE PHYSICIAN DECIDED TO USE THE CTAG DEVICE. IN A 22FR INTRODUCER SHEATH, THE CTAG DEVICE WAS ABLE TO REACH THE FIRST STENT FROM COOK BRANCHED DEVICE. HOWEVER, DURING THE FIRST STEP OF DEPLOYMENT, THE DEPLOYMENT LINE BROKE, MAKING THE DEPLOYMENT OF THE ENDOPROSTHESIS DIFFICULT. THE ENDOPROSTHESIS WAS PARTIALLY DEPLOYED (ABOUT 75% OF ITS FULL DIAMETER) BEFORE THE BREAKAGE. THE PHYSICIAN ATTEMPTED TO PULL THE DEPLOYMENT HANDLE IN ORDER TO DEPLOY THE DEVICE FURTHER. HOWEVER, DUE TO THE BLOOD FLOW, THE CTAG GRAFT SLID DOWN IN THE ABDOMINAL AORTA AND COVERED THE DIGESTIVE ARTERIES (CELIAC ARTERY AND SUPERIOR MESENTERIC ARTERY) AND PART OF THE RENAL ARTERIES. IT WAS DECIDED TO OPEN THE PATIENT WITH A LAPAROTOMY ACCESS, REMOVE THE CTAG DEVICE, AND ANOTHER CTAG DEVICE WAS USED TO TREAT THE PATIENT. ONE WEEK LATER, THE PATIENT PASSED AWAY DUE TO COMPLICATIONS IN THE AORTIC ARCH/STROKE. Manufacturer Narrative: A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. PRODUCT HISTORY REVIEW IS ONGOING. ENGINEERING EVALUATION OF THE DEVICE IS ANTICIPATED (DEVICE CURRENTLY IN TRANSIT). THE INITIAL REPORTER HAS BEEN CONTACTED IN ORDER TO RECEIVE MORE INFORMATION ON THE EVENT AND PATIENT DETAILS. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY HAVE NOT BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: UPDATED B2, B5, D10, G3, H1-2. UPDATED H6: ADDED HEALTH EFFECT - CLINICAL CODE E0133. ADDED HEALTH EFFECT - IMPACT CODE F02. ADDED COMPONENT CODE G04122. D10: CONCOMITANT MEDICAL PRODUCTS: (RELEVANT ASSOCIATED DEVICES USED WITH THE DEVICE BEING REPORTED ON): A BRANCHED CUSTOM-MADE COOK DEVICE, AND A SECOND COOK STENT GRAFT (NOT IMPLANTED). AS PART OF OUR ROUTINE QUALITY PROCEDURES, EACH BATCH OF DEVICES UNDERGOES COMPREHENSIVE QUALITY CONTROL TESTING AND INSPECTIONS PRIOR TO RELEASE FOR DISTRIBUTION. GORE REVIEWED THE MANUFACTURING RECORDS ASSOCIATED WITH THE REPORTED LOT NUMBER AND VERIFIED THAT ALL RELEASE CRITERIA HAD BEEN MET.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22094983&pc=DSQ | 2916596-2025-03445 | 2025-05-17 04:00 | Death | THORATEC CORPORATION | 2025-05-28 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH PATIENT OUTCOME THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22095097&pc=BZD | 2518422-2025-045351 | 2021-11-02 04:00 | Death | RESPIRONICS, INC. | 2025-05-28 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRM FOAM PARTICLES IN THE DEVICE EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22095121&pc=DSQ | 2916596-2025-03148 | 2025-05-03 04:00 | Death | THORATEC CORPORATION | 2025-05-28 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Mechanical Problem; Decreased Pump Speed; Adverse Event Without Identified Device or Use Problem | Fatigue; Muscle Weakness; No Clinical Signs, Symptoms or Conditions | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS CURRENTLY ADMITTED AND BEGAN TO HAVE LOW SPEED ADVISORIES, AND REPLACE SYSTEM CONTROLLER ALARMS. THE PATIENT WAS NOTED TO NOT HAVE GAINED ANY WEIGHT. ADDITIONALLY IT WAS NOTED THAT THERE WERE 2 AREAS OF EXTERNAL SLEEVE BREAKDOWN THAT WAS TAPED WITH RESCUE TAPE, WHICH WOULD REQUIRE X-RAYS. LOG FILES WERE SUBMITTED FOR REVIEW AND CAPTURED INTERMITTENT SPEED DROPS ON (B)(6) 2025 AT 3:20 AND ON (B)(6) 2025 AT 3:44. THESE ALARMS OCCURRED WHILE THE PATIENT WAS CONNECTED TO THE POWER MODULE. IT WAS NOTED THAT FOR THE LAST 24 HOURS OF THE LOG FILE, WHILE THE PATIENT WAS ON BATTERY POWER, THE PATIENT HAD NO PUMP SPEED ISSUES. THE LOW SPEED ADVISORIES MAY HAVE BEEN CAUSED BY SHORT TO SHIELD (STS). IMAGES OF THE DRIVELINE WERE UNREMARKABLE.Event Description: IT WAS LATER REPORTED THROUGH PATIENT OUTCOME THAT THE PATIENT PASSED AWAY DUE FAILURE TO THRIVE. THE PATIENT WAS NOTED TO HAVE HAD INCREASING WEAKNESS DUE TO THE PATIENTS PROLONGED HOSPITAL STAY AND ADVANCED AGE. THE PATIENTS FAMILY ELECTED FOR COMPASSIONATE WITHDRAWAL OF SUPPORT. THE DEATH WAS NOT CONSIDERED DEVICE RELATED, THE DEVICE WAS NOTED TO HAVE OPERATING AS EXPECTED, AND THE DEVICE WOULD NOT RETURN FOR EVALUATION. Manufacturer Narrative: SECTION H6: CLINICAL CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILE CONFIRMED THE REPORTED LOW SPEED EVENTS. ALTHOUGH A SPECIFIC CAUSE COULD NOT BE CONCLUSIVELY DETERMINED AS NO PRODUCT WAS RETURNED FOR EVALUATION, THE ABNORMAL PUMP OPERATION CAPTURED IN THE SUBMITTED LOG FILE APPEARED TO BE CONSISTENT WITH POTENTIAL DRIVELINE WIRE COMPROMISE. THE REPORTED COSMETIC DAMAGE TO THE PATIENT¿S DRIVELINE WAS UNABLE TO BE CONFIRMED VIA THE SUBMITTED X-RAY IMAGES. FURTHERMORE, A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SUBMITTED LOG FILE CONTAINED EVENTS FROM (B)(6) 2025 AT 03:20:05 ¿ (B)(6) 2025 AT 07:47:47, PER THE TIMESTAMPS. SEVERAL LOW SPEED EVENTS, AS WELL AS A DRIVELINE FAULT ASSOCIATED WITH A YELLOW WRENCH ADVISORY, WERE CAPTURED ON (B)(6) 2025. ADDITIONALLY OF NOTE, SEVERAL ELEVATIONS IN POWER AND FLOW WERE CAPTURED THROUGHOUT THE MAJORITY OF THE FILE. THESE EVENTS ALL OCCURRED WHILE THE PATIENT WAS OPERATING ON THE POWER MODULE. BASED ON PREVIOUS COMPLAINT HISTORY, THE ABNORMAL PUMP OPERATION APPEARED TO BE CONSISTENT WITH A POTENTIAL SHORT TO SHIELD DRIVELINE ISSUE. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) AND DRIVELINE SERIAL NUMBER (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE II LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. SECTION 6 OF THE IFU ENTITLED "PATIENT CARE AND MANAGEMENT" ADDRESSES HOW TO CARE FOR THE DRIVELINE; HOWEVER, ALL HEARTMATE II DRIVELINES HAVE THE POTENTIAL FOR WIRE/SHIELD BREAKDOWN TO OCCUR DEPENDENT ON DURATION OF USE AND PATIENT HANDLING. IN ADDITION, SECTION 6 (UNDER "PUMP PERFORMANCE MONITORING") OUTLINES INDICATIONS OF DRIVELINE DAMAGE AS WELL AS HOW TO RESPOND TO SUCH EVENTS. SECTION 7 ENTITLED "ALARMS AND TROUBLESHOOTING" OUTLINES ALL SYSTEM CONTROLLER ALARMS AND THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. THE HEARTMATE II LVAS PATIENT HANDBOOK CONTAINS A SECTION ON ¿CARING FOR THE DRIVELINE.¿ THE SECTION ENTITLED ¿ALARMS AND TROUBLESHOOTING¿ OUTLINES ALL SYSTEM CONTROLLER ALARMS AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. A SECTION ON HANDLING EMERGENCIES IS ALSO PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22095808&pc=OZD | 1220648-2025-28765 | 2025-05-05 04:00 | Death | ABIOMED, INC. | 2025-05-28 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA 5.5 SUFFERED A PUNCTURE OF THE IMPELLA THROUGH THE LEFT VENTRICLE WALL. THE PERFORATION WAS REPAIRED WITH A STITCH. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING: SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22095823&pc=OBJ | 2124215-2025-29804 | 2025-05-07 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-28 04:00 | OBJ | OPTICROSS HD | Detachment of Device or Device Component; Adverse Event Without Identified Device or Use Problem; Physical Resistance/Sticking | Foreign Body In Patient | K173284 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT DEATH OCCURRED. THE 90% STENOSED TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED LEFT CIRCUMFLEX (LCX) ARTERY. THE OPTICROSS HD IMAGING CATHETER WAS SELECTED FOR ULTRASOUND EXAMINATION OF THE TARGET LESION. AFTER COMPLETING THE ULTRASOUND IMAGING, THE PHYSICIAN ENCOUNTERED RESISTANCE WHILE ATTEMPTING TO WITHDRAW THE CATHETER. UPON APPLYING FORCE, THE CATHETER BROKE, LEAVING APPROXIMATELY 10 CM OF THE DISTAL PORTION INSIDE THE PATIENT. THIS COMPLICATION LED TO RAPID DETERIORATION OF THE PATIENT CONDITION, PREVENTING THE CONTINUATION OF THE PROCEDURE. DUE TO THE COMBINED CIRCUMSTANCES, THE PATIENT DID NOT RECOVER AND PASSED AWAY IN THE INTENSIVE CARE UNIT (ICU).Event Description: IT WAS REPORTED THAT PATIENT DEATH OCCURRED. THE 90% STENOSED TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED LEFT CIRCUMFLEX (LCX) ARTERY. THE OPTICROSS HD IMAGING CATHETER WAS SELECTED FOR ULTRASOUND EXAMINATION OF THE TARGET LESION. AFTER COMPLETING THE ULTRASOUND IMAGING, THE PHYSICIAN ENCOUNTERED RESISTANCE WHILE ATTEMPTING TO WITHDRAW THE CATHETER. UPON APPLYING FORCE, THE CATHETER BROKE, LEAVING APPROXIMATELY 10 CM OF THE DISTAL PORTION INSIDE THE PATIENT. THIS COMPLICATION LED TO RAPID DETERIORATION OF THE PATIENT CONDITION, PREVENTING THE CONTINUATION OF THE PROCEDURE. DUE TO THE COMBINED CIRCUMSTANCES, THE PATIENT DID NOT RECOVER AND PASSED AWAY IN THE INTENSIVE CARE UNIT (ICU). Manufacturer Narrative: THE DEVICE WAS RETURNED FOR ANALYSIS. VISUAL AND MICROSCOPIC INSPECTION REVEALED THE TELESCOPE, SHEATH, AND IMAGING WINDOW WERE KINKED AND THE DISTAL SECTION OF THE IMAGING WINDOW WAS STRETCHED AND DETACHED. THE DETACHED PORTION, INCLUDING THE TIP, WAS NOT RETURNED FOR ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22095937&pc=MUZ | 1644487-2025-10179 | 2025-05-01 04:00 | Death | LIVANOVA USA, INC. | 2025-05-28 04:00 | MUZ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY IN THEIR SLEEP DUE TO STROKE. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE.Event Description: IT WAS LATER REPORTED BY THE CAREGIVER THAT OFFICIAL CAUSE OF DEATH NOTED ON THE DEATH CERTIFICATE WAS STROKE. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE.Event Description: IT WAS LATER REPORTED BY THE PHYSICIAN THAT THE PATIENT PASSED DUE TO A STROKE. THE PROVIDER STATED THAT THE CAUSE WAS NOT RELATED TO VNS. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22095965&pc=DTB | 2649622-2025-14430 | 2025-05-08 04:00 | Death | MPRI | 2025-05-28 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Failure to Capture; Over-Sensing; Pacing Problem; Under-Sensing | Insufficient Information; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC RESPIRATORY FAILURE. THE RIGHT ATRIAL (RA) LEAD EXHIBITED NO CAPTURE AND UNDERS ENSING TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF AT/AF EVENT(S) AT TIMES AND INAPPROPRIATE SENSING. IT WAS ALSO REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED NO CAPTURE AND HIGH THRESHOLDS. THE LEADS REMAIN IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT IS DECEASED. IT WAS FURTHER NOTED THAT THE RA AND RV LEAD WERE REPROGRAMMED. IT WAS MENTIONED THAT SUSPECTED METABOLIC ACIDOSIS AND HYPERKALEMIA WERE CONTRIBUTORS IN PART TO LACK OF CAPTURE. IT STATES REASONS FOR DEATH ARE UNRELATED TO PACEMAKER FUNCTION.Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC RESPIRATORY FAILURE. THE RIGHT ATRIAL (RA) LEAD EXHIBITED NO CAPTURE AND UNDERSENSING TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF AT/AF EVENT(S) AT TIMES AND INAPPROPRIATE PACING. IT WAS ALSO REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED NO CAPTURE AND HIGH THRESHOLDS. THE LEADS REMAIN IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. IT WAS FURTHER REPORTED THAT THE PATIENT IS DECEASED. IT WAS FURTHER NOTED THAT THE RA AND RV LEAD WERE REPROGRAMMED. IT WAS MENTIONED THAT SUSPECTED METABOLIC ACIDOSIS AND HYPERKALEMIA WERE CONTRIBUTORS IN PART TO LACK OF CAPTURE. IT STATES REASONS FOR DEATH ARE UNRELATED TO PACEMAKER FUNCTION. Manufacturer Narrative: CORRECTED B5 AND H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22095978&pc=NVN | 2649622-2025-14431 | 2025-05-08 04:00 | Death | MPRI | 2025-05-28 04:00 | NVN | SELECTSECURE MRI SURESCAN | Failure to Capture; High Capture Threshold | Insufficient Information; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC RESPIRATORY FAILURE. THE RIGHT ATRIAL (RA) LEAD EXHIBITED NO CAPTURE AND UNDERSENSING TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF AT/AF EVENT(S) AT TIMES AND INAPPROPRIATE SENSING. IT WAS ALSO REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED NO CAPTURE AND HIGH THRESHOLDS. THE LEADS REMAIN IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT IS DECEASED. IT WAS FURTHER NOTED THAT THE RA AND RV LEAD WERE REPROGRAMMED. IT WAS MENTIONED THAT SUSPECTED METABOLIC ACIDOSIS AND HYPERKALEMIA WERE CONTRIBUTORS IN PART TO LACK OF CAPTURE. IT STATES REASONS FOR DEATH ARE UNRELATED TO PACEMAKER FUNCTION.Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC RESPIRATORY FAILURE. THE RIGHT ATRIAL (RA) LEAD EXHIBITED NO CAPTURE AND UNDERSENSING TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF AT/AF EVENT(S) AT TIMES AND INAPPROPRIATE PACING. IT WAS ALSO REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED NO CAPTURE AND HIGH THRESHOLDS. THE LEADS REMAIN IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. IT WAS FURTHER REPORTED THAT THE PATIENT IS DECEASED. IT WAS FURTHER NOTED THAT THE RA AND RV LEAD WERE REPROGRAMMED. IT WAS MENTIONED THAT SUSPECTED METABOLIC ACIDOSIS AND HYPERKALEMIA WERE CONTRIBUTORS IN PART TO LACK OF CAPTURE. IT STATES REASONS FOR DEATH ARE UNRELATED TO PACEMAKER FUNCTION. Manufacturer Narrative: CORRECTED B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22096010&pc=OZD | 1220648-2025-28768 | 2025-05-04 04:00 | Death | ABIOMED, INC. | 2025-05-28 04:00 | OZD | IMPELLA CP | Positioning Failure; Restricted Flow rate; Deformation Due to Compressive Stress | Hematoma; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A .018 GUIDEWIRE KINKED DURING INSERTION FOR IMPELLA CP IMPLANT. THE WIRE WAS REPLACED WITH ANOTHER GUIDEWIRE FROM A SECOND SET AND THE SUBSEQUENT IMPLANT WAS SUCCESSFUL. OOZING WAS NOTED AT THE ACCESS SITE LEADING TO FEMOSTOP PLACEMENT TO MANAGE THE CONDITION. SUCTION ALARMS APPEARED THE FOLLOWING DAY AND THE PIGTAIL WAS FOUND TO BE UNDER THE PAPILLARY MUSCLE. TROUBLESHOOTING RESULTED IN SUBOPTIMAL POSITIONING WITH ON THE LOWER END FOR THE SET P-LEVEL. DUE TO THE PATIENT'S STABILITY, THE DECISION WAS MADE TO CONTINUE WITH THE CURRENT POSITIONING/FLOWS. ALTHOUGH THE OOZING HAD STOPPED, A RESULTING HEMATOMA HAD DOUBLED IN SIZE. FIVE UNITS OF BLOOD WERE GIVEN TO TREAT THE HEMATOMA, HOWEVER SIGNIFICANT BLEEDING THEN DEVELOPED AT THE SITE FOLLOWING ADDITIONAL REPOSITIONING OF THE DEVICE. SUCTION ALSO REMAINED INTERMITTENT, FOR WHICH VOLUME AND POSITIONING WERE MANAGED. PLANS WERE MADE TO ESCALATE TO IMPELLA 5.5 SUPPORT, BUT THE PATIENT CONTINUED TO DETERIORATE AND LATER PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING, POSITIONING ISSUES, DELIVERY ISSUES, PRODUCT DAMAGE (GUIDEWIRE DAMAGE - PERIPHERAL VESSEL, AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. LOW PUMP FLOW: DATA LOGS WERE REVIEWED AND LT LOGS SHOW SUCTION ALARMS OCCURRING INTERMITTENTLY DURING SUPPORT. MULTIPLE INSTANCES OF PLACEMENT SIGNAL (PS) INCREASES WITH DROP IN MOTOR CURRENT (MC) AND PUMP FLOW. NO POSITIONING ALARMS OR MOTOR CURRENT SPIKES TO INDICATE INGESTION WERE OBSERVED. CLINICAL DETAILS NOTED THAT SUCTION ALARMS OCCURRED MULTIPLE TIMES THROUGHOUT SUPPORT. CLINICAL STATES GOOD FLOWS INITIALLY, BUT THEN LATER ON DAY 1 OF SUPPORT THE PUMP WAS REPOSITIONED INTO LESS THAN IDEAL POSITION WITH FLOWS ON LOW END FOR THE P-LEVEL. ON DAY 2 OF SUPPORT PATIENT HAD SUCTION ALARMS, WAS GIVEN VOLUME, AND THE PUMP WAS REPOSITIONED WITH IMPROVEMENT IN FLOWS. SUCTION ALARMS STATED ON DAY 3 WITH PUMP STATED TO BE SHALLOW. THE ROOT CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY IMPROPER POSITION BASED ON CLINICAL DETAILS INDICATING LOW FLOW ISSUES OCCURRING CONCURRENTLY WITH POSITIONING ISSUES. POSITIONING ISSUES: THE ROOT CAUSE OF THE POSITIONING ISSUES WAS MOST LIKELY DUE TO PATIENT MOVEMENT AS THE PUMP WAS NOTED TO BE OUT OF POSITION AFTER PATIENT WAS TRANSFERRED TO UNIT AND AFTER PATIENT WAS TRANSFERRED HOSPITALS. ACCESS SITE BLEEDING: THE ROOT CAUSE OF THE ACCESS SITE BLEEDING COULD NOT BE DEFINITIVELY DETERMINED AS MULTIPLE FACTORS, INCLUDING PATIENT MOVEMENT AND TRANSFER AND GUIDEWIRE KINKING, OCCURRED PRIOR TO THE OOZING AND HEMATOMA. PRODUCT DAMAGE (GUIDEWIRE DAMAGE): CLINICAL DETAILS NOTED THAT FIRST 0.018 GUIDEWIRE ATTEMPTED TO BE INSERTED WAS KINKED AND GUIDEWIRE WAS REPLACED WITH ANOTHER 0.018 GUIDEWIRE. THE ROOT CAUSE OF THE GUIDEWIRE DAMAGE COULD NOT BE DEFINITIVELY DETERMINED AS NO PRODUCT WAS RETURNED FOR EVALUATION AND LACK OF CLINICAL DETAILS. DELIVERY ISSUES: CLINICAL DETAILS NOTED THAT GUIDEWIRE KINKED DURING INSERTION. GUIDEWIRE WAS REMOVED AND A NEW GUIDEWIRE WAS USED FOR PUMP INSERTION. THE ROOT CAUSE OF THE DELIVERY ISSUES WAS MOST LIKELY DUE TO A GUIDEWIRE KINK PER CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22096245&pc=PNJ | 9612164-2025-02622 | 2025-05-20 04:00 | Death | MEDTRONIC IRELAND | 2025-05-28 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Great Vessel Perforation | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING IMPLANT OF LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT DIED OF A VASCULAR COMPLICATION RELATED TO TRANS-ARTERIAL ACCESS SITE FROM THE LEFT SUBCLAVIAN ARTERY. THE LEADLESS IPG REMAINS IN THE PATIENT. THE PATIENTS ARTERY WAS TORN AND THE ARTERY WAS TRANSECTED AND RETRACTED INTO THE CHEST. THERE WAS CONSIDERABLE BLOOD LOSS FROM THE TORN ARTERY AND AFTER SOME TIME THE PATIENT ARRESTED AND DIED. Manufacturer Narrative: CORRECTION TO H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22096523&pc=OZD | 1220648-2025-28771 | 2025-05-05 04:00 | Death | ABIOMED, INC. | 2025-05-28 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A 63 YEAR OLD MALE PATIENT EXPERIENCED SOME OOZING AT THEIR ACCESS SITE DURING IMPELLA CP SUPPORT. HEPARIN LEVELS WERE REDUCED AND THE DRESSING WAS CHANGED TO MANAGE THE CONDITION. SUPPORT CONTINUED FOLLOWING THE INTERVENTION. THREE DAYS LATER, CARE WAS WITHDRAWN AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME.Event Description: ADDITIONALLY VENTRICULAR TACHYCARDIA (VT) WAS ADDED DUE TO CARDIOVERSIONS WERE REQUIRED FOR AN UNKNOWN INDICATION. Manufacturer Narrative: THE ROOT CAUSE OF ACCESS SITE BLEEDING IS DETERMINED TO BE ANTI-COAGULATION MANAGEMENT BECAUSE THE HEMOSTASIS WAS ACHIEVED BY REDUCING THE HEPARIN AND DRESSING CHANGE. Manufacturer Narrative: B5 ADDITIONAL INFORMATION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28771. H6 HEALTH EFFECT - CLINICAL CODE 2095 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28771. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22096664&pc=MNS | 2518422-2025-106101 | 2025-05-21 04:00 | Death | RESPIRONICS, INC. | 2025-05-28 04:00 | MNS | DREAMSTATION BIPAP S/T 30 W/HEATED HUMIDIFIER, GB | Adverse Event Without Identified Device or Use Problem | Unspecified Respiratory Problem | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAS SADLY PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22096710&pc=NGV | 2124215-2025-34314 | 2025-05-21 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-28 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Unspecified Infection; Cardiac Perforation; Multiple Organ Failure; Pericardial Effusion | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT PERICARDIAL EFFUSION OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED AND A 31MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS SELECTED TO BE USED. THE APPENDAGE WAS HEAVILY PECTINATED AND SHALLOW. TWO TRANSEPTAL PUNCTURES (TSP) WERE NEEDED. AFTER THE SECOND TSP, AFTER MANY REDEPLOYMENTS AND PRIOR TO ATTEMPTING ANOTHER WATCHMAN CLOSURE DEVICE, THE PATIENTS' BLOOD PRESSURE DROPPED, AND AN EFFUSION WAS NOTED ON TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). THERE WAS A PERICARDIAL EFFUSION AT THE DISTAL END OF THE LAA. THE PATIENT WAS SENT TO CARDIAC SURGERY WHERE A STERNOTOMY WAS PERFORMED. THE PATIENT IS CURRENTLY INTUBATED AT THE HOSPITAL BUT EXPECTED TO FULLY RECOVER.Event Description: IT WAS REPORTED THAT PERICARDIAL EFFUSION OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED AND A 31MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS SELECTED TO BE USED. THE APPENDAGE WAS HEAVILY PECTINATED AND SHALLOW. TWO TRANSEPTAL PUNCTURES (TSP) WERE NEEDED. AFTER THE SECOND TSP, AFTER MANY REDEPLOYMENTS AND PRIOR TO ATTEMPTING ANOTHER WATCHMAN CLOSURE DEVICE, THE PATIENTS' BLOOD PRESSURE DROPPED, AND AN EFFUSION WAS NOTED ON TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). THERE WAS A PERICARDIAL EFFUSION AT THE DISTAL END OF THE LAA. THE PATIENT WAS SENT TO CARDIAC SURGERY WHERE A STERNOTOMY WAS PERFORMED. THE PATIENT IS CURRENTLY INTUBATED AT THE HOSPITAL BUT EXPECTED TO FULLY RECOVER. IT WAS FURTHER REPORTED THAT PATIENT DEATH OCCURRED ON (B)(6) 2025. THE PATIENT DEVELOPED LACTIC ACIDOSIS TWO DAYS AFTER THE STERNOTOMY AND HAD MULTIPLE ORGAN FAILURE.Event Description: IT WAS REPORTED THAT PERICARDIAL EFFUSION OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED AND A 31MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS SELECTED TO BE USED. THE APPENDAGE WAS HEAVILY PECTINATED AND SHALLOW. TWO TRANSEPTAL PUNCTURES (TSP) WERE NEEDED. AFTER THE SECOND TSP, AFTER MANY REDEPLOYMENTS AND PRIOR TO ATTEMPTING ANOTHER WATCHMAN CLOSURE DEVICE, THE PATIENTS' BLOOD PRESSURE DROPPED, AND AN EFFUSION WAS NOTED ON TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). THERE WAS A PERICARDIAL EFFUSION AT THE DISTAL END OF THE LAA. THE PATIENT WAS SENT TO CARDIAC SURGERY WHERE A STERNOTOMY WAS PERFORMED. THE PATIENT IS CURRENTLY INTUBATED AT THE HOSPITAL BUT EXPECTED TO FULLY RECOVER. IT WAS FURTHER REPORTED THAT PATIENT DEATH OCCURRED ON (B)(6) 2025. THE PATIENT DEVELOPED LACTIC ACIDOSIS TWO DAYS AFTER THE STERNOTOMY AND HAD MULTIPLE ORGAN FAILURE. IT WAS FURTHER REPORTED THAT THE PATIENT DEVELOPED A SEVERE INFECTION, AND THAT A PERFORATION CAUSED THE OPEN-HEART SURGERY TO BE NECESSARY. THIS EVENT WAS ALSO REPORTED UNDER REPORT NUMBER 2124215-2025-38817. Manufacturer Narrative: H2 CORRECTION: THIS EVENT WAS ALSO REPORTED UNDER REPORT NUMBER 2124215-2025-38817. B5 DESCRIBE EVENT OR PROBLEM: UPDATED WITH ADDITIONAL INFORMATION RECEIVED. H6: PATIENT CODE ADDED. Manufacturer Narrative: B5 DESCRIBE EVENT OR PROBLEM: UPDATED WITH ADDITIONAL INFORMATION RECEIVED. H6: PATIENT CODE ADDED. B2: DATE OF DEATH ADDED. H6: IMPACT CODE ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22096899&pc=DSQ | 2916596-2025-03566 | 2025-05-05 04:00 | Death | THORATEC CORPORATION | 2025-05-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE PUMP WAS EXPLANTED. THE SYSTEM CONTROLLER WAS NOT COMMUNICATING WITH THE SYSTEM MONITOR SO THE DEVICE WAS UNABLE TO BE INTERROGATED.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE CAUSE OF DEATH WAS ACUTE INTRACEREBRAL HEMORRHAGE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED WITH THE EXCEPTION OF NOT BEING ABLE TO TRANSMIT ANY DATA FROM THE SYSTEM CONTROLLER TO THE SYSTEM MONITOR. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22096943&pc=DSQ | 2916596-2025-03028 | 2023-01-01 05:00 | Death | THORATEC CORPORATION | 2025-05-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿FACTORS ASSOCIATED WITH READMISSIONS FOLLOWING LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION¿ THAT HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH HEART FAILURE EXASPERATION, MAJOR INFECTION, BLEEDING, UNEXPECTED SURGICAL/MEDICAL INTERVENTION, ARRHYTHMIA, NEUROLOGICAL DYSFUNCTION, KIDNEY DYSFUNCTION AND TRANSPLANT. THIS SINGLE-CENTER RETROSPECTIVE STUDY EVALUATED 151 PATIENTS WHO WERE IMPLANTED WITH A HM3 LEFT VENTRICULAR ASSIST DEVICE(LVAD) BETWEEN 2017 AND 2023 WHO WERE IMPLANTED WITH A DESTINATION THERAPY TREATMENT PLAN. THE AVERAGE AGE OF THE STUDY GROUP WAS 63.7 +/- 11.4 YEARS. 123 OF THE 151 PATIENTS WERE MALE. PATIENTS WHO RECEIVED A TRANSPLANT OR LVAD EXCHANGE AT ANY POINT FOR ANY REASON WERE EXCLUDED FROM THE STUDY THE GOAL OF THE STUDY WAS TO EVALUATE THE INCIDENCE AND FACTORS ASSOCIATED WITH READMISSION.12 PATIENTS WERE NOTED TO HAVE PASSED AWAY SHORTLY AFTER LVAD IMPLANT. 41 PATIENTS PASSED AWAY TOTAL. 38 PATIENTS HAD POSTOPERATIVE MAJOR BLEEDING, 6 PATIENTS HAD POSTOPERATIVE ISCHEMIC STROKE, 31 PATIENTS HAD POSTOPERATIVE RIGHT HEART FAILURE REQUIRING RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT, TEMPORARY HEMODIALYSIS WAS REQUIRED IN 23 PATIENTS POSTOPERATIVELY AND 27 PATIENTS REQUIRED A POSTOPERATIVE TRACHEOSTOMY. SOME OF THE PATIENTS WERE NOTED TO HAVE HAD TO BE ADMITTED TO THE INTENSIVE CARE UNIT (ICU) DUE TO THEIR POSTOPERATIVE COMPLICATIONS. THE 139 REMAINING PATIENTS EXPERIENCED 456 SEPARATE READMISSIONS WITH A MEDIAN FOLLOW-UP OF 590 (303-1002) DAYS AND A TOTAL FOLLOW-UP OF 270.7 PATIENT-YEARS FOR A RATE OF 1.7 READMISSIONS PER PATIENT YEAR. 108 OF THE PATIENTS HAD TO BE READMITTED AT LEAST ONE TIME. PATIENTS WERE READMITTED FOR HEART FAILURE EXACERBATION (27.0%), MAJOR INFECTION (17.8%) KIDNEY DYSFUNCTION (5.7%), CARDIAC ARRHYTHMIA (9.9%), PLANNED PROCEDURES (8.6%), NEUROLOGICAL DYSFUNCTION (4.8%) VARIOUS LVAD ALARMS/MALFUNCTIONS (3.7%), ANTICOAGULATION ADJUSTMENT (0.7%) OTHER UNKNOWN REASONS (8.3%), AND MAJOR BLEEDING (13.6%). PATIENTS WHO WERE READMITTED WITHIN 6 MONTHS WERE MORE LIKELY TO HAVE NEW YORK HEART ASSOCIATION CLASS III OR IV SYMPTOMS COMPARED TO THOSE WHO HAD NOT BEEN READMITTED 39.1% VERSUS 11.8% VERSUS 15%, P =0.0008. FOLLOWING LVAD IMPLANTATION, READMISSIONS WITHIN 6 MONTHS OF SURGERY ARE ASSOCIATED WITH WORSE FUNCTIONAL STATUS AT 6 MONTHS AND 1 YEAR POSTOPERATIVELY. INCREASED BODY MASS INDEX AND PREOPERATIVE BILIRUBIN WERE ASSOCIATED WITH MORE FREQUENT READMISSIONS. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B2: SPECIFIC DEATH DATE IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. EISENGA, JOHN B; MCCULLOUGH, KYLE A; AFZAL, AASIM; DIMAIO, J MICHAEL; MOUBARAK, GHADI; ET AL. THE JOURNAL OF SURGICAL RESEARCH308: 202-208. (APR 2025)VENTRICULAR ASSIST DEVICE. THE JOURNAL OF THORACIC AND CARDIOVASCULAR SURGERY, 167(4), 1322¿1330.E6. HTTPS://DOI.ORG/10.1016/J.JTCVS.2022.09.051. BAYLOR SCOTT AND WHITE, THE HEART HOSPITAL, PLANO, TEXAS. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22096947&pc=DSQ | 2916596-2025-03569 | 2025-05-04 04:00 | Death | THORATEC CORPORATION | 2025-05-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Septic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE PATIENT OUTCOME, THE PATIENT PASSED AWAY DUE TO SEPTIC SHOCK, REFRACTORY VASCOPLEGIA. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE PUMP WILL NOT BE EXPLANTED. ADDITIONAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22096952&pc=OZD | 1220648-2025-28775 | 2025-05-05 04:00 | Death | ABIOMED, INC. | 2025-05-28 04:00 | OZD | IMPELLA CP | Migration or Expulsion of Device | Cardiac Arrest | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA CP WAS PLACED ON A PATIENT WITH LATERAL WALL MI. THE IMPELLA WAS SUCCESSFULLY PLACED. DURING SUPPORT IT WAS REPORTED THAT THE PATIENT WENT INTO CARDIAC ARRESTED AFTER ARRIVAL TO THE UNIT. DURING CPR THE IMPELLA MIGRATED INTO THE AORTA AND WHILE USING ECHO GUIDANCE THEY COULD NOT CROSS BACK INTO THE VALVE. THE IMPELLA WAS PLACED IN THE DESCENDING AORTA AND PLACED ON P1 AS THE TEAM TRIED TO CONNECT WITH FAMILY TO DISCUSS GOALS OF CARE. ULTIMATELY, THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. POSITIONING ISSUES: THE CAUSE OF THE POSITIONING ISSUES WAS MOST LIKELY USE RELATED AS THE PUMP CAME OUT OF POSITION DURING CARDIOPULMONARY RESUSCITATION (CPR). VT/VF: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22097119&pc=OZD | 1220648-2025-28776 | 2025-05-02 04:00 | Death | ABIOMED, INC. | 2025-05-28 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Tachycardia; Thrombocytopenia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS), AND THE FOLLOWING DAY, THE PATIENT DEVELOPED DECREASED PULSES IN THE RIGHT LOWER EXTREMITY. THE IMPELLA CP WAS EXPLANTED AND AN IMPELLA 5.5 WAS IMPLANTED FOR ESCALATION OF CARE. DURING THIS TIME A FASCIOTOMY WAS PERFORMED TO RESOLVE THE LIMB ISCHEMIA, AND TWO UNITS OF BLOOD WERE GIVEN TO THE PATIENT FOLLOWING THE INTERVENTION. STRONG PULSES WERE NOTED STATUS POST FASCIOTOMY. THE PATIENT WAS NOTED TO BE IN STABLE CONDITION FOLLOWING THE EVENT AND CONTINUED ON IMPELLA 5.5 MCS. HOWEVER, THE PATIENT'S NEUROLOGICAL STATUS WAS NOT FULLY INTACT, PATIENT CONFUSED AND WAS GIVEN HALDOL. FOUR DAYS LATER, THE PATIENT EXPERIENCED HEPARIN INDUCED THROMBOCYTOPENIA (HIT), AND THE NEXT DAY A RUN OF VENTRICULAR TACHYCARDIA, WHICH WAS RESOLVED WITHOUT INTERVENTION. HOWEVER, THE PATIENT WAS STILL HIT POSITIVE. THE PATIENT'S NEUROLOGICAL STATUS WAS STILL NOT FULLY INTACT; THE PATIENT REMAINED CONFUSED. APPROXIMATELY NINE DAYS LATER THE PATIENT WOULD EXPIRE. CARE WAS WITHDRAWN. Manufacturer Narrative: MEDICAL SAFETY REVIEWED THE EVENT AND GIVEN THE DETAILS OF EVENT, IMPELLA 5.5 RELATED EVENTS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA 5.5 AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME. THE IMPELLA 5.5 IS ONE OF TWO DEVICES ASSOCIATED WITH THE REPORTED EVENT. REFER TO MEDICAL DEVICE REPORT FOR IMPELLA CP SERIAL NUMBER (B)(6) WITH DATE OF EVENT 38-APR-2025 AND PATIENT IDENTIFIER ENDING IN 3700. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22097317&pc=PNJ | 9612164-2025-02627 | 2023-01-01 05:00 | Death | MEDTRONIC IRELAND | 2025-05-28 04:00 | PNJ | MICRA | Use of Device Problem; Appropriate Term/Code Not Available | Pulmonary Embolism; Hematoma; Hemorrhage/Bleeding; Cardiac Tamponade; Thrombosis/Thrombus; Cardiovascular Insufficiency; Pericarditis; Insufficient Information | P150033 |  | 1 | Event Description: DATA WAS REVIEWED REGARDING IMPLANTS AND OUTCOMES OF LEADLESS IMPLANTABLE PULSE GENERATORS (IPGS). THEY DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE, THIRTY-DAY MORTALITY. THERE WERE PATIENTS WHO EXPERIENCED ACUTE, THIRTY-DAY COMPLICATIONS WHICH INCLUDED VASCULAR ANEURYSMS. THERE WERE PATIENTS WHO EXPERIENCED DEEP VEIN THROMBOSIS, PULMONARY EMBOLISMS, CARDIAC TAMPONADE, POST IMPLANT HEMATOMAS AND HEMORRHAGES, PERICARDITIS, AND UNKNOWN VASCULAR COMPLICATIONS. THERE WERE DISLODGMENTS OF LEADLESS IPGS AND OTHER UNKNOWN MALFUNCTIONS IN WHICH THERE WERE REVISIONS, REPLACEMENTS, AND OTHER INTERVENTIONS. THE STATUS OF THE DEVICES, DELIVERY SYSTEMS, AND INTRODUCERS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: D1: MICRA, D4: MODEL #: MC1VR01-DELSYS/ EXPIRATION DATE: UNKNOWN / SERIAL#: UNKNOWN, UDI #: ASKU, D9: NO, DEV RTN TO MFR? NO. H4: MFG DATE: UNKNOWN. H5: YES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22097322&pc=DYB | 9612164-2025-02628 | 2023-01-01 05:00 | Death | MEDTRONIC IRELAND | 2025-05-28 04:00 | DYB | MICRA | Use of Device Problem; Appropriate Term/Code Not Available | Pulmonary Embolism; Hematoma; Hemorrhage/Bleeding; Cardiac Tamponade; Thrombosis/Thrombus; Cardiovascular Insufficiency; Pericarditis; Insufficient Information | K132030 |  | 1 | Event Description: DATA WAS REVIEWED REGARDING IMPLANTS AND OUTCOMES OF LEADLESS IMPLANTABLE PULSE GENERATORS (IPGS). THEY DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE, THIRTY-DAY MORTALITY. THERE WERE PATIENTS WHO EXPERIENCED ACUTE, THIRTY-DAY COMPLICATIONS WHICH INCLUDED VASCULAR ANEURYSMS. THERE WERE PATIENTS WHO EXPERIENCED DEEP VEIN THROMBOSIS, PULMONARY EMBOLISMS, CARDIAC TAMPONADE, POST IMPLANT HEMATOMAS AND HEMORRHAGES, PERICARDITIS, AND UNKNOWN VASCULAR COMPLICATIONS. THERE WERE DISLODGMENTS OF LEADLESS IPGS AND OTHER UNKNOWN MALFUNCTIONS IN WHICH THERE WERE REVISIONS, REPLACEMENTS, AND OTHER INTERVENTIONS. THE STATUS OF THE DEVICES, DELIVERY SYSTEMS, AND INTRODUCERS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22097454&pc=OZD | 1220648-2025-28780 | 2025-04-28 04:00 | Death | ABIOMED, INC. | 2025-05-28 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding; Sepsis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT IN AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT, ESCALATED TO AN IMPELLA 5.5 DEVICE AND EXPERIENCED A HEMATOMA AT BOTH ACCESS SITES WITH PERSISTENT BLEEDING. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. THE PATIENT PRESENTED WITH CHEST PAIN AND LEFT HEART CATHETERIZATION REVEALED MULTI-VESSEL CORONARY ARTERY DISEASE AND ELEVATED END DIASTOLIC PRESSURE. THE DECISION WAS MADE TO PLACE IMPELLA CP VIA THE RIGHT FEMORAL ARTERY, WHICH WAS SUCCESSFULLY COMPLETED. THE PATIENT WAS PLACED ON P-6 SUPPORT LEVEL, NOTED TO BE HEMODYNAMICALLY STABLE AWAITING LIFE FLIGHT FOR TRANSFER TO A HIGHER LEVEL OF CARE FACILITY AND CARDIOVASCULAR CONSULTATION. APPROXIMATELY, 16 DAYS LATER, THE IMPELLA CP DEVICE WAS EXPLANTED; THE PATIENT WAS ESCALATED TO AN IMPELLA 5.5 VIA THE LEFT AXILLARY ARTERY FOR ESCALATION IN THERAPY. FOLLOWING, A LARGE HEMATOMA TO LEFT CHEST WALL ENCOMPASSING ENTIRE PECTORAL AND SHOULDER AREAS AND EXTENDING TO LEFT UPPER ARM HAD DEVELOPED. A RIGHT FEMORAL HEMATOMA AS WELL HAD DEVELOPED AT SITE OF IMPELLA CP EXPLANT. HEMOSTATIC AGENTS WERE USED IN THE POCKETS WHILE IN THE OPERATING ROOM AND LONG PERIODS OF MANUAL PRESSURE WERE ALSO APPLIED. FOUR "JUMBO" PLATELETS WERE GIVEN WHILE IN THE OPERATING ROOM AND AS WELL PROTAMINE WAS ADMINISTERED. THE PATIENT WAS CONTINUED ON THERAPY WITH THE IMPELLA 5.5 WITH POSSIBLE SEPSIS INDICATED APPROXIMATELY SIX DAYS LATER. THE NEXT DAY, THERE WAS NO COUGH, NO GAG, PUPILS UNREACTIVE, AND HEMODYNAMICS WERE DECLINING. THE PATIENT WOULD EXPIRE THE FOLLOWING DAY; CARE WAS WITHDRAWN. Manufacturer Narrative: MEDICAL SAFETY REVIEW OF THE EVENT NOTED LARGE HEMATOMA AT FEMORAL ACCESS SITE REQUIRING INTERVENTION IS SERIOUS INJURY FOR THE IMPELLA CP; THERE WAS ESCALATION OF CARE TO IMPELLA 5.5 AND ADVERSE EVENTS OCCURRING (THAT INCLUDES POSSIBLE SEPSIS) CONTRIBUTED MORE TO THE PATIENT¿S DEMISE. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA 5.5 DEVICE USED AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME. THE IMPELLA 5.5 IS ONE OF TWO DEVICES ASSOCIATED WITH THE EVENT. REFER TO MEDICAL DEVICE REPORT FOR IMPELLA CP SERIAL NUMBER (B)(6) WITH DATE OF EVENT 28-APR-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMARTASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: POTENTIAL ADVERSE EVENTS ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS: ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ Manufacturer Narrative: THE INVESTIGATION OF THE REPORTED ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22097479&pc=NFK | 1282497-2025-00365 | 2025-03-14 04:00 | Death | COVIDIEN LLC | 2025-05-28 04:00 | NFK | NA | Detachment of Device or Device Component | Cardiac Arrest; Hemorrhage/Bleeding; Ventricular Fibrillation; Asystole | K092797 |  | 1 | Event Description: ACCORDING TO THE REPORTER, DURING LONG-TERM DIALYSIS, THE VENOUS PORT WAS DISCONNECTED, WHICH LED TO SIGNIFICANT BLOOD LOSS. THE PATIENT BECAME PULSELESS, AND CPR (CARDIOPULMONARY RESUSCITATION) STARTED. THEY NOTED THE RHYTHM WAS VFIB (VENTRICULAR FIBRILLATION) AND THEN ASYSTOLE. THE PATIENT UNDERWENT A MASSIVE BLOOD TRANSFUSION PROTOCOL. ROSC (RETURN OF SPONTANEOUS CIRCULATION) WAS UNABLE TO BE ACHIEVED. THE PATIENT PASSED AWAY.Event Description: ACCORDING TO THE REPORTER, ON THE DAY OF THE EVENT, THE PATIENT WAS IN THE ER (EMERGENCY ROOM) FOR EMESIS, ELEVATED WBC (WHITE BLOOD CELLS), AND HGB (HEMOGLOBIN) 6.4. BASED ON LAB WORK, DIALYSIS WAS NEEDED, AND THE PATIENT WAS TAKEN TO THE DIALYSIS UNIT. THE BLOOD FLOW FOR DIALYSIS WAS 400, WHICH WAS CONSISTENT WITH NORMAL TREATMENT PRESCRIPTION. ACCORDING TO THEIR RECORDS, THE BLOODLINE CONNECTION WAS SECURE AND HAD A CONNECTOR ON IT THAT SECURED THE BLOODLINES TO THE CATHETER (WAS REFERRED TO AS A SECURE LOCK). SHORTLY AFTER THE START OF LONG-TERM DIALYSIS, THE DIALYSIS MACHINE ALARMED. THE NURSE RESPONDED TO THE MACHINE ALARM, AND IT WAS FOUND THAT THE VENOUS PORT/HUB HAD DISCONNECTED, WHICH LED TO SIGNIFICANT BLOOD LOSS. THE RAPID RESPONSE TEAM WAS ENGAGED. THE MACHINE WAS MAINTAINED BY THE 3RD PARTIES BIOMED. FOR THE HOSPITAL, IT STATED THE MACHINE WAS NOT A CONCERN; IT DID WHAT IT WAS SUPPOSED TO AND ALARMED ACCORDING TO THE 3RD PARTY. THE MACHINES WERE MAINTAINED BY A 3RD PARTY, AS THEY WERE THE 3RD PARTY¿S EQUIPMENT. THE HOSPITAL DID NOT HAVE ACCESS TO VERIFY THE MACHINE ALARMED. THE PATIENT WAS NON-COMPLIANT, AND THERE WAS NO INFORMATION ON HOW THE CATHETER HAD CARED FOR/MAINTAINED OUTSIDE OF DIALYSIS SESSIONS. THE PATIENT WAS KNOWN TO MOVE A BIT DURING THE TREATMENT, SO THERE MIGHT HAVE BEEN SOME TENSION. NO CRACKS OR CONCERNS WITH THE CATHETER UPON INSPECTION. THEY INVESTIGATED IF THE CATHETER SHOULD HAVE BEEN REMOVED AND REPLACED EARLIER BUT COULD NOT FIND ANYTHING THAT INDICATED A USAGE/TIMING LIMIT. THE PATIENT BECAME PULSELESS, AND CPR (CARDI OPULMONARY RESUSCITATION) STARTED. THEY NOTED THE RHYTHM WAS VFIB (VENTRICULAR FIBRILLATION) AND THEN ASYSTOLE. THE PATIENT UNDERWENT A MASSIVE BLOOD TRANSFUSION PROTOCOL. ROSC (RETURN OF SPONTANEOUS CIRCULATION) WAS UNABLE TO BE ACHIEVED. THE PATIENT PASSED AWAY ON THE SAME DAY AS THE EVENT.Event Description: ACCORDING TO THE REPORTER, ON THE DAY OF THE EVENT, THE PATIENT WAS IN THE ER (EMERGENCY ROOM) FOR EMESIS, ELEVATED WBC (WHITE BLOOD CELLS), AND HGB (HEMOGLOBIN) 6.4. BASED ON LAB WORK, DIALYSIS WAS NEEDED, AND THE PATIENT WAS TAKEN TO THE DIALYSIS UNIT. THE BLOOD FLOW FOR DIALYSIS WAS 400, WHICH WAS CONSISTENT WITH NORMAL TREATMENT PRESCRIPTION. ACCORDING TO THEIR RECORDS, THE BLOODLINE CONNECTION WAS SECURE AND HAD A CONNECTOR ON IT THAT SECURED THE BLOODLINES TO THE CATHETER (WAS REFERRED TO AS A SECURE LOCK). SHORTLY AFTER THE START OF LONG-TERM DIALYSIS, THE DIALYSIS MACHINE ALARMED. THE NURSE RESPONDED TO THE MACHINE ALARM, AND IT WAS FOUND THAT THE VENOUS PORT/HUB HAD DISCONNECTED, WHICH LED TO SIGNIFICANT BLOOD LOSS. THE RAPID RESPONSE TEAM WAS ENGAGED. THE MACHINE WAS MAINTAINED BY THE 3RD PARTIES BIOMED. FOR THE HOSPITAL, IT STATED THE MACHINE WAS NOT A CONCERN; IT DID WHAT IT WAS SUPPOSED TO AND ALARMED ACCORDING TO THE 3RD PARTY. THE MACHINES WERE MAINTAINED BY A 3RD PARTY, AS THEY WERE THE 3RD PARTY¿S EQUIPMENT. THE HOSPITAL DID NOT HAVE ACCESS TO VERIFY THE MACHINE ALARMED. IT WAS CONFIRMED THAT THE CATHETER HAD PREVIOUSLY UNDERGONE A REPAIR ON (B)(6) 2024. THE LOCUM PROVIDER¿S NOTE DID NOT PROVIDE SPECIFIC INFORMATION ABOUT THE REPAIR PROCEDURE, BUT THERE WAS NO NEED FOR BOTH L IMBS TO BE REPAIRED SINCE IT WAS THE CLAMP PART FROM THE CATHETER THAT WAS REPAIRED. THE DISCONNECTION OF THE LUER ADAPTER OCCURRED ON THE VENOUS LIMB OF THE CATHETER, ALIGNING WITH THE SIDE WHERE THE CLAMP WAS REPAIRED IN NOVEMBER. THE HOSPITAL NOTED THAT THE ONLY AVAILABLE REPAIR KIT AT THE TIME OF THE NOVEMBER REPAIR WAS THE ONE NOW IDENTIFIED AS THE KIT FROM WHICH THE LUER ADAPTER THAT DISCONNECTED IS BELIEVED TO HAVE ORIGINATED. THERE WAS NO DOCUMENTATION TO INDICATE THE CATHETER HAD BEEN REPAIRED OR ALTERED AGAIN AFTER THAT DATE. THE PATIENT WAS NON-COMPLIANT, AND THERE WAS NO INFORMATION ON HOW THE CATHETER CARED FOR/MAINTAINED OUTSIDE OF DIALYSIS SESSIONS. THE PATIENT HAD A PRE-EXISTING CONDITION OF CHRONIC PAIN, AND STAFF REPORTED THAT THE PATIENT CONSTANTLY MOVED AROUND IN THE CHAIR/BED DURING DIALYSIS TREATMENT, WHICH MIGHT HAVE CONTRIBUTED TO SOME TENSION ON THE CATHETER. NO CRACKS OR OBVIOUS DAMAGE TO THE CATHETER WERE OBSERVED UPON INSPECTION BY STAFF. THEY INVESTIGATED IF THE CATHETER SHOULD HAVE BEEN REMOVED AND REPLACED EARLIER BUT COULD NOT FIND ANYTHING THAT INDICATED A USAGE/TIMING LIMIT. THE PATIENT BECAME PULSELESS, AND CPR (CARDIOPULMONARY RESUSCITATION) STARTED. THEY NOTED THE RHYTHM WAS VFIB (VENTRICULAR FIBRILLATION) AND THEN ASYSTOLE. THE PATIENT UNDERWENT A MASSIVE BLOOD TRANSFUSION PROTOCOL. ROSC (RETURN OF SPONTANEOUS CIRCULATION) WAS UNABLE TO BE ACHIEVED. THE PATIENT PASSED AWAY ON THE SAME DAY AS THE EVENT. Manufacturer Narrative: H3. EVALUATION SUMMARY: MEDTRONIC CONDUCTED AN INVESTIGATION BASED UPON ALL INFORMATION RECEIVED. THE DEVICE WAS NOT RETURNED, BUT A PHOTO WAS AVAILABLE FOR EVALUATION. VISUAL INSPECTION NOTED A DETACHED VENOUS PORT AND CLIP, WITH VISIBLE SIGNS OF USE SUCH AS BLOOD. THERE APPEARED TO BE NO OTHER ABNORMALITIES WITH THE DEVICE. IT WAS REPORTED THAT LUER ADAPTER DETACHED FROM THE EXTENSION TUBE. THE REPORTED ISSUE WAS CONFIRMED. THE MOST LIKELY CAUSE COULD NOT BE ESTABLISHED FROM THE INFORMATION AVAILABLE. THE MANUFACTURING RECORDS FOR EACH DEVICE ARE THOROUGHLY REVIEWED PRIOR TO RELEASE TO ENSURE THAT IT MEETS ALL MEDTRONIC QUALITY SPECIFICATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: E1(FACILITY NAME, FAX NUMBER). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B2 (OUTCOME ATTRIBUTED TO ADVERSE EVENT, DATE OF DEATH), B5, D6A, G3, H11. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22097672&pc=DSQ | 3007042319-2025-01191 | 2025-05-22 04:00 | Death | HEARTWARE, INC. | 2025-05-28 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Obstruction of Flow; Pumping Problem | Chest Pain; Thrombosis/Thrombus | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT PRESENTED TO THE EMERGENCY ROOM WITH CHEST PAIN. THE PATIENT WAS FOUND TO HAVE A THROMBUS IN THE PUMP INFLOW CANNULA, WHICH WAS ASSOCIATED WITH A SUDDEN REDUCTION IN DEVICE FLOW AND SUSPECTED INFLOW CANNULA OBSTRUCTION. HIGH WATT AND LOW FLOW ALARMS WERE TRIGGERED, INDICATING POTENTIAL INFLOW CANNULA OBSTRUCTION. THE PATIENT WAS ON COMPLIANT WITH ANTICOAGULATION THERAPY, AS EVIDENCED BY AN INTERNATIONAL NORMALIZED RATIO (INR) AROUND 1, AND HAD A HISTORY OF COCAINE USE AS A RISK FACTOR FOR THROMBUS FORMATION. BLOOD TESTS, A COMPUTED TOMOGRAPHY (CT) SCAN (WHICH WAS NEGATIVE), AND A TRANSESOPHAGEAL ECHOCARDIOGRAM WERE CONDUCTED, THOUGH THE TEE RESULTS WERE NOT AVAILABLE. AFTER INTUBATION, THE VAD FLOW DECREASED TO ZERO, AND THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) FOR THREE DAYS BUT DID NOT RECOVER AND THE PATIENT DIED. THE CAUSE OF DEATH WAS DETERMINED TO BE VAD THROMBOSIS AND CARDIOGENIC SHOCK. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. ADDITIONAL PRODUCTS: D4: SERIAL OR LOT#: (B)(6). H3: YES H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. PRODUCT EVENT SUMMARY: (B)(6) AND THE CONTROLLER ((B)(6)) WERE NOT RETURNED FOR EVALUATION. NO PERFORMANCE ALLEGATIONS WERE MADE AGAINST THE CONTROLLER. A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS REVEALED AN INCREASE IN POWER CONSUMPTION AND ESTIMATED FLOWS BEGINNING ON (B)(6) 2025, LEADING TO PARAMETERS ABOVE NORMAL OPERATING RANGE, FOLLOWED BY A SUDDEN DECREASE IN PARAMETERS ON (B)(6) 2025. IN ADDITION, FOUR (4) HIGH WATT ALARMS AND TWENTY-SEVEN (27) LOW FLOW ALARMS WERE LOGGED SINCE (B)(6) 2025. OF NOTE, SEVERAL VAD SPEED AND ALARM LIMIT CHANGES WERE LOGGED SINCE (B)(6) 2025 AT 16:57:32. REVIEW OF THE EVENT LOG FILE REVEALED A CONTROLLER POWER DOWN AND ONE (1) VAD DISCONNECT ALARM WERE LOGGED ON (B)(6) 2025 AT 11:20:01, INDICATING A PHYSICAL DISCONNECTION OF THE DRIVELINE FROM THE CONTROLLER. AS A RESULT, THE REPORTED HIGH POWER AND LOW FLOW EVENTS WERE CONFIRMED. THE REPORTED THROMBUS EVENT COULD NOT BE CONFIRMED DUE TO INSUFFICIENT EVIDENCE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING VAD THROMBOSIS AND CARDIOGENIC SHOCK. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE AVAILABLE INFORMATION AND HISTORICAL REVIEW OF SIMILAR EVENTS, THE MOST LIKELY ROOT CAUSE OF THE HIGH-POWER EVENT CAN BE ATTRIBUTED TO EXTERNAL FACTORS SUCH AS THROMBUS FORMATION/INGESTION. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE REPORTED LOW FLOW EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. THE MOST LIKELY ROOT CAUSE OF VAD DISCONNECT ALARM CAN BE ATTRIBUTED TO PHYSICAL DISCONNECTION OF THE DRIVELINE FROM THE CONTROLLER. PER THE INSTRUCTIONS FOR USE, DEVICE THROMBUS, PAIN, DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF THROMBUS EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS, AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL, AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22097884&pc=PNJ | 9612164-2025-02631 | 2023-01-01 05:00 | Death | MEDTRONIC IRELAND | 2025-05-28 04:00 | PNJ | MICRA AV | Use of Device Problem; Appropriate Term/Code Not Available | Pulmonary Embolism; Cardiac Arrest; Fistula; Hematoma; Hemorrhage/Bleeding; Cardiac Tamponade; Cardiac Perforation; Thrombosis/Thrombus; Cardiovascular Insufficiency; Pericarditis; Insufficient Information | P150033 |  | 1 | Event Description: DATA WAS REVIEWED REGARDING IMPLANTS AND OUTCOMES OF LEADLESS IMPLANTABLE PULSE GENERATORS (IPGS). THEY DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE, THIRTY-DAY MORTALITY. THERE WERE PATIENTS WHO EXPERIENCED ACUTE, THIRTY-DAY COMPLICATIONS WHICH INCLUDED PUNCTURE SITE EVENTS SUCH AS ARTERIOVENOUS FISTULA AND VASCULAR ANEURYSMS. THERE WERE PATIENTS WHO EXPERIENCED DEEP VEIN THROMBOSIS, PULMONARY EMBOLISMS, CARDIAC PERFORATION, CARDIAC TAMPONADE, STENOSIS DUE TO DEVICE OR IMPLANT, POST IMPLANT HEMATOMAS OR HEMORRHAGES, INTRAOPERATIVE CARDIAC ARREST, PERICARDITIS, AND UNKNOWN VASCULAR COMPLICATIONS. THERE WERE DISLODGMENTS OF LEADLESS IPGS AND OTHER UNKNOWN MALFUNCTIONS IN WHICH THERE WERE REVISIONS, REPLACEMENTS, AND OTHER INTERVENTIONS. THE STATUS OF THE DEVICES, DELIVERY SYSTEMS, AND INTRODUCERS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: D1: MICRA, D4:MODEL #: MC1AVR1-DELSYS/ EXPIRATION DATE: UNKNOWN / SERIAL#: UNKNOWN, UDI #: ASKU, D9: NO, DEV RTN TO MFR? NO, H4: MFG DATE: UNKNOWN, H5: YES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22097892&pc=DYB | 9612164-2025-02632 | 2023-01-01 05:00 | Death | MEDTRONIC IRELAND | 2025-05-28 04:00 | DYB | MICRA | Use of Device Problem; Appropriate Term/Code Not Available | Pulmonary Embolism; Cardiac Arrest; Fistula; Hematoma; Hemorrhage/Bleeding; Cardiac Tamponade; Cardiac Perforation; Thrombosis/Thrombus; Cardiovascular Insufficiency; Pericarditis; Insufficient Information | K132030 |  | 1 | Event Description: DATA WAS REVIEWED REGARDING IMPLANTS AND OUTCOMES OF LEADLESS IMPLANTABLE PULSE GENERATORS (IPGS). THEY DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE, THIRTY-DAY MORTALITY. THERE WERE PATIENTS WHO EXPERIENCED ACUTE, THIRTY-DAY COMPLICATIONS WHICH INCLUDED PUNCTURE SITE EVENTS SUCH AS ARTERIOVENOUS FISTULA AND VASCULAR ANEURYSMS. THERE WERE PATIENTS WHO EXPERIENCED DEEP VEIN THROMBOSIS, PULMONARY EMBOLISMS, CARDIAC PERFORATION, CARDIAC TAMPONADE, STENOSIS DUE TO DEVICE OR IMPLANT, POST IMPLANT HEMATOMAS OR HEMORRHAGES, INTRAOPERATIVE CARDIAC ARREST, PERICARDITIS, AND UNKNOWN VASCULAR COMPLICATIONS. THERE WERE DISLODGMENTS OF LEADLESS IPGS AND OTHER UNKNOWN MALFUNCTIONS IN WHICH THERE WERE REVISIONS, REPLACEMENTS, AND OTHER INTERVENTIONS. THE STATUS OF THE DEVICES, DELIVERY SYSTEMS, AND INTRODUCERS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22098305&pc=DSQ | 3007042319-2025-01196 | 2025-05-20 04:00 | Death | HEARTWARE, INC. | 2025-05-28 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Adverse Event Without Identified Device or Use Problem | Fall; Head Injury; Intracranial Hemorrhage | P100047 |  | 1 | Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS ADMITTED TO WATCH FOR WORSENING.Event Description: IT WAS FURTHER REPORTED THAT AS THE PATIENT WAS RECOVERING, STARTING ON ANTICOAGULANTS HAD A RECURRENCE OF HEAD BLEED AND COMFORT CARE WAS INITIATED. THE VAD WAS TURNED OFF AND THE PATIENT PASSED AWAY.Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE (VAD) EXPERIENCED A FALL AT AN OUTSIDE FACILITY AND SUSTAINED A HEAD INJURY. ACCORDING TO THE PATIENT'S DAUGHTER, A HEAD COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED AND A SMALL INTRACRANIAL HEMORRHAGE WAS DETECTED. THE PATIENT WAS REPORTED TO BE ALIVE WITH INJURY. THE VAD REMAIN IN USE. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. CONTINUATION OF D10: 5071-53 LEAD IMPLANT DATE: (B)(6) 2015, 305U227 TISSUE VALVE IMPLANT DATE: (B)(6) 2015, 5071-53 LEAD IMPLANT DATE: (B)(6) 2015. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION RECEIVED. UPDATED B2, B5 AND H1. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT RECEIVED ANTICOAGULANTS AFTER EXPERIENCING AN INTRACRANIAL HEMORRHAGE; THE DECISION WAS MADE TO END THE VAD THERAPY, AND THE DEVICE WAS TURNED OFF. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, BLEEDING AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22098712&pc=OZD | 1220648-2025-28789 | 2025-04-30 04:00 | Death | ABIOMED, INC. | 2025-05-28 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Atrial Fibrillation; Hemolysis; Sepsis; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A 52-YEAR-OLD MALE PATIENT WITH CONTINUED MONOPHASIC FLOW TO RIGHT LOWER EXTREMITY (IMPELLA CP LEG) AND CONTINUED PRESSOR/INOTROPES/LOW EJECTION FRACTION, LOW PULSATILITY. DECISION WAS MADE TO ESCALATE TO IMPELLA 5.5 GIVEN PERSISTENT CARDIOGENIC SHOCK. THE 5.5 WAS SUCCESSFULLY IMPLANTED AND CP EXPLANTED. DURING 5.5 SUPPORT, IT WAS NOTED OF SOME SUCTION EVENTS AND PINK-TINGED URINE. IN ADDITION, THE PATIENT EXPERIENCED ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE (RVR). HOWEVER, IT WAS DOCUMENTED THAT AFTER REPOSITIONING, THAT HEMOLYSIS AND SUCTION RESOLVED. THE PATIENT RECEIVED 2 UNITS OF BLOOD FOR A GASTROINTESTINAL (GI) BLEED AND HEPARIN WAS PLACED ON HOLD. DURING THIS TIME, THE PATIENT EXPERIENCED INTERMITTENT FEVER. DUE TO THE SEPTIC COMPONENT, THE PATIENT WAS GIVEN ADDITION OF VASOPRESSIN AND LACTIC ACID. FAMILY ELECTED TO WITHDRAW CARE AND THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS, LOW PUMP FLOW, VF/VT/AF/AT, VASCULAR DAMAGE - PERIPHERAL VESSEL, AND INFECTION/SEPSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. HEMOLYSIS: THE CAUSE OF THE HEMOLYSIS ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED, BASED ON DATA LOG REVIEW AND GIVEN CLINICAL DETAILS. LOW PUMP FLOW: THE CAUSE OF THE LOW PUMP FLOW ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED, BASED ON DATA LOG REVIEW AND GIVEN CLINICAL DETAILS. VF/VT/AF/AT: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. VASCULAR DAMAGE - PERIPHERAL VESSEL: THE CAUSE OF THE VASCULAR DAMAGE ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND SUFFICIENT CLINICAL INFORMATION WAS PROVIDED. THE RELATION BETWEEN GI BLEED AND IMPELLA WAS UNKNOWN. INFECTION/SEPSIS: BECAUSE SUFFICIENT EVIDENCE HAS NOT BEEN PROVIDED, THE ROOT CAUSE OF THE INFECTION CANNOT BE DETERMINED. INSTRUCTIONS FOR USE: HEMOLYSIS: SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ INFECTION/SEPSIS: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & CP WITH SMARTASSIST SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ B1 PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28789. E4 SHOULD HAVE BEEN LEFT BLANK ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28789 AS IT IS UNKNOWN IF THE INITIAL REPORTER ALSO SENT THE REPORT TO THE FOOD AND DRUG ADMINISTRATION. H10 INSTRUCTIONS FOR USE WERE INCOMPLETE ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28789. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ THE MEDICAL SAFETY OFFICER (MSO) REVIEWED THE PATIENT'S OUTCOME AND DETERMINED IMPELLA 5.5 RELATED EVENTS (INCLUDING BUT NOT LIMITED TO LOW FLOW, SEPSIS ETC.) ARE OF A MORE DIRECT ISSUE AND THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THIS DEVICE USED AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22099349&pc=NPT | 2025587-2025-03927 | 2025-03-24 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-28 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Pneumonia | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 12 DAYS AFTER THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT HAD AN INCREASED INFLAMMATORY RESPONSE AND DEVELOPED PNEUMONIA. SUBSEQUENTLY, THE PATIENT WAS HOSPITALIZED AND WAS ADMINISTERED BACTERIAL AND FUNGAL INFECTION MEDICATION, HOWEVER; 18 DAYS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO INABILITY TO CONTROL THE INFECTION.Event Description: IT WAS REPORTED THAT APPROXIMATELY 12 DAYS AFTER THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT HAD AN INCREASED INFLAMMATORY RESPONSE AND DEVELOPED PNEUMONIA. SUBSEQUENTLY, THE PATIENT WAS HOSPITALIZED AND WAS ADMINISTERED BACTERIAL AND FUNGAL INFECTION MEDICATION, HOWEVER; 18 DAYS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO INABILITY TO CONTROL THE INFECTION. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS CONFIRMED TO BE PNEUMONIA. THE VALVE AND DCS CAN BE EXCLUDED AND CONTRIBUTING CAUSES TO THE DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. ADDED SECOND PARAGRAPH. D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22099615&pc=PYX | 1220648-2025-28794 | 2025-05-06 04:00 | Death | ABIOMED, INC. | 2025-05-28 04:00 | PYX | IMPELLA RP FLEX | Failure to Advance | Cardiac Arrest | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AT THE BEGINNING OF IMPELLA RP FLEX PLACEMENT, THE PHYSICIAN HAD VENOUS AND PULMONARY ARTERY ACCESS WITH WIRE AND WAS JUST ABOUT TO START WITH THE IMPELLA INTRODUCER KIT WHEN THE PATIENT CODED AND EXPIRED. PER MEDICAL SAFETY OFFICER REVIEW: THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE [GUIDEWIRE] USED AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. GIVEN THE INFORMATION AT HAND, ACCESS WAS MADE WITH THE GUIDEWIRE, USED AT THE TIME OF THE PATIENT'S DEMISE AND NO ALTERNATIVE CAUSE FOR THE EVENT IS IDENTIFIED AS A LIKELY CAUSE OF SUCH EVENT. INFORMATION IS LIMITED AND CAUSE OF DEATH DETAILS ARE UNNOTED. Manufacturer Narrative: A.1, A.4 AND A.5 ARE UNKNOWN. D.6A AND D.6B IMPLANT AND EXPLANT DATES ARE UNKNOWN. E.1 NAME PREFIX/TITLE, FIRST NAME AND LAST NAME ARE UNKNOWN. PER MEDICAL SAFETY OFFICER REVIEW: THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE [GUIDEWIRE] USED AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. GIVEN THE INFORMATION AT HAND, ACCESS WAS MADE WITH THE GUIDEWIRE, USED AT THE TIME OF THE PATIENT'S DEMISE AND NO ALTERNATIVE CAUSE FOR THE EVENT IS IDENTIFIED AS A LIKELY CAUSE OF SUCH EVENT. INFORMATION IS LIMITED AND CAUSE OF DEATH DETAILS ARE UNNOTED. THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER, SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿DO NOT ADVANCE OR WITHDRAW THE IMPELLA RP FLEX WITH SMARTASSIST SYSTEM CATHETER AGAINST RESISTANCE WITHOUT USING FLUOROSCOPY TO DETERMINE THE CAUSE OF THE RESISTANCE. DOING SO COULD RESULT IN SEPARATION OF THE CATHETER OR GUIDEWIRE TIP, DAMAGE TO THE CATHETER OR VESSEL, OR PERFORATION.¿ Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF DELIVERY ISSUE IS DETERMINED TO BE PATIENT CONDITION BECAUSE PATIENT UNFORTUNATELY EXPIRED BEFORE DEVICE AND THE INSERTION WAS ABORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22099650&pc=DSQ | 2916596-2025-03419 | 2025-05-08 04:00 | Death | THORATEC CORPORATION | 2025-05-28 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AT HOME ON HOSPICE ON (B)(6) 2025. THE PATIENT DIED FROM NON-CURABLE BACTEREMIA. THE VENTRICULAR ASSIST DEVICE (VAD) WAS DECOMMISSIONED, AND THE VAD FUNCTIONED AS INTENDED. THE PATIENT HAD A CHRONIC INFECTION AND WAS ON LIFELONG ANTIBIOTICS. THEY FOUGHT FOR A YEAR AND GOT TIRED, ELECTING TO STOP ALL TREATMENT AND START AT HOME HOSPICE WHERE THEY PASSED AWAY. THE SOURCE OF BACTEREMIA WAS SUSPECTED FROM THE DRIVELINE EXIT SITE. THERE WAS DRIVELINE EXIT SITE DRAINAGE, WITH POSITIVE WOUND AND BLOOD CULTURES. THE INFECTION WAS TREATED WITH ANTIBIOTICS. THE PATIENT HAD A DRIVELINE INFECTION, BUT THIS WAS NOT A DEVICE RELATED DEATH. THERE WAS NO PUMP MALFUNCTION AND IT WAS TURNED OFF AT THE END OF LIFE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22100252&pc=LCO | 1217183-2025-00032 | 2025-05-23 04:00 | Death | INSTRUMENTATION LABORATORY CO. | 2025-05-28 04:00 | LCO | HEMOSIL HIT- AB(PF4-H) | False Negative Result | Stroke/CVA; Thrombosis/Thrombus | K153137 |  | 1 | Event Description: A COMPLAINT HAS BEEN FILED FOR A DISCREPANT HIT TEST RESULT. THE PATIENT SAMPLE ID (B)(6), TESTED ON THE ACL TOP USING REAGENT/CALIBRATOR PN 00020014600, YIELDED A NEGATIVE RESULT. THE SAME SAMPLE, WHEN SENT TO MAYO CLINIC, RETURNED A POSITIVE SRA RESULT (CUT-OFF: >0.400 POSITIVE). THE SAMPLE WAS INITIALLY SENT TO MAYO AFTER PHARMACY QUESTIONED THE NEGATIVE RESULT, GIVEN THE PATIENT'S CLINICAL HISTORY OF A LEFT APICAL THROMBUS, STROKE, AND SUBSEQUENT DEATH UNDER PALLIATIVE CARE. THE SITE HAS NOW DISCONTINUED HIT TESTING ON THE ACL TOP AND IS SENDING SAMPLES TO QUEST. Manufacturer Narrative: QC AND MANUFACTURING DATA FOR HEMOSIL HIT-AB(PF4-H) LOT B37250 WERE REVIEWED AND NO ISSUES OR TRENDS WERE IDENTIFIED WITH THE IMPACTED LOT. REVIEW OF SUBMITTED CALIBRATION REPORT CONFIRMS THAT HIT-AB LOT B37250 WAS CALIBRATED ON 4/17/2025 APRIL 17, 2025, USING THE CORRECT TARGET VALUE OF 5.5 U/ML, AND THE RESULTS WERE PROPRELY VALIDATED. QC RESULTS WERE PERFORMED AFTER THE CALIBRATION WAS VALIDATED AND WERE CONFIRMED WITHIN ACCEPTABLE RANGE FOR HIT_AB CONTROLS LOT B37286. REVIEW OF THE SUBMITTED QC LEVEY JENNINGS CHARTS FOR 05/01/2025 TO 05/19/2025 SHOWS THE HIT ANTIBODIES FOR ACL TOP 350, HIT LOW AND HIGH LOT # B37286 REPORTED ACCEPTABLE RESULTS ON 05/08/2025 PRIOR TO SAMPLE TESTING. REVIEW OF THE SERVICE HISTORY SHOWS CALL 2005283457 ON 5/27/2025 - THE FIELD SERVICE ENGINEER VERIFIED THE INSTRUMENT FOR THE HIT DISCREPANT RESULTS. PREVENTATIVE MAINTENANCE FOR THE INSTRUMENT WAS LAST COMPLETED ON 4/25/25. NO INSTRUMENT ISSUES FOUND AT THAT TIME. REVIEWED GENERAL LOG LIST, CLEANED AND INSPECTED INSTRUMENT; NO ISSUES FOUND. FLUIDIC PRECISION PASSES BOTH ARMS. FLOW RATE PERFORMED AND PASSED. NO LEAKS FOUND. OPTICAL READING UNIT (ORU) REFERENCE READINGS IN RANGE AND STABLE. LIQUID LEVEL DETECTION (LLD) COUNTS FOR BOTH PROBES WERE NORMAL AND STABLE. COORDINATE DELTAS WERE LOW AND STABLE. NO INSTRUMENT RELATED ISSUES WERE IDENTIFIED. SINCE THE AFFECTED SAMPLE IS NOT AVAILABLE TO PERFORM TESTING, NO FURTHER INVESTIGATIONAL TESTING CAN BE DONE TO DETERMINE THE ROOT CAUSE OF THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE (IFU) OF THE PRODUCT REPORTS A TOTAL PERCENTAGE AGREEMENT OF 82% IN METHOD COMPARISON AND 90% IN EXTERNAL VALIDATION, VS ELISA ASSAY, 87% VS SRA. THESE FIGURES SUGGEST THAT DISCREPANCIES MAY ARISE IN SPECIFIC SAMPLES WHEN COMPARING HEMOSIL HIT-AB(PF4-H) WITH OTHER METHODOLOGIES. FOR THIS REASON, THE IFU STATES THAT CONFIRMATION OF RESULTS WITH A FUNCTIONAL TEST IS RECOMMENDED AND THAT THE ASSAY RESULTS SHOULD BE USED WITH OTHER INFORMATION, INCLUDING THE CLINICAL CONTEXT IN FORMING A DIAGNOSIS. BASED ON THIS ASSESSMENT, NO REMEDIAL ACTION IS REQUIRED. THIS INCIDENT WILL BE MONITORED THROUGH TRENDING ANALYSIS. Manufacturer Narrative: THIS IS AN INITIAL REPORT. A FOLLOW-UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22100307&pc=DSQ | 2916596-2025-03110 | 2025-04-30 04:00 | Death | THORATEC CORPORATION | 2025-05-28 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST ON (B)(6) 2025. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE SITE BELIEVED THAT THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: SECTION D4: LOT NUMBER WAS NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU CAN ALSO BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22100332&pc=NKE | 2182208-2025-02601 | 2017-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-05-28 04:00 | NKE | MEDTRONIC CRT-P | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: DATA WAS REVIEWED REGARDING IMPLANTS AND OUTCOMES OF BIVENTRICULAR PACING VERSUS CONDUCTION SYSTEM PACING. THEY DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE, THIRTY-DAY MORTALITY. THERE WERE PATIENTS IN BOTH GROUPS WHO EXPERIENCED HEART FAILURE HOSPITALIZATIONS, EMBOLISMS, THROMBOSIS, CARDIAC EFFUSION, PERFORATION, AND UNKNOWN DEVICE RELATED COMPLICATIONS. PATIENTS UNDERWENT UNKNOWN INTERVENTIONS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: A LIST OF POTENTIAL DEVICE AND LEAD MODELS ARE IN ATTACHED EXCEL FILE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22100339&pc=BZD | 2518422-2025-105844 | 2025-04-28 04:00 | Death | RESPIRONICS, INC. | 2025-05-28 04:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem |  | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK FILTER, BREATHING ISSUE, AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THERE ARE NO REPORTABLE MALFUNCTION. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22100370&pc=NIK | 2182208-2025-02602 | 2017-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-05-28 04:00 | NIK | MEDTRONIC CRT-D | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: DATA WAS REVIEWED REGARDING IMPLANTS AND OUTCOMES OF BIVENTRICULAR PACING VERSUS CONDUCTION SYSTEM PACING. THEY DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE, THIRTY-DAY MORTALITY. THERE WERE PATIENTS IN BOTH GROUPS WHO EXPERIENCED HEART FAILURE HOSPITALIZATIONS, EMBOLISMS, THROMBOSIS, CARDIAC EFFUSION, PERFORATION, AND UNKNOWN DEVICE RELATED COMPLICATIONS. PATIENTS UNDERWENT UNKNOWN INTERVENTIONS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22100658&pc=NVZ | 2182208-2025-02604 | 2017-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-05-28 04:00 | NVZ | MEDTRONIC IPG | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: DATA WAS REVIEWED REGARDING IMPLANTS AND OUTCOMES OF BIVENTRICULAR PACING VERSUS CONDUCTION SYSTEM PACING. THEY DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE, THIRTY-DAY MORTALITY. THERE WERE PATIENTS IN BOTH GROUPS WHO EXPERIENCED HEART FAILURE HOSPITALIZATIONS, EMBOLISMS, THROMBOSIS, CARDIAC EFFUSION, PERFORATION, AND UNKNOWN DEVICE RELATED COMPLICATIONS. PATIENTS UNDERWENT UNKNOWN INTERVENTIONS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: A LIST OF POTENTIAL DEVICE AND LEAD MODELS ARE IN ATTACHED EXCEL FILE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22100670&pc=DXY | 2182208-2025-02605 | 2017-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-05-28 04:00 | DXY | MEDTRONIC LEAD | Appropriate Term/Code Not Available | Insufficient Information |  |  | 1 | Event Description: DATA WAS REVIEWED REGARDING IMPLANTS AND OUTCOMES OF BIVENTRICULAR PACING VERSUS CONDUCTION SYSTEM PACING. THEY DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE, THIRTY-DAY MORTALITY. THERE WERE PATIENTS IN BOTH GROUPS WHO EXPERIENCED HEART FAILURE HOSPITALIZATIONS, EMBOLISMS, THROMBOSIS, CARDIAC EFFUSION, PERFORATION, AND UNKNOWN DEVICE RELATED COMPLICATIONS. PATIENTS UNDERWENT UNKNOWN INTERVENTIONS. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: A LIST OF POTENTIAL DEVICE AND LEAD MODELS ARE IN ATTACHED EXCEL FILE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22100866&pc=NPT | 2015691-2025-04209 | 2008-01-01 05:00 | Death | EDWARDS LIFESCIENCES | 2025-05-28 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFSCIENCES AFFILIATE IN ISRAEL THROUGH REVIEW OF MEDICAL ARTICLE "WHEN BRAIN AND HEART COLLIDE: A DEEPER DIVE INTO TREATMENT PATHWAYS OF STROKE COMPLICATING TAVI" , CORRESPONDING AUTHOR MARK KHEIFETS, THE FOLLOWING EVENTS WAS/WERE IDENTIFIED AS PERTAINING TO AN EDWARDS DEVICE: A 86-Y-O PATIENT IMPLANTED WITH A 29MM SAPIEN 3 VALVE WAS DIAGNOSED WITH SEVERE STROKE DUE TO MIDDLE CEREBRAL ARTERY (MCA) OCCLUSION 1.5 HOURS AFTER THE PROCEDURE. THE PATIENT RECEIVED CONSERVATIVE TREATMENT AND PASSED AWAY 13 DAYS LATER. THE AIMS OF THIS STUDY WERE TO DESCRIBE THE INCIDENCE, CHARACTERISTIC, AND OUTCOMES OF PATIENTS PRESENTING WITH ACUTE ISCHEMIC STROKE COMPLICATING TAVI (AISCT), AND TO COMPARE CONSERVATIVE MANAGEMENT (CM) VS. NEUROINTERVENTION (NI), IN ELIGIBLE PATIENTS IN THE (B)(6) CENTER BETWEEN 2008 AND 2021. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED TO REFERENCE RELATED REPORTS. THIS IS ONE OF FIVE MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. PLEASE REFERENCE RELATED MANUFACTURER REPORTS: 2015691-2025-04205. 2015691-2025-04206. 2015691-2025-04207. 2015691-2025-04208. 2015691-2025-04209. Manufacturer Narrative: E1 PHONE NUMBER: (B)(6). THIS IS ONE OF FIVE MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. REFERENCE ARTICLE KHEIFETS, MARK, ET AL. "WHEN BRAIN AND HEART COLLIDE: A DEEPER DIVE INTO TREATMENT PATHWAYS OF STROKE COMPLICATING TAVI." CARDIOVASCULAR INTERVENTION AND THERAPEUTICS (2025): 1-12. THE DATES OF THE EVENTS ARE UNKNOWN; HOWEVER, THE ARTICLE PROVIDED THE FOLLOWING DATE RANGE ON PAGE 1: FROM 2008 UNTIL 2021. THEREFORE, 01-JAN-2021 WAS USED AS THE OCCURRENCE DATE. PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRAPROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT FACTORS (AGE) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22101746&pc=NPT | 2025587-2025-03932 | 2024-11-29 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-28 04:00 | NPT | MEDTRONIC TRANSCATHETER VALVE | Patient Device Interaction Problem | Insufficient Information | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING PERI-PROCEDURAL OUTCOME ACCORDING TO VARC-3 CRITERIA AND HEMODYNAMIC MID-TERM FOLLOW-UP AFTER VALVE-IN-VALVE TRANSCATHETER AORTIC VALVE REPLACEMENT FOR DETERIORATED AORTIC BIOPROSTHESIS. THE STUDY POPULATION INCLUDED 41 PATIENTS WHO UNDERWENT VALVE-IN-VALVE TRANSCATHETER AORTIC VALVE IMPLANTATION (VIV-TAVI). MULTIPLE MANUFACTURER¿S DEVICES WERE IMPLANTED IN THE STUDY POPULATION; 10 PATIENTS WERE IMPLANTED WITH A MEDTRONIC MOSAIC (N=6) OR FREESTYLE (N=4) SURGICAL BIOPROSTHESIS AND 29 PATIENTS WERE IMPLANTED WITH A MEDTRONIC COREVALVE, EVOLUT R, EVOLUT PRO OR EVOLUT PRO+ TRANSCATHETER BIOPROSTHETIC VALVE. DEATHS OCCURRED IN THE STUDY POPULATION INCLUDING ONE DEEMED A ¿VALVE-RELATED MORTALITY¿ WHICH OCCURRED DURING THE TWO-YEAR FOLLOW-UP PERIOD. NO FURTHER DETAILS WERE PROVIDED ON THIS DEATH. AS FOR THE REMAINING PATIENT DEATHS THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. AMONG ALL SURGICAL VALVE PATIENTS, CLINICAL OBSERVATIONS INCLUDED: UNSPECIFIED STRUCTURAL VALVE DYSFUNCTION, AND MIGRATION OF A FREESTYLE VALVE. AMONG ALL TRANSCATHETER VALVE PATIENTS, CLINICAL OBSERVATIONS INCLUDED: PATIENT-PROSTHESIS MISMATCH, HIGH GRADIENTS 27 MMHG, STROKE, CORONARY OCCLUSION, ARRHYTHMIA REQUIRING PERMANENT PACEMAKER IMPLANT, RENAL FAILURE, AND MYOCARDIAL INFARCTION. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: TAKAGI ET AL. PERI-PROCEDURAL OUTCOME ACCORDING TO VARC-3 CRITERIA AND HEMODYNAMIC MID-TERM FOLLOW-UP AFTER VALVE-IN-VALVE TRANSCATHETER AORTIC VALVE REPLACEMENT FOR FAILED AORTIC BIOPROSTHESIS. CARDIOVASC INTERV THER. 2025 JAN;40(1):164-176. DOI: 10.10 07/S12928-024-01063-9. EPUB 2024 NOV 29. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22102271&pc=NIK | 2017865-2025-64931 | 2025-05-07 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-28 04:00 | NIK | ENTRANT HF | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.Event Description: NEW INFORMATION WAS NOTES THAT PATIENT HAD LOW BLOOD PRESSURE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: CORRECTION: TO B5 FOR THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. STATEMENT SHOULD HAVE BEEN ADDED AND H6 FOR HYPOTENSION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22102282&pc=NPU | 2015691-2025-04220 | 2025-02-20 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-05-28 04:00 | NPU | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER VALVE | Perivalvular Leak | Arrhythmia; Asystole | P140031 |  | 1 | Event Description: A REPORT WAS RECEIVED BY THE EDWARDS IMPLANT PATIENT REGISTRY INDICATING THE PATIENT EXPIRED 6 DAYS POST INDEX PROCEDURE. THE PATIENT HAD PREVIOUSLY UNDERGONE A MITRACLIP BUT HAD RESIDUAL SEVERE MITRAL REGURGITATION. THE PATIENT WAS IN CARDIOGENIC SHOCK, ON MULTIPLE VASOACTIVE MEDICATIONS, AS WELL AS AN IABP. THE PATIENT WAS NOT AN OPERATIVE CANDIDATE; HOWEVER, IT WAS DECIDED TO IMPLANT A SAPIEN VALVE MEDIAL TO A MITRA CLIP, AS WELL AS A MELODY VALVE INSIDE A CP STENT (SIDE-BY-SIDE) TO OCCUPY THE PATIENT'S LARGE MITRAL ORIFICE. THE PROCEDURE WAS DEEMED AN EMERGENCY. BOTH VALVES WERE SUCCESSFULLY DEPLOYED. AN INTRA OP TEE REVEALED TRACE VALVULAR LEAK THROUGH BOTH THE SAPIEN AND THE MELODY VALVES. THERE WERE A MULTITUDE OF SMALL JETS OF PVL THAT IN TOTAL WERE LIKELY MODERATE REGURGITATION. THE MVMG WAS 1 MMHG. THE RESIDUAL MR WAS DEEMED DIFFICULT TO QUANTIFY, BUT LIKELY MODERATE DUE TO ITS ORIGIN BEING BETWEEN THE TWO VALVES AND LIKELY THROUGH THE MITRA CLIP. DURING THE PROCEDURE, THE PATIENT EXPERIENCED 2 INTRA-OP VF CARDIAC ARRESTS TX WITH DEFIB X 2 WITH 2-MINTUES DOWN TIME WHICH REQUIRED EPINEPHRINE IV. THE TEAM WAS ABLE TO SUCCESSFULLY WEAN THE PATIENT FROM BYPASS WITH ONLY VASOACTIVE SUPPORT AND THE IAPB. IT WAS NOTED THERE WAS DISLODGEMENT OF THE TEMPORARY PACER AND DURING THAT TIME THE PATIENT WAS COMPLETELY ASYSTOLIC WITH NO RELIABLE ESCAPE RHYTHM. A TEMPORARY PACER WAS PLACED IN THE RT IJ AND SECURED. EP WAS TO EVALUATE FOR PPM AT A LATER DATE. CPB WAS WEANED OFF SUCCESSFULLY AT THE END OF THE PROCEDURE. THE PATIENT WAS ABLE TO BE TRANSFERRED TO THE CCU IN CRITICAL CONDITION. AN ECHO TAKEN ON POD-2 REVEALED MILD TO MODERATE MITRAL REGURGITATION, LIKELY PARAVALVULAR WITH A MVMG OF 5 MMHG. FOLLOWING THE INTERVENTION, THE PATIENT NEVER FOLLOWED COMMANDS. EEG FINDINGS WERE INDICATIVE OF A SEVERE DEGREE OF ENCEPHALOPATHY ETIOLOGY, NONSPECIFIC AND ONGOING TOXIC METABOLIC ABNORMALITIES AND ONGOING ANALGESIA ARE AT LEAST PARTIALLY CONTRIBUTING. THE PATIENT WENT INTO MULTISYSTEM ORGAN FAILURE AND EXTREME CARDIOGENIC SHOCK. THE CV TEAM WAS UNABLE TO WEAN THE PATIENT OFF OF INOTROPES OR THE IABP. ULTIMATELY, THE FAMILY REQUESTED TO TRANSITION TO COMFORT CARE AND THE PATIENT WAS PRONOUNCED DEAD WITH THE FAMILY AT THE BEDSIDE. AFTER REVIEW OF THE MEDICAL RECORDS PROVIDED, THE TMVR PROCEDURAL PERFORMED AT THE SAME TIME OF THE PLACEMENT OF THE MELODY VALVE IN A CP STENT WAS AN EMERGENCY OFF-LABEL CASE. MODERATE PVL WAS NOTED S/P TMVR. THE PATIENT ALSO DEVELOPED A LIFE-THREATENING ARRHYTHMIA FOR WHICH A TEMPORARY PACEMAKER REMAINED IN PLACE BUT WOULD REQUIRE A PPM. AS AN EEG CONFIRMED A SEVERE DEGREE OF BRAIN ENCEPHALOPATHY AND ONGOING TOXIC METABOLIC ABNORMALITIES, THE FAMILY ULTIMATELY DECIDED ON COMFORT CARE MEASURES AND PATIENT PASSED AWAY. THE EWS BIOPROSTHESIS MAY HAVE CONTRIBUTED TO THE PATIENT DEATH. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PARAVALVULAR LEAK (PVL) IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE AND THE LANDING ZONE, AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE. SOME PVL IS NOT UNCOMMON POST-DEPLOYMENT. MANY CASES ARE MILD TO MODERATE, AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS. OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION. THE PATIENT SCREENING MANUAL AND THE PROCEDURE DIDACTIC IDENTIFY SEVERAL PROCEDURAL AND ANATOMICAL FACTORS WHICH COULD CONTRIBUTE TO PVL, INCLUDING DEVICE MALPOSITION, INACCURATE MEASUREMENT OF THE LANDING ZONE, UNEVEN DISTRIBUTION OF CALCIUM ON THE LANDING ZONE, BULKY OR SEVERE CALCIFICATION, AN ELLIPTICAL ANNULUS SHAPE, AND VALVE UNDER-SIZING. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). THE DEVICE TRAINING MANUALS INSTRUCTS THE OPERATOR ON PROPER IMAGING SCREENING REQUIREMENTS, INCLUDING THE USE OF GOOD QUALITY ECHOCARDIOGRAPHY AND/OR COMPUTED TOMOGRAPHY (CT) TO APPROPRIATELY MEASURE THE LANDING ZONE, ASSESS THE CONTENT AND DISTRIBUTION OF CALCIUM, AND OTHER PATIENT ANATOMICAL FACTORS. MULTIPLE IMAGING MODALITIES SHOULD BE CONSIDERED DURING VALVE SIZE SELECTION. CONTRAINDICATIONS, IMPORTANT CONSIDERATIONS WHEN ASSESSING THE VALVE, AND CHOOSING THE PROPER TRANSCATHETER HEART VALVE ARE ALSO DISCUSSED. THE DEVICE TRAINING MANUALS ALSO INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING PATIENTS SCREENING AND PROCEDURAL CONSIDERATIONS. DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES ARE ALSO INCLUDED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THE OFF-LABEL USE COULD HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. THE DEVICE WAS USED IN OFF-LABEL IMPLANTATION IN THE NATIVE MITRAL POSITION. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO NATIVE MITRAL VALVE PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22102530&pc=LPB | 2029046-2025-01736 | 2025-05-01 04:00 | Death | BIOSENSE WEBSTER INC | 2025-05-28 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade | P030031 |  | 1 | Event Description: THE REPORT INDICATED THAT AFTER COMPLETING A CARDIAC ABLATION PROCEDURE WITH A THERMOCOOL SMARTTOUCH SF (STSF), THE PATIENT EXPERIENCED PERICARDIAL EFFUSION TREATED WITH PERICARDIOCENTESIS TO DRAIN BLOOD. THE PATIENT WAS TAKEN TO THE OPERATING ROOM (OR) FOR SURGERY. A PERICARDIAL EFFUSION WAS NOTICED WHILE USING INTRACARDIAC ECHOCARDIOGRAPHY (ICE), TOWARDS THE END OF THE CASE, AFTER ABLATION WAS COMPLETED. THE INTERVENTION TEAM USED ANOTHER FORM OF ULTRASOUND TO CONFIRM. THE PERICARDIOCENTESIS PROCEDURE WAS COMPLETED TO DRAIN THE BLOOD, AND THE PATIENT WAS TAKEN TO THE OPERATING ROOM (OR) FOR SURGERY. THE STATUS OF THE PATIENT WAS UNKNOWN AT THE TIME OF INITIAL REPORTING. THE SOUNDSTAR® CATHETER WAS IN THE BODY WHEN THE ISSUE WAS NOTICED. THE PHYSICIAN¿S OPINION OF THE EVENT WAS THAT NEITHER CATHETER, STSF NOR SOUNDSTAR CAUSED THE ISSUE. HOWEVER, THE ABLATOR WAS IN THE AREA IN QUESTION (LEFT VENTRICLE) AND THE SOUNDSTAR WAS IN THE RIGHT ATRIUM/RIGHT VENTRICLE. ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THAT THE PATIENT EXPIRED. HOWEVER, THE DATE OF DEATH, AS WELL AS THE CAUSE OF DEATH IS UNKNOWN. THE PHYSICIAN¿S OPINION ON THE CAUSE OF CARDIAC TAMPONADE WAS PATIENT CONDITION RELATED. NO RELEVANT TESTS/LABORATORY DATA, AND OTHER RELEVANT HISTORY/PREEXISTING CONDITION REPORTED WAS VALVE REPLACEMENT. NO TRANSSEPTAL PUNCTURE SITE WAS PERFORMED DURING THE PROCEDURE. Manufacturer Narrative: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22085767&pc=DSP | 2248146-2025-0000287 | 2025-04-22 04:00 | Death | DATASCOPE CORP. | 2025-05-27 04:00 | DSP | MEGA 7.5 FR. 40CC IAB (CHINA) | Break | Embolism/Embolus; Insufficient Information; No Clinical Signs, Symptoms or Conditions | K120868 |  | 1 | Event Description: IT WAS REPORTED THAT DURING INTRA AORTIC BALLOON (IAB) THERAPY, AFTER SUCCESSFUL PUNCTURE, THE GUIDEWIRE BROKE DURING THE IABP CATHETER PLACEMENT.Event Description: IT WAS REPORTED THAT DURING ROUTINE INTRA-AORTIC BALLOON (IAB) THERAPY INVOLVING STANDARD DISINFECTION, DRAPING, ADMINISTRATION OF LOCAL LIDOCAINE ANESTHESIA, AND SELDINGER PUNCTURE OF THE RIGHT FEMORAL ARTERY, VASCULAR ACCESS WAS SUCCESSFULLY OBTAINED. HOWEVER, DURING THE INSERTION OF THE IAB CATHETER, THE GUIDEWIRE FRACTURED. THE FIRST IAB CATHETER LUMEN WAS NOTED TO BE FOLDED, AND A REPLACEMENT BALLOON WAS SUBSEQUENTLY INSERTED. FOLLOWING THE PLACEMENT OF THE SECOND IAB CATHETER, IT WAS NOT CONNECTED TO THE INTRA-AORTIC BALLOON PUMP (IABP). THE PATIENT PASSED AWAY SHORTLY THEREAFTER.Event Description: IT WAS REPORTED THAT DURING ROUTINE INTRA-AORTIC BALLOON (IAB) THERAPY INVOLVING STANDARD DISINFECTION, DRAPING, ADMINISTRATION OF LOCAL LIDOCAINE ANESTHESIA, AND SELDINGER PUNCTURE OF THE RIGHT FEMORAL ARTERY, VASCULAR ACCESS WAS SUCCESSFULLY OBTAINED. HOWEVER, DURING THE INSERTION OF THE IAB CATHETER, THE GUIDEWIRE FRACTURED. THE FIRST IAB CATHETER LUMEN WAS NOTED TO BE FOLDED, AND A REPLACEMENT BALLOON WAS SUBSEQUENTLY INSERTED. FOLLOWING THE PLACEMENT OF THE SECOND IAB CATHETER, IT WAS NOT CONNECTED TO THE INTRA-AORTIC BALLOON PUMP (IABP). THE PATIENT PASSED AWAY SHORTLY THEREAFTER.Event Description: N.A.Event Description: IT WAS REPORTED THAT DURING ROUTINE PREPARATION FOR INTRA-AORTIC BALLOON PUMP (IABP) THERAPY¿INCLUDING DISINFECTION, DRAPING, ADMINISTRATION OF LOCAL LIDOCAINE ANESTHESIA, AND SELDINGER PUNCTURE OF THE RIGHT FEMORAL ARTERY, THE GUIDEWIRE FRACTURED DURING THE PLACEMENT OF THE IABP CATHETER, DESPITE SUCCESSFUL PUNCTURE. Manufacturer Narrative: UPDATED FIELDS: B4, G3, G6, H2, H11. CORRECTED FIELDS: B1, B2, B5, H1, H6 (HEALTH EFFECT CLINICAL CODE AND IMPACT CODE). Manufacturer Narrative: UPDATED FIELDS: B4, D9, G1 (CONTACT PERSON MFG SITE), G3, G6, H2, H3, H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS). CORRECTED FIELD: D10, H6 (HEALTH EFFECT CLINICAL CODE). NO DECON OR VISUAL INSPECTION PERFORMED SINCE PRODUCT WAS NOT RETURNED AND COULD NOT BE EVALUATED. THEREFORE, WE WERE UNABLE TO CONFIRM OR DETERMINE A ROOT CAUSE FOR THE REPORTED FAILURE. Manufacturer Narrative: UPDATED FIELDS: B2, B4, G3, G6, H2, H11. CORRECTED FIELD: H6 (HEALTH EFFECT CLINICAL CODE, TYPE OF INVESTIGATION, COMPONENT CODES). Manufacturer Narrative: UPDATED FIELDS: B4, G3, G6, H1, H2, H11. CORRECTED FIELDS: B1, B5, D4 (EXP DATE AND UDI), H1, H4, H6 (HEALTH EFFECT CLINICAL CODE AND IMPACT CODE). Manufacturer Narrative: THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER SO WE ARE UNABLE TO COMPLETE AN EVALUATION. IF PROVIDED WE WILL SEND A SUPPLEMENTAL REPORT WITH OUR ADDITIONAL FINDINGS. COMPLAINT RECORD ID (B)(4). DUE TO CHARACTER RESTRICTIONS IN BLOCK E1 (B)(6) HOSPITAL OF (B)(6) UNIVERSITY. Manufacturer Narrative: UPDATED FIELDS: A4, B1, B2, B4, B5, E1 (INITIAL REPORTER, EVENT SITE EMAIL), G3, G4 (PMA/510K), G6, H1, H2, H6 (HEALTH EFFECT IMPACT CODES) CORRECTED FIELDS: D4 (VERSION OR MODEL, EXP. DATE, UNIQUE IDENTIFIER), D7A, H4. DUE TO CHARACTER LIMIT IN BLOCK E1 EVENT SITE ADDRESS: NO.(B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22086006&pc=DTB | 2649622-2025-14200 | 2025-05-06 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-05-27 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion; Asystole | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PRESENTED TO THE ELECTROPHYSIOLOGY LAB FOR A DUAL CHAMBER PLACEMENT. A PACING LEAD WAS PLACED IN AN ACCEPTABLE FLUOROSCOPIC LOCATION ON THE RIGHT VENTRICLE (RV) SEPTUM AND ACCEPTABLE ELECTRICAL MEASUREMENTS WERE OBTAINED. THIS LEAD WAS SUTURED IN PLACE. THE RIGHT ATRIAL (RA) LEAD WAS INSERTED INTO THE RIGHT ATRIUM. THE PATIENT WENT ASYSTOLIC AND THEIR BLOOD PRESSURE DECREASED. PACING WAS STARTED FROM THE LEFT BUNDLE LEAD WHILE THE PATIENT WAS STABILIZED. THE PATIENT'S INTRINSIC RHYTHM RETURNED AND THE BLOOD PRESSURE STABILIZED AND THE DOCTOR RETURNED TO PLACING THE ATRIAL LEAD. THE ATRIAL LEAD WAS PLACED AND ACCEPTABLE ELECTRICAL MEASUREMENTS WERE OBTAINED. THE PHYSICIAN WAS SEWING DOWN THE ATRIAL LEAD WHEN THE PATIENT'S HEMODYNAMICS CRASHED AGAIN. THE PATIENT'S BLOOD PRESSURE DROPPED SIGNIFICANTLY. A PULSE CHECK WAS PERFORMED AND DESPITE RV PACING, THE PATIENT HAD NO PALPABLE PULSE. CHEST COMPRESSIONS BEGAN AND A FULL CODE WAS PERFORMED. PERICARDIOCENTESIS WAS PERFORMED AND BLOOD WAS WITHDRAWN FROM THE PERICARDIAL SAC. THE PATIENT WAS CODED FOR APPROXIMATELY FIFTY MINUTES WITHOUT REGAINING THE PATIENT'S PULSE OR INDEPENDENT RESPIRATION. IT WAS REPORTED THAT THE PATIENT DIED. A PACEMAKER DEVICE WAS NEVER OPENED. THE PATIENT'S POCKET WAS NEVER CLOSED.Event Description: IT WAS FURTHER REPORTED THAT THE RA LEAD PERFORATED THE ANTERIOR WALL OF THE ATRIUM DURING LEAD PLACEMENT LEADING TO PERICARDIAL EFFUSION AND HEMODYNAMIC INSTABILITY. THE PATIENT'S DEATH WAS NOT BELIEVED TO BE RELATED TO THE RV LEAD. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22086012&pc=NVN | 2649622-2025-14201 | 2025-05-06 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-05-27 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Asystole | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PRESENTED TO THE ELECTROPHYSIOLOGY LAB FOR A DUAL CHAMBER PLACEMENT. A PACING LEAD WAS PLACED IN AN ACCEPTABLE FLUOROSCOPIC LOCATION ON THE RIGHT VENTRICLE (RV) SEPTUM AND ACCEPTABLE ELECTRICAL MEASUREMENTS WERE OBTAINED. THIS LEAD WAS SUTURED IN PLACE. THE RIGHT ATRIAL (RA) LEAD WAS INSERTED INTO THE RIGHT ATRIUM. THE PATIENT WENT ASYSTOLIC AND THEIR BLOOD PRESSURE DECREASED. PACING WAS STARTED FROM THE LEFT BUNDLE LEAD WHILE THE PATIENT WAS STABILIZED. THE PATIENT'S INTRINSIC RHYTHM RETURNED AND THE BLOOD PRESSURE STABILIZED AND THE DOCTOR RETURNED TO PLACING THE ATRIAL LEAD. THE ATRIAL LEAD WAS PLACED AND ACCEPTABLE ELECTRICAL MEASUREMENTS WERE OBTAINED. THE PHYSICIAN WAS SEWING DOWN THE ATRIAL LEAD WHEN THE PATIENT'S HEMODYNAMICS CRASHED AGAIN. THE PATIENT'S BLOOD PRESSURE DROPPED SIGNIFICANTLY. A PULSE CHECK WAS PERFORMED AND DESPITE RV PACING, THE PATIENT HAD NO PALPABLE PULSE. CHEST COMPRESSIONS BEGAN AND A FULL CODE WAS PERFORMED. PERICARDIOCENTESIS WAS PERFORMED AND BLOOD WAS WITHDRAWN FROM THE PERICARDIAL SAC. THE PATIENT WAS CODED FOR APPROXIMATELY FIFTY MINUTES WITHOUT REGAINING THE PATIENT'S PULSE OR INDEPENDENT RESPIRATION. IT WAS REPORTED THAT THE PATIENT DIED. A PACEMAKER DEVICE WAS NEVER OPENED. THE PATIENT'S POCKET WAS NEVER CLOSED.Event Description: IT WAS FURTHER REPORTED THAT THE RA LEAD PERFORATED THE ANTERIOR WALL OF THE ATRIUM DURING LEAD PLACEMENT LEADING TO PERICARDIAL EFFUSION AND HEMODYNAMIC INSTABILITY. THE PATIENT'S DEATH WAS NOT BELIEVED TO BE RELATED TO THE RV LEAD. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22086651&pc=OZD | 1220648-2025-28708 | 2025-05-01 04:00 | Death | ABIOMED, INC. | 2025-05-27 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 AND EXPERIENCED BLEEDING AT THE IMPELLA ACCESS SITE. HEPARIN WAS HELD. THE NEXT DAY HEPARIN WAS STARTED AND IT WAS NOTED THAT THE PATIENT CODED TWICE OVER THE NIGHT BUT RETURN OF SPONTANEOUS CIRCULATION WAS ABLE TO BE RETURNED QUICKLY. THREE DAYS LATER, THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA AFTER TURNING THE TRANSVENOUS PACEMAKER DOWN TO 50. THE PATIENT BRIEFLY CODED AND RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED. THE PATIENT UNFORTUNATELY HAD NO BRIDGE TO LEFT VENTRICLE ASSIST DEVICE OR TRANSPLANT. THE FAMILY MADE THE DECISION FOR COMFORT CARE THREE DAYS LATER, AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: B2 SELECTION OF REQUIRED INTERVENTION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28708. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ACCESS SITE BLEEDING: AS PER THE CLINICAL INFORMATION, THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS IMPROPER ANTICOAGULATION MANAGEMENT OF THE PATIENT SINCE THE PATIENT WAS ON THE BLOOD THINNERS WITH HIGH ACTIVATED CLOTTING TIME (ACT) VALUES AT IMPLANT WHICH WAS 262. VT/VF: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22086992&pc=OZD | 1220648-2025-28712 | 2025-05-01 04:00 | Death | ABIOMED, INC. | 2025-05-27 04:00 | OZD | IMPELLA 5.5 | Positioning Failure | Arrhythmia; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT HAD AN ADDITIONAL STITCH ADDED TO THE CHEST INCISION WHILE IN ICU. IT APPEARED THAT THE STITCHES APPLIED IN THE OR WEREN¿T HOLDING THE INCISION CLOSED COMPLETELY. ADDITIONALLY, THE PATIENT WITH A HISTORY OF VENTRICULAR TACHYCARDIA (VT), EXPERIENCED VT WHILE ON SUPPORT, REQUIRING CPR, EPINEPHRINE AND SHOCK X1. THE IMPELLA WAS PULLED BACK ABOUT 0.5 CM UNDER ECHO GUIDANCE. ELECTROPHYSIOLOGIST WAS CONSULTED AND THE PLAN WAS TO INTERROGATE IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR AND PLAN FOR VT ABLATION. IT WAS FURTHER REPORTED THAT THE FAMILY DECIDED TO WITHDRAW CARE, AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA USED AS AN ASSOCIATED FACTOR TO THE EVENT OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL, VF/VT/AF/AT, AND POSITIONING ISSUES HAS BEEN COMPLETED. PRODUCT WAS RETURNED FOR REVIEW. VISUAL INSPECTION FOUND NO ISSUES ON THE PUMP OR ISSUES WITH REPOSITIONING UNIT. POSITIONING ISSUES: THE DATA LOGS SHOW POSITIONING ALARMS THROUGHOUT THE CASE. THE ROOT CAUSE OF POSITIONING ISSUE WAS UNABLE TO BE DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED AND NO ABNORMALITIES FOUND ON RETURNED PUMP. VASCULAR DAMAGE: THE ROOT CAUSE OF VASCULAR DAMAGE WAS UNABLE TO BE DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED. VT: THE ROOT CAUSE OF VT WAS UNABLE TO BE DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED. D9 DEVICE RETURNED TO MANUFACTURER DATE ADDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22087420&pc=OZD | 1220648-2025-28718 | 2025-05-05 04:00 | Death | ABIOMED, INC. | 2025-05-27 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT HAD AN IMPELLA CP PUMP PLACED VIA THE AXILLARY ARTERY FOR CARDIOGENIC SHOCK. THERE WAS AN ACCESS SITE BLEED AT THE AXILLARY THAT PROMPTED TWO UNITS OF BLOOD TO BE GIVEN TO THE PATIENT. THE NEXT DAY THE FAMILY/MEDICAL TEAM MADE THE CHOICE TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. IT WAS NOTED THAT DURING THE TIME OF BLEEDING, THE ACTIVATED CLOTTING TIME (ACT) VALUE WAS 233 WHICH IS ABOVE THE RECOMMENDED RANGE DURING THE USE OF IMPELLA. THE ROOT CAUSE OF ACCESS SITE BLEEDING WAS DETERMINED TO BE ANTICOAGULATION MANAGEMENT BECAUSE THE ACT WERE NOT MAINTAINED IN THE RECOMMENDED THERAPEUTIC RANGE DURING THE TIME OF BLEEDING. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22087516&pc=DSQ | 2916596-2025-03311 | 2025-05-12 04:00 | Death | THORATEC CORPORATION | 2025-05-27 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Tachycardia | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE VENTRICULAR ASSIST DEVICE (VAD) WAS TURNED OFF AS PART OF WITHDRAW OF CARE. THE PUMP OPERATED AS INTENDED. THE CAUSE OF DEATH WAS VENTRICULAR TACHYCARDIA AND WAS NOT DEVICE RELATED. THE PATIENT HAD BEEN ADMITTED FOR UNSUCCESSFUL IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR (ICD) SHOCKS REQUIRING EXTERNAL DEFIBRILLATION. THE ICD WAS IMPLANTED PRIOR TO THE VAD. AMIODARONE AND LIDOCAINE DRIPS WERE INITIATED. THE PATIENT DECIDED TO TRANSITION TO COMFORT CARE/DO NOT RESUSCITATE (DNR). THE ICD WAS TURNED OFF, AND THE DRIPS AND VAD WERE TURNED OFF. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE LEFT VENTRICULAR ASSIST DEVICE (LVAD), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS ADVERSE EVENTS, INCLUDING ARRHYTHMIA AND DEATH, WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS. THE IFU ALSO LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22087735&pc=NIQ | 2024168-2025-05365 | 2024-08-02 04:00 | Death | ABBOTT VASCULAR | 2025-05-27 04:00 | NIQ | XIENCE PRIME¿ | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P110019 |  | 1 | Event Description: PATIENT ID: (B)(6). IT WAS REPORTED THE INITIAL PROCEDURE WAS PERFORMED ON (B)(6) 2024. A 3.0X18MM XIENCE PRIME STENT WAS IMPLANTED IN THE PROXIMAL LEFT TIBIOPERONEAL TRUNK ARTERY WITHOUT ISSUE. THE PATIENT WAS REPORTED TO HAVE DIED ON (B)(6) 2024. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE XIENCE PRIME, XIENCE PRIME SMALL VESSEL (SV), AND XIENCE PRIME LONG LENGTH (LL), EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE (IFU) AS A KNOWN PATIENT EFFECT OF CORONARY STENTING PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22087940&pc=DQY | 2124215-2025-32414 | 2019-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-27 04:00 | DQY | TRAPPER EXCHANGE DEVICE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: BOSTON SCIENTIFIC PERFORMED A RETROSPECTIVE DATA ANALYSIS ON TRAPPER USING THE PINC AI HEALTHCARE DATABASE FROM PREMIER. PATIENTS ARE IDENTIFIED THROUGH USE OF TRAPPER AS IDENTIFIED IN CHARGE MASTER BILLING. ADVERSE EVENTS WERE IDENTIFIED THROUGH THE RESPECTIVE CPT/ICD-10 CODING. THE PROCEDURES WERE PERFORMED FROM (B)(6) 2019 TO (B)(6) 2024. THESE ANALYSES ARE BEING PERFORMED AS A SPECIFIC ACTIVITY TO ASSESS SAFETY AND PERFORMANCE RATES ASSOCIATED WITH THE SUBJECT DEVICE. TRAPPER OUTCOMES WERE ASSESSED AGAINST TRAPIT AND RUNTHROUGH. RATES OF THE ADVERSE EVENTS INCLUDING DEATH, MYOCARDIAL INFARCTION, PERICARDIAL EFFUSION, CARDIAC TAMPONADE, ACUTE RENAL FAILURE, CONTRAST ALLERGY, BLEEDING, EMBOLISM AND STROKE ARE REPORTED. A TOTAL OF 1879 PATIENTS UNDERWENT A PROCEDURE USING TRAPPER. THE FOLLOWING IS A SUMMARY OF THOSE RESULTS OVER 5 YEARS (JANUARY 2019 - MARCH 2024): MORTALITY RATES TRAPPER CASES: 59 OUT OF 1879 PATIENTS RESULTING IN A RATE OF 3.14%, COMPARED TO THE COMPETITOR RATE OF 2.41%-3.06%. MORTALITY RATES FOR TRAPPER CASES ARE HIGHER THAN THAT OF COMPETITORS. THE 95% CONFIDENCE INTERVAL FOR TRAPPER CASES OVERLAPS WITH THE COMPETITORS 95% CONFIDENCE INTERVAL. RATES FOR MORTALITY ARE THEREFORE COMPARABLE TO THOSE OF THE COMPETITIVE PRODUCTS AND ARE THEREFORE ACCEPTABLE. Manufacturer Narrative: B2 DATE OF DEATH AND B3 DATE OF EVENT: DATA WAS PROVIDED FOR EVENTS RECORDED 01JAN2019-31MAR2024. 01JAN2019 SELECTED AS THE EARLIEST POSSIBLE DATE OF EVENT AND DATE OF DEATH. DATA OBTAINED FOR THIS STUDY IS DE-IDENTIFIED BEFORE BEING ACCESSED BY BOSTON SCIENTIFIC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO OUR ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE STUDY DATA DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC AND NO FURTHER INFORMATION IS AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22088763&pc=DSQ | 2916596-2025-03486 | 2025-04-01 04:00 | Death | THORATEC CORPORATION | 2025-05-27 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. IT WAS UNKNOWN IF THE DEATH WAS DEVICE RELATED. THE SITE BELIEVED THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: SECTIONS B2 AND B3: DATE OF DEATH AND DATE OF EVENT WERE ESTIMATED. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE ACCOUNT ULTIMATELY COMMUNICATED THAT NO FURTHER INFORMATION REGARDING THE PATIENT OUTCOME, PATIENT DETAILS, OR PUMP SERIAL NUMBER WILL BE PROVIDED. THE PUMP WILL NOT BE RETURNING FOR EVALUATION. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22088774&pc=NPT | 2015691-2025-04130 | 2025-05-06 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-27 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY OUR CANADIAN AFFILIATES, DURING IMPLANT OF A 29MM SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE THE VALVE WAS DEPLOYED AT 31CC VOLUME AS PLANNED, HOWEVER, AFTER DEPLOYMENT, THE VALVE APPEARED TO BE CONSTRAINED BY THE CALCIUM, TO WHICH THE PHYSICIAN WANTED TO ADD AN EXTRA 2CC TO POST-DILATE THE VALVE. THE CARDIAC SURGEON SUGGESTED THAT OVERSIZING WITH NOMINAL VOLUME WOULD BE TOO AGGRESSIVE GIVEN THE OVERALL FRAILTY OF THE PATIENT AND AGGRESSIVE ROOT FEATURES. THE PHYSICIAN PROCEEDED TO ADD 2CC'S OF VOLUME AND POST-DILATED THE VALVE AT NOMINAL VOLUME. THE PATIENT IMMEDIATELY BECAME PROFOUNDLY HYPOTENSIVE AND UNRESPONSIVE. THE ROOT ANGIOGRAM SHOWED EXTRAVASATION NEAR THE LEFT CORONARY CUSP (LCC), SUGGESTING AN AORTIC ROOT RUPTURE. THE PATIENT WAS NOT A RESUSCITATION CANDIDATE, NO LIFE-SUPPORTING MEASURES WERE ATTEMPTED, AND THE PATIENT PASSED AWAY. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR AN UPDATE TO THE INVESTIGATION. THE FOLLOWING SECTIONS OF THIS REPORT HAS BEEN UPDATED: UPDATE TO B4, G3, G6, H2, H6, H11. THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PATIENT CONDITION (HIGH-RISK ANATOMICAL PROFILE) AND PROCEDURAL FACTORS (POST-DILATATION) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE AN AORTIC ROOT RUPTURE DUE TO EXCESSIVE SIZING IN CONJUNCTION WITH HIGH RISK ROOT FEATURES COULD HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22088788&pc=DTZ | 8010762-2025-0000229 | 2025-05-20 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-05-27 04:00 | DTZ | TUBING SET | Pressure Problem | Insufficient Information | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN USA DURING TREATMENT. IT WAS REPORTED THAT SUPPORT WAS INITIATED, AND THE VENOUS PRESSURE DID NOT READ. THE CUSTOMER STATED THAT THE VENOUS PRESSURE WAS DISPLAYED DURING SET UP AND PRIMING. NEW INFORMATION WAS RECEIVED ON B)(6) 2025 THAT THE PATIENT PASSED AWAY AND THE HLS SET WAS DISCARDED BY THE CUSTOMER. AS THERE WAS A PRESSURE READING ISSUE DURING TREATMENT, WHICH CAN LEAD TO A PUMP STOP, IF THE INTERVENTION WAS SET BY THE USER AND THE PATIENT PASSED AWAY, A REPORT IS REQUIRED. COMPLAINT ID # (B)(4).Event Description: COMPLAINT ID# (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: THE EVENT OCCURRED IN USA DURING TREATMENT. IT WAS REPORTED THAT SUPPORT WAS INITIATED, AND THE VENOUS PRESSURE DID NOT READ. THE CUSTOMER STATED THAT THE VENOUS PRESSURE WAS DISPLAYED DURING SET UP AND PRIMING. NEW INFORMATION WAS RECEIVED ON 2025-05-23 THAT THE PATIENT PASSED AWAY. FURTHER THEY CONFIRMED THE HLS SET TO BE DISCARDED.THE PATIENT PASSED AWAY. AS THERE WAS A PRESSURE READING ISSUE DURING TREATMENT, WHICH CAN LEAD TO A PUMP STOP, IF THE INTERVENTION WAS SET BY THE USER, A REPORT IS REQUIRED. THE EXACT ROOT CAUSE REMAINS UNKNOWN, AS THE PRODUCT WAS DISCARDED AND THE REQUESTED INFORMATION WAS NOT MADE AVAILABLE BY THE CUSTOMER. HOWEVER, A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-07-21 WITH FOLLOWING CONCLUSION: "ACCORDING TO THE COMPLAINT NARRATIVE, THE INCIDENT IN SCOPE INVOLVED THE VENOUS PRESSURE NOT READING AFTER INITIATION OF ARTERIOVENOUS SUPPORT. THE CUSTOMER REPORTED THAT THE VENOUS PRESSURE VALUE WAS PRESENT AND READING APPROPRIATELY DURING SETUP AND PRIMING. NO ADDITIONAL ALARMS OR DEVICE MALFUNCTIONS WERE MENTIONED. THE PRODUCT WAS DISCARDED FOLLOWING THE EVENT AND WAS NOT MADE AVAILABLE FOR RETURN OR EVALUATION. DUE TO THE ABSENCE OF RETURNED PRODUCT AND THE LIMITED INFORMATION PROVIDED BY THE CUSTOMER, THE ROOT CAUSE OF THE REPORTED ISSUE, CANNOT BE DEFINITIVELY DETERMINED. HOWEVER, A STRUCTURED ANALYSIS OF POTENTIAL CONTRIBUTING FACTORS, BROADLY CATEGORIZED INTO EXTERNAL FACTORS (RELATED TO SETUP, CLINICAL USE, OR THE ENVIRONMENT) AND INTERNAL FACTORS (RELATED TO THE DEVICE OR SYSTEM COMPONENTS THEMSELVES) MAY BE PROPOSED. EXTERNAL FACTORS REFER TO CONDITIONS OR VARIABLES INTRODUCED DURING CLINICAL SETUP OR USE THAT ARE NOT INHERENT TO THE DEVICE. IN THIS CASE, IT IS PLAUSIBLE THAT THE PRESSURE READING MAY HAVE BEEN LOST DUE TO MECHANICAL STRESS APPLIED TO THE HLS SET DURING HANDLING OR TRANSPORT, WHICH COULD LEAD TO INTERNAL DISRUPTION OF SIGNAL TRANSMISSION WITHIN THE SENSOR CABLE ASSEMBLY. SIMILARLY, FLUID INGRESS INTO THE SENSOR CABLE, DURING OR AFTER PRIMING, MAY DISRUPT ELECTRICAL CONTINUITY AND MAY CAUSE LOSS OF PRESSURE SIGNAL. INTERNAL DEVICE-RELATED FACTORS INVOLVE POTENTIAL ISSUES WITHIN THE SENSOR OR SYSTEM INTERFACES. ONE THEORETICAL CAUSE COULD BE AN ISOLATED MALFUNCTION OR CALIBRATION DRIFT OF THE VENOUS PRESSURE SENSOR AFFECTING ONLY THAT SPECIFIC PVEN SENSOR. HOWEVER, IN THE ABSENCE OF DEVICE RETURN OR FURTHER INFORMATION ON THE SEQUENCE OF EVENTS, THIS REMAINS SPECULATIVE. ANOTHER POSSIBLE CAUSE WOULD BE A DISRUPTION IN THE ELECTRICAL CONNECTION BETWEEN THE SENSOR AND THE CONSOLE, SUCH AS A DISCONNECTION OR BREAK IN THE SENSOR CABLE. THIS EXPLANATION IS CONSIDERED UNLIKELY IN THIS CASE BECAUSE THE VENOUS PRESSURE SENSOR IS PART OF A BUNDLED SENSOR CABLE CONNECTED TO THE OXYGENATOR THAT ALSO TRANSMITS OTHER PRESSURE READINGS, SUCH AS PINT AND PART PRESSURE, AS WELL AS ARTERIAL TEMPERATURE DATA. IF THE CABLE HAD BEEN PHYSICALLY DISCONNECTED OR DAMAGED, IT WOULD BE EXPECTED THAT ALL SENSOR READINGS WOULD BE LOST SIMULTANEOUSLY, NOT JUST PVEN. SINCE NO LOSS OF ADDITIONAL PARAMETERS WAS REPORTED, A GLOBAL CABLE FAILURE APPEARS INCONSISTENT WITH THE EVENT DESCRIPTION. LIKEWISE, A SOFTWARE OR SIGNAL ACQUISITION FAULT WITHIN THE ECMO CONSOLE MAY (GENERALLY) AFFECT MULTIPLE CHANNELS AND TRIGGER A SYSTEM ALERT, NEITHER OF WHICH WAS NOTED BY THE CUSTOMER. IN CONCLUSION, WHILE A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITHOUT FURTHER INFORMATION OR PRODUCT EVALUATION, THE CIRCUMSTANCES OF THE EVENT ARE MORE CONSISTENT WITH EXTERNAL, USER RELATED FACTORS THAT IMPACTED THE CIRCUIT AFTER PRIMING. INTERNAL SYSTEM OR COMPONENT FAILURE IS CONSIDERED LESS LIKELY GIVEN THE ISOLATED NATURE OF THE REPORTED SIGNAL LOSS. AS SUCH, THE ISSUE REMAINS NON-REPRODUCIBLE AND INCONCLUSIVE. REGARDING PATIENT OUTCOME, THE CUSTOMER REPORTED THAT THE PATIENT EXPIRED. NO ADDITIONAL INFORMATION ON PATIENT STATUS, COMORBIDITIES, OR ANY POTENTIAL RELATIONSHIP BETWEEN THIS INCIDENT AND THE PATIENT OUTCOME WAS PROVIDED. THEREFORE, NO CAUSAL ASSOCIATION BETWEEN THE FAILURE OF VENOUS PRESSURE MONITORING AND THE PATIENT¿S EXPIRATION CAN BE ESTABLISHED." ACCORDING TO THE INSTRUCTION FOR USE (HLS SET, CHAPTER "PREPARATION AND INSTALLATION") THE PRESSURE SENSORS HAVE TO BE CHECKED BEFORE PRIMING. IN GENERAL, IT SHOULD BE NOTED THAT PRE-PRIMING CAN HAVE A NEGATIVE IMPACT ON THE SENSOR FUNCTION AND THE ASSOCIATED PRESSURE VALUES AND CAN LEAD TO THE OUTPUT OF UNEXPECTED / IMPLAUSIBLE PRESSURE VALUES SHORTLY BEFORE USE, AS INDICATED IN CHAPTER "PRIMING THE SYSTEM", WHERE IT SAYS "ONLY FILL THE SYSTEM SHORTLY BEFORE USE AND IN ACCORDANCE WITH THE PRIMING INSTRUCTIONS. CALIBRATION VALUES ARE LOST WHEN THE DRIVE UNIT IS SWITCHED OFF". FURTHER THE CARDIOHELP HAS A FLOW/BUBBLE SENSOR AND A VENOUS PROBE TO MEASURE AND CONTROL THE BLOOD FLOW AND PARAMETERS. IF THE MEASURED VALUES ARE ABOVE HIGH LIMIT OR BELOW LOW LIMIT OF THE SET LIMITS THE SYSTEM GENERATES A VISUAL AND ACOUSTICAL ALARM. THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON 2025-07-21. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARDS TO THE REPORTED FAILURE WERE FOUND. BASED ON THE RESULTS AND THE PROVIDED PHOTOGRAPHICAL EVIDENCE PROVIDED BY THE CUSTOMER, THE REPORTED FAILURE "PRESSURE READING ISSUE" COULD BE CONFIRMED. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22088931&pc=NGV | 2135147-2025-02794 | 2014-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-05-27 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200049 |  | 1 | Event Description: THE ARTICLE, "LEFT ATRIAL APPENDAGE OCCLUSION IN VULNERABLE OLDEST-OLD PATIENTS ¿ EFFECTIVENESS AND SAFETY IN A LARGE SINGLE-CENTER REGISTRY", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO INVESTIGATE THE EFFECTIVENESS AND SAFETY OF A SIMPLIFIED NON-COMPLEX LEFT ATRIAL APPENDAGE OCCLUSION (LAAO) PROCEDURE. DEVICES INCLUDED IN THE STUDY WERE AMPLATZER AMULET, WATCHMAN, WATCHMAN FLX, AND LAMBRE. THE ARTICLE CONCLUDED THAT NON-COMPLEX LAAO PROCEDURE ACHIEVED A STATE-OF-THE-ART TECHNICAL SUCCESS RATE AND A LOW COMPLICATION RATE IN VULNERABLE OLDEST-OLD PATIENTS. CARE SHOULD BE TAKEN TO IDENTIFY PATIENTS WITH SUFFICIENT LIFE EXPECTANCY TO BENEFIT FROM LAAO. THE TREND TOWARDS MORE ISCHEMIC STROKES DURING FOLLOW-UP IS AN IMPORTANT SIGNAL FOCUSING ON CARDIOEMBOLIC STROKE ETIOLOGIES BESIDES LAA THROMBI IN THESE PATIENTS. [THE PRIMARY AND CORRESPONDING AUTHOR WAS CHRISTIAN FASTNER, DEPARTMENT OF CARDIOLOGY, HAEMOSTASEOLOGY AND MEDICAL INTENSIVE CARE, UNIVERSITY MEDICAL CENTRE MANNHEIM, MEDICAL FACULTY MANNHEIM, HEIDELBERG UNIVERSITY, THEODOR-KUTZER-UFER 1-3, 68167, MANNHEIM, GERMANY, WITH CORRESPONDING EMAIL: CHRISTIAN.FASTNER@UMM.DE] THE TIME FRAME OF THE STUDY BEGAN IN 2014 WITH NO SPECIFIED END DATE. A TOTAL OF 230 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 132 (58.4%) RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 79.0 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED ARTERIAL HYPERTENSION, DIABETES MELLITUS, CORONARY HEART DISEASE (PRIOR BYPASS SURGERY, MYOCARDIAL INFARCTION), PERIPHERAL ARTERIAL DISEASE, TRANSIENT ISCHEMIC ATTACH, ISCHEMIC/HEMORRHAGIC STROKE, CHRONIC KIDNEY DISEASE, CHRONIC LIVER DISEASE, ALCOHOL ABUSE, PRIOR BLEEDING, PRIOR PULMONARY VEIN ISOLATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF LEFT ATRIAL APPENDAGE OCCLUSION IN VULNERABLE OLDEST-OLD PATIENTS ¿ EFFECTIVENESS AND SAFETY IN A LARGE SINGLE-CENTER REGISTRY WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING ARTERIAL HYPERTENSION, DIABETES MELLITUS, CORONARY HEART DISEASE (PRIOR BYPASS SURGERY, MYOCARDIAL INFARCTION), PERIPHERAL ARTERIAL DISEASE, TRANSIENT ISCHEMIC ATTACH, ISCHEMIC/HEMORRHAGIC STROKE, CHRONIC KIDNEY DISEASE, CHRONIC LIVER DISEASE, ALCOHOL ABUSE, PRIOR BLEEDING, PRIOR PULMONARY VEIN ISOLATION. SOME OF THE COMPLICATIONS REPORTED WERE STROKE, PERICARDIAL EFFUSION, CARDIAC TAMPONADE, EMBOLISM, RENAL FAILURE, MYOCARDIAL INFARCTION, TRANSIENT ISCHEMIC ATTACK AND DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: LEFT ATRIAL APPENDAGE OCCLUSION IN VULNERABLE OLDEST-OLD PATIENTS ¿ EFFECTIVENESS AND SAFETY IN A LARGE SINGLE-CENTER REGISTRY. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: LEFT ATRIAL APPENDAGE OCCLUSION IN VULNERABLE OLDEST-OLD PATIENTS ¿ EFFECTIVENESS AND SAFETY IN A LARGE SINGLE-CENTER REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22088946&pc=HHW | 1216677-2025-00030 | 2025-04-03 04:00 | Death | COOPERSURGICAL, INC. | 2025-05-27 04:00 | HHW | MILEX | Use of Device Problem | Rectovaginal Fistula | K904774 |  | 1 | Event Description: IT WAS REPORTED THAT THE GELHORN PESSARY CREATED A HOLE BETWEEN THE PATIENT'S VAGINA AND THEIR RECTUM WHICH ALLOWED STOOL TO PASS THROUGH AND OUT OF THE VAGINAL OPENING. PATIENT IS SCHEDULED FOR SURGERY TO CORRECT THE FISTULA. MXPGE2-14 MILEX 2025-05-0000391.Event Description: IT WAS REPORTED THAT THE GELHORN PESSARY CREATED A HOLE BETWEEN THE PATIENT'S VAGINA AND THEIR RECTUM WHICH ALLOWED STOOL TO PASS THROUGH AND OUT OF THE VAGINAL OPENING. PATIENT IS SCHEDULED FOR SURGERY TO CORRECT THE FISTULA. UPDATE: JUNE 15, 2025: THE DANISH MEDICINE AGENCY HAS BEEN INFORMED FROM THE HEALTHCARE FACILITY THAT THE PATIENT INVOLVED IN THE INCIDENT HAS PASSED AWAY.Event Description: NO ADDITIONAL INFORMATION. Manufacturer Narrative: G2: DENMARK DEVICE LOCATION IS UNKNOWN. ONCE THE INVESTIGATION HAS BEEN COMPLETED, A FOLLOW-UP MDR WILL BE SUBMITTED. Manufacturer Narrative: UPDATED: B4, G3, G6, H2, H3, H6, H11. DISTRIBUTION HISTORY: THE LOT NUMBER WAS NOT PROVIDED, SO A MANUFACTURE DATE, DHR NUMBER AND SHIP DATE ARE NOT AVAILABLE. MANUFACTURING RECORD REVIEW: A REVIEW OF THE DEVICE HISTORY RECORD COULD NOT BE PERFORMED BECAUSE THE LOT NUMBER WAS NOT PROVIDED. HOWEVER, IT SHOULD BE NOTED AT THE TIME OF MANUFACTURE, RECORDS FROM EACH LOT ARE THOROUGHLY REVIEWED TO ENSURE THAT PRODUCTS ARE RELEASED MEETING ALL COOPERSURGICAL QUALITY RELEASE SPECIFICATIONS. A LIST OF ALL MXPGE2-1/4 UNITS THAT WERE SHIPPED TO DENMARK WAS ASSEMBLED. ALL DHR'S WERE REVIEWED AND NO NON-CONFORMITIES, RELATED TO THE COMPLAINT CONDITION, WERE NOTED. THE CORRECT MATERIAL WAS USED, THE METHOD SHEETS WERE FOLLOWED AND THE OPERATORS WERE TRAINED TO THEM. PARTS WERE INSPECTED AND FOUND TO BE ACCEPTABLE. INCOMING INSPECTION REVIEW: THE FOLLOWING LOT NUMBERS FOR PART # MIL0373 WERE INSPECTED AS THESE WERE USED IN WORKORDERS THAT WERE SHIPPED TO DENMARK: (B)(4). NO NON-CONFORMANCES WERE DETECTED IN THESE LOT NUMBER. HISTORICAL COMPLAINT REVIEW: A REVIEW OF THE 2-YEAR COMPLAINT HISTORY SHOWED NO SIMILAR REPORTED COMPLAINT CONDITIONS. THIS IS AN ISOLATED INCIDENT. PRODUCT RECEIPT: THE COMPLAINT PRODUCT HAS NOT BEEN RETURNED TO COOPERSURGICAL. VISUAL EVALUATION: EVALUATION OF THE COMPLAINT PRODUCT COULD NOT BE COMPLETED AS THE COMPLAINT PRODUCT HAS NOT BEEN RETURNED TO COOPERSURGICAL. IF THE PRODUCT SHOULD BE RETURNED AT A LATER DATE, IT WILL BE EVALUATED, AND ANY FINDINGS WILL BE APPENDED TO THIS INVESTIGATION. FUNCTIONAL EVALUATION: EVALUATION OF THE COMPLAINT PRODUCT COULD NOT BE COMPLETED AS THE COMPLAINT PRODUCT HAS NOT BEEN RETURNED. IF THE PRODUCT SHOULD BE RETURNED AT A LATER DATE, IT WILL BE EVALUATED, AND ANY FINDINGS WILL BE APPENDED TO THIS INVESTIGATION. ANY RELEVANT OR CLINICAL INFORMATION: PRODUCT SURVEILLANCE MADE THREE GOOD FAITH ATTEMPTS WITH THE CUSTOMER TO TRY AND OBTAIN FURTHER CLINICAL DETAILS SURROUNDING THE EVENT. ROOT CAUSE: THE MILEX PESSARY COMPLIES WITH THE NEW REGULATION (EU) 2017/745 OF THE EUROPEAN PARLIAMENT AND THE COUNCIL ON MEDICAL DEVICES THEREFORE, THERE IS NO INDICATION OF ANY BIOCOMPATIBILITY ISSUES THAT COULD HAVE CAUSED A LESION WITHIN THE PATIENT AS INDICATED IN COMPLAINT. PRODUCT LABELING INCLUDES INFORMATION INDICATING THE IMPORTANCE OF PERIODIC REMOVAL OF THE PESSARY BY THE HEALTHCARE PROVIDER FOR INSPECTION OF THE PESSARY AND THE VAGINA TO REDUCE THE LIKELIHOOD OF CERTAIN HARMS. HEALTHCARE PROVIDERS ARE ALSO INSTRUCTED TO COMMUNICATE THE IMPORTANCE OF THIS TO THEIR PATIENTS. THE SIZE IS INTENDED TO BE SELECTED ACCORDING TO THE FEEL AND COMFORT OF THE PATIENT AND DURING CONSULTATION WITH THE PHYSICIAN AS INDICATED IN THE IFU. THIS SHAPE IS INDICATED FOR PATIENTS WITH THIRD DEGREE PROLAPSE OF PROCIDENTIA. THE CONTROLS TO MINIMIZE THE RISK HAVE ALREADY BEEN PUT IN PLACE AND THE REMAINING RISK IS DEEMED ACCEPTABLE. AMONG THESE CONTROLS WE HAVE THE IFU INDICATING: FREQUENT VISITS TO THE PHYSICIAN TO REVIEW THE PATIENT'S CONDITION. THE PESSARY IS NOT MEANT TO LEAVE FOR LONG PERIODS OF TIME. A LACERATION LIKE THE ONE DESCRIBED IN THE COMPLAINT IS HIGHLY LIKELY TO HAVE CAUSED SOME KIND OF BLEEDING. IFU INDICATES THAT IF THERE IS BLEEDING, THE PATIENT NEEDS TO CONSULT WITH A PHYSICIAN. ROOT CAUSE IS NOT APPLICABLE AS THE COMPLAINT CONDITION WAS NOT CONFIRMED. COOPERSURGICAL WILL CONTINUE TO MONITOR THIS COMPLAINT CONDITION FOR TRENDS. Manufacturer Narrative: UPDATED: B4, B2, G3, H1, H6. SUPPLEMENTAL SUBMISSION IS BEING MADE DUE TO RECEIVING AN UPDATE ON THE PATIENT EXPIRING. INVESTIGATION IS ONGOING. FOLLOW-UP MDR WILL BE FILED UPON COMPLETION OF INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22088996&pc=BZD | 2518422-2025-106070 | 2024-08-30 04:00 | Death | RESPIRONICS, INC. | 2025-05-27 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR PRO C-FLEX+ DEVICE. THERE WAS A REPORT OF PATIENT DEATH. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22089032&pc=DSQ | 2916596-2025-03408 | 2025-04-30 04:00 | Death | THORATEC CORPORATION | 2025-05-27 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Chronic Obstructive Pulmonary Disease (COPD) | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO PULMONARY DISEASE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION.Event Description: IT WAS LATER REPORTED THAT THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE WAS ALSO SAID TO HAVE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU REV. D, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22089263&pc=NPT | 2015691-2025-04140 | 2025-05-09 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-27 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Migration or Expulsion of Device | Cardiac Perforation | P140031 |  | 1 | Event Description: AS REPORTED BY OUR GERMAN AFFILIATES, DURING IMPLANT OF A 29MM SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IMPLANT IN THE AORTIC POSITION VIA TRANSFEMORAL APPROACH IN A PATIENT WITH A PRE-EXISTING NON-EDWARDS BIOPROSTHESIS, THE VALVE WAS IMPLANTED WITH AN ADDITIONAL 3 ML OF VOLUME. DURING THE IMPLANTATION PHASE, A 2:1 CAPTURE OCCURRED. AT THIS POINT, THE BALLOON/VALVE WAS ALREADY PARTIALLY OPENED, NECESSITATING THE COMPLETION OF THE IMPLANT. HOWEVER, DUE TO INSUFFICIENT CAPTURE AND SHAKING OF THE VALVE DURING POSITIONING, THE VALVE WAS IMPLANTED TOO DEEP INTO THE LEFT VENTRICULAR OUTFLOW TRACT (LVOT). THE NC SIDE OF THE VALVE HAD ITS UPPER FRAME INSIDE THE ANNULUS REGION, WHEREAS THE LC SIDE WAS ALREADY POSITIONED IN THE LVOT. A SECOND VALVE WAS PREPARED, BUT DURING THE IMPLANTATION OF THIS VALVE, THEY WERE UNABLE TO FIXATE THE FORMER VALVE AND SEAL AT THE ANNULUS LEVEL. IN THE TEE, IT WAS VISIBLE THAT THE VALVES MIGRATED STEPWISE FURTHER INTO THE VENTRICLE. ADDITIONALLY, A PERFORATION OF THE ANTERIOR MITRAL LEAFLET (AML) BASE WAS VISIBLE. THE PATIENT UNDERWENT OPEN-HEART SURGERY. THERE WAS NO MALFUNCTION OF THE EDWARDS DEVICES. THE PATIENT PASSED AWAY DURING THE SURGICAL INTERVENTION. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR AN ADDITION TO THE INVESTIGATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED OR CORRECTED: B4, G3, G6, H2, H6, H11. THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), VALVE MIGRATION REQUIRING INTERVENTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH TRANSCATHETER VALVE REPLACEMENT (TVR). ACCORDING TO THE LITERATURE REVIEW, VALVE MIGRATION RESULTS WHEN FORCES ACTING ON THE TRANSCATHETER HEART VALVE (THV) OVERCOME THE STRENGTH OF ATTACHMENT OF THE VALVE TO THE ANNULUS. STENT VALVES ARE SUBJECTED TO ANTEGRADE EJECTION FORCES DURING SYSTOLE. LESS-THAN-SEVERE AND NON-UNIFORMLY DISTRIBUTED CALCIFICATION OF THE LEAFLETS, INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION CAN CONTRIBUTE TO VALVE MIGRATION. ADDITIONALLY, RESIDUAL OVERHANGING LEAFLETS CAN EXERT DOWNWARDS FORCE DURING DIASTOLE, CAUSING MIGRATION OF THE THV TOWARDS THE VENTRICLE. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR MIGRATION (I.E., MINIMAL LEAFLET CALCIFICATION), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE PROCEDURAL FACTORS (PREVIOUS VALVE IMPLANTED TOO LOW) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: THIS IS TWO OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), VALVE MIGRATION REQUIRING INTERVENTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH TRANSCATHETER VALVE REPLACEMENT (TVR). ACCORDING TO THE LITERATURE REVIEW, VALVE MIGRATION RESULTS WHEN FORCES ACTING ON THE TRANSCATHETER HEART VALVE (THV) OVERCOME THE STRENGTH OF ATTACHMENT OF THE VALVE TO THE ANNULUS. STENT VALVES ARE SUBJECTED TO ANTEGRADE EJECTION FORCES DURING SYSTOLE. LESS-THAN-SEVERE AND NON-UNIFORMLY DISTRIBUTED CALCIFICATION OF THE LEAFLETS, INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION CAN CONTRIBUTE TO VALVE MIGRATION. ADDITIONALLY, RESIDUAL OVERHANGING LEAFLETS CAN EXERT DOWNWARDS FORCE DURING DIASTOLE, CAUSING MIGRATION OF THE THV TOWARDS THE VENTRICLE. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR MIGRATION (I.E., MINIMAL LEAFLET CALCIFICATION), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. DUE TO A LIMITED AMOUNT OF INFORMATION RECEIVED, IT IS UNKNOWN WHAT COULD HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22089312&pc=DTZ | 3013876692-2025-00041 | 2025-05-20 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-05-27 04:00 | DTZ | TUBING SET | Pressure Problem | Insufficient Information | K112360 |  | 1 | Event Description: THE EVENT OCCURRED IN USA DURING TREATMENT. IT WAS REPORTED THAT SUPPORT WAS INITIATED, AND THE VENOUS PRESSURE DID NOT READ. THE CUSTOMER STATED THAT THE VENOUS PRESSURE WAS DISPLAYED DURING SET UP AND PRIMING. NEW INFORMATION WAS RECEIVED ON 2025-05-23 THAT THE PATIENT PASSED AWAY AND THE HLS SET WAS DISCARDED BY THE CUSTOMER. AS THERE WAS A PRESSURE READING ISSUE DURING TREATMENT, WHICH CAN LEAD TO A PUMP STOP, IF THE INTERVENTION WAS SET BY THE USER AND THE PATIENT PASSED AWAY, A REPORT IS REQUIRED. COMPLAINT ID #: (B)(4). Event Description: COMPLAINT ID# (B)(4). Manufacturer Narrative: THE EVENT OCCURRED IN USA DURING TREATMENT. IT WAS REPORTED THAT SUPPORT WAS INITIATED, AND THE VENOUS PRESSURE DID NOT READ. THE CUSTOMER STATED THAT THE VENOUS PRESSURE WAS DISPLAYED DURING SET UP AND PRIMING. NEW INFORMATION WAS RECEIVED ON 2025-05-23 THAT THE PATIENT PASSED AWAY. FURTHER THEY CONFIRMED THE HLS SET TO BE DISCARDED. THE PATIENT PASSED AWAY. AS THERE WAS A PRESSURE READING ISSUE DURING TREATMENT, WHICH CAN LEAD TO A PUMP STOP, IF THE INTERVENTION WAS SET BY THE USER, A REPORT IS REQUIRED. THE EXACT ROOT CAUSE REMAINS UNKNOWN, AS THE PRODUCT WAS DISCARDED AND THE REQUESTED INFORMATION WAS NOT MADE AVAILABLE BY THE CUSTOMER. HOWEVER, A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-07-21 WITH FOLLOWING CONCLUSION: "ACCORDING TO THE COMPLAINT NARRATIVE, THE INCIDENT IN SCOPE INVOLVED THE VENOUS PRESSURE NOT READING AFTER INITIATION OF ARTERIOVENOUS SUPPORT. THE CUSTOMER REPORTED THAT THE VENOUS PRESSURE VALUE WAS PRESENT AND READING APPROPRIATELY DURING SETUP AND PRIMING. NO ADDITIONAL ALARMS OR DEVICE MALFUNCTIONS WERE MENTIONED. THE PRODUCT WAS DISCARDED FOLLOWING THE EVENT AND WAS NOT MADE AVAILABLE FOR RETURN OR EVALUATION. DUE TO THE ABSENCE OF RETURNED PRODUCT AND THE LIMITED INFORMATION PROVIDED BY THE CUSTOMER, THE ROOT CAUSE OF THE REPORTED ISSUE, CANNOT BE DEFINITIVELY DETERMINED. HOWEVER, A STRUCTURED ANALYSIS OF POTENTIAL CONTRIBUTING FACTORS, BROADLY CATEGORIZED INTO EXTERNAL FACTORS (RELATED TO SETUP, CLINICAL USE, OR THE ENVIRONMENT) AND INTERNAL FACTORS (RELATED TO THE DEVICE OR SYSTEM COMPONENTS THEMSELVES) MAY BE PROPOSED. EXTERNAL FACTORS REFER TO CONDITIONS OR VARIABLES INTRODUCED DURING CLINICAL SETUP OR USE THAT ARE NOT INHERENT TO THE DEVICE. IN THIS CASE, IT IS PLAUSIBLE THAT THE PRESSURE READING MAY HAVE BEEN LOST DUE TO MECHANICAL STRESS APPLIED TO THE HLS SET DURING HANDLING OR TRANSPORT, WHICH COULD LEAD TO INTERNAL DISRUPTION OF SIGNAL TRANSMISSION WITHIN THE SENSOR CABLE ASSEMBLY. SIMILARLY, FLUID INGRESS INTO THE SENSOR CABLE, DURING OR AFTER PRIMING, MAY DISRUPT ELECTRICAL CONTINUITY AND MAY CAUSE LOSS OF PRESSURE SIGNAL. INTERNAL DEVICE-RELATED FACTORS INVOLVE POTENTIAL ISSUES WITHIN THE SENSOR OR SYSTEM INTERFACES. ONE THEORETICAL CAUSE COULD BE AN ISOLATED MALFUNCTION OR CALIBRATION DRIFT OF THE VENOUS PRESSURE SENSOR AFFECTING ONLY THAT SPECIFIC PVEN SENSOR. HOWEVER, IN THE ABSENCE OF DEVICE RETURN OR FURTHER INFORMATION ON THE SEQUENCE OF EVENTS, THIS REMAINS SPECULATIVE. ANOTHER POSSIBLE CAUSE WOULD BE A DISRUPTION IN THE ELECTRICAL CONNECTION BETWEEN THE SENSOR AND THE CONSOLE, SUCH AS A DISCONNECTION OR BREAK IN THE SENSOR CABLE. THIS EXPLANATION IS CONSIDERED UNLIKELY IN THIS CASE BECAUSE THE VENOUS PRESSURE SENSOR IS PART OF A BUNDLED SENSOR CABLE CONNECTED TO THE OXYGENATOR THAT ALSO TRANSMITS OTHER PRESSURE READINGS, SUCH AS PINT AND PART PRESSURE, AS WELL AS ARTERIAL TEMPERATURE DATA. IF THE CABLE HAD BEEN PHYSICALLY DISCONNECTED OR DAMAGED, IT WOULD BE EXPECTED THAT ALL SENSOR READINGS WOULD BE LOST SIMULTANEOUSLY, NOT JUST PVEN. SINCE NO LOSS OF ADDITIONAL PARAMETERS WAS REPORTED, A GLOBAL CABLE FAILURE APPEARS INCONSISTENT WITH THE EVENT DESCRIPTION. LIKEWISE, A SOFTWARE OR SIGNAL ACQUISITION FAULT WITHIN THE ECMO CONSOLE MAY (GENERALLY) AFFECT MULTIPLE CHANNELS AND TRIGGER A SYSTEM ALERT, NEITHER OF WHICH WAS NOTED BY THE CUSTOMER. IN CONCLUSION, WHILE A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITHOUT FURTHER INFORMATION OR PRODUCT EVALUATION, THE CIRCUMSTANCES OF THE EVENT ARE MORE CONSISTENT WITH EXTERNAL, USER RELATED FACTORS THAT IMPACTED THE CIRCUIT AFTER PRIMING. INTERNAL SYSTEM OR COMPONENT FAILURE IS CONSIDERED LESS LIKELY GIVEN THE ISOLATED NATURE OF THE REPORTED SIGNAL LOSS. AS SUCH, THE ISSUE REMAINS NON-REPRODUCIBLE AND INCONCLUSIVE. REGARDING PATIENT OUTCOME, THE CUSTOMER REPORTED THAT THE PATIENT EXPIRED. NO ADDITIONAL INFORMATION ON PATIENT STATUS, COMORBIDITIES, OR ANY POTENTIAL RELATIONSHIP BETWEEN THIS INCIDENT AND THE PATIENT OUTCOME WAS PROVIDED. THEREFORE, NO CAUSAL ASSOCIATION BETWEEN THE FAILURE OF VENOUS PRESSURE MONITORING AND THE PATIENT¿S EXPIRATION CAN BE ESTABLISHED." ACCORDING TO THE INSTRUCTION FOR USE (HLS SET, CHAPTER "PREPARATION AND INSTALLATION") THE PRESSURE SENSORS HAVE TO BE CHECKED BEFORE PRIMING. IN GENERAL, IT SHOULD BE NOTED THAT PRE-PRIMING CAN HAVE A NEGATIVE IMPACT ON THE SENSOR FUNCTION AND THE ASSOCIATED PRESSURE VALUES AND CAN LEAD TO THE OUTPUT OF UNEXPECTED / IMPLAUSIBLE PRESSURE VALUES SHORTLY BEFORE USE, AS INDICATED IN CHAPTER "PRIMING THE SYSTEM", WHERE IT SAYS "ONLY FILL THE SYSTEM SHORTLY BEFORE USE AND IN ACCORDANCE WITH THE PRIMING INSTRUCTIONS. CALIBRATION VALUES ARE LOST WHEN THE DRIVE UNIT IS SWITCHED OFF". FURTHER THE CARDIOHELP HAS A FLOW/BUBBLE SENSOR AND A VENOUS PROBE TO MEASURE AND CONTROL THE BLOOD FLOW AND PARAMETERS. IF THE MEASURED VALUES ARE ABOVE HIGH LIMIT OR BELOW LOW LIMIT OF THE SET LIMITS THE SYSTEM GENERATES A VISUAL AND ACOUSTICAL ALARM. THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON 2025-07-21. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARD TO THE REPORTED FAILURE WERE FOUND. BASED ON THE RESULTS AND THE PROVIDED PHOTOGRAPHICAL EVIDENCE PROVIDED BY THE CUSTOMER, THE REPORTED FAILURE "PRESSURE READING ISSUE" COULD BE CONFIRMED. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22089344&pc=OZD | 1220648-2025-28735 | 2025-05-05 04:00 | Death | ABIOMED, INC. | 2025-05-27 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemolysis; Hemorrhage/Bleeding; Mitral Perforation | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND EXHIBITED ACCESS SITE BLEEDING. HEMOSTASIS WAS ACHIEVED BY REPOSITIONING THE SHEATH. SUPPORT WAS CONTINUED. ALSO, THERE WAS LOW PUMP FLOW AND THE PATIENT'S PLASMA FREE HEMOGLOBIN SPIKED. THE LOW PUMP FLOW ERROR RESOLVED WITHOUT INTERVENTION. THE PATIENT THEN EXPERIENCED MITRAL VALVE RUPTURE. IT IS UNCLEAR IF THE RUPTURE WAS A RESULT OF IMPELLA USE OR AN ACUTE MYOCARDIAL INFARCTION. ALSO, A BLEED OCCURRED WHICH APPEARED TO COME FROM THE IMPELLA SIDEPORT. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF HEART VALVE DAMAGE, HEMOLYSIS, ACCESS SITE BLEEDING, AND LOW PUMP FLOW HAS BEEN COMPLETED. HEART VALVE DAMAGE: THE ROOT CAUSE OF THE HEART VALVE DAMAGE WAS NOT DETERMINED AS LIMITED CLINICAL INFORMATION RELATED TO FAILURE WAS PROVIDED. HEMOLYSIS: THE RETURNED PRODUCT WAS INSPECTED AND THERE WERE NO PHYSICAL ABNORMALITIES RELATED TO HEMOLYSIS SUCH AS BIOMATERIAL WRAPPED AROUND THE IMPELLER OR DAMAGED IMPELLER BLADES. THE DATA LOGS FROM THE CASE DO NOT SHOW ANY MOTOR CURRENT SPIKES OR POSITIONING ISSUES. CLINICAL DETAILS NOTE SUCTION ALARMS PFHGB SPIKED TO 300 BUT DECREASED TO 13 WITHIN 24H. CLINICAL DETAILS DO NOT CONFIRM PUMP OUT OF POSITION. THE PUMP WAS HEMOLYSIS TESTED AND PASSED. THE ROOT CAUSE OF HEMOLYSIS WAS NOT DETERMINED AS THERE WERE NO DATA LOG ABNORMALITIES, ABNORMALITIES ON RETURNED PRODUCT AND PUMP PASSED HEMOLYSIS TESTING; LIMITED CLINICAL INFORMATION WAS PROVIDED. ACCESS SITE BLEEDING: ¿CLINICAL MENTION A BLEED WITH NO APPARENT FOCUS (LARGE VOLUME OF BRIGHT RED BLOOD IN THE BED. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING WAS NOT DETERMINED AS LIMITED CLINICAL INFORMATION WAS PROVIDED. LOW PUMP FLOW: THE RETURNED PUMP DID NOT HAVE ANY BIOMATERIAL OR CANNULA KINKS. THE DATA LOGS CONFIRM SUCTION ALARMS. THERE WERE NO MOTOR CURRENT SPIKES OR POSITION RELATED ISSUES. CLINICAL MENTION SUCTION ALARMS (DIASTOLIC) AND THIS WAS ASSOCIATED WITH INTRAVASCULAR HYPOVOLEMIA AND SIRS. THE PUMP WAS TESTED IN THE LAB AND LOW PUMP FLOW DID NOT REPRODUCE. THE ROOT CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY PATIENT CONDITION AS NO LOG ABNORMALITIES WERE OBSERVED, NO ISSUES OR REPRODUCTION ON THE RETURNED PRODUCT AND CLINICAL DETAILS MENTION SUCTION WAS ASSOCIATED WITH INTRAVASCULAR HYPOVOLEMIA AND SIRS. D9 DEVICE RETURNED TO MANUFACTURER DATE ADDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ SECTION: IMPELLA CP WITH SMARTASSIST CATHETER INSERTION ¿IMPELLA CATHETER PERFORMANCE WILL BE COMPROMISED IF CORRECT PLACEMENT CANNOT BE CONFIRMED. WHILE OTHER IMAGING TECHNIQUES, SUCH AS TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE), PORTABLE C-ARM FLUOROSCOPY, OR CHEST X-RAY CAN HELP CONFIRM THE POSITION OF THE IMPELLA CATHETER AFTER PLACEMENT, THESE METHODS DO NOT ALLOW VISUALIZATION OF THE ENTIRE CATHETER ASSEMBLY AND ARE INADEQUATE FOR RELIABLY PLACING THE IMPELLA CATHETER ACROSS THE AORTIC VALVE¿. SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER IMPELLA CATHETER INLET TO THE AORTA ¿IF THE INLET AREA OF THE IMPELLA CATHETER IS TOO CLOSE TO OR ENTANGLED IN THE PAPILLARY MUSCLE AND/OR SUBANNULAR STRUCTURES SURROUNDING THE MITRAL VALVE, IT CAN AFFECT MITRAL VALVE FUNCTION AND NEGATIVELY IMPACT CATHETER FLOW.¿ SECTION: UNDERSTANDING AND MANAGING IMPELLA CATHETER POSITION ALARMS ¿IF THE IMPELLA CATHETER IS COMPLETELY OUT OF THE VENTRICLE, DO NOT ATTEMPT TO REPOSITION THE CATHETER ACROSS THE VALVE WITHOUT A GUIDEWIRE.¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22089685&pc=LGW | 3008514029-2025-00415 | 2025-05-02 04:00 | Death | NEVRO CORP. | 2025-05-27 04:00 | LGW | SENZA | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Sepsis; Insufficient Information | P130022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO A SEPSIS INFECTION. A MEDICAL ASSESSMENT FROM THE PHYSICIAN REGARDING THE CAUSE OF DEATH WAS NOT AVAILABLE. THERE WERE NO REPORTS OF DEVICE-RELATED ISSUES FROM THE PATIENT PRIOR TO THE PASSING. Manufacturer Narrative: A REVIEW OF THE DEVICE'S MANUFACTURING RECORDS WAS COMPLETED, INCLUDING BATCH HISTORY AND DEVICE HISTORY; THERE WERE NO DEVIATIONS OR NONCONFORMITIES ASSOCIATED WITH THE REPORTED EVENT. THE DEVICE WAS NOT RETURNED FOR EVALUATION, WHICH LIMITED THE ABILITY TO CONDUCT ANY PHYSICAL, FUNCTIONAL, AND/OR ROOT CAUSE ANALYSIS. NEVRO SUBMITS THIS REPORT IN COMPLIANCE WITH FDA¿S MEDICAL DEVICE REPORTING REGULATIONS UNDER 21 CFR PART 803. NEVRO HAS COMPLIED WITH REGULATORY INVESTIGATION REQUIREMENTS AND IS SUBMITTING ALL INFORMATION THAT IS REASONABLY KNOWN TO US AT THIS TIME. HOWEVER, WE MAY NOT HAVE BEEN ABLE TO CONFIRM THIS INFORMATION OR COMPLETE THE INVESTIGATION WITHIN THE TIMEFRAME FOR FILING THIS REPORT. WE MAY HAVE GIVEN NO RESPONSE OR AN INCOMPLETE RESPONSE TO CERTAIN QUESTIONS BECAUSE WE DO NOT CURRENTLY HAVE INFORMATION AVAILABLE TO PROVIDE A COMPLETE RESPONSE. IF WE LATER OBTAIN ANY REQUIRED INFORMATION THAT WAS NOT AVAILABLE AT THE TIME OF THIS INITIAL REPORT, WE WILL SUBMIT A SUPPLEMENTAL REPORT. THIS REPORT IS NOT AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT WAS DEFECTIVE, THAT IT MALFUNCTIONED, OR THAT IT CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THIS REPORT. WE MAY CONCLUDE THAT THE DEVICE HAD NO DEFECT, DID NOT MALFUNCTION, OR DID NOT CAUSE OR CONTRIBUTE TO A REPORTABLE EVENT. SOME OF THE ITEMS ON THIS FORM INCLUDE FORCED-CHOICE TERMS USED BY THE FDA FOR REPORTING PURPOSES THAT DO NOT NECESSARILY REFLECT NEVRO¿S CONCLUSIONS ABOUT THE CAUSES OR NATURE OF THE EVENT. THIS STATEMENT SHOULD BE INCLUDED WITH ANY FREEDOM OF INFORMATION ACT RESPONSE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22089885&pc=OZD | 1220648-2025-28739 | 2025-05-03 04:00 | Death | ABIOMED, INC. | 2025-05-27 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED CP PUMP WAS PLACED TO SUPPORT A PATIENT ADMITTED IN FOR SURGERY AND PREOPERATIVELY THE CP WAS PLACED. AFTER 34 HOURS OF SUPPORT THE PATIENT WENT INTO PULSELESS VT AND THE PATIENT EXPIRED WHEN CARE WAS WITHDRAWN. THE PATIENT EXPIRED ON THE CP SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22090019&pc=OZD | 1220648-2025-28742 | 2025-05-03 04:00 | Death | ABIOMED, INC. | 2025-05-27 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Atrial Fibrillation; Hemolysis; Hemorrhage/Bleeding; Ectopic Heartbeat | P140003 |  | 1 | Event Description: A PATIENT WAS FOUND TO NOT HAVE OBSTRUCTIVE DISEASE. THE DECISION WAS MADE TO PLACE IMPELLA CP FOR HEMODYNAMIC SUPPORT. DURING SUPPORT, THE PATIENT WAS STILL HAVING ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE. LEVOPHED WAS HELD AND THEY WERE USING NEO-SYNEPHRINE FOR BLOOD PRESSURE SUPPORT THE RATE WAS IMPROVING WITH ADJUSTMENT. ALSO THE PATIENT WAS GIVEN ONE AND A HALF LITER OF FLUIDS FOR HEMOLYSIS AND DECLINING RENAL FUNCTION. THEY ALSO REPORTED THAT HEPARIN WAS ON HOLD DUE TO BLEEDING AT THE SITE. ALSO THE PATIENT WAS GIVEN ALBUMIN AND A 500 MILLILITER BOLUS TO RESOLVE SUCTION. ADDITIONALLY, PERFORMANCE LEVEL WAS DECREASED DUE TO VENTRICULAR ECTOPY AND CONCERNS FOR SUCTION/HEMOLYSIS. THE PATIENT REMAINED ON SUPPORT FOR 7 DAYS IN TOTAL.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: "ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA." SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: "REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE." SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)." SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: "EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT." IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: "OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE." SECTION: HEMOLYSIS: "HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS". "PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS." SECTION: SUCTION: "SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE." Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING, HEMOLYSIS, VF/VT/AF/AT, AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ACCESS SITE BLEEDING - MAJOR: THE ROOT CAUSE OF THE ACCESS SITE BLEEDING WAS MOST LIKELY DUE TO ANTICOAGULATION MANAGEMENT AS HEPARIN WAS PAUSED AND BLEEDING WAS ABLE TO RESOLVE WITHOUT ANY OTHER TROUBLESHOOTING. HEMOLYSIS: DATA LOGS WERE REVIEWED AND SUCTION ALARMS OCCURRED ONLY ON FIRST DAY OF SUPPORT. NO MOTOR CURRENT SPIKES OR PERSISTENT POSITIONING ALARMS WERE NOTED IN LOGS. CLINICAL DETAILS NOTED THAT THE PATIENT WAS EXPERIENCING RED URINE AFTER PUMP INSERTION. PATIENT RECEIVED MULTIPLE ROUNDS OF FLUID AND BOLUSES AND RED URINE CLEARED UP. THE ROOT CAUSE OF THE HEMOLYSIS WAS MOST LIKELY PATIENT CONDITION RELATED AS DATA LOGS DID NOT SHOW ANY SIGNS OF INGESTION OR IMPROPER POSITIONING OF THE PUMP AND CLINICAL DETAILS NOTED THAT ADDITIONAL FLUIDS HELP REDUCE RED URINE. LOW PUMP FLOW: DATA LOGS WERE REVIEWED AND THE LOG SHOWED SUCTION ALARMS OCCURRED ON FIRST DAY OF SUPPORT ON (B)(6) 2025 WITH SLIGHT INCREASE IN PS. PUMP FLOWS PRIOR TO SUCTION ALARMS SHOWED FLOWS LOWER THAN SET P-LEVEL OF P-9. WHEN PUMP WAS TURNED DOWN TO P-8, PUMP FLOWS WERE WITHIN NORMAL P-LEVEL RANGE FOR A FEW DAYS BEFORE PUMP FLOWS AGAIN DROPPED BELOW SET P-LEVEL RANGE ON (B)(6) 2025. WHEN PUMP WAS TURNED DOWN TO P-7 ON (B)(6) 2025, PUMP FLOWS WERE BELOW SET P-LEVEL RANGE FOR REMINDER OF DATA LOGS. NO MOTOR CURRENT SPIKES NOTED AT TIME OF LOW FLOWS. CLINICAL DETAILS NOTED THAT PATIENT HAD SUCTION ALARMS OCCURRING AND FLUIDS WERE GIVEN AND SUCTION ALARMS RESOLVED. THE ROOT CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY PATIENT CONDITION RELATED AS CLINICAL DETAILS NOTED THAT ADDITIONAL FLUIDS RESOLVED SUCTION ALARMS. VF/VT: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. B2 OUTCOME DEATH AND DATE OF DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28742. B5 PATIENT INFORMATION REGARDING OUTCOME OF DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28742. H1 TYPE OF REPORT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28742. H6 HEALTH EFFECT - IMPACT CODE 1802 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28742.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22090124&pc=OZD | 1220648-2025-28743 | 2025-05-03 04:00 | Death | ABIOMED, INC. | 2025-05-27 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Tachycardia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 USING BEST PRACTICES. DURING SUPPORT, IT WAS NOTED THAT THE VENTRICULAR TACHYCARDIA REQUIRING ONE SHOCK AND REQUIRED REINTUBATION. ADDITIONALLY, THE PATIENT HAD A GASTROINTESTINAL BLEED. ONE UNIT OF BLOOD WAS GIVEN OVERNIGHT, IN ADDITION TO ANOTHER UNIT GIVEN DURING THE DAY ALONG WITH TWO UNITS OF PLASMA. IT WAS NOTED THAT THE PATIENT WAS EXPERIENCING INTESTINAL ISSUES. HOWEVER, A FULL BODY COMPUTED TOMOGRAPHY SCAN DID NOT REVEAL ANY SIGNIFICANT FINDINGS. THEY WERE EVALUATING FOR LIVER VERSUS BOWEL/GASTRIC ISCHEMIA DUE TO PERSISTENT LACTATE SPILL. ULTIMATELY, CARE WAS WITHDRAWN DUE TO ISCHEMIC BOWEL AND STOMACH ISSUES AND THE PATIENT EXPIRED SHORTLY AFTER SUPPORT WAS RAMPED DOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22090281&pc=OZD | 1220648-2025-28744 | 2025-05-03 04:00 | Death | ABIOMED, INC. | 2025-05-27 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON A SECOND IMPELLA CP PUMP AND CARDIOPULMONARY RESUSCITATION (CPR) WAS ON GOING. THE IMPELLA CP AT P8 HAD A CONSTANT SUCTION ALARM. P LEVEL WAS WEANED UNTIL SUCTION RESOLVED. ONLY ABLE TO FLOW 2.0 LITERS PER MINUTE WITH THE IMPELLA CP DUE TO SUCTION AT HIGHER P LEVEL. TEAM BLOUSING FLUIDS AND VERIFYING THE POSITION. NO RETURN OF SPONTANEOUS CIRCULATION WITH HOLDING CPR. SYSTEMIC PRESSURES WERE NOT ADEQUATE DESPITE IMPELLA CP SUPPORT AND ADVANCED CARDIAC LIFE SUPPORT MEDICATIONS. THE DOCTOR MADE THE DECISION TO STOP RESUSCITATION EFFORTS. THE PATIENT EXPIRED IN THE PROCEDURAL AREA. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22090684&pc=OZD | 1220648-2025-28747 | 2025-02-18 05:00 | Death | ABIOMED, INC. | 2025-05-27 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation | P140003 |  | 1 | Event Description: IT WAS REPORTED VIA ABIOMED'S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND EXPERIENCED ATRIAL ARRYTHMIA "POSSIBLY" RELATED TO THE PUMP. THE LOQI STATES: "AFTER PCI WITH IMPELLA PATIENT WAS FOUND TO BE IN NEW ONSET ATRIAL FIBRILLATION WITH RVR VIA EKG. PATIENT WAS STARTED ON AN AMIODARONE DRIP TRANSITIONING TO AN ORAL DOSE. PLAN WAS FOR THE PATIENT TO UNDERGO A CARDIOVERSION WITH ANESTHESIA ON (B)(6) 2025 BUT PATIENT EXPIRED THE DAY BEFORE DUE TO CARDIAC ARREST." THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22090855&pc=MVK | 3008642652-2025-05228 | 2025-03-27 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-27 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED EIGHTEEN APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 21:09:16 ON (B)(6) 2025. AT 11:58:38 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 11:59:14, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS PACED RHYTHM @ 100 BPM. AT 12:00:37, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS PACED RHYTHM @ 120 BPM /VF. AT 12:01:41, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS PACED RHYTHM @ 100 BPM. AT 12:02:27, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH PACEMAKER SPIKES. AT 12:02:58, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH PACEMAKER SPIKES. POST SHOCK RHYTHM WAS PACED RHYTHM @ 100 BPM. AT 12:03:25, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 130 BPM . POST SHOCK RHYTHM WAS VT @ 160 BPM WITH PACEMAKER SPIKES/VF WITH PACEMAKER SPIKES. AT 12:03:57, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH PACEMAKER SPIKES. POST SHOCK RHYTHM WAS PACED RHYTHM @ 90 BPM WITH MOTION ARTIFACT. AT 12:05:48, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 12:06:20, THE PATIENT RECEIVED THE SEVENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS PACED RHYTHM @ 80 BPM. AT 12:08:07, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 12:08:37, THE PATIENT RECEIVED THE EIGHTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 100 BPM/VF. AT 12:09:14, THE PATIENT RECEIVED THE NINTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 100 BPM/VF. AT 12:09:48, THE PATIENT RECEIVED THE TENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 110 BPM/VF. AT 12:10:17, THE PATIENT RECEIVED THE ELEVENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 110 BPM/VF. AT 12:10:49, THE PATIENT RECEIVED THE TWELFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 110 BPM/VF. AT 12:18:23, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH VARYING AMPLITUDES/CPR/MOTION ARTIFACT. AT 12:19:46, THE PATIENT RECEIVED THE THIRTEENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH VARYING AMPLITUDES AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VF WITH VARYING AMPLITUDES AND CPR/MOTION ARTIFACT. AT 12:20:20, THE PATIENT RECEIVED THE FOURTEENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH VARYING AMPLITUDES AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS VT @ 100 BPM. AT 12:21:38, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH VARYING AMPLITUDES/CPR/MOTION ARTIFACT. AT 12:23:02, THE PATIENT RECEIVED THE FIFTEENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH VARYING AMPLITUDES AND CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE FOR 4 SECONDS TRANSITIONING TO VT @ 100 BPM. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 12:29:22, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 100 BPM. AT 12:30:14, THE PATIENT RECEIVED THE SIXTEENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 100 BPM. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 90 BPM. AT 12:30:45, THE PATIENT RECEIVED THE SEVENTEENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 100 BPM. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 90 BPM DEGRADING TO VF WITH VARYING AMPLITUDES/CPR/MOTION ARTIFACT. AT 12:32:57, THE PATIENT RECEIVED THE EIGHTEENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH VARYING AMPLITUDES/CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH CPR/ELECTRODE LEAD FALL OFF. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE DEVICE WAS SHUT DOWN AT 14:16:12 ON (B)(6) 2025. PER HOLTER MONITOR THE PATIENT WAS IN AN IDIOVENTRICULAR RHYTHM/PACED RHYTHM @ 40 BPM WITH CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF AT THE TIME OF THE DEVICE SHUTDOWN. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22090860&pc=MVK | 3008642652-2025-05227 | 2025-05-01 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-27 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED SEVEN INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 17:38:49 ON (B)(6) 2025. AT 03:21:19 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS NSVT. AT 03:22:06, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 65 BPM WITH PVCS/NSVT. AT 03:22:34, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA @ 120 BPM WITH PVCS/NSVT. AT 03:23:21, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS NSVT. AT 03:23:51, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 60 BPM WITH PVCS/NSVT. AT 03:24:18, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS RHYTHM @ 80 BPM WITH PVCS/NSVT. POST SHOCK RHYTHM WAS NSVT. AT 03:24:57, THE PATIENT RECEIVED THE FIFTH INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 80 BPM WITH PVCS/NSVT. AT 03:25:31, THE PATIENT RECEIVED THE SIXTH INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS NSVT. AT 03:25:53, THE PATIENT RECEIVED THE SEVENTH INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS NSVT. THE DEVICE WAS SHUT DOWN AT 04:15:08 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22091123&pc=NVN | 2017865-2025-64795 | 2011-02-20 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-27 04:00 | NVN | TENDRIL SDX LEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PULSE GENERATOR AND RIGHT VENTRICULAR (RV), AND RIGHT ATRIAL (RA) LEADS WERE EXTRACTED DUE TO PATIENT DEATH. THE IMMEDIATE CAUSE OF DEATH WAS INDICATED AS CARDIAC INSUFFICIENCY. THERE WAS NO ALLEGATION OF MALFUNCTION. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED FOR ANALYSIS. ELECTRICAL TESTING THAT COULD BE MEASURED DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22091141&pc=NVZ | 2017865-2025-64793 | 2011-02-20 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-27 04:00 | NVZ | ZEPHYR XL DR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P880086 |  | 1 | Event Description: IT WAS REPORTED THAT THE PULSE GENERATOR AND RIGHT VENTRICULAR (RV), AND RIGHT ATRIAL (RA) LEADS WERE EXTRACTED DUE TO PATIENT DEATH. THE IMMEDIATE CAUSE OF DEATH WAS INDICATED AS CARDIAC INSUFFICIENCY. THERE WAS NO ALLEGATION OF MALFUNCTION. Manufacturer Narrative: CORRECTION: H6 AND B2 - "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES)" WAS INCORRECTLY SELECTED IN PREVIOUS REPORT. Manufacturer Narrative: THE DEVICE HAD NO TELEMETRY AND BATTERY AT END OF SERVICE (EOS) LEVEL UPON RECEIPT. DEVICE IMAGE COULD NOT BE SAVED, AND DATA TREND COULD NOT BE PLOTTED DUE TO LACK OF DEVICE IMAGE. THE DEVICE WAS EVALUATED UNDER NOMINAL CONDITIONS, AFTER REPLACING THE BATTERY, AND RESULTS INDICATED NORMAL FUNCTIONALITY. FURTHER ANALYSIS AT VARIOUS PULSE AMPLITUDE FOUND NORMAL CURRENT AND VOLTAGE LEVEL. LONGEVITY ASSESSMENT COULD NOT BE PERFORMED BECAUSE DEVICE IMAGE COULD NOT BE SAVED AND THERE WERE NO OTHER SOURCE OF INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22091155&pc=NVN | 2017865-2025-64794 | 2011-02-20 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-27 04:00 | NVN | TENDRIL SDX LEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PULSE GENERATOR AND RIGHT VENTRICULAR (RV), AND RIGHT ATRIAL (RA) LEADS WERE EXTRACTED DUE TO PATIENT DEATH. THE IMMEDIATE CAUSE OF DEATH WAS INDICATED AS CARDIAC INSUFFICIENCY. THERE WAS NO ALLEGATION OF MALFUNCTION. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF A PARTIAL LEAD WAS RETURNED IN ONE PIECE. VISUAL AND X-RAY INSPECTIONS OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22091175&pc=DSQ | 2916596-2025-03208 | 2022-01-01 05:00 | Death | THORATEC CORPORATION | 2025-05-27 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED, ¿MODEL FOR END STAGE LIVER DISEASE EXCLUDING INTERNATIONAL NORMALIZED RATIO PREDICTS SEVERE RIGHT VENTRICULAR FAILURE AFTER HEARTMATE 3 IMPLANTATION IN A CONTEMPORARY COHORT¿, THAT RIGHT VENTRICULAR FAILURE (RVF) AFTER LEFT VENTRICULAR ASSIST DEVICES (LVAD) WAS ASSOCIATED WITH SIGNIFICANT MORBIDITY AND MORTALITY. THE MODEL FOR END STAGE LIVER DISEASE EXCLUDING INTERNATIONAL NORMALIZED RATIO (MELD- XI), IS A PREVIOUSLY VALIDATED MEASURE OF LIVER DYSFUNCTION DEVELOPED FOR PATIENTS WHO MAY BE ON ANTICOAGULATION, WAS CALCULATED AS 5.11×LN(BILIRUBIN)+11.76×LN(CREATININE)+9.44.14 THE EUROMACS (EUROPEAN REGISTRY FOR PATIENTS WITH MECHANICAL CIRCULATORY SUPPORT) SCORE WAS CALCULATED AS RIGHT ATRIAL PRESSURE TO PULMONARY CAPILLARY WEDGE PRESSURE RATIO (RAP/PCWP)>0.54 (2 POINTS)+HEMOGLOBIN =10 G/DL (1 POINT)+MULTIPLE INTRAVENOUS INOTROPES (2.5 POINTS)+INTERMACS (INTERAGENCY REGISTRY FOR MECHANICALLY ASSISTED CIRCULATORY SUPPORT) PROFILE 1 TO 3 (2 POINTS)+SEVERE RV DYSFUNCTION ON ECHOCARDIOGRAPHY (2 POINTS).8 THE MICHIGAN RVF RISK SCORE WAS CALCULATED AS VASOPRESSOR REQUIREMENT (4 POINTS)+ASPARTATE TRANSAMINASE =80 IU/L (2 POINTS)+BILIRUBIN =2.0 MG/DL (2.5 POINTS)+CREATININE =2.3 MG/DL (3 POINTS). THE LIVER IS THE FIRST ORGAN UPSTREAM OF THE RIGHT- SIDED HEART CHAMBERS AND IS SENSITIVE TO ELEVATED CENTRAL VENOUS PRESSURES. CONGESTIVE HEPATOPATHY IS A RECOGNIZED COMPLICATION OF HEART FAILURE (HF) THAT CAN LEAD TO HEPATIC DYSFUNCTION AND IN SOME CASES CARDIAC CIRRHOSIS. (MELD- XI) IS A MODIFICATION OF THE ORIGINAL MELD THAT WAS CREATED FOR USE IN PATIENTS REQUIRING ANTICOAGULATION, WHICH IS COMMON AMONG PATIENTS WITH HF. ELEVATED MELD- XI HAS BEEN ASSOCIATED WITH WORSE SURVIVAL IN HOSPITALIZED PATIENTS WITH HF AND LVAD RECIPIENTS AND MAY PREDICT RVF IN A NON- HM3 POPULATION. THE STUDY RETROSPECTIVELY ANALYZED ALL ADULT PATIENTS OVER THE AGE OF 18 WHO RECEIVED A HEART MATE 3 (HM3) LVAD AS INITIAL IMPLANTATION BETWEEN 2015 AND 2022 AT WASHINGTON UNIVERSITY IN ST. LOUIS AND MASSACHUSETTS GENERAL HOSPITAL. EXCLUDED WERE PATIENTS WHO UNDERWENT PUMP EXCHANGE OR HAD INSUFFICIENT CLINICAL DATA. THE PRIMARY OUTCOME WAS SEVERE RVF ACCORDING TO THE INTERMACS DEFINITION, CHARACTERIZED BY CLINICAL SYMPTOMS OF ELEVATED FILLING PRESSURES AND NEEDING RV ASSIST DEVICE IMPLANTATION, >14 DAYS OF INOTROPES, OR >48 HOURS OF PULMONARY VASODILATORS FOLLOWING HM3 IMPLANTATION. SECONDARY OUTCOMES WERE IN- HOSPITAL DEATH AND A COMPOSITE OF POST DISCHARGE DEATH OR PUMP EXCHANGE. THE STUDY INCLUDED 246 HM3 RECIPIENTS, OF WHOM 28% WERE BLACK AND 20% WERE FEMALE. IN TOTAL, 74 (30%) OF PATIENTS EXPERIENCED SEVERE RVF BY THE INTERMACS DEFINITION. WHEN COMPARING THOSE WHO DEVELOPED SEVERE RVF TO THOSE WHO DID NOT, BASELINE DEMOGRAPHICS AND MEDICAL COMORBIDITIES WERE SIMILAR. HOWEVER, AT TIME OF IMPLANTATION, PATIENTS WHO DEVELOPED SEVERE RVF WERE MORE LIKELY TO BE INTERMACS PROFILE 1 OR 2 (50% VERSUS 26%, P<0.00. ASSESSED WHETHER PREOPERATIVE MELD- XI CAN PREDICT MORTALITY AFTER HM3 IMPLANTATION. IN THE COHORT, 45 (18%) OF PATIENTS DIED WITHIN 1 YEAR OF FOLLOW- UP WITH MEDIAN TIME- TO- DEATH OF 32 DAYS. WHEN STRATIFIED BY MELD- XI =14 OR OR<14 A PRIORI, WE FOUND THAT MELD- XI =14 WAS ASSOCIATED WITH SIGNIFICANTLY WORSE IN- HOSPITAL MORTALITY AFTER LVAD IMPLANTATION (LOG- RANK P=0.002). MELD- XI DID NOT PREDICT THE RISK OF DEATH OR PUMP EXCHANGE AT 1- YEAR FOLLOW- UP AMONG PATIENTS WHO SURVIVED INDEX HOSPITALIZATION (LOG- RANK P=0.9. THE OCCURRENCE OF SEVERE RVF AFTER LVAD INCREASES THE RISK OF IN- HOSPITAL MORTALITY AS WELL AS SIGNIFICANT MORBIDITY INCLUDING THE NEED FOR RENAL REPLACEMENT THERAPY, STROKE, AND GASTROINTESTINAL BLEEDING.2¿4,27¿29 THUS, IDENTIFICATION OF PATIENTS AT INCREASED RISK OF RVF IS PARAMOUNT IN THE EVALUATION OF PATIENTS FOR LVAD CANDIDACY, SHARED DECISION- MAKING TO PROCEED WITH HM3 IMPLANTATION, AND PERIOPERATIVE PLANNING. IN THE PRESENT STUDY, IT IS SHOWN THAT HIGHER MELD- XI IS ASSOCIATED WITH AN INCREASED LIKELIHOOD OF SEVERE RVF. IN CONCLUSION, MELD- XI IS INDEPENDENTLY ASSOCIATED WITH AN INCREASED RISK OF SEVERE RVF AND IN-HOSPITAL MORTALITY AFTER HM3 IMPLANTATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEMS IN USE WERE NOT PROVIDED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF DEATH ESTIMATED AS DATE OF EVENT THE DATE OF EVENT HAS BEEN ENTERED AS 01JAN2022 AS DATA WAS COLLECTED BETWEEN JAN2015 AND JAN2022. AUTHOR INFORMATION: LAMBERT, DAVID S., ET AL. MODEL FOR END STAGE LIVER DISEASE EXCLUDING INTERNATIONAL NORMALIZED RATIO PREDICTS SEVERE RIGHT VENTRICULAR FAILURE AFTER HEARTMATE 3 IMPLANTATION IN A CONTEMPORARY COHORT. JOURNAL OF THE AMERICAN HEART ASSOCIATION VOL14, ISSUE 7 APRIL 2025. HTTP://DX.DOI.ORG/10.1161/JAHA.124.037553 STUDY SITES: DEPARTMENT OF MEDICINE WASHINGTON UNIVERSITY SCHOOL OF MEDICINE IN ST. LOUIS SAINT LOUIS MO USA DEPARTMENT OF CARDIAC SURGERY MASSACHUSETTS GENERAL HOSPITAL BOSTON MA USA CARDIOVASCULAR DIVISION WASHINGTON UNIVERSITY SCHOOL OF MEDICINE IN ST. LOUIS SAINT LOUIS MO USA DIVISION OF CARDIOLOGY MASSACHUSETTS GENERAL HOSPITAL/HARVARD MEDICAL SCHOOL BOSTON MA USA NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22091596&pc=NPU | 2015691-2025-04162 | 2025-05-07 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-27 04:00 | NPU | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES AFFILAITE IN ITALY, REGARDING AN IMPLANT OF A 29MM SAPIEN 3 VALVE IN MITRAL POSITION BY TRANSEPTAL APPROACH AS A VALVE-IN-RING. AT THE END OF THE PROCEDURE, THE PATIENT DEVELOPED VENTRICULAR FIBRILLATION DUE TO AN ANTERIOR MYOCARDIAL INFARCTION (LAD OCCLUSION). THE PATIENT WAS RESUSCITATED AND THE CORONARY ARTERY WAS REOPENED, BUT THE PATIENT PASSED AWAY DURING THE NIGHT. Manufacturer Narrative: E1 PHONE NUMBER (B)(6). PER THE INSTRUCTIONS FOR USE (IFU), THROMBUS FORMATION, PLAQUE DISLODGEMENT, AND EMBOLIZATION THAT MAY RESULT IN MYOCARDIAL INFARCTION (MI) ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS NEAR CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PERCUTANEOUS CORONARY INTERVENTION). MI RELATED TO THE TVR PROCEDURE, AND THE VALVE IMPLANT WILL MANIFEST INTRA-PROCEDURALLY OR IN THE IMMEDIATE POST-OPERATIVE PERIOD AND IS TYPICALLY DUE TO A COMBINATION OF PATIENT AND/OR PROCEDURAL FACTORS. AS DEFINED IN THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC) PUBLICATION ON TVR COMPLICATIONS, THE PERI PROCEDURAL INTERVAL IS INCLUSIVE OF ALL EVENTS THAT BEGIN WITHIN 72 HOURS OF THE INDEX PROCEDURE. MI EVENTS THAT OCCUR AFTER THE PERI-PROCEDURAL PERIOD (MORE THAN 72HRS) ARE TYPICALLY RELATED TO THE PATIENTS UNDERLYING CORONARY DISEASE. MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR MI DURING THE TVR PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE, AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THE CAUSE OF THE EVENT COULD NOT BE DETERMINED. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22092262&pc=NVZ | 2017865-2025-64808 | 2025-05-06 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-27 04:00 | NVZ | ASSURITY RF DR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P880086 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22092264&pc=NVN | 2017865-2025-64811 | 2025-05-06 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-27 04:00 | NVN | TENDRIL ST | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22092277&pc=NVN | 2017865-2025-64812 | 2025-05-06 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-27 04:00 | NVN | TENDRIL ST | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22092773&pc=BZD | 2518422-2025-106039 | 2022-09-14 04:00 | Death | RESPIRONICS, INC. | 2025-05-27 04:00 | BZD | MW, DREAMWEAR FULL CUSHION, RP | Adverse Event Without Identified Device or Use Problem | Dyspnea; Valvular Stenosis | K142554 |  | 1 | Event Description: THE MANUFACTURER RECEIVED A REPORT FROM A PATIENT REPRESENTATIVE THAT THE PATIENT HAD UNDERGONE A TRANSCATHETER MITRAL TENDYNE VALVE REPLACEMENT ON (B)(6) 2022 WITH NO COMPLICATIONS. AFTER APPROXIMATELY 10 DAYS POST-SURGERY, THE PATIENT BEGAN TO HAVE TROUBLE BREATHING, REQUIRING TWO SEPARATE VISITS TO THE HOSPITAL. THE PATIENT WAS SUBSEQUENTLY DISCHARGED, HOWEVER, ON (B)(6) 2022, HE PASSED AWAY AT HIS HOME. THE REPORTER STATES THAT THE PATIENT USED PHILIPS-BRANDED CPAP EQUIPMENT (DREAMWEAVER NIV INTERFACE) FOR APPROXIMATELY 6 YEARS. ON (B)(6) 2022, THE REPORTER RECEIVED A FIELD SAFETY NOTICE (FSN) RELATED TO NON-INVASIVE INTERFACE MAGNET USE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22093410&pc=DTE | 2182208-2025-02576 | 2025-04-30 04:00 | Death | MEDTRONIC, INC. | 2025-05-27 04:00 | DTE | DUAL CHAMBER TEMPORARY PACEMAKER | No Device Output |  | K150246 |  | 1 | Event Description: IT WAS REPORTED THAT THE EXTERNAL PULSE GENERATOR (EPG) STOPPED PACING DURING THE TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE CAUSING THE VALVE TO BE DEPLOYED IN AN UNINTENDED LOCATION. THIS RESULTED IN VALVE MIGRATION AND THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22084587&pc=PNJ | 9612164-2025-02586 | 2025-05-06 04:00 | Death | MEDTRONIC IRELAND | 2025-05-26 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC PERFORATION IN THE RIGHT VENTRICLE DURING A PROCEDURE TO IMPLANT A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). ECHOCARDIOGRAM RIGHT VENTRICLE GRAM, PERIOCARDIOCENTESIS AND RESUSCITATIVE MEASURES WERE PERFORMED. IT WAS REPORTED THAT THE PATIENT DIED. CAUSE OF DEATH WAS PROVIDED AS PERFORATION FROM THE LEADLESS IPG. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6). D9: <(><<)>NO>>. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22084592&pc=DYB | 9612164-2025-02587 | 2025-05-06 04:00 | Death | MEDTRONIC IRELAND | 2025-05-26 04:00 | DYB | MICRA | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC PERFORATION IN THE RIGHT VENTRICLE DURING A PROCEDURE TO IMPLANT A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). ECHOCARDIOGRAM RIGHT VENTRICLE GRAM, PERIOCARDIOCENTESIS AND RESUSCITATIVE MEASURES WERE PERFORMED. IT WAS REPORTED THAT THE PATIENT DIED. CAUSE OF DEATH WAS PROVIDED AS PERFORATION FROM THE LEADLESS IPG. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: MC2AVR1 LEADLESS IPG; IMPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22085020&pc=NVZ | 2017865-2025-64740 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-26 04:00 | NVZ | ELLIPSE DR ICD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. FURTHER INFORMATION WAS REQUESTED HOWEVER, WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22085021&pc=NVY | 2017865-2025-64742 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-26 04:00 | NVY | RIATA ACTIVE FIXATION | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. FURTHER INFORMATION WAS REQUESTED HOWEVER, WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22085022&pc=NVN | 2017865-2025-64741 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-26 04:00 | NVN | TENDRIL SDX LEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. FURTHER INFORMATION WAS REQUESTED HOWEVER, WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22085273&pc=DSQ | 2916596-2025-03153 | 2025-05-04 04:00 | Death | THORATEC CORPORATION | 2025-05-26 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Dyspnea; Paralysis; Coma; Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS AT HOME AND PASSED AWAY ON (B)(6) 2025. THE PATIENT DIED OF AN URGENT READMISSION DUE TO A CEREBRAL HEMORRHAGE. THE HEARTMATE 3 BLOOD PUMP WAS NOT RESECTED BECAUSE NO AUTOPSY WAS REQUESTED.Event Description: IT WAS ADDITIONALLY REPORTED THAT A LOUD NOISE WAS HEARD AND UPON RUSHING TO THE SCENE, THE PATIENT WAS FOUND WITH RIGHT SIDED PARALYSIS. EMERGENCY SERVICES WERE CALLED. AT THE TIME OF THE EMERGENCY CALL, THE PATIENT WAS CONFUSED BUT HAD NO ISSUES WITH CONSCIOUSNESS LEVEL. HOWEVER DURING TRANSPORT, THERE WAS A RAPID DECLINE IN CONSCIOUSNESS AND WORSENING RESPIRATORY STATUS. THE EMERGENCY DEPARTMENT, NEUROLOGY, AND INTENSIVE CARE UNIT LEAD WERE CONTACTED. UPON ARRIVAL, THEY WERE III-300 ON THE JAPAN COMA SCALE (JCS). THE PATIENT HAD SPONTANEOUS BREATHING BUT SHOWED AGONAL (JAW) BREATHING. OXYGENATION AND VENTILATION WERE MAINTAINED WITH APPROXIMATELY 10L/MIN VIA RESERVOIR MASK. ANISOCORIA (UNEQUAL PUPILS) WAS OBSERVED. COMPUTED TOMOGRAPHY (CT) SCAN SHOWED EXTENSIVE SUBCORTICAL HEMORRHAGE IN THE LEFT CEREBRAL HEMISPHERE WITH MIDLINE SHIFT. CONSULTATIONS WERE MADE WITH THE ON-CALL NEUROLOGIST AND NEUROSURGEON. DUE TO THE HEMORRHAGE BEING IN THE DOMINANT HEMISPHERE, THERE WAS A HIGH LIKELIHOOD OF LASTING SEQUELAE SUCH AS HEMIPLEGIA AND APHASIA. NEUROLOGICAL PROGNOSIS WAS DEEMED POOR. DURING EMERGENCY TRANSPORT, THE SON WAS CONTACTED AND EXPRESSED A WISH FOR FULL-CODE TREATMENT. HOWEVER, THE ADVANCE DIRECTIVE INDICATED THAT INVASIVE TREATMENTS WERE NOT DESIRED. AFTER EXPLAINING THE CT FINDINGS AND OTHER DETAILS AGAIN (AS DOCUMENTED IN THE INFORMED CONSENT RECORD), THE FAMILY DECIDED TO RESPECT THE ADVANCE DIRECTIVE. WHILE THE SON WAS AT THE HOSPITAL, OTHER FAMILY MEMBERS WERE CONTACTED TO CONFIRM AGREEMENT WITH THE TREATMENT PLAN. ULTIMATELY, INFORMED CONSENT WAS OBTAINED, AND THE CODE STATUS WAS FINALIZED. TREATMENT INCLUDED: K2 (VITAMIN K), KCENTRA (PROTHROMBIN COMPLEX CONCENTRATE) REVERSAL, NICARDIPINE FOR BLOOD PRESSURE CONTROL, GLYCEROL, CONTINUATION OF HEART FAILURE MEDICATIONS, ELECTROLYTE AND OTHER ADJUSTMENTS. THE CAUSE OF DEATH WAS CEREBRAL HEMORRHAGE. THERE WERE NO NOTABLE CHANGES TO THE PATIENT CONDITION OR ANTICOAGULATION STATUS. THE PATIENT EXPERIENCED IMPAIRED CONSCIOUSNESS, AND THEY WERE TREATED WITH ANTIHYPERTENSIVE AGENTS AND DIURETICS. THE ONSET DATE OF THE CEREBRAL HEMORRHAGE (B)(6) 2025. IT COULD NOT BE RULED OUT THAT THE CEREBRAL HEMORRHAGE WAS DEVICE RELATED, BUT THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE PATIENT'S DEATH AND DEVICE FUNCTION. THE CEREBRAL HEMORRHAGE OCCURRED WHILE THE PATIENT WAS TAKING WARFARIN. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081925&pc=OZO | 2032227-2025-180764 | 2025-04-17 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-24 04:00 | OZO | 630G INSULIN PUMP MMT-1715K 630G BLACK MG | Adverse Event Without Identified Device or Use Problem | High Blood Pressure/ Hypertension; Cardiovascular Insufficiency | P150001 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS HYPERTENSIVE ATHEROSCLEROSIS, CARDIOVASCULAR DISEASE, AND DIABETES. THE CUSTOMER¿S BLOOD GLUCOSE WAS UNKNOWN. THE CUSTOMER WAS WEARING THE INSULIN PUMP AT THE TIME OF DEATH WAS UNKNOWN. THE EVENT INVOLVED PRODUCT(S) MMT-1715K, MMT-242A, MMT-332A. TROUBLESHOOTING WAS NOT PERFORMED. NO PRODUCT RETURN IS REQUIRED FOR MMT-1715K, MMT-242A, MMT-332A. Manufacturer Narrative: HEALTH EFFECT-CLINICAL CODE-1728, HEALTH EFFECT-IMPACT CODE-2199. THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22075697&pc=OZP | 2032227-2025-179933 | 2025-02-13 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-23 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P160017 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS DIABETES AND BAD HEART. THE CUSTOMER¿S BLOOD GLUCOSE WAS UNKNOWN. IT WAS UNKNOWN WHETHER THE CUSTOMER WAS WEARING THE INSULIN PUMP AT THE TIME OF DEATH. THE EVENT INVOLVED PRODUCT(S) MMT-1884, UNOMEDICAL, UNK_RESERVOIR. TROUBLESHOOTING WAS NOT PERFORMED. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884, UNOMEDICAL, UNK_RESERVOIR. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22076414&pc=DTB | 2649622-2025-13987 | 2025-04-13 04:00 | Death | MPRI | 2025-05-23 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Sepsis | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED SEPSIS SEQUELAE. EVENTUALLY, IT WAS NOTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE FULL LEAD WAS RETURNED, ANALYZED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22076417&pc=DTB | 2649622-2025-13988 | 2025-04-13 04:00 | Death | MPRI | 2025-05-23 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Sepsis | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED SEPSIS SEQUELAE. EVENTUALLY, IT WAS NOTED THAT THE PATIENT PASSED AWAY Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE FULL LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22076842&pc=PNJ | 9612164-2025-02553 | 2025-01-01 05:00 | Death | MEDTRONIC IRELAND | 2025-05-23 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Tachycardia; Ventricular Fibrillation; Loss of consciousness | P150033 |  | 1 | Event Description: IT WAS REPORTED POST OPERATIVELY AFTER THE IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), PATIENT HAD DECREASED LEVEL OF CONSCIOUSNESS AND WEAK RESPONSES TO CALLS, INDICATING ACIDOSIS. THE ATRIOVENTRICULAR CONDUCTION FAILED LEADING TO VENTRICULAR PACING DEPENDENCY. THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA (VT), VENTRICULAR FIBRILLATION (VF) AND EVENTUALLY PASSED AWAY. IT WAS DETERMINED THAT THE LEADLESS IPG IMPLANT PROCEDURE WAS NOT THE CAUSE. THE PATIENT IS A PARTICIPANT IN THE POST APPROVAL CLINICAL S URVEILLANCE PRODUCT SURVEILLANCE REGISTRY. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [(B)(6)]). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22077188&pc=FPO | 0001831750-2025-00701 | 2025-03-11 04:00 | Death | STRYKER MEDICAL-KALAMAZOO | 2025-05-23 04:00 | FPO | 6506 MTS PWRPRO COT MID CONFIG | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT THE DEVICE WAS IN AN AMBULANCE ACCIDENT. THE PATIENT INVOLVED PASSED AWAY. NO PRODUCT MALFUNCTION WAS ALLEGED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22077206&pc=OZD | 1220648-2025-28671 | 2025-04-28 04:00 | Death | ABIOMED, INC. | 2025-05-23 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Cardiac Arrest; Ischemia; Hematuria | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS. AFTER THE IMPLANT, SAME DAY, THE PATIENT HAD SUCTION ALARMS. THE PERFORMANCE LEVEL WAS TURN DOWN TO RECTIFY THE ALARM. THE PATIENT WAS GIVEN FLUID AND CONTINUED TO MONITOR. THEREAFTER, A SMALL AMOUNT OF BLOOD WAS IN THE PATIENT'S URINE, ALONG WITH THE PATIENT NOW HAVING A COLD LEG AND INABILITY TO FEEL PULSES (THAT WERE PRESENT EARLIER). A DISCUSSION WAS MADE TO REMOVE THE IMPELLA CP DEVICE AT BEDSIDE AND USE A FEMSTOP TO CREATE HEMOSTASIS. SUBSEQUENTLY, THE IMPELLA CP WAS REMOVED, AND AT SOME POINT, THEREAFTER, CARDIOPULMONARY RESUSCITATION HAD BEEN INITIATED WITH RETURN OF SPONTANEOUS CIRCULATION ACHIEVED. THE NEXT STEPS FOR THE PATIENT WERE NOTED TO BE UNKNOWN. HOWEVER, THE PATIENTS OUTCOME NOTED THE PATIENT HAD EVENTUALLY EXPIRED AFTER A TOTAL OF APPROXIMATELY FIVE HOURS ON IMPELLA MCS. Manufacturer Narrative: THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA RELATED EVENTS AS AN ASSOCIATED FACTOR. VERY LIMITED DETAILS SURROUNDING THE PATIENT¿S DEMISE. THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW, LIMB ISCHEMIA, AND HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. LOW PUMP FLOW: DATA LOGS CONFIRM SUCTION ALARMS AND LOW FLOWS EVEN AFTER REDUCING THE P-LEVEL. FLOWS 2.8L/MIN AT P7. ONCE SUCTIONS ALARMS WERE OFF, P-LEVEL INCREASED TO P8 AND STILL LOW FLOWS OF 3.0L/MIN WERE OBSERVED. PLACEMENT SIGNAL WAS ALSO SEEN TO BE TRENDING UPWARDS AFTER SUCTION EVENTS. THEREFORE, THE ROOT CAUSE OF THE LOW PUMP FLOW ISSUE WAS NOT DETERMINED SINCE INSUFFICIENT CLINICAL INFORMATION WAS PROVIDED AND DATA LOGS WERE INCONCLUSIVE. LIMB ISCHEMIA: DUE TO INSUFFICIENT CLINICAL DETAILS AND INCONCLUSIVE DATA LOG INVESTIGATION, THE ROOT CAUSE OF LIMB ISCHEMIA WAS NOT DETERMINED. HEMOLYSIS: DATA LOGS WERE INVESTIGATED. ¿IMPELLA IN AORTA¿ ALARMS WERE OBSERVED AT 10:10PM RIGHT BEFORE THE PUMP WAS EXPLANTED. NO SUCTION OR MOTOR CURRENT SPIKES AROUND THAT TIME. THEREFORE, DUE TO INSUFFICIENT CLINICAL DETAILS AND INCONCLUSIVE DATA LOG INVESTIGATION, THE ROOT CAUSE OF THE HEMOLYSIS ISSUE WAS NOT DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22077381&pc=PNJ | 9612164-2025-02567 | 2025-04-28 04:00 | Death | MEDTRONIC IRELAND | 2025-05-23 04:00 | PNJ | MICRA¿ AV2 | Use of Device Problem; Activation, Positioning or Separation Problem; Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT INTRA OPERATIVELY DURING THE PLACEMENT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), AN ATTEMPT WAS MADE TO IMPLANT WITH THE HELP OF SNARE. THE DEVICE SLIPPED DURING DEPLOYMENT AND A RAPID DROP IN BLOOD PRESSURE WAS OBSERVED. CARDIAC PERFORATION, PERICARDIAL EFFUSION AND CARDIAC TAMPONADE WERE NOTED. BLOOD PRESSURE DECREASED FURTHER. PERICARDIOCENTESIS AND DRAINAGE WERE PERFORMED. BLOOD PRESSURE TEMPORARILY INCREASED. HOWEVER, THE PATIENT WENT INTO SHOCK AGAIN AND THE TEMPORARY PACEMAKER FAILED TO CAPTURE IT LEADING TO THE INSERTION OF PERCUTANEOUS CARDIOPULMONARY SUPPORT (PCPS). EVENTUALLY, IT WAS NOTED THAT THE PATIENT PASSED AWAY. THE PATIENT IS A PARTICIPANT IN THE POST APPROVAL CLINICAL SURVEILLANCE PRODUCT SURVEILLANCE REGISTRY. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: [MICRA], PRODUCT ID: [MC2AVR1-DELSYS], (SERIAL: [(B)(6)]. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION H6: ANNEX A. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22077393&pc=OZD | 1220648-2025-28673 | 2025-05-02 04:00 | Death | ABIOMED, INC. | 2025-05-23 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT AN IMPELLA CP WAS IMPLANTED IN A PATIENT WITH ST-ELEVATED MYOCARDIAL INFARCTION AND LEFT HEART CATHETERIZATION. THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION THAT PROGRESSED TO PULSELESS ELECTRICAL ACTIVITY. THE PATIENT WAS TREATED WITH INOTROPES AND VASOPRESSORS, CARDIOPULMONARY RESUSCITATION, AND DEFIBRILLATION BUT ULTIMATELY EXPIRED ON SUPPORT. ADDITIONAL INFORMATION WAS RECEIVED. THE PATIENT EXPIRED ON SUPPORT SO THE PUMP IS NOT AVAILABLE. THE PATIENT HAD VENTRICULAR ISSUES TO PULSELESS ELECTRICAL ACTIVITY PRIOR TO THE IMPELLA AS WELL AS AFTER. THE PATIENT WAS IN E SHOCK. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22077596&pc=NVN | 2182208-2025-02524 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-05-23 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiogenic Shock; Mitral Valve Insufficiency/ Regurgitation | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING CONDUCTION SYSTEM PACING (CSP). THE AUTHORS DESCRIBED ONE PATIENT DEATH FROM HEART FAILURE DUE TO AN ISCHEMIC EVENT. THE PATIENT HAD AN ACUTE MYOCARDIAL INFARCTION IN THE TERRITORY OF THE LEFT ANTERIOR DESCENDING ARTERY WHICH LEDTO SEVERE MITRAL INSUFFICIENCY AND RESULTED IN CARDIOGENIC SHOCK AND DEATH. THERE WAS ONE HIS BUNDLE PACING (HBP) LEAD WHICH EXHIBITED AN EARLY DISLODGEMENT. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/73 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: EARLY IMPROVEMENT IN CARDIAC FUNCTION AND DYSSYNCHRONY AFTER PHYSIOLOGICAL UPGRADING IN PACING-INDUCED CARDIOMYO PATHY. PACING AND CLINICAL ELECTROPHYSIOLOGY. 2025; 48:256¿261. DOI: 10.1111/PACE.15126 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22077709&pc=QFG | 3004464228-2025-22453 | 2025-05-14 04:00 | Death | INSULET CORPORATION | 2025-05-23 04:00 | QFG | OMNIPOD 5 AUTOMATED INSULIN DELIVERY SYSTEM | Insufficient Information | Diabetic Ketoacidosis | K203768 |  | 1 | Event Description: IT HAS BEEN REPORTED BY THE PATIENT'S BROTHER, THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT'S BROTHER STATES THE PATIENT WENT INTO DIABETIC KETOACIDOSIS (DKA) AND PASSED AWAY. THE AUTOPSY REPORT STATES DKA AS THE CAUSE OF DEATH. Manufacturer Narrative: PER THE REPORTER, THE PATIENT EXPIRED (B)(6) 2025. AFTER A REVIEW OF CLOUD DATA, THE LAST POD ACTIVATED WAS ON (B)(6) 2025 AND WAS DEACTIVATED (B)(6) 2025. BASED ON THE AVAILABLE INFORMATION, THE DEVICE WAS NOT IN USE THE DAYS PRECEDING OR AT THE TIME OF DEATH. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, THE CASE WILL BE REOPENED AND REASSESSED. CLOUD DATA WAS REVIEWED FOR THIS USER FOR THE PERIOD OF (B)(6) 2025. DATA WAS AVAILABLE UP TO (B)(6) 2025. THE LAST POD USED WAS ACTIVATED AT 22:22 ON (B)(6) 2025. THE POD WAS PAIRED WITH A DEXCOM G6 SENSOR AND THE SYSTEM WAS USED IN AUTOMATED MODE. IN THE DAYS PRIOR, THE SYSTEM WAS USED IN MANUAL MODE WITH NO SENSOR PAIRED. THE AVAILABLE DATA SHOWED THAT THE AUTOMATED ALGORITHM WAS ABLE TO APPROPRIATELY ADAPT THE MICROBOLUS AMOUNTS BASED ON THE ESTIMATED GLUCOSE VALUES AND USER INPUTS. THE POD STARTED RECEIVING ABERRANT SENSOR VALUES FROM THE SENSOR ON (B)(6) 2025. THE SYSTEM RESPONDED APPROPRIATELY TO THESE ABERRANT VALUES, PUTTING THE SYSTEM INTO AUTOMATED MODE: LIMITED AFTER FOUR CONTINUOUS CYCLES AND GENERATING MISSING SENSOR VALUES MESSAGES AFTER ONE HOUR OF CONTINUOUSLY ABERRANT VALUES. THE POD WAS DEACTIVATED AT 22:12 ON (B)(6) 2025 AND DATA ENTRIES FOR THE REMAINING TWO DAYS CONSIST OF NO ACTIVE POD MESSAGES AND TOTAL DAILY INSULIN SUMMARIES. NO EVIDENCE OF ISSUES WERE OBSERVED WITH THE POD OR CONTROLLER ON THE LAST DAYS OF USE BASED ON THE AVAILABLE CLOUD DATA.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22077913&pc=DYE | 2025587-2025-03852 | 2025-04-25 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-23 04:00 | DYE | MOSAIC MITRAL BIOPROSTHETIC HEART VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Laceration(s) | P990064 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING THE IMPLANT PROCEDURE OF THIS 25MM MITRAL BIOPROSTHETIC VALVE, IT WAS EXPLANTED AND REPLACED WITH ANOTHER MANUFACTURERS MECHANICAL VALVE PRODUCT, THEN THE PATIENT DIED. IT WAS FURTHER REPORTED THAT A LEFT ATRIAL TUMORECTOMY, LEFT ATRIAL RAPHE AND LEFT ATRIAL APPENDAGE CLOSURE WERE PERFORMED CONCOMITANTLY DURING THE IMPLANT PROCEDURE. IT WAS THEN NOTED THAT WHEN THE PATIENT WAS WEANED FROM CARDIOPULMONARY BYPASS, AIR WAS RELEASED, LEADING TO CARDIAC ARREST. THE PATIENT WAS PUT BACK ON CARDIOPULMONARY BYPASS, AND A LACERATION WAS CONFIRMED IN THE LEFT VENTRICLE. THE AREA WAS SUTURED AND THE BIOPROSTHETIC VALVE WAS EXPLANTED AND REPLACED WITH A MECHANICAL VALVE OF ANOTHER MANUFACTURER DUE TO CONCERN OF THE BIOPROSTHETIC VALVE STENT HEIGHT. PULSATION WAS RESUMED, BUT THERE WAS A LOT OF BLEEDING WHICH WAS DIFFICULT TO STOP. THE PATIENT DIED. IT IS UNKNOWN AT THIS TIME IF THE BIOPROSTHETIC VALVE CAUSED OR CONTRIBUTED TO THE LACERATION AND SUBSEQUENT BLEEDING. THE EXACT CAUSE OF DEATH WAS NOT RECEIVED. TH EREFORE, RELATEDNESS OF THE DEATH TO THE BIOPROSTHETIC VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE LACERATION OCCURRED WHILE THE MOSAIC VALVE WAS FULLY IMPLANTED, PRIOR TO EXPLANT. IT WAS STATED THAT THE PHYSICIAN DOES NOT BELIEVE THAT THE 25MM MITRAL BIOPROSTHETIC VALVE CAUSED OR CONTRIBUTED TO THIS DEATH OR LACERATION. Manufacturer Narrative: B5: UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22077965&pc=DSQ | 2916596-2025-03397 | 2025-05-12 04:00 | Death | THORATEC CORPORATION | 2025-05-23 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE PATIENT OUTCOME; THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: A2: PATIENT AGE NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078049&pc=NPT | 2025587-2025-03856 | 2025-05-01 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-23 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information; Device Stenosis | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 4 YEARS AND 8 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED THAT REVEALED THE VALVE FAILED DUE TO AN UNKNOWN REASON. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS REPORTED THAT APPROXIMATELY 4 YEARS AND 8 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED THAT REVEALED THE VALVE FAILED DUE TO AN UNKNOWN REASON. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE FAILURE OF THE VALVE WAS AORTIC STENOSIS. THE PATIENT PASSED AWAY BEFORE TREATMENT OR INTERVENTION WAS PERFORMED. THE CAUSE OF DEATH IS UNSPECIFIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B1. B2. B5. ADDED SECOND PARAGRAPH. H1. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078067&pc=DTB | 2649622-2025-14076 | 2025-05-20 04:00 | Death | MPRI | 2025-05-23 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Heart Failure/Congestive Heart Failure | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT¿S IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS REMOVED DUE INFECTION. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED ONE WEEK POST-SYSTEM REMOVAL. FOLLOW-UP WITH THE PATIENT'S CLINIC WAS UNABLE TO CONFIRM A CAUSE OF DEATH. PER AN ONLINE OBITUARY, THE PATIENT EXPIRED DUE TO HEART FAILURE. Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: DDBC3D1 ICD IMPLANTED: (B)(6) 2018 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: H2; H3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE FULL LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078077&pc=LWQ | 2135147-2025-02709 | 2019-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-05-23 04:00 | LWQ | REGENT HEART VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P810002 |  | 1 | Event Description: THE ARTICLE, "A RETROSPECTIVE ANALYSIS OF STANDARDIZED GRADIENT CALCULATIONS FOR EVALUATING PATIENT¿PROSTHESIS MISMATCH FOLLOWING MECHANICAL AORTIC VALVE REPLACEMENT", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO CALCULATE PREOPERATIVE AND POSTOPERATIVE STANDARDIZED TRANSVALVULAR GRADIENTS IN PATIENTS UNDERGOING MECHANICAL AVR AND COMPARE THESE STANDARDIZED VALUES WITH CONVENTIONAL TRANSVALVULAR GRADIENT MEASUREMENTS. DEVICES INCLUDED IN THE STUDY WERE ABBOTT MASTERS HP MECHANICAL HEART VALVE AND ABBOTT REGENT MECHANICAL HEART VALVE. THE ARTICLE CONCLUDED STANDARDIZED GRADIENT CALCULATIONS PROVIDE A MORE OBJECTIVE, RELIABLE, AND PATIENT-SPECIFIC ASSESSMENT OF PROSTHETIC VALVE FUNCTION BY MINIMIZING INTERPATIENT VARIABILITY. THIS APPROACH IMPROVES THE DETECTION OF PATIENT¿PROSTHESIS MISMATCH AND OPTIMIZES POSTOPERATIVE HEMODYNAMIC EVALUATION, POTENTIALLY LEADING TO BETTER PROSTHESIS SELECTION AND SURGICAL DECISION-MAKING. [THE PRIMARY AND CORRESPONDING AUTHOR WAS MUHAMMET FETHI SAGLAM, DEPARTMENT OF CARDIOVASCULAR SURGERY, FACULTY OF MEDICINE, ANKARA YILDIRIM BEYAZIT UNIVERSITY, 06000 ANKARA, TÜRKIYE, WITH CORRESPONDING EMAIL: DR.M.FETHISAGLAM@GMAIL.COM] THE TIME FRAME OF THE STUDY WAS FROM 2019 TO 2024. A TOTAL OF 115 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 85 RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 55.25 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES MELLITUS, CHRONIC KIDNEY DISEASE, AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF A RETROSPECTIVE ANALYSIS OF STANDARDIZED GRADIENT CALCULATIONS FOR EVALUATING PATIENT¿PROSTHESIS MISMATCH FOLLOWING MECHANICAL AORTIC VALVE REPLACEMENT WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES MELLITUS, CHRONIC KIDNEY DISEASE, AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE. SOME OF THE COMPLICATION REPORTED WAS DEATH; THIS COMPLICATION IS ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: A RETROSPECTIVE ANALYSIS OF STANDARDIZED GRADIENT CALCULATIONS FOR EVALUATING PATIENT¿PROSTHESIS MISMATCH FOLLOWING MECHANICAL AORTIC VALVE REPLACEMENT. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: A RETROSPECTIVE ANALYSIS OF STANDARDIZED GRADIENT CALCULATIONS FOR EVALUATING PATIENT¿PROSTHESIS MISMATCH FOLLOWING MECHANICAL AORTIC VALVE REPLACEMENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078079&pc=LWS | 2649622-2025-14077 | 2025-05-20 04:00 | Death | MPRI | 2025-05-23 04:00 | LWS | SPRINT QUATTRO LEAD | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Heart Failure/Congestive Heart Failure | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT¿S IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS REMOVED DUE INFECTION. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED ONE WEEK POST-SYSTEM REMOVAL. FOLLOW-UP WITH THE PATIENT'S CLINIC WAS UNABLE TO CONFIRM A CAUSE OF DEATH. PER AN ONLINE OBITUARY, THE PATIENT EXPIRED DUE TO HEART FAILURE. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PARTIAL LEAD WAS RETURNED IN SEGMENTS, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: DDBC3D1 ICD IMPLANTED: (B)(6) 2018. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078083&pc=LWS | 3004209178-2025-08903 | 2025-05-20 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-23 04:00 | LWS | EVERA¿ S DR | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Heart Failure/Congestive Heart Failure | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT¿S IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SYSTEM WAS REMOVED DUE INFECTION. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED ONE WEEK POST-SYSTEM REMOVAL. FOLLOW-UP WITH THE PATIENT'S CLINIC WAS UNABLE TO CONFIRM A CAUSE OF DEATH. PER AN ONLINE OBITUARY, THE PATIENT EXPIRED DUE TO HEART FAILURE. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: H2; H3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078116&pc=HSN | 3010667733-2025-00646 | 2025-04-28 04:00 | Death | WRIGHT MEDICAL TECHNOLOGY INC | 2025-05-23 04:00 | HSN | UNKNOWN TAR IMPLANT | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Insufficient Information |  |  | 1 | Event Description: THE DIGITAL STRYKER PROPHECY TEAM RECEIVED A CT SCAN INDICATING THAT THE PATIENT MAY REQUIRE A REVISION SURGERY OF TAR FOR REASONS THAT ARE NOT AVAILABLE AT THE TIME OF THIS REPORT.Event Description: THE DIGITAL STRYKER PROPHECY TEAM RECEIVED A CT SCAN INDICATING THAT THE PATIENT MAY REQUIRE A REVISION SURGERY OF TAR FOR REASONS THAT ARE NOT AVAILABLE AT THE TIME OF THIS REPORT. Manufacturer Narrative: CORRECTION - PLEASE REFER H6 DEVICE CODE. THE REPORTED EVENT COULD NOT BE CONFIRMED, SINCE THE DEVICE WAS NOT RETURNED FOR EVALUATION, AND NO OTHER EVIDENCE WAS PROVIDED. THE DEVICE INSPECTION WAS NOT POSSIBLE AS THE PRODUCT WAS NOT RETURNED FOR INVESTIGATION. THE DEVICE HISTORY RECORD COULD NOT BE REVIEWED BECAUSE THE AFFECTED DEVICE WAS NOT RETURNED, AND THE LOT NUMBER WAS NOT COMMUNICATED. A REVIEW OF THE LABELING WAS NOT POSSIBLE BECAUSE THE CATALOG NUMBER AND LOT NUMBER WERE NOT COMMUNICATED. INDICATIONS OF MATERIAL, MANUFACTURING, OR DESIGN RELATED PROBLEMS WERE UNABLE TO BE IDENTIFIED AS THE CATALOG NUMBER AND LOT NUMBER WERE NOT COMMUNICATED. HOWEVER, FAILURE OF TOTAL ANKLE REPLACEMENT TAR OVER TIME IS A KNOWN COMPLICATION. IN CASE OF A PLANNED REVISION PROCEDURE, A MEDICAL OPINION AIMING TO DETERMINE THE ROOT CAUSE OF THE FAILURE IS PART OF THE INTERNAL INVESTIGATION PROCESS. SUFFICIENT RADIOLOGICAL AND CLINICAL INFORMATION MUST BE PROVIDED TO ENABLE A MEANINGFUL CLINICAL ASSESSMENT AND TO IDENTIFY POSSIBLE CAUSES OF FAILURE. IN CASES WHERE A CT SCAN IS THE ONLY AVAILABLE CLINICAL SOURCE, THIS ASSESSMENT IS LIMITED. NO CONCLUSIVE STATEMENT CAN BE PROVIDED, BECAUSE KEY CLINICAL INFORMATION E.G. CLINICAL STATUS, EXACT SYMPTOMS AND RANGE OF MOTION OF THE PATIENT, ASSESSMENT OF THE TREATING PHYSICIAN IS MISSING. THE ROOT CAUSES OF RADIOGRAPHIC FINDINGS SUCH AS RADIOLUCENT AREAS AND BONE CYSTS ARE OFTEN COMPLEX AND MULTIFACTORIAL AND CANNOT BE DETERMINED SCIENTIFICALLY WITHOUT THIS DECISIVE EVIDENCE. DUE TO THESE LIMITATIONS, WE ARE UNABLE IN SUCH CASES TO PROVIDE STATEMENTS ABOUT THE PATIENT, THE PROCEDURE AND OR THE DEVICE IN RELATION TO CAUSE OF THE FAILURE. MORE DETAILED INFORMATION ABOUT THE COMPLAINT EVENT AS WELL AS THE AFFECTED DEVICE MUST BE AVAILABLE TO DETERMINE THE ROOT CAUSE OF THE COMPLAINT EVENT. IF THE DEVICE IS RETURNED OR ANY FURTHER INFORMATION IS PROVIDED, THE INVESTIGATION REPORT WILL BE REASSESSED. Manufacturer Narrative: ADDITIONAL INFORMATION HAS BEEN REQUESTED AND IF RECEIVED, WILL BE PROVIDED IN THE SUPPLEMENTAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078144&pc=NIQ | 2124215-2025-33441 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | NIQ | SYNERGY MEGATRON | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, HEART FAILURE, CARDIAC ARREST, CARDIOGENIC SHOCK THE FOLLOWING EVENT IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE THE PROVIDED DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED. Manufacturer Narrative: DUE TO SAFE HARBOR IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078149&pc=NIQ | 2124215-2025-33442 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | NIQ | SYNERGY MEGATRON | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, HEART FAILURE, CARDIAC ARREST, CARDIOGENIC SHOCK. THE FOLLOWING EVENT IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE THE PROVIDED DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078250&pc=NPT | 2135147-2025-02716 | 2017-09-01 04:00 | Death | ABBOTT MEDICAL | 2025-05-23 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Perivalvular Leak; Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Myocardial Infarction; Renal Failure; Heart Block; Insufficient Information | P190023 |  | 1 | Event Description: THE ARTICLE, "TRANSCATHETER AORTIC VALVE IMPLANTATION VERSUS SURGERY IN LOW-RISK PATIENTS: IN-HOSPITAL AND MID-TERM OUTCOME", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO COMPARE POST PROCEDURAL OUTCOMES AND MID-TERM MORTALITY OF LOW-RISK PATIENTS TREATED BY TRANS-FEMORAL TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) OR SURGICAL AORTIC VALVE REPLACEMENT (AVR) FOR SEVERE AORTIC STENOSIS. DEVICES INCLUDED IN THE STUDY WERE SAPIEN, EVOLUT, PORTICO/NAVITOR, CARPENTIER¿EDWARDS MAGNA EASE, AVALUS, AND EPIC. THE ARTICLE CONCLUDED THAT TAVI PATIENTS DEMONSTRATED A HIGHER MID-TERM MORTALITY AND A HIGHER INCIDENCE OF POST-PROCEDURAL CONDUCTION ABNORMALITIES AND PARA-VALVULAR LEAK (PVL) WHICH REMAIN A CONCERN IN LOW-RISK PATIENTS. [THE PRIMARY AND CORRESPONDING AUTHOR WAS LODO VITTORIA, AZIENDA OSPEDALIERA ORDINE MAURIZIANO DI TORINO, LARGO FILIPPO TURATI 6, 10128, TURIN, ITALY, WITH CORRESPONDING: VITTORIA.LODO90@GMAIL.COM] THE TIME FRAME OF THE STUDY WAS FROM SEPTEMBER 2017 TO DECEMBER 2021. A TOTAL OF 156 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH IT WAS NOT REPORTED HOW MANY RECEIVED AN ABBOTT DEVICE. FOR THE AVR GROUP, THE AVERAGE AGE WAS 79 YEARS AND THE MAJORITY GENDER WAS MALE. FOR THE TAVI GROUP, THE AVERAGE AGE WAS 79 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, DYSLIPIDEMIA, SMOKING, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, PERIPHERAL ARTERY DISEASE, CHRONIC KIDNEY DISEASE, RENAL REPLACEMENT THERAPY, HISTORY OF CEREBROVASCULAR EVENT, HISTORY OF CORONARY DISEASE, HISTORY OF MALIGNANCY, POOR MOBILITY, HEART BLOCK, AND HEART FAILURE. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: TRANSCATHETER AORTIC VALVE IMPLANTATION VERSUS SURGERY IN LOW-RISK PATIENTS: IN-HOSPITAL AND MID-TERM OUTCOMES. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF NAVITOR WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES, DYSLIPIDEMIA, SMOKING, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, PERIPHERAL ARTERY DISEASE, CHRONIC KIDNEY DISEASE, RENAL REPLACEMENT THERAPY, HISTORY OF CEREBROVASCULAR EVENT, HISTORY OF CORONARY DISEASE, HISTORY OF MALIGNANCY, POOR MOBILITY, HEART BLOCK, AND HEART FAILURE. COMPLICATIONS REPORTED INCLUDED DEATH, SURGICAL INTERVENTION (PACEMAKER), ACUTE KIDNEY INJURY (RENAL FAILURE), MYOCARDIAL INFARCTION, HEART BLOCK, ATRIAL FIBRILLATION, PARAVALVULAR LEAK; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: TRANSCATHETER AORTIC VALVE IMPLANTATION VERSUS SURGERY IN LOW-RISK PATIENTS: IN-HOSPITAL AND MID-TERM OUTCOMES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078372&pc=LWR | 2135147-2025-02719 | 2017-09-01 04:00 | Death | ABBOTT MEDICAL | 2025-05-23 04:00 | LWR | EPIC¿ VALVE (AORTIC) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P040021 |  | 1 | Event Description: THE ARTICLE, "TRANSCATHETER AORTIC VALVE IMPLANTATION VERSUS SURGERY IN LOW-RISK PATIENTS: IN-HOSPITAL AND MID-TERM OUTCOME", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO COMPARE POST PROCEDURAL OUTCOMES AND MID-TERM MORTALITY OF LOW-RISK PATIENTS TREATED BY TRANS-FEMORAL TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) OR SURGICAL AORTIC VALVE REPLACEMENT (AVR) FOR SEVERE AORTIC STENOSIS. DEVICES INCLUDED IN THE STUDY WERE SAPIEN, EVOLUT, PORTICO/NAVITOR, CARPENTIER¿EDWARDS MAGNA EASE, AVALUS, AND EPIC. THE ARTICLE CONCLUDED THAT TAVI PATIENTS DEMONSTRATED A HIGHER MID-TERM MORTALITY AND A HIGHER INCIDENCE OF POST-PROCEDURAL CONDUCTION ABNORMALITIES AND PARA-VALVULAR LEAK (PVL) WHICH REMAIN A CONCERN IN LOW-RISK PATIENTS. [THE PRIMARY AND CORRESPONDING AUTHOR WAS LODO VITTORIA, AZIENDA OSPEDALIERA ORDINE MAURIZIANO DI TORINO, LARGO FILIPPO TURATI 6, 10128, TURIN, ITALY, WITH CORRESPONDING: VITTORIA.LODO90@GMAIL.COM] THE TIME FRAME OF THE STUDY WAS FROM SEPTEMBER 2017 TO DECEMBER 2021. A TOTAL OF 156 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH IT WAS NOT REPORTED HOW MANY RECEIVED AN ABBOTT DEVICE. FOR THE AVR GROUP, THE AVERAGE AGE WAS 79 YEARS AND THE MAJORITY GENDER WAS MALE. FOR THE TAVI GROUP, THE AVERAGE AGE WAS 79 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, DYSLIPIDEMIA, SMOKING, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, PERIPHERAL ARTERY DISEASE, CHRONIC KIDNEY DISEASE, RENAL REPLACEMENT THERAPY, HISTORY OF CEREBROVASCULAR EVENT, HISTORY OF CORONARY DISEASE, HISTORY OF MALIGNANCY, POOR MOBILITY, HEART BLOCK, AND HEART FAILURE. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: TRANSCATHETER AORTIC VALVE IMPLANTATION VERSUS SURGERY IN LOW-RISK PATIENTS: IN-HOSPITAL AND MID-TERM OUTCOMES. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF EPIC STENTED PORCINE HEART VALVE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES, DYSLIPIDEMIA, SMOKING, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, PERIPHERAL ARTERY DISEASE, CHRONIC KIDNEY DISEASE, HISTORY OF CEREBROVASCULAR EVENT, HISTORY OF CORONARY DISEASE, HISTORY OF MALIGNANCY, POOR MOBILITY, HEART BLOCK, AND HEART FAILURE. COMPLICATIONS REPORTED INCLUDED DEATH, SURGICAL INTERVENTION (PACEMAKER), ACUTE KIDNEY INJURY (RENAL FAILURE), MYOCARDIAL INFARCTION, HEART BLOCK, ATRIAL FIBRILLATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: TRANSCATHETER AORTIC VALVE IMPLANTATION VERSUS SURGERY IN LOW-RISK PATIENTS: IN-HOSPITAL AND MID-TERM OUTCOMES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078475&pc=DSQ | 3007042319-2025-01177 | 2025-05-18 04:00 | Death | HEARTWARE, INC. | 2025-05-23 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem; Pumping Problem | Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE (VAD) WAS FOUND UNRESPONSIVE AT HOME AND TRANSPORTED BY EMERGENCY MEDICAL SERVICES (EMS) TO A LOCAL EMERGENCY ROOM. THE PATIENT WAS THEN TAKEN TO (B)(6) HOSPITAL, WHERE THEY WERE PRONOUNCED DEAD. THE PUMP EXHIBITED LOW FLOWS ATTRIBUTED TO THE PATIENT'S CONDITION, WHICH WAS DESCRIBED AS AN UNKNOWN MEDICAL EVENT. REVIEW OF THE LOG FILE INDICATED NO MALFUNCTION OF THE DEVICE. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINS. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS REVEALED A SHARP DECREASE IN POWER CONSUMPTION AND ESTIMATED FLOWS BEGINNING ON (B)(6) 2025, LEADING TO PARAMETERS BELOW NORMAL OPERATING RANGE, AND SIX (6) LOW FLOW ALARMS LOGGED SINCE ON (B)(6) 2025. AS A RESULT, THE REPORTED LOW FLOW EVENT WAS CONFIRMED. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT WAS FOUND UNRESPONSIVE AT HOME AND SUBSEQUENTLY PRONOUNCED DEAD AT THE HOSPITAL FOLLOWING LOW FLOWS DUE TO AN UNKNOWN MEDICAL EVENT. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE REPORTED LOW FLOW EVENT MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078493&pc=OOB | 2124215-2025-33453 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIAC ARREST, CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078674&pc=MKJ | 1218058-2025-00093 | 2025-05-15 04:00 | Death | BIO-DETEK INCORPORATED | 2025-05-23 04:00 | MKJ | ONESTEP CPR (AP) ELECTRODE, 8 PER CASE | Unable to Obtain Readings; Failure of Device to Self-Test | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE WAS UNABLE TO OBTAIN AN ECG SIGNAL VIA ELECTRODE PADS. COMPLAINANT INDICATED AFTER THE PATIENT EVENT THESE PADS CAUSED THE ASSOCIATED DEFIBRILLATOR (R SERIES, SN (B)(6)) TO FAIL THE SELF-TEST DURING TESTING. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE PRODUCT WAS NOT RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. THE AVAILABLE DATA LOGS ONLY GO UP TO (B)(6) 2025, AND DO NOT INCLUDE THE DATE OF THE REPORTED EVENT, MAKING THE LOGS INCONCLUSIVE. A REVIEW OF THE ONESTEP CPR ELECTRODES (LOT 1425) WAS PERFORMED, INCLUDING DEVICE HISTORY RECORD (DHR), DEVICE MASTER RECORD (DMR) LOG, AND PRE- RETAIN EVALUATIONS, ALL OF WHICH SHOWED NO DESCREPANCIES. THE ELECTRODES WERE PROPERLY LABELED, BUILT TO SPECIFICATION, AND PASSED ELECTRICAL TESTING. THERE WAS NO FAULT FOUND. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE PRODUCT FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078799&pc=BZD | 2518422-2025-106017 | 2025-05-13 04:00 | Death | RESPIRONICS, INC. | 2025-05-23 04:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING HUSBAND PASSED AWAY FROM CANCER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078817&pc=NWX | 2124215-2025-33493 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | NWX | WOLVERINE CORONARY CUTTING BALLOON | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Unspecified Kidney or Urinary Problem |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIAC ARREST, CARDIOGENIC SHOCK, NEW REQUIREMENT FOR DIALYSIS THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078838&pc=OOB | 2124215-2025-33499 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiac Tamponade; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION HEART FAILURE, CARDIAC TAMPONADE, CARDIOGENIC SHOCK THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078982&pc=OOB | 2124215-2025-33510 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED: DEATH DURING PROCEDURE. DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIAC ARREST. THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR REIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22078996&pc=NIQ | 2124215-2025-33509 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | NIQ | SYNERGY MEGATRON | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION HEART FAILURE, CARDIAC ARREST, CARDIAC ARREST, MYOCARDIAL INFARCTION, CARDIOGENIC SHOCK. THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22079099&pc=LOX | 2124215-2025-33531 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH DURING PROCEDURE DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIAC ARREST THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22079156&pc=LWQ | 2135147-2025-02738 | 2015-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-05-23 04:00 | LWQ | MASTERS SERIES MECHANICAL HEART VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P810002 |  | 1 | Event Description: THE ARTICLE, "SEX-RELATED DIFFERENCES IN PATIENT CHARACTERISTICS, PRACTICE PATTERNS AND OUTCOMES AFTER BIOPROSTHETIC AND MECHANICAL AORTIC VALVE REPLACEMENT", WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE, MULTICENTER STUDY THAT EXAMINED SEX-BASED DIFFERENCES IN BASELINE CHARACTERISTICS, PRACTICE PATTERNS, AND MID-TERM SURVIVAL FOLLOWING AORTIC VALVE REPLACEMENT (AVR). DEVICES INCLUDED IN THIS STUDY WERE SJM MASTERS, SJM REGENT, MEDTRONIC ATS OPEN PIVOT, CYROLIFE ON-X, LIVANOVA CARBOMEDICS, MEDTRONIC HANCOCK II, ABBOTT EPIC, ABBOTT TRIFECTA GT, LIVANOVA CROWN PRT, LIVANOVA SOLOSMART, AND LIVANOVA PERCEVAL S. THE ARTICLE CONCLUDED THAT AT THE TIME OF REFERRAL FOR AVR, FEMALE PATIENTS WERE SIGNIFICANTLY OLDER AND HAD WORSE CLINICAL PROFILES THAN MALE PATIENTS. DESPITE THE HIGHER RATES OF BIOPROSTHETIC VALVE IMPLANTATION AND PPM IN FEMALES, MID-TERM SURVIVAL RATES WERE NOT SIGNIFICANTLY DIFFERENT. IN CONTRAST, MALES, ESPECIALLY THOSE UNDER 65, SHOWED HIGHER MORTALITY FOLLOWING BIOPROSTHETIC AVR. THESE FINDINGS UNDERSCORE THE NEED FOR FURTHER RESEARCH FOCUSING ON THE SEX-BASED TREATMENT DETERMINATES IN AVR. [THE PRIMARY AND CORRESPONDING AUTHOR WAS MILOS MATKOVIC, DEPARTMENT FOR CARDIAC SURGERY, CLINICAL CENTER OF SERBIA, KOSTE TODOROVICA 8, 11000 BELGRADE, SERBIA, WITH CORRESPONDING EMAIL: DR.MATKO@HOTMAIL.COM] THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2015 TO JANUARY 2020. A TOTAL OF 1249 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 803 (64.3%) RECEIVED AN ABBOTT DEVICE. THE MAJORITY GENDER WAS MALE. IN THE MALE GROUP, THE AVERAGE AGE WAS 64.4 YEARS. IN THE FEMALE GROUP, THE AVERAGE AGE WAS 67.2 YEARS. COMORBIDITIES INCLUDED HYPERTENSION, HYPERLIPIDEMIA, CHRONIC OBSTRUCTIVE PULMONARY DISORDER, CHRONIC KIDNEY DISEASE, DIABETES, PERIPHERAL VASCULAR DISEASE, PRIOR CEREBROVASCULAR ACCIDENT Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MASTERS SERIES MECHANICAL HEART VALVE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, HYPERLIPIDEMIA, CHRONIC OBSTRUCTIVE PULMONARY DISORDER, CHRONIC KIDNEY DISEASE, DIABETES, PERIPHERAL VASCULAR DISEASE, PRIOR CEREBROVASCULAR ACCIDENT.. COMPLICATIONS REPORTED INCLUDED DEATH, SURGICAL INTERVENTION (PACEMAKER), STROKE (CEREBROVASCULAR ACCIDENT); THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
B2: DATE OF DEATH IS ESTIMATED.
B3: DATE OF EVENT IS ESTIMATED.
D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
LITERATURE ATTACHMENT: SEX-RELATED DIFFERENCES IN PATIENT CHARACTERISTICS, PRACTICE PATTERNS AND OUTCOMES AFTER BIOPROSTHETIC AND MECHANICAL AORTIC VALVE REPLACEMENT
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22079186&pc=LWQ | 2135147-2025-02735 | 2015-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-05-23 04:00 | LWQ | REGENT HEART VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P810002 |  | 1 | Event Description: THE ARTICLE, "SEX-RELATED DIFFERENCES IN PATIENT CHARACTERISTICS, PRACTICE PATTERNS AND OUTCOMES AFTER BIOPROSTHETIC AND MECHANICAL AORTIC VALVE REPLACEMENT", WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE, MULTICENTER STUDY THAT EXAMINED SEX-BASED DIFFERENCES IN BASELINE CHARACTERISTICS, PRACTICE PATTERNS, AND MID-TERM SURVIVAL FOLLOWING AORTIC VALVE REPLACEMENT (AVR). DEVICES INCLUDED IN THIS STUDY WERE SJM MASTERS, SJM REGENT, MEDTRONIC ATS OPEN PIVOT, CYROLIFE ON-X, LIVANOVA CARBOMEDICS, MEDTRONIC HANCOCK II, ABBOTT EPIC, ABBOTT TRIFECTA GT, LIVANOVA CROWN PRT, LIVANOVA SOLOSMART, AND LIVANOVA PERCEVAL S. THE ARTICLE CONCLUDED THAT AT THE TIME OF REFERRAL FOR AVR, FEMALE PATIENTS WERE SIGNIFICANTLY OLDER AND HAD WORSE CLINICAL PROFILES THAN MALE PATIENTS. DESPITE THE HIGHER RATES OF BIOPROSTHETIC VALVE IMPLANTATION AND PPM IN FEMALES, MID-TERM SURVIVAL RATES WERE NOT SIGNIFICANTLY DIFFERENT. IN CONTRAST, MALES, ESPECIALLY THOSE UNDER 65, SHOWED HIGHER MORTALITY FOLLOWING BIOPROSTHETIC AVR. THESE FINDINGS UNDERSCORE THE NEED FOR FURTHER RESEARCH FOCUSING ON THE SEX-BASED TREATMENT DETERMINATES IN AVR. [THE PRIMARY AND CORRESPONDING AUTHOR WAS MILOS MATKOVIC, DEPARTMENT FOR CARDIAC SURGERY, CLINICAL CENTER OF SERBIA, KOSTE TODOROVICA 8, 11000 BELGRADE, SERBIA, WITH CORRESPONDING EMAIL: DR.MATKO@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS FROM (B)(6) 2015 TO (B)(6) 2020. A TOTAL OF 1249 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 803 (64.3%) RECEIVED AN ABBOTT DEVICE. THE MAJORITY GENDER WAS MALE. IN THE MALE GROUP, THE AVERAGE AGE WAS 64.4 YEARS. IN THE FEMALE GROUP, THE AVERAGE AGE WAS 67.2 YEARS. COMORBIDITIES INCLUDED HYPERTENSION, HYPERLIPIDEMIA, CHRONIC OBSTRUCTIVE PULMONARY DISORDER, CHRONIC KIDNEY DISEASE, DIABETES, PERIPHERAL VASCULAR DISEASE, PRIOR CEREBROVASCULAR ACCIDENT. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF REGENT HEART VALVE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, HYPERLIPIDEMIA, CHRONIC OBSTRUCTIVE PULMONARY DISORDER, CHRONIC KIDNEY DISEASE, DIABETES, PERIPHERAL VASCULAR DISEASE, PRIOR CEREBROVASCULAR ACCIDENT.. COMPLICATIONS REPORTED INCLUDED DEATH, SURGICAL INTERVENTION (PACEMAKER), STROKE (CEREBROVASCULAR ACCIDENT); THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: SEX-RELATED DIFFERENCES IN PATIENT CHARACTERISTICS, PRACTICE PATTERNS AND OUTCOMES AFTER BIOPROSTHETIC AND MECHANICAL AORTIC VALVE REPLACEMENT
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22079220&pc=OZD | 1220648-2025-28692 | 2025-05-02 04:00 | Death | ABIOMED, INC. | 2025-05-23 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Thrombocytopenia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT AN IMPELLA 5.5 WAS IMPLANTED IN A PATIENT ADMITTED WITH MULTIPLE CARDIAC COMORBIDITIES. THE PUMP WAS SUPPORTING ALONG WITH ECMO DURING TRANSFER FROM COMMUNITY TO TERTIARY MEDICAL CENTER. DURING THE SUPPORT THROMBOCYTOPENIA CONCERNS WERE MADE AS WELL AS BLEEDING IN THE LUNGS. THE PATIENT HAD HEPARIN REMOVED FROM SYSTEMIC INFUSION AND SODIUM BICARB WAS ADDED TO THE PURGE. THE PATIENT EXPIRED AFTER SEVEN DAYS OF SUPPORT. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS. ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22079275&pc=LWR | 2135147-2025-02741 | 2015-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-05-23 04:00 | LWR | EPIC¿ VALVE (AORTIC) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P040021 |  | 1 | Event Description: THE ARTICLE, "SEX-RELATED DIFFERENCES IN PATIENT CHARACTERISTICS, PRACTICE PATTERNS AND OUTCOMES AFTER BIOPROSTHETIC AND MECHANICAL AORTIC VALVE REPLACEMENT", WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE, MULTICENTER STUDY THAT EXAMINED SEX-BASED DIFFERENCES IN BASELINE CHARACTERISTICS, PRACTICE PATTERNS, AND MID-TERM SURVIVAL FOLLOWING AORTIC VALVE REPLACEMENT (AVR). DEVICES INCLUDED IN THIS STUDY WERE SJM MASTERS, SJM REGENT, MEDTRONIC ATS OPEN PIVOT, CYROLIFE ON-X, LIVANOVA CARBOMEDICS, MEDTRONIC HANCOCK II, ABBOTT EPIC, ABBOTT TRIFECTA GT, LIVANOVA CROWN PRT, LIVANOVA SOLOSMART, AND LIVANOVA PERCEVAL S. THE ARTICLE CONCLUDED THAT AT THE TIME OF REFERRAL FOR AVR, FEMALE PATIENTS WERE SIGNIFICANTLY OLDER AND HAD WORSE CLINICAL PROFILES THAN MALE PATIENTS. DESPITE THE HIGHER RATES OF BIOPROSTHETIC VALVE IMPLANTATION AND PPM IN FEMALES, MID-TERM SURVIVAL RATES WERE NOT SIGNIFICANTLY DIFFERENT. IN CONTRAST, MALES, ESPECIALLY THOSE UNDER 65, SHOWED HIGHER MORTALITY FOLLOWING BIOPROSTHETIC AVR. THESE FINDINGS UNDERSCORE THE NEED FOR FURTHER RESEARCH FOCUSING ON THE SEX-BASED TREATMENT DETERMINATES IN AVR. [THE PRIMARY AND CORRESPONDING AUTHOR WAS MILOS MATKOVIC, DEPARTMENT FOR CARDIAC SURGERY, CLINICAL CENTER OF SERBIA, KOSTE TODOROVICA 8, 11000 BELGRADE, SERBIA, WITH CORRESPONDING EMAIL: DR.MATKO@HOTMAIL.COM] THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2015 TO JANUARY 2020. A TOTAL OF 1249 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 803 (64.3%) RECEIVED AN ABBOTT DEVICE. THE MAJORITY GENDER WAS MALE. IN THE MALE GROUP, THE AVERAGE AGE WAS 64.4 YEARS. IN THE FEMALE GROUP, THE AVERAGE AGE WAS 67.2 YEARS. COMORBIDITIES INCLUDED HYPERTENSION, HYPERLIPIDEMIA, CHRONIC OBSTRUCTIVE PULMONARY DISORDER, CHRONIC KIDNEY DISEASE, DIABETES, PERIPHERAL VASCULAR DISEASE, PRIOR CEREBROVASCULAR ACCIDENT. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF EPIC STENTED PORCINE HEART VALVE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, HYPERLIPIDEMIA, CHRONIC OBSTRUCTIVE PULMONARY DISORDER, CHRONIC KIDNEY DISEASE, DIABETES, PERIPHERAL VASCULAR DISEASE, PRIOR CEREBROVASCULAR ACCIDENT. COMPLICATIONS REPORTED INCLUDED DEATH, SURGICAL INTERVENTION (PACEMAKER), STROKE (CEREBROVASCULAR ACCIDENT); THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: SEX-RELATED DIFFERENCES IN PATIENT CHARACTERISTICS, PRACTICE PATTERNS AND OUTCOMES AFTER BIOPROSTHETIC AND MECHANICAL AORTIC VALVE REPLACEMENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22079363&pc=OOB | 2124215-2025-33553 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH DURING PROCEDURE DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIAC ARREST THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22079367&pc=NIQ | 2124215-2025-33552 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | NIQ | SYNERGY MEGATRON | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED: DEATH AT DISCHARGE, DECEASED, ACUTE MYOCARDIAL INFARCTION, HEART FAILURE, CARDIAC ARREST, CARDIOGENIC SHOCK. THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22079369&pc=QZI | 2124215-2025-33540 | 2025-02-21 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. DURING A PULSED FIELD ABLATION (PFA) PROCEDURE, A FARAWAVE PFA CATHETER WAS SELECTED FOR USE. THIRTY-TWO (32) DAYS POST PFA PROCEDURE, THE PATIENT PASSED AWAY. NO FURTHER INFORMATION IS AVAILABLE ON THE CAUSE OF DEATH. THERE IS NO INDICATION THAT THE DEVICE WILL BE RETURNED FOR LABORATORY ANALYSIS. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS UNABLE TO BE OBTAINED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE FULL UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22079397&pc=LWR | 2135147-2025-02743 | 2015-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-05-23 04:00 | LWR | TRIFECTA¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100029 |  | 1 | Event Description: THE ARTICLE, "SEX-RELATED DIFFERENCES IN PATIENT CHARACTERISTICS, PRACTICE PATTERNS AND OUTCOMES AFTER BIOPROSTHETIC AND MECHANICAL AORTIC VALVE REPLACEMENT", WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE, MULTICENTER STUDY THAT EXAMINED SEX-BASED DIFFERENCES IN BASELINE CHARACTERISTICS, PRACTICE PATTERNS, AND MID-TERM SURVIVAL FOLLOWING AORTIC VALVE REPLACEMENT (AVR). DEVICES INCLUDED IN THIS STUDY WERE SJM MASTERS, SJM REGENT, MEDTRONIC ATS OPEN PIVOT, CYROLIFE ON-X, LIVANOVA CARBOMEDICS, MEDTRONIC HANCOCK II, ABBOTT EPIC, ABBOTT TRIFECTA GT, LIVANOVA CROWN PRT, LIVANOVA SOLOSMART, AND LIVANOVA PERCEVAL S. THE ARTICLE CONCLUDED THAT AT THE TIME OF REFERRAL FOR AVR, FEMALE PATIENTS WERE SIGNIFICANTLY OLDER AND HAD WORSE CLINICAL PROFILES THAN MALE PATIENTS. DESPITE THE HIGHER RATES OF BIOPROSTHETIC VALVE IMPLANTATION AND PPM IN FEMALES, MID-TERM SURVIVAL RATES WERE NOT SIGNIFICANTLY DIFFERENT. IN CONTRAST, MALES, ESPECIALLY THOSE UNDER 65, SHOWED HIGHER MORTALITY FOLLOWING BIOPROSTHETIC AVR. THESE FINDINGS UNDERSCORE THE NEED FOR FURTHER RESEARCH FOCUSING ON THE SEX-BASED TREATMENT DETERMINATES IN AVR. [THE PRIMARY AND CORRESPONDING AUTHOR WAS MILOS MATKOVIC, DEPARTMENT FOR CARDIAC SURGERY, CLINICAL CENTER OF SERBIA, KOSTE TODOROVICA 8, 11000 BELGRADE, SERBIA, WITH CORRESPONDING EMAIL: DR.MATKO@HOTMAIL.COM]. THE TIME FRAME OF THE STUDY WAS FROM (B)(6) 2015 TO (B)(6) 2020. A TOTAL OF 1249 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 803 (64.3%) RECEIVED AN ABBOTT DEVICE. THE MAJORITY GENDER WAS MALE. IN THE MALE GROUP, THE AVERAGE AGE WAS 64.4 YEARS. IN THE FEMALE GROUP, THE AVERAGE AGE WAS 67.2 YEARS. COMORBIDITIES INCLUDED HYPERTENSION, HYPERLIPIDEMIA, CHRONIC OBSTRUCTIVE PULMONARY DISORDER, CHRONIC KIDNEY DISEASE, DIABETES, PERIPHERAL VASCULAR DISEASE, PRIOR CEREBROVASCULAR ACCIDENT. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRIFECTA VALVE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, HYPERLIPIDEMIA, CHRONIC OBSTRUCTIVE PULMONARY DISORDER, CHRONIC KIDNEY DISEASE, DIABETES, PERIPHERAL VASCULAR DISEASE, PRIOR CEREBROVASCULAR ACCIDENT.. COMPLICATIONS REPORTED INCLUDED DEATH, SURGICAL INTERVENTION (PACEMAKER), STROKE (CEREBROVASCULAR ACCIDENT); THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: SEX-RELATED DIFFERENCES IN PATIENT CHARACTERISTICS, PRACTICE PATTERNS AND OUTCOMES AFTER BIOPROSTHETIC AND MECHANICAL AORTIC VALVE REPLACEMENT
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22079426&pc=OOB | 2124215-2025-33559 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIOGENIC SHOCK THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22079433&pc=OOB | 2124215-2025-33560 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIAC ARREST THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22079439&pc=MCX | 2124215-2025-33561 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | MCX | ROTAPRO | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIAC ARREST. THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22079441&pc=BZD | 2518422-2025-045304 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-23 04:00 | BZD | REMSTAR HEATED HUMIDIFIER PKG | Degraded | Cancer |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22079474&pc=OZD | 1220648-2025-28695 | 2025-05-01 04:00 | Death | ABIOMED, INC. | 2025-05-23 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT, THE NURSE REPORTING A PLAN TO STOP ANTICOAGULATION DUE TO BLEEDING AND OOZING AT SITES. THE NURSE WAS REMINDED OF RECOMMENDATION TO RESTART ANTICOAGULATION TO MAINTAIN ACT 160-180 ONCE BLEEDING IS UNDER CONTROL. THE PATIENT CONTINUED ON SUPPORT. WHEN PATIENT RETURNED FROM COMPUTED TOMOGRAPHY SCAN TO ASSESS NEURO STATUS, HE WENT INTO CARDIAC ARREST AND RETURN OF SPONTANEOUS CIRCULATION WAS NEVER ACHIEVED. THE PATIENT SUBSEQUENTLY EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA IFU: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080187&pc=LOX | 2124215-2025-33616 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080196&pc=DSQ | 2916596-2025-03083 | 2025-04-29 04:00 | Death | THORATEC CORPORATION | 2025-05-23 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-ORGAN FAILURE AFTER A PNEUMOCOCCI INFECTION. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE WAS WORKING AS EXPECTED FOR THE DURATION OF SUPPORT. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS REPORTED THAT THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE FURTHER PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, PUMP POCKET), MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE HEARTMATE 3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION. THE IFU ALSO PROVIDES SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080203&pc=GCJ | 2955842-2025-21943 | 2025-03-27 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-05-23 04:00 | GCJ | WECK VISTA | Insufficient Information | Pulmonary Embolism; Bowel Perforation | K152663 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED BENIGN HYSTERECTOMY WITH SALPINGECTOMY, THE PATIENT EXPIRED. THE SURGEON REPORTED THAT THERE WAS NO DEVICE MALFUNCTION OR ANY ISSUES DURING THE PROCEDURE. ON POSTOPERATIVE DAY (POD) 1, A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED DUE TO THE PATIENT'S PAIN LEVEL AND TO RULE OUT A PERFORATION. THE CT SCAN RESULTS SHOWED INTRA-ABDOMINAL AIR; HOWEVER, A DEFINITIVE STATEMENT REGARDING PERFORATION OF A HOLLOW ORGAN WAS NOT POSSIBLE WITH CERTAINTY. ON POD 2 THE PATIENT UNDERWENT AN OPEN LAPAROTOMY WHERE AN INJURY TO THE COLON WAS DISCOVERED; A BOWEL RESECTION WITH END-TO-END ANASTOMOSIS WAS PERFORMED. THE SURGEON REPORTED THAT THE INJURY LIKELY OCCURRED DURING PLACEMENT OF THE TROCAR FOR THE FIRST PORT AT THE UMBILICUS, WHICH WAS PERFORMED USING THE VERESS TECHNIQUE WITHOUT VISUALIZATION. THE PATIENT WAS REPORTEDLY OBESE, HAD STRONG ADHESIONS FROM A PAST SURGERY, AND HAD A THICK ABDOMINAL WALL. THE PATIENT EXPIRED LATER THE SAME DAY, WITH THE OFFICIAL CAUSE STATED TO BE A PULMONARY EMBOLISM. NO ADDITIONAL INFORMATION WAS PROVIDED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: SECTION H11 ADDITIONAL INFORMATION: A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE FIVE SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. Manufacturer Narrative: THE FOLLOWING CONCOMITANT DEVICES WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, FENESTRATED BIPOLAR FORCEPS, PROGRASP FORCEPS, AND SYNCHROSEAL. A SITE REVIEW FOUND NO RELATED COMPLAINTS FOR ANY OF THE PRODUCTS USED. A DEVICE HISTORY RECORD REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE SHOWED NO NON-CONFORMANCES WERE IDENTIFIED. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT DIED 2 DAYS FOLLOWING A HYSTERECTOMY WITH SALPINGECTOMY. FREE AIR AND INTRABDOMINAL FLUID WERE NOTED ON A CT SCAN LEADING TO AN EXPLORATION ON POST OP DAY 2 WHERE THE INJURY WAS DISCOVERED. THE SURGEON SPECULATED THE BOWEL INJURY WAS LIKELY THE RESULT OF A TROCAR AT THE INITIAL PORT PLACEMENT. ADDITIONAL INFORMATION SUGGESTING A PULMONARY EMBOLISM ALSO OCCURRED ALTHOUGH NO INFORMATION AS TO HOW THAT WAS DIAGNOSED IS PROVIDED. THE BOWEL INJURY FROM THE PORT PLACEMENT LIKELY PLAYED A MAJOR ROLE IN THIS PATIENT¿S DEATH ALTHOUGH IT'S POSSIBLE THE PULMONARY EMBOLISM DID AS WELL. NO ALLEGATION IS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SECTION A2: ONLY THE PATIENT YEAR OF BIRTH WAS PROVIDED, ESTIMATE OF (B)(6) 1971 USED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080321&pc=MIH | 3011063223-2025-00092 |  | Death | ENDOLOGIX | 2025-05-23 04:00 | MIH | AFX2 | Obstruction of Flow | Obstruction/Occlusion | P040002 |  | 1 | Event Description: THE REPORTED INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH REVIEW OF AN INDUSTRY RELEVANT JOURNAL ARTICLE. DOI: 10.1177/15266028241284364. IT WAS REPORTED THAT 535 PATIENTS WERE ENROLLED FROM 43 ITALIAN AND SPANISH CENTERS BETWEEN NOVEMBER 2019 AND AUGUST 2021. AFX2 BIFURCATED STENT GRAFT SYSTEM WERE IMPLANTED IN PATIENTS TO TREAT ABDOMINAL AORTIC ANEURYSMS (AAAS), PENETRATING AORTIC ULCERS (PAU), OR ISOLATED INFRARENAL AORTIC DISSECTIONS. IT WAS REPORTED WITHIN THE JOURNAL ARTICLE ONE PATIENT EXPERIENCED FIBRINOLYSIS AND WAS TREATED WITH PERCUTANEOUS TRANSLUMINAL ANGIOPLASTY (PTA), AND STENTING. HOWEVER, RESULTED IN DEATH. THE EXACT HOSPITAL, CASE DATE, EVENT DETAILS, LOT NUMBERS, AND CATALOG NUMBERS ARE UNKNOWN. NO ADDITIONAL INFORMATION WILL BECOME AVAILABLE REGARDING THIS REPORT DUE TO THE ABSENCE OF PATIENT IDENTIFIERS. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. THE REPORTED ADVERSE EVENT/INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH AN ONGOING REVIEW OF INDUSTRY RELEVANT JOURNAL ARTICLES WHERE PATIENT RELATED OUTCOMES ARE PRESENTED ANONYMOUSLY, AND PATIENT SPECIFIC DEVICE INFORMATION IS UNAVAILABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE COMPLETED. THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION REMAINS UNKNOWN. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT COULD NOT BE COMPLETED AS THE REPORTED EVENT WAS FOUND DURING PUBLICATION REVIEW WHERE THE PATIENT INFORMATION IS ANONYMOUS. DEVICE, USE, PROCEDURE, AND/OR ANATOMY RELATEDNESS TO THIS ADVERSE EVENT/INCIDENT COULD NOT BE EVALUATED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080369&pc=QEZ | 1319211-2025-00025 | 2025-05-07 04:00 | Death | ANGIODYNAMICS, INC. | 2025-05-23 04:00 | QEZ | ALPHAVAC | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Cardiac Arrest | K240397 |  | 1 | Event Description: AN ANGIODYNAMICS TERRITORY MANAGER (TM) REPORTED THAT A PATIENT, WITH SEVERAL COMORBIDITIES, EXPIRED DURING A PULMONARY EMBOLISM (PE) REMOVAL PROCEDURE, AFTER USE OF AN ALPHAVAC MULTIPURPOSE MECHANICAL ASPIRATOR F18 C85. THE PROCEDURE WAS PERFORMD IN ACCORDANCE WITH THE DEVICE'S IFU AND ALL APPEARED NORMAL ACCORDING TO THE PHYSICIANS; HOWEVER, THE ALPHAVAC WAS REMOVED TO INSPECT IF THE CANNULA OR SHEATH WERE OBSTRUCTED; NEITHER WERE, SO THE GORE DRYSEAL WAS REMOVED. AFTER REMOVAL OF THE DRYSEAL, THE PATIENT'S PRESSURES BEGAN TO DROP AND THEY DECOMPENSATED [CARDIAC ISSUE]. AT THIS TIME, AN ACLS PROTOCOL WAS ACTIVATED. THE PHYSICIANS CHECKED THE PATIENT FOR A PLEURAL EFFUSION, PERICARDIAL EFFUSION AND HEMATOMA; HOWEVER, NONE WERE FOUND. AFTER SEVERAL RESUSCITATION ATTEMPTS, THE PATIENT ULTIMATELY EXPIRED. THIS PROCEDURE WAS THE FIRST ALPHAVAC PROCEDURE FOR BOTH ATTENDING PHYSICIANS; HOWEVER, ONE OF THE ATTENDINGS IS A REGULAR ANGIOVAC USER. THE PHYSICIANS DO NOT BELIEVE THAT ANY ANGIODYNAMICS DEVICE CONTRIBUTED, IN ANY WAY, TO THE PATIENT'S EXPIRATION. Manufacturer Narrative: THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF THE PATIENT DECOMPENSATED [CARDIAC ISSUES] AND EXPIRED AFTER REMOVAL OF THE ALPHAVAC CANNULA/SHEATH AND GORE DRYSEAL SHEATH CANNOT BE CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THIS SERIOUS ADVERSE EVENT. NO ALPHAVAC PRODUCT WAS RETURNED TO ANGIODYNAMICS FOR EVALUATION SINCE THERE WAS NO REPORT OF DEVICE MALFUNCTION OR PERFORMANCE ISSUE DURING USE IN THE PROCEDURE. THE DEVICE HISTORY RECORD (DHR) REVIEW OF THE PACKAGING/CANNULA/SHEATH LOTS REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. THE PHYSICIANS STATED THEY DO NOT BELIEVE THAT ANY ANGIODYNAMICS' DEVICE CONTRIBUTED, IN ANY WAY, TO THE PATIENT'S EXPIRATION. THERE WAS NO REPORT OF ALPHAVAC DEVICE MALFUNCTION OR USE ERROR DURING THE PROCEDURE AND PATIENT EXPIRATION IS POTENTIALLY DUE TO THEIR CO-MORBIDITIES. PULMONARY EMBOLISM, CARDIAC ARREST AND DEATH ARE LISTED IN THE DEVICE DIRECTIONS FOR USE (DFU) AS POSSIBLE COMPLICATIONS/ADVERSE EVENTS DUE TO USE OF EMBOLECTOMY SYSTEM. LABELING REVIEW: DIRECTIONS FOR USE (10903689) IS PROVIDED WITH THIS DEVICE AND CONTAINS THE FOLLOWING STATEMENTS: INDICATIONS FOR USE THE CANNULA IS INDICATED FOR: THE NON-SURGICAL REMOVAL OF THROMBI OR EMBOLI FROM THE VENOUS VASCULATURE. ASPIRATION OF CONTRAST MEDIA AND OTHER FLUIDS FROM THE VENOUS VASCULATURE. THE CANNULA IS INTENDED FOR USE IN THE VENOUS SYSTEM AND FOR THE TREATMENT OF PULMONARY EMBOLISM. INTENDED USE: THE ALPHAVAC SYSTEM IS INTENDED TO BE USED WITH COMMONLY AVAILABLE VASCULAR ACCESS TOOLS (E.G., GUIDEWIRE, VASCULAR INTRODUCER, ETC.) TO FACILITATE THE REMOVAL OF THROMBOEMBOLI, BUT NOT LIMITED TO, THROMBUS, EMBOLUS, OR CLOT DURING MINIMALLY INVASIVE PERCUTANEOUS PROCEDURES. WARNINGS: SELECTION OF THE PATIENT AS A CANDIDATE FOR USE WITH THIS DEVICE AND FOR SUCH PROCEDURES AS IT IS INTENDED IS THE PHYSICIANS' SOLE RESPONSIBILITY. THE OUTCOME IS DEPENDENT ON MANY VARIABLES INCLUDING, PATIENT PATHOLOGY, SURGICAL PROCEDURE, AND PROCEDURE/TECHNIQUE. THE BENEFITS OF USE OF THIS DEVICE MUST BE WEIGHED AGAINST THE RISKS INCLUDING RISKS OF SYSTEMIC ANTICOAGULATION AND MUST BE ASSESSED BY THE PRESCRIBING PHYSICIAN. ADVERSE EVENTS: THIS DEVICE, AS DO ALL EMBOLECTOMY SYSTEMS AND DEVICES, HAS POSSIBLE SIDE EFFECTS, WHICH INCLUDE, BUT ARE NOT LIMITED TO INFECTIONS, BLOOD LOSS, THROMBUS FORMATION, EMBOLIC EVENTS, VESSEL, VENTRICULAR OR VALVULAR DAMAGE AND COMPLICATIONS OF PERCUTANEOUS OR SURGICAL INSERTION TECHNIQUES. THESE MAY OCCUR IF THE DIRECTIONS FOR USE ARE NOT FOLLOWED. POSSIBLE COMPLICATIONS INCLUDE THOSE NORMALLY ASSOCIATED WITH LARGE BORE SURGICAL AND/OR PERCUTANEOUS VESSEL CATHETERIZATION/CANNULATION, ANTICOAGULATION, AND APPLICATION OF INTRAVASCULAR INTRODUCER SYSTEMS WHICH INCLUDE BUT ARE NOT LIMITED TO: ARRHYTHMIAS, CARDIAC ARREST, DEATH, DISTAL EMBOLIZATION OF THROMBUS, PULMONARY EMBOLISM, TACHYCARDIA. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT CONTINUES TO BE MONITORED FOR TRENDS. REFERENCE (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080522&pc=NPT | 2025587-2025-03871 | 2024-10-06 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-23 04:00 | NPT | EVOLUT FX DCS | Device Dislodged or Dislocated; Activation Failure; Patient Device Interaction Problem | Hematoma; Low Blood Pressure/ Hypotension; Foreign Body In Patient | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE INVOLVING THE EVOLUTFX-29 IN A PATIENT WITH SEVERE CALCIFICATION IN THE AORTIC ANNULUS, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED 2 MM BELOW THE SMALLEST DIAMETER TO AVOID ANNULUS RUPTURE. UPON THE INITIAL DEPLOYMENT ATTEMPT AT APPROXIMATELY 4 MM BELOW THE VALVE PLANE, THE FIRST THIRD OF THE VALVE WAS SIGNIFICANTLY UNDER-EXPANDED ("TULIP-SHAPED"). THE VALVE WAS RECAPTURED INTO THE DELIVERY CATHETER SYSTEM (DCS), AND ON THE THIRD ATTEMPT, THE VALVE WAS IMPLANTED AND REMAINED AT A DEPTH OF 4MM. HOWEVER, THE VALVE REMAINED UNDER-EXPANDED IN THE INITIAL THIRD AFTER REMOVING THE NOSECONE OF THE DCS. DUE TO THE NARROWED VALVE OUTLET, IT WAS OBSERVED THAT THE NOSECONE PULLED THE VALVE FROM ONE OF THE POSTS, DESPITE THE NOSECONE BEING CORRECTLY WITHDRAWN. A DISLODGEMENT OCCURRED. THE IMPLANTING PHYSICIANS SPENT APPROXIMATELY TWO HOURS TRYING TO "UNTIE" THE VALVE IN DIFFERENT WAYS; HOWEVER, WERE UNSUCCESSFUL, AND A SECOND VALVE WAS DECIDED TO BE IMPLANTED. WHEN THE SECOND VALVE AND DCS WERE INSERTED, THE FIRST VALVE DISLODGED AGAIN. IT WAS ATTEMPTED TO RAISE THE VALVE AGAIN WITH A SNARE, AND THEN DECIDED TO INFLATE IT WITH A BALLOON, HOUSING THE GUIDEWIRE SECURELY IN THE VALVE PLANE. THE BALLOON WAS INFLATED INSIDE THE DISLODGED VALVE WITH THE ATTEMPT TO DIRECT THE SYSTEM TO A MORE FAVORABLE ORIENTATION. ULTIMATELY, THE SECOND VALVE WAS SUCCESSFULLY IMPLANTED. THE LEFT ACCESS SITE, WHICH WAS USED AS THE DCS ACCESS SITE, WAS NEGLECTED DUE TO MANIPULATION, RESULTING IN THE 14FR MEDTRONIC INTRODUCER SHEATH COMING OUT, NECESSITATING THE PLACEMENT OF ANOTHER 14FR INTRODUCER TO LOCATE THE BALLOON. BOTH ACCESS SITES HAD 14FR INTRODUCERS. FOLLOWING IMPLANT, THE LEFT ACCESS SITE DEVELOPED A HEMATOMA, AND THE ACCESS COULD NOT BE CLOSED BY THE IMPLANTING PHYSICIAN. THE PATIENT WAS REFERRED FOR SURGICAL TREATMENT FOR THE ACCESS SITE CLOSURE, AND LEFT THE CATHETERIZATION LABORATORY WITH LOW BLOOD PRESSURE REQUIRING MEDICATION. THE PATIENT SUBSEQUENTLY DEVELOPED SHOCK AND DIED 4 HOURS LATER. PER THE PHYSICIAN, NEITHER THE MEDTRONIC DCS NOR SHEATH CONTRIBUTED TO THE HEMATOMA; HOWEVER, BOTH VALVES AND BOTH DCS CANNOT BE EXCLUDED AS CONTRIBUTING FACTORS TO THE DEATH.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE INVOLVING THE EVOLUTFX-29 IN A PATIENT WITH SEVERE CALCIFICATION IN THE AORTIC ANNULUS, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED 2 MM BELOW THE SMALLEST DIAMETER TO AVOID ANNULUS RUPTURE. UPON THE INITIAL DEPLOYMENT ATTEMPT AT APPROXIMATELY 4 MM BELOW THE VALVE PLANE, THE FIRST THIRD OF THE VALVE WAS SIGNIFICANTLY UNDER-EXPANDED ("TULIP-SHAPED"). THE VALVE WAS RECAPTURED INTO THE DELIVERY CATHETER SYSTEM (DCS), AND ON THE THIRD ATTEMPT, THE VALVE WAS IMPLANTED AND REMAINED AT A DEPTH OF 4MM. HOWEVER, THE VALVE REMAINED UNDER-EXPANDED IN THE INITIAL THIRD AFTER REMOVING THE NOSECONE OF THE DCS. DUE TO THE NARROWED VALVE OUTLET, IT WAS OBSERVED THAT THE NOSECONE PULLED THE VALVE FROM ONE OF THE POSTS, DESPITE THE NOSECONE BEING CORRECTLY WITHDRAWN. A DISLODGEMENT OCCURRED. THE IMPLANTING PHYSICIANS SPENT APPROXIMATELY TWO HOURS TRYING TO "UNTIE" THE VALVE IN DIFFERENT WAYS; HOWEVER, WERE UNSUCCESSFUL, AND A SECOND VALVE WAS DECIDED TO BE IMPLANTED. WHEN THE SECOND VALVE AND DCS WERE INSERTED, THE FIRST VALVE DISLODGED AGAIN. IT WAS ATTEMPTED TO RAISE THE VALVE AGAIN WITH A SNARE, AND THEN DECIDED TO INFLATE IT WITH A BALLOON, HOUSING THE GUIDEWIRE SECURELY IN THE VALVE PLANE. THE BALLOON WAS INFLATED INSIDE THE DISLODGED VALVE WITH THE ATTEMPT TO DIRECT THE SYSTEM TO A MORE FAVORABLE ORIENTATION. ULTIMATELY, THE SECOND VALVE WAS SUCCESSFULLY IMPLANTED. THE LEFTACCESS SITE, WHICH WAS USED AS THE DCS ACCESS SITE, WAS NEGLECTED DUE TO MANIPULATION, RESULTING IN THE 14FR MEDTRONIC INTRODUCER SHEATH COMING OUT, NECESSITATING THE PLACEMENT OF ANOTHER 14FR INTRODUCER TO LOCATE THE BALLOON. BOTH ACCESS SITES HAD 14FR INTRODUCERS. FOLLOWING IMPLANT, THE LEFT ACCESS SITE DEVELOPED A HEMATOMA, AND THE ACCESS COULD NOT BE CLOSED BY THE IMPLANTING PHYSICIAN. THE PATIENT WAS REFERRED FOR SURGICAL TREATMENT FOR THE ACCESS SITE CLOSURE, AND LEFT THE CATHETERIZATION LABORATORY WITH LOW BLOOD PRESSURE REQUIRING MEDICATION. THE PATIENT SUBSEQUENTLY DEVELOPED SHOCK AND DIED 4 HOURS LATER. PER THE PHYSICIAN, NEITHER THE MEDTRONIC DCS NOR SHEATH CONTRIBUTED TO THE HEMATOMA; HOWEVER, BOTH VALVES AND BOTH DCS CANNOT BE EXCLUDED AS CONTRIBUTING FACTORS TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE CAUSE OF DEATH WAS RELATED TO THE LEFT ACCESS SITE INJURY WHICH REQUIRED SURGICAL CLOSURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE LEFT ACCESS SITE WAS USED FOR IMPLANT. A PRE BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED. THE VALVE WAS IMPLANTED INTO THE NATIVE ANNULUS. THE CUSPAL OVERLAP TECHNIQUE WAS USED. TRAINING GUIDANCE FOR SLOWLY DEPLOYMENT WAS FOLLOWED. PRIOR TO SLOWLY WITHDRAWING THE DCS, THE NOSE CONE WAS CENTERED WITHIN THE FRAME INFLOW BY SLIGHTLY RETRACTING THE GUIDEWIRE. THE DCS WAS NOT TWISTED OR TORQUED AT ANY POINT DURING THE DEPLOYMENT. PER THE PHYSICIAN, THE CAUSE OF THE DISLODGEMENT WAS THAT THE FIRST VALVE IMPLANTED CAUSE THE SECOND VALVE TO DISPLACE THE OTHER. PER THE PHYSICIAN, THE DIFFICULTY CLOSING THE FEMORAL ACCESS MAY HAVE IMPACTED THE EVENT. PER THE PHYSICIAN, THERE WAS NOTHING OF NOTE IN TERMS OF PATIENT ANATOMY THAT CONTRIBUTED TO THE DISLODGMENT. Manufacturer Narrative: CONTINUATION OF D10. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME EVOLUT FX VALVE; PRODUCT ID: EVOLUTFX-29 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2024; EXPLANT DATE OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012275343); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID: EVOLUTFX-29 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2024; EXPLANT DATE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. ADDED SECOND PARAGRAPH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080581&pc=FKX | 0008030665-2025-01163 | 2025-05-19 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-05-23 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Dyspnea; Rash; Dizziness | K181108 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) WAS HOSPITALIZED AND EXPIRED ON (B)(6) 2025 APPROXIMATELY 26 MINUTES AFTER CCPD THERAPY BEGAN. FOLLOW-UP WITH THE PATIENT¿S INPATIENT PROGRAM MANAGER (IPM) REVEALED THE PATIENT PRESENTED TO THE EMERGENCY ROOM ON (B)(6) 2025 WITH COMPLAINTS OF DYSPNEA, DIZZINESS, AND A RASH. ALTHOUGH THE TIMELINE/SPECIFICS OF THE SERIOUS ADVERSE EVENTS ARE LARGELY UNKNOWN (PATIENT¿S ELECTRONIC MEDICAL RECORD CLOSED), THE IPM WAS ABLE TO RECALL THE PATIENT WAS DIAGNOSED WITH AN ALMOST COMPLETE ATRIOVENTRICULAR (AV) BLOCK (WHICH WOULD HAVE REQUIRED A PACEMAKER), AND A SEVERE CASE OF ADULT-ONSET HAND/FOOT/MOUTH DISEASE. THE IPM STATED THE PATIENT HAD A DO NOT RESUSCITATE (DNR) ORDER IN PLACE AND PASSED AWAY ON (B)(6) 2025. THE PATIENT WAS UNDERGOING CCPD AT THE TIME OF HER PASSING (26 MINUTES INTO TREATMENT), HOWEVER THE IPM WAS CONFIDENT THE CARDIAC ARREST OCCURRED DUE TO THE PATIENT¿S CARDIAC COMORBID CONDITIONS WHILE THE BODY WAS FIGHTING THE HAND/FOOT/MOUTH INFECTION. PER THE IPM, THE EVENTS WERE UNRELATED TO ANY FRESENIUS PRODUCT(S) AND/OR DEVICE(S) DEFICIENCY OR MALFUNCTION. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINIC REVIEW: BASED ON THE AVAILABLE INFORMATION, THE PATIENT¿S LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM HAVING A CAUSAL OR CONTRIBUTORY ROLE IN THE ADVERSE EVENTS EXPERIENCED BY THE PATIENT. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A SERIOUS INJURY, PATIENT DEATH, OR OTHER SERIOUS ADVERSE EVENT(S) RELATED TO A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) OCCURRED. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080588&pc=NPT | 2025587-2025-03872 | 2024-10-06 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-23 04:00 | NPT | EVOLUT FX DCS | Device Dislodged or Dislocated; Patient Device Interaction Problem | Hematoma; Low Blood Pressure/ Hypotension | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE INVOLVING THE EVOLUTFX-29 IN A PATIENT WITH SEVERE CALCIFICATION IN THE AORTIC ANNULUS, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED 2 MM BELOW THE SMALLEST DIAMETER TO AVOID ANNULUS RUPTURE. UPON THE INITIAL DEPLOYMENT ATTEMPT AT APPROXIMATELY 4 MM BELOW THE VALVE PLANE, THE FIRST THIRD OF THE VALVE WAS SIGNIFICANTLY UNDER-EXPANDED ("TULIP-SHAPED"). THE VALVE WAS RECAPTURED INTO THE DELIVERY CATHETER SYSTEM (DCS), AND ON THE THIRD ATTEMPT, THE VALVE WAS IMPLANTED AND REMAINED AT A DEPTH OF 4MM. HOWEVER, THE VALVE REMAINED UNDER-EXPANDED IN THE INITIAL THIRD AFTER REMOVING THE NOSECONE OF THE DCS. DUE TO THE NARROWED VALVE OUTLET, IT WAS OBSERVED THAT THE NOSECONE PULLED THE VALVE FROM ONE OF THE POSTS, DESPITE THE NOSECONE BEING CORRECTLY WITHDRAWN. A DISLODGEMENT OCCURRED. THE IMPLANTING PHYSICIANS SPENT APPROXIMATELY TWO HOURS TRYING TO "UNTIE" THE VALVE IN DIFFERENT WAYS; HOWEVER, WERE UNSUCCESSFUL, AND A SECOND VALVE WAS DECIDED TO BE IMPLANTED. WHEN THE SECOND VALVE AND DCS WERE INSERTED, THE FIRST VALVE DISLODGED AGAIN. IT WAS ATTEMPTED TO RAISE THE VALVE AGAIN WITH A SNARE, AND THEN DECIDED TO INFLATE IT WITH A BALLOON, HOUSING THE GUIDEWIRE SECURELY IN THE VALVE PLANE. THE BALLOON WAS INFLATED INSIDE THE DISLODGED VALVE WITH THE ATTEMPT TO DIRECT THE SYSTEM TO A MORE FAVORABLE ORIENTATION. ULTIMATELY, THE SECOND VALVE WAS SUCCESSFULLY IMPLANTED. THE LEFT ACCESS SITE, WHICH WAS USED AS THE DCS ACCESS SITE, WAS NEGLECTED DUE TO MANIPULATION, RESULTING IN THE 14FR MEDTRONIC INTRODUCER SHEATH COMING OUT, NECESSITATING THE PLACEMENT OF ANOTHER 14FR INTRODUCER TO LOCATE THE BALLOON. BOTH ACCESS SITES HAD 14FR INTRODUCERS. FOLLOWING IMPLANT, THE LEFT ACCESS SITE DEVELOPED A HEMATOMA, AND THE ACCESS COULD NOT BE CLOSED BY THE IMPLANTING PHYSICIAN. THE PATIENT WAS REFERRED FOR SURGICAL TREATMENT FOR THE ACCESS SITE CLOSURE AND LEFT THE CATHETERIZATION LABORATORY WITH LOW BLOOD PRESSURE REQUIRING MEDICATION. THE PATIENT SUBSEQUENTLY DEVELOPED SHOCK AND DIED 4 HOURS LATER. PER THE PHYSICIAN, NEITHER THE MEDTRONIC DCS NOR SHEATH CONTRIBUTED TO THE HEMATOMA; HOWEVER, BOTH VALVES AND BOTH DCS CANNOT BE EXCLUDED AS CONTRIBUTING FACTORS TO THE DEATH.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER HEART VALVE PROCEDURE INVOLVING THE EVOLUTFX-29 IN A PATIENT WITH SEVERE CALCIFICATION IN THE AORTIC ANNULUS, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED 2 MM BELOW THE SMALLEST DIAMETER TO AVOID ANNULUS RUPTURE. UPON THE INITIAL DEPLOYMENT ATTEMPT AT APPROXIMATELY 4 MM BELOW THE VALVE PLANE, THE FIRST THIRD OF THE VALVE WAS SIGNIFICANTLY UNDER-EXPANDED ("TULIP-SHAPED"). THE VALVE WAS RECAPTURED INTO THE DELIVERY CATHETER SYSTEM (DCS), AND ON THE THIRD ATTEMPT, THE VALVE WAS IMPLANTED AND REMAINED AT A DEPTH OF 4MM. HOWEVER, THE VALVE REMAINED UNDER-EXPANDED IN THE INITIAL THIRD AFTER REMOVING THE NOSECONE OF THE DCS. DUE TO THE NARROWED VALVE OUTLET, IT WAS OBSERVED THAT THE NOSECONE PULLED THE VALVE FROM ONE OF THE POSTS, DESPITE THE NOSECONE BEING CORRECTLY WITHDRAWN. A DISLODGEMENT OCCURRED. THE IMPLANTING PHYSICIANS SPENT APPROXIMATELY TWO HOURS TRYING TO "UNTIE" THE VALVE IN DIFFERENT WAYS; HOWEVER, WERE UNSUCCESSFUL, AND A SECOND VALVE WAS DECIDED TO BE IMPLANTED. WHEN THE SECOND VALVE AND DCS WERE INSERTED, THE FIRST VALVE DISLODGED AGAIN. IT WAS ATTEMPTED TO RAISE THE VALVE AGAIN WITH A SNARE, AND THEN DECIDED TO INFLATE IT WITH A BALLOON, HOUSING THE GUIDEWIRE SECURELY IN THE VALVE PLANE. THE BALLOON WAS INFLATED INSIDE THE DISLODGED VALVE WITH THE ATTEMPT TO DIRECT THE SYSTEM TO A MORE FAVORABLE ORIENTATION. ULTIMATELY, THE SECOND VALVE WAS SUCCESSFULLY IMPLANTED. THE LEFTACCESS SITE, WHICH WAS USED AS THE DCS ACCESS SITE, WAS NEGLECTED DUE TO MANIPULATION, RESULTING IN THE 14FR MEDTRONIC INTRODUCER SHEATH COMING OUT, NECESSITATING THE PLACEMENT OF ANOTHER 14FR INTRODUCER TO LOCATE THE BALLOON. BOTH ACCESS SITES HAD 14FR INTRODUCERS. FOLLOWING IMPLANT, THE LEFT ACCESS SITE DEVELOPED A HEMATOMA, AND THE ACCESS COULD NOT BE CLOSED BY THE IMPLANTING PHYSICIAN. THE PATIENT WAS REFERRED FOR SURGICAL TREATMENT FOR THE ACCESS SITE CLOSURE, AND LEFT THE CATHETERIZATION LABORATORY WITH LOW BLOOD PRESSURE REQUIRING MEDICATION. THE PATIENT SUBSEQUENTLY DEVELOPED SHOCK AND DIED 4 HOURS LATER. PER THE PHYSICIAN, NEITHER THE MEDTRONIC DCS NOR SHEATH CONTRIBUTED TO THE HEMATOMA; HOWEVER, BOTH VALVES AND BOTH DCS CANNOT BE EXCLUDED AS CONTRIBUTING FACTORS TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE CAUSE OF DEATH WAS RELATED TO THE LEFT ACCESS SITE INJURY WHICH REQUIRED SURGICAL CLOSURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE LEFT ACCESS SITE WAS USED FOR IMPLANT. A PRE BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED. THE VALVE WAS IMPLANTED INTO THE NATIVE ANNULUS. THE CUSPAL OVERLAP TECHNIQUE WAS USED. TRAINING GUIDANCE FOR SLOWLY DEPLOYMENT WAS FOLLOWED. PRIOR TO SLOWLY WITHDRAWING THE DCS, THE NOSE CONE WAS CENTERED WITHIN THE FRAME INFLOW BY SLIGHTLY RETRACTING THE GUIDEWIRE. THE DCS WAS NOT TWISTED OR TORQUED AT ANY POINT DURING THE DEPLOYMENT. PER THE PHYSICIAN, THE CAUSE OF THE DISLODGEMENT WAS THAT THE FIRST VALVE IMPLANTED CAUSE THE SECOND VALVE TO DISPLACE THE OTHER. PER THE PHYSICIAN, THE DIFFICULTY CLOSING THE FEMORAL ACCESS MAY HAVE IMPACTED THE EVENT. PER THE PHYSICIAN, THERE WAS NOTHING OF NOTE IN TERMS OF PATIENT ANATOMY THAT CONTRIBUTED TO THE DISLODGMENT. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME EVOLUT FX VALVE; PRODUCT ID EVOLUTFX-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2024; OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012275343); PRODUCT TYPE: 0195-HEART VALVES; NON-IMPLANTABLE DEVICE PRODUCT ID EVOLUTFX-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2024 SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. ADDED SECOND PARAGRAPH. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080662&pc=LOX | 2124215-2025-33621 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH DURING PROCEDURE DECEASED,ACUTE MYOCARDIAL INFARCTION THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080663&pc=LOX | 2124215-2025-33622 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED,ACUTE MYOCARDIAL INFARCTION CARDIAC ARREST THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080664&pc=LOX | 2124215-2025-33625 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiac Tamponade; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED,ACUTE MYOCARDIAL INFARCTION HEART FAILURE,CARDIAC TAMPONADE,CARDIOGENIC SHOCK THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080666&pc=LOX | 2124215-2025-33643 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED,ACUTE MYOCARDIAL INFARCTION CARDIAC ARREST THE FOLLOWING EVENT, BASED ON DATA FROM THE CATH PCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATH PCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATH PCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. (B)(6) OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080667&pc=LOX | 2124215-2025-33634 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED,ACUTE MYOCARDIAL INFARCTION HEART FAILURE,CARDIAC ARREST,CARDIAC ARREST,MYOCARDIAL INFARCTION,CARDIOGENIC SHOCK. THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. (B)(6) OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080674&pc=LOX | 2124215-2025-33629 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED, DEATH AT DISCHARGE, DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIAC ARREST. THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080676&pc=LOX | 2124215-2025-33641 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: AGENT PAS REGISTRY IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED,ACUTE MYOCARDIAL INFARCTION HEART FAILURE,CARDIAC ARREST,CARDIAC ARREST,MYOCARDIAL INFARCTION,CARDIOGENIC SHOCK THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080678&pc=LOX | 2124215-2025-33633 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED,ACUTE MYOCARDIAL INFARCTION CARDIAC ARREST THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080682&pc=LOX | 2124215-2025-33628 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED,ACUTE MYOCARDIAL INFARCTION HEART FAILURE,CARDIAC ARREST,CARDIAC ARREST,MYOCARDIAL INFARCTION,CARDIOGENIC SHOCK THE FOLLOWING EVENT, BASED ON DATA FROM THE CATH PCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATH PCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATH PCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. (B)(6) OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080684&pc=LOX | 2124215-2025-33626 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH DURING PROCEDURE DECEASED,ACUTE MYOCARDIAL INFARCTION THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080687&pc=LOX | 2124215-2025-33617 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiac Tamponade; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED,ACUTE MYOCARDIAL INFARCTION HEART FAILURE,CARDIAC TAMPONADE,CARDIOGENIC SHOCK THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080689&pc=LOX | 2124215-2025-33640 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: AGENT PAS REGISTRY. IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION CARDIAC ARREST THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080690&pc=NIQ | 2124215-2025-33605 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | NIQ | SYNERGY MEGATRON | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiogenic Shock |  |  | 1 | Event Description: AGENT PAS REGISTRY. IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION CARDIOGENIC SHOCK THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080905&pc=NKM | 2135147-2025-02770 | 2010-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-05-23 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100009 |  | 1 | Event Description: THE ARTICLE "TRANSCATHETER EDGE-TO-EDGE REPAIR IN ANATOMICALLY COMPLEX DEGENERATIVE MITRAL REGURGITATION (DMR): 3-YEAR OUTCOMES FROM A REAL-WORLD REGISTRY" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE OUTCOMES BY SPECIFIC ANATOMICAL CRITERIA IN PATIENTS WITH SEVERE DMR AND COMPLEX VALVE ANATOMY ENROLLED IN THE PROSPECTIVE MITRAL REGISTRY. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT IN PATIENTS WITH ANATOMICALLY COMPLEX DMR, ADVANCES IN PROCEDURAL TECHNIQUES FOR MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR (M-TEER) HAVE ALLOWED SUCCESSFUL TREATMENT WITH SUSTAINED MR REDUCTION AND IMPROVED LONG-TERM SURVIVAL. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS NICOLETA NITA AT UNIVERSITY MEDICAL CENTER, ULM, GERMANY WITH CORRESPONDING EMAIL *NICOLETA.NITA@UNIKLINIK-ULM.DE*.] THE TIME FRAME OF THE STUDY WAS JANUARY 2010 AND JUNE 2021. A TOTAL OF 304 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 80 YEARS AND MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED HEART FAILURE HOSPITALIZATION, PRIOR CARDIAC SURGICAL INTERVENTION, PRIOR CARDIAC TRANSCATHETER INTERVENTION, CORONARY ARTERY DISEASE, MYOCARDIAL INFARCTION, HYPERTENSION, DIABETES, ATRIAL FIBRILLATION, PERIPHERAL ARTERY DISEASE, COPD, CHRONIC RENAL FAILURE, PREVIOUS CANCER, PACEMAKER, DEGENERATIVE MITRAL REGURGITATION. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HOSPITALIZATION, MAJOR ADVERSE CARDIAC AND CEREBROVASCULAR EVENTS (CARDIOVASCULAR DEATH, STROKE, MYOCARDIAL INFARCTION AND SEVERE BLEEDING), REINTERVENTION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HEART FAILURE HOSPITALIZATION, PRIOR CARDIAC SURGICAL INTERVENTION, PRIOR CARDIAC TRANSCATHETER INTERVENTION, CORONARY ARTERY DISEASE, MYOCARDIAL INFARCTION, HYPERTENSION, DIABETES, ATRIAL FIBRILLATION, PERIPHERAL ARTERY DISEASE, COPD, CHRONIC RENAL FAILURE, PREVIOUS CANCER, PACEMAKER, DEGENERATIVE MITRAL REGURGITATION. COMPLICATIONS REPORTED INCLUDED DEATH, HOSPITALIZATION, MAJOR ADVERSE CARDIAC AND CEREBROVASCULAR EVENTS (CARDIOVASCULAR DEATH, STROKE, MYOCARDIAL INFARCTION AND SEVERE BLEEDING), REINTERVENTION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: AS THIS IS FROM A LITERATURE REVIEW, THE DEATH DATE IS ESTIMATED. B3: AS THIS IS FROM A LITERATURE REVIEW, THE EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: TRANSCATHETER EDGE-TO-EDGE REPAIR IN ANATOMICALLY COMPLEX DEGENERATIVE MITRAL REGURGITATION: 3-YEAR OUTCOMES FROM A REAL-WORLD REGISTRY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080949&pc=PNJ | 2017865-2025-64663 | 2025-05-05 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-23 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT ATRIUM | Failure to Capture; Adverse Event Without Identified Device or Use Problem; Interrogation Problem | Insufficient Information | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC. IT WAS NOTED THAT THE PATIENT'S LEADLESS PACEMAKER EXPERIENCED DIFFICULTY TO BE INTERROGATED VIA CONDUCTIVE TELEMETRY. INTERFERENCE WAS BELIEVED TO BE THE CAUSE OF THE ISSUE. LOSS OF CAPTURE WAS NOTED. THE PATIENT HAD DECEASED DUE TO UNKNOWN CAUSES. THERE WERE NO ALLEGATIONS THAT THE PATIENT'S DEATH WAS RELATED TO THEIR LEADLESS PACEMAKER SYSTEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22080950&pc=OZD | 1220648-2025-28707 | 2025-01-22 05:00 | Death | ABIOMED, INC. | 2025-05-23 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Unspecified Kidney or Urinary Problem | P140003 |  | 1 | Event Description: IT WAS REPORTED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY THAT THE PATIENT EXPERIENCED ACCESS SITE BLEEDING AND WORSENING KIDNEY INJURY DURING IMPELLA CP SUPPORT. THE BLEED WAS BELIEVED TO HAVE A PROBABLY RELATIONSHIP TO IMPELLA USE WHILE THE KIDNEY INJURY WAS THOUGHT TO HAVE A POSSIBLE RELATIONSHIP TO IMPELLA SUPPORT. THE PATIENT RECEIVED TWO UNITS OF PACKED RED BLOOD CELLS AND 1 UNIT OF CRYO IN RESPONSE TO THE BLEED. THE PATIENT'S ACIDOSIS WAS MANAGED WITH SODIUM BICARBONATE PUSHED AND HIGH PRESSORS. THE PATIENT UNFORTUNATELY DID NOT RECOVER AND EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081071&pc=OOB | 2124215-2025-33719 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH DURING PROCEDURE DECEASED, ACUTE MYOCARDIAL INFARCTION THE FOLLOWING EVENT, BASED ON DATA FROM THE CATH PCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATH PCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATH PCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. (B)(6) OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081072&pc=NIQ | 2124215-2025-33702 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | NIQ | SYNERGY MEGATRON | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION HEART FAILURE, CARDIAC ARREST, CARDIAC ARREST, MYOCARDIAL INFARCTION, CARDIOGENIC SHOCK. THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081075&pc=OOB | 2124215-2025-33716 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiac Tamponade; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED,ACUTE MYOCARDIAL INFARCTION HEART FAILURE,CARDIAC TAMPONADE,CARDIOGENIC SHOCK THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081080&pc=OOB | 2124215-2025-33711 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081085&pc=OOB | 2124215-2025-33730 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION CARDIAC ARREST. THE FOLLOWING EVENT, BASED ON DATA FROM THE CATH PCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATH PCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATH PCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. (B)(6) OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081086&pc=OOB | 2124215-2025-33731 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION CARDIAC ARREST. THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081087&pc=LOX | 2124215-2025-33606 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED,ACUTE MYOCARDIAL INFARCTION CARDIAC ARREST THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081089&pc=OOB | 2124215-2025-33738 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED,ACUTE MYOCARDIAL INFARCTION CARDIAC ARREST THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081091&pc=OOB | 2124215-2025-33712 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED,ACUTE MYOCARDIAL INFARCTION CARDIAC ARREST THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081092&pc=OOB | 2124215-2025-33739 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED, DEATH AT DISCHARGE, DECEASED, ACUTE MYOCARDIAL INFARCTION, HEART FAILURE, CARDIAC ARREST, CARDIAC ARREST, MYOCARDIAL, INFARCTION, CARDIOGENIC SHOCK. THE FOLLOWING EVENT, BASED ON DATA FROM THE CATHPCI REGISTRY, IS BEING REPORTED AS A POSSIBLE DUPLICATE OF AN EVENT THAT WAS ALREADY KNOWN TO BSC AND REPORTED AS AN MDR WHEN THE EVENT OCCURRED. BSC REPORTS THE EVENTS IN COMPLIANCE TO 21 CFR 803.3. BECAUSE CATHPCI REGISTRY DATA IS DE-IDENTIFIED BEFORE BEING TRANSMITTED TO BSC, THERE ARE SIGNIFICANT LIMITATIONS TO BSC'S ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED TO BSC. INITIAL REPORTER IS NOT PROVIDED DUE TO THE TERMS OF THE CATHPCI REGISTRY. Manufacturer Narrative: DUE TO SAFE HARBOR REIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081168&pc=DRA | 2135147-2025-02772 | 2025-05-01 04:00 | Death | ABBOTT MEDICAL | 2025-05-23 04:00 | DRA | STEERABLE GUIDE CATHETER | Adverse Event Without Identified Device or Use Problem | Perforation; Pericardial Effusion | K172394 |  | 1 | Event Description: IT WAS REPORTED THIS WAS A MITRACLIP PROCEDURE TO TREAT FUNCTINOAL MITRAL REGURGITATION (MR) WITH A GRADE OF 4+. ONE CLIP WAS PREPARED AND ADVANCED INTO THE ANATOMY. THE CLIP WAS DEPLOYED, REDUCING MR TO <1. IT WAS LATER FOUND THAT THE PATIENT HAD A SEPTAL INJURY AND PERICARDIAL EFFUSION. ULTIMATELY, THE PATIENT PASSED AWAY THE SAME DAY.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A MITRACLIP DEVICE WAS SELECTED FOR IMPLANT TO TREAT FUNCTIONAL MITRAL REGURGITATION (MR) WITH A GRADE OF 4+. ONE CLIP WAS PREPARED AND ADVANCED INTO THE ANATOMY. THE CLIP WAS DEPLOYED, REDUCING MR TO <1. AFTER THE CLIP WAS RELEASED, IT WAS FOUND THAT THE PATIENT HAD SEPTAL INJURY AND PERICARDIAL EFFUSION. IT WAS NOTED THAT THE PATIENT HAD SYMPTOMS OF LOW BLOOD PRESSURE. A PERICARDIOCENTESIS WAS PERFORMED. IT WAS NOTED THAT SGC AND MITRACLIP CONTRIBUTED TO THE PERICARDIAL EFFUSION. THE PATIENT PASSED AWAY THE SAME DAY IN CRITICAL CARE UNIT (CCU). IT WAS NOTED THAT THE PATIENT DID NOT HAVE PRE-EXISTING CONDITIONS THAT CAUSED OR CONTRIBUTED TO THE DEATH OF THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: IN THIS CASE, THERE WAS NO FUNCTIONAL TESTING THAT WAS NEEDED FOR THE RETURNED DEVICE AS THERE WAS NO REPORTED DEVICE MALFUNCTION ASSOCIATED WITH THE SGC. A REVIEW OF THE LOT HISTORY RECORD IDENTIFIED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE INFORMATION REVIEWED, A CAUSE FOR THE REPORTED DEATH/EXPIRED, PERFORATION, AND PERICARDIAL EFFUSION CANNOT BE DETERMINED. DEATH/EXPIRED, PERFORATION, AND PERICARDIAL EFFUSION ARE LISTED IN THE INSTRUCTIONS FOR USE (IFU) AS A KNOWN POSSIBLE COMPLICATION ASSOCIATED WITH MITRACLIP PROCEDURES. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081193&pc=LWS | 2124215-2025-33754 | 2025-01-09 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-23 04:00 | LWS | ENDOTAK RELIANCE G | Inappropriate/Inadequate Shock/Stimulation; High Capture Threshold | Low Blood Pressure/ Hypotension; Respiratory Failure; Shock from Patient Lead(s) | P910073/S043 |  | 1 | Event Description: IT WAS REPORTED THAT THIS RIGHT VENTRICULAR (RV) LEAD WAS EXPLANTED DUE TO AN UNKNOWN PRODUCT ANOMALY. OTHER THAN THE PROCEDURE, NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, THIS RV LEAD HAS NOT BEEN RETURNED TO BOSTON SCIENTIFIC FOR FURTHER ANALYSIS. ADDITIONAL INFORMATION WAS RECEIVED. THE LOCAL SALES REPRESENTATIVE CONTACTED THE HOSPITAL AND LEARNED THAT THE RV LEAD WAS PARTIALLY EXPLANTED DUE TO HIGH PACING THRESHOLD MEASUREMENTS AND BECAUSE IT CONTRIBUTED TO INAPPROPRIATE SHOCKS BEING DELIVERED TO THE PATIENT. THE PATIENT'S CHRONIC NON-BOSTON SCIENTIFIC RIGHT ATRIAL (RA) LEAD WAS ALSO EXPLANTED AND THE PATIENT WAS IMPLANTED WITH A NEW, NON-BOSTON SCIENTIFIC CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM. HOWEVER, ACUTE RESPIRATORY FAILURE AND HYPOTENSION WERE REPORTED AND WERE BELIEVED TO BE ASSOCIATED WITH THE EXTRACTION AND UPGRADE PROCEDURE. THE PATIENT SUBSEQUENTLY DIED TWO WEEKS AFTER THE PROCEDURE. IT WAS NOTED THE HOSPITAL WAS IN POSSESSION OF THE EXPLANTED PRODUCTS AND IT WAS UNKNOWN IF THEY WOULD BE RETURNING THE RV LEAD TO BOSTON SCIENTIFIC FOR ANALYSIS. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED PRIOR TO IMPLEMENTATION OF THE UDI REGULATION AND AS A RESULT, THE COMPLETE UDI IS NOT AVAILABLE. APPLICABLE US PRODUCT DATA HAS BEEN INCLUDED IN THIS REPORT. IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081267&pc=NKM | 2135147-2025-02775 | 2025-05-01 04:00 | Death | ABBOTT MEDICAL | 2025-05-23 04:00 | NKM | MITRACLIP | Adverse Event Without Identified Device or Use Problem | Hematoma; Pericardial Effusion; Unspecified Tissue Injury | P100009 |  | 1 | Event Description: IT WAS REPORTED THIS WAS A MITRACLIP PROCEDURE TO TREAT DEGENERATIVE MITRAL REGURGITATION (MR) WITH A GRADE OF 4, A SMALL ATRIUM AND A FLAIL. THREE CLIPS WERE SUCCESSFULLY IMPLANTED. TO FURTHER REDUCE MR, AN ADDITIONAL CLIP WAS INSERTED. IT WAS NOTED AT SOME POINT DURING STEERING AND DEPLOYMENT, TISSUE DAMAGE OCCURRED. THE CLIP WAS THEN IMPLANTED, REDUCING MR TO A GRADE OF <1. UPON REMOVING THE STEERABLE GUIDE CATHETER (SGC), A PERICARDIAL EFFUSION WAS OBSERVED AND PERICARDIOCENTESIS WAS PERFORMED. HOWEVER, THE PERICARDIAL EFFUSION WAS NOT RESOLVED AND THE PATIENT WAS TRANSFERRED TO THE OR, WHERE THE EFFUSION WAS STOPPED. THERE WAS NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE.Event Description: IT WAS REPORTED THIS WAS A MITRACLIP PROCEDURE TO TREAT DEGENERATIVE MITRAL REGURGITATION (MR) WITH A GRADE OF 4, A SMALL ATRIUM AND A FLAIL. THREE CLIPS WERE SUCCESSFULLY IMPLANTED. TO FURTHER REDUCE MR, AN ADDITIONAL CLIP WAS INSERTED. IT WAS NOTED AT SOME POINT DURING STEERING AND DEPLOYMENT, TISSUE DAMAGE OCCURRED. THE CLIP WAS THEN IMPLANTED, REDUCING MR TO A GRADE OF <1. UPON REMOVING THE STEERABLE GUIDE CATHETER (SGC), A PERICARDIAL EFFUSION WAS OBSERVED AND PERICARDIOCENTESIS WAS PERFORMED. HOWEVER, THE PERICARDIAL EFFUSION WAS NOT RESOLVED AND THE PATIENT WAS TRANSFERRED TO THE OPERATING ROOM, WHERE THE EFFUSION WAS STOPPED. THERE WAS NO CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE. ADDITIONAL INFORMATION RECEIVED ON 27 MAY 2025, THAT THE PATIENT PASSED AWAY AFTER SURGERY ON (B)(6) 2025. THE CAUSE OF DEATH WAS NOTED TO HAVE BEEN RELATED TO RIGHT SIDED HEART FAILURE ASSOCIATED WITH A PERICARDIAL TAP THAT PUNCTURED A VESSEL ON THE RIGHT SIDE OF THE HEART. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR THE REPORTED HEMATOMA, TISSUE INJURY, AND PERICARDIAL EFFUSION. THE CAUSE OF DEATH WAS RIGHT SIDED HEART FAILURE ASSOCIATED WITH THE PERICARDIAL TAP THAT PUNCTURED A VESSEL ON THE RIGHT HEART. HEMATOMA, TISSUE INJURY, DEATH, AND PERICARDIAL EFFUSION ARE LISTED IN THE INSTRUCTIONS FOR USE IS A KNOWN POSSIBLE COMPLICATION ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED HOSPITALIZATION, UNEXPECTED MEDICAL INTERVENTION, AND SURGICAL INTERVENTION WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081283&pc=OMP | 8043484-2025-00056 | 2024-09-07 04:00 | Death | SMITH & NEPHEW MEDICAL LTD. | 2025-05-23 04:00 | OMP | RENASYS TOUCH NON CONNECT 4TH ED DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: ON (B)(6) 2024, A PHYSICIAN TREATED AN ABDOMINAL WOUND OF A HOSPICE PATIENT WITH THE DEVICE. THE PATIENT DIED IN THE DAYS AFTER THE TREATMENT. ACCORDING TO THE PHYSICIAN, A CAREGIVER WHO WAS PRESENT DURING THE TREATMENT HAS RAISED CONCERNS VIS-A-VIS THE GERMAN INVESTIGATING AUTHORITIES AS TO WHETHER THE WOUND THERAPY WAS INDICATED/PROPERLY CARRIED OUT AND THE DEVICE WAS PROPERLY FUNCTIONING. AT THIS POINT IN TIME, THE MANUFACTURER IS NOT AWARE OF ANY PROBLEM WITH THE DEVICE NOR THAT THE DEATH OCCURRED IN CONNECTION WITH THE USE OF THE DEVICE. THE NOTIFICATION IS MADE AS A PURELY PRECAUTIONARY MEASURE. Manufacturer Narrative: H3, H6: THE DEVICE WAS RETURNED FOR EVALUATION IN THE SERVICE CENTER. THE VISUAL AND FUNCTIONAL EVALUATION DID NOT REVEAL A PHYSICAL OR FUNCTIONAL FAILURE. THE EXACT CAUSE OF THE PATIENT'S DEATH REMAINS UNKNOWN DUE TO INSUFFICIENT INFORMATION, THOUGH POSSIBLE FACTORS INCLUDE A POST-PROCEDURAL COMPLICATION OR PRE-EXISTING MEDICAL CONDITIONS REQUIRING HOSPICE CARE. SPECIFIC DETAILS RELATED TO THE PATIENT'S ABDOMINAL WOUND WERE NOT PROVIDED; HOWEVER, THE PHYSICIAN ORDERED NEGATIVE PRESSURE WOUND THERAPY (NPWT) AND WAS PRESENT DURING ITS INITIAL APPLICATION. AS THE CLINICIAN USER MANUAL FOR THE RENASYS TOUCH, INSTRUCTS PRIOR TO PLACEMENT, "THE MEDICAL PROFESSIONAL TREATING THE WOUND MUST ASSESS HOW TO BEST USE THE SYSTEM FOR AN INDIVIDUAL WOUND. IT IS IMPORTANT TO CAREFULLY ASSESS BOTH THE WOUND AND THE PATIENT TO ENSURE THE CLINICAL INDICATIONS FOR NPWT ARE MET". THE ATTENDING PHYSICIAN STATED THAT THERE WAS NO CONNECTION BETWEEN THE RENASYS DEVICE AND THE PATIENT¿S DEATH. A DEATH CERTIFICATE HAS YET TO BE PROVIDED, LIMITING FURTHER ASSESSMENT, BUT THERE IS NO INDICATION THAT THE PATIENT¿S DEATH WAS RELATED TO THE SMITH & NEPHEW DEVICE. ANY FURTHER IMPACT ON THE PATIENT BEYOND WHAT HAS BEEN REPORTED COULD NOT BE ASSESSED. THE REVIEW OF THE DEVICE'S PRODUCTION RECORDS SHOWED NO MANUFACTURING ISSUES THAT COULD HAVE CONTRIBUTED TO THE REPORTED INCIDENT. NO SIMILAR EVENTS WERE FOUND IN THE COMPLAINT HISTORY FOR THIS PART NUMBER DURING THE LAST 14 MONTHS. THERE HAVE BEEN NO PREVIOUS ESCALATED ACTIONS FOR THE PART NUMBER AND OUTCOME REPORTED. AT THIS TIME, WE HAVE NO REASON TO SUSPECT THAT THE PRODUCT FAILED TO MEET ANY PRODUCT SPECIFICATIONS AT THE TIME OF MANUFACTURE. BASED ON THIS INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED. THE REPORTED DEVICE WILL BE RETAINED BY THE MANUFACTURER. Manufacturer Narrative: INTERNAL COMPLAINT REFERENCE: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081374&pc=NKM | 2135147-2025-02781 | 2025-05-01 04:00 | Death | ABBOTT MEDICAL | 2025-05-23 04:00 | NKM | MITRACLIP | Adverse Event Without Identified Device or Use Problem | Perforation; Pericardial Effusion | P100009 |  | 1 | Event Description: IT WAS REPORTED THIS WAS A MITRACLIP PROCEDURE TO TREAT FUNCTINOAL MITRAL REGURGITATION (MR) WITH A GRADE OF 4+. ONE CLIP WAS PREPARED AND ADVANCED INTO THE ANATOMY. THE CLIP WAS DEPLOYED, REDUCING MR TO <1. IT WAS LATER FOUND THAT THE PATIENT HAD A SEPTAL INJURY AND PERICARDIAL EFFUSION. ULTIMATELY, THE PATIENT PASSED AWAY THE SAME DAY.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A MITRACLIP DEVICE WAS SELECTED FOR IMPLANT TO TREAT FUNCTIONAL MITRAL REGURGITATION (MR) WITH A GRADE OF 4+. ONE CLIP WAS PREPARED AND ADVANCED INTO THE ANATOMY. THE CLIP WAS DEPLOYED, REDUCING MR TO <1. AFTER THE CLIP WAS RELEASED, IT WAS FOUND THAT THE PATIENT HAD SEPTAL INJURY AND PERICARDIAL EFFUSION. IT WAS NOTED THAT THE PATIENT HAD SYMPTOMS OF LOW BLOOD PRESSURE. A PERICARDIOCENTESIS WAS PERFORMED. IT WAS NOTED THAT SGC AND MITRACLIP CONTRIBUTED TO THE PERICARDIAL EFFUSION. THE PATIENT PASSED AWAY THE SAME DAY IN CRITICAL CARE UNIT (CCU). IT WAS NOTED THAT THE PATIENT DID NOT HAVE PRE-EXISTING CONDITIONS THAT CAUSED OR CONTRIBUTED TO THE DEATH OF THE PATIENT. NO ADDITIONAL INFORMATION PROVIDED. Manufacturer Narrative: IN THIS CASE, THERE WAS NO FUNCTIONAL TESTING THAT WAS NEEDED FOR THE RETURNED DEVICE AS THERE WAS NO REPORTED DEVICE MALFUNCTION ASSOCIATED WITH THE CDS. A REVIEW OF THE LOT HISTORY RECORD IDENTIFIED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE INFORMATION REVIEWED, A CAUSE FOR THE REPORTED DEATH/EXPIRED, PERFORATION, AND PERICARDIAL EFFUSION CANNOT BE DETERMINED. DEATH/EXPIRED, PERFORATION, AND PERICARDIAL EFFUSION ARE LISTED IN THE INSTRUCTIONS FOR USE (IFU) AS A KNOWN POSSIBLE COMPLICATION ASSOCIATED WITH MITRACLIP PROCEDURES. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081837&pc=NIK | 2017865-2025-64689 | 2025-05-07 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-23 04:00 | NIK | GALLANT HF | Failure to Convert Rhythm | Arrhythmia | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD EXPERIENCED VENTRICULAR FIBRILLATION, AND THE IMPLANTED SYSTEM INCLUDING THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR AND THE RIGHT VENTRICULAR LEAD FAILED TO CONVERT THE RHYTHM, DESPITE AT THE MAX OUTPUT. THE PATIENT PASSED AWAY. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081838&pc=NVY | 2017865-2025-64690 | 2025-05-07 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-23 04:00 | NVY | OPTISURE ACTIVE FIXATION, DF-4 CONNECTOR | Failure to Convert Rhythm | Arrhythmia | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD EXPERIENCED VENTRICULAR FIBRILLATION, AND THE IMPLANTED SYSTEM INCLUDING THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR AND THE RIGHT VENTRICULAR LEAD FAILED TO CONVERT THE RHYTHM, DESPITE AT THE MAX OUTPUT. THE PATIENT PASSED AWAY. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22081839&pc=MKJ | 0003015876-2025-01074 | 2025-04-26 04:00 | Death | PHYSIO-CONTROL, INC. - 3015876 | 2025-05-23 04:00 | MKJ | LIFEPAK® CR2 DEFIBRILLATOR | Failure to Analyze Signal | No Clinical Signs, Symptoms or Conditions | P170018 |  | 1 | Event Description: THE CUSTOMER CONTACTED STRYKER TO REPORT THEIR DEVICE WOULD NOT ANALYZE THE RHYTHM. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THE PATIENT INVOLVED DID NOT SURVIVE HOWEVER IT IS UNKNOWN IF THE DEVICE CAUSED OR CONTRIBUTED. ANOTHER DEVICE WAS USED.Event Description: THE CUSTOMER CONTACTED STRYKER TO REPORT THEIR DEVICE WOULD NOT ANALYZE THE RHYTHM. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THE PATIENT INVOLVED DID NOT SURVIVE HOWEVER IT IS UNKNOWN IF THE DEVICE CAUSED OR CONTRIBUTED. ANOTHER DEVICE WAS USED. Manufacturer Narrative: STRYKER CONTINUES TO INVESTIGATE THE REPORTED FAILURE AND WILL SUBMIT A SUPPLEMENTAL REPORT ON THIS EVENT TO THE FDA AS PROVIDED BY 21 CFR 803.56. THE CUSTOMER HAS RECEIVED A REPLACEMENT DEVICE. A CLINICAL REVIEW WAS COMPLETED AND THERE IS INSUFFICIENT INFORMATION AVAILABLE TO DETERMINE IF THE DEVICE CONTRIBUTED TO THE PATIENT OUTCOME. Manufacturer Narrative: STRYKER EVALUATED THE DEVICE AND WAS ABLE TO VERIFY THE REPORTED ISSUE THROUGH REVIEW OF THE SOFTWARE LOGS. INVESTIGATION FOUND THE ELECTRODE TRAY WAS NOT DEPLOYED PROPERLY, THE RED LOOPS WERE NOT PULLED ON THE ELECTRODE PADS. USE ERROR WAS DETERMINED TO BE THE CAUSE OF THE REPORTED ISSUE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22067478&pc=OZP | 2032227-2025-179116 | 2025-04-22 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-22 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hyperglycemia; High Blood Pressure/ Hypertension; Diabetic Ketoacidosis | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER WAS DECEASED DUE TO CARDIAC ARREST. THE EVENT INVOLVED PRODUCT(S)MMT-1884, MMT-332A AND UNOMEDICAL. TROUBLESHOOTING WAS PERFORMED AND REPORTING PARTY REPORTED THAT CUSTOMER WAS HOSPITALIZED DUE TO HIGH BLOOD SUGAR AND DIABETIC KETOACIDOSIS. CUSTOMER ALSO HAD OTHER HEALTH ILLNESS SUCH AS CHRONIC KIDNEY DISEASE, HYPERTENSION, AND HIGH CHOLESTEROL. IT WAS UNKNOWN WHETHER THE CUSTOMER USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF REPORTED EVENT AND CUSTOMER WAS USING THE AUTO MODE/SMART GUARD FEATURE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-1884, MMT-332A AND UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22068850&pc=OZD | 1220648-2025-28629 | 2025-04-17 04:00 | Death | ABIOMED, INC | 2025-05-22 04:00 | OZD | IMPELLA CP | Failure to Advance; Patient-Device Incompatibility | Cardiovascular Insufficiency; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT IMPLANTATION OF AN IMPELLA CP WAS ATTEMPTED BUT DUE TO CALCIFICATION OF THE LOWER LIMBS PLACEMENT OF THE PUMP WAS ABORTED.Event Description: PER THE MEDICAL SAFETY OFFICER REVIEW THE DEVICE WAS UNABLE TO BE INSERTED; CALCIFICATION LOWER LIMB. DELAY IN/INABILITY TO DELIVERY MCS VERY LIKELY DUE TO PATIENT'S UNDERLINE ANATOMY. HOWEVER, THERE REMAINS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR TO THE OUTCOME OF DEATH. Manufacturer Narrative: B1 SELECTION OF ADVERSE EVENT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. B2 SELECTIONS DEATH AND DATE OR DEATH WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. B5 REVISED TO INCLUDE MSO STATEMENT AND THE PATIENT OUTCOME THAT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. H1 TYPE OF REPORT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. H6 CODES 4582 AND 2199 WERE ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. NEW CODED ADDED TO HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION SINCE THE PATIENT HAD CALCIFICATION IN THE LOWER LIMBS WAS SO SEVERE THAT BOTH INSERTIONS WERE DIFFICULT. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22069584&pc=CAW | 2518422-2025-045233 | 2021-08-31 04:00 | Death | RESPIRONICS, INC. | 2025-05-22 04:00 | CAW | EVERFLO INTL OPI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K061261 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22069616&pc=QEW | MW5170705 | 2023-09-04 04:00 | Death | PENUMBRA INC. | 2025-05-22 04:00 | QEW | SEP12, SEPARATOR 12 | Adverse Event Without Identified Device or Use Problem | Perforation |  |  | 1 | Event Description: (B)(6). (B)(6) INFORMED OF A SERIOUS PRODUCT COMPLICATION THAT OCCURRED (B)(6) 2023 -- INVOLVING PENUMBRA'S SEP12 SEPARATOR AND LIGHTNING12 CATHETER. AN 84-YEAR-OLD PATIENT (B)(6) SUFFERED A FATAL PULMONARY ARTERY PERFORATION DURING THE PROCEDURE. THE INTERVENTIONAL CARDIOLOGIST PERFORMING THE CASE CLEARLY INDICATED THAT THE PENUMBRA DEVICE WAS RESPONSIBLE FOR THE PERFORATION. THIS WAS SUPPORTED BY ANGIOGRAPHIC IMAGING CAPTURED DURING THE CASE. A (B)(6) EXPLICITLY ATTRIBUTED THE PATIENT'S DEATH TO PENUMBRA'S DEVICE. (B)(6). CONCERNED ABOUT HOW TO PROCEED AND FEARING THE POSSIBILITY OF MISHANDLING THE SITUATION, (B)(6) FOR GUIDANCE. RATHER THAN ASSISTING ME WITH (B)(6) REPORTING PROTOCOLS, (B)(6) INSTRUCTED ME NOT TO REPORT THE INCIDENT. (B)(4). (B)(6) DISMISSED THE PHYSICIAN'S CREDIBILITY, DESPITE THEIR STRONG CLINICAL REPUTATION AND EXTENSIVE EXPERIENCE. (B)(6). THIS INTERACTION IS DOCUMENTED IN TIMELINE EVIDENCE AND CONTEMPORANEOUS MESSAGES. (B)(6). THIS EVENT REFLECTS A BROADER PATTERN AT PENUMBRA IN WHICH (B)(6) WERE DISCOURAGED FROM REPORTING COMPLICATIONS. (B)(6). PENUMBRA'S FAILURE TO INVESTIGATE OR REPORT THIS FATALITY DEMONSTRATES A SYSTEMIC BREAKDOWN IN POST-MARKET SURVEILLANCE, NONCOMPLIANCE WITH FDA ADVERSE EVENT REPORTING OBLIGATIONS, AND A DELIBERATE EFFORT TO CONCEAL SAFETY RISKS. THIS PLACED FUTURE PATIENTS IN DANGER, ERODED TRUST AMONG HEALTHCARE PROVIDERS, AND UNDERMINED REGULATORY OVERSIGHT. I AM SUBMITTING THIS REPORT NOW BECAUSE I BELIEVE THE COMPANY'S (B)(6) HANDLING OF THE INCIDENT WAS NOT ONLY INADEQUATE -- IT WAS NEGLIGENT, (B)(6), AND IN DIRECT VIOLATION OF PATIENT SAFETY STANDARDS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22070329&pc=MKJ | 3004123209-2025-00155 | 2025-05-09 04:00 | Death | HEARTSINE TECHNOLOGIES LTD | 2025-05-22 04:00 | MKJ | UNKNOWN_MEDICAL - BELFAST_PRODUCT | Failure to Sense | No Clinical Signs, Symptoms or Conditions | P160008 |  | 1 | Event Description: THE CUSTOMER CONTACTED HEARTSINE TO REPORT THAT THEIR DEVICE WOULD NOT RECOGNIZE A PATIENT ATTACHED TO THE PADS DURING A CARDIAC EVENT. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THIS PATIENT IN THIS EVENT IS DECEASED. A CLINICAL REVIEW WILL BE COMPLETED TO DETERMINE DEVICE CONTRIBUTION. Manufacturer Narrative: HEARTSINE HAS REQUESTED THE RETURN OF THE DEVICE FOR INVESTIGATION. UPON COMPLETION, THE CONCLUSIONS WILL BE SUBMITTED IN A FOLLOW-UP REPORT. THE HDF 3500 DEVICE IS NOT AVAILABLE FOR DISTRIBUTION IN THE UNITED STATES BUT IS SIMILAR TO THE SAM 350P AND 450P WHICH IS DISTRIBUTED IN THE UNITED STATES. HEARTSINE CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. NO RESPONSE HAS BEEN RECEIVED FROM THE CUSTOMER. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22070350&pc=NIK | 9614453-2025-02030 | 2025-05-13 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-05-22 04:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Failure to Deliver Shock/Stimulation; Device Sensing Problem | Arrhythmia | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT A FEW DAYS AFTER THE IMPLANT OF THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM AND AN ATRIO VENTRICULAR NODE ABLATION, THE PATIENT PASSED AWAY. ON THAT DAY, IT WAS OBSERVED THAT THE CRT-D HAD INAPPROPRIATELY DETECTED THE ARRHYTHMIA AS NON-SUSTAINED VENTRICULAR TACHYCARDIA RESULTING DELAYED AND INAPPROPRIATELY WITHHELD THERAPY. IT WAS NOTED THAT THE CRT-D HAD INITIALLY DETECTED AND TREATED THE ARRHYTHMIA APPROPRIATELY. IT WAS FURTHER REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) FOR ELEVATED SENSING INTEGRITY COUNTER (SIC). Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22070365&pc=LWS | 2649622-2025-13937 | 2025-05-13 04:00 | Death | MPRI | 2025-05-22 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing | Arrhythmia | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT A FEW DAYS AFTER THE IMPLANT OF THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) SYSTEM AND AN ATRIO VENTRICULAR NODE ABLATION, THE PATIENT PASSED AWAY. ON THAT DAY, IT WAS OBSERVED THAT THE CRT-D HAD INAPPROPRIATELY DETECTED THE ARRHYTHMIA AS NON-SUSTAINED VENTRICULAR TACHYCARDIA RESULTING DELAYED AND INAPPROPRIATELY WITHHELD THERAPY. IT WAS NOTED THAT THE CRT-D HAD INITIALLY DETECTED AND TREATED THE ARRHYTHMIA APPROPRIATELY. IT WAS FURTHER REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) FOR ELEVATED SENSING INTEGRITY COUNTER (SIC). Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED OVERSENSING DUE TO NON-PHYSIOLOGIC SIGNALS/SENSING INTEGRITY COUNTER. ANALYSIS OF THE DEVICE MEMORY INDICATED THE CRITERIA FOR THE RIGHT VENTRICULAR LEAD INTEGRITY ALERT WERE MET. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22070438&pc=LWS | 2124215-2025-33028 | 2025-04-29 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-22 04:00 | LWS | RESONATE HF ICD DR | Failure to Convert Rhythm; Inaccurate Synchronization | Ventricular Fibrillation | P960040 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) PATIENT PASSED AWAY ON (B)(6) 2025 WHILE IN A PARKING LOT. THE PATIENT RECEIVED THERAPY ON THE DAY OF DEATH. IN THE FIRST EPISODE, A TOTAL OF 4 SHOCKS WERE DELIVERED, WITH THE FOURTH SHOCK APPEARING TO CONVERT THE VENTRICLES. TECHNICAL SERVICES REVIEW OF THE SECOND EPISODE INDICATED THAT THE FIRST SHOCKS LOOKED APPROPRIATE FOR WHAT APPEARED TO BE A DUAL TACHYCARDIA. THE PATIENT ENTERED RAPID VENTRICULAR RESPONSE WHICH THEN ACCELERATED BACK INTO DUAL TACHYCARDIA. ATTEMPT 5 (41 JOULE SHOCK) ACCELERATED THE PATIENT FROM VENTRICULAR TACHYCARDIA (VT) INTO VENTRICULAR FIBRILLATION (VF). A SUBSEQUENT SHOCK THEN CONVERTED THE RHYTHM. THE PATIENT'S ARRHYTHMIA THEN RETURNED, WITH ADDITIONAL SHOCKS DELIVERED. THE SECOND-TO-LAST SHOCK AGAIN ACCELERATED THE PATIENT FROM VT TO VF AND THE FINAL SHOCK DID NOT CONVERT THE PATIENT. ICD THERAPY WAS THEN EXHAUSTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22070483&pc=NGV | 2135147-2025-02687 | 2022-10-31 04:00 | Death | ABBOTT MEDICAL | 2025-05-22 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Endocarditis; Cardiac Enzyme Elevation; Fatigue; Fever; Unspecified Infection | P200049 |  | 1 | Event Description: THE ARTICLE, "INFECTIVE ENDOCARDITIS OF A LEFT ATRIAL APPENDAGE CLOSURE DEVICE: A CASE REPORT AND LITERATURE REVIEW" WAS REVIEWED. IT WAS REPORTED THAT ON AN UNKNOWN DATE, A AMPLATZER AMULET WAS IMPLANTED IN A PATIENT WITH ARTERIAL HYPERTENSION, PULMONARY HYPERTENSION, SEVERE TRICUSPID REGURGITATION, PERMANENT AF, AND RECURRENT BLEEDING ON ORAL ANTICOAGULATION. 20 MONTHS LATER, THE PATIENT PRESENTED WITH FEVER AND FATIGUE LASTING FOR TWO DAYS ACCOMPANIED BY CHILLS AND DECREASED ALERTNESS. ON ADMISSION, THE INITIAL WORK-UP REVEALED AN ELEVATED WHITE BLOOD CELL COUNT, C-REACTIVE PROTEIN LEVEL, AND HIGH PROCALCITONIN LEVEL. GALL BLADDER WALL THICKENING WITHOUT CHOLECYSTOLITHIASIS WAS ALSO NOTED. BLOOD CULTURES WERE PERFORMED AND SHOWED STAPHYLOCOCCUS AUREUS, FOR WHICH INTRAVENOUS ANTIBIOTIC THERAPY WAS STARTED. A TRANSESOPHAGEAL ECHOCARDIOGRAPHY WAS THEN PERFORMED WHICH SHOWED A LAA CLOSURE DEVICE¿ASSOCIATED MOBILE, ECHO-DENSE MASS, CONSISTENT WITH INFECTIOUS VEGETATION. LAAO ENDOCARDITIS WAS DIAGNOSED AS A RESULT. WITHIN 3 DAYS OF ADMISSION, THE PATIENTS CLINICAL CONDITION IMPROVED, THE WHITE BLOOD CELL COUNT NORMALIZED, THE CRP AND PROCALCITONIN BLOOD LEVELS DECLINED, BUT DID NOT NORMALIZE. REPEATED SETS OF BLOOD CULTURES OBTAINED 5 DAYS AFTER INITIATING THE ANTIBIOTICS WERE NEGATIVE. IT WAS RECOMMENDED THE COMPLETE REMOVAL OF THE DEVICE, BUT THE PATIENT REFUSED. THE PATIENT SUBSEQUENTLY DIED TWO WEEKS LATER AS A RESULT OF PROGRESSIVE ENDOCARDITIS AND MULTIPLE PRE-EXISTING CO-MORBIDITIES. (B)(6). Manufacturer Narrative: AS REPORTED IN A RESEARCH ARTICLE, AN AMPLATZER AMULET WAS IMPLANTED IN A PATIENT WITH ARTERIAL HYPERTENSION, PULMONARY HYPERTENSION, SEVERE TRICUSPID REGURGITATION, PERMANENT AF, AND RECURRENT BLEEDING ON ORAL ANTICOAGULATION. 20 MONTHS LATER, THE PATIENT PRESENTED WITH FEVER AND FATIGUE LASTING FOR TWO DAYS ACCOMPANIED BY CHILLS AND DECREASED ALERTNESS. ON ADMISSION, THE INITIAL WORK-UP REVEALED AN ELEVATED WHITE BLOOD CELL COUNT, C-REACTIVE PROTEIN LEVEL, AND HIGH PROCALCITONIN LEVEL. LAAO ENDOCARDITIS WAS DIAGNOSED AS A RESULT. THE COMPLETE REMOVAL OF THE DEVICE WAS RECOMMENDED, BUT THE PATIENT REFUSED. THE PATIENT SUBSEQUENTLY EXPIRED TWO WEEKS LATER AS A RESULT OF PROGRESSIVE ENDOCARDITIS AND MULTIPLE PRE-EXISTING CO-MORBIDITIES. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED PATIENT EFFECTS OF FEVER, FATIGUE, ELEVATED WHITE BLOOD CELL COUNT, AND INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. B2 DEATH DATE: DEATH DATE IS ESTIMATED. B3 EVENT DATE: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22070611&pc=FSA | 9681684-2025-00043 | 2025-04-23 04:00 | Death | ARJOHUNTLEIGH MAGOG INC. | 2025-05-22 04:00 | FSA | MAXI SKY 2 | Detachment of Device or Device Component | Bone Fracture(s); Hematoma |  |  | 1 | Event Description: ARJO BECAME AWARE OF THE PATIENT FALL FROM THE SLING USED WITH MAXI SKY 2 CEILING LIFT. ACCORDING TO THE INFORMATION PROVIDED BY THE CUSTOMER, THE SPREADER BAR GOT HOOKED ON THE LIFTING POLE OF THE BED, WHICH CAUSED IT TO TILT AND LED TO THE 2 POINTS OF THE SLING SLIPPED OUT AND THE PATIENT FELL ON THE FLOOR. AS A RESULT, A PATIENT SUSTAINED THE SUBDURAL HEMATOMA AND FRACTURED SHOULDER. ON 13 MAY, WE HAVE LEARNT THAT THIS SUBSEQUENTLY RESULTED IN THE PATIENT'S DEATH. CEILING LIFT AND SLING EVALUATION REVEALED NO FAULTS. Manufacturer Narrative: THE INVESTIGATION IS PENDING COMPLETION. ONCE IT IS COMPLETED, THE SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: IT WAS REPORTED THAT THE PATIENT FELL FROM THE SLING USED WITH THE MAXI SKY 2 CEILING LIFT. DURING THE TRANSFER, THE RIGHT SIDE OF THE SPREADER BAR TOUCHED THE BED POLE. THE SPREADER BAR MOVED INTO A VERTICAL POSITION, WHICH ALLEGEDLY CAUSED TWO ATTACHMENT POINTS OF THE SLING TO SLIP OUT OF THE SPREADER BAR HOOK, RESULTING IN A PATIENT FALLING. THE PATIENT SUSTAINED A SUBDURAL HEMATOMA AND A FRACTURE OF THE LEFT SHOULDER. ALMOST 3 WEEKS LATER ARJO WAS INFORMED THAT THE PATIENT HAD PASSED AWAY. THE LIFT AND THE SLING REMAINED IN USE AFTER THE INCIDENT. THE DEVICE WAS INSPECTED AND NO MALFUNCTION WAS DETECTED. THE DEVICE OPERATED CORRECTLY. THE 2-HOOK SPREADER BAR IS A COMPONENT USED TO ATTACH A LOOP SLING FOR A RESIDENT TRANSFER. THE SPREADER BAR CONSISTS OF A SINGLE ROD WITH ROUNDED HOOKS ON BOTH ENDS, WITH SECURED LOCKING LATCHES, DESIGNED TO HOLD THE SLING LOOPS IN PLACE. THE PHOTOGRAPHIC EVIDENCES PROVIDED BY THE CUSTOMER WERE REVIEWED AND FOUND THAT IT IS UNLIKELY THAT THE LOOPS DETACH FROM THE HOOK AS THE PATIENT WEIGHT AND GRAVITY WOULD KEEP THE LOOP IN PLACE, EVEN THOUGH THE LOCKING MECHANISM WAS IN AN OPEN POSITION. ADDITIONAL SIMULATION PERFORMED USING RANDOM SAMPLE, AT THE MANUFACTURING SITE, SHOWED THAT THE DETACHMENT IS ALSO UNLIKELY FOR THE SECOND HOOK, NOT ONLY BECAUSE OF GRAVITY, BUT ALSO DUE TO A BUILT-IN LOCKING FEATURE. THE LOCKING MECHANISM WAS STILL IN CLOSED POSITION. THEREFORE, IT HAS BEEN SUGGESTED THAT THE PATIENT COULD SLIDE OUT OF THE SLING, AS THE VERTICAL POSITION OF THE SPREADER BAR IS NOT STABLE AND ANY UNEXPECTED PATIENT MOVEMENTS FURTHER INCREASES INSTABILITY, RAISING THE LIKELIHOOD OF THE PATIENT SLIDING OUT OF THE SLING. THE HYPOTHESIS IS THAT THE PATIENT SLID OUT OF THE SLING FIRST CAUSING THE LOOP TO LEAVE THE SPREADER BAR HOOK. MAXI SKY 2 CEILING LIFT INSTRUCTION FOR USE (IFU) GUIDES THROUGH PROPER TRANSFER USING THE LIFT. THE INSTRUCTION STATES TO ¿BEFORE TRANSFERRING A PATIENT, MAKE SURE THERE ARE NO OBSTACLES IN THE TRANSFER PATH. INJURIES COULD OCCUR.¿ BASED ON THE AVAILABLE INFORMATION AND THE ANALYSIS PROVIDED BY THE MANUFACTURER, IT IS POSSIBLE TO DETERMINE THE ROOT CAUSE OF THE SPREADER BAR MOVING TO A VERTICAL POSITION, AS RELATED TO THE OBSTACLE. ACCORDING TO IFU THE CAREGIVER SHALL ENSURE THERE IS NO OBSTACLE ON THE LIFT PATH. FROM THAT PERSPECTIVE THIS SITUATION IS RELATED TO A HANDLING ERROR. IT IS UNLIKELY THAT THE LOOPS DETACH FROM THE HOOK, EVEN IF THE SPREADER BAR IS IN VERTICAL POSITION. THIS INCIDENT IS A SINGULAR OCCURRENCE. ARJO DEVICE WAS USED FOR A PATIENT TREATMENT WHEN THE EVENT OCCURRED AND FROM THAT PERSPECTIVE IT PLAYED A ROLE IN THE EVENT. THE DEVICE WAS PERFORMING AS INTENDED. NO MALFUNCTION WAS FOUND DURING THE DEVICE EVALUATION. THIS ADVERSE EVENT WAS REPORTED DUE TO CUSTOMER ALLEGATION OF LOOP DETACHMENT AND PATIENT FALLING FROM THE SLING. Manufacturer Narrative: THE COLLECTING INFORMATION IS ONGOING. ADDITIONAL INFORMATION WILL BE PROVIDED UPON INVESTIGATION CONCLUSION. UDI: (B)(4). THE DEVICE IS DISTRIBUTED OUTSIDE THE US AND NO GUDID INFORMATION EXISTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22070612&pc=FSA | 1419652-2025-00026 | 2025-04-23 04:00 | Death | ARJOHUNTLEIGH MAGOG INC. | 2025-05-22 04:00 | FSA | MAXI SKY 2 | Detachment of Device or Device Component | Bone Fracture(s); Hematoma |  |  | 1 | Event Description: ARJO BECAME AWARE OF THE PATIENT FALL FROM THE SLING USED WITH MAXI SKY 2 CEILING LIFT. ACCORDING TO THE INFORMATION PROVIDED BY THE CUSTOMER, THE SPREADER BAR GOT HOOKED ON THE LIFTING POLE OF THE BED, WHICH CAUSED IT TO TILT AND LED TO THE 2 POINTS OF THE SLING SLIPPED OUT AND THE PATIENT FELL ON THE FLOOR. AS A RESULT, A PATIENT SUSTAINED THE SUBDURAL HEMATOMA AND FRACTURED SHOULDER. ON 13 MAY, WE HAVE LEARNT THAT THIS SUBSEQUENTLY RESULTED IN THE PATIENT'S DEATH. CEILING LIFT AND SLING EVALUATION REVEALED NO FAULTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22070937&pc=BZD | 2518422-2025-105978 | 2025-04-08 04:00 | Death | RESPIRONICS, INC. | 2025-05-22 04:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22070979&pc=OZD | 1220648-2025-28648 | 2025-05-01 04:00 | Death | ABIOMED, INC. | 2025-05-22 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ON IMPELLA 5.5 SUPPORT HAD A HEAD CT SHOWING MULTIPLE AREAS OF INFARCTIONS. NEUROSURGEON WAS CONSULTED FOR EMERGENT CRANIOTOMY, BUT THE FAMILY DECIDED TO WITHDRAW CARE, AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22071010&pc=OZD | 1220648-2025-28649 | 2025-04-29 04:00 | Death | ABIOMED, INC. | 2025-05-22 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS), AND DURING THE APPROXIMATELY 14-HOUR MCS, THE PATIENT EXPERIENCED BLEEDING ISSUES AT THE PUNCTURE SIDE. HEMOSTASIS WAS ACHIEVED WITH FEM STOP AND THE LEG WAS NOTED WELL PERFUSED. IT IS UNKNOWN IF BLOOD PRODUCTS WERE ADMINISTERED AND/OR IF ANTICOAGULATION WAS SUSPENDED. FURTHER DETAILS AROUND THE BLEEDING EVENT WERE UNNOTED. HOWEVER, SOME TIME THEREAFTER THE BLEEDING EVENT (SAME DAY), THE PATIENT WAS INDICATED TO HAVE EXPIRED WHILE ON SUPPORT. NO DETAILS SURROUNDING THE DEMISE OUTCOME WERE PROVIDED. Manufacturer Narrative: GIVEN THE IMPELLA-RELATED EVENT WITH INSUFFICIENT DETAILS AND DEMISE OUTCOME, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP DEVICE USED AS AN ASSOCIATED FACTOR IN THE OVERALL OUTCOME. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22071155&pc=MIH | 3011063223-2025-00073 |  | Death | ENDOLOGIX | 2025-05-22 04:00 | MIH | AFX | Insufficient Information | Aneurysm | P040002 |  | 1 | Event Description: THE REPORTED INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH REVIEW OF AN INDUSTRY RELEVANT JOURNAL ARTICLE. THERE ARE TWO JOURNAL ARTICLES THAT ARE THE SAME SUBSET OF PATIENTS. DOI.ORG/10.1177/15266028241255541. DOI.ORG/10.1177/15266028251317910. IT WAS REPORTED THAT 872 PATIENTS WERE INITIALLY IMPLANTED IN THE ITALY WITH THE AFX BIFURCATED STENT GRAFT SYSTEM (DEVICE ITERATION; AFX STRATA, DURAPLY, AFX2 ARE UNKNOWN) BETWEEN 1994 AND 2020. IT WAS REPORTED WITHIN THE JOURNAL ARTICLE EIGHT (8) PATIENTS EXPERIENCED AAA (ABDOMINAL AORTIC ANEURYSM) RELATED DEATHS. THE EXACT HOSPITAL, CASE DATE, EVENT DETAILS, LOT NUMBERS, AND CATALOG NUMBERS ARE UNKNOWN. NO ADDITIONAL INFORMATION WILL BECOME AVAILABLE REGARDING THIS INITIAL/FINAL REPORT DUE TO THE ABSENCE OF PATIENT IDENTIFIERS. NOTE: AS THE EXACT DATE OF INCIDENT IS UNKNOWN, THE BEST ESTIMATED DATE WAS SELECTED. THIS REPORT ADDRESSES PATIENT 1 OF 8. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. THE REPORTED ADVERSE EVENT/INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH AN ONGOING REVIEW OF INDUSTRY RELEVANT JOURNAL ARTICLES WHERE PATIENT RELATED OUTCOMES ARE PRESENTED ANONYMOUSLY, AND PATIENT SPECIFIC DEVICE INFORMATION IS UNAVAILABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE COMPLETED. THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION REMAINS UNKNOWN. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT COULD NOT BE COMPLETED AS THE REPORTED EVENT WAS FOUND DURING PUBLICATION REVIEW WHERE THE PATIENT INFORMATION IS ANONYMOUS. DEVICE, USE, PROCEDURE, AND/OR ANATOMY RELATEDNESS TO THIS ADVERSE EVENT/INCIDENT COULD NOT BE EVALUATED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22071250&pc=MIH | 3011063223-2025-00075 |  | Death | ENDOLOGIX | 2025-05-22 04:00 | MIH | AFX | Insufficient Information | Aneurysm | P040002 |  | 1 | Event Description: THE REPORTED INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH REVIEW OF AN INDUSTRY RELEVANT JOURNAL ARTICLE. THERE ARE TWO JOURNAL ARTICLES THAT ARE THE SAME SUBSET OF PATIENTS. DOI.ORG/10.1177/15266028241255541. DOI.ORG/10.1177/15266028251317910. IT WAS REPORTED THAT 872 PATIENTS WERE INITIALLY IMPLANTED IN THE ITALY WITH THE AFX BIFURCATED STENT GRAFT SYSTEM (DEVICE ITERATION; AFX STRATA, DURAPLY, AFX2 ARE UNKNOWN) BETWEEN 1994 AND 2020. IT WAS REPORTED WITHIN THE JOURNAL ARTICLE EIGHT (8) PATIENTS EXPERIENCED AAA (ABDOMINAL AORTIC ANEURYSM) RELATED DEATHS. THE EXACT HOSPITAL, CASE DATE, EVENT DETAILS, LOT NUMBERS, AND CATALOG NUMBERS ARE UNKNOWN. NO ADDITIONAL INFORMATION WILL BECOME AVAILABLE REGARDING THIS INITIAL/FINAL REPORT DUE TO THE ABSENCE OF PATIENT IDENTIFIERS. NOTE: AS THE EXACT DATE OF INCIDENT IS UNKNOWN, THE BEST ESTIMATED DATE WAS SELECTED. THIS REPORT ADDRESSES PATIENT 2 OF 8. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. THE REPORTED ADVERSE EVENT/INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH AN ONGOING REVIEW OF INDUSTRY RELEVANT JOURNAL ARTICLES WHERE PATIENT RELATED OUTCOMES ARE PRESENTED ANONYMOUSLY, AND PATIENT SPECIFIC DEVICE INFORMATION IS UNAVAILABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE COMPLETED. THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION REMAINS UNKNOWN. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT COULD NOT BE COMPLETED AS THE REPORTED EVENT WAS FOUND DURING PUBLICATION REVIEW WHERE THE PATIENT INFORMATION IS ANONYMOUS. DEVICE, USE, PROCEDURE, AND/OR ANATOMY RELATEDNESS TO THIS ADVERSE EVENT/INCIDENT COULD NOT BE EVALUATED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22071342&pc=MIH | 3011063223-2025-00086 |  | Death | ENDOLOGIX | 2025-05-22 04:00 | MIH | AFX | Insufficient Information | Aneurysm | P040002 |  | 1 | Event Description: THE REPORTED INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH REVIEW OF AN INDUSTRY RELEVANT JOURNAL ARTICLE. THERE ARE TWO JOURNAL ARTICLES THAT ARE THE SAME SUBSET OF PATIENTS. DOI.ORG/10.1177/15266028241255541. DOI.ORG/10.1177/15266028251317910 . IT WAS REPORTED THAT 872 PATIENTS WERE INITIALLY IMPLANTED IN THE ITALY WITH THE AFX BIFURCATED STENT GRAFT SYSTEM (DEVICE ITERATION; AFX STRATA, DURAPLY, AFX2 ARE UNKNOWN) BETWEEN 1994 AND 2020. IT WAS REPORTED WITHIN THE JOURNAL ARTICLE EIGHT (8) PATIENTS EXPERIENCED AAA (ABDOMINAL AORTIC ANEURYSM) RELATED DEATHS. THE EXACT HOSPITAL, CASE DATE, EVENT DETAILS, LOT NUMBERS, AND CATALOG NUMBERS ARE UNKNOWN. NO ADDITIONAL INFORMATION WILL BECOME AVAILABLE REGARDING THIS INITIAL/FINAL REPORT DUE TO THE ABSENCE OF PATIENT IDENTIFIERS. THIS REPORT ADDRESSES PATIENT 3 OF 8. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. THE REPORTED ADVERSE EVENT/INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH AN ONGOING REVIEW OF INDUSTRY RELEVANT JOURNAL ARTICLES WHERE PATIENT RELATED OUTCOMES ARE PRESENTED ANONYMOUSLY, AND PATIENT SPECIFIC DEVICE INFORMATION IS UNAVAILABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE COMPLETED. THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION REMAINS UNKNOWN. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT COULD NOT BE COMPLETED AS THE REPORTED EVENT WAS FOUND DURING PUBLICATION REVIEW WHERE THE PATIENT INFORMATION IS ANONYMOUS. DEVICE, USE, PROCEDURE, AND/OR ANATOMY RELATEDNESS TO THIS ADVERSE EVENT/INCIDENT COULD NOT BE EVALUATED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22071360&pc=MIH | 3011063223-2025-00087 |  | Death | ENDOLOGIX | 2025-05-22 04:00 | MIH | AFX | Insufficient Information | Aneurysm | P040002 |  | 1 | Event Description: THE REPORTED INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH REVIEW OF AN INDUSTRY RELEVANT JOURNAL ARTICLE. THERE ARE TWO JOURNAL ARTICLES THAT ARE THE SAME SUBSET OF PATIENTS. DOI.ORG/10.1177/15266028241255541. DOI.ORG/10.1177/15266028251317910. IT WAS REPORTED THAT 872 PATIENTS WERE INITIALLY IMPLANTED IN THE ITALY WITH THE AFX BIFURCATED STENT GRAFT SYSTEM (DEVICE ITERATION; AFX STRATA, DURAPLY, AFX2 ARE UNKNOWN) BETWEEN 1994 AND 2020. IT WAS REPORTED WITHIN THE JOURNAL ARTICLE EIGHT (8) PATIENTS EXPERIENCED AAA (ABDOMINAL AORTIC ANEURYSM) RELATED DEATHS. THE EXACT HOSPITAL, CASE DATE, EVENT DETAILS, LOT NUMBERS, AND CATALOG NUMBERS ARE UNKNOWN. NO ADDITIONAL INFORMATION WILL BECOME AVAILABLE REGARDING THIS INITIAL/FINAL REPORT DUE TO THE ABSENCE OF PATIENT IDENTIFIERS. THIS REPORT ADDRESSES PATIENT 4 OF 8. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. THE REPORTED ADVERSE EVENT/INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH AN ONGOING REVIEW OF INDUSTRY RELEVANT JOURNAL ARTICLES WHERE PATIENT RELATED OUTCOMES ARE PRESENTED ANONYMOUSLY, AND PATIENT SPECIFIC DEVICE INFORMATION IS UNAVAILABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE COMPLETED. THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION REMAINS UNKNOWN. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT COULD NOT BE COMPLETED AS THE REPORTED EVENT WAS FOUND DURING PUBLICATION REVIEW WHERE THE PATIENT INFORMATION IS ANONYMOUS. DEVICE, USE, PROCEDURE, AND/OR ANATOMY RELATEDNESS TO THIS ADVERSE EVENT/INCIDENT COULD NOT BE EVALUATED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22071394&pc=MIH | 3011063223-2025-00091 |  | Death | ENDOLOGIX | 2025-05-22 04:00 | MIH | AFX | Insufficient Information | Aneurysm | P040002 |  | 1 | Event Description: THE REPORTED INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH REVIEW OF AN INDUSTRY RELEVANT JOURNAL ARTICLE. THERE ARE TWO JOURNAL ARTICLES THAT ARE THE SAME SUBSET OF PATIENTS. DOI.ORG/10.1177/15266028241255541. DOI.ORG/10.1177/15266028251317910. IT WAS REPORTED THAT 872 PATIENTS WERE INITIALLY IMPLANTED IN THE ITALY WITH THE AFX BIFURCATED STENT GRAFT SYSTEM (DEVICE ITERATION; AFX STRATA, DURAPLY, AFX2 ARE UNKNOWN) BETWEEN 1994 AND 2020. IT WAS REPORTED WITHIN THE JOURNAL ARTICLE EIGHT (8) PATIENTS EXPERIENCED AAA (ABDOMINAL AORTIC ANEURYSM) RELATED DEATHS. THE EXACT HOSPITAL, CASE DATE, EVENT DETAILS, LOT NUMBERS, AND CATALOG NUMBERS ARE UNKNOWN. NO ADDITIONAL INFORMATION WILL BECOME AVAILABLE REGARDING THIS INITIAL/FINAL REPORT DUE TO THE ABSENCE OF PATIENT IDENTIFIERS. THIS REPORT ADDRESSES PATIENT 5 OF 8. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. THE REPORTED ADVERSE EVENT/INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH AN ONGOING REVIEW OF INDUSTRY RELEVANT JOURNAL ARTICLES WHERE PATIENT RELATED OUTCOMES ARE PRESENTED ANONYMOUSLY, AND PATIENT SPECIFIC DEVICE INFORMATION IS UNAVAILABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE COMPLETED. THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION REMAINS UNKNOWN. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT COULD NOT BE COMPLETED AS THE REPORTED EVENT WAS FOUND DURING PUBLICATION REVIEW WHERE THE PATIENT INFORMATION IS ANONYMOUS. DEVICE, USE, PROCEDURE, AND/OR ANATOMY RELATEDNESS TO THIS ADVERSE EVENT/INCIDENT COULD NOT BE EVALUATED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22071414&pc=MIH | 3011063223-2025-00088 |  | Death | ENDOLOGIX | 2025-05-22 04:00 | MIH | AFX | Insufficient Information | Aneurysm | P040002 |  | 1 | Event Description: THE REPORTED INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH REVIEW OF AN INDUSTRY RELEVANT JOURNAL ARTICLE. THERE ARE TWO JOURNAL ARTICLES THAT ARE THE SAME SUBSET OF PATIENTS. DOI.ORG/10.1177/15266028241255541. DOI.ORG/10.1177/15266028251317910. IT WAS REPORTED THAT 872 PATIENTS WERE INITIALLY IMPLANTED IN THE ITALY WITH THE AFX BIFURCATED STENT GRAFT SYSTEM (DEVICE ITERATION; AFX STRATA, DURAPLY, AFX2 ARE UNKNOWN) BETWEEN 1994 AND 2020. IT WAS REPORTED WITHIN THE JOURNAL ARTICLE EIGHT (8) PATIENTS EXPERIENCED AAA (ABDOMINAL AORTIC ANEURYSM) RELATED DEATHS. THE EXACT HOSPITAL, CASE DATE, EVENT DETAILS, LOT NUMBERS, AND CATALOG NUMBERS ARE UNKNOWN. NO ADDITIONAL INFORMATION WILL BECOME AVAILABLE REGARDING THIS INITIAL/FINAL REPORT DUE TO THE ABSENCE OF PATIENT IDENTIFIERS. THIS REPORT ADDRESSES PATIENT 6 OF 8. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. THE REPORTED ADVERSE EVENT/INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH AN ONGOING REVIEW OF INDUSTRY RELEVANT JOURNAL ARTICLES WHERE PATIENT RELATED OUTCOMES ARE PRESENTED ANONYMOUSLY, AND PATIENT SPECIFIC DEVICE INFORMATION IS UNAVAILABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE COMPLETED. THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION REMAINS UNKNOWN. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT COULD NOT BE COMPLETED AS THE REPORTED EVENT WAS FOUND DURING PUBLICATION REVIEW WHERE THE PATIENT INFORMATION IS ANONYMOUS. DEVICE, USE, PROCEDURE, AND/OR ANATOMY RELATEDNESS TO THIS ADVERSE EVENT/INCIDENT COULD NOT BE EVALUATED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22071573&pc=OZD | 1220648-2025-28654 | 2025-05-03 04:00 | Death | ABIOMED, INC. | 2025-05-22 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT THE IMPELLA 5.5 WAS PLACED ON A ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION INDUCED VENTRICULAR SEPTAL DEFECT PATIENT. IT WAS REPORTED WHILE ON SUPPORT, THE TEAM NOTICED UNEVEN PUPILS AND DID EMERGENT COMPUTED TOMOGRAPHY SCAN OF WHICH SHOWED GLOBAL ANOXIA AND THE FAMILY DECIDED TO WITHDRAW CARE. THE PATIENT EXPIRED AND THE RELATIONSHIP BETWEEN THE IMPELLA 5.5 AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22071597&pc=MIH | 3011063223-2025-00089 |  | Death | ENDOLOGIX | 2025-05-22 04:00 | MIH | AFX | Insufficient Information | Aneurysm | P040002 |  | 1 | Event Description: THE REPORTED INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH REVIEW OF AN INDUSTRY RELEVANT JOURNAL ARTICLE. THERE ARE TWO JOURNAL ARTICLES THAT ARE THE SAME SUBSET OF PATIENTS. DOI.ORG/10.1177/15266028241255541. DOI.ORG/10.1177/15266028251317910. IT WAS REPORTED THAT 872 PATIENTS WERE INITIALLY IMPLANTED IN THE ITALY WITH THE AFX BIFURCATED STENT GRAFT SYSTEM (DEVICE ITERATION; AFX STRATA, DURAPLY, AFX2 ARE UNKNOWN) BETWEEN 1994 AND 2020. IT WAS REPORTED WITHIN THE JOURNAL ARTICLE EIGHT (8) PATIENTS EXPERIENCED AAA (ABDOMINAL AORTIC ANEURYSM) RELATED DEATHS. THE EXACT HOSPITAL, CASE DATE, EVENT DETAILS, LOT NUMBERS, AND CATALOG NUMBERS ARE UNKNOWN. NO ADDITIONAL INFORMATION WILL BECOME AVAILABLE REGARDING THIS INITIAL/FINAL REPORT DUE TO THE ABSENCE OF PATIENT IDENTIFIERS. THIS REPORT ADDRESSES PATIENT 7 OF 8. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. THE REPORTED ADVERSE EVENT/INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH AN ONGOING REVIEW OF INDUSTRY RELEVANT JOURNAL ARTICLES WHERE PATIENT RELATED OUTCOMES ARE PRESENTED ANONYMOUSLY, AND PATIENT SPECIFIC DEVICE INFORMATION IS UNAVAILABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE COMPLETED. THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION REMAINS UNKNOWN. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT COULD NOT BE COMPLETED AS THE REPORTED EVENT WAS FOUND DURING PUBLICATION REVIEW WHERE THE PATIENT INFORMATION IS ANONYMOUS. DEVICE, USE, PROCEDURE, AND/OR ANATOMY RELATEDNESS TO THIS ADVERSE EVENT/INCIDENT COULD NOT BE EVALUATED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22071614&pc=MIH | 3011063223-2025-00090 |  | Death | ENDOLOGIX | 2025-05-22 04:00 | MIH | AFX | Insufficient Information | Aneurysm | P040002 |  | 1 | Event Description: THE REPORTED INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH REVIEW OF AN INDUSTRY RELEVANT JOURNAL ARTICLE. THERE ARE TWO JOURNAL ARTICLES THAT ARE THE SAME SUBSET OF PATIENTS. DOI.ORG/10.1177/15266028241255541. DOI.ORG/10.1177/15266028251317910. IT WAS REPORTED THAT 872 PATIENTS WERE INITIALLY IMPLANTED IN THE ITALY WITH THE AFX BIFURCATED STENT GRAFT SYSTEM (DEVICE ITERATION; AFX STRATA, DURAPLY, AFX2 ARE UNKNOWN) BETWEEN 1994 AND 2020. IT WAS REPORTED WITHIN THE JOURNAL ARTICLE EIGHT (8) PATIENTS EXPERIENCED AAA (ABDOMINAL AORTIC ANEURYSM) RELATED DEATHS. THE EXACT HOSPITAL, CASE DATE, EVENT DETAILS, LOT NUMBERS, AND CATALOG NUMBERS ARE UNKNOWN. NO ADDITIONAL INFORMATION WILL BECOME AVAILABLE REGARDING THIS INITIAL/FINAL REPORT DUE TO THE ABSENCE OF PATIENT IDENTIFIERS. THIS REPORT ADDRESSES PATIENT 8 OF 8. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. THE REPORTED ADVERSE EVENT/INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH AN ONGOING REVIEW OF INDUSTRY RELEVANT JOURNAL ARTICLES WHERE PATIENT RELATED OUTCOMES ARE PRESENTED ANONYMOUSLY, AND PATIENT SPECIFIC DEVICE INFORMATION IS UNAVAILABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE COMPLETED. THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION REMAINS UNKNOWN. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT COULD NOT BE COMPLETED AS THE REPORTED EVENT WAS FOUND DURING PUBLICATION REVIEW WHERE THE PATIENT INFORMATION IS ANONYMOUS. DEVICE, USE, PROCEDURE, AND/OR ANATOMY RELATEDNESS TO THIS ADVERSE EVENT/INCIDENT COULD NOT BE EVALUATED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22071657&pc=OZD | 1220648-2025-28655 | 2025-05-03 04:00 | Death | ABIOMED, INC. | 2025-05-22 04:00 | OZD | IMPELLA CP | Positioning Failure | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE IMPELLA WAS REPOSITIONED DUE TO PATIENT MOVEMENT. THE PATIENT LATER WENT INTO VENTRICULAR TACHYCARDIA (VT). THE PATIENT WAS GIVEN A BOLUS OF AMIODARONE. IT IS UNKNOWN IF THE POSITION OF THE PUMP CONTRIBUTED TO THE PATIENT'S VT. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: H6 HEALTH EFFECT - IMPACT CODE 4644 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28655. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES AND VF/VT/AF/AT HAS BEEN COMPLETED. THE PRODUCT AND DATA WERE NOT RETURNED FOR REVIEW. BASED ON CLINICAL DESCRIPTION, THE ROOT CAUSE OF POSITIONING ISSUE WAS MOST LIKELY PATIENT MOVEMENT SINCE THE ISSUE OCCURRED WHEN TURNING THE PATIENT. THE VF/VT/AF/AT COULD NOT BE DETERMINED WITHOUT ADDITIONAL CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22071788&pc=OZD | 1220648-2025-28656 | 2025-05-01 04:00 | Death | ABIOMED, INC. | 2025-05-22 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT DURING IMPELLA CP SUPPORT OF A 57-YEAR-OLD MALE, THE PATIENT DEVELOPED A GI BLEED AND HEPARIN WAS DISCONTINUED. ADDITIONALLY, IT WAS REPORTED THAT THE PATIENT WAS NOT DOING WELL, OCCASIONAL INTERMITTENT VAGAL, BRADYCARDIA AND ALTERNATING VENTRICULAR TACHYCARDIA. PATIENT WAS ON VASOPRESSIN. PATIENT WAS MADE A DO NOT RESUSCITATE. ULTIMATELY CASE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. H1 TYPE OF REPORT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28656. H6 HEALTH EFFECT - IMPACT CODE 1802 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28656.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22071978&pc=BZD | 2518422-2025-045275 | 2024-06-25 04:00 | Death | RESPIRONICS, INC. | 2025-05-22 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION IT IS OBSERVED THAT THERE ARE 0 ERRORS FOUND IN LOGS, NO FAULTS FOUND, AND ALL TESTING PASSED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22072267&pc=NJE | 3008114965-2025-00465 | 2024-10-17 04:00 | Death | MEDOS INTERNATIONAL SARL | 2025-05-22 04:00 | NJE | CERENOVUS ENTERPRISE | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | H60001 |  | 1 | Event Description: THE EVENT WAS REPORTED VIA THE ICAD STUDY IN CHINA. THE 53-YEAR-OLD FEMALE PATIENT (SUBJECT (B)(6)) WITH THE FOLLOWING MEDICAL HISTORY: HYPERLIPIDEMIA: FROM (B)(6) 2015, TO PRESENT, ATORVASTATIN WAS GIVEN ¦ HYPERTENSION: (B)(6) 2024 TO PRESENT, TREATED WITH CANDESARTAN TABLETS. CEREBRAL INFARCTION ON AN UNKNOWN DATE IN (B)(6) 2024, THIS WAS TREATED WITH ASPIRIN AND CLOPIDOGREL. CEREBRAL INFARCTION ON (B)(6) 2024; THE PATIENT WAS GIVEN ASPIRIN AND CLOPIDOGREL. THE PATIENT HAD NO ALLERGIC HISTORY. IT WAS REPORTED THAT THE PATIENT HAD BEEN REGULARLY TAKING ASPIRIN TABLETS AND CLOPIDOGREL ANTICOAGULANT DRUGS SINCE (B)(6) 2024 AND DISCONTINUED ANTICOAGULANT DRUGS ON (B)(6) 2024. THE PATIENT SIGNED THE INFORMED CONSENT FORM ON (B)(6) 2024, OF "IMPLANTATION OF INTRACRANIAL STENT IN PATIENTS WITH SEVERE SYMPTOMATIC ATHEROSCLEROTIC INTRACRANIAL ARTERIAL STENOSIS (CERENOVUS ENTERPRISE 2INTRACRANIAL STENT): A CHINESE MULTICENTER, PROSPECTIVE, SINGLE-ARM TARGET VALUE STUDY" WITH SCREENING NUMBER OF 1107. ON (B)(6) 2024, THE PATIENT UNDERWENT A VASCULAR STENT PLACEMENT PROCEDURE USING A 4MM X 23MM ENTERPRISE 2 STENT (ENCR402312 / 6920558). HEPARIN WAS ADMINISTERED DURING THE PROCEDURE. PRE-OPERATIVE ANGIOGRAPHY REVEALED SEVERE STENOSIS OF THE RIGHT INTERNAL CAROTID ARTERY C6 SEGMENT. THE LENGTH OF THE STENOSIS WAS ABOUT 9.93 MM, THE DIAMETER OF STENOSIS WAS 0.86 MM, THE DIAMETER OF THE NORMAL SITE AT THE PROXIMAL END OF THE STENOSIS WAS 2.94 MM. THE DIAMETER OF THE DISEASED VESSEL AT THE DISTAL END OF THE STENOSIS WAS 2.92 MM AND STENOSIS DEGREE OF ABOUT 70.7%. TEN MINUTES AFTER INTRAOPERATIVE STENT PLACEMENT, ANGIOGRAPHY SHOWED GOOD, SUSPECTED DISSECTION SHADOW AT THE DISTAL END OF THE STENT, AND 5 ML OF TIROFIBAN WAS INJECTED INTO THE ARTERY, AND ANGIOGRAPHY SHOWED RESIDUAL STENOSIS OF ABOUT 40%. FIVE MINUTES LATER, POST-DILATION WAS PERFORMED WITH BALLOON CATHETER. REPEAT ANGIOGRAPHY REVEALED GOOD STENT APPOSITION WITHOUT STENT THROMBOSIS, THE DIAMETER OF THE NARROWEST PART OF THE DISEASED VESSEL WAS 2.18 MM, PROXIMAL 2.86 MM, DISTAL 2.78 MM, AND THE RESIDUAL STENOSIS WAS ABOUT 23.8%, AND THE STENT HAD COMPLETELY COVERED THE LESION WAS SUCCESSFULLY RELEASED. THE PROCEDURE WAS COMPLETED, AND THE PATIENT WAS SENT TO ICU FOR OBSERVATION AND TREATMENT. ADDITIONAL INFORMATION WAS RECEIVED ON 18-OCT-2024. THE INFORMATION INDICATED THAT THE PATIENT EXPERIENCED A CEREBRAL HEMORRHAGE ON (B)(6) 2024 AFTER THE PROCEDURE. AT 18:00, THE PATIENT¿S PUPILS WERE DILATED COMPARED TO BEFORE, AND LIGHT REFLEX DISAPPEARED. HEAD COMPUTED TOMOGRAPHY (CT) PLAIN SCAN SHOWED BLEEDING IN THE RIGHT BASAL GANGLIA, WHICH RUPTURED INTO THE VENTRICLE, WITH A SMALL AMOUNT OF SUBARACHNOID HEMORRHAGE. THE CONSULTATION SHOWED SURGICAL INDICATIONS. AT 21:25, THE RIGHT CRANIOTOMY EVACUATION OF INTRACRANIAL HEMATOMA + DECOMPRESSIVE CRANIECTOMY + LEFT TREPANATION AND EXTERNAL VENTRICULAR DRAINAGE WAS PERFORMED. AFTER THAT, THE PATIENT RETURNED TO INTENSIVE CARE UNIT (ICU) AFTER SURGERY. THE OXYGENATION WAS MAINTAINED WELL AFTER SUBSEQUENT TRACHEAL INTUBATION AND MECHANICAL VENTILATION, AND THE LIGHT REFLEX OF BILATERAL MYDRIASIS DISAPPEARED. SYMPTOMATIC TREATMENT SUCH AS ANTI-PLATELET, LIPID REGULATION AND BLOOD PRESSURE CONTROL WERE PERFORMED, AND THE PATIENT HAD SEVERE CONDITION AND POOR PROGNOSIS. ON (B)(6) 2024 20:08, THE TEST ITEM POINT OF CARE TESTING (POCT) ELECTROLYTE ANALYSIS (ARTERIAL) SHOWED ABNORMALLY HIGH SODIUM ION RESULT, WHICH WAS DIAGNOSED AS HYPERNATRAEMIA, AND THE PATIENT WAS GIVEN SODIUM REDUCTION TREATMENT. DYNAMIC REEXAMINATION OF SERUM SODIUM LEVEL WAS PERFORMED THE PRINCIPAL INVESTIGATOR (PI) ASSESSED THE EVENT AS A SERIOUS, SEVERE ADVERSE EVENT, AND AS BEING POSSIBLY UNRELATED TO THE INVESTIGATIONAL DEVICE AND POSSIBLY RELATED TO THE SURGICAL PROCEDURE. THE EVENT WAS TREATED WITH MEDICATION AND REQUIRED AN IN-PATIENT SURGERY. THE EVENT WAS NOT CLASSIFIED AS UNANTICIPATED ADVERSE DEVICE EFFECT (UADE). THE EVENT WAS NOT CLASSIFIED AS AN ISCHEMIC STROKE BUT WAS CLASSIFIED AS A SYMPTOMATIC CEREBRAL HEMORRHAGE. THE EVENT REQUIRED PROLONG INPATIENT HOSPITALIZATION. THE EVENT RESULTED IN DEATH, WHICH WAS CLASSIFIED AS AN ¿ALL-CAUSE DEATH: CEREBROVASCULAR DEATH.¿ NO DEVICE DEFICIENCIES WERE REPORTED, AND THE DEVICE REMAINS IMPLANTED. ON (B)(6) 2024, THE PATIENT HAD HIGH URINE VOLUME, CONSIDERING DIABETES INSIPIDUS, AND WAS GIVEN HYPOPHYSIN FOR TREATMENT. RECENTLY, THE PATIENT DEVELOPED HYPERNATREMIA AND DIABETES INSIPIDUS, CONSIDERING SEVERE BRAIN INJURY. DEHYDRATION DRUGS WERE SUSPENDED CURRENTLY. THE PATIENT'S ALBUMIN WAS ELEVATED TO NORMAL LEVEL. ALBUMIN SUPPLEMENTATION WAS DISCONTINUED TODAY. CEREBRAL VASOSPASM CONTINUED TO BE PREVENTED. TEMPORARY STEROIDS COULD BE USED TO REDUCE OEDEMA IF NECESSARY. THE PATIENT HAD NO FEVER, AND INFLAMMATORY INDICATORS HAVE ELEVATED, CONSIDERING INEFFECTIVE ANTI-INFECTION CURRENTLY. CEFTAZIDIME WAS DISCONTINUED FOR ANTI-INFECTION TODAY AND CHANGED TO PIPERACILLIN AND TAZOBACTAM FOR INJECTION FOR ANTI-INFECTION. THE PATIENT CONTINUED TO PAY ATTENTION TO CHANGES IN INFLAMMATORY INDICATORS AND BODY TEMPERATURE AND ADJUSTED THE ANTIBIOTIC REGIMEN AGAIN WHEN NECESSARY. OTHER TREATMENTS WERE THE SAME AS BEFORE, AND SYMPTOMATIC AND SUPPORTIVE TREATMENTS SUCH AS ACID SUPPRESSION AND STOMACH PROTECTION, NUTRITION AND FLUID REPLACEMENT, AND WAKE-PROMOTING WERE GIVEN. ON (B)(6) 2024, THE PATIENT WAS DIAGNOSED WITH HYPERNATRAEMIA AND HYPOKALAEMIA. SYMPTOMATIC TREATMENT WITH POTASSIUM SUPPLEMENTS AND SODIUM REDUCTION WAS GIVEN. THE PATIENT HAD CENTRAL DIABETES INSIPIDUS, PITUITARY MAINTENANCE AND DESMOPRESSIN WAS ADDED. NOW THE INFLAMMATORY INDICATORS CONTINUED TO SIGNIFICANTLY INCREASE, CONSIDERING SEVERE INFECTION. THE ANTI-INFECTION REGIMEN WAS ADJUSTED. CONSULTATION WAS ON (B)(6) 2024. CT SCAN SHOWED DIFFUSE BRAIN SWELLING. THE FEVER CONTINUED FROM (B)(6) 2024. ON (B)(6) 2024, THE PATIENT'S BODY TEMPERATURE STILL FLUCTUATED, AND THE DYNAMIC RE-EXAMINATION OF INFECTION INDICATORS SHOWED A DECREASED TREND. ON (B)(6) 2024, BRONCHOSCOPY WAS PERFORMED AND SUBMITTED FOR EXAMINATION. ON (B)(6) 2024, THE PATIENT SHOWED A TENDENCY OF INCREASED SERUM SODIUM, FLUCTUATING BODY TEMPERATURE, INCREASED INFLAMMATORY INDICATORS, AND WAS ALERT TO THE POSSIBILITY OF INTRACRANIAL INFECTION. SPUTUM CULTURE ON (B)(6) 2024 WAS NOT ABNORMAL. COAGULATION WAS IMPROVING AFTER THE TRANSFUSION. ON (B)(6) 2024, NA + RESULT: 168 MMOL/L. HYPOPHYSIN WAS GIVEN TO IMPROVE DIABETES INSIPIDUS, AND STERILIZED WATER WAS USED TO REDUCE SODIUM. THE PATIENT WAS REPORTED TO BE IN A PERSISTENT COMA, CONSIDERING NEUROGENIC SHOCK. FAMILY MEMBERS ASKED FOR AUTOMATIC DISCHARGE AND WAS INFORMED THAT THE PATIENT WOULD HAVE REPEATED CONDITION, AGGRAVATED AND LIFE-THREATENING AT ANY TIME AFTER DISCHARGE, AND THE FAMILY MEMBERS SIGNED THE INFORMED CONSENT FORM AND WERE DISCHARGED HOME. THE PI MAINTAINED THE JUDGMENT THAT IT WAS POSSIBLY RELATED TO THE SURGERY, POSSIBLY UNRELATED TO THE INVESTIGATIONAL DEVICE, AND THE SEVERITY WAS SEVERE. ON (B)(6) 2024, THE RESEARCHER CONDUCTED A FOLLOW-UP VISIT ONE MONTH AFTER THE PATIENT'S SURGERY. THE PATIENT'S FAMILY INFORMED THE RESEARCHER THAT THE PATIENT HAD PASSED AWAY ON (B)(6) 2024, AFTER BEING DISCHARGED FROM THE HOSPITAL AND RETURNING HOME ON (B)(6) 2024. ON (B)(6) 2024, A CLINICAL EVENT COMMITTEE (CEC) WAS RECEIVED REGARDING THE CEREBRAL HEMORRHAGE EVENT. THE CEC ASSESSED THE EVENT AS BEING POSSIBLY UNRELATED TO THE STUDY DEVICE AND PROBABLY RELATED TO THE PROCEDURE. OTHER POSSIBLE REASONS FOR THE EVENT WERE REPORTED TO BE ¿1. HAEMORRHAGE AFTER CEREBRAL INFARCTION; 2. HYPERPERFUSION CEREBRAL HAEMORRHAGE; 3. CAUSED BY ANTIPLATELET DRUGS; AND 4. STENT DISPLACEMENT CAUSED BY POST-STENT DILATION.¿ BASED ON THE ADDITIONAL INFORMATION RECEIVED ON 18-OCT-2024 AND 24-DEC2024, THE EVENT MEETS USFDA REPORTABILITY CRITERIA UNDER 21 CFR 803 WITH A CLASSIFICATION OF ¿DEATH.¿ Manufacturer Narrative: MANUFACTURER¿S REF. NO: (B)(4). THE PURPOSE OF THIS MDR SUBMISSION IS TO CORRECT THE AWARENESS DATE THAT WAS DOCUMENTED IN THE INITIAL 3500A MEDWATCH REPORT. [CORRECTION]: ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2024. THE INFORMATION INDICATED THAT THE PATIENT EXPERIENCED A CEREBRAL HEMORRHAGE ON (B)(6) 2024 AFTER THE PROCEDURE. ON (B)(6) 2024, A CLINICAL EVENT COMMITTEE (CEC) WAS RECEIVED REGARDING THE CEREBRAL HEMORRHAGE EVENT. THE CEC ASSESSED THE EVENT AS BEING POSSIBLY UNRELATED TO THE STUDY DEVICE AND PROBABLY RELATED TO THE PROCEDURE. OTHER POSSIBLE REASONS FOR THE EVENT WERE REPORTED TO BE ¿1. HAEMORRHAGE AFTER CEREBRAL INFARCTION; 2. HYPERPERFUSION CEREBRAL HAEMORRHAGE; 3. CAUSED BY ANTIPLATELET DRUGS; AND 4. STENT DISPLACEMENT CAUSED BY POST-STENT DILATION.¿ BASED ON THE ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2024 AND 24-DEC-2024, THE EVENT MEETS USFDA REPORTABILITY CRITERIA UNDER 21 CFR 803 WITH A CLASSIFICATION OF ¿DEATH.¿ THE MANUFACTURER WILL SUBMIT A SUPPLEMENTAL REPORT IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT. ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. Manufacturer Narrative: MANUFACTURER¿S REF. NO: (B)(4). INFORMATION REGARDING PATIENT DATE OF BIRTH AND GENDER WERE NOT PROVIDED. SECTION E.1: THE INITIAL REPORTER PHONE: (B)(6). THE INITIAL REPORTER EMAIL ADDRESS WAS NOT AVAILABLE / REPORTED. BASED ON COMPLAINT INFORMATION, THE DEVICE REMAINS IMPLANTED AND IS THUS NOT AVAILABLE FOR EVALUATION. LAKE REGION MEDICAL PERFORMED A REVIEW OF THE DEVICE HISTORY RECORDS RELATIVE TO THE MANUFACTURING, INSPECTION, AND PACKAGING OF THE LOT 6920558. THE HISTORY RECORD INDICATES THIS PRODUCT WAS FINAL INSPECTION TESTED AT LAKE REGION MEDICAL AND WAS DETERMINED TO BE ACCEPTABLE. THERE WAS NO DEVICE PERFORMANCE ISSUE OR DEVICE MALFUNCTION REPORTED DURING THE PROCEDURE. BASED ON THE MANUFACTURING DOCUMENTATION REVIEW, THERE IS NO INDICATION THAT THE EVENT IS RELATED TO THE DEVICE MANUFACTURING PROCESS. PER THE ADDITIONAL INFORMATION RECEIVED ON 18-OCT-2024 AND 24-DEC-2024; REGARDING THE ADVERSE EVENT OF ¿CEREBRAL HEMORRHAGE¿, THIS EVENT IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE USE OF THE ENTERPRISE2 VASCULAR RECONSTRUCTION DEVICE AND IS LISTED IN THE INSTRUCTIONS FOR USE (IFU) AS SUCH. THERE WERE NO ALLEGED QUALITY ISSUES/ MALFUNCTIONS RELATED TO THE DEVICE, AS THE DEVICE PERFORMED AS INTENDED. THE CEC ASSESSED THE EVENT AS BEING POSSIBLY UNRELATED TO THE STUDY DEVICE AND PROBABLY RELATED TO THE PROCEDURE. BASED ON THIS ASSESSMENT, THE RELATIONSHIP BETWEEN EVENT TO THE USED ENTERPRISE2 STUDY DEVICE AS A CONTRIBUTING FACTOR CANNOT BE EXCLUDED. THEREFORE, THIS EVENT WILL BE CONSERVATIVELY REPORTED TO THE US FDA UNDER REPORTING CRITERIA 21 CFR 803 WITH THE CLASSIFICATION OF ¿DEATH.¿ THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. AS PART OF THE POST MARKET SURVEILLANCE PROGRAM, INFORMATION FROM THIS COMPLAINT IS TRENDED FOR STATISTICAL SIGNALS AND CORRECTIVE / PREVENTIVE ACTION MAY BE TRIGGERED AT A LATER TIME. SINCE THERE WAS NO EVIDENCE TO SUGGEST THE EVENT WAS RELATED TO A MANUFACTURING OR DESIGN ISSUE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS BLANK; THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION. THE MANUFACTURER WILL SUBMIT A SUPPLEMENTAL REPORT IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT. ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22072296&pc=NJE | 3008114965-2025-00466 | 2023-10-11 04:00 | Death | MEDOS INTERNATIONAL SARL | 2025-05-22 04:00 | NJE | CERENOVUS ENTERPRISE | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage |  |  | 1 | Event Description: THE EVENT WAS REPORTED VIA THE ICAD STUDY IN CHINA. THE 51-YEAR-OLD MALE PATIENT (SUBJECT (B)(6): ) WITH A MEDICAL HISTORY OF HYPERTENSION FOR 2 YEARS (BLOOD PRESSURE WAS NOT MONITORED), DIABETES FOR HALF A YEAR (BLOOD GLUCOSE WAS NOT MONITORED), CEREBRAL INFARCTION, DYSLIPIDEMIA, HYPERHOMOCYSTEINEMIA, INTERNAL CAROTID ARTERY STENOSIS (SEVERE IN THE RIGHT POSTERIOR COMMUNICATING SEGMENT AND MILD IN MULTIPLE SITES IN THE LEFT SIDE), VERTEBRAL ARTERY STENOSIS (MILD IN BILATERAL SIDES) WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2023 DUE TO ¿"NUMBNESS OF THE LEFT SIDE OF THE BODY WITH DEVIATED MOUTH AND LISP FOR MORE THAN 10 DAYS AND DIAGNOSED WITH ¿INTERNAL CAROTID ARTERY STENOSIS CEREBRAL INFARCTION AND INTERNAL CAROTID ARTERY STENOSIS (SEVERE IN THE RIGHT POSTERIOR COMMUNICATING SEGMENT AND MILD IN MULTIPLE SITES IN THE LEFT SIDE).¿ THE INFORMED CONSENT FORM WAS SIGNED ON (B)(6) 2023. THE PATIENT UNDERWENT A VASCULAR STENT PLACEMENT USING AN UNKNOWN ENTERPRISE II VASCULAR RECONSTRUCTION DEVICE (PRODUCT CODE / LOT # UNKNOWN) ON (B)(6) 2023. ON (B)(6) 2023, THE PATIENT EXPERIENCED A CEREBRAL HEMORRHAGE, WHICH BECAME KNOWN TO THE SITE ON (B)(6) 2023 AND TO THE SPONSOR ON (B)(6) 2023. THE PRINCIPAL INVESTIGATOR (PI) ASSESSED THE EVENT AS A SERIOUS, SEVERE ADVERSE EVENT THAT WAS UNRELATED TO THE STUDY DEVICE AND UNRELATED TO THE SURGICAL PROCEDURE. THE EVENT REQUIRED AN IN-PATIENT SURGERY AND HOSPITALIZATION. THE EVENT WAS NOT CLASSIFIED AS UNANTICIPATED ADVERSE DEVICE EFFECT (UADE). THE EVENT WAS CLASSIFIED AS A SYMPTOMATIC CEREBRAL HEMORRHAGE. THE EVENT RESULTED IN DEATH, CLASSIFIED AS AN ¿ALL-CAUSE DEATH: CEREBROVASCULAR DEATH.¿ THE PATIENT EXPIRED ON (B)(6) 2024. NO DEVICE DEFICIENCIES WERE REPORTED, AND THE DEVICE REMAINED. THE SUBJECT WAS WITHDRAWN FROM THE TRIAL DUE TO THE ADVERSE EVENT. THE FOLLOWING ADDITION EVENT DETAILS WERE REPORTED: ON (B)(6) 2023, THE PI PERFORMED A 30-DAY POSTOPERATIVE TELEPHONE VISIT. THE PATIENT¿S FAMILY INFORMED THE PI THAT THE PATIENT HAD DIED ON THE EVENING OF (B)(6) 2023, AT HOME. THE PATIENT¿S FAMILY REPORTED THAT THE PATIENT WOKE UP IN THE MORNING OF (B)(6) 2023 AND SUDDENLY BECAME UNCONSCIOUS, UNABLE TO ANSWER THE CALL, NAUSEA, VOMITING AND OTHER SYMPTOMS, AND WAS URGENTLY ADMITTED TO THE LOCAL HOSPITAL. THE PATIENT¿S BLOOD PRESSURE WAS MEASURED AT 200/120 MMHG. THE CRANIAL CT SUGGESTED THAT THE RIGHT FRONTOTEMPORAL LOBE-BASAL GANGLIA CEREBRAL HEMORRHAGE AND BROKE INTO THE VENTRICULAR SYSTEM. AFTER COMPLETING THE RELEVANT INVESTIGATIONS, THE PATIENT UNDERWENT THE PROCEDURE OF "REMOVAL OF INTRACEREBRAL HEMATOMA, DECOMPRESSION OF CRANIAL BONE DEBRIDEMENT FLAP, AND LATERAL VENTRICLE EXTERNAL DRAINAGE" UNDER GENERAL ANESTHESIA AT 10:30 ON (B)(6) 2023 IN AN EMERGENCY CLINIC. AFTER THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO THE DEPARTMENT OF INTENSIVE CARE MEDICINE AT 15:15 ON (B)(6) 2023, DUE TO CRITICAL CONDITION, AND WAS THEN GIVEN MECHANICAL VENTILATION, REHYDRATION, ELEVATION OF PRESSURE, HEMOSTASIS, ANTI-INFECTION, NERVE NOURISHMENT, AND SYMPTOMATIC TREATMENT. ON (B)(6) 2023, CRANIOCEREBRAL CT RE-EXAMINATION SHOWED THAT THERE WAS MORE BLOOD IN THE CEREBRAL VENTRICLES THAN BEFORE. AT 10:40 ON (B)(6) 2023, THE SUBJECT UNDERWENT "CONICAL CRANIAL HEMATOMA DRAINAGE" UNDER LOCAL ANESTHESIA. AFTER THE OPERATION, THE SUBJECT CONTINUED TO BE GIVEN MECHANICAL VENTILATION, FLUID REHYDRATION, NERVE NOURISHMENT, MAINTENANCE OF STABLE INTERNAL ENVIRONMENT, TRACHEOTOMY, LUMBAR LARGE POOL TUBE DRAINAGE, AND SYMPTOMATIC TREATMENT. ON (B)(6) 2023, CT EXAMINATION WAS PERFORMED, SUGGESTING THAT THE RIGHT FRONTOTEMPORAL LOBE-BASAL GANGLIA AREA CEREBRAL HEMORRHAGE AND BREAK INTO THE VENTRICLE SYSTEM WHICH WAS CHANGED AFTER CRANIOTOMY AND DRAINAGE SURGERY, AND DYNAMIC REVIEW WAS RECOMMENDED. AFTER ACTIVE TREATMENT THE PATIENT'S CONDITION DID NOT RECOVER, WITH PERSISTENT COMA, BILATERAL PUPIL DILATATION, VENTILATOR WAS USED TO MAINTAIN THE RESPIRATION. THE CONDITION IS CRITICAL. AFTER CONSULTING THE LOCAL DOCTORS, COMBINED WITH THE CONDITION AT THAT TIME, THE PATIENT¿S FAMILY DECIDED TO GIVE UP THE TREATMENT AFTER DELIBERATION AND THE PATIENT WAS DISCHARGED ON (B)(6) 2023. DISCHARGE EXAMINATION: P 99 BEATS/MIN, R 14 BEATS/MIN, BP 130/90 MMHG, DEEP COMA, UNCOOPERATIVE EXAMINATION. THE SUBJECT DIED ON THE EVENING OF (B)(6) 2023. IN CONJUNCTION WITH THE PATIENT¿S POST-OPERATIVE RECOVERY, THE PATIENT¿S COAGULATION AND PLATELET COUNTS WERE NORMAL PRE-OPERATIVELY AND THE PATIENT WAS TAKING ANTIPLATELET AGGREGATION MEDICATIONS REGULARLY POSTOPERATIVELY. BECAUSE OF THE PATIENT'S CLOPIDOGREL RESISTANCE, THE PATIENT WAS GIVEN TICAGRELOR 40 MG PO BID. CONSIDERING THAT THIS EVENT MAY BE RELATED TO THE PATIENT'S EMOTIONAL STATE AFTER RETURNING HOME FROM A PARTY WITH RELATIVES ON THE EVENING OF (B)(6) 2023, WHICH ELEVATED THE PATIENT¿S BLOOD PRESSURE; IT MAY ALSO BE RELATED TO THE PATIENT'S USE OF TICAGRELOR. IN ADDITION TO THE POSTOPERATIVE DAY'S REVIEW, THE CT PATIENT HAD NO BLEEDING AND THE PATIENT'S CONDITION WAS STABLE ON THE FIFTH DAY OF THE POSTOPERATIVE PERIOD. BASED ON THE REASONS MENTIONED ABOVE, IT IS DETERMINED THAT THIS EVENT IS NOT RELATED TO THE RESEARCH DEVICE, AND IT IS NOT RELATED TO THE SURGICAL OPERATION. ON (B)(6) 2023, A CLINICAL EVENT COMMITTEE (CEC) WAS RECEIVED RELATED TO THE CEREBRAL HEMORRHAGE EVENT. THE CEC ASSESSED THE EVENT AS BEING POSSIBLY UNRELATED TO THE STUDY DEVICE AND POSSIBLY RELATED TO THE PROCEDURE. OTHER POSSIBLE CAUSES FOR THE EVENT WERE REPORTED AS, ¿MAY BE RELATED TO ANTI-PLATELET DRUGS, POOR BLOOD PRESSURE CONTROL AND EMOTIONAL FACTORS.¿ BASED ON THE CEC ADJUDICATION RECEIVED ON 14-DEC-2023, THE EVENT MEETS USFDA REPORTABILITY CRITERIA UNDER 21 CFR 803 WITH A CLASSIFICATION OF ¿DEATH.¿ Manufacturer Narrative: MANUFACTURER¿S REF. NO: (B)(4). INFORMATION REGARDING PATIENT IDENTIFIER, DATE OF BIRTH, AGE, SEX, GENDER, WEIGHT, RACE, AND ETHNICITY WERE NOT PROVIDED. SECTION D.4: THE PRODUCT CATALOG AND LOT NUMBERS ARE NOT AVAILABLE / NOT REPORTED. THE UNIQUE IDENTIFIER (UDI) AND EXPIRATION DATE OF THE DEVICE ARE NOT AVAILABLE. SECTION E.1: THE NAME, PHONE AND EMAIL ADDRESS OF THE INITIAL REPORTER ARE NOT AVAILABLE / REPORTED. SECTION H.4: THE DEVICE MANUFACTURE DATE IS NOT KNOWN AS THE DEVICE LOT NUMBER IS NOT AVAILABLE / NOT REPORTED. BASED ON COMPLAINT INFORMATION, THE DEVICE REMAINS IMPLANTED AND IS THUS NOT AVAILABLE FOR EVALUATION. THE DEVICE LOT NUMBER WAS NOT AVAILABLE. THE MANUFACTURING DOCUMENTATION REVIEW COULD NOT BE PERFORMED WITHOUT THE LOT NUMBER. THERE WAS NO DEVICE PERFORMANCE ISSUE OR DEVICE MALFUNCTION REPORTED DURING THE PROCEDURE. CEREBRAL HEMORRHAGE IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE ENTERPRISE2 VASCULAR RECONSTRUCTION DEVICE AND IS LISTED IN THE INSTRUCTIONS FOR USE (IFU) AS SUCH. THERE WAS NO ALLEGED QUALITY ISSUES RELATED TO THE USED DEVICE, AS THE DEVICE PERFORMED AS INTENDED. THE PI ASSESSED THE EVENT AS BEING UNRELATED TO THE ENTERPRISE2 STUDY DEVICE AND UNRELATED TO THE PROCEDURE. HOWEVER, A CEC ADJUDICATION ASSESSED THE EVENT AS POSSIBLY UNRELATED TO THE STUDY DEVICE AND POSSIBLY RELATED TO THE PROCEDURE. OTHER POSSIBLE CAUSES FOR THE EVENT WERE REPORTED TO BE ¿ANTI-PLATELET DRUGS, POOR BLOOD PRESSURE CONTROL AND EMOTIONAL FACTORS.¿ BASED ON THE CEC ADJUDICATION, THE CORRELATING RELATIONSHIP BETWEEN THE EVENT TO THE USED DEVICE AS A CONTRIBUTING FACTOR CANNOT BE RULED OUT ENTIRELY. THEREFORE, THIS EVENT MEETS US FDA REPORTING CRITERIA UNDER 21 CFR 803 WITH THE CLASSIFICATION OF ¿DEATH.¿ THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. AS PART OF THE POST MARKET SURVEILLANCE PROGRAM, INFORMATION FROM THIS COMPLAINT IS TRENDED FOR STATISTICAL SIGNALS AND CORRECTIVE / PREVENTIVE ACTION MAY BE TRIGGERED AT A LATER TIME. SINCE THERE WAS NO EVIDENCE TO SUGGEST THE EVENT WAS RELATED TO A MANUFACTURING OR DESIGN ISSUE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. THE MANUFACTURER WILL SUBMIT A SUPPLEMENTAL REPORT IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT. ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22072807&pc=MVK | 3008642652-2025-05104 | 2025-04-30 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-22 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FOUR INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 17:50:59 ON (B)(6) 2025. AT 18:52:02 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION ARTIFACT. AT 18:52:39, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH FLUTTER WAVES. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 18:53:06, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 18:53:33, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 18:54:04, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. THE ELECTRODE BELT WAS DISCONNECTED AT 19:21:16 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22072808&pc=MVK | 3008642652-2025-05105 | 2025-05-04 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-22 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FOUR APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 22:52:15 ON (B)(6) 2025. AT 04:03:02 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 190 BPM WITH MOTION ARTIFACT SELF-CONVERTING TO SVT @ 210 BPM WITH PVCS, LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT. AT 04:05:00, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 200 BPM WITH LOW AMPLITUDE CARDIAC SIGNAL AND MOTION ARTIFACT. POST SHOCK RHYTHM WAS VT FROM 100-220 BPM WITH LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT. AT 04:05:30, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT @ 210 BPM WITH MOTION ARTIFACT. AT 04:06:20, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 210 BPM WITH MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. POST SHOCK RHYTHM WAS VT @ 120 BPM WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. AT 04:07:14, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 190 BPM WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. POST SHOCK RHYTHM WAS VT FROM 140-210 BPM WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. AT 04:08:16, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. POST SHOCK RHYTHM WAS VT FROM 140-210 BPM WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. THE ELECTRODE BELT WAS DISCONNECTED AT 04:08:29 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22072815&pc=DSQ | 2916596-2025-03119 | 2025-02-26 05:00 | Death | THORATEC CORPORATION | 2025-05-22 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Septic Shock; Heart Failure/Congestive Heart Failure; Tricuspid Valve Insufficiency/ Regurgitation; Renal Impairment; Peripheral Edema | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO HEART FAILURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT TRICUSPID REGURGITATION GRADUALLY DEVELOPED POST OPERATIVELY. ON (B)(6) 2025, THE PATIENT WAS ADMITTED TO THE CORONARY CARE UNIT (CCU) DUE TO SHOCK FROM RIGHT HEART FAILURE WITH SYMPTOMS OF ASCITES AND PERIPHERAL EDEMA AND CATECHOLAMINE THERAPY WAS INITIATED. ON (B)(6) 2025, THE PATIENT WAS INTUBATED AND PLACED ON MECHANICAL VENTILATION. ON (B)(6) 2025, EMERGENCY SURGERY WAS PERFORMED FOR TRICUSPID VALVE REPLACEMENT (TVR) AND THROMBECTOMY AT THE AORTIC VALVE POSITION. RIGHT HEART FAILURE PERSISTED POSTOPERATIVELY. ON (B)(6) 2025, CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) WAS INITIATED. ON (B)(6) 2025, THE PATIENT DEVELOPED SEPTIC SHOCK DUE TO PSEUDOMONAS AERUGINOSA. ON (B)(6) 2025, VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (VV-ECMO) WAS INITIATED. AS THE PATIENT WAS ON VENTRICULAR ASSIST DEVICE (VAD) SUPPORT, CAUSAL RELATIONSHIP COULD NOT BE DENIED. THE DEVICE OPERATED AS EXPECTED AT THE TIME OF DEATHEvent Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE RIGHT HEART FAILURE DID NOT EXIST BEFORE IMPLANT. ALTHOUGH THIS WAS NOT CONSIDERED TO BE DEVICE RELATED, IT COULD NOT BE RULED OUT. DIASTOLIC DISFUNCTION AND WORSENING TRICUSPID REGURGITATION LED TO THE ONSET OF THE RIGHT HEART FAILURE, WHICH WAS FURTHER EXACERBATED BY SEPTIC SHOCK. THE PATIENT EXPERIENCED SYMPTOMS OF SHORTNESS OF BREATH AND DYSPNEA. ON (B)(6) 2025, THE PATIENT HAD POSITIVE BACTERIAL BLOOD CULTURES, WHICH WERE TREATED WITH PHARMALOGICAL THERAPY ONLY.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT BLOOD CULTURES WERE POSITIVE FOR ENTEROBACTER. THE PATIENT DID NOT HAVE KIDNEY FAILURE BEFORE LEFT VENTRICULAR ASSIST DEVICE (LVAD). HOWEVER, THE LVAD WAS NOT A FACTOR TO CAUSE OR CONTRIBUTE TO KIDNEY FAILURE. A 3CM LONG THROMBUS WAS FOUND IN THE LEFT CORONARY CUSP OF THE AORTIC VALVE DURING TRICUSPID VALVE REPLACEMENT. HOWEVER, NO IMAGES WERE AVAILABLE. NO PARTICULAR SYMPTOMS WERE NOTED DUE TO THE THROMBUS. THE PATIENT'S RIGHT HEART FAILURE WAS FURTHER AGGRAVATED BY SEPTIC SHOCK. THE INTERNATIONAL NORMALIZED RATIO (INR) WAS CONTROLLED BUT THE PATIENT WAS BLEEDING DUE TO THROMBOCYTOPENIA (MAX 19000).Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE HEART FAILURE WAS CONSIDERED TO BE PROGRESSION OF THE PRIMARY DISEASE AND INFECTION. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: CORRECTED DATA: SECTION H6: DUE TO SYSTEM LIMITATIONS, THE FOLLOWING WAS NOT INCLUDED PREVIOUSLY: HEALTH EFFECT - CLINICAL CODE: 1816 DYSPNEA. HEALTH EFFECT - CLINICAL CODE: 2596 ASCITES. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22072816&pc=MVK | 3008642652-2025-05106 | 2025-02-19 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-22 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 13:57:39 ON (B)(6) 2025. AT 18:16:14 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 18:16:47, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. THE ELECTRODE BELT WAS DISCONNECTED AT 18:17:06 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22072817&pc=MVK | 3008642652-2025-05107 | 2025-05-03 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-22 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 15:14:05 ON (B)(6) 2025. AT 18:40:16, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 18:41:27, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM. AT 18:41:53, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 30 BPM. THE DEVICE WAS SHUT DOWN AT 18:46:30 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22072836&pc=MVK | 3008642652-2025-05108 | 2025-04-26 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-22 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 06:28:33 ON (B)(6) 2025. AT 20:41:29 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 20:44:03, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 20 BPM WITH CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 21:07:13 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22072920&pc=MVK | 3008642652-2025-05109 | 2025-03-04 05:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-22 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FIVE APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 10:52:03 ON (B)(6) 2025. AT 01:22:38 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS IDIOVENTRICULAR RHYTHM @ 40 BPM DEGRADING TO VF. AT 01:23:50, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 01:24:29, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 01:25:16, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 01:25:49, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH MOTION ARTIFACT AND ELECTRODE LEAD FALL OFF. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 01:29:45, AN ARRHYTHMIA WAS DETECTED. ECG WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 01:30:48, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VT @ 200 BPM WITH MOTION ARTIFACT. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. THE ELECTRODE BELT WAS DISCONNECTED AT 01:58:39 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22072982&pc=KPO | 1528807-2025-00004 | 2024-10-15 04:00 | Death | NIPRO RENAL SOLUTIONS USA CORPORATION | 2025-05-22 04:00 | KPO | MEDICALYTE | Device Contaminated During Manufacture or Shipping |  | K012547 |  | 1 | Event Description: THE PATIENT STARTED DIALYSIS TREATMENT AND THE NURSE NOTICED CONTAMINATION IN THE BICARBONATE DIALYSATE. THE TREATMENT WAS STOPPED AND BLOOD CULTURES PROVIDED ELEVATED RESULTS. THE PATIENT WAS GIVEN ANTIBIOTICS, AND LATER PASSED AWAY. THE RN STATED THE DEATH DID NOT APPEAR TO BE RELATED TO THE CONTAMINATED PRODUCT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22073417&pc=LWR | 3004478276-2025-00136 | 2025-04-23 04:00 | Death | CORCYM S.R.L. | 2025-05-22 04:00 | LWR | PERCEVAL PLUS SUTURELESS AORTIC HEART VALVE | Leak/Splash | Aortic Valve Insufficiency/ Regurgitation | P150011S013 |  | 1 | Event Description: MANUFACTURER WAS INFORMED OF THE FOLLOWING EVENT: REPORTEDLY, ON (B)(6) 2025, PATIENT WAS PREPARED FOR BYPASS AND AORTIC VALVE REPLACEMENT. AFTER COMPLETING MULTIPLE BYPASSES, WHICH WERE REPORTED AS CHALLENGING DUE TO VERY DELICATE TISSUE, THE AORTA WAS OPENED, THE LEAFLETS AND CALCIUM WERE REMOVED, AND THE ANNULUS WAS SIZED AS A SMALL PERCEVAL PLUS VALVE. THE VALVE WAS COLLAPSED, IMPLANTED, AND THE BALLOONING WAS PERFORMED. THEN, THEY STARTED TO REWARM THE PATIENT AND COME OFF BYPASS, HOWEVER IT WAS SEEN ON THE ECHO THAT THERE WERE SEVERAL LEAKS AROUND THE VALVE, AND IT WAS SITTING ABOVE THE ANNULUS IN THE SINUSES. THEY WENT BACK ON BYPASS AND REMOVED THE VALVE. AS SUCH, THE ANNULUS WAS CHECKED AND A BIT MORE CALCIUM WAS REMOVED. THE VALVE WAS LOADED ON THE HOLDER AGAIN AND REIMPLANTED TAKING MORE CAREFUL STEPS TO ENSURE PROPER IMPLANT AND DEPLOYMENT OF THE VALVE. FURTHER CHECKS FOR VALVE SEATING POSITION WERE PERFORMED, AND BALLOONING WAS COMPLETED. THEN, THEY STARTED TO REWARM AND SLOWLY COME OFF BYPASS. REPORTEDLY, ON ECHO EVALUATION THE VALVE HAD NO LEAKS AND WAS SEEN TO BE IN THE CORRECT POSITION WITH GOOD COAPTATION OF THE LEAFLETS. THE PATIENT'S RIGHT VENTRICLE WAS SAID TO BE BALLOONING OUT AND HEMODYNAMICS AND BLOOD PRESSURE WERE NOT STABLE. THUS, THEY WENT BACK ON BYPASS TO GAIN CONTROL AND ATTEMPT TO COME OFF SLOWER WITH MANAGING THE VOLUMES. IN TOTAL, PATIENT WAS ON BYPASS FOR OVER 4 HOURS AND CROSS CLAMP FOR OVER 1 HOUR. IT WAS REPORTED THAT THE PATIENT ULTIMATELY PASSED AWAY. PATIENT MEDICAL HISTORY WAS REPORTED AS: CORONARY ARTERY DISEASE, AORTIC STENOSIS, END STAGE RENAL DISEASE, OBESITY, OBSTRUCTIVE SLEEP APNEA, TYPE 2 DIABETES, HYPERTENSION, HYPOTHYROIDISM, DYSLIPIDEMIA. Manufacturer Narrative: THE MANUFACTURING AND MATERIAL RECORDS FOR THE DEVICE INVOLVED IN THIS EVENT, AS THEY PERTAIN TO THE REPORTED ISSUE, WERE RETRIEVED AND REVIEWED BY THE MANUFACTURER¿S QUALITY ENGINEERING. THE RESULTS CONFIRMED THAT THIS PROSTHESIS SATISFIED ALL REQUIRED MATERIAL, VISUAL, AND PERFORMANCE STANDARDS AT THE TIME OF MANUFACTURE AND RELEASE. MANUFACTURER IS FOLLOWING UP TO GATHER MORE INFORMATION REGARDING THE REPORTED EVENT. A FOLLOW UP REPORT WILL BE PROVIDED UPON RECEIPT OF ANY FURTHER INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22073464&pc=DYE | 2025587-2025-03845 | 2025-03-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-22 04:00 | DYE | MOSAIC AORTIC BIOPROSTHETIC HEART VALVE | Insufficient Information; Patient Device Interaction Problem | Fever; Low Blood Pressure/ Hypotension; Sepsis; Chills; Cardiogenic Shock; Insufficient Information | P990064 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT 13 YEARS AND 6 MONTHS POST IMPLANT OF THIS 23MM AORTIC BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, THE RELATEDNESS OF THE DEATH TO THE AORTIC BIOPROSTHETIC VALVE CAN NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE OF DEATH WAS CARDIOGENIC SHOCK. IT WAS STATED THAT PRIOR TO THE PATIENTS DEATH, THE PATIENT PRESENTED TO THE HOSPITAL DUE TO FEVER AND CHILLS. THE PATIENT WAS ADMITTED FOR SEPSIS AND FOUND TO HAVE STREP BACTEREMIA. IT WAS REPORTED THAT THE PATIENT WAS TREATED WITH ANTIBIOTICS, DEVELOPED HYPOTENSION AND A TRANSTHORACIC ECHOCARDIOGRAM (TTE) PERFORMED SHOWED NO EVIDENCE OF ENDOCARDITIS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5, B7 AND H6 UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22073717&pc=DSQ | 2916596-2025-03037 | 2025-04-24 04:00 | Death | THORATEC CORPORATION | 2025-05-22 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI SYSTEM ORGAN FAILURE. THE PATIENT WAS TOO SICK. THIS WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED.Event Description: THE DEVICE OPERATED AS EXPECTED. THE TREATMENT FOR MULTI SYSTEM ORGAN FAILURE WAS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. G AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. G ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTI ORGAN FAILURE (INCLUDING RESPIRATORY FAILURE, RIGHT HEART FAILURE, HEPATIC DYSFUNCTION, AND RENAL DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22074303&pc=OBJ | 2124215-2025-31098 | 2025-04-24 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-22 04:00 | OBJ | OPTICROSS? HD | Poor Quality Image | Air Embolism; Cardiac Arrest; Non specific EKG/ECG Changes | K173284 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. THE STENOSED TARGET LESION WAS LOCATED IN THE DISTAL LEFT MAIN CORONARY ARTERY. THE OPTICROSS HD IMAGING CATHETER WAS ADVANCED FOR INTRAVASCULAR ULTRASOUND (IVUS) OF THE TARGET LESION. AS THE PHYSICIAN ATTEMPTED TO PASS THE IVUS CATHETER ACROSS THE LESION, IMAGING WAS ACTIVATED AND AN ARTIFACT WAS OBSERVED ON THE SCREEN. THE TECHNICIAN WAS ASKED WHETHER FLUSHING WAS OCCURRING, AND THEY CONFIRMED THAT IT WAS. THE TECHNICIAN WAS THEN INSTRUCTED TO STOP FLUSHING, AND THE PHYSICIAN AND TECHNICIAN DECIDED TO REMOVE AND RE-PREP THE CATHETER. THE PHYSICIAN NOTED ELEVATED ST-SEGMENTS, AND CONTRAST INJECTION REVEALED A SLOW-FLOW PHENOMENON. A DECISION WAS MADE TO ASPIRATE THE LEFT ANTERIOR DESCENDING ARTERY (LAD). ASPIRATION WAS PERFORMED, WHICH RESTORED SOME FLOW TO THE PROXIMAL LAD. HOWEVER, THE PATIENTS CONDITION WORSENED, AND THEY WENT INTO CARDIAC ARREST WITH PULSELESS ELECTRICAL ACTIVITY (PEA). CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED, AND A PERCUTANEOUS LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS INSERTED. CORONARY ANGIOGRAPHY PERFORMED AFTER IMPELLA PLACEMENT SHOWED NORMAL CORONARY FLOW. DESPITE EFFORTS, THE PATIENT DIED FOLLOWING A PROLONGED PEA ARREST. POST-PROCEDURE INSPECTION OF THE CATHETER REVEALED AIR IN THE CLEAR LUMEN. IT WAS NOTED THAT THE SUPPLIED EXTENSION TUBING HAD NOT BEEN USED; INSTEAD, TUBING FROM A DIFFERENT MANUFACTURER HAD BEEN CONNECTED TO THE BOSTON SCIENTIFIC STOPCOCK AND SYRINGES. THE PHYSICIAN BELIEVED THAT IMPROPER CATHETER PREPARATION LED TO AN AIR EMBOLISM. THE PROCEDURE WAS NOT COMPLETED DUE TO THIS EVENT.Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. THE STENOSED TARGET LESION WAS LOCATED IN THE DISTAL LEFT MAIN CORONARY ARTERY. THE OPTICROSS HD IMAGING CATHETER WAS ADVANCED FOR INTRAVASCULAR ULTRASOUND (IVUS) OF THE TARGET LESION. AS THE PHYSICIAN ATTEMPTED TO PASS THE IVUS CATHETER ACROSS THE LESION, IMAGING WAS ACTIVATED AND AN ARTIFACT WAS OBSERVED ON THE SCREEN. THE TECHNICIAN WAS ASKED WHETHER FLUSHING WAS OCCURRING, AND THEY CONFIRMED THAT IT WAS. THE TECHNICIAN WAS THEN INSTRUCTED TO STOP FLUSHING, AND THE PHYSICIAN AND TECHNICIAN DECIDED TO REMOVE AND RE-PREP THE CATHETER. THE PHYSICIAN NOTED ELEVATED ST-SEGMENTS, AND CONTRAST INJECTION REVEALED A SLOW-FLOW PHENOMENON. A DECISION WAS MADE TO ASPIRATE THE LEFT ANTERIOR DESCENDING ARTERY (LAD). ASPIRATION WAS PERFORMED, WHICH RESTORED SOME FLOW TO THE PROXIMAL LAD. HOWEVER, THE PATIENTS CONDITION WORSENED, AND THEY WENT INTO CARDIAC ARREST WITH PULSELESS ELECTRICAL ACTIVITY (PEA). CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED, AND A PERCUTANEOUS LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS INSERTED. CORONARY ANGIOGRAPHY PERFORMED AFTER IMPELLA PLACEMENT SHOWED NORMAL CORONARY FLOW. DESPITE EFFORTS, THE PATIENT DIED FOLLOWING A PROLONGED PEA ARREST. POST-PROCEDURE INSPECTION OF THE CATHETER REVEALED AIR IN THE CLEAR LUMEN. IT WAS NOTED THAT THE SUPPLIED EXTENSION TUBING HAD NOT BEEN USED; INSTEAD, TUBING FROM A DIFFERENT MANUFACTURER HAD BEEN CONNECTED TO THE BOSTON SCIENTIFIC STOPCOCK AND SYRINGES. THE PHYSICIAN BELIEVED THAT IMPROPER CATHETER PREPARATION LED TO AN AIR EMBOLISM. THE PROCEDURE WAS NOT COMPLETED DUE TO THIS EVENT. Manufacturer Narrative: G4: PREMARKET / 510(K) # K173284, K213593. Manufacturer Narrative: THE DEVICE WAS RETURNED FOR ANALYSIS. VISUAL INSPECTION REVEALED NO ISSUES, AND THE DEVICE APPEARED TO BE IN GOOD CONDITION. THE FLUSHING ACCESSORIES WERE NOT RETURNED FOR ANALYSIS. FUNCTIONAL INSPECTION OF THE MOTOR DRIVE UNIT (MDU) AND ILAB SYSTEM SHOWED THAT THE CATHETER FLUSHED NORMALLY WHEN THE TELESCOPE WAS FULLY RETRACTED AND FULLY ADVANCED. AN IMPEDANCE TEST SHOWED NO ELECTRICAL ISSUES. A CLEAR SQUARE IMAGE APPEARED ON THE ILAB SYSTEM DURING IMAGE CHARACTERIZATION TESTING. DURING THE FLUSH TEST, NO AIR BUBBLES REMAINED INSIDE THE CATHETER AFTER FLUSHING, AND NO SIGNS OF LEAKAGE WERE OBSERVED. G4: PREMARKET / 510(K) # K173284, K213593. H6 IMPACT CODES: CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22075042&pc=EOQ | 3021325287-2025-00012 | 2025-04-28 04:00 | Death | NOAH MEDICAL CORP. | 2025-05-22 04:00 | EOQ | GALAXY SYSTEM | Adverse Event Without Identified Device or Use Problem | Bronchial Hemorrhage | K223144 |  | 1 | Event Description: IT WAS REPORTED A PATIENT'S DEATH OCCURRED FOLLOWING A GALAXY SYSTEM BIOPSY PROCEDURE FOR A LESION IN THE LEFT LOWER LOBE. THE PATIENT WAS INITIALLY REPORTED TO HAVE BLED IN ACCORDANCE WITH TYPICAL BIOPSY PROTOCOL, WITH BLEEDING APPEARING TO BE CONTROLLED PRIOR TO THE GALAXY SYSTEM BEING UNDOCKED. SUBSEQUENTLY, A HEMOTHORAX WAS DISCOVERED, FORMING A CONDUIT THROUGH NECROTIC TISSUE INTO THE PLEURAL SPACE. CODE PROTOCOL WAS INITIATED, AND CONSULTATIONS WERE CONDUCTED WITH AN INTERVENTIONAL RADIOLOGIST AND THORACIC SURGEON. IT IS IMPORTANT TO NOTE THAT THE PATIENT HAD A PRE-EXISTING CHEST TUBE FROM A HEALED PNEUMOTHORAX DUE TO A NECROTIC LESION IN THE LEFT LUNG PRIOR TO THE PROCEDURE AND WAS ALSO UNDERGOING CHEMOTHERAPY. NO MALFUNCTIONS WERE REPORTED AND THE SCOPE WAS DISCARDED. EFFORTS WERE MADE TO DETERMINE THE DEFINITIVE CAUSE OF DEATH; HOWEVER, THE PHYSICIAN DID NOT PROVIDE ADDITIONAL COMMENTS AND THE FAMILY HAS CHOSEN NOT TO PURSUE AN AUTOPSY REPORT. THE PATIENT WAS ALSO DESIGNATED AS DO NOT RESUSCITATE (DNR). Manufacturer Narrative: SYSTEM AND BRONCHOSCOPE MANUFACTURING RECORDS WERE REVIEWED AND FOUND NO ISSUES ASSOCIATED WITH THIS CASE. REVIEW OF THE VIDEO LOGS REVEALED BLOOD OBSERVED DURING REGISTRATION AT 0:44. AT 00:56, BLEEDING AND TISSUE TRAUMA WERE NOTED IN THE LEFT LOWER LOBE AIRWAYS, INDICATING THAT THESE ISSUES WERE PRESENT IN THE PATIENT PRIOR TO THE GALAXY PROCEDURE. THE USER CONTINUED NAVIGATION, WITH BLEEDING OBSERVED IN THE SAME AREA AS THE TARGET LESION. THE USER WAS ABLE TO COMPLETE ALL BIOPSIES AND CONCLUDE THE PROCEDURE. IT WAS CONFIRMED BY A NOAH REPRESENTATIVE ON-SITE THAT THE EXCESSIVE BLEEDING WAS FIRST IDENTIFIED BY THE USER DURING STANDARD BRONCHOALVEOLAR LAVAGE (BAL) FOLLOWING THE GALAXY PROCEDURE. NO USE ERRORS WERE IDENTIFIED. THE SCOPE PASSED ALL THE MANUFACTURING PROCESS AND FINAL INSPECTIONS. THE PRODUCT WAS MANUFACTURED AS DESIGNED. FAILURE ANALYSIS OF THE SCOPE COULD NOT BE PERFORMED AS THE SCOPE WAS DISCARDED. THE PHYSICIAN DOES NOT ATTRIBUTE THE DEATH TO THE GALAXY SYSTEM, INSTEAD STATING THAT IT WAS CAUSED BY BLEEDING FROM THE BIOPSY. THE PATIENT'S PRE-EXISTING CONDITION CONTRIBUTED TO AN INCREASED RISK. THIS CASE IS BEING REPORTED BASED ON THE FOLLOWING CONCLUSIONS: A PATIENT'S DEATH OCCURRED FOLLOWING A GALAXY-ASSISTED BIOPSY PROCEDURE, AND STANDARD CODE PROTOCOLS WERE INITIATED. THE PHYSICIAN DOES NOT ATTRIBUTE THE DEATH TO THE GALAXY SYSTEM, INSTEAD STATING THAT IT WAS CAUSED BY BLEEDING FROM THE BIOPSY. ADDITIONALLY, THE PATIENT HAD A PRE-EXISTING CHEST TUBE FROM A HEALED PNEUMOTHORAX DUE TO A NECROTIC LESION IN THE LEFT LUNG PRIOR TO THE PROCEDURE AND WAS ALSO UNDERGOING CHEMOTHERAPY. NO MALFUNCTIONS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22075115&pc=NAY | 2955842-2025-21600 | 2025-03-29 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-05-22 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Coagulation Disorder; Ischemia; Septic Shock | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED PANCREATODUODENECTOMY PROCEDURE, THE PATIENT EXPERIENCED MULTIPLE POST-OPERATIVE COMPLICATIONS. THE PATIENT ULTIMATELY EXPIRED. THE SURGEON REPORTED THAT THERE WAS NO DA VINCI MALFUNCTION OR ISSUES RELATED TO THE DA VINCI PRODUCTS DURING THE PROCEDURE. IT WAS STATED THAT THE ROBOTIC APPROACH WAS NOT FEASIBLE TO RESECT A SECOND NEUROENDOCRINE TUMOR LOCATED IN THE MESENTERY DUE TO THE ROBOT DOCKING ISSUES [POSSIBLY FROM PORT SITE POSITIONING]; THEREFORE, THE DA VINCI SYSTEM WAS UNDOCKED AND THE PROCEDURE WAS CONVERTED TO A LAPAROSCOPIC APPROACH. DURING THE LAPAROSCOPIC DISSECTION, AN ESTIMATED BLOOD LOSS OF 1300ML OCCURRED. THE ROBOTIC SYSTEM WAS NOT UTILIZED FOR THE LYMPHADENECTOMY PORTION OF THE PROCEDURE WHERE AN ADDITIONAL 700 ML OF BLOOD LOSS OCCURRED, RESULTING IN THE PATIENT EXPERIENCING TACHYCARDIA AND HYPOTENSION. POSTOPERATIVE COMPLICATIONS BEGAN 36 HOURS LATER WHEN THE PATIENT DEVELOPED MESENTERIC ISCHEMIA, SECONDARY TO THE RESECTION OF THE NEUROENDOCRINE MESENTERIC TUMOR. A LAPAROSCOPIC ENTERECTOMY AND AN ENTERO-ANASTOMOSIS WAS PERFORMED. THE PATIENT LATER DEVELOPED A PANCREATIC FISTULA WITH A SUBHEPATIC COLLECTION, REQUIRING TWO ADDITIONAL LAPAROSCOPIC PROCEDURES FOR REVISION AND DRAINAGE. ON POSTOPERATIVE DAY (POD)19, THE PATIENT DEVELOPED SEPTIC SHOCK AND LEAKAGE OF SECRETIONS THROUGH ONE OF THE ROBOTIC PORT SITES. EXPLORATORY LAPAROTOMY FOUND A FISTULA AT THE ENTERO-ENTERO ANASTOMOSIS. THE PATIENT¿S CONDITION FURTHER DETERIORATED, RESULTING IN BLOOD DYSCRASIA AND INSTABILITY REQUIRING A SECOND EXPLORATORY LAPAROTOMY. THE PATIENT EXPIRED ON POD 20. THE CAUSE OF THE DEATH RECORDED ON THE DEATH CERTIFICATE WAS SEPTIC AND HYPOVOLEMIC SHOCK. AN AUTOPSY WAS NOT PERFORMED. THE SURGEON STATED THAT THE POST-OPERATIVE CASCADE OF COMPLICATIONS, SECONDARY TO THE NEUROENDOCRINE MESENTERIC TUMOR RESECTION, WAS THOUGHT TO BE THE CAUSE OF DEATH.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE, LARGE NEEDLE DRIVER, CADIERE FORCEPS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT DIED FOLLOWING A NUMBER OF PROCEDURES WITH SEVERAL COMPLICATIONS. THE INDEX OPERATION, A PANCREATICODUODENECTOMY, WAS INITIALLY INTENDED TO BE PERFORMED USING ROBOTIC SURGERY. HOWEVER, THIS WAS CONVERTED TO LAPAROSCOPIC AT SOME POINT DURING THE PROCEDURE TO ADDRESS AN ADDITIONAL MESENTERIC TUMOR. HOW MUCH OF THIS INDEX OPERATION WAS COMPLETED ROBOTICALLY IS NOT PROVIDED. THE CUSTOMER REPORTED THE MESENTERIC TUMOR RESECTION WAS PERFORMED WITH A THIRD PARTY DEVICE AND WAS NOT DONE ROBOTICALLY. THIS SECTION OF BOWEL BECAME ISCHEMIC REQUIRING ADDITIONAL SURGERY. ADDITIONAL COMPLICATIONS INCLUDED A PANCREATIC FISTULA, SUBHEPATIC FLUID COLLECTION, SEPSIS, AND SHOCK. THE PATIENT ULTIMATELY DIED ON POST OPERATIVE DAY 20 AFTER ADDITIONAL PROCEDURES. HOW THEY DIED AND THE SEQUENCE OF EVENTS THAT LED TO THE DEATH ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. Manufacturer Narrative: SECTION B5 ADDITIONAL INFORMATION: THE RESECTIONS AND ANASTOMOSES WERE PERFORMED ROBOTICALLY USING A THIRD-PARTY STAPLER, LIGASURE, AND A DA VINCI ROBOTIC MARYLAND BIPOLAR FORCEPS. THE PANCREATICOJEJUNOSTOMY AND HEPATICOJEJUNOSTOMY WERE SUTURED, WHILE THE GASTROJEJUNOSTOMY WAS STAPLED. THE PROCEDURE WAS THEN CONVERTED FROM ROBOTIC TO LAPAROSCOPY PRIOR TO PERFORMING LYMPHADENECTOMY AND RESECTING THE SECOND MESENTERIC TUMOR. THIS DECISION WAS MADE DUE TO THE TUMOR'S LOCATION, WHICH WOULD HAVE NECESSITATED REPOSITIONING ALL ROBOTIC PORT SITES. DURING THE LYMPHADENECTOMY, APPROXIMATELY 700 ML OF RETROPERITONEAL BLEEDING OCCURRED, RESULTING IN THE PATIENT EXPERIENCING TACHYCARDIA AND HYPOTENSION. HEMOSTASIS WAS ACHIEVED, AND THE LYMPHADENECTOMY WAS SUCCESSFULLY COMPLETED. THE MESENTERIC TUMOR, INITIALLY IDENTIFIED AS A MESENTERIC LYMPH NODE, WAS RESECTED. PATHOLOGY CONFIRMED IT TO BE A NEUROENDOCRINE TUMOR CONCURRENT WITH DUODENAL CARCINOMA. THE RESECTION RESULTED IN APPROXIMATELY 1 LITER OF BLOOD LOSS DUE TO THE TUMOR'S LOCATION AND ADHERENCE. HEMOSTASIS WAS ACHIEVED USING A LIGASURE DEVICE AND SUTURING. THIRTY-SIX HOURS POST-OPERATIVELY, THE PATIENT DEVELOPED MESENTERIC ISCHEMIA, SECONDARY TO THE RESECTION OF THE NEUROENDOCRINE MESENTERIC TUMOR. A LAPAROSCOPIC PROCEDURE WAS PERFORMED FOR AN ENTERECTOMY AND AN ENTEROANASTOMOSIS USING A THIRD-PARTY STAPLER INSTRUMENT. THE PATIENT LATER DEVELOPED A PANCREATIC FISTULA WITH A SUBHEPATIC COLLECTION, WHICH REQUIRED TWO ADDITIONAL LAPAROSCOPIC PROCEDURES FOR REVISION AND DRAINAGE. ON POST-OPERATIVE DAY 19, THE PATIENT DEVELOPED SEPTIC SHOCK AND WAS EXPERIENCING A LEAKAGE OF SECRETIONS THROUGH ONE OF THE ROBOTIC PORT SITES. AN EXPLORATORY LAPAROTOMY FOUND A FISTULA HAD FORMED AT THE ENTERO-ETEROANASTOMOSIS. LATER THAT DAY, THE PATIENT¿S CONDITION FURTHER DETERIORATED, RESULTING IN BLOOD DYSCRASIA AND INSTABILITY. A SECOND EXPLORATORY LAPAROTOMY WAS PERFORMED REVEALING A HEMOPERITONEUM, THOUGH THE SOURCE OF THE BLEEDING COULD NOT BE IDENTIFIED. DESPITE THESE EFFORTS, THE SEQUENCE OF EVENTS LEAD TO THE PATIENT DYING THE NEXT MORNING, ON POST-OPERATIVE DAY 20. A REVIEW OF THE ADDITIONAL EVENT INFORMATION PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT HAD A COMPLEX AND CHALLENGING TUMOR IN WHICH THEY UNDERWENT AN INITIAL RESECTION WHICH INCORPORATED ROBOTIC INSTRUMENTATION AS WELL AS SEVERAL 3RD PARTY DEVICES. A MESENTERIC EXTENSION OF THE TUMOR IN THE FORM OF A LARGE LYMPH NODE NECESSITATED A BOWEL RESECTION WHICH WAS PERFORMED USING A 3RD PARTY DEVICE AND THAT BOWEL WENT ON TO BECOME ISCHEMIC REQUIRING ADDITIONAL EXPLORATION AND RESECTION. A LEAK ALSO DEVELOPED BETWEEN AN ANASTOMOSIS FOLLOWING THE SECOND OPERATION AND FURTHER COMPLICATIONS INCLUDING A FISTULA OCCURRED. THE ULTIMATE CAUSE OF DEATH WAS SEPTIC AND HYPOVOLEMIC SHOCK. HOW THESE COMPLICATIONS MAY OR MAY NOT HAVE BEEN RELATED TO THE INITIAL SURGERY AND ANY ROBOTIC DEVICE IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22075496&pc=NAY | 2955842-2025-21601 | 2025-04-23 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-05-22 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Oversedation; Respiratory Arrest | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED CHOLECYSTECTOMY PROCEDURE, THE PATIENT EXPIRED. THE SURGEON STATED THAT THE ROBOTIC PROCEDURE WAS COMPLETED WITHOUT ANY INTRAOPERATIVE COMPLICATIONS; THERE WERE NO DA VINCI-RELATED ERRORS OR ISSUES. POSTOPERATIVELY, WHILE THE PATIENT WAS RECOVERING IN THE POST ANESTHESIA CARE UNIT (PACU), NARCOTICS WERE ADMINISTERED FOR PAIN. DUE TO A MISADMINISTRATION OF NARCOTICS IN THE PACU, THE PATIENT EXPIRED. THE PARTNER OF THE SURGEON CLARIFIED THAT THE PATIENT WAS RECOVERING IN THE OUTPATIENT SURGICAL CENTER¿S PACU AND WAS ADMINISTERED FENTANYL. THE SURGICAL CENTER'S PACU DID NOT HAVE NARCAN AVAILABLE AT THE TIME OF THE CODE. PARAMEDICS WERE CALLED AND ADMINISTERED NARCAN UPON ARRIVAL. ADVANCED RESUSCITATION ACTIVITIES AND CHEST COMPRESSIONS WERE CONTINUED FOR TWO HOURS. AN INITIAL POST-PROCEDURE ULTRASOUND SHOWED NO SIGNS OF BLOOD WITHIN THE ABDOMEN. AT AN UNSPECIFIED TIME DURING RESUSCITATION, A REPEAT ULTRASOUND DID SHOW SIGNS OF BLOOD IN THE ABDOMEN. IT WAS PLANNED TO TRANSFER THE PATIENT TO THE HOSPITAL'S MAIN OPERATING ROOM FOR A REOPERATION DUE TO THE FINDINGS; HOWEVER, THE PATIENT EXPIRED IN THE PACU AND THE SUBSEQUENT PROCEDURE DID NOT OCCUR. IT WAS STATED THAT THE NEW BLEEDING WITHIN THE ABDOMEN WAS BELIEVED TO BE DUE TO THE CHEST COMPRESSIONS, POSSIBLY CRACKED RIBS. Manufacturer Narrative: REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES FOUND THERE WERE NO INTERRUPTIONS OR ERROR EVENTS IMPACTING ANY OF THE PROCEDURES. THE SYSTEM WAS FUNCTIONING TO MANUFACTURING SPECIFICATIONS. THE DAY AFTER THE EVENT, THE CUSTOMER REQUESTED A FIELD SERVICE ENGINEER (FSE) TO CHECK OUT THE SYSTEM EVEN THOUGH ¿THEY DID NOT BELIEVE THE ROBOT HAD ANYTHING TO DO WIGHT THE ADVERSE EVENT¿ TO VERIFY THE EQUIPMENT. THE FSE TESTED THE SYSTEM, INCLUDING THE E-100 AND ERBE GENERATORS WITH NO ISSUES FOUND. ELECTRICAL SAFETY TESTING OF ALL THE CARTS AND GENERATORS WAS ALSO PERFORMED, FOLLOWED BY A SECOND SYSTEM TEST DRIVE. NO ISSUES WERE FOUND. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, PERMANENT CAUTERY HOOK, MARYLAND BIPOLAR FORCEPS, CADIERE FORCEPS, PROGRASP FORCEPS, MEDIUM-LARGE CLIP APPLIERS X2. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING OR SUBSEQUENT TO THIS PROCEDURE. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICES USED DURING THE REPORTED EVENT FOUND NO RELATED NON-CONFORMANCES. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT UNDERWENT A CHOLECYSTECTOMY, WHERE NO INTRAOPERATIVE COMPLICATIONS OR ISSUES WERE REPORTED. POSTOPERATIVELY, THE PATIENT RECEIVED FENTANYL AND CODED. THE TEMPORAL RELATIONSHIP OF THE CARDIOVASCULAR COLLAPSE SUGGESTS THE EVENT WAS PRECIPITATED BY THE NARCOTIC AS AN INITIAL ULTRASOUND SHOWED NO EVIDENCE OF FLUID IN THE ABDOMEN. ALTHOUGH A SUBSEQUENT ULTRASOUND EVENTUALLY SHOWED FLUID IN THE ABDOMEN, THIS FINDING WAS SPECULATED TO BE A RESULT OF THE RESUSCITATION EFFORTS. CONTINUED RESUSCITATION EFFORTS WERE UNSUCCESSFUL, AND THE PATIENT EXPIRED. THE CAUSE OF DEATH WAS SPECULATED TO HAVE BEEN PRECIPITATED BY THE ADMINISTRATION OF THE FENTANYL NARCOTIC. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22075552&pc=NAY | 2955842-2025-21939 | 2025-04-13 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-05-22 04:00 | NAY | DAVINCI X | Adverse Event Without Identified Device or Use Problem | Ischemia; Sepsis | K171294 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT DA VINCI-ASSISTED RECTOSIGMOIDECTOMY AND PLANNED CHOLECYSTECTOMY PROCEDURE. THERE WERE NO INTRAOPERATIVE COMPLICATIONS DURING THE PROCEDURE WHICH WAS SUCCESSFULLY COMPLETED. THE PATIENT SUBSEQUENTLY EXPIRED ON POST-OPERATIVE DAY (POD) 6 FROM COMPLICATIONS. IT WAS REPORTED THAT BY POD 4, THE PATIENT EXHIBITED SIGNS OF SEPSIS. A CT SCAN SHOWED FREE FLUID IN THE PELVIS, AND ALTHOUGH THERE WAS NO FISTULA DETECTED, THE FINDINGS WERE SUSPECTED TO BE CAUSED BY FOURNIER'S GANGRENE. THE PATIENT UNDERWENT AN EXPLORATORY LAPAROTOMY "BY VIDEO" FOR PELVIC DRAINAGE. TWO DAYS LATER, AS A RESULT OF THE FOURNIER'S GANGRENE, THE PATIENT WAS RETURNED TO THE SURGICAL CENTER ON MECHANICAL VENTILATION AND WITH VASOACTIVE INFUSIONS (NOREPINEPHRINE AND VASOPRESSIN) FOR EMERGENCY DEBRIDEMENT OF THE SCROTAL SAC. THE PATIENT EXPIRED DURING THE PROCEDURE. AN AUTOPSY WAS NOT PERFORMED. THE CAUSE OF DEATH RECORDED ON THE DEATH CERTIFICATE WAS FOURNIER'S GANGRENE. ADDITIONAL DETAILS REGARDING THE SEQUENCE OF EVENTS AND TREATMENT WERE REQUESTED; NO RESPONSE HAS BEEN RECEIVED. Manufacturer Narrative: SEVERAL PROCEDURES WERE PERFORMED ON THE PROVIDED INITIAL ROBOTIC PROCEDURE DATE. INSUFFICIENT INFORMATION IS AVAILABLE FOR PROCEDURE VERIFICATION; THEREFORE, NO DA VINCI SYSTEM OR INSTRUMENT LOGS ARE AVAILABLE FOR REVIEW. CONFIRMATION OF THE SPECIFIC PROCEDURE TIME HAS BEEN REQUESTED. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT UNDERWENT 2 CONCURRENT ROBOTIC PROCEDURES, A CHOLECYSTECTOMY AND A RECTOSIGMOIDECTOMY. THE PATIENT WAS ON INOTROPIC SUPPORT AND MECHANICAL VENTILATION FOLLOWING THE PROCEDURE, BUT THE SPECIFIC DETAILS OF WHY THIS LEVEL OF SUPPORT WAS REQUIRED IS NOT PROVIDED. THE PATIENT THEN DEVELOPED SEPSIS AND FOURNIER¿S GANGRENE WHICH IS A CONDITION WITH HIGH MORTALITY. HOW THEY DEVELOPED THIS CONDITION IS NOT PROVIDED. THE PATIENT EVENTUALLY DIED AS A RESULT OF THESE COMPLICATIONS. HOW THESE EVENTS WERE RELATED TO THE INITIAL PROCEDURE IS NOT PROVIDED. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22058760&pc=PNJ | 9612164-2025-02504 | 2025-04-24 04:00 | Death | MEDTRONIC IRELAND | 2025-05-21 04:00 | PNJ | MICRA¿ VR2 | Failure to Capture; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT INTRA OPERATIVELY DURING THE PLACEMENT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), NO CAPTURE COULD BE OBTAINED AT MULTIPLE LOCATIONS. THE LEADLESS IPG WAS ATTEMPTED NOT USED AND REPLACED. LATER, A NEW LEADLESS IPG SYSTEM WAS IMPLANTED SUCCESSFULLY WITHOUT ANY COMPLICATIONS. ONE WEEK LATER, IT WAS NOTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS IDENTIFIED TO BE PULSELESS ELECTRICAL ACTIVITY (PEA) ARREST, HYPOTENSION AND SEVERE HYPOXIA. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2VR01-DELSYS] (SERIAL: [(B)(6)]). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ID MC2VR01-DELSYS PRODUCT EVENT SUMMARY: A PARTIAL PORTION OF THE DELIVERY SYSTEM WAS RETURNED AND ANALYZED. THE DELIVERY SYSTEM TETHER WAS BROKEN/CUT/PULLED APART. VISUAL ANALYSIS OF THE DELIVERY SYSTEM INDICATED DAMAGE DURING USE. THE ANALYST NOTE ONLY A PIECE OF THE TETHER WAS RETURNED IN THE RECAPTURE FEATURE OF THE DEVICE (B)(6). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MC2VR01-DELSYS SN:(B)(6). CORRECTION H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22058763&pc=PNJ | 9612164-2025-02505 | 2025-04-24 04:00 | Death | MEDTRONIC IRELAND | 2025-05-21 04:00 | PNJ | MICRA¿ VR2 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT INTRA OPERATIVELY DURING THE PLACEMENT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), NO CAPTURE COULD BE OBTAINED AT MULTIPLE LOCATIONS. THE LEADLESS IPG WAS ATTEMPTED NOT USED AND REPLACED. LATER, A NEW LEADLESS IPG SYSTEM WAS IMPLANTED SUCCESSFULLY WITHOUT ANY COMPLICATIONS. ONE WEEK LATER, IT WAS NOTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS IDENTIFIED TO BE PULSELESS ELECTRICAL ACTIVITY (PEA) ARREST, HYPOTENSION AND SEVERE HYPOXIA. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2VR01-DELSYS] (SERIAL: [(B)(6)]). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22058783&pc=PNJ | 9612164-2025-02506 | 2025-04-28 04:00 | Death | MEDTRONIC IRELAND | 2025-05-21 04:00 | PNJ | MICRA AV | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT INTRA OPERATIVELY DURING THE PLACEMENT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE PATIENT EXPERIENCED CARDIAC PERFORATION, PERICARDIAL TAMPONADE AND EFFUSION. PERICARDIOCENTESIS AND DRAINAGE WERE PERFORMED. THE LEADLESS IPG SYSTEM WAS ATTEMPTED NOT USED AND THE IMPLANT ATTEMPT WAS ABANDONED. EVENTUALLY IT WAS NOTED THAT THE PATIENT PASSED AWAY. THE PATIENT IS A PARTICIPANT IN THE POST APPROVAL CLINICAL SURVEILLANCE PRODUCT SURVEILLANCE REGISTRY. Manufacturer Narrative: CORRECTION H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC1AVR1-DELSYS] (SERIAL: [(B)(6)]). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22058787&pc=DYB | 9612164-2025-02507 | 2025-04-28 04:00 | Death | MEDTRONIC IRELAND | 2025-05-21 04:00 | DYB | MICRA | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion; Insufficient Information | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT INTRA OPERATIVELY DURING THE PLACEMENT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE PATIENT EXPERIENCED CARDIAC PERFORATION, PERICARDIAL TAMPONADE AND EFFUSION. PERICARDIOCENTESIS AND DRAINAGE WERE PERFORMED. THE LEADLESS IPG SYSTEM WAS ATTEMPTED NOT USED AND THE IMPLANT ATTEMPT WAS ABANDONED. EVENTUALLY IT WAS NOTED THAT THE PATIENT PASSED AWAY. THE PATIENT IS A PARTICIPANT IN THE POST APPROVAL CLINICAL SURVEILLANCE PRODUCT SURVEILLANCE REGISTRY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22059192&pc=NVZ | 3008973940-2025-03598 | 2025-04-13 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-05-21 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Sepsis | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED SEPSIS SEQUELAE. EVENTUALLY, IT WAS NOTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22059218&pc=NVZ | 9614453-2025-02006 | 2025-04-24 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-05-21 04:00 | NVZ | ASTRA S DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY TWO DAYS AFTER THE IMPLANT OF THE LEADS. ADDITIONAL INFORMATION OBTAINED SUGGESTED THAT THE CAUSE OF DEATH COULD NOT BE IDENTIFIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22059414&pc=OZD | 1220648-2025-28591 | 2025-05-14 04:00 | Death | ABIOMED, INC. | 2025-05-21 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ESCALATED FROM INTRA-AORTIC BALLOON PUMP (IABP) TO IMPELLA 5.5 IN ANTICIPATION OF VENTRICULAR SEPTAL PERFORATION (VSP) CLOSURE. DURING IMPELLA 5.5 MANAGEMENT, THE PURGE PRESSURE ROSE, AND THE PURGE FLOW RATE DROPPED WITH THE PRESSURE RISE, CAUSING PURGE-RELATED ALARMS. THERE WERE NO PROBLEMS WITH THE PUMP FLOW RATE OR MOTOR CURRENT CONSUMPTION. ALTHOUGH THE PURGE SYSTEM WAS REPLACED AND CHECKED FOR BENDS, NO IMPROVEMENT WAS FOUND, SO THE IMPELLA 5.5 DEVICE WAS REPLACED. IT WAS NOTED THAT WHEN THE EXISTING IMPELLA 5.5 WAS REMOVED, A THROMBUS WAS FOUND TO HAVE ADHERED TO THE OUTFLOW SITE. THE DEVICE WAS REPLACED WITHOUT ANY ISSUES AND WITHOUT AFFECTING THE PATIENT'S HEMODYNAMICS. THE PATIENT IS IN STABLE CONDITION.Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED WHILE ON SUPPORT, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH.". IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ IMPELLA 5.5 WITH SMARTASSIST: SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE AND THROMBUS HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. THE DATA LOG SHOWS A 30-MINUTE GRADUAL RISE IN PURGE PRESSURE, FOLLOWED BY A HIGH PURGE PRESSURE ALARM WITHOUT ANY OBSERVED MOTOR CURRENT SPIKE. PURGE PRESSURE WAS SEEN TO RECOVER WITH A GRADUAL TREND AFTER 12 HOURS. THE CAUSE OF THE HIGH PURGE PRESSURE ISSUE WAS MOST LIKELY BIOMATERIAL BUILDUP, BASED ON DATA LOG REVIEW. THE ROOT CAUSE OF THE THROMBUS WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. B2 DEATH AND DATE OF DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28591. B5 PATIENT OUTCOME WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28591. H1 TYPE OF REPORT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28591. H6 HEALTH EFFECT - IMPACT CODE 1802 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28591.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22059777&pc=CBK | 2518422-2025-105916 | 2024-12-23 05:00 | Death | RESPIRONICS, INC. | 2025-05-21 04:00 | CBK | TRILOGY EVO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT PASSED AWAY WHILE USING THE TRILOGY EVO DEVICE. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CUSTOMER STATED THE DEATH WAS EXPECTED. THE PATIENT WAS ON HOSPICE, SO DID NOT WISH TO COMPLETE THIS OR HAVE THE UNIT SENT IN DUE TO THE EXPECTED DEATH. CUSTOMER DID NOT WISH TO ANSWER ANY FURTHER QUESTIONS MORE THAN WHAT THEY HAD ALREADY DONE SO. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING AND WILL MAKE ATTEMPTS TO RETRIEVE THE DEVICE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. PMS DECISION: THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USER ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED. BASED ON AVAILABLE INFORMATION AT THIS TIME, THE EXPECTED PATIENT DEATH AS THE PATIENT WAS ON HOSPICE CARE IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. NO FURTHER ACTION IS REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22060177&pc=ITI | 2530130-2025-00846 | 2025-05-07 04:00 | Death | ZHEJIANG INNUOVO REHABILITATION DEVICES CO.,LTD | 2025-05-21 04:00 | ITI | PRIDE MOBILITY PRODUCTS | Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | K220747 |  | 1 | Event Description: CONSUMER'S WIFE ALLEGES HER HUSBAND ALLEGEDLY FELL OUT OF THE CHAIR WENT TO HOSPITAL AND PASSED AWAY. SHE SAID IT WAS NOT DUE TO THE CHAIR, HE WAS VERY WEAK WHICH CAUSED HIM TO FALL.Event Description: CONSUMER'S WIFE ALLEGES HER HUSBAND ALLEGEDLY FELL OUT OF THE CHAIR WENT TO HOSPITAL AND PASSED AWAY. SHE SAID IT WAS NOT DUE TO THE CHAIR, HE WAS VERY WEAK WHICH CAUSED HIM TO FALL. Manufacturer Narrative: THE DEVICE HAS NOT YET BEEN MADE AVAILABLE FOR EVALUATION. SHOULD FURTHER INFORMATION OR THE DEVICE BECOME AVAILABLE, A FOLLOW-UP REPORT WILL THEN BE ISSUED. Manufacturer Narrative: DEVICE NOT ATTRIBUTED TO EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22060224&pc=DSQ | 2916596-2025-03212 | 2025-05-07 04:00 | Death | THORATEC CORPORATION | 2025-05-21 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Anemia | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ON HOME HOSPICE AND PASSED AWAY DUE TO CHRONIC ANEMIA AND FAILURE TO THRIVE (FTT). THE OUTCOME WAS NOT CONSIDERED TO HAVE BEEN DEVICE OR THERAPY RELATED. THE PUMP HAD OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ CONTAINS INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22060373&pc=NPT | 2124215-2025-30640 | 2020-09-03 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | NPT | LOTUS EDGE? VALVE SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P180029 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. IN (B)(6) 2020, THE PATIENT UNDERWENT A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE AND WAS SUCCESSFULLY IMPLANTED WITH A 25MM LOTUS EDGE VALVE. TWENTY-TWO (22) DAYS POST PROCEDURE THE PATIENT DIED OF UNREPORTED CAUSES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22060628&pc=NPT | 2025587-2025-03805 | 2025-05-14 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-21 04:00 | NPT | EVOLUT FX PLUS VALVE | Malposition of Device; Device Dislodged or Dislocated; Patient Device Interaction Problem; Migration | Foreign Body In Patient; Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT A TRANSCATHETER AORTIC VALVE WAS PLACED IN A BICUSPID VALVE WITH MINIMAL CALCIFICATION IN THE NON-CORONARY CUSP (NCC) AND MIGRATED INTO THE ASCENDING AORTA. IT WAS SUGGESTED THAT THE MIGRATION OCCURRED BECAUSE THE VALVE WAS PLACED TOO HIGH (1 MM ON THE NCC). A SECOND VALVE WAS POSITIONED LOWER AT 3 MM WITH AN ACCEPTABLE OUTCOME. TRACE-MILD PARAVALVULAR LEAK (PVL) WAS OBSERVED AND A LARGE RAPHE. THE PATIENT DIED DUE TO MULTIPLE AORTIC DISSECTIONS, WITH A POSSIBLE CAUSE BEING THE EMBOLIZED VALVE SNARED UP INTO THE ASCENDING AORTA.Event Description: IT WAS REPORTED THAT A TRANSCATHETER AORTIC VALVE WAS PLACED IN A BICUSPID VALVE WITH MINIMAL CALCIFICATION IN THE NON-CORONARY CUSP (NCC) AND MIGRATED INTO THE ASCENDING AORTA. IT WAS SUGGESTED THAT THE MIGRATION OCCURRED BECAUSE THE VALVE WAS PLACED TOO HIGH (1 MM ON THE NCC). A SECOND VALVE WAS POSITIONED LOWER AT 3 MM WITH AN ACCEPTABLE OUTCOME. TRACE-MILD PARAVALVULAR LEAK (PVL) WAS OBSERVED AND A LARGE RAPHE. THE PATIENT DIED DUE TO MULTIPLE AORTIC DISSECTIONS, WITH A POSSIBLE CAUSE BEING THE EMBOLIZED VALVE SNARED UP INTO THE ASCENDING AORTA. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT FOLLOWING THE VALVE IMPLANT PROCEDURE, THE PATIENT WAS TRANSFERRED TO INTERMEDIATE CARE WHICH IS WHEN THE DISSECTION WAS FIRST NOTED. ADDITIONALLY, IT WAS NOTED THE PATIENT HAD MULTIPLE STROKES, THEREFORE, NO FURTHER INTERVENTION WAS MADE.Event Description: IT WAS REPORTED THAT A TRANSCATHETER AORTIC VALVE WAS PLACED IN A BICUSPID VALVE WITH MINIMAL CALCIFICATION IN THE NON-CORONARY CUSP (NCC) AND MIGRATED INTO THE ASCENDING AORTA. IT WAS SUGGESTED THAT THE MIGRATION OCCURRED BECAUSE THE VALVE WAS PLACED TOO HIGH (1 MM ON THE NCC). A SECOND VALVE WAS POSITIONED LOWER AT 3 MM WITH AN ACCEPTABLE OUTCOME. TRACE-MILD PARAVALVULAR LEAK (PVL) WAS OBSERVED AND A LARGE RAPHE. THE PATIENT DIED DUE TO MULTIPLE AORTIC DISSECTIONS, WITH A POSSIBLE CAUSE BEING THE EMBOLIZED VALVE SNARED UP INTO THE ASCENDING AORTA. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT FOLLOWING THE VALVE IMPLANT PROCEDURE, THE PATIENT WAS TRANSFERRED TO INTERMEDIATE CARE WHICH IS WHEN THE DISSECTION WAS FIRST NOTED. ADDITIONALLY, IT WAS NOTED THE PATIENT HAD MULTIPLE STROKES, AND; THEREFORE, NO FURTHER INTERVENTION WAS MADE. ADDITIONAL INFORMATION WAS RECEIVED THAT THE LOCATION OF THE DISSECTIONS WERE LIMITED TO THE ASCENDING AORTA. NO PRE-IMPLANT OR POST-IMPLANT BALLOON DILATIONS WERE PERFORMED. THE FIRST VALVE WAS PULLED BACK TO THE ASCENDING AORTA BEFORE IMPLANTING THE SECOND VALVE.Event Description: IT WAS REPORTED THAT A TRANSCATHETER AORTIC VALVE WAS PLACED IN A BICUSPID VALVE WITH MINIMAL CALCIFICATION IN THE NON-CORONARY CUSP (NCC) AND MIGRATED INTO THE ASCENDING AORTA. IT WAS SUGGESTED THAT THE MIGRATION OCCURRED BECAUSE THE VALVE WAS PLACED TOO HIGH (1 MM ON THE NCC). A SECOND VALVE WAS POSITIONED LOWER AT 3 MM WITH AN ACCEPTABLE OUTCOME. TRACE-MILD PARAVALVULAR LEAK (PVL) WAS OBSERVED AND A LARGE RAPHE. THE PATIENT DIED DUE TO MULTIPLE AORTIC DISSECTIONS, WITH A POSSIBLE CAUSE BEING THE EMBOLIZED VALVE SNARED UP INTO THE ASCENDING AORTA. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT FOLLOWING THE VALVE IMPLANT PROCEDURE, THE PATIENT WAS TRANSFERRED TO INTERMEDIATE CARE WHICH IS WHEN THE DISSECTION WAS FIRST NOTED. ADDITIONALLY, IT WAS NOTED THE PATIENT HAD MULTIPLE STROKES, AND; THEREFORE, NO FURTHER INTERVENTION WAS MADE. ADDITIONAL INFORMATION WAS RECEIVED THAT THE LOCATION OF THE DISSECTIONS WERE LIMITED TO THE ASCENDING AORTA. NO PRE-IMPLANT OR POST-IMPLANT BALLOON DILATIONS WERE PERFORMED. THE FIRST VALVE WAS PULLED BACK TO THE ASCENDING AORTA BEFORE IMPLANTING THE SECOND VALVE. ADDITIONAL INFORMATION WAS ADDED TO CLARIFY THE "MIGRATION" OCCURRED PRIOR TO THE CLOSE OF THE VASCULAR ACCESS AT THE END OF THE PROCEDURE AND, THEREFORE, WAS A VALVE DISLODGEMENT, NOT A MIGRATION.Event Description: IT WAS REPORTED THAT A TRANSCATHETER AORTIC VALVE WAS PLACED IN A BICUSPID VALVE WITH MINIMAL CALCIFICATION IN THE NON-CORONARY CUSP (NCC) AND MIGRATED INTO THE ASCENDING AORTA. IT WAS SUGGESTED THAT THE MIGRATION OCCURRED BECAUSE THE VALVE WAS PLACED TOO HIGH (1 MM ON THE NCC). A SECOND VALVE WAS POSITIONED LOWER AT 3 MM WITH AN ACCEPTABLE OUTCOME. TRACE-MILD PARAVALVULAR LEAK (PVL) WAS OBSERVED AND A LARGE RAPHE. THE PATIENT DIED DUE TO MULTIPLE AORTIC DISSECTIONS, WITH A POSSIBLE CAUSE BEING THE EMBOLIZED VALVE SNARED UP INTO THE ASCENDING AORTA. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT FOLLOWING THE VALVE IMPLANT PROCEDURE, THE PATIENT WAS TRANSFERRED TO INTERMEDIATE CARE WHICH IS WHEN THE DISSECTION WAS FIRST NOTED. ADDITIONALLY, IT WAS NOTED THE PATIENT HAD MULTIPLE STROKES, AND; THEREFORE, NO FURTHER INTERVENTION WAS MADE. ADDITIONAL INFORMATION WAS RECEIVED THAT THE LOCATION OF THE DISSECTIONS WERE LIMITED TO THE ASCENDING AORTA. NO PRE-IMPLANT OR POST-IMPLANT BALLOON DILATIONS WERE PERFORMED. THE FIRST VALVE WAS PULLED BACK TO THE ASCENDING AORTA BEFORE IMPLANTING THE SECOND VALVE. ADDITIONAL INFORMATION WAS ADDED TO CLARIFY THE "MIGRATION" OCCURRED PRIOR TO THE CLOSE OF THE VASCULAR ACCESS AT THE END OF THE PROCEDURE AND, THEREFORE, WAS A VALVE DISLODGEMENT, NOT A MIGRATION. Manufacturer Narrative: CORRECTED DATA: H6. CONCLUSION: THE SUSPECT VALVE REMAINED IN THE PATIENT AT THE TIME THIS INVESTIGATION WAS COMPLETED; THEREFORE, NO PRODUCT ANALYSIS COULD BE PERFORMED. BASED ON THE AVAILABLE INFORMATION AND INVESTIGATION ACTIVITIES, THE REPORTED EVENT MAY HAVE BEEN RELATED TO THE MEDTRONIC DEVICE, BUT THE EXACT RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED EVENT CANNOT BE DETERMINED DEFINITIVELY. THE COMPLAINT INVESTIGATION DETERMINED THIS EVENT IS FORESEEN IN RISK MANAGEMENT AND IS INCLUDED IN MONITORING/TRENDING. THE DEVICE HISTORY RECORD WAS REVIEWED AND SHOWED THIS PRODUCT MET ALL MANUFACTURING SPECIFICATIONS FOR PRODUCT RELEASED FOR DISTRIBUTION. NO ISSUES WERE IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. A MEDICAL SAFETY ASSESSMENT WAS PERFORMED FOR THIS EVENT. BASED ON REVIEW OF AVAILABLE INFORMATION, THE SEQUENCE OF EVENTS, AND MEDICAL EXPERTISE THE EVENT OF DEVICE DISLODGEMENT WAS LIKELY RELATED TO THE DEVICE AND PROCEDURE. IN SUMMARY, THE MEDICAL SAFETY ASSESSMENT DETERMINED THE LIKELY CONTRIBUTING FACTOR WAS A TOO HIGH PLACEMENT OF THE VALVE WHICH CAUSED THE DEVICE TO DISLODGE AORTICALLY IN THE SETTING OF A NATIVE BICUSPID VALVE. THE AORTIC DISSECTION(S) WERE LIKELY RELATED TO THE VALVE WHEN IT MIGRATED INTO THE ASCENDING AORTA AND/OR WHEN IT WAS SNARED FURTHER INTO THE ASCENDING AORTA. THE PATIENT¿S DEATH WAS ASSESSED AS LIKELY RELATED TO THE COMPLICATIONS FROM THE REPORTED MULTIPLE DISSECTIONS. THE DEVICE INSTRUCTIONS FOR USE (IFU) LISTS DISLODGEMENT AND DISSECTION AS POTENTIAL RISKS ASSOCIATED WITH THE TREATMENT PROCEDURES. THERE IS NO IDENTIFIED INADEQUACY WITH THE IFU IN RELATION TO THIS EVENT. POTENTIAL FACTORS THAT CAN INFLUENCE A DISLODGED VALVE INCLUDE TENSION APPLIED ON THE DCS DURING POSITIONING, CALCIFICATION LEVELS IN THE NATIVE VESSEL, COMPLIANCE OF THE AORTA AND NATIVE VESSELS, INADEQUATE DEPLOYMENT DUE TO IMPROPER SIZING BETWEEN VALVE AND ANNULUS, IRREGULAR PATIENT ANATOMY, INCOMPLETE FRAME EXPANSION, AND A NUMBER OF OTHERS. HOWEVER, A CONCLUSIVE CAUSE COULD NOT BE DETERMINED. VASCULAR COMPLICATIONS, SUCH AS DISSECTION ARE LISTED IN THE DEVICE IFU AS POTENTIAL ADVERSE EFFECTS AND CAN OCCUR DURING ANY PERCUTANEOUS INTERVENTION. THESE EVENTS ARE TYPICALLY RELATED TO PATIENT FACTORS (TORTUOUS ANATOMY, CALCIFICATION, ETC.), AND/OR PROCEDURAL EFFECTS (SHEATH USED, USER TECHNIQUE, GUIDEWIRE MANAGEMENT, ETC.). THE VALVE DISLODGEMENT MAY HAVE BEEN THE CONTRIBUTING FACTOR TO DISSECTION. UNFORTUNATELY, THIS COULD NOT BE CONFIRMED BY MEDTRONIC WITHOUT A COPY OF THE ECHO OR IMAGING FILM FOR REVIEW, THE FAILURE MECHANISM OF THE DEVICE CANNOT BE ESTABLISHED, AND THE CONCLUSIVE CAUSE OF THE REPORTED DISSECTION CANNOT BE DETERMINED. PATIENT DEATH IS A POTENTIAL RISK ASSOCIATED WITH THE IMPLANTATION OF A BIOPROSTHESIS VALVE PER THE DEVICE IFU. A PROCEDURE- OR VALVE-RELATED DEATH IS AN INHERENT RISK WHEN THE PATIENT CONDITION IS SUCH THAT A TAV IS NEEDED TO SUSTAIN CARDIAC FUNCTION, AND IT CAN OCCUR DESPITE AN IDEAL IMPLANT PROCEDURE OR DEVICE FUNCTIONALITY. THE DEATH IS ASSESSED AS LIKELY RELATED TO THE COMPLICATIONS FROM THE REPORTED MULTIPLE DISSECTIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4 H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED DATA: B5. A FOURTH PARAGRAPH WAS ADDED TO CLARIFY THE VALVE DISLODGED AND DIDN'T MIGRATE. ALTHOUGH BOTH SCENARIOS WOULD BE REPORTED AS A DEVICE MALFUNCTION WHERE THE CHANCE OF DEATH OR SERIOUS INJURY IS NOT REMOTE IF TO RECUR, A DIFFERENT ANNEX A CODE - MEDICAL DEVICE PROBLEM IS ASSIGNED TO EACH OF THESE MALFUNCTIONS. LET THIS SUPPLEMENTAL 004 ACCURATELY REFLECT A DISLODGED VALVE. H6. LET THIS REPORT ACCURATELY REFLECT THE ANNEX A CODE FOR DISLODGEMENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2. DATE OF DEATH B5. A SECOND PARAGRAPH WAS ADDED. SECTION D OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID: EVFXPLUS-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: EVOLUT FX DCS, PRODUCT ID: D-EVOLUTFX-2329, LOT: 0012665445, USE BY DATE: 2027-02-11, UDI: (B)(4). SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012617235); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012676208); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVFXPLUS-29 (UNKNOWN SERIAL); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22060638&pc=BZD | 2518422-2025-045231 | 2021-09-27 04:00 | Death | RESPIRONICS, INC. | 2025-05-21 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22060688&pc=LWP | 2017865-2025-53933 | 2024-10-06 04:00 | Death | ABBOTT | 2025-05-21 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. THERE IS NO CLEAR INFORMATION AS TO WHETHER THE DEATH WAS DEVICE-RELATED, HOWEVER THE DEVICE WAS EXPLANTED. ADDITIONAL INFORMATION WAS REQUESTED YET NOT RECEIVED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22060802&pc=DYB | 9612164-2025-02524 | 2025-04-24 04:00 | Death | MEDTRONIC IRELAND | 2025-05-21 04:00 | DYB | SENTRANT INTRODUCER SHEATH | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Perforation of Vessels | K123990 |  | 1 | Event Description: A SENTRANT SHEATH WAS USED IN A TAVI PROCEDURE BEFORE THE PRE-DILATION OF THE AORTIC VALVE. IT WAS NOTED THAT DUE TO SEVERE CALCIFICATION OF THE FEMORAL ILIAC PATHWAY, THE ARTERIES HAD A SMALLER DIAMETER THAN RECOMMENDED. IT WAS STATED THE VASCULAR SPECIALIST WAS AWARE OF THIS BUT STILL DECIDED TO ATTEMPT THE PROCEDURE VIA THE RIGHT FEMORAL ARTERY. DURING THE PROCEDURE, AFTER PREPARING THE PRO PLUS 23 LOT: J304417 IN D-EVPROP23-29, THE 14F SENTRANT WAS INTRODUCED INTO THE RIGHT FEMORAL ARTERY WITH DIFFICULTY. THE VASCULAR TEAM PERFORMED ANGIOPLASTY IN THE COMMON ILIAC AND RIGHT EXTERNAL ILIAC WITH A 7X40 BALLOON, PRE-DILATED THE AORTIC VALVE WITH A NON-MDT 18X40 BALLOON. THE 14F SENTRANT WAS REMOVED AND THE VASCULAR TEAM WAS UNABLE TO PROGRESS WITH THE VALVE DELIVERY SYSTEM THROUGH THE RIGHT FEMORAL ARTERY, SO A NEW ANGIOPLASTY WAS PERFORMED WITH A 7X120 BALLOON IN THE COMMON AND EXTERNAL ILIAC ARTERY AND COMMON FEMORAL ARTERY. AFTER IMAGE CONTROL, THE DOCTORS SHOWED BLOOD EXTRAVASATION IN THE EXTERNAL ILIAC AND FEMORAL ARTERY AND SEVERE HYPOTENSION. A 7X58 COVERED STENT WAS IMPLANTED IN THE EXTERNAL ILIAC, AND OPEN BYPASS SURGERY WITH GRAFT WAS CHOSEN. THE PATIENT WAS TREATED WITH 5 RED BLOOD CELL CONCENTRATES, OROTRACHEAL INTUBATION, NOREPINEPHRINE IN A CONTINUOUS PUMP, AND BYPASS SURGERY. IT WAS DECIDED NOT TO IMPLANT THE VALVE DUE TO VASCULAR SEVERITY. DESPITE THESE INTERVENTIONS, THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT, HAD TWO CARDIORESPIRATORY ARRESTS, AND DIED. ALTHOUGH THERE WAS NO ALLEGED PRODUCT ISSUE AGAINST THE SENTRANT SHEATH AND THE DEVICE DID NOT PRESENT ANY DEFECTS, IT IS BELIEVED THAT THE SENTRANT SHEATH MAY HAVE CONTRIBUTED TO THE REPORTED EVENTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22060933&pc=MKJ | 1220908-2025-01708 | 2025-04-27 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-05-21 04:00 | MKJ | X SERIES | Unable to Obtain Readings | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE WAS UNABLE TO OBTAIN AN ECG SIGNAL. COMPLAINANT INDICATED THAT AFTER THE PATIENT EVENT, DURING DEVICE TESTING, THE DEVICE WAS UNABLE TO RECOGNIZE THE ATTACHED ELECTRODE PADS. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED, HOWEVER IT WAS NOT A RESULT OF THE REPORTED MALFUNCTION. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS RECEIVED THE PRODUCT AND WILL BE PROVIDING A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS COMPLETED. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. REVIEW OF THE DEVICE LOG CONFIRMED CPR ACTIVITY WITH VALID PATIENT IMPEDANCE DETECTED AND LEADS PROPERLY CONNECTED. THE DEVICE WAS SET TO LEAD I VIEW, SO ECG WAVEFORMS WERE NOT DISPLAYED, SWITCHING TO PADS VIEW WOULD HAVE DISPLAYED ACTIVE ECG. IT WAS DETERMINED AFTER CONVERSATION WITH THE CUSTOMER THAT THIS DEVICE HAD A DIFFERENT USER CONFIGURATION THAN MOST THEIR DEVICES. THIS DEVICE WAS CONFIGURED TO START UP IN LEAD I, BUT MOST OTHER DEVICES PADS. THE CUSTOMER WILL BE CHECKING USER CONFIGURATIONS TO ASSURE THEY ARE ALL THE SAME TO REDUCE CONFUSION. ALL DEVICE TESTS PASSED, WITH NO ISSUES FOUND IN FUNCTIONALITY, ECG ACQUISITION, OR USER INTERFACE. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS FOR REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22061322&pc=LWS | 2649622-2025-13825 | 2025-05-09 04:00 | Death | MPRI | 2025-05-21 04:00 | LWS | SPRINT QUATTRO SECURE MRI SURESCAN | Over-Sensing; Defibrillation/Stimulation Problem; Under-Sensing | Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA) AND OVERSENSING AND INTERMITTENT UNDERSENSING WERE OBSERVED. IT WAS ALSO NOTED THAT THE LIA EXTENDED THE NUMBER OF INTERVALS NEEDED TO DETECT WHICH MAY HAVE DELAYED THERAPY FOR VENTRICULAR FIBRILLATION (VF). SUBSEQUENTLY, THE PATIENT DIED.Event Description: IT WAS FURTHER REPORTED THAT THE LEAD INTEGRITY ALERT (LIA) TRIGGERED FOR HIGH RATE, NON-SUSTAINED EPISODES AND SHORT VENTRICULAR INTERVALS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANTED: (B)(6) 2010. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22061690&pc=BZD | 2518422-2025-045232 | 2022-08-16 04:00 | Death | RESPIRONICS, INC. | 2025-05-21 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22061772&pc=NEZ | 3008062894-2025-00011 | 2025-04-28 04:00 | Death | CAPSOVISION, INC | 2025-05-21 04:00 | NEZ | CAPSOCAM SV-3 | Adverse Event Without Identified Device or Use Problem | Foreign Body In Patient | K242643 |  | 1 | Event Description: ON (B)(6) 2025 - WE WERE NOTIFIED OF A PATIENT WHO PASSED AWAY. FRM-0089C CAPSULE INCIDENT QUESTIONNAIRE WAS SENT TO THE CUSTOMER TO GATHER ADDITIONAL INFORMATION. FRM-0089C WAS RECEIVED INDICATING PATIENT HAD PRE-EXISTING CONDITIONS AND THAT THE PATIENT CAME IN A "WHEELCHAIR DUE TO BILATERAL AMPUTEE" AND THAT A "NURSE FROM THE PATIENT'S LONG-TERM FACILITY WHO VERIFIED BOWEL PREP. PATIENT THEN SWALLOWED THE VIDEO CAPSULE EASILY WITH A SMALL AMOUNT OF WATER. PATIENT THEN LEFT VIA MEDICAL TRANSPORT IN SAME CONDITION AS ARRIVAL". THE DATE OF DEATH WAS NOTED AS (B)(6) 2025. AFTER REQUESTING THE REASON FOR THE CAUSE OF THE DEATH WE WERE INFORMED THAT THE CAUSE WAS UNKNOWN. THE PATIENT HAD MULTIPLE COMORBIDITIES, SHE LIVED A LONG-TERM CARE FACILITY AND THEY DID NOT KNOW THE EXACT CAUSE OF DEATH." ON (B)(6) 2025 - WE WERE INFORMED THAT THE CAPSULE PROCEDURE WAS ADMINISTERED DUE TO THE PATIENT BEING ANEMIC. TO DATE, WE ARE STILL WAITING FOR ADDITIONAL INFORMATION TO IDENTIFY THE CAUSE OF THE DEATH. Manufacturer Narrative: ON (B)(6) 2025 - WE WERE NOTIFIED OF A PATIENT WHO PASSED AWAY. FRM-0089C CAPSULE INCIDENT QUESTIONNAIRE WAS SENT TO THE CUSTOMER TO GATHER ADDITIONAL INFORMATION. FRM-0089C WAS RECEIVED INDICATING PATIENT HAD PRE-EXISTING CONDITIONS AND THAT THE PATIENT CAME IN A "WHEELCHAIR DUE TO BILATERAL AMPUTEE" AND THAT A "NURSE FROM THE PATIENT'S LONG-TERM FACILITY WHO VERIFIED BOWEL PREP. PATIENT THEN SWALLOWED THE VIDEO CAPSULE EASILY WITH A SMALL AMOUNT OF WATER. PATIENT THEN LEFT VIA MEDICAL TRANSPORT IN SAME CONDITION AS ARRIVAL". THE DATE OF DEATH WAS NOTED AS (B)(6) 2025. AFTER REQUESTING THE REASON FOR THE CAUSE OF THE DEATH WE WERE INFORMED THAT THE CAUSE WAS UNKNOWN. THE PATIENT HAD MULTIPLE COMORBIDITIES, SHE LIVED A LONG-TERM CARE FACILITY AND THEY DID NOT KNOW THE EXACT CAUSE OF DEATH." ON (B)(6) 2025 - WE WERE INFORMED THAT THE CAPSULE PROCEDURE WAS ADMINISTERED DUE TO THE PATIENT BEING ANEMIC. TO DATE, WE ARE STILL WAITING FOR ADDITIONAL INFORMATION TO IDENTIFY THE CAUSE OF THE DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22062326&pc=DSQ | 2916596-2025-03062 | 2025-05-01 04:00 | Death | THORATEC CORPORATION | 2025-05-21 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. WHETHER THE DEATH WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE HOSPITAL DID NOT COMMUNICATE DETAILS OR CAUSE OF DEATH. THIS WAS NOT CONSIDERED DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE PROVIDED INFORMATION INDICATED THE PATIENT EXPIRED. ALTHOUGH THE CAUSE OF DEATH WAS REPORTED AS UNKNOWN, THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22062366&pc=OZD | 1220648-2025-28612 | 2025-05-01 04:00 | Death | ABIOMED, INC. | 2025-05-21 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Atrial Fibrillation; Tachycardia; Thrombocytopenia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THERE WERE SUCTION EVENTS. THE PATIENT RECEIVED FOUR LITERS OF LACTATED RINGER¿S SOLUTION AND 500 CUBIC CENTIMETERS OF ALBUMIN. SUPPORT WAS CONTINUED. ADDITIONALLY, THE PATIENT HAD ATRIAL FIBRILLATION AND VENTRICULAR TACHYCARDIA (VT) DURING SUPPORT. THE PATIENT HAD TO BE CARDIOVERTED DUE TO THE VT. AMIODARONE AND LIDOCAINE WERE GIVEN TO THE PATIENT. FURTHERMORE, THERE WAS CONCERN FOR HEPARIN-INDUCED THROMBOCYTOPENIA. ANTICOAGULATION WAS SWITCHED TO BIVAL. THE PATIENT REMAINS ON SUPPORT.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF THROMBOCYTOPENIA, VF/VT/AF/AT, AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. DATA LOG REVIEWED: MANY SUCTION ALARMS OBSERVED DURING IMPELLA SUPPORT AROUND DAYS SUCTIONS ISSUE REPORTED. NO MC SPIKE OUTSIDE OF P-LEVEL CHANGES. NO POSITIONING OR PLACEMENT SIGNAL ISSUES SEEN. NO DROP IN FLOW OBSERVED. LOW PUMP FLOW: THE CAUSE OF THE LOW PUMP FLOW MOST LIKELY WAS PATIENT CONDITION RELATED WITH FLUID BOLUS ADMINISTRATION/REDUCTION IN P-LEVEL RESOLVING SUCTION ALARMS. VF/VT/AF/AT: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. THROMBOCYTOPENIA: THE CAUSE OF THE THROMBOCYTOPENIA ISSUE WAS NOT DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. B2 UPDATED TO INCLUDE DEATH AND DATE OF DEATH. B5 UPDATED WITH PATIENT OUTCOME. D6B EXPLANT DATE ADDED. H1 TYPE OF EVENT UPDATED TO DEATH. H6 HEALTH EFFECT - IMPACT CODE 1802 ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22062482&pc=MNS | 2518422-2025-105912 | 2024-10-02 04:00 | Death | RESPIRONICS, INC. | 2025-05-21 04:00 | MNS | BIPAP AUTOSV ADVANCED SYSTEM ONE 60 SERIES | Degraded | Insufficient Information | K090248 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE SPOUSE OF THE PATIENT STATED THE REPLACEMENT DEVICE WAS NO LONGER NEEDED. THE SPOUSE OF THE PATIENT INSTANTLY BECAME SAD AND DEPRESSED AS THOUGH THE PATIENT HAD PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22062659&pc=OZD | 1220648-2025-28615 | 2025-04-30 04:00 | Death | ABIOMED, INC. | 2025-05-21 04:00 | OZD | IMPELLA 5.5 | Patient-Device Incompatibility | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE PATIENT HAD HIGH-RISK CARDIAC SURGERY WHICH THE PATIENT NEVER RECOVERED HEMODYNAMICALLY ON VA ECMO FOR ONE WEEK, SO IMPELLA 5.5 WAS ATTEMPTED. THE AORTIC VALVE (AOV) WAS NOT OPEN AND DIFFICULT ANATOMY MADE CROSSING AOV VERY CHALLENGING BUT WAS SUCCESSFUL. IMPELLA PLACED OVER 035 SUPER STIFF AMPLATZ AND WAS UNSUCCESSFUL AT CROSSING AOV DESPITE SEVERAL MEDICAL MODALITIES TO FORCE EJECTION. DURING THE CASE THE PATIENT HAD MULTIPLE EPISODES OF VENTRICULAR FIBRILLATION (VF), REQUIRING OVER 20 SHOCKS. THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ B2 DATE OF DEATH UNKNOWN. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. DELIVERY ISSUES: BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF DELIVERY ISSUE IS DETERMINED TO BE PATIENT CONDITION BECAUSE PUMP WAS UNABLE TO PASS THROUGH VASCULATURE DUE TO DIFFICULT ANATOMY OF THE PATIENT. VT/VF: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ IMPELLA 5.5 WITH SMARTASSIST WITH AUTOMATED IMPELLA CONTROLLER: SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION: SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER: SECTION: DIRECT AORTIC INSERTION: SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK WITH AUTOMATED IMPELLA CONTROLLER: SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING IMPELLA WITH AUTOMATED IMPELLA CONTROLLER: SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION: SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER: SECTION: DIRECT AORTIC INSERTION: SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ H10 INCOMPLETE INSTRUCTIONS FOR USE WERE PROVIDED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28615.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22062779&pc=DSQ | 2916596-2025-03362 | 2024-12-23 05:00 | Death | THORATEC CORPORATION | 2025-05-21 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO UNKNOWN ETIOLOGY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WILL NOT BE EXPLANTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22062790&pc=OWB | 3006544299-2025-00418 | 2025-05-12 04:00 | Death | MEDTRONIC NAVIGATION, INC. (BOXBOROUGH) | 2025-05-21 04:00 | OWB | O-ARM O2 IMAGING SYSTEM | Poor Quality Image; Unexpected/Unintended Radiation Output | Cardiac Arrest | K200074 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION REGARDING AN IMAGING SYSTEM BEING USED FOR SACROILIAC AND THORACOLUMBAR PROCEDURE. IT WAS REPORTED THAT DURING A SPINAL FUSION CASE, THE SPINE SMART DOSE IMAGES APPEARED OVEREXPOSED. THE ANTERIOR PORTION OF THE VERTEBRAL BODY APPEARED WHITE AND CLINICAL CONSULTANT (CC) HAD TO ADJUST THE LEVEL AND WIDTH TO PROCEED WITH THE CASE. THEY TOOK A POST-OP HD SCAN WHICH HAD BETTER IMAGE QUALITY. CC NOTED THAT THE PATIENT HAD DIMINISHED ANATOMY DUE TO CANCER, AND THE PATIENT WAS HEAVY. THIS ISSUE OCCURRED INTRAOPERATIVELY AND DID NOT CAUSE ANY SURGICAL DELAY. ADDITIONAL INFORMATION STATING THAT THE CC REPORTED THAT WHEN THE SURGEON REMOVED THE FRAME FROM THE PATIENT, THE PATIENT CODED AND PASSED AWAY. CC WAS UNSURE HOW OR WHY THE PATIENT CODED. CC WAS NOT PRESENT WHEN THE PATIENT CODED, BUT CC BELIEVED THAT SPINE REPRESENTATIVE WAS IN THE VICINITY WHEN IT HAPPENED. Manufacturer Narrative: H3, H6: THE MANUFACTURE REPRESENTATIVE WENT TO THE SITE TO TEST THE IMAGING SYSTEM AND NO FAULT FOUND. MDR CODES: B01, C19, D14. H3: THE SOFTWARE INVESTIGATION FOUND THAT THIS WAS A CONFIRMED SOFTWARE ISSUE. MDR CODES: B01, C10, D02. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H3,H6) NO PARTS HAVE BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID: M018081C001, VERSION #: 4.3.0. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22062861&pc=DSQ | 2916596-2025-03128 | 2025-04-30 04:00 | Death | THORATEC CORPORATION | 2025-05-21 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. THE DEATH WAS NOT RELATED TO THE DEVICE OR DEVICE THERAPY. THERE WAS NO MALFUNCTION OF THE PUMP REPORTED. THE CAUSE OF DEATH AND/OR DETAILS LEADING UP TO DEATH WAS NOT COMMUNICATED BY THE CENTER PER PATIENT DATA PRIVACY LAWS. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. IT WAS REPORTED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED DUE TO PATIENT DATA PRIVACY. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22062879&pc=DSQ | 2916596-2025-03121 | 2025-02-03 05:00 | Death | THORATEC CORPORATION | 2025-05-21 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE HOSPITAL STATED THAT THE CAUSE OF DEATH WAS UNKNOWN. THE DEATH WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED AND THE DEVICE WAS NOT EXPLANTED. THE DEVICE WOULD NOT RETURN. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22063047&pc=MIH | 3011063223-2025-00068 |  | Death | ENDOLOGIX | 2025-05-21 04:00 | MIH | AFX | Material Puncture/Hole |  | P040002 |  | 1 | Event Description: THE REPORTED INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH REVIEW OF AN INDUSTRY RELEVANT JOURNAL ARTICLE. THERE ARE TWO JOURNAL ARTICLES THAT ARE THE SAME SUBSET OF PATIENTS. DOI.ORG/10.1177/15266028241255541. DOI.ORG/10.1177/15266028251317910. IT WAS REPORTED THAT 872 PATIENTS WERE INITIALLY IMPLANTED IN THE ITALY WITH THE AFX BIFURCATED STENT GRAFT SYSTEM (DEVICE ITERATION; AFX STRATA, DURAPLY, AFX2 ARE UNKNOWN) BETWEEN 1994 AND 2020. IT WAS REPORTED WITHIN THE JOURNAL ARTICLE THREE (3) OF THESE PATIENTS EXPERIENCED TYPE III ENDOLEAK (UNKNOWN SUB-TYPE) AND DEATH. THE EXACT CASE DATE, EVENT DETAILS, LOT NUMBERS, AND CATALOG NUMBERS ARE UNKNOWN. NO ADDITIONAL INFORMATION WILL BECOME AVAILABLE REGARDING THIS INITIAL/FINAL REPORT DUE TO THE ABSENCE OF PATIENT IDENTIFIERS. NOTE: AS THE EXACT DATE OF INCIDENT IS UNKNOWN, THE BEST ESTIMATED DATE WAS SELECTED. THIS REPORT ADDRESSES PATIENT 1 OF 3. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. THE REPORTED ADVERSE EVENT/INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH AN ONGOING REVIEW OF INDUSTRY RELEVANT JOURNAL ARTICLES WHERE PATIENT RELATED OUTCOMES ARE PRESENTED ANONYMOUSLY, AND PATIENT SPECIFIC DEVICE INFORMATION IS UNAVAILABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE COMPLETED. THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION REMAINS UNKNOWN. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT COULD NOT BE COMPLETED AS THE REPORTED EVENT WAS FOUND DURING PUBLICATION REVIEW WHERE THE PATIENT INFORMATION IS ANONYMOUS. DEVICE, USE, PROCEDURE, AND/OR ANATOMY RELATEDNESS TO THIS ADVERSE EVENT/INCIDENT COULD NOT BE EVALUATED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22063054&pc=MIH | 3011063223-2025-00066 |  | Death | ENDOLOGIX | 2025-05-21 04:00 | MIH | AFX2 | Material Puncture/Hole |  | P040002 |  | 1 | Event Description: THE REPORTED INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH REVIEW OF AN INDUSTRY RELEVANT JOURNAL ARTICLE. THERE ARE TWO JOURNAL ARTICLES THAT ARE THE SAME SUBSET OF PATIENTS. DOI.ORG/10.1177/15266028241255541. DOI.ORG/10.1177/15266028251317910. IT WAS REPORTED THAT 872 PATIENTS WERE INITIALLY IMPLANTED IN THE ITALY WITH THE AFX BIFURCATED STENT GRAFT SYSTEM (DEVICE ITERATION; AFX STRATA, DURAPLY, AFX2 ARE UNKNOWN) BETWEEN 1994 AND 2020. IT WAS REPORTED WITHIN THE JOURNAL ARTICLE THREE (3) OF THESE PATIENTS EXPERIENCED TYPE III ENDOLEAK (UNKNOWN SUB- TYPE) AND DEATH. THE EXACT CASE DATE, EVENT DETAILS, LOT NUMBERS, AND CATALOG NUMBERS ARE UNKNOWN. NO ADDITIONAL INFORMATION WILL BECOME AVAILABLE REGARDING THIS INITIAL/FINAL REPORT DUE TO THE ABSENCE OF PATIENT IDENTIFIERS. NOTE: AS THE EXACT DATE OF INCIDENT IS UNKNOWN, THE BEST ESTIMATED DATE WAS SELECTED. THIS REPORT ADDRESSES PATIENT 3 OF 3. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. THE REPORTED ADVERSE EVENT/INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH AN ONGOING REVIEW OF INDUSTRY RELEVANT JOURNAL ARTICLES WHERE PATIENT RELATED OUTCOMES ARE PRESENTED ANONYMOUSLY, AND PATIENT SPECIFIC DEVICE INFORMATION IS UNAVAILABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE COMPLETED. THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION REMAINS UNKNOWN. THE USER FACILITY WAS UNABLE TO PROVIDE THIS INFORMATION. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT COULD NOT BE COMPLETED AS THE REPORTED EVENT WAS FOUND DURING PUBLICATION REVIEW WHERE THE PATIENT INFORMATION IS ANONYMOUS. DEVICE, USE, PROCEDURE, AND/OR ANATOMY RELATEDNESS TO THIS ADVERSE EVENT/INCIDENT COULD NOT BE EVALUATED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22063059&pc=MIH | 3011063223-2025-00074 |  | Death | ENDOLOGIX | 2025-05-21 04:00 | MIH | AFX | Material Puncture/Hole |  | P040002 |  | 1 | Event Description: THE REPORTED INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH REVIEW OF AN INDUSTRY RELEVANT JOURNAL ARTICLE. THERE ARE TWO JOURNAL ARTICLES THAT ARE THE SAME SUBSET OF PATIENTS. DOI.ORG/10.1177/15266028241255541. DOI.ORG/10.1177/15266028251317910. IT WAS REPORTED THAT 872 PATIENTS WERE INITIALLY IMPLANTED IN THE ITALY WITH THE AFX BIFURCATED STENT GRAFT SYSTEM (DEVICE ITERATION; AFX STRATA, DURAPLY, AFX2 ARE UNKNOWN) BETWEEN 1994 AND 2020. IT WAS REPORTED WITHIN THE JOURNAL ARTICLE THREE (3) OF THESE PATIENTS EXPERIENCED TYPE III ENDOLEAK (UNKNOWN SUB-TYPE) AND DEATH. THE EXACT CASE DATE, EVENT DETAILS, LOT NUMBERS, AND CATALOG NUMBERS ARE UNKNOWN. NO ADDITIONAL INFORMATION WILL BECOME AVAILABLE REGARDING THIS INITIAL/FINAL REPORT DUE TO THE ABSENCE OF PATIENT IDENTIFIERS. NOTE: AS THE EXACT DATE OF INCIDENT IS UNKNOWN, THE BEST ESTIMATED DATE WAS SELECTED. THIS REPORT ADDRESSES PATIENT 2 OF 3. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. THE REPORTED ADVERSE EVENT/INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH AN ONGOING REVIEW OF INDUSTRY RELEVANT JOURNAL ARTICLES WHERE PATIENT RELATED OUTCOMES ARE PRESENTED ANONYMOUSLY, AND PATIENT SPECIFIC DEVICE INFORMATION IS UNAVAILABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE COMPLETED. THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION REMAINS UNKNOWN. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT COULD NOT BE COMPLETED AS THE REPORTED EVENT WAS FOUND DURING PUBLICATION REVIEW WHERE THE PATIENT INFORMATION IS ANONYMOUS. DEVICE, USE, PROCEDURE, AND/OR ANATOMY RELATEDNESS TO THIS ADVERSE EVENT/INCIDENT COULD NOT BE EVALUATED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22063060&pc=MIH | 3011063223-2025-00081 | 2025-05-01 04:00 | Death | ENDOLOGIX | 2025-05-21 04:00 | MIH | AFX2 | Off-Label Use; Material Puncture/Hole; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Ruptured Aneurysm; Embolism/Embolus | P040002 |  | 1 | Event Description: THE PATIENT WAS INITIALLY TREATED WITH IMPLANT OF AN AFX2 BIFURCATED STENT GRAFT, AN AFX VELA SUPRARENAL, AN AFX LIMB STENT GRAFT AND AN OVATION IX EXTENDER. APPROXIMATELY TWO (2) YEARS POST INITIAL PROCEDURE, THE PATIENT EMERGENTLY PRESENTED WITH ABDOMINAL PAIN. THE PHYSICIAN IDENTIFIED A POTENTIAL ENDOLEAK (INDETERMINATE ORIGIN) DURING A COMPUTED TOMOGRAPHY (CT) SCAN. THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL FOR OPEN SURGERY TO REPAIR ENDOLEAK. THE PATIENT WAS HEMODYNAMICALLY STABLE AND EXPIRED DUE TO UNKNOWN REASON. NOTE: THE AFX2 SYSTEM WAS EVALUATED USING ENDOLOGIX COMPONENTS. THE SAFETY AND EFFECTIVENESS OF OTHER STENT OR STENT GRAFT DEVICES USED IN CONJUNCTION WITH THE AFX2 SYSTEM HAVE NOT BEEN ESTABLISHED. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. AN EVALUATION OF THE MANUFACTURING RECORD WAS COMPLETED. A REVIEW OF THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION SHOWS THE DEVICE TO HAVE BEEN PROPERLY MANUFACTURED AND RELEASED IN ACCORDANCE WITH THE DEVICE MASTER RECORD. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT WAS COMPLETED. AN EXAMINATION OF MEDICAL RECORDS AND/OR MEDICAL IMAGING RECEIVED BY ENDOLOGIX SHOWS THAT THE INDETERMINATE ENDOLEAK WAS DETERMINED TO BE A TYPE IA ENDOLEAK AND A TYPE IIIB ENDOLEAK. THE PREEXISTING RUPTURED AORTA, OPEN SURGICAL INTERVENTION AND DEATH COMPLAINTS ARE CONFIRMED. THIS IS MODERATELY CONSISTENT WITH THE REPORTED ADVERSE EVENT/INCIDENT. THE INITIAL PROCEDURE IS OFF LABEL DUE TO CONCOMITANT PRODUCT USAGE (OVATION LIMB). THE SAFETY AND EFFECTIVENESS OF OTHER STENT OR STENT GRAFT DEVICES USED IN CONJUNCTION WITH THE AFX2 SYSTEM HAVE NOT BEEN ESTABLISHED. IT IS UNCLEAR IF THIS CONTRIBUTED TO THE REPORTED EVENT. THE SUPERIOR STENT MARGIN OF THE MAIN BODY WAS POSITIONED AT A 25.1° ANGULATION, RESULTING IN POOR APPOSITION BETWEEN THE MAIN BODY AND THE PROXIMAL CUFF. THIS COULD HAVE CONTRIBUTED TO THE REPORTED TYPE IA ENDOLEAK. THE PATIENT HAD PREEXISTING RENAL CELL CARCINOMA WITH LUNG METASTASES. BOTH THE UNDERLYING MALIGNANCY AND TREATMENT WITH IMMUNOTHERAPY COULD CONTRIBUTED TO AN INCREASED RISK OF DISSEMINATED INTRAVASCULAR COAGULATION (DIC). THE PREEXISTING RUPTURE LIKELY CONTRIBUTED TO A HYPERCOAGULABLE STATE THAT LED TO DIC. THE THROMBUS IN THE RIGHT ATRIUM WAS LIKELY A RESULT OF THE HYPERCOAGULABLE STATE ASSOCIATED WITH DISSEMINATED INTRAVASCULAR COAGULATION (DIC). THE THROMBOSIS IN THE HEART, COMBINED WITH HYPOXIA, LIKELY CONTRIBUTED TO THE DEVELOPMENT OF PULSELESS ELECTRICAL ACTIVITY (PEA), WHICH ULTIMATELY RESULTED IN THE PATIENT¿S DEATH MAKING THIS PROCEDURE RELATED. THE CLINAL ASSESSMENT DETERMINED THAT THERE WAS ANEURYSM ENLARGEMENT OF 7.5MM THAT WAS NOT INCLUDED IN THE EVENT AS REPORTED. THE ANEURYSM ENLARGEMENT WAS DISCOVERED DURING REVIEW OF THE CT SCAN DATED ON (B)(6) 2025. THE DEATH IS MOST LIKELY ANATOMY RELATED. DEVICE, USER, PROCEDURE OR ANATOMY RELATEDNESS OF THE TYPE IA AND IIIB ENDOLEAK COMPLAINTS COULD NOT BE DETERMINED. THE COMPLAINT IS MOST LIKELY ANATOMY RELATED. THE PROCEDURE RELATED HARMS ARE ABNORMAL BLOOD LOSS (FROM DIC COAGULOPATHY), EMBOLISM (BLOOD CLOT RIGHT ATRIUM), HEMODYNAMIC INSTABILITY, CARDIAC ARREST AND DEATH. THE FINAL PATIENT STATUS WAS REPORTED AS DECEASED DURING PROCEDURE. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS. CORRECTIONS: G3: DATE RECEIVED BY MANUFACTURER HAS BEEN UPDATED. H6: RESULT CODE: REMOVE CODE 3233. H6: CONCLUSION CODE: REMOVE CODE 11. H6: MEDICAL DEVICE PROBLEM CODES: REMOVE CODE 2993. Manufacturer Narrative: THE DEVICES INVOLVED IN THIS EVENT WILL NOT BE RETURNED FOR EVALUATION. PATIENT MEDICAL RECORDS AND IMAGING STUDIES WILL BE REQUESTED FOR FURTHER EVALUATION BY A CLINICAL SPECIALIST. IF ADDITIONAL INFORMATION PERTINENT TO THE INCIDENT IS OBTAINED, A FOLLOW-UP REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22063080&pc=DXY | 2182208-2025-02483 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-05-21 04:00 | DXY | MEDTRONIC LEAD | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Unspecified Infection |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING SUBCUTANEOUS IMPLANTABLE CARDIOVERTER-DEFIBRILLATORS (S-ICDS) COMPARED WITH TRANSVENOUS IMPLANTABLE CARDIOVERTER-DEFIBRILLATORS (TV-ICDS). THE AUTHORS DESCRIBED FOUR PATIENT DEATHS IN THE TV-ICD GROUP; THE CAUSES OF DEATH WERE HEART FAILURE, CEREBRAL HEMORRHAGE, OR INFECTION. THERE WERE PATIENTS WHO EXPERIENCED INAPPROPRIATE THERAPY DUE TO ATRIAL TACHYARRHYTHMIAS. THERE WERE PROCEDURE-RELATED COMPLICATIONS WHICH INCLUDED HEMATOMA, INFECTION, OR WOUND DEHISCENCE REQUIRING SURGICAL INTERVENTION AND LEAD-RELATED COMPLICATIONS WHICH INCLUDED LEAD FRACTURE OR LEAD DISLODGEMENT. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: LONG-TERM SAFETY AND EFFICACY OF SUBCUTANEOUS IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR COMPARED WITH TRANSVENOUS I MPLANTABLE CARDIOVERTER-DEFIBRILLATOR IN PROPENSITY SCORE-MATCHED PATIENTS FROM JAPAN. JOURNAL OF ARRHYTHMIA. 2025; 41:E70063. DOI: 10.1002/JOA3.70063 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22063087&pc=DXY | 2182208-2025-02482 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-05-21 04:00 | DXY | MEDTRONIC ICD | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Unspecified Infection |  |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING SUBCUTANEOUS IMPLANTABLE CARDIOVERTER-DEFIBRILLATORS (S-ICDS) COMPARED WITH TRANSVENOUS IMPLANTABLE CARDIOVERTER-DEFIBRILLATORS (TV-ICDS). THE AUTHORS DESCRIBED FOUR PATIENT DEATHS IN THE TV-ICD GROUP; THE CAUSES OF DEATH WERE HEART FAILURE, CEREBRAL HEMORRHAGE, OR INFECTION. THERE WERE PATIENTS WHO EXPERIENCED INAPPROPRIATE THERAPY DUE TO ATRIAL TACHYARRHYTHMIAS. THERE WERE PROCEDURE-RELATED COMPLICATIONS WHICH INCLUDED HEMATOMA, INFECTION, OR WOUND DEHISCENCE REQUIRING SURGICAL INTERVENTION AND LEAD-RELATED COMPLICATIONS WHICH INCLUDED LEAD FRACTURE OR LEAD DISLODGEMENT. THE STATUS OF THE DEVICES AND LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: LONG-TERM SAFETY AND EFFICACY OF SUBCUTANEOUS IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR COMPARED WITH TRANSVENOUS IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR IN PROPENSITY SCORE-MATCHED PATIENTS FROM JAPAN. JOURNAL OF ARRHYTHMIA. 2025; 41: E70063. DOI: 10.1002/JOA3.70063. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22063127&pc=DSQ | 2916596-2025-03363 | 2025-05-09 04:00 | Death | THORATEC CORPORATION | 2025-05-21 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. AN AUTOPSY WAS NOT PERFORMED AND THE DEVICE WAS NOT EXPLANTED. THIS WAS NOT DEVICE OR THERAPY RELATED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT'S DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22063298&pc=DSQ | 2916596-2025-03238 | 2025-04-22 04:00 | Death | THORATEC CORPORATION | 2025-05-21 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT EXPIRED. NO ADDITIONAL INFORMATION WAS PROVIDED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU, REV. D IS CURRENTLY AVAILABLE. CHAPTER 1, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22063344&pc=NPT | 2025587-2025-03816 | 2025-05-14 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-21 04:00 | NPT | EVOLUT FX PLUS VALVE | Perivalvular Leak; Malposition of Device; Insufficient Information; Patient Device Interaction Problem | Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT A TRANSCATHETER AORTIC VALVE WAS PLACED IN A BICUSPID VALVE WITH MINIMAL CALCIFICATION IN THE NON-CORONARY CUSP (NCC) AND MIGRATED INTO THE ASCENDING AORTA. IT WAS SUGGESTED THAT THE MIGRATION OCCURRED BECAUSE THE VALVE WAS PLACED TOO HIGH (1 MM ON THE NCC). A SECOND VALVE WAS POSITIONED LOWER AT 3 MM WITH AN ACCEPTABLE OUTCOME. TRACE-MILD PARAVALVULAR LEAK (PVL) WAS OBSERVED AND A LARGE RAPHE. THE PATIENT DIED DUE TO MULTIPLE AORTIC DISSECTIONS, WITH A POSSIBLE CAUSE BEING THE EMBOLIZED VALVE SNARED UP INTO THE ASCENDING AORTA.Event Description: IT WAS REPORTED THAT A TRANSCATHETER AORTIC VALVE WAS PLACED IN A BICUSPID VALVE WITH MINIMAL CALCIFICATION IN THE NON-CORONARY CUSP (NCC) AND MIGRATED INTO THE ASCENDING AORTA. IT WAS SUGGESTED THAT THE MIGRATION OCCURRED BECAUSE THE VALVE WAS PLACED TOO HIGH (1 MM ON THE NCC). A SECOND VALVE WAS POSITIONED LOWER AT 3 MM WITH AN ACCEPTABLE OUTCOME. TRACE-MILD PARAVALVULAR LEAK (PVL) WAS OBSERVED AND A LARGE RAPHE. THE PATIENT DIED DUE TO MULTIPLE AORTIC DISSECTIONS, WITH A POSSIBLE CAUSE BEING THE EMBOLIZED VALVE SNARED UP INTO THE ASCENDING AORTA. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT FOLLOWING THE VALVE IMPLANT PROCEDURE, THE PATIENT WAS TRANSFERRED TO INTERMEDIATE CARE WHICH IS WHEN THE DISSECTION WAS FIRST NOTED. ADDITIONALLY, IT WAS NOTED THE PATIENT HAD MULTIPLE STROKES, AND; THEREFORE, NO FURTHER INTERVENTION WAS MADE.Event Description: IT WAS REPORTED THAT A TRANSCATHETER AORTIC VALVE WAS PLACED IN A BICUSPID VALVE WITH MINIMAL CALCIFICATION IN THE NON-CORONARY CUSP (NCC) AND MIGRATED INTO THE ASCENDING AORTA. IT WAS SUGGESTED THAT THE MIGRATION OCCURRED BECAUSE THE VALVE WAS PLACED TOO HIGH (1 MM ON THE NCC). A SECOND VALVE WAS POSITIONED LOWER AT 3 MM WITH AN ACCEPTABLE OUTCOME. TRACE-MILD PARAVALVULAR LEAK (PVL) WAS OBSERVED AND A LARGE RAPHE. THE PATIENT DIED DUE TO MULTIPLE AORTIC DISSECTIONS, WITH A POSSIBLE CAUSE BEING THE EMBOLIZED VALVE SNARED UP INTO THE ASCENDING AORTA. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT FOLLOWING THE VALVE IMPLANT PROCEDURE, THE PATIENT WAS TRANSFERRED TO INTERMEDIATE CARE WHICH IS WHEN THE DISSECTION WAS FIRST NOTED. ADDITIONALLY, IT WAS NOTED THE PATIENT HAD MULTIPLE STROKES, AND; THEREFORE, NO FURTHER INTERVENTION WAS MADE. ADDITIONAL INFORMATION WAS RECEIVED THAT THE LOCATION OF THE DISSECTIONS WERE LIMITED TO THE ASCENDING AORTA. NO PRE-IMPLANT OR POST-IMPLANT BALLOON DILATIONS WERE PERFORMED. THE FIRST VALVE WAS PULLED BACK TO THE ASCENDING AORTA BEFORE IMPLANTING THE SECOND VALVE. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: EVOLUT FX DCS, PRODUCT ID: D-EVOLUTFX-2329, LOT: 0012676208, USE BY DATE: 2027-02-18, UDI: (B)(4). SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012617235); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID: D-EVOLUTFX-2329 (LOT: 0012665445); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID: EVFXPLUS-29 (K087424); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: ON (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED: H6 - ADDED CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2. DATE OF DEATH B5. A SECOND PARAGRAPH WAS ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22063548&pc=OZD | 1220648-2025-28622 | 2025-04-30 04:00 | Death | ABIOMED, INC. | 2025-05-21 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD A IMPELLA CP PUMP SUPPORT ONGOING FOR A PATIENT ADMITTED IN WITH KNOWN DISEASE AND UNDERGOING A HIGH-RISK PERCUTANEOUS CORONARY INTERVENTION. THE PUMP WAS SUPPORTING WHEN THERE WERE SIGNS OF HEMOLYSIS, URINE BURGUNDY AND DECREASED TO 10/HR. THE CARE WAS WITHDRAWN AND PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS: INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES: MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22063884&pc=OZD | 1220648-2025-28623 | 2025-04-30 04:00 | Death | ABIOMED, INC. | 2025-05-21 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD A IMPELLA CP PUMP PLACED TO SUPPORT A PATIENT ADMITTED IN WITH MULTIPLE CARDIAC AND SYSTEMIC COMORBIDITIES WITH NEW CARDIOGENIC SHOCK. THE REGISTERED NURSE REPORTED THAT THE PATIENT HAD RUN OF A FLUTTER/WIDE COMPLEX VENTRICULAR TACHYCARDIA REQUIRING CARDIOVERSION. THE PATIENTS PRESSURE DROPPED FOLLOWING AND THE NURSE DOUBLED VASOPRESSIN AND LEVOPHED DRIPS, ADDING AMIODARONE AND DC¿D DOBUTAMINE. AN ECHOCARDIOGRAM REPEATED BUT NOT RESULTED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22064111&pc=DSQ | 2916596-2025-02982 | 2025-04-18 04:00 | Death | THORATEC CORPORATION | 2025-05-21 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Headache; Hemorrhagic Stroke; Thrombosis/Thrombus | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A COMPUTED TOMOGRAPHY ANGIOGRAM (CTA) OF THE CHEST, ABDOMEN, AND PELVIS, WHICH REFLECTED CLOT MATERIAL IN THE OUTFLOW GRAFT. NO ALARMS WERE NOTED AND PATIENT¿S PARAMETERS WERE WITHIN THEIR USUAL RANGE. THERE WERE NO UNUSUAL EVENTS RECORDED IN THE LOG FILE EVENT HISTORY AND THE MECHANICAL CIRCULATORY SUPPORT (MCS) EQUIPMENT OPERATED AS INTENDED. ON (B)(6) 2025, THE PATIENT WENT TO AN INPATIENT REHAB FACILITY AND WAS TRANSFERRED AND ADMITTED TO THE HOSPITAL ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT WAS PLACED ON A HEPARIN DRIP UPON RECEIPT OF THE CTA RESULTS. ON (B)(6) 2025, THE PATIENT HAD A SEVERE HEADACHE. ON (B)(6) 2025, THE PATIENT HAD STROKE-LIKE SYMPTOMS, SO THEY WERE TAKEN FOR A CT SCAN, WHICH REVEALED A LARGE HEMORRHAGIC STROKE. THE PATIENT'S FAMILY ELECTED TO PURSUE HOSPICE MEASURES AND THE PATIENT PASSED AWAY ON (B)(6) 2025. THERE WERE NO SYMPTOMS OF THROMBUS OBSERVED AND IT WAS UNDETERMINED WHETHER THE THROMBUS WAS RESOLVED. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: THE REPORT OF SUSPECTED OUTFLOW GRAFT THROMBUS COULD NOT BE CONFIRMED AS NO PRODUCT WAS RETURNED FOR EVALUATION AND REVIEW OF THE SUBMITTED LOG FILES CAPTURED THE PUMP FUNCTIONING AS INTENDED AT THE FIXED SPEED. NO UNUSUAL EVENTS OR ALARMS WERE NOTED. FURTHER, A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED. THE PATIENT'S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED, AND THE DEVICE REPORTEDLY FUNCTIONED AS INTENDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING STROKE, PUMP THROMBOSIS, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 5 OF THE IFU, "SURGICAL PROCEDURES", INSTRUCTS TO INSPECT THE VENTRICULAR CHAMBER FOR MURAL THROMBI AND CROSSING TRABECULAE FOLLOWING REMOVAL OF THE CORE AND TO ADDRESS ONE OR BOTH, AS NEEDED. THIS SECTION ALSO INSTRUCTS TO INSPECT THE VENTRICLE AND REMOVE ANY PREVIOUSLY FORMED CLOTS THAT MAY CAUSE EMBOLISM OR ANY TRABECULAE THAT MAY IMPEDE FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22064390&pc=FKX | 0008030665-2025-01143 | 2025-04-24 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-05-21 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K222318 |  | 1 | Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED TO FRESENIUS CUSTOMER SUPPORT THAT THE PATIENT EXPIRED ON THE MACHINE AND CALLED TO SCHEDULE A PICKUP OF THE CYCLER AND MODEM. THE PATIENT WENT TO BED THAT NIGHT AND NEVER WOKE UP. PER THE PDRN, THE DOCTOR BELIEVES IT IS RELATED TO A CARDIAC ISSUE. FOLLOW-UP WITH THE PATIENTS¿ PDRN CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING HOURS OF (B)(6) 2025. THE PATIENT WAS DISCOVERED BY A FAMILY MEMBER WHEN THEY ENTERED THE BEDROOM TO ASSIST THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA INDICATED THE LIBERTY SELECT CYCLER STALLED ON DRAIN CYCLE 2, HOWEVER THERE WERE NO TREATMENT RELATED CONCERNS. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING AND TRANSPORTED THE PATIENT DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO HER EXTENSIVE CARDIAC HISTORY. PER THE PDRN, THE EVENTS WERE UNRELATED TO ANY FRESENIUS PRODUCT(S) AND/OR DEVICE(S) DEFICIENCY OR MALFUNCTION.Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED TO FRESENIUS CUSTOMER SUPPORT THAT THE PATIENT EXPIRED ON THE MACHINE AND CALLED TO SCHEDULE A PICKUP OF THE CYCLER AND MODEM. THE PATIENT WENT TO BED THAT NIGHT AND NEVER WOKE UP. PER THE PDRN, THE DOCTOR BELIEVES IT IS RELATED TO A CARDIAC ISSUE. FOLLOW-UP WITH THE PATIENTS¿ PDRN CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING HOURS OF (B)(6) 2025. THE PATIENT WAS DISCOVERED BY A FAMILY MEMBER WHEN THEY ENTERED THE BEDROOM TO ASSIST THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA INDICATED THE LIBERTY SELECT CYCLER STALLED ON DRAIN CYCLE 2, HOWEVER THERE WERE NO TREATMENT RELATED CONCERNS. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING AND TRANSPORTED THE PATIENT DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO HER EXTENSIVE CARDIAC HISTORY. PER THE PDRN, THE EVENTS WERE UNRELATED TO ANY FRESENIUS PRODUCT(S) AND/OR DEVICE(S) DEFICIENCY OR MALFUNCTION. Manufacturer Narrative: CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND SUBSEQUENT DEATH, AS THE PATIENT WAS UNDERGOING HD THERAPY WHEN HE EXPIRED. THE PDRN REPORTED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO THE PATIENT¿S EXTENSIVE CARDIAC HISTORY. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, G4, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED AND SHOWED SIGNS OF PHYSICAL DAMAGE: HEATER TRAY DAMAGED; DENT. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. A (AS-RECEIVED WITH REDUCED DWELL) SIMULATED THERAPY WAS INITIATED AND COMPLETED ON THE CYCLER WITHOUT COMPLICATION. THE CYCLER UNDERWENT SYSTEM AIR LEAK AND VALVE ACTUATION TESTING AND WAS FOUND TO MEET PRODUCT SPECIFICATIONS. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22064649&pc=OZD | 1220648-2025-28627 | 2025-05-01 04:00 | Death | ABIOMED, INC. | 2025-05-21 04:00 | OZD | IMPELLA CP | Positioning Failure | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT, THERE WERE NOT POSITIONING ISSUES DUE TO PATIENT MOVEMENT. THE PUMP WAS RE-POSITIONED. THE PATIENT CONTINUED ON SUPPORT AND THE PATIENTS CONDITION WAS NOTED TO BE STABLE, BUT POOR. ADDITIONAL INFORMATION NOTED CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE CAUSE OF THE POSITIONING ISSUES MOST LIKELY WAS DUE TO PATIENT MOVEMENT. B1 PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28627.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22064853&pc=NGV | 2124215-2025-32716 | 2025-03-28 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | NGV | WATCHMAN LAA CLOSURE DEVICE & DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: REPORTED VIA LITERATURE ARTICLE IT WAS REPORTED DEATH(S) OCCURRED. THE NATIONAL CARDIOVASCULAR DATA REGISTRY LEFT ATRIAL APPENDAGE OCCLUSION (LAAO) REGISTRY DATA WAS USED FOR INPATIENT MEDICARE CLAIMS DATA USING DETERMINISTIC MATCHING AND ESTABLISHED A COHORT OF PATIENTS >/=65 YEARS OF AGE WHO UNDERWENT LAAO BETWEEN JANUARY 2016 AND DECEMBER 2019. EVENT FREE SURVIVAL WAS CALCULATED USING THE KAPLAN-MEIER METHOD FOR DEATH AND THE CUMULATIVE INCIDENCE FUNCTION, ACCOUNTING FOR COMPETING RISK OF DEATH, FOR STROKE. RESULTS: A TOTAL OF 34,975 PATIENTS WITH AN AVERAGE AGE 77.5 YEARS (SD 6.5) AND 41.8% FEMALE WERE INCLUDED. THE AVERAGE CHA(2) DS(2)-VASC SCORE WAS 4.7 (SD 1.4). MEDIAN FOLLOW-UP WAS 2.8 YEARS (IQR 2.2-3.7 YEARS). THE INCIDENCE OF ANY STROKE AND ISCHEMIC STROKE WERE 1.80 (95% CI 1.66-94) AND 1.47 (95% CI 1.35-1.59) PER 100 PERSON-YEAR, RESPECTIVELY. THE FIVE-YEAR RATES OF ANY STROKE AND ISCHEMIC STROKE WERE 5.70% (95% CI 5.37-6.05) AND 5.03% (95% CI 4.71-5.36%), RESPECTIVELY. THE INCIDENCE OF DEATH WAS 10.37 (10.05-10.68) PER 100 PERSON-YEARS AND THE FIVE-YEAR MORTALITY RATE WAS 44.34% (95% CI 43.45-45.23%). CONCLUSION: LONG-TERM ISCHEMIC STROKE RATES IN A LARGE NATIONALLY REPRESENTATIVE COHORT OF MEDICARE PATIENTS FOLLOWING WATCHMAN LAAO WERE LOW DESPITE A HIGH EMBOLIC RISK. CITATION: HIGGINS, A, ZIMMERMAN, S, WANG, Y. ET AL. LONG-TERM OUTCOMES FOLLOWING LEFT ATRIAL APPENDAGE OCCLUSION IN MEDICARE BENEFICIARIES: OUTCOMES FROM THE NATIONAL CARDIOVASCULAR DATA REGISTRY. JACC. 2025 APR, 85 (12_SUPPLEMENT) 869. HTTPS://DOI.ORG/10.1016/S0735-1097(25)01353-1 Manufacturer Narrative: B3 EVENT DATE: ESTIMATED AS 03/28/2025 AS THE ONLINE PUBLISHED DATE FOR THE ARTICLE IS NOTED AS MARCH 28, 2025, CITATION: HIGGINS, A, ZIMMERMAN, S, WANG, Y. ET AL. LONG-TERM OUTCOMES FOLLOWING LEFT ATRIAL APPENDAGE OCCLUSION IN MEDICARE BENEFICIARIES: OUTCOMES FROM THE NATIONAL CARDIOVASCULAR DATA REGISTRY. JACC. 2025 APR, 85 (12_SUPPLEMENT) 869. HTTPS://DOI.ORG/10.1016/S0735-1097(25)01353-1.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065065&pc=OOB | 2124215-2025-32568 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIOGENIC SHOCK Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065082&pc=HGM | 9610816-2025-000426 | 2016-05-09 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-05-21 04:00 | HGM | AVALON FM30 FETAL MONITOR | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Hypoxia; Thrombosis/Thrombus | K140535 |  | 1 | Event Description: IT WAS REPORTED FOLLOWING AN INFANT DEATH, THE MOTHER REPORTED CONCERNS RELATED TO PREVIOUS RECALLS RELATED TO THE DEVICE. IT WAS INDICATED THE EVENT OCCURRED IN 2016. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. Manufacturer Narrative: THE MANUFACTURING DATE FOR THE REPORTED DEVICE IS PRIOR TO 24SEPT2016 SO NO UDI REQUIRED. THE REPORTER'S NAME AND EMAIL WERE PROVIDED; HOWEVER, DUE TO PRIVACY LAWS WERE NOT INCLUDED IN THE REPORT. E1 REPORTER PHONE#: (B)(6). PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065087&pc=BZD | 2518422-2025-105942 | 2025-05-01 04:00 | Death | RESPIRONICS, INC. | 2025-05-21 04:00 | BZD | REMSTAR PLUS CFLEX ASSEMBLY | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. A RELATIVE OF THE PATIENT REPORTED THE PATIENT PASSED AWAY. THE REPORTER INQUIRED IF THE DEVICE WAS A PART OF THE RECALL. IT WAS CONFIRMED THE DEVICE WAS INCLUDED IN THE RECALL. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THIS DEVICE WAS MANUFACTURED 01/28/2011 WHICH IS PRIOR TO UDI REQUIREMENT IMPLEMENTATION. THIS DEVICE DOES NOT HAVE A UDI AND NO UDI WILL BE LISTED IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065169&pc=OOB | 2124215-2025-32569 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065176&pc=LOX | 2124215-2025-32572 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED, DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, HEART FAILURE, CARDIAC ARREST, MYOCARDIAL INFARCTION, CARDIOGENIC SHOCK. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065180&pc=LOX | 2124215-2025-32573 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED, DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065187&pc=LOX | 2124215-2025-32578 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI)# AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065239&pc=OOB | 2124215-2025-32575 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH DURING PROCEDURE DECEASED, ACUTE MYOCARDIAL INFARCTION, HEART FAILURE, CARDIOGENIC SHOCK Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065241&pc=LOX | 2124215-2025-32579 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED, DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065254&pc=NWX | 2124215-2025-32583 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | NWX | WOLVERINE CORONARY CUTTING BALLOON | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED. DEATH AT DISCHARGE DECEASED, OTHER CARDIOVASCULAR REASON HEART FAILURE, MYOCARDIAL INFARCTION, CARDIOGENIC SHOCK. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065256&pc=OOB | 2124215-2025-32576 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH DURING PROCEDURE DECEASED, ACUTE MYOCARDIAL INFARCTION, HEART FAILURE, CARDIOGENIC SHOCK Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065262&pc=MGB | 2024168-2025-05305 | 2021-01-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-05-21 04:00 | MGB | PERCLOSE PROGLIDE¿ | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P960043 |  | 1 | Event Description: THE ARTICLE AIMED EVALUATE THE SAFETY AND EFFICACY OF THE POST-CLOSURE TECHNIQUE USING PROGLIDE DEVICE DURING VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA-ECMO) DECANNULATION. 170 PATIENT UNDERWENT VA-ECMO BETWEEN JANUARY 2021 AND JUNE 2023. ALL PATIENTS HAD THEIR FEMORAL ARTERY PUNCTURE SITES CLOSED USING THE POST-CLOSURE TECHNIQUE UPON DECANNULATION. IT WAS REPORTED THE PROGLIDE DEVICE EXPERIENCED SUTURE BREAKS AND FAILURE TO ACHIEVE HEMOSTASIS, RESULTING IN MEDICATION INTERVENTION. THE PROGLIDE MAY HAVE ALSO CAUSED OR CONTRIBUTED TO LIMB ARTERIAL THROMBOSIS, PSEUDO-ANEURYSM, HEMATOMA, MANUAL COMPRESSION, PROLONGED HOSPITALIZATION, MEDICAL INTERVENTION, AND DEATH. THE STUDY CONCLUDED THAT THE SUTURE-MEDIATED CLOSURE OF THE FEMORAL ARTERY DURING VA-ECMO DECANNULATION IS A PROMISING THERAPEUTIC STRATEGY. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED ¿EFFICACY AND SAFETY OF PERCUTANEOUS POST-CLOSURE TECHNIQUE DURING THE DECANNULATION FOR VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION.¿ Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. A REVIEW OF THE PRODUCTION RECORD AND COMPLAINT HISTORY OF THE REPORTED LOT COULD NOT BE CONDUCTED BECAUSE THE PART AND LOT NUMBER WERE NOT PROVIDED. BASED ON THE CASE INFORMATION, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT OF DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE PATIENT EFFECT OF DEATH IS LISTED IN THE ELECTRONIC PROGLIDE INSTRUCTIONS FOR USE (IFU), AS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE USE OF THE DEVICE. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. B2, B3: ESTIMATED DATE D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS REPORTED IN THE ARTICLE ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. ATTACHED ARTICLE, TITLED ¿EFFICACY AND SAFETY OF PERCUTANEOUS POST-CLOSURE TECHNIQUE DURING THE DECANNULATION FOR VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065276&pc=OOB | 2124215-2025-32584 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIOGENIC SHOCK Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065280&pc=LOX | 2124215-2025-32582 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065286&pc=FDT | 9610595-2025-09249 | 2025-03-10 04:00 | Death | AIZU OLYMPUS CO., LTD. | 2025-05-21 04:00 | FDT | EVIS DUODENOVIDEOSCOPE | Adverse Event Without Identified Device or Use Problem | Pancreatitis | K954451 |  | 1 | Event Description: OLYMPUS REVIEWED THE FOLLOWING LITERATURE TITLED "COMPARISON OF MILD AND MODERATE TO SEVERE DEGREE PANCREATITIS AFTER ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY.¿ PANCREATITIS IS THE MOST COMMON AND DEVASTATING ADVERSE EVENT OF ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY (ERCP). POST-ERCP PANCREATITIS (PEP) IS MOSTLY MILD, BUT SOME CAN PROGRESS TO MORE SEVERE CONDITIONS WITH LETHAL OUTCOMES. THE GOAL OF THIS RETROSPECTIVE STUDY WAS TO CLARIFY PREDICTORS OF SEVERITY BY COMPARING MILD AND MODERATE TO SEVERE PEP. DATA WAS COMPILED FROM A TOTAL OF 155 ELIGIBLE PATIENTS WHO WERE DIAGNOSED WITH PEP AND COMPARED THOSE WHO HAD MILD PEP VERSUS THOSE WHO HAD MODERATE TO SEVERE PEP. SPECIFICALLY, BASELINE PROFILES, INTRAPROCEDURAL DATA, AND POST-ERCP OUTCOMES WERE COMPARED BETWEEN THE TWO GROUPS. THE STUDY CONCLUDED THAT DIFFICULT CANNULATION WITH SELECTIVE BILIARY CANNULATION TIME GREATER THAN 10 MINUTES WAS A SIGNIFICANT RISK FACTOR FOR MODERATE TO SEVERE PEP. IT WAS NOTED THAT PROLONGED CANNULATION MAY CAUSE MORE MECHANICAL TRAUMA TO THE PAPILLA, RESULTING IN MORE SEVERE EDEMATOUS OBSTRUCTION OF THE PANCREATIC DUCT (PD) AND WORSENING OF SUBSEQUENT PEP. ADDITIONALLY, THE STUDY FOUND CONCURRENT POST-ERCP COMPLICATIONS, PARTICULARLY MICRO-PERFORATION IS A SIGNIFICANT PREDICTIVE FACTOR FOR MODERATE TO SEVERE PEP. THE MECHANISM OF MICRO-PERFORATION, RATHER THAN THE MICRO-PERFORATION ITSELF, CONTRIBUTES TO THE INCREASED SEVERITY OF PEP. IT WAS NOTED THAT CARELESS MANIPULATION OF GUIDEWIRES OR ERCP INSTRUMENTS, SUCH AS FORCEPS, BASKETS, MECHANICAL LITHOTRIPTERS, OR STENTS, CAN LEAD TO PANCREATIC INJURY AND POTENTIALLY WORSEN THE SEVERITY OF PEP, EVEN IF SELECTIVE BILIARY CANNULATION IS SUCCESSFULLY AND SMOOTHLY ACHIEVED. IT WAS REPORTED THERE WAS ONE PEP RELATED DEATH WITHIN THE MODERATE TO SEVERE PEP GROUP. NO FURTHER DETAILS REGARDING THE PATIENT DEATH WERE PROVIDED WITHIN THE LITERATURE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION IDENTIFIED IN THE ARTICLE.Event Description: NO ADDITIONAL INFORMATION RECEIVED FROM CUSTOMER. Manufacturer Narrative: THIS REPORT IS RELATED TO THE FOLLOWING LINKED PATIENT IDENTIFIERS: (B)(6). THE INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED OR IF ADDITIONAL INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: THIS REPORT IS BEING SUPPLEMENTED TO PROVIDE ADDITIONAL INFORMATION BASED ON THE APPROVED FINAL INVESTIGATION. UPDATED FIELDS: D8, H6, H11. BASED ON THE RESULTS OF THE INVESTIGATION, AND SINCE THE DEVICE WAS NOT RETURNED FOR EVALUATION, THE DEFINITIVE ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. OLYMPUS WILL CONTINUE TO MONITOR THE FIELD PERFORMANCE OF THIS DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065293&pc=FDT | 9610595-2025-09250 | 2025-03-10 04:00 | Death | AIZU OLYMPUS CO., LTD. | 2025-05-21 04:00 | FDT | EVIS LUCERA DUODENOVIDEOSCOPE | Adverse Event Without Identified Device or Use Problem | Pancreatitis | N/A |  | 1 | Event Description: OLYMPUS REVIEWED THE FOLLOWING LITERATURE TITLED "COMPARISON OF MILD AND MODERATE TO SEVERE DEGREE PANCREATITIS AFTER ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY.¿ PANCREATITIS IS THE MOST COMMON AND DEVASTATING ADVERSE EVENT OF ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY (ERCP). POST-ERCP PANCREATITIS (PEP) IS MOSTLY MILD, BUT SOME CAN PROGRESS TO MORE SEVERE CONDITIONS WITH LETHAL OUTCOMES. THE GOAL OF THIS RETROSPECTIVE STUDY WAS TO CLARIFY PREDICTORS OF SEVERITY BY COMPARING MILD AND MODERATE TO SEVERE PEP. DATA WAS COMPILED FROM A TOTAL OF 155 ELIGIBLE PATIENTS WHO WERE DIAGNOSED WITH PEP AND COMPARED THOSE WHO HAD MILD PEP VERSUS THOSE WHO HAD MODERATE TO SEVERE PEP. SPECIFICALLY, BASELINE PROFILES, INTRAPROCEDURAL DATA, AND POST-ERCP OUTCOMES WERE COMPARED BETWEEN THE TWO GROUPS. THE STUDY CONCLUDED THAT DIFFICULT CANNULATION WITH SELECTIVE BILIARY CANNULATION TIME GREATER THAN 10 MINUTES WAS A SIGNIFICANT RISK FACTOR FOR MODERATE TO SEVERE PEP. IT WAS NOTED THAT PROLONGED CANNULATION MAY CAUSE MORE MECHANICAL TRAUMA TO THE PAPILLA, RESULTING IN MORE SEVERE EDEMATOUS OBSTRUCTION OF THE PANCREATIC DUCT (PD) AND WORSENING OF SUBSEQUENT PEP. ADDITIONALLY, THE STUDY FOUND CONCURRENT POST-ERCP COMPLICATIONS, PARTICULARLY MICRO-PERFORATION IS A SIGNIFICANT PREDICTIVE FACTOR FOR MODERATE TO SEVERE PEP. THE MECHANISM OF MICRO-PERFORATION, RATHER THAN THE MICRO-PERFORATION ITSELF, CONTRIBUTES TO THE INCREASED SEVERITY OF PEP. IT WAS NOTED THAT CARELESS MANIPULATION OF GUIDEWIRES OR ERCP INSTRUMENTS, SUCH AS FORCEPS, BASKETS, MECHANICAL LITHOTRIPTERS, OR STENTS, CAN LEAD TO PANCREATIC INJURY AND POTENTIALLY WORSEN THE SEVERITY OF PEP, EVEN IF SELECTIVE BILIARY CANNULATION IS SUCCESSFULLY AND SMOOTHLY ACHIEVED. IT WAS REPORTED THERE WAS ONE PEP RELATED DEATH WITHIN THE MODERATE TO SEVERE PEP GROUP. NO FURTHER DETAILS REGARDING THE PATIENT DEATH WERE PROVIDED WITHIN THE LITERATURE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION IDENTIFIED IN THE ARTICLE.Event Description: NO ADDITIONAL INFORMATION RECEIVED FROM CUSTOMER. Manufacturer Narrative: THIS REPORT IS BEING SUPPLEMENTED TO PROVIDE ADDITIONAL INFORMATION BASED ON THE APPROVED FINAL INVESTIGATION. UPDATED FIELDS: D8, H6, H11. BASED ON THE RESULTS OF THE INVESTIGATION, AND SINCE THE DEVICE WAS NOT RETURNED FOR EVALUATION, THE DEFINITIVE ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. OLYMPUS WILL CONTINUE TO MONITOR THE FIELD PERFORMANCE OF THIS DEVICE. Manufacturer Narrative: THIS REPORT IS RELATED TO THE FOLLOWING LINKED PATIENT IDENTIFIERS: (B)(6). THE LITERATURE DESCRIBED "JF-260V" AND "TJF-260V", SO OLYMPUS SELECTED ¿TJF-260V" AS THE REPRESENTATIVE MODEL. THE INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED OR IF ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065536&pc=OOB | 2124215-2025-32860 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, HEART FAILURE, MYOCARDIAL INFARCTION, CARDIOGENIC SHOCK Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065538&pc=LOX | 2124215-2025-32846 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065552&pc=QZI | 2124215-2025-32543 | 2025-04-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER THE PATIENT PASSED AWAY. AN UNSPECIFIED TIME FOLLOWING THE PROCEDURE THE PATIENT "DECOMPENSATED AND THE FAMILY WAS SLOW TO REACT". AFTER SOME TIME THE PATIENT WAS INTUBATED AND STARTED TO IMPROVE. THE PATIENT WAS THEN MOVED TO COMFORT CARE FOR UNSPECIFIED REASONS AND THE PATIENT PASSED AWAY. THE PHYSICIAN DOES NOT BELIEVE THE PATIENT'S DEATH WAS RELATED TO THE PROCEDURE AS THERE WERE OTHER COMPLICATING FACTORS PRESENT AT THE TIME. HOWEVER, THE CAUSE OF DEATH WAS NOT CONFIRMED. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT FURTHER INFORMATION WAS UNAVAILABLE PER THE FACILITY. BECAUSE THE PRODUCT IS UNKNOWN, THUS WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT, PATIENT, OR INCIDENT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065553&pc=LOX | 2124215-2025-32847 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, HEART FAILURE, CARDIAC ARREST, CARDIOGENIC SHOCK Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22065554&pc=OOB | 2124215-2025-32859 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIOGENIC SHOCK. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22066097&pc=OOB | 2124215-2025-32869 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED, DEATH AT DISCHARGE DECEASED. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22066100&pc=LOX | 2124215-2025-32879 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, HEART FAILURE, CARDIAC ARREST, CARDIOGENIC SHOCK Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22066102&pc=OOB | 2124215-2025-32870 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Ischemia Stroke |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIAC ARREST, CARDIOGENIC SHOCK, ISCHEMIC STROKE Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22066112&pc=MCX | 2124215-2025-32889 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | MCX | ROTAPRO | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH AT DISCHARGE DECEASED, ACUTE MYOCARDIAL INFARCTION, HEART FAILURE, CARDIAC ARREST, CARDIOGENIC SHOCK. Manufacturer Narrative: DUE TO SAFE HARBOR DE-IDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22066116&pc=LOX | 2124215-2025-32888 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED, DEATH AT DISCHARGE DECEASED, Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22066133&pc=LOX | 2124215-2025-32878 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED, DEATH AT DISCHARGE DECEASED, Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22066292&pc=OOB | 2124215-2025-32895 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-21 04:00 | OOB | AGENT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Myocardial Infarction; Cardiogenic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT THE FOLLOWING EVENTS OCCURRED DEATH DURING PROCEDURE DECEASED, ACUTE MYOCARDIAL INFARCTION, CARDIAC ARREST, CARDIOGENIC SHOCK. Manufacturer Narrative: DUE TO SAFE HARBOR DEIDENTIFICATION RULES WITH THE REGISTRY DATA OWNER (ACCF), DATA IN THE EVENT DATE AND DATE OF DEATH FIELDS ARE LIMITED TO THE YEAR ONLY. JANUARY 1 OF THE GIVEN YEAR WAS USED TO CAPTURE THIS. UNDER THE TERMS AND CONDITIONS OF THE REGISTRY, ANONYMIZED DATA WAS PROVIDED. NO PRODUCTS WERE RETURNED AS PART OF THE REGISTRY. IT CANNOT BE DETERMINED IF THESE EVENTS HAVE BEEN PREVIOUSLY REPORTED OR IF THE DEVICES WERE RETURNED FOR ANALYSIS AS PART OF A PREVIOUSLY REPORTED SPONTANEOUS COMPLAINT. HOWEVER, THE EVENTS REPORTED WERE ANTICIPATED IN NATURE AS DEFINED BY THE "POTENTIAL ADVERSE EVENTS" LIST IN THE FDA-APPROVED INSTRUCTIONS FOR USE (IFU). DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22049479&pc=NWX | 3007835716-2025-00051 |  | Death | GOODMAN CO., LTD. | 2025-05-20 04:00 | NWX | APERTA NSE CORONARY DILATATION CATHETER | Adverse Event Without Identified Device or Use Problem | Ischemia; Perforation of Vessels |  |  | 1 | Event Description: THE PATIENT WITH A SYMPTOM OF CHEST PAIN AT REST, AN EMERGENT CAG WAS DONE, AND IT REVEALED A SEVERE STENOSIS OF THE RIGHT CORONARY ARTERY WITH DELAYED FLOW AND A SEVERE STENOSIS OF THE MIDDLE PART OF THE LEFT ANTERIOR DESCENDING AND THE OBTUSE MARGIN BRANCH. THE RIGHT CORONARY ARTERY LESION WAS JUDGED TO BE THE CULPRIT SITE, AND THE REVASCULARIZATION FOR THE RIGHT CORONARY ARTERY WAS PERFORMED. BECAUSE THE CARDIAC FUNCTION OF THE PATIENT WAS VERY LOW AND THE SYMPTOM OF HEART FAILURE WAS PERSISTENT, THE ADDITIONAL INTERVENTION FOR THE MIDDLE PART OF THE LEFT ANTERIOR DESCENDING AND OBTUSE MARGINAL BRANCH WAS PERFORMED 2 DAYS AFTER THE FIRST INDEX PROCEDURE. A DRUG-COATED BALLOON (DCB) WAS USED FOR TREATING THE LEFT ANTERIOR DESCENDING LESION AND A GOOD ANGIOGRAPHICAL RESULT WAS OBTAINED. THEN, PCI FOR THE OBTUSE MARGINAL BRANCH WAS PERFORMED, AND THE PRE-PROCEDURAL IVUS IMAGE SHOWED AN ECCENTRIC, APPROXIMATELY 270-DEGREE, CALCIFIED PLAQUE AND A THIN PLAQUE AT THE RESIDUAL SITE, WHOSE LENGTH WAS 6.1 MM. PRE-DILATION USING A APERTA NSE CORONARY DILATATION CATHETER (APN25013) WAS PERFORMED AT NOMINAL PRESSURE AND THE VASCULAR RUPTURE WAS SEEN. THEREFORE, THE COVERED STENT WAS IMMEDIATELY IMPLANTED IN THE OBTUSE MARGINAL BRANCH. FINALLY, A COMPLETE REVASCULARIZATION FOR THE LEFT ANTERIOR DESCENDING AND OBTUSE MARGINAL BRANCH WERE ACCOMPLISHED, THE PATIENT DIED OF NON-OCCLUSIVE MESENTERIC ISCHAEMIA DURING THE PERIOD OF HOSPITALIZATION. A TEMPORARY DECREASE IN BLOOD PRESSURE DURING PCI MAY PARTIALLY EXPLAIN THIS DEVASTATING EVENT. THIS INFORMATION IS TAKEN FROM THE LITERATURE. (EUROPEAN HEART JOURNAL - CASE REPORTS (2023) 7, 1-5). Manufacturer Narrative: THE PRODUCT WAS NOT RETURNED AND THERE WAS NOT ENOUGH INFORMATION AVAILABLE. THEREFORE, IT WAS NOT POSSIBLE TO CONDUCT A DETAILED INVESTIGATION OF THE PRODUCT OR TO CHECK THE MANUFACTURING AND INSPECTION RECORDS. IT IS CONSIDERED THAT THE CONDITION OF THE LESION INFLUENCED THIS INCIDENT, LEADING TO DAMAGE TO THE BLOOD VESSELS WHEN THE PRODUCT WAS EXPANDED, BUT THE DETAILED CAUSE COULD NOT BE IDENTIFIED. THE INSTRUCTIONS FOR USE (IFU) PROVIDE GENERAL INSTRUCTIONS FOR USE, WARNINGS AND PRECAUTIONS RELATED TO THE DEVICE, AND INFORMATION TO MAKE USERS AWARE OF POTENTIAL COMPLICATIONS AND OTHER EVENTS SIMILAR TO THIS EVENT THAT MAY OCCUR. THIS REPORT DOES NOT IMPLY AN ADMISSION BY ANYONE THAT THE PRODUCTS MENTIONED IN THIS REPORT ARE DEFECTIVE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22050144&pc=BZD | 2518422-2025-045106 | 2023-04-07 04:00 | Death | RESPIRONICS, INC. | 2025-05-20 04:00 | BZD | DREAM STATION AUTO BIPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. .THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22051444&pc=LWS | 3004209178-2025-08635 | 2025-03-05 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-05-20 04:00 | LWS | EVERA MRI S DR SURESCAN | Failure to Deliver Shock/Stimulation | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT¿S SPOUSE INDICATED THEY WERE TOLD THAT IF THE HEART WAS ABOUT THE HAVE A FAILURE, THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WOULD ACT AND AVOID FOR THAT TO HAPPEN. ADDITIONALLY, THE PATIENT¿S SPOUSE INDICATED THE PATIENT MAY ALSO SHOW A PHYSICAL REACTION DUE TO A SHOCK IN SUCH A SCENARIO. THE CAUSE OF DEATH WAS REQUESTED BUT NOT RECEIVED. ADDITIONALLY, IT WAS REPORTED THAT THE REMOTE MONITOR MISSED LAST SCHEDULED TRANSMISSION. NO TROUBLESHOOTING TAKEN. SCHEDULED TRANSMISSION DID NOT TAKE PLACE AS NEW REMOTE MONITOR WAS NOT SETUP. THE MONITOR IS EXPECTED TO BE RETURNED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22051729&pc=PNJ | MW5170595 | 2025-05-09 04:00 | Death | ABBOTT MEDICAL | 2025-05-20 04:00 | PNJ | AVEIR LEADLESS RV PACER | Insufficient Information | Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: PATIENT CAME TO CATH LAB FOR LEADLESS PACEMAKER IMPLANT. DURING THE PROCEDURE, BLOOD PRESSURE BEGAN TO TREND DOWNWARDS AND PATIENT WAS URGENTLY INTUBATED. CODE BLUE WAS CALLED, RESUSCITATION PER ACLS PROTOCOL AND MTP WAS INITIATED. SURGEON ARRIVED IN THE ROOM AND OPENED CHEST. ECMO WAS CANNULATED AND STARTED. TIME OF DEATH CALLED AT 12:22.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22051730&pc=MVK | 3008642652-2025-04941 | 2025-04-19 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-20 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 01:48:03 ON (B)(6) 2025. AT 17:43:10, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 70 BPM WITH PAUSES, MOTION ARTIFACT AND LOW AMPLITUDE CARDIAC SIGNAL. AT 17:44:17, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL/CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS RHYTHM @ 60 BPM WITH PAUSES, MOTION ARTIFACT AND LOW AMPLITUDE CARDIAC SIGNAL. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM WITH PAUSES, MOTION ARTIFACT AND LOW AMPLITUDE CARDIAC SIGNAL. AT 19:00:00, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION/TACTILE ARTIFACT. AT 19:00:54, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL/CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH MOTION/TACTILE ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION/TACTILE ARTIFACT. AT 19:01:26, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL/CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH MOTION/TACTILE ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND MOTION/TACTILE ARTIFACT. THE DEVICE WAS SHUT DOWN AT 19:01:34 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22051764&pc=DXF | MW5170596 | 2025-05-09 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-20 04:00 | DXF | VERSACROSS RF WIRE | Use of Device Problem | Cardiac Perforation; Pericardial Effusion |  |  | 1 | Event Description: DURING AN ATRIAL FIBRILLATION (AFIB) CASE, A PERFORATION WAS NOTICED. THE INJURY WAS DISCOVERED WHEN A VERY SIGNIFICANT EFFUSION WAS OBSERVED ON INTRACARDIAC ECHO (ICE) POST TRANSSEPTAL. THEY LOOKED AT ICE BEFORE GOING TRANSSEPTAL AND AGAIN AFTER GOING TRANSSEPTAL AND NOTICED A SIGNIFICANT AMOUNT OF FLUID. THE INJURY WAS CONFIRMED ON ICE. A PERICARDIOCENTESIS WAS ATTEMPTED ON THE PATIENT. THEY PULLED OUT OVER A LITER OF FLUID. THE PATIENT CONTINUED TO BLEED. THE PATIENTS CHEST WAS OPENED UP AND THEY NOTICED 1 PERFORATION IN THE RIGHT VENTRICLE (RV) THAT THEY BELIEVED MAY HAVE BEEN CAUSED BY THE TAP FOR THE PERICARDIOCENTESIS. THE LEFT APPENDAGE HAD A COUPLE OF PERFORATIONS AS WELL THAT THEY BELIEVE WAS FROM THE DILATOR. THE PATIENT WAS PUT ON BYPASS AT 10:30AM. THE LAST KNOWN STATUS OF THE PATIENT WAS DECEASED. A REPROCESSED STERILMED SOUNDSTAR CATHETER AND AN OCTARAY CATHETER WERE USED DURING THE PROCEDURE. NO ABLATION WAS PERFORMED DURING THE PROCEDURE. THEY ONLY MAPPED THE CASE. THE CALLER STATED THEY ASSUMED THE FARADRIVE CATHETER AND THE VERSACROSS WIRE WERE THE CAUSE OF THE INJURY. A FARADRIVE SHEATH AND A BOSTON CORONARY SINUS (CS) WERE IN THE BODY WHEN THE EVENT OCCURRED. THE CALLER STATED THEY BELIEVE THE FARADRIVE DILATOR PERFORATED THE LEFT APPENDAGE WHILE GOING TRANSSEPTAL. FARAPULSE GENERATOR WAS IN USE. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22051818&pc=OZD | 1220648-2025-28555 | 2025-04-26 04:00 | Death | ABIOMED, INC. | 2025-05-20 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: IT WAS REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THERE WERE SUCTION ALARMS. VOLUME WAS GIVEN AND THE POSITION WAS CHECKED. HOWEVER, FURTHER DETAILS STATED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE PURGE CASSETTE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW HAS BEEN COMPLETED. THE CAUSE OF THE LOW PUMP FLOW WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED, DATA LOGS WERE NOT RECOVERED, AND INSUFFICIENT CLINICAL DETAILS. ONLY THE CASSETTE WAS RETURNED FOR INVESTIGATION, NOT THE PUMP.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22051994&pc=OZD | 1220648-2025-28558 | 2025-04-28 04:00 | Death | ABIOMED, INC. | 2025-05-20 04:00 | OZD | IMPELLA 5.5-CA | Positioning Failure | Sepsis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE IMPELLA 5.5 HAD NUMEROUS SUCTION ALARMS OCCURRING OVER P-7. A DECISION WAS MADE TO USE TEE TO DETERMINE POSITION OF IMPELLA GIVEN CVP WAS 11. ON TEE, THE POSITION WAS NOTED TO BE 6.4CM DEEP. THE PHYSICIAN ELECTED TO PULL THE DEVICE BACK ABOUT 1.5 CM UNDER TEE GUIDANCE, MEASURED AT 4.7 CM. IT WAS FURTHER REPORTED THAT FAULTY POSITIONING METRIC WAS CAUSING THE AIC TO SENSE THAT THE IMPELLA 5.5 WAS IN THE AORTA. CONFIRMATION OF POSITION UNDER TEE REVEALED THE IMPELLA WAS IN THE CORRECT SPOT AND WAS NOT ABUTTED AGAINST OTHER STRUCTURES IN THE HEART. REPOSITIONING WAS DONE TO OPTIMAL POSITION DID NOT HAVE AN EFFECT ON THE ALARM. THE ALARM WAS RESOLVED BY DISABLING POSITION METRICS. IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." . IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ THE PATIENT OUTCOME WAS REVIEWED BY ABIOMED'S MEDICAL SAFETY OFFICER. THERE IS NO ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR TO THE OUTCOME. THERE WAS REPORT OF SEPSIS, AND THE IMPELLA DEVICE CANNOT BE RULED OUT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22052017&pc=OZP | MW5170598 |  | Death | MEDTRONIC MINIMED | 2025-05-20 04:00 | OZP | GUARDIAN GLUCOSE SENSOR | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: PATIENT'S DAUGHTER CALLED IN TO INFORM US THAT PATIENT HAS PASSED AWAY. REFERENCE REPORT #MW5170599, #MW5170600.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22052050&pc=OZP | MW5170599 |  | Death | MEDTRONIC MINIMED | 2025-05-20 04:00 | OZP | GUARDIAN GLUCOSE TRANSMITTER | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: PATIENT'S DAUGHTER CALLED IN TO INFORM US THAT PATIENT HAS PASSED AWAY. REFERENCE REPORT #MW5170598, #MW5170600.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22052063&pc=NBW | MW5170600 |  | Death | ROCHE DIAGNOSTICS OPERATIONS, INC. | 2025-05-20 04:00 | NBW | ACCU-CHECK GUIDE TEST STRIP | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: PATIENT'S DAUGHTER CALLED IN TO INFORM US THAT PATIENT HAS PASSED AWAY. REFERENCE REPORT #MW5170598, #MW5170599.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22052180&pc=IYN | 3019216-2025-000199 | 2025-01-09 05:00 | Death | PHILIPS ULTRASOUND | 2025-05-20 04:00 | IYN | INNOSIGHT | Poor Quality Image | Insufficient Information | K192254 |  | 1 | Event Description: IT WAS REPORTED THAT A THROMBUS WAS NOT VISUALIZED ON THE INNOSIGHT ULTRASOUND SYSTEM DURING A PORTABLE FOCUSED ECHO SCAN AND PATHOLOGY. A REPEAT ECHO WAS PERFORMED AT A LATER DATE USING A DIFFERENT SYSTEM. THE THROMBUS WAS VISUALIZED. THE PATIENT RECEIVED A BONE MARROW TRANSPLANT AND SUBSEQUENTLY DIED. AN INTERNAL SITE INVESTIGATION IS CURRENTLY UNDERWAY TO DETERMINE WHETHER THE LACK OF VISUALIZATION OF THE SUBSTRATE CONTRIBUTED TO CAUSE OF DEATH. PHILIPS HAS STARTED AN INVESTIGATION, AND UPON COMPLETION, A FOLLOW-UP REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22052257&pc=NVZ | 3008973940-2025-03568 | 2024-10-08 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-05-20 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Dyspnea; Fatigue; Hemorrhage/Bleeding; Hypovolemic Shock; Tricuspid Valve Insufficiency/ Regurgitation; Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED SEVERE TRICUSPID REGURGITATION WITH SIGNIFICANT PROGRESSION OVER THE PAST SEVERAL MONTHS WITH FATIGUE AND DYSPNEA ON EXERTION. IT WAS ALSO NOTED THAT THE PATIENT WAS "FINDING IT DIFFICULT TO PERFORM ACTIVITIES OF DAILY LIVING". DURING A PROCEDURE FOR SEVERE TRICUSPID REGURGITATION THE PATIENT EXPERIENCED SEVERE HEMORRHAGIC SHOCK DUE TO AN UNCONTROLLED INTRAOPERATIVE HEMORRHAGE. THE PATIENT DIED. THE PATIENT IS A PARTICIPANT IN THE POST APPROVAL CLINICAL SURVEILLANCE PRODUCT SURVEILLANCE REGISTRY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22052275&pc=DTB | 2649622-2025-13664 | 2024-10-08 04:00 | Death | MPRI | 2025-05-20 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Dyspnea; Fatigue; Hemorrhage/Bleeding; Hypovolemic Shock; Tricuspid Valve Insufficiency/ Regurgitation; Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED SEVERE TRICUSPID REGURGITATION WITH SIGNIFICANT PROGRESSION OVER THE PAST SEVERAL MONTHS WITH FATIGUE AND DYSPNEA ON EXERTION. IT WAS ALSO NOTED THAT THE PATIENT WAS "FINDING IT DIFFICULT TO PERFORM ACTIVITIES OF DAILY LIVING". DURING A PROCEDURE FOR SEVERE TRICUSPID REGURGITATION THE PATIENT EXPERIENCED SEVERE HEMORRHAGIC SHOCK DUE TO AN UNCONTROLLED INTRAOPERATIVE HEMORRHAGE. THE PATIENT DIED. THE PATIENT IS A PARTICIPANT IN THE POST APPROVAL CLINICAL SURVEILLANCE PRODUCT SURVEILLANCE REGISTRY. Manufacturer Narrative: CORRECTED: C. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: G3 FIELD FROM PREVIOUS FOLLOW-UP REPORT 001 WAS OMITTED. THE OMITTED DATE SHOULD HAVE BEEN (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22052301&pc=NVN | 2649622-2025-13665 | 2024-10-08 04:00 | Death | MPRI | 2025-05-20 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Dyspnea; Fatigue; Hemorrhage/Bleeding; Hypovolemic Shock; Tricuspid Valve Insufficiency/ Regurgitation; Insufficient Information | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED SEVERE TRICUSPID REGURGITATION WITH SIGNIFICANT PROGRESSION OVER THE PAST SEVERAL MONTHS WITH FATIGUE AND DYSPNEA ON EXERTION. IT WAS ALSO NOTED THAT THE PATIENT WAS "FINDING IT DIFFICULT TO PERFORM ACTIVITIES OF DAILY LIVING". DURING A PROCEDURE FOR SEVERE TRICUSPID REGURGITATION THE PATIENT EXPERIENCED SEVERE HEMORRHAGIC SHOCK DUE TO AN UNCONTROLLED INTRAOPERATIVE HEMORRHAGE. THE PATIENT DIED. THE PATIENT IS A PARTICIPANT IN THE POST APPROVAL CLINICAL SURVEILLANCE PRODUCT SURVEILLANCE REGISTRY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22052581&pc=OZD | 1220648-2025-28565 | 2025-04-29 04:00 | Death | ABIOMED, INC. | 2025-05-20 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA 5.5 WAS USED TO SUPPORT A PATIENT WITH NON-ST SEGMENT ELEVATION MYOCARDIAL INFARCTION AND DURING A HIGH RISK CARDIAC PROCEDURE. THE PATIENT HAD DECOMPENSATED PRIOR PUMP IMPLANTATION REQUIRING RESUSCITATION AND A BALLOON PUMP PLACED. DURING SUPPORT, THE PATIENT EXPERIENCED RUNS OF VENTRICULAR TACHYCARDIA (VT). BLOOD PRESSURE AND FLOWS REMAINED STABLE. ON A DIFFERENT DATE IT WAS AGAIN NOTED OF EPISODES OF VT, THAT REQUIRED SHOCK AND THAT PATIENT WAS ON LIDOCAINE AND AMIODARONE. OXYGEN WAS INCREASED. THE PATIENT WAS UNSTABLE AND REQUIRED REINTUBATION. THE PATIENT WAS TAKEN DOWN TO THE CATHETERIZATION LAB TO ATTEMPT TO OPEN THE LEFT ANTERIOR DESCENDING ARTERY WITH NO SUCCESS. ULTIMATELY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22052586&pc=LWP | 2124215-2025-32287 | 2025-04-28 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-20 04:00 | LWP | INGENIO | Signal Artifact/Noise; Over-Sensing; Pacing Problem; Incorrect, Inadequate or Imprecise Result or Readings | Dizziness; Asystole | N970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT WITH DIZZINESS. UPON INTERROGATION THE PACEMAKER WAS OBSERVED TO BE IN SAFETY MODE AND PACING INHIBITION HAD OCCURRED, WITH PAUSES NOTED ON TELEMETRY. THE PATIENT HAD MULTIPLE MEDICAL COMORBIDITIES, INCLUDING DEMENTIA, AND WAS TRANSITIONED TO HOSPICE DUE TO HIS INTRACTABLE SYMPTOMS. THE PATIENT SUBSEQUENTLY PASSED AWAY (DATE OF DEATH NOT REPORTED). IT WAS NOTED THAT THE PATIENT HAD KNOWN PACEMAKER DEPENDENCE AND THAT PRIOR TO THE DEVICE ENTERING SAFETY MODE HE HAD BEEN STABLE IN THE OUTPATIENT SETTING. IT WAS NOT REPORTED WHETHER THE DEVICE WOULD BE RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22053596&pc=DSQ | 2916596-2025-03020 | 2025-04-25 04:00 | Death | THORATEC CORPORATION | 2025-05-20 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hypoxia; Cardiogenic Shock; Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIOGENIC AND RIGHT VENTRICLE SHOCK ACUTE HYPOXIC HYPERCARBIC RESPIRATORY FAILURE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. Manufacturer Narrative: CORRECTED DATA: SECTION D3: MANUFACTURER INFORMATION CORRECTED. SECTION G1: MANUFACTURING SITE CORRECTED. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION REGARDING THIS EVENT WERE SUBMITTED TO THE ACCOUNT; HOWEVER, NO RESPONSE HAS BEEN RECEIVED, AND HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), HAS NOT BEEN RETURNED FOR EVALUATION AT THIS TIME. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH, RIGHT HEART FAILURE, AND RESPIRATORY FAILURE. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE¿) ALSO OUTLINES INDICATIONS OF RIGHT HEART FAILURE AS WELL AS POSSIBLE TREATMENTS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22053695&pc=PNJ | 2017865-2025-64130 | 2025-04-29 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-20 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT ATRIUM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC FOR DUAL CHAMBER AVEIR IMPLANT. DURING THE PROCEDURE, THE ATRIAL LP WAS NOTED TO EXHIBIT LOW PACING IMPEDANCE AND HIGH CAPTURE THRESHOLD, INDICATING POOR POSITIONING ON FIXATION. IT WAS DECIDED THAT THE ATRIAL DEVICE SHOULD BE REPOSITIONED. THE DEVICE DID NOT RELEASE FROM THE ATRIAL WALL, AND AFTER PULLING ON THE DEVICE, IT WAS OBSERVED THAT THE HELIX BECAME STRETCHED. A NEW ATRIAL DEVICE WAS ATTEMPTED. THE PATIENT¿S BLOOD PRESSURE STARTED TO CRASH. THE PATIENT WAS FOUND TO BE IN CARDIAC TAMPONADE. A PERICARDIOCENTESIS WAS PERFORMED. THE PATIENT STABILIZED AND THE PROCEDURE CONTINUED. AS THE DEVICE TOUCHED THE ATRIAL WALL, THE PATIENT CRASHED. THE PATIENT WAS PRONOUNCED DECEASED ON (B)(6) 2025. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22053700&pc=PNJ | 2017865-2025-64129 | 2025-04-29 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-20 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT ATRIUM | Positioning Failure; Stretched; Low impedance; Capturing Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC FOR DUAL CHAMBER AVEIR IMPLANT. DURING THE PROCEDURE, THE ATRIAL LP WAS NOTED TO EXHIBIT LOW PACING IMPEDANCE AND HIGH CAPTURE THRESHOLD, INDICATING POOR POSITIONING ON FIXATION. IT WAS DECIDED THAT THE ATRIAL DEVICE SHOULD BE REPOSITIONED. THE DEVICE DID NOT RELEASE FROM THE ATRIAL WALL, AND AFTER PULLING ON THE DEVICE, IT WAS OBSERVED THAT THE HELIX BECAME STRETCHED. A NEW ATRIAL DEVICE WAS ATTEMPTED. THE PATIENT¿S BLOOD PRESSURE STARTED TO CRASH. THE PATIENT WAS FOUND TO BE IN CARDIAC TAMPONADE. A PERICARDIOCENTESIS WAS PERFORMED. THE PATIENT STABILIZED AND THE PROCEDURE CONTINUED. AS THE DEVICE TOUCHED THE ATRIAL WALL, THE PATIENT CRASHED. THE PATIENT WAS PRONOUNCED DECEASED ON (B)(6) 2025. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22053711&pc=PNJ | 2017865-2025-64131 | 2025-04-29 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-20 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC FOR DUAL CHAMBER AVEIR IMPLANT. DURING THE PROCEDURE, THE ATRIAL LP WAS NOTED TO EXHIBIT LOW PACING IMPEDANCE AND HIGH CAPTURE THRESHOLD, INDICATING POOR POSITIONING ON FIXATION. IT WAS DECIDED THAT THE ATRIAL DEVICE SHOULD BE REPOSITIONED. THE DEVICE DID NOT RELEASE FROM THE ATRIAL WALL, AND AFTER PULLING ON THE DEVICE, IT WAS OBSERVED THAT THE HELIX BECAME STRETCHED. A NEW ATRIAL DEVICE WAS ATTEMPTED. THE PATIENT¿S BLOOD PRESSURE STARTED TO CRASH. THE PATIENT WAS FOUND TO BE IN CARDIAC TAMPONADE. A PERICARDIOCENTESIS WAS PERFORMED. THE PATIENT STABILIZED AND THE PROCEDURE CONTINUED. AS THE DEVICE TOUCHED THE ATRIAL WALL, THE PATIENT CRASHED. THE PATIENT WAS PRONOUNCED DECEASED ON (B)(6) 2025. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22053719&pc=DSQ | 2916596-2025-02961 | 2025-04-16 04:00 | Death | THORATEC CORPORATION | 2025-05-20 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE DEVICE WAS NOT EXPLANTED. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CAUSE OF DEATH WAS REPORTED AS UNKNOWN. AN AUTOPSY WAS NOT PERFORMED. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION REGARDING THIS EVENT WERE SUBMITTED TO THE ACCOUNT; HOWEVER, NO RESPONSE HAS BEEN RECEIVED AT THIS TIME. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22053734&pc=DSQ | 2916596-2025-02997 | 2024-07-01 04:00 | Death | THORATEC CORPORATION | 2025-05-20 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: EVENT DETAILS: IT WAS REPORTED IN THE RESEARCH ARTICLE ¿MITIGATING POST-OPERATIVE RIGHT VENTRICULAR DYSFUNCTION AFTER LEFT VENTRICULAR ASSIST DEVICE: THE RV PROTECTION STUDY¿ THAT THE HEARTMATE 3 MAY BE RELATED TO RIGHT HEART FAILURE AND DEATH. A PROSPECTIVE, RANDOMIZED, CONTROLLED TRIAL WAS PERFORMED AT A SINGLE CENTER AND WERE PROSPECTIVELY RANDOMIZED IN A 1:1 FASHION TO EITHER THE RV-PROTECTION STRATEGY OR USUAL CARE FOR POST-OPERATIVE LVAD MANAGEMENT. PARTICIPANTS WERE FOLLOWED FOR UP TO 24 WEEKS. ELIGIBLE PARTICIPANTS WERE AT LEAST 18 YEARS OF AGE AND HAD ADVANCED HEART FAILURE, WITH LEFT VENTRICULAR EJECTION FRACTION (LVEF) < 25% AND/OR ON INOTROPIC THERAPY, WHO WERE SCHEDULED TO UNDERGO LVAD IMPLANTATION WITHOUT THE INTENTION OF UPFRONT PERIOPERATIVE RVAD SUPPORT. BETWEEN (B)(6) 2023 AND (B)(6) 2024, A TOTAL OF 60 LVAD RECIPIENTS WERE SCREENED, AND 20 PARTICIPANTS WERE ENROLLED. THE 20 PARTICIPANTS WERE RANDOMIZED PRE-LVAD IMPLANTATION: 10 TO THE USUAL-CARE ARM AND 10 TO THE RV-PROTECTION STRATEGY. ALL PATIENTS WHO WERE RANDOMIZED RECEIVED THE ASSIGNED INTERVENTION. THE RV PROTECTION STRATEGY CONSISTED OF A COMPREHENSIVE UPFRONT RV-MANAGEMENT APPROACH TARGETING RV AFTERLOAD (INHALED NITRIC OXIDE [INO] FOR 48 HOURS, PULMONARY CAPILLARY WEDGE PRESSURE [PCWP] < 18 MMHG), RV PRELOAD (CENTRAL VENOUS PRESSURE [CVP] 8 14 MMHG), RV CORONARY PERFUSION (MAP GOAL 70 90 MMHG, HEMOGLOBIN [HGB] > 8 G/DL), RV CONTRACTILITY (INTRAVENOUS [IV] INOTROPES), RATE AND RHYTHM CONTROL (HEART RATE [HR] > 100 BPM, NORMAL SINUS RHYTHM), VENTILATORY MANAGEMENT (OXYGEN SATURATION [SPO2] > 95%, PARTIAL PRESSURE OF CARBON DIOXIDE [PACO2] < 30, POSITIVE END-EXPIRATORY PRESSURE [PEEP] 5), AND RV GEOMETRY (ECHO-GUIDED SEPTAL POSITION MIDLINE OR SLIGHTLY RIGHTWARD) ON POST-OPERATIVE DAYS (PODS) 1 AND 3. PARTICIPANTS IN THE RV-PROTECTION STRATEGY WERE ADHERENT TO ALL 7 INTERVENTIONS DURING THEIR TIME IN THE ICU, ESPECIALLY IN THE FIRST 48 TO 72 HOURS OTHER THAN 2 PATIENTS WHOSE HEART RATES WERE SLIGHTLY LOWER THAN 100 BPM. ONLY 3 PATIENTS HAD UPFRONT INO IN THE USUAL-CARE MANAGEMENT, WHEREAS 3 MORE PATIENTS WERE STARTED ON INO LATER IN THEIR COURSES AFTER THERE WAS CONCERN FOR RV DYSFUNCTION. IN ADDITION, ONLY 50% TO 60% OF PARTICIPANTS FELL WITHIN MANY OF THE OTHER PARAMETERS THAT ARE THOUGHT TO BE STANDARD, INCLUDING RV PRELOAD, RV PERFUSION AND RATE/RHYTHM CONTROL. 1 PATIENT REQUIRED IV INOTROPIC THERAPY DUE TO RVF UNTIL SHE EXPIRED AFTER 24WEEKS, ANOTHER PATIENT EXPERIENCED IN-HOSPITAL MORTALITY SECONDARY TO RVF WITHIN THE FIRST 30 DAYS, AND THE THIRD PATIENT REQUIRED RVAD PLACEMENT WITHIN THE FIRST 30 DAYS. THERE WERE NO SIGNIFICANT DIFFERENCES IN BASELINE CHARACTERISTICS BETWEEN THE 2 GROUPS. THE MEDIAN AGE WAS 60 YEARS, 75% OF PATIENTS WERE MALE, 50% OF PATIENTS WERE BLACK, 50% HAD PREDOMINANTLY ISCHEMIC CARDIOMYOPATHY, AND 50% WERE ON INOTROPIC THERAPY PRIOR TO LVAD IMPLANTATION. THE PRIMARY OUTCOME WAS SURVIVAL FREE FROM ANY HDRES (EG, RVF AND/OR AI) OR HRAES (NONSURGICAL BLEEDING, THROMBOEMBOLIC EVENT, PUMP THROMBOSIS, OR NEUROLOGICAL EVENT). OF NOTE, THE HDRES CONSISTED OF ONLY RVF; NO PATIENTS HAD MILD/MODERATE OR GREATER AI OVER THE 24-WEEKFOLLOW-UP. SECONDARY OUTCOMES INCLUDED SEVERE RVF BY THE INTERMACS DEFINITION, EARLY/LATE RVF BY THE 2020 ACADEMIC RESEARCH CONSORTIUM (ARC) DEFINITION, AND CHANGES IN HEMODYNAMICS AND TTE PARAMETERS. PARTICIPANTS RECEIVING A NOVEL, COMPREHENSIVE, UP FRONT RV PROTECTION STRATEGY POST-LVAD IMPLANTATION HAD SIGNIFICANTLY GREATER SURVIVAL RATES FREE FROM HDRES OR HRAES AT 24WEEKS. Manufacturer Narrative: SECTION B: DATE OF DEATH ESTIMATED AS DATE OF EVENT. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM AND THE SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A, D4: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT IS APPROXIMATE AS THE DATA WERE COLLECTED BETWEEN 18MAR2023 AND 1JUL2024. KANELIDIS, ET AL. MITIGATING POST-OPERATIVE RIGHT VENTRICULAR DYSFUNCTION AFTER LEFT VENTRICULAR ASSIST DEVICE: THE RV PROTECTION STUDY. JOURNAL OF CARDIAC FAILURE. HTTPS://DOI.ORG/10.1016/J.CARDFAIL.2025.01.017 CHICAGO,IL;COLUMBUS,OH;ANDSINGAPORE, SINGAPORE NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22054112&pc=NVN | 2017865-2025-64138 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-20 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22054116&pc=NVN | 2017865-2025-64135 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-20 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22054139&pc=LWP | 2017865-2025-64134 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-20 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO PATIENT EXPIRED. THE DEVICE WAS RECEIVED WITH A BATTERY LEVEL WITHIN THE NORMAL RANGE OF OPERATION. ELECTRICAL, LONGEVITY, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22054346&pc=MVK | 3008642652-2025-04965 | 2025-04-21 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-20 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO APPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 15:28:31 ON (B)(6) 2025. AT 03:21:06 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS RHYTHM @ 70 BPM DEGRADING TO VF WITH MOTION ARTIFACT. AT 03:21:42, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM DEGRADING TO ASYSTOLE. AT 03:23:49, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS BRADYCARDIA @ 40 BPM WITH PVCS DEGRADING TO VF. AT 03:24:19, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM. AT 03:26:19, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS BRADYCARDIA @ 40 BPM WITH PVCS DEGRADING TO VF WITH CPR/MOTION ARTIFACT, LOW/VARYING AMPLITUDE SIGNAL AND ELECTRODE LEAD FALLOFF. THE DEVICE PROPERLY DETECTED VF. LOW/VARYING AMPLITUDE SIGNAL AND CPR/MOTION ARTIFACT PREVENTED THE LIFEVEST FROM DELIVERING ANOTHER TREATMENT. THE ELECTRODE BELT WAS DISCONNECTED AT 04:03:40 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22054425&pc=MOZ | 3007093114-2025-00006 | 2025-01-01 05:00 | Death | ANIKA | 2025-05-20 04:00 | MOZ | ORTHOVISC | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THIS REPORT IS BEING SUBMITTED IN PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ANIKA OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, ANIKA, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO A POTENTIAL PATIENT EVENT DOCUMENTED ON THIS REPORT. A MEDWATCH REPORT (MW5169518) WAS RECEIVED FROM THE FDA. A PATIENT OF UNKNOWN AGE AND DEMOGRAPHICS REPORTEDLY EXPIRED. THE CAUSE OF THE PATIENT DEATH HAS NOT BEEN REPORTED. ORTHOVISC (LOT NUMBER WAS NOT PROVIDED) WAS LISTED AS THE DEVICE. THE DATE OF THE ORTHOVISC TREATMENT AND THE PATIENT DEATH WAS NOT REPORTED ON THE MEDWATCH. PATIENT COMORBIDITIES IS UNKNOWN. THERE WAS NO REPORT OF A DEVICE MALFUNCTION. CONTACT INFORMATION WAS NOT PROVIDED SO A REQUEST FOR ADDITIONAL INFORMATION COULD NOT BE PERFORMED. THE DISTRIBUTOR COPIED THE REPORT AND DID NOT INDICATE THIS CASE WAS REPORTED TO THEM. Manufacturer Narrative: THE REPORTED EVENT IS NOT CONFIRMED. ADDITIONAL INFORMATION COULD NOT BE OBTAINED DUE TO THE CONTACT INFORMATION NOT BEING PROVIDED ON THE MEDWATCH REPORT. THE LOT NUMBER WAS NOT PROVIDED, THEREFORE A REVIEW OF THE BATCH RECORD COULD NOT BE PERFORMED. A THREE-YEAR RETROSPECTIVE REVIEW OF THE NCRS THIS PRODUCT WAS PERFORMED. THERE WAS NO NONCONFORMANCE'S DOCUMENTED THAT WAS RELATED TO THE REPORTED EVENT. A REVIEW OF THE RECENT STABILITY STUDY REPORT FOR ORTHOVISC WAS PERFORMED. THE CONCLUSION FOR THE PRODUCT STATED THAT THE PRODUCT HAS MET ALL REQUIRED SPECIFICATIONS. THE CLINICAL REVIEW OF THIS CASE DETERMINED THAT THE CAUSE NOT ESTABLISHED BASED ON INSUFFICIENT INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON RECEIPT OF NEW AND RELEVANT INFORMATION. THIS CASE WILL BE MONITORED AND TRENDED FOR FUTURE ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22054851&pc=OZD | 1220648-2025-28581 | 2025-04-29 04:00 | Death | ABIOMED, INC. | 2025-05-20 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT PRESENTED WITH TRIPLE VESSEL DISEASE AND NEEDS IMPELLA CP SUPPORT. WHILE ON SUPPORT THE PATIENT WAS HAVING DIFFUSE BLEEDING FROM MULTIPLE ACCESS SITES, INCLUDING HEMATOMA TO IMPELLA SITE. IMPELLA WAS SUTURED TO SKIN AND PATIENT TRANSPORTED BACK TO THE INTENSIVE CARE UNIT (ICU). WHILE TRANSPORTING TO ICU, PATIENT LOST NATIVE HEART PULSATILITY AND IMPELLA FLOWS DROPPED. UPON ARRIVAL TO ICU, UNABLE TO FIND PALPABLE PULSE AND CARDIOPULMONARY RESUSCITATION INITIATED. PATIENT CODED IN ICU ROOM FOR APPROXIMATELY 30 MINUTES. RETURN OF SPONTANEOUS CIRCULATION OBTAINED. DURING SUPPORT IT WAS ALSO NOTED OF BLEEDING CONTINUED, NOT JUST FROM IMPELLA SITE, BUT ALL ACCESS SITES. IN ADDITION, TO UNABLE TO DOPPLER BILATERAL LOWER EXTREMITY PULSES. SURGICAL TEAM NOTIFIED THAT IF PATIENT STABILIZED, ESCALATION TO IMPELLA 5.5 MAY BE NECESSARY. HOWEVER CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER. ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22055047&pc=MKJ | 3030677-2025-001370 | 2022-11-04 04:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-05-20 04:00 | MKJ | HEARTSTART | Failure to Deliver Shock/Stimulation | Cardiac Arrest | P160029 |  | 1 | Event Description: IT HAS BEEN REPORTED TO PHILIPS THAT THE DEVICE MAY NOT BE FUNCTIONING AS EXPECTED. THE PATIENT DID NOT SURVIVE. Manufacturer Narrative: THE INITIAL "DATE RECEIVED BY MANUFACTURER" ON EMDR 6083778 WITH MFR REPORT # 3030677-2025-001370 SHOULD BE 15MAY2025. Manufacturer Narrative: UPDATED MODEL, AND CATALOG NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22055265&pc=OZD | 1220648-2025-28585 | 2025-04-29 04:00 | Death | ABIOMED, INC. | 2025-05-20 04:00 | OZD | IMPELLA CP | Positioning Failure; Device Sensing Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemolysis; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ON IMPELLA CP SUPPORT ARRIVED AT ANOTHER FACILITY WITH RED URINE AND THE CP WAS REPOSITIONED FROM 6.5CM TO 4.5CM. 30 MIN POST REPOSITIONING THE PATIENT WAS ASYSTOLIC. DESPITE MULTIPLE ROUNDS OF ADVANCED CARDIAC LIFE SUPPORT, PULSE WAS RESUMED BRIEFLY AND IMPELLA REPOSITIONED UNDER ECHO DESPITE OPTICAL SENSOR FAILURE AFTER INITIAL COMPRESSIONS. THE PATIENT WENT ASYSTOLIC AGAIN, AND HEALTHCARE TEAM UNABLE TO REGAIN PERFUSABLE RHYTHM. THE PATIENT EXPIRED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT COULD HAVE LIKELY LIMB ISCHEMIA. THE TEAM WAS WORKING ON DETERMINING A PLAN OF ACTION WHEN THE PATIENT ARRESTED, SO NO INTERVENTIONS HAD A CHANCE TO BE PERFORMED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES, VF/VT/AF/AT, HEMOLYSIS, LIMB ISCHEMIA, AND OPTICAL SIGNAL ISSUE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE HEMOLYSIS WAS MOST LIKELY DUE TO IMPROPER POSITIONING BASED ON THE CLINICAL DETAILS PROVIDED. THE CLINICAL STATED THAT THE IMPELLA WAS MEASURING 6.5CM. IT IS UNKNOWN THE ROOT CAUSE OF THE POSITIONING ISSUES DUE TO LACK OF CLINICAL DETAILS. THE CLINICAL STATED THAT THERE WAS OPTICAL SENSOR FAILURE AFTER COMPRESSIONS. HOWEVER, THERE WAS NO PRODUCT OR DATA LOGS RETURNED SO THE ROOT CAUSE COULD NOT BE DETERMINED. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF LIMB ISCHEMIA AND VT/VF WAS NOT DETERMINED. B2 REQUIRE INTERVENTION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. B5 INFORMATION ADDED THAT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER. H6 HEALTH EFFECT - CLINICAL CODE 1942, HEALTH EFFECT - IMPACT CODE 4647 AND COMPONENT CODE 3038 WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22055349&pc=PYX | 1220648-2025-28587 | 2025-04-29 04:00 | Death | ABIOMED, INC. | 2025-05-20 04:00 | PYX | IMPELLA RP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED THE PATIENT WAS ON IMPELLA RP FLEX SUPPORT AND DURING SUPPORT, THE PATIENT HAD OOZING. AN ADDITIONAL SUTURE WAS PLACED AND THE OOZING RESOLVED. SUPPORT CONTINUED. IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED ON SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA RP FLEX WITH SMARTASSIST INSTRUCTIONS FOR USE & CLINICAL REFERENCE MANUAL SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP FLEX SMARTASSIST SYSTEM CATHETER USE OF THE REPOSITIONING SHEATH AND THE 23 FR PEEL-AWAY INTRODUCER ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE INTERNAL JUGULAR VEIN. CLOSE AND DRESS THE WOUND.¿ IMPELLA RP WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS AND CAUTIONS ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE INTERNAL JUGULAR VEIN. CLOSE AND DRESS THE WOUND ¿BE SURE THAT THE STOPCOCK ON THE REPOSITIONING SHEATH IS ALWAYS KEPT IN THE CLOSED POSITION. SIGNIFICANT BLEED BACK CAN RESULT IF THE STOPCOCK IS OPEN.¿ ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE FEMORAL VEIN. CLOSE AND DRESS THE WOUND.¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF ACCESS SITE BLEEDING WAS MOST LIKELY USE ISSUE SINCE TIGHTENING THE SUTURE AND ADDING ADDITIONAL SUTURE RESOLVED THE BLEED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22055357&pc=DRA | 2124215-2025-31606 | 2025-04-18 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-20 04:00 | DRA | FARADRIVE | Use of Device Problem | Air Embolism; Arrhythmia; Non specific EKG/ECG Changes; Brain Injury |  |  | 1 | Event Description: IT WAS REPORTED THAT A USER ERROR OCCURRED WITH IRRIGATION OF THE SHEATH AND THE PATIENT EXPERIENCED AIR EMBOLISM, ST SEGMENT ELEVATION, BRAIN DAMAGE, AND VENTRICULAR TACHYCARDIA (VT). THE PATIENT DIED. DURING A PULSE FIELD ABLATION (PFA) PULMONARY VEIN ISOLATION (PVI) PROCEDURE TO TREAT PAROXYSMAL ATRIAL FIBRILLATION A FARADRIVE SHEATH WAS USED. UPON GAINING TRANSSEPTAL ACCESS 20CC OF BLOOD WAS ASPIRATED FROM THE SHEATH. THE FLUSH LINE WAS CONNECTED TO A PRESSURE BAG AND THEN CONNECTED TO THE SHEATH BEFORE INSERTION OF A NON-BOSTON SCIENTIFIC MAPPING CATHETER. A LEFT ATRIAL VOLTAGE MAP WAS ACQUIRED USING THE MAPPING CATHETER, AND THEN IT WAS EXCHANGED FOR THE FARAWAVE CATHETER. THE CATHETER WAS CONFIRMED TO BE DRIPPING WHEN CONNECTED TO A PRESSURE BAG, AND ANOTHER 20CC WERE ASPIRATED PRIOR TO ADVANCING THE CATHETER INTO THE LEFT ATRIUM. AFTER COMPLETING ABLATION, THE FARAWAVE CATHETER WAS REMOVED AND 20 CC WAS ASPIRATED AGAIN, THEN THE MAPPING CATHETER WAS REINSERTED TO ACQUIRE A POST-ABLATION VOLTAGE MAP. DURING POST-MAPPING AIR BUBBLES WERE OBSERVED IN THE FLUSH LINE CONNECTED TO THE SHEATH, AND IT WAS FOUND THAT THE HEPARINIZED SALINE BAG CONNECTED TO THE SHEATH FLUSH LINE WAS EMPTY, WHICH CONSTITUTED A USER ERROR. THE LINE TO THE PATIENT WAS PROMPTLY CLOSED. AIR WAS IDENTIFIED IN THE LEFT SIDE OF THE HEART ON ULTRASOUND IMAGING. AT THIS TIME AN ST SEGMENT ELEVATION WAS OBSERVED ON THE V1 LEAD AND THE PATIENT WENT INTO VT. A STROKE ALERT TEAM WAS THEN ACTIVATED. A COMPUTED TOMOGRAPHY (CT) SCAN OF THE BRAIN WAS PERFORMED FINDING NEUROLOGIC BRAIN DAMAGE AND 100CC OF AIR. THE PATIENT WAS SENT TO THE NEUROLOGICAL INTENSIVE CARE UNIT (ICU) AND PLACED IN THE TRENDELENBURG POSITION ON THE BED. ALL ABLATIONS WERE COMPLETED BEFORE ONSET OF THE COMPLICATIONS, SO THE PROCEDURE WAS COMPLETED. TWO DAYS AFTER THE PROCEDURE THE PATIENT DIED IN THE ICU.Event Description: IT WAS REPORTED THAT A USER ERROR OCCURRED WITH IRRIGATION OF THE SHEATH AND THE PATIENT EXPERIENCED AIR EMBOLISM, ST SEGMENT ELEVATION, BRAIN DAMAGE, AND VENTRICULAR TACHYCARDIA (VT). THE PATIENT DIED. DURING A PULSE FIELD ABLATION (PFA) PULMONARY VEIN ISOLATION (PVI) PROCEDURE TO TREAT PAROXYSMAL ATRIAL FIBRILLATION A FARADRIVE SHEATH WAS USED. UPON GAINING TRANSSEPTAL ACCESS 20CC OF BLOOD WAS ASPIRATED FROM THE SHEATH. THE FLUSH LINE WAS CONNECTED TO A PRESSURE BAG AND THEN CONNECTED TO THE SHEATH BEFORE INSERTION OF A NON-BOSTON SCIENTIFIC MAPPING CATHETER. A LEFT ATRIAL VOLTAGE MAP WAS ACQUIRED USING THE MAPPING CATHETER, AND THEN IT WAS EXCHANGED FOR THE FARAWAVE CATHETER. THE CATHETER WAS CONFIRMED TO BE DRIPPING WHEN CONNECTED TO A PRESSURE BAG, AND ANOTHER 20CC WERE ASPIRATED PRIOR TO ADVANCING THE CATHETER INTO THE LEFT ATRIUM. AFTER COMPLETING ABLATION, THE FARAWAVE CATHETER WAS REMOVED AND 20 CC WAS ASPIRATED AGAIN, THEN THE MAPPING CATHETER WAS REINSERTED TO ACQUIRE A POST-ABLATION VOLTAGE MAP. DURING POST-MAPPING AIR BUBBLES WERE OBSERVED IN THE FLUSH LINE CONNECTED TO THE SHEATH, AND IT WAS FOUND THAT THE HEPARINIZED SALINE BAG CONNECTED TO THE SHEATH FLUSH LINE WAS EMPTY, WHICH CONSTITUTED A USER ERROR. THE LINE TO THE PATIENT WAS PROMPTLY CLOSED. AIR WAS IDENTIFIED IN THE LEFT SIDE OF THE HEART ON ULTRASOUND IMAGING. AT THIS TIME AN ST SEGMENT ELEVATION WAS OBSERVED ON THE V1 LEAD AND THE PATIENT WENT INTO VT. A STROKE ALERT TEAM WAS THEN ACTIVATED. A COMPUTED TOMOGRAPHY (CT) SCAN OF THE BRAIN WAS PERFORMED FINDING NEUROLOGIC BRAIN DAMAGE AND 100CC OF AIR. THE PATIENT WAS SENT TO THE NEUROLOGICAL INTENSIVE CARE UNIT (ICU) AND PLACED IN THE TRENDELENBURG POSITION ON THE BED. ALL ABLATIONS WERE COMPLETED BEFORE ONSET OF THE COMPLICATIONS, SO THE PROCEDURE WAS COMPLETED. TWO DAYS AFTER THE PROCEDURE THE PATIENT DIED IN THE ICU. IT WAS FURTHER REPORTED THAT PATIENT DEATH WAS DETERMINED BY NEUROLOGICAL CRITERIA, AND IT IS REGARDED AS THE OFFICIAL CAUSE OF DEATH. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE BECAUSE THE DEVICE AND ITS PACKAGING WERE DISCARDED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. Manufacturer Narrative: ADDITIONAL INFORMATION WAS ADDED TO BLOCK B5 DESCRIBE EVENT OR PROBLEM AND H11 ADDITIONAL MFR NARRATIVE. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION; HOWEVER, THE USER ERROR WAS CONFIRMED BY REFERENCE TO THE DEVICE INSTRUCTIONS FOR USE (IFU). DUE TO THE POST-MAPPING OBSERVATION OF THE EMPTY SALINE BAG CONNECTED TO THE FLUSH LINE, THIS CONSTITUTED AS USER ERROR: THE IFU STATES, "ALWAYS MAINTAIN A CONSTANT STERILE, HEPARINIZED SALINE INFUSION TO PREVENT COAGULATION WITHIN THE SHEATH WHICH MAY RESULT IN EMBOLISM. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE BECAUSE THE DEVICE AND ITS PACKAGING WERE DISCARDED. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22056115&pc=BZD | 2518422-2025-045210 | 2022-07-23 04:00 | Death | RESPIRONICS, INC. | 2025-05-20 04:00 | BZD | REP DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A REP DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22056143&pc=MHX | 9610816-2025-000423 | 2025-04-24 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-05-20 04:00 | MHX | PATIENT INFORMATION CENTER IX | No Audible Alarm | Bradycardia; Asystole | K211900 |  | 1 | Event Description: IT WAS REPORTED THE ALARM DID NOT SOUND AND THE INTENSIVE CARE UNIT (ICU) PATIENT PASSED AWAY. Manufacturer Narrative: A PHILIPS REMOTE SERVICE ENGINEER (RSE) SPOKE WITH THE CUSTOMER. THE EVENT TOOK PLACE ON (B)(6) 2025 AT BED D137. WHILE A SPECIFIC TIME FOR THE FAILURE WAS NOT PROVIDED IN THE ALLEGATION, THE REMOTE INVESTIGATION FOCUSED ON THE TIMELINE OF 13:22 THROUGH THE PATIENT DISCONNECTION FROM THE MONITOR AT 13:58. THE PATIENT WAS BEING MONITORED FOR ECG AND FOR SPO2. PER THE LOG REVIEW, 'DURING THE OBSERVATION PERIOD, A TOTAL OF EIGHT EXTREME BRADYCARDIA RED ALARMS WERE RECORDED, ACCOMPANIED BY SEVERAL YELLOW ALARMS, IN A CONTEXT MARKED BY REPEATED TECHNICAL ALERTS CONCERNING THE SPO2 SENSOR (SIGNAL DIFFICULT TO CAPTURE, AGITATED PATIENT, POORLY POSITIONED SENSOR). THE FIRST RED ALARMS OCCURRED FROM 13:22:23, WITH A SUCCESSION OF CRITICAL EVENTS UNTIL 13:26:02. NO MANUAL ACKNOWLEDGMENT WAS PERFORMED DURING THIS INITIAL SEQUENCE. THE FIRST MANUAL ACKNOWLEDGMENT OF A RED ALARM AT THE CENTRAL UNIT WAS RECORDED AT 13:27:41, FOR AN ALARM GENERATED AROUND 13:26. THE ASYSTOLE RED ALARMS BEGAN AROUND 13:29, WITH A MANUAL ACKNOWLEDGMENT AT THE CENTRAL UNIT 30 SECONDS LATER. OTHER REPETITIONS OF THE ASYSTOLE ALARM WERE OBSERVED, NOTABLY AT 1:32 P.M., 1:35 P.M., 1:39 P.M., 1:42 P.M. AND 1:46 P.M., WITH SUCCESSIVE ACKNOWLEDGMENTS AT THE CENTRAL UNIT BUT ALSO AT THE MONITOR IN THE PATIENT'S ROOM, THE FIRST BEING NOTED AT 1:40:54 AND THE SECOND AT 1:46 P.M. THE SPO2 SENSOR, OFTEN DISCONNECTED OR POORLY POSITIONED, COMPLICATED THE QUALITY OF MONITORING. THE EPISODE ENDED WITH THE PATIENT BEING DISCONNECTED AT 1:58 P.M., PUTTING AN END TO THE ASYSTOLE ALARMS.' THE SPO2 SENSOR, WHICH WAS OFTEN DISCONNECTED OR POORLY POSITIONED, COMPLICATED THE QUALITY OF THE MONITORING. THE EPISODE ENDED WITH THE PATIENT BEING DISCONNECTED AT 13:58, ENDING THE ASYSTOLE ALARMS. A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) AND A PHILIPS CLINICAL PRODUCT SPECIALIST (CPS) REVIEWED THE PROVIDED INFORMATION AND CONFIRMED THE RSE FINDINGS. THE PSE ALSO REVIEWED PIC IX TECHNICAL LOG DATA FROM THE "SURVEILLANCE" HOST. THE PSE NOTED SEVERAL THREE STAR (***) XBRADY ALARMS AND AN ASYSTOLIE SENT TO THE RECORDER. THE TIMESTAMPS CORRELATE WITH THE SOUND MANAGER ENTRIES FROM THE PIC, INDICATING THE RED ALARM SOUND BEING PLAYED AT THE CENTRAL STATION. THE REPORTED ISSUE WAS NOT CONFIRMED. THE CAUSE OF THE DEATH AND THE RELATION TO THE DEVICE REMAINS UNKNOWN DUE TO LIMITED INFORMATION FROM THE CUSTOMER. THE ORIGINAL EVENT DESCRIPTION INDICATES THE PATIENT WAS NOT TREATED IN TIME. BASED ON THIS INFORMATION, THERE DOES NOT APPEAR TO BE A DEVICE MALFUNCTION WHICH CAUSED OR CONTRIBUTED TO THE REPORTED EVENT; HOWEVER, ALARM MANAGEMENT, CLINICAL WORKFLOW, OR OTHER USE ISSUE MAY HAVE BEEN FACTORS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22056229&pc=DSQ | 2916596-2025-02855 | 2021-01-01 05:00 | Death | THORATEC CORPORATION | 2025-05-20 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE ARTICLE 'OUTCOMES IN HEARTMATE 3 RECIPIENTS WITH PRE-EXISTING ATRIAL FIBRILLATION' THAT A RETROSPECTIVE STUDY WAS DONE FROM A SINGLE CENTER FROM 2015 TO 2021 THAT INCLUDED 54 PATIENTS WHO WERE IMPLANTED WITH HEARTMATE 3; 25 PATIENTS HAD NO PRIOR ATRIAL FIBRILLATION (AF) AND 29 HAD AF PRIOR TO IMPLANT. POST-IMPLANT ATRIAL FIBRILLATION (POAF) INCIDENCE WAS 28% IN THE PRE-IMPLANT ATRIAL FIBRILLATION PRAF GROUP AND 31% IN THOSE WITHOUT PRAF (P=0.81). THERE WAS NO DIFFERENCE IN TIME TO POAF ONSET (HR 1.06, P=0.92) OR 5-YEAR SURVIVAL (HR 1.29, P=0.58) BETWEEN GROUPS. 6 PATIENTS WERE TRANSPLANTED, 2 THAT HAD NO PRAF AND 4 THAT HAD PRAF (P=0.5). 22 PATIENTS PASSED AWAY, 8 THAT HAD NO PRAF AND 14 THAT HAD PRAF (P=0.23). Manufacturer Narrative: SECTION E3: OCCUPATION CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A SPECIFIC CAUSE FOR THE PATIENT DEATHS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION, AND A DIRECT CORRELATION WITH THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) COULD NOT BE CONCLUSIVELY ESTABLISHED. NO DEVICES WERE RETURNED FOR THIS EVALUATION. THE DEVICE SERIAL NUMBERS WERE NOT PROVIDED AND WERE UNABLE TO BE DETERMINED THROUGH THIS EVALUATION. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. SECTION A: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. REQUEST FOR THIS INFORMATION HAS BEEN MADE. DEVICE WAS IMPLANTED AT TIME OF EVENT. B3: DATE OF EVENT HAS BEEN ENTERED AS 01JAN2021 AS PATIENTS WERE IMPLANTED BETWEEN 2015 AND 2021. DATE OF DEATH WAS ESTIMATED AS DATE OF EVENT D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. UPPALAPATI, L., ET AL. (2025). OUTCOMES IN HEART MATE 3 RECIPIENTS WITH PRE-EXISTING ATRIAL FIBRILLATION. JOURNAL OF THE AMERICAN COLLEGE OF CARDIOLOGY, 85(12), 1378. HTTPS://DOI.ORG/10.1016/S0735-1097(25)01862-5.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22056364&pc=DSQ | 2916596-2025-02978 | 2025-04-20 04:00 | Death | THORATEC CORPORATION | 2025-05-20 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22056566&pc=DSQ | 2916596-2025-03237 | 2025-04-23 04:00 | Death | THORATEC CORPORATION | 2025-05-20 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR RESPIRATORY FAILURE AND PASSED AWAY THE FOLLOWING DAY. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENTS; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CUSTOMER. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU), REV. A IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A3A: PATIENT GENDER NOT PROVIDED. SECTION D4: EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22056591&pc=NVZ | 2017865-2025-64199 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-20 04:00 | NVZ | ASSURITY SR RF | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P880086 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASE AND CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION, SUSPICION, DOUBT, AND/OR NO INFORMATION SUGGESTING THE DEATH WAS PRODUCT-RELATED. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO THE PATIENT'S DEATH. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. THE RESULTS OF THE ELECTRICAL AND STRESS/ENVIRONMENTAL TESTS WERE PERFORMED TO INDICATE THAT THE PACER IS EXHIBITING NORMAL DEVICE CHARACTERISTICS. A LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22057056&pc=LWS | 2017865-2025-64210 | 2024-12-15 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-05-20 04:00 | LWS | ABBOTT ICD | No Pacing | Arrhythmia; Fatigue |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR PNEUMONIA, HERNIA. THE PATIENT WAS INTUBATED AND LETHARGY AND HYPOXIA WAS NOTED. ARRYTHMIA WAS NOTED BUT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS NOT PACING. TEMPORARY PACEMAKER WAS PLACED. PATIENT HAD BLEEDING WITHIN THE PERITONEAL CAVITY. THE PATIENT PASSED AWAY. THE DEATH WAS NOT BELIEVED TO BE DEVICE RELATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22057193&pc=BZD | 2518422-2025-105905 | 2025-04-22 04:00 | Death | RESPIRONICS, INC. | 2025-05-20 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE HAD FOUND THE DEVICE AFTER THE PATIENT PASSED AWAY A FEW YEARS AGO AND WOULD LIKE IT REPLACED. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22057533&pc=MVK | 3008642652-2025-04980 | 2025-04-15 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-05-20 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing; Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED ELEVEN APPROPRIATE TREATMENTS AND THREE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 20:53:30 ON (B)(6) 2025. AT 01:07:38 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 01:08:14, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 240 BPM. AT 01:08:46, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SINUS BRADYCARDIA @ 50 BPM WITH NSVT. POST SHOCK RHYTHM WAS NSVT. AT 01:09:17, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 01:09:50, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS NSVT. AT 01:10:24, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 240 BPM DEGRADING BACK TO VF WITH CPR/MOTION ARTIFACT. AT 01:22:02, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH MOTION ARTIFACT AND LOW AMPLITUDE CARDIAC SIGNAL. AT 01:22:42, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH MOTION ARTIFACT AND LOW AMPLITUDE CARDIAC SIGNAL. POST SHOCK RHYTHM WAS VT @ 120 BPM. AT 01:23:07, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VF. AT 01:23:37, THE PATIENT RECEIVED THE SEVENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 01:24:00, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST SHOCK RHYTHM WAS VF. AT 01:24:29, THE PATIENT RECEIVED THE EIGHTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 01:24:51, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. AT 01:25:29, THE PATIENT RECEIVED THE NINTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 90 BPM. AT 01:25:51, THE PATIENT RECEIVED THE TENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH LOW CARDIAC SIGNAL. POST SHOCK RHYTHM WAS ASYSTOLE. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 01:26:13, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, TRANSITIONING TO VF WITH LOW CARDIAC SIGNAL. AT 01:26:54, THE PATIENT RECEIVED THE ELEVENTH APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS VF WITH LOW/VARYING CARDIAC SIGNAL. POST SHOCK RHYTHM WAS VF WITH LOW/VARYING CARDIAC SIGNAL. THE ELECTRODE BELT WAS DISCONNECTED AT 01:40:36 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22057953&pc=MKJ | 0003015876-2025-01031 | 2025-04-08 04:00 | Death | PHYSIO-CONTROL, INC. - 3015876 | 2025-05-20 04:00 | MKJ | LIFEPAK 15 DEFIBRILLATOR/MONITOR | Device Sensing Problem | Cardiac Arrest | K142430 |  | 1 | Event Description: THE CUSTOMER CONTACTED STRYKER TO REPORT THAT THEIR DEVICE WAS CONNECTED TO A PATIENT IN CARDIAC ARREST AND FAILED TO FLAG THE PATIENT'S ECG RHYTHM AS VENTRICULAR FIBRILLATION. AS A RESULT, DEFIBRILLATION THERAPY WAS NOT DELIVERED TO THE PATIENT. THE PATIENT IS DECEASED AND THE CUSTOMER IS UNSURE IF THE DEVICE CONTRIBUTED TO THE PATIENT'S OUTCOME. Manufacturer Narrative: STRYKER CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. STRYKER REQUESTS PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. THE CUSTOMER PROVIDED STRYKER WITH ALL THE AVAILABLE PATIENT INFORMATION. PATIENT FIELDS IN WHICH INFORMATION WAS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK. SECTION A1 PATIENT IDENTIFIER CHARACTER LIMIT REACHED. PATIENT IDENTIFIER IS (B)(6) STRYKER CONTINUES TO INVESTIGATE THE REPORTED FAILURE AND WILL SUBMIT A SUPPLEMENTAL REPORT ON THIS EVENT TO THE FDA AS PROVIDED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22042395&pc=MIH | 9612164-2025-02460 | 2025-05-02 04:00 | Death | MEDTRONIC IRELAND | 2025-05-19 04:00 | MIH | VALIANT CAPTIVIA - FF | Leak/Splash | Low Blood Pressure/ Hypotension; Unspecified Infection | P100040 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS A CONCERN ABOUT THE POROSITY OF THE VALIANT CAPTIVIA STENT GRAFTS FOLLOWING THE TREATMENT OF A 64MM CONTAINED RUPTURE TAA . ONE DAY POST THE INDEX PROCEDURE, THE PATIENT BECAME HYPOTENSIVE, PROMPTING A RETURN TO THE OPERATING ROOMFOR RELINING THE IMPLANTED GRAFT WITH A NON MEDTRONIC ((CTAG) STENT GRAFT. THE PHYSICIAN SUSPECTED A TYPE IV ENDOLEAK OR TYPE IIIB ENDOLEAK. THE ENDOLEAK WAS DESCRIBED AS A BLUSHING. THE ENDOLEAK WAS NOTED IN THE AREA OF THE STENT GRAFTS WHICH OVERLAPPED. IT WAS CONFIRMED THERE WAS ADEQUATE OVERLAP. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: IT WAS REPORTED THE PATIENT HAS SINCE PASSED AWAY, THE CAUSE OF DEATH IS UNKNOWN DUE TO NO AUTOPSY BEING PERFORMED. THE PATIENT WAS BELIEVED TO HAVE AN INFECTION IN THE FOLLOW UP SCANS WHICH WAS CONFIRMED WITH BLOOD CULTURES AND INCREASED WHITE CELL COUNTS. B.2 YEAR VALID ONLY MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B2: DATE OF DEATH: MONTH AND YEAR VALID ONLY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: FILM ANALYSIS: THE REPORTED ENDOLEAK WAS CONFIRMED ON THE FILMS PROVIDED; HOWEVER, THE EXACT TYPE AND CAUSE OF THE ENDOLEAK COULD NOT BE DETERMINED. THE REPORTED PATIENT DEATH AND POSSIBLE INFECTION COULD NOT BE ASSESSED. ANGIOGRAMS (INCLUDING ENDOLEAK INTERROGATION) TAKEN DURING THE INTERVENTION PROCEDURE COULD HAVE ALLOWED FOR A MORE COMPREHENSIVE DETERMINATION OF THE ENDOLEAK, BUT THESE WERE NOT PROVIDED FOR REVIEW. ALTHOUGH A TYPE IIIA JUNCTIONAL ENDOLEAK CANNOT BE FULLY RULED OUT THERE APPEARED TO BE SUFFICIENT COMPONENT OVERLAP BETWEEN THE VALIANT CAPTIVA STENT GRAFTS TO PREVENT THIS. A TYPE III FABRIC ENDOLEAK WOULD BE LESS LIKELY TO HAVE OCCURRED AT INDEX AND ANALYSIS OF THE RETURNED FILMS DID NOT REVEAL ANY OBVIOUS ANATOMICAL ANOMALIES THAT MAY HAVE LED TO THIS, E.G CALCIFICATION, BUT ITS POSSIBLE OCCURRENCE CAN NOT BE RULED OUT. IT IS ALSO POSSIBLE THAT THE CONTRAST MAY BE RESIDUAL AFTER SNORKELING. A TYPE IA ENDOLEAK WAS NOT CONSIDERED TO BE PRESENT AS ADEQUATE PROXIMAL SEAL WAS OBSERVED. A TYPE IV ENDOLEAK IS UNLIKELY DUE TO THE DENSITY OF THE CONTRAST BUT ITS POSSIBLE OCCURRENCE ALSO CAN NOT BE RULED OUT. THE REPORTED INTERVENTION WITH STENT GRAFT RELINING WOULD HAVE RESOLVED SEVERAL DIFFERENT ENDOLEAK TYPES. ANALYSIS OF THE RETURNED FILMS DID NOT REVEAL ANY OUT OF SPECIFICATION STENT GRAFT INTEGRITY ISSUES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: BRAND NAME: VALIANT CAPTIVIA FF; PRODUCT ID: VAMF3632C150TU, SERIAL/LOT #: (B)(6), UBD: 12-DEC-2026, UDI#: (B)(4). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22042491&pc=FRN | 2016493-2025-81050 | 2024-04-25 04:00 | Death | CAREFUSION SD | 2025-05-19 04:00 | FRN | ALARIS SYSTEM | Use of Device Problem | Insufficient Information | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS AN OVER INFUSION OF HEPARIN TO A PATIENT ADMITTED WITH UPPER BACK PAIN, WHO "CODED" FOLLOWING ADMISSION AND WAS ADMITTED TO ICU WITH POSTERIOR SEGMENT ELEVATION MYOCARDIAL INFARCTION. THE HEPARIN (HEPARIN 25,000 UNITS/500ML) WAS INTENDED TO INFUSE AT A RATE OF 16ML/HOUR (800 UNITS/HOUR), HOWEVER THE MEDICATION BAG WAS FOUND TO BE EMPTY UNEXPECTEDLY QUICKLY. THE HEPARIN INFUSION WAS STARTED AT 0401, AND THE ENTIRE 500ML BAG WAS NOTED TO BE EMPTY AT 0520. THE PATIENT'S PARTIAL THROMBOPLASTIN TIME WAS 25.3 SECONDS ON THE INITIAL BLOOD DRAW AT 0120, THEN JUMPED TO >169 SECONDS FOLLOWING THE OVER INFUSION. THE PATIENT DID NOT RECEIVE THE ANTIDOTE BECAUSE THEIR BLOOD PRESSURE WAS TOO LOW, ACCORDING TO THE INTENSIVIST. THE PATIENT EXPIRED LATER THAT MORNING. THE CLINICIAN INDICATED THE CAUSE OF DEATH WAS MULTIFACTORIAL. THE CLINICIAN INDICATED THE PATIENT HAD A LARGE POSTERIOR SEGMENT ELEVATION MYOCARDIAL INFARCTION, "CODED", ACHIEVED RETURN OF SPONTANEOUS CIRCULATION, AND WAS ON MULTIPLE VASOPRESSORS. A REPORT WAS RECEIVED FROM HEALTH CANADA'S CANADA VIGILANCE PROGRAM WHICH STATES, "THE PUMP WAS PROGRAMED FOR 800UTS/HR. AND SHOULD HAVE INFUSED AT 16ML/HR. BUT THE PATIENT RECEIVED AN ENTIRE BAG OF HEPARIN IN APPROXIMATELY 1 HOUR. THE CONSEQUENCES OF THAT WERE AN EXTREMELY HIGH PTT. THE PATIENT WAS NOT ABLE TO RECEIVE THE ANTIDOTE DUE TO CLINICAL CONDITION (LOW BP) AS DECIDED BY ICU. THE PATIENT PASSED BUT THE CAUSE OF THIS WAS MULTIFACTORIAL".Event Description: IT WAS REPORTED THAT THERE WAS AN OVER INFUSION OF HEPARIN TO A PATIENT ADMITTED WITH UPPER BACK PAIN, WHO "CODED" FOLLOWING ADMISSION AND WAS ADMITTED TO ICU WITH POSTERIOR SEGMENT ELEVATION MYOCARDIAL INFARCTION. THE HEPARIN (HEPARIN 25,000 UNITS/500ML) WAS INTENDED TO INFUSE AT A RATE OF 16ML/HOUR (800 UNITS/HOUR), HOWEVER THE MEDICATION BAG WAS FOUND TO BE EMPTY UNEXPECTEDLY QUICKLY. THE HEPARIN INFUSION WAS STARTED AT 0401, AND THE ENTIRE 500ML BAG WAS NOTED TO BE EMPTY AT 0520. THE PATIENT'S PARTIAL THROMBOPLASTIN TIME WAS 25.3 SECONDS ON THE INITIAL BLOOD DRAW AT 0120, THEN JUMPED TO >169 SECONDS FOLLOWING THE OVER INFUSION. THE PATIENT DID NOT RECEIVE THE ANTIDOTE BECAUSE THEIR BLOOD PRESSURE WAS TOO LOW, ACCORDING TO THE INTENSIVIST. THE PATIENT EXPIRED LATER THAT MORNING. THE CLINICIAN INDICATED THE CAUSE OF DEATH WAS MULTIFACTORIAL. THE CLINICIAN INDICATED THE PATIENT HAD A LARGE POSTERIOR SEGMENT ELEVATION MYOCARDIAL INFARCTION, "CODED", ACHIEVED RETURN OF SPONTANEOUS CIRCULATION, AND WAS ON MULTIPLE VASOPRESSORS. A REPORT WAS RECEIVED FROM HEALTH CANADA'S CANADA VIGILANCE PROGRAM WHICH STATES, "THE PUMP WAS PROGRAMED FOR 800UTS/HR. AND SHOULD HAVE INFUSED AT 16ML/HR. BUT THE PATIENT RECEIVED AN ENTIRE BAG OF HEPARIN IN APPROXIMATELY 1 HOUR. THE CONSEQUENCES OF THAT WERE AN EXTREMELY HIGH PTT. THE PATIENT WAS NOT ABLE TO RECEIVE THE ANTIDOTE DUE TO CLINICAL CONDITION (LOW BP) AS DECIDED BY ICU. THE PATIENT PASSED BUT THE CAUSE OF THIS WAS MULTIFACTORIAL." Manufacturer Narrative: ADDITIONAL INFORMATION: OTHER RELEVANT HISTORY, DEVICE EVAL BY MANUFACTURER?, IMDRF ANNEX B CODE, REMEDIAL ACTION REQUIRED, REMEDIAL ACTION # AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY: THE REPORT OF AN OVER-INFUSION OF HEPARIN WAS CONFIRMED THROUGH LOG ANALYSIS AND ATTRIBUTED TO A USER PROGRAMMING ERROR, WHICH REVEALED A DISCREPANCY BETWEEN THE REPORTED HEPARIN CONCENTRATION OF 25000U/500 ML AND THE ENTERED CONCENTRATION OF 800 U/500 ML. THE CONCENTRATION ENTRY ERROR RESULTED IN A PUMP CALCULATED RATE OF 500 ML/H WHEN A DOSE OF 800 U/H WAS ENTERED. THE REPORT INDICATED THE DESIRED RATE WAS 16 ML/HOUR. THE EVENT LOG REVIEW COULD NOT CONFIRM THE REPORT THAT ON THE DATE (B)(6) 2024 AT THE TIME OF 0401 AN INFUSION HEPARIN WAS STARTED AND AT THE TIME OF 0520 THE HEPARIN BAG WAS FOUND EMPTY. THE INSPECTION AND TESTING PROCESS DID NOT FIND ANY EXISTING MALFUNCTIONS. THE SUSPECT PUMP MODULE WAS FOUND TO BE INFUSING WITHIN SPECIFICATIONS WITH NO OBSERVANCES OF UNREGULATED FLOW OCCURRING. THE SEAR WAS ALSO FOUND TO BE PROPERLY CLOSING THE ADMINISTRATION SET SAFETY CLAMP WHEN THE DOOR WAS OPENED. AFTER UNREGULATED FLOW TESTING, THE SUSPECT PUMP MODULE WAS PROGRAMMED TO PERFORM A ONE-HOUR SYSTEM LEVEL RATE ACCURACY TEST AT THE INCIDENT INFUSION RATE OF 16ML/HR AND TEST RESULTS MEASURED THE ACTUAL RATE AT 15.43ML/HR (-3.55% ERROR) INDICATING THE DEVICE IS WITHIN SPECIFICATION AFTER RATE CALIBRATION HAVING NO INCIDENTS OF UNREGULATED FLOW OCCURRING DURING THE TESTING. NO ERRORS WERE OBSERVED ON BOTH PCU AND SUSPECT PUMP MODULE ON THE REPORTED DATE OF THE EVENT. REVIEW OF THE SUSPECT PUMP MODULE ERROR LOG SHOWED NO MALFUNCTIONS RELEVANT TO THE FACILITY¿S COMPLAINT. ON THE DATE (B)(6) 2024 AT THE TIME OF 2:25 AM, THE CLINICIAN WAS RECORDED TO SELECT HEPARIN (ACS) (DRUGID=251) FROM GUARDRAIL DRUGS. THE HEPARIN SELECTION WAS A CUSTOM CONCENTRATION WITH THE CLINICIAN ENTERING A DRUG AMOUNT OF 800 UNITS AND A DILUENT VOLUME 500 ML. THE INTENDED CONCENTRATION HAD BEEN 25000U/500 ML. THE ENTRY ERROR RESULTED IN A SOFT GUARDRAIL CONCENTRATION ALERT WHICH THE USER ACCEPTED AND CONTINUED PROGRAMMING THE INFUSION. THE DOSE WAS ENTERED CORRECTLY AT 800 U/HR BUT BECAUSE OF THE CONCENTRATION ERROR THE INFUSION RATE CALCULATED TO 500 ML/H INSTEAD OF THE INTENDED 16 ML/H. THE INFUSION STARTED AT 2:27 AM BUT WAS PAUSED A SHORT TIME LATER BY THE CLINICIAN WITH A PRIMARY VOLUME INFUSED OF 4.171 ML. AT 2:31 AM THE SUSPECT WAS PROGRAMED TO CONTINUE THE INFUSION WITH A VOLUME TO BE INFUSED OF 495.004ML. THE INFUSION FINISHED AT 3:33 AM WITH A REPORTED PRIMARY VOLUME INFUSED OF 511.403ML ML AFTER INFUSING FOR 61 MINUTES. THE ADMINISTRATION SET WAS REQUESTED BUT WAS NOT RETURNED; THEREFORE, IT COULD NOT BE INSPECTED OR TESTED. ROOT CAUSE: THE ROOT CAUSE FOR THE REPORT OF OVER INFUSION OF HEPARIN WAS ABLE TO BE IDENTIFIED FROM THE LOG ANALYSIS AND IS BEING ATTRIBUTED TO USER PROGRAMMING ERROR. THE SUSPECT PUMP MODULE WAS FOUND TO BE INFUSING WITHIN SPECIFICATION, AND THE SEAR WAS FOUND TO BE PROPERLY CLOSING THE SAFETY CLAMP WHEN THE ADMINISTRATION SET IS REMOVED. Manufacturer Narrative: A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE FAILURE INVESTIGATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22042527&pc=QSK | 9612515-2025-00050 | 2025-04-04 04:00 | Death | VASCUTEK LTD. | 2025-05-19 04:00 | QSK | THORAFLEX HYBRID | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Cardiac Arrest; Stroke/CVA | P210006 |  | 1 | Event Description: EVENT DATE: 04 APR 2025. IMPLANT DATE: ON (B)(6) 2025. EVENT REPORTED FROM EXTEND STUDY: (B)(4) (SITE: (B)(4) FOR PATIENT (B)(6). FIRST EVENT OF CARDIAC ARREST ON POST OPERATIVE DAY 0 (ON (B)(6) 2025), SITE REPORTED A SERIOUS ADVERSE EVENT (SAE) OF CARDIAC ARREST - FURTHER COMMENTS DETAIL "PATIENT CODED POST-OPERATIVELY IN THE ICU". SECOND EVENT ON POST OPERATIVE DAY 0 (ON (B)(6) 2025), SITE REPORTED A SERIOUS ADVERSE EVENT (SAE) OF STROKE, ANOXIC BRAIN INJURY - FURTHER COMMENTS DETAIL "BILATERAL, MULTIFOCAL POSTERIOR CIRCULATION STROKES, ANOXIC BRAIN INJURY POST-CARDIAC ARREST". CARDIAC ARREST OCCURRED POST-OPERATIVELY ON (B)(6) 2025, AS DETAILED IN SEPARATE EVENT INTAKE FORM. THIRD EVENT ON POST OPERATIVE DAY 16 (ON (B)(6) 2025), SITE REPORTED A SERIOUS ADVERSE EVENT (SAE) OF CARDIOPULMONARY ARREST - FURTHER COMMENTS DETAIL "CARDIOPULMONARY ARREST DUE TO ANOXIC BRAIN INJURY AND RESPIRATORY INSUFFICIENCY. PATIENT HAD BEEN MOVED TO PALLIATIVE CARE AND MADE DNR (DO NOT RESUSCITATE)". THERE WAS NO DEVICE FAILURE/DEFICIENCY REPORTED. SAE NOTED AS ONGOING AT TIME OF REPORTING, BUT SUBSEQUENT SAE DETAILS SUBJECT OUTCOME OF DEATH ON (B)(6) 2025.Event Description: EVENT DATE: 04 APR 2025. IMPLANT DATE: ON (B)(6) 2025. EVENT REPORTED FROM EXTEND STUDY: (B)(4) (SITE: (B)(4) FOR PATIENT: (B)(6). EMAIL WAS RECEIVED ON 20 MAY 2025 UPDATING THE DATE OF THE SECOND EVENT - STROKE FROM ON (B)(6) 2025. FIRST EVENT OF CARDIAC ARREST ON POST OPERATIVE DAY 0 ON (B)(6) 2025), SITE REPORTED A SERIOUS ADVERSE EVENT (SAE) OF CARDIAC ARREST. FURTHER COMMENTS DETAIL "PATIENT CODED POST-OPERATIVELY IN THE ICU". SECOND EVENT ON POST OPERATIVE DAY 0 ON (B)(6) 2025), SITE REPORTED A SERIOUS ADVERSE EVENT (SAE) OF STROKE, ANOXIC BRAIN INJURY. FURTHER COMMENTS DETAIL "BILATERAL, MULTIFOCAL POSTERIOR CIRCULATION STROKES, ANOXIC BRAIN INJURY POST-CARDIAC ARREST". CARDIAC ARREST OCCURRED POST-OPERATIVELY ON (B)(6) 2025, THIRD EVENT ON POST OPERATIVE DAY 16 ON (B)(6) 2025), SITE REPORTED A SERIOUS ADVERSE EVENT (SAE) OF CARDIOPULMONARY ARREST. FURTHER COMMENTS DETAIL "CARDIOPULMONARY ARREST DUE TO ANOXIC BRAIN INJURY AND RESPIRATORY INSUFFICIENCY. PATIENT HAD BEEN MOVED TO PALLIATIVE CARE AND MADE DNR (DO NOT RESUSCITATE)". THERE WAS NO DEVICE FAILURE/DEFICIENCY REPORTED. SAE NOTED AS ONGOING AT TIME OF REPORTING, BUT SUBSEQUENT SAE DETAILS SUBJECT OUTCOME OF DEATH ON (B)(6) 2025. UPDATED NARRATIVE. ON REVIEW BY THE MEDICAL MONITOR FURTHER EVENTS WERE RECEIVED FOR THIS PATIENT ON (B)(6) 2025 WHICH UPDATED THE NARRATIVE TO THE FOLLOWING: THE PATIENT UNDERWENT A THORAFLEX HYBRID IMPLANT PROCEDURE ON (B)(6) 2025. ON THE SAME DAY, THE PATIENT SUFFERED A CARDIAC ARREST (CLINICAL STUDY SITE INVESTIGATOR CONSIDERS THE SAE AS UNANTICIPATED, POSSIBLY RELATED TO PROCEDURE, POSSIBLY RELATED TO DEVICE, AND POSSIBLY RELATED TO PRE-EXISTING CONDITION. SAE IS CONSIDERED SEVERE), BILATERAL, MULTIFOCAL POSTERIOR CIRCULATION STROKES, ANOXIC BRAIN INJURY (CLINICAL STUDY SITE INVESTIGATOR CONSIDERS THE SAE AS UNANTICIPATED, POSSIBLY RELATED TO PROCEDURE, POSSIBLY RELATED TO DEVICE, AND POSSIBLY RELATED TO PRE-EXISTING CONDITION. SAE IS CONSIDERED SEVERE) POST-CARDIAC ARREST AND REQUIRED PROLONGED MECHANICAL VENTILATION (CLINICAL STUDY SITE INVESTIGATOR CONSIDERED THE EVENT OF PROLONGED MECHANICAL VENTILATION POSSIBLY RELATED TO THE DEVICE AND A PRE-EXISTING CONDITION, RELATED TO THE PROCEDURE, ANTICIPATED, NOT A SERIOUS ADVERSE EVENT BUT SEVERE IN SEVERITY). ON (B)(6) 2025, THE PATIENT DEVELOPED ATRIAL FIBRILLATION (CLINICAL STUDY SITE INVESTIGATOR CONSIDERS THE AE AS ANTICIPATED, POSSIBLY RELATED TO PROCEDURE, POSSIBLY RELATED TO DEVICE, NOT RELATED TO A DEVICE DEFICIENCY AND POSSIBLY RELATED TO PRE-EXISTING CONDITION. AE IS CONSIDERED MILD IN SEVERITY) AND BY ON (B)(6) 2025, THE PATIENT EXPERIENCED ACUTE KIDNEY INJURY (CLINICAL STUDY SITE INVESTIGATOR CONSIDERS THE AE AS ANTICIPATED, POSSIBLY RELATED TO PROCEDURE, POSSIBLY RELATED TO DEVICE, NOT RELATED TO A DEVICE DEFICIENCY AND POSSIBLY RELATED TO PRE-EXISTING CONDITION. AE IS CONSIDERED MODERATE IN SEVERITY) WHICH WAS UPDATED ON 28 MAY 25 TO ONGOING AT THE TIME OF DEATH. INTAKE FORMS STATE ALL OF THE EVENTS REPORTED WERE NOT RELATED TO A DEVICE FAILURE /DEFICIENCY. THIS REPORT IS BEING SUBMITTED AS FOLLOW UP #1 FOR MANUFACTURING REPORT NUMBER 9612515-2025-00050 TO PROVIDE EVENT CLOSURE INFORMATION FOR TAG0229. Manufacturer Narrative: MANUFACTURER NARRATIVE: CLINICAL CODE: 1762 -CARDIAC ARREST - FIRST EVENT OF CARDIAC ARREST ON POST OPERATIVE DAY 0 ((B)(6) 2025), SITE REPORTED A SERIOUS ADVERSE EVENT (SAE) OF CARDIAC ARREST - FURTHER COMMENTS DETAIL "PATIENT CODED POST-OPERATIVELY IN THE ICU". 1770 - STROKE -SECOND EVENT ON POST OPERATIVE DAY 0 ((B)(6) 2025), SITE REPORTED A SERIOUS ADVERSE EVENT (SAE) OF STROKE, ANOXIC BRAIN INJURY - FURTHER COMMENTS DETAIL "BILATERAL, MULTIFOCAL POSTERIOR CIRCULATION STROKES, ANOXIC BRAIN INJURY POST-CARDIAC ARREST". CARDIAC ARREST OCCURRED POST-OPERATIVELY ON (B)(6) 2025, AS DETAILED IN SEPARATE EVENT INTAKE FORM. THIRD EVENT ON POST OPERATIVE DAY 16 ((B)(6) 2025), SITE REPORTED A SERIOUS ADVERSE EVENT (SAE) OF CARDIOPULMONARY ARREST - FURTHER COMMENTS DETAIL "CARDIOPULMONARY ARREST DUE TO ANOXIC BRAIN INJURY AND RESPIRATORY INSUFFICIENCY. PATIENT HAD BEEN MOVED TO PALLIATIVE CARE AND MADE DNR (DO NOT RESUSCITATE)". 4882 - KIDNEY INJURY ADDITIONAL INFORMATION RECEIVED ON THIS EVENT REPORTED - ACUTE KIDNEY INJURY ALONG WITH ATRIAL FIBRILLATION REPORTED FOR THIS PATIENT. IMPACT CODE: 1802 - DEATH - PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO CARDIOPULMONARY ARREST. DEVICE PROBLEM CODE: 2993 - ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM- NO DEVICE FAILURE /DEFICIENCY HAS BEEN REPORTED FULL BATCH REVIEW PERFORMED FROM BASE FABRIC TO FINISHED PRODUCT WAS PERFORMED WHICH CONFIRMED THE DEVICE WAS MANUFACTURED TO SPECIFICATION. WITH NO ISSUES FOUND COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4110 - TREND ANALYSIS: A 4-YEAR REVIEW OF SIMILAR COMPLAINTS THORAFLEX HYBRID SALES JAN 21 - APR 25 VS MEDICAL EVENT > STROKE ON (B)(6) GAVE AN OCCURRENCE RATE OF (B)(4) NO TREND REQUIRING ACTION HAS BEEN IDENTIFIED. 4111 - COMMUNICATION INTERVIEW: ADDITIONAL INFORMATION WAS RECEIVED -THE PATIENT UNDERWENT ELECTIVE AORTIC SURGERY INVOLVING THE DEPLOYMENT OF A THORAFLEX¿ HYBRID DEVICE. THE DISTAL LANDING ZONE OF THE IMPLANT WAS MEASURED AT 32 MM IN DIAMETER, AND NO OVERSIZING WAS APPLIED DURING THE PROCEDURE. THIS MEASUREMENT WAS BASED ON THE ORIGINAL VESSEL DIAMETER, AS NO DISSECTION WAS PRESENT IN THE REGION OF THE DISTAL LANDING ZONE. THE ANATOMY OF THE AORTA AT THE DISTAL LANDING ZONE WAS UNREMARKABLE, WITH NO SIGNIFICANT CURVATURE OR OTHER ANATOMICAL FEATURES OF NOTE THAT COULD HAVE IMPACTED DEVICE DEPLOYMENT START TIME OF PROCEDURE 09.00, END TIME OF PROCEDURE 18.44. DURATION OF PROCEDURE 584 MINUTES. 3331 - ANALYSIS OF PRODUCTION RECORDS: FULL BATCH REVIEW PERFORMED FROM BASE FABRIC TO FINISHED PRODUCT FOR DEVICE ID: 25882308 WHICH CONFIRMED THE BATCH WAS MANUFACTURED TO SPECIFICATION. 4117 - DEVICE NOT RETURNED: DEVICE REMAINS IMPLANTED. INVESTIGATION FINDINGS: 213 - NO DEVICE PROBLEM FOUND: DEVICE WAS MANUFACTURED TO SPECIFICATION AND NO DEVICE DEFICIENCY WAS REPORTED. INVESTIGATION CONCLUSION: 4315 - CAUSE NOT ESTABLISHED: NO ROOT CAUSE WAS IDENTIFIED. ADDITIONAL INFORMATION. SECTION B5: UPDATED NARRATIVE TO INCLUDE ADDITIONAL INFORMATION ON THIS EVENT. Manufacturer Narrative: MANUFACTURER NARRATIVE: CLINICAL CODE: 1762 -CARDIAC ARREST - FIRST EVENT OF CARDIAC ARREST ON POST OPERATIVE DAY 0 (04-APR-2025), SITE REPORTED A SERIOUS ADVERSE EVENT (SAE) OF CARDIAC ARREST - FURTHER COMMENTS DETAIL "PATIENT CODED POST-OPERATIVELY IN THE ICU". 1770 - STROKE -SECOND EVENT ON POST OPERATIVE DAY 0 (04-APR-2025), SITE REPORTED A SERIOUS ADVERSE EVENT (SAE) OF STROKE, ANOXIC BRAIN INJURY - FURTHER COMMENTS DETAIL "BILATERAL, MULTIFOCAL POSTERIOR CIRCULATION STROKES, ANOXIC BRAIN INJURY POST-CARDIAC ARREST". CARDIAC ARREST OCCURRED POST-OPERATIVELY ON (B)(6) 2025, AS DETAILED IN SEPARATE EVENT INTAKE FORM. THIRD EVENT ON POST OPERATIVE DAY 16 (20-APR-2025), SITE REPORTED A SERIOUS ADVERSE EVENT (SAE) OF CARDIOPULMONARY ARREST - FURTHER COMMENTS DETAIL "CARDIOPULMONARY ARREST DUE TO ANOXIC BRAIN INJURY AND RESPIRATORY INSUFFICIENCY. PATIENT HAD BEEN MOVED TO PALLIATIVE CARE AND MADE DNR (DO NOT RESUSCITATE)". IMPACT CODE: 1802 - DEATH - PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO CARDIOPULMONARY ARREST. DEVICE PROBLEM CODE: 3190 - INSUFFICIENT INFORMATION - NO DEVICE FAILURE /DEFICIENCY HAS BEEN REPORTED. COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4110 - TREND ANALYSIS: A 4-YEAR REVIEW OF SIMILAR COMPLAINTS THORAFLEX HYBRID SALES JAN 21 - APR 25 VS MEDICAL EVENT > STROKE JAN 21 - APR 25 GAVE AN OCCURRENCE RATE OF (B)(4) NO TREND REQUIRING ACTION HAS BEEN IDENTIFIED. 4111 - COMMUNICATION INTERVIEW: ADDITIONAL INFORMATION. 3331 - ANALYSIS OF PRODUCTION RECORDS: FULL BATCH REVIEW PERFORMED FROM BASE FABRIC TO FINISHED PRODUCT FOR DEVICE ID- 25882308 WHICH CONFIRMED THE BATCH WAS MANUFACTURED TO SPECIFICATION. 4117 - DEVICE NOT RETURNED: DEVICE REMAINS IMPLANTED. INVESTIGATION FINDINGS: 3233 - RESULTS PENDING COMPLETION OF INVESTIGATION. INVESTIGATION CONCLUSION: 11 - CONCLUSION NOT YET AVAILABLE AS INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22043022&pc=LJT | 3006260740-2025-03206 | 2025-04-28 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-05-19 04:00 | LJT | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Pain; Thrombosis/Thrombus; Insufficient Information | K072549 |  | 1 | Event Description: IT WAS REPORTED THROUGH LITIGATION PROCESS THAT SOMETIME LATER PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY DEVELOPED WITH PULMONARY EMBOLISM, BLOOD CLOT AND OTHER SEVERE PAIN WITH ONGOING MEDICAL TREATMENT. LATER THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS OF (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22043040&pc=MKJ | 0003015876-2025-01009 | 2025-04-18 04:00 | Death | PHYSIO-CONTROL, INC. - 3015876 | 2025-05-19 04:00 | MKJ | LIFEPAK 35 MONITOR/DEFIBRILLATOR | Failure to Sense | No Clinical Signs, Symptoms or Conditions | P160026 |  | 1 | Event Description: A CUSTOMER CONTACTED STRYKER TO REPORT THAT THEIR DEVICE INDICATED NO PADS CONNECTED, WHILE BEING ATTACHED TO THE PATIENT. A BACK UP SET OF PADS WERE AVAILABLE AND WERE USED, HOWEVER THE ISSUE PERSISTED. THE CUSTOMER ADVISED THAT ANOTHER DEVICE WAS AVAILABLE AND WAS USED WITH THE SAME PADS, AND THEY WERE ABLE TO SUCCESSFULLY DEFIBRILLATE THE PATIENT. THE PATIENT ASSOCIATED WITH THE REPORTED EVENT DID NOT SURVIVE. Manufacturer Narrative: STRYKER CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. STRYKER REQUESTS PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. THE CUSTOMER PROVIDED STRYKER WITH THE AVAILABLE PATIENT INFORMATION. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK. STRYKER PERFORMED A CLINICAL REVIEW AND IT WAS DETERMINED THAT THE DEVICE USE MAY HAVE CONTRIBUTED TO THE PATIENT OUTCOME STRYKER CONTINUES TO INVESTIGATE THE REPORTED FAILURE AND WILL SUBMIT A SUPPLEMENTAL REPORT ON THIS EVENT TO THE FDA AS PROVIDED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22043067&pc=DSQ | 2916596-2025-03327 | 2025-05-04 04:00 | Death | THORATEC CORPORATION | 2025-05-19 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE ANY ADDITIONAL INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. G IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22043428&pc=DSQ | 2916596-2025-02956 | 2025-04-10 04:00 | Death | THORATEC CORPORATION | 2025-05-19 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ALTHOUGH THE CAUSE OF DEATH WAS REPORTED AS UNKNOWN, THE PATIENT'S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE CUSTOMER. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION", LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22043661&pc=OZD | 1220648-2025-28514 | 2025-04-26 04:00 | Death | ABIOMED, INC. | 2025-05-19 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ON IMPELLA 5.5 DEVELOPED GASTROINTESTINAL BLEEDING. SYSTEMIC HEPARIN WAS HELD. IT WAS FURTHER REPORTED THAT THE FAMILY DECIDED TO MAKE PATIENT COMFORT CARE AND THE PATIENT EXPIRED. PER MEDICAL SAFETY OFFICER REVIEW THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22043962&pc=NPT | 2025587-2025-03754 | 2025-03-15 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-19 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Septic Shock | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT'S PROGRESS WAS NOTED TO BE GOOD. THE PATIENT WAS DISCHARGED HOME, HOWEVER, THE PATIENT'S FAMILY REQUESTED CARDIAC REHABILITATION. 7 DAYS AFTER THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO THE HOSPITAL FOR THE CARDIAC REHABILITATION. THE FOLLOWING DAY, NO NOTABLE ABNORMALITIES WERE OBSERVED ON BLOOD TESTS OR AN ECHOCARDIOGRAM. AROUND NOON OF THE NEXT DAY, THE PATIENT DEVELOPED A FEVER OF 39 DEGREES CELSIUS. AT NIGHT PULSELESS ELECTRICAL ACTIVITY OCCURRED AND ADVANCED CARDIOVASCULAR LIFE SUPPORT WAS PERFORMED. THE PATIENT COULD NOT BE RESUSCITATED. THE CAUSE OF DEATH WAS NOTED TO BE SEPTIC SHOCK. Manufacturer Narrative: UPDATED DATA: D4., H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22043974&pc=NPU | 2015691-2025-03931 | 2025-03-09 05:00 | Death | EDWARDS LIFESCIENCES | 2025-05-19 04:00 | NPU | US S3 UR VALVE 29MM | Perivalvular Leak | Hemolysis | P140031 |  | 1 | Event Description: AS REPORTED BY IMPLANT PATIENT REGISTRY, A PATIENT UNDERWENT A TMVR PROCEDURE WITH A 29MM SAPIEN 3 ULTRA RESILIA VALVE IN THE NATIVE MITRAL POSITION. ON POST OP DAY 2, THE PATIENT PASSED AWAY AND CAUSE OF DEATH WAS REPORTED AS CARDIOGENIC SHOCK. AFTER A REVIEW OF THE MEDICAL RECORDS, THE PATIENT UNDERWENT TMVR AND TAVR WITH MULTIPLE PVL CLOSURES. THERE REMAINED 'SIGNIFICANT PVL' WITHOUT GOOD OPTIONS FOR REPAIR. ON POST OP DAY 1, TTE REVEALED MEAN MITRAL GRADIENT OF 5MMHG WITH SEVERE PVL, LVOT GRADIENT WAS 94MMHG. TAVR VALVE GRADIENTS WERE 4 TO 8MMHG. THE PATIENT WAS BEING SUPPORTED IN ICU WITH MULTIPLE VASOPRESSORS AND COULD NOT BE WEANED OFF. ADDITIONAL TMVR PROCEDURES WERE NOT POSSIBLE DUE TO SMALL LV SIZE AND RISK FOR LVOT OBSTRUCTION. ADDITIONALLY, THE PATIENT BEGAN TO DEVELOP HEMOLYSIS ATTRIBUTED TO THE PVL. THE PATIENT WAS ON CRRT AND CONTINUED TO DETERIORATE. CODE BLUE WAS CALLED AND THE FAMILY WAS CALLED TO THE BEDSIDE. THE PATIENT WAS ALLOWED TO DIE A NATURAL DEATH AND EXPIRED ON POD#2. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PARAVALVULAR LEAK (PVL) IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE AND THE LANDING ZONE, AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE. SOME PVL IS NOT UNCOMMON POST-DEPLOYMENT. MANY CASES ARE MILD TO MODERATE, AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS. OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION. THE PATIENT SCREENING MANUAL AND THE PROCEDURE DIDACTIC IDENTIFY SEVERAL PROCEDURAL AND ANATOMICAL FACTORS WHICH COULD CONTRIBUTE TO PVL, INCLUDING DEVICE MALPOSITION, INACCURATE MEASUREMENT OF THE LANDING ZONE, UNEVEN DISTRIBUTION OF CALCIUM ON THE LANDING ZONE, BULKY OR SEVERE CALCIFICATION, AN ELLIPTICAL ANNULUS SHAPE, AND VALVE UNDER-SIZING. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. HEMOLYSIS IS THE DESTRUCTION OF RED BLOOD CELLS. THERE ARE MULTIPLE EXTRINSIC AND INTRINSIC FACTORS INCLUDING USE OF EXTRACORPOREAL CIRCULATION, TRANSFUSION, INFECTION, TUMORS, AUTOIMMUNE DISORDERS, MEDICATION SIDE EFFECTS, A METABOLIC ABNORMALITY, AND RED BLOOD CELL MEMBRANE INSTABILITY. NORMAL RED BLOOD CELLS (ERYTHROCYTES) HAVE A LIFESPAN OF ABOUT 120 DAYS. AFTER THE LIFESPAN IS OVER THE RED BLOOD CELLS BREAK DOWN AND ARE REMOVED FROM THE CIRCULATION BY THE SPLEEN. IN SOME MEDICAL CONDITIONS, OR BECAUSE OF TAKING CERTAIN MEDICATIONS, THIS BREAKDOWN OF RED BLOOD CELLS IS INCREASED. MEDICATIONS THAT HAVE BEEN ASSOCIATED WITH HEMOLYSIS INCLUDE ACETAMINOPHEN, PENICILLIN, AND OTHER PAIN MEDICATIONS. RED CELLS MAY BE DESTROYED DUE TO DEFECTS IN THE CELLS THEMSELVES. ANOTHER CAUSE OF HEMOLYTIC ANEMIA IS BREAK DOWN DUE TO MECHANICAL DAMAGE, SUCH AS FROM ARTIFICIAL HEART VALVES OR HEART-LUNG BYPASS; OR THEY AS A RESULT OF TURBULENT BLOOD FLOW PATTERN AND ERYTHROCYTE DESTRUCTION CAUSED BY PVL OR CENTRAL REGURGITATION. THE DEVICE WAS USED IN OFF-LABEL IMPLANTATION IN THE NATIVE MITRAL POSITION. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO NATIVE MITRAL PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). THE DEVICE TRAINING MANUALS INSTRUCTS THE OPERATOR ON PROPER IMAGING SCREENING REQUIREMENTS, INCLUDING THE USE OF GOOD QUALITY ECHOCARDIOGRAPHY AND/OR COMPUTED TOMOGRAPHY (CT) TO APPROPRIATELY MEASURE THE LANDING ZONE, ASSESS THE CONTENT AND DISTRIBUTION OF CALCIUM, AND OTHER PATIENT ANATOMICAL FACTORS. MULTIPLE IMAGING MODALITIES SHOULD BE CONSIDERED DURING VALVE SIZE SELECTION. CONTRAINDICATIONS, IMPORTANT CONSIDERATIONS WHEN ASSESSING THE VALVE, AND CHOOSING THE PROPER TRANSCATHETER HEART VALVE ARE ALSO DISCUSSED. THE DEVICE TRAINING MANUALS ALSO INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING PATIENTS SCREENING AND PROCEDURAL CONSIDERATIONS. DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES ARE ALSO INCLUDED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST PATIENT FACTORS (MAC, CAD, AND DM2) AND/OR THE MECHANISMS DESCRIBED ABOVE MAY HAVE CAUSED OR CONTRIBUTED TO THE PVL WITH SUBSEQUENT HEMOLYSIS/HEMOLYTIC ANEMIA. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22044025&pc=GKT | 3004548776-2025-00257 | 2024-06-19 04:00 | Death | FENWAL INTERNATIONAL INC. | 2025-05-19 04:00 | GKT | FENWAL PLASMACELL-C DISPOSABLE SET FOR USE WITH SPIKESMART SYSTEM | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | BK110072 |  | 1 | Event Description: ON (B)(6) 2025 BIOLIFE PLASMA SERVICES RECEIVED TELEPHONE NOTIFICATION OF A DONOR DEATH. THE CHIEF MEDICAL EXAMINER IN (B)(6) NOTED THE DONOR HAD DIED AT HIS HOME ON (B)(6) 2024. THE CHIEF MEDICAL EXAMINER RELATED THAT NO CAUSE OF DEATH WAS KNOWN AND INITIATED A REQUEST FOR BIOLIFE RECORDS. A REVIEW OF DONOR RECORDS SHOWS THE DONOR LAST DONATED ON (B)(6) 2024. THE BIOLIFE CENTER PHYSICIAN CONTACTED THE MEDICAL EXAMINER'S OFFICE ON (B)(6) 2025 AND WAS INFORMED THAT WHILE THE CAUSE OF DEATH IS NOT KNOWN THAT A HISTORY OF NAUSEA, VOMITING AND CEPHALGIA WERE NOTED AFTER PLASMAPHERESIS AND THAT UNSPECIFIED HYPONATREMIA WAS ALSO NOTED. UPON NOTIFICATION OF THE DONOR'S DEATH, THE PLASMAPHERESIS DEVICE WAS TAKEN OUT OF SERVICE. THE PLASMAPHERESIS DEVICE LOGS AND SINGLE USE COLLECTIONS MATERIALS RECORDS WERE INSPECTED AND NO ABERRANT FINDINGS WERE NOTED. THE PREVENTATIVE MAINTENANCE TPDS WAS PERFORMED ON (B)(6) 2025 AND FOUND THE INSTRUMENT TO BE WORKING AS INTENDED. IN CONCLUSION, THE DEATH OF THE DONOR, OCCURRING SHORTLY AFTER A PLASMA DONATION AT BIOLIFE PLASMA SERVICES, THE REPORTED SYMPTOMS OF NAUSEA, VOMITING, CEPHALGIA, AND UNSPECIFIED HYPONATREMIA NOTED BY THE MEDICAL EXAMINER. HOWEVER, A DEFINITIVE CAUSAL RELATIONSHIP BETWEEN THE PLASMAPHERESIS PROCEDURE AND THE DONOR'S DEATH CANNOT BE DETERMINED AT THIS TIME DUE TO THE LACK OF DETAILED AUTOPSY RECORDS. THE INVESTIGATION OF THE DONATION PROCEDURE AND EQUIPMENT REVEALED NO IMMEDIATE ANOMALIES. PER BIOLIFE, THE DONOR DID NOT RELATE A HISTORY OF PRIOR ELECTROLYTE ABNORMALITIES BUT DID NOTE A RESOLVED HISTORY OF KIDNEY STONES. AS NO FURTHER MEDICAL OR AUTOPSY RECORDS ARE CURRENTLY AVAILABLE, NO CAUSAL RELATION BETWEEN THE PLASMAPHERESIS PROCEDURE AND THE FATALITY CAN BE DETERMINED. FURTHER INVESTIGATION BY FRESENIUS KABI IS PENDING. GIVEN THAT THE DONOR DIED ONE DAY AFTER THE FINAL UNCOMPLICATED DONATION AND THE CAUSE OF THE DEATH IS CURRENTLY UNKNOWN, WE ARE REPORTING THIS INCIDENT CONSERVATIVELY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22044277&pc=OZD | 1220648-2025-28519 | 2025-04-27 04:00 | Death | ABIOMED, INC. | 2025-05-19 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate; Deformation Due to Compressive Stress | Stroke/CVA; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE AXILLARY SHEATH FOR 5.5 WAS DAMAGED OUT OF THE PACKAGE. ONCE PLACED INTO THE 10MM PLATINUM GRAFT AND TWO GRAFT LOCKS PLACED THERE WAS SIGNIFICANT BLEEDING FROM PERFORATION POINTS DOWN THE SIDES OF THE HUB. NEW AXILLARY SHEATH PLACED, ISSUE RESOLVED. THE PATIENT RECEIVED 3 UNITS PRBCS. ADDITIONALLY, THE PATIENTS' OUTPUTS DROPPED OVERNIGHT AND SUCTION ALARM OCCURRED ON P8. LV SEEMED TO SUCKED DOWN. DOWN TO P7 AND FLOLAN STARTED TO HELP RIGHT TO LEFT FLOW. PATIENT RECOVERED AND WAS EVENTUALLY WEANED OFF LEVO. THE PATIENT CONTINUES ON SUPPORT. ADDITIONAL INFORMATION RECEIVED INDICATING THAT THE PATIENT HAD A RIGHT MCA STROKE. CARE WAS ULTIMATLY WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF INTRODUCER DAMAGE ¿ MECHANICAL, ACCESS SITE BLEEDING AND LOW PUMP FLOW HAS BEEN COMPLETED. INTRODUCER DAMAGE - MECHANICAL: DURING IMPLANTATION PROCEDURE, AXILLARY SHEATH FOR 5.5 WAS DAMAGED OUT OF THE PACKAGE. ONCE PLACED INTO THE 10MM PLATINUM GRAFT AND TWO GRAFT LOCKS PLACED THERE WAS SIGNIFICANT BLEEDING FROM PERFORATION POINTS DOWN THE SIDES OF THE HUB. NEW AXILLARY SHEATH PLACED, ISSUE RESOLVED. THE AXILLARY SHEATH WAS RETURNED AND INSPECTED. VISUAL INSPECTION REVEALED DRIED BLOOD AT THE HEMOSTATIC VALVE OF THE SHEATH WHERE THE TWO HALVES WOULD PEEL AWAY. THE SHEATH WAS LEAK TESTED WHERE FLUID WOULD FLOW OUT OF THE VALVE AT 105.6MMHG, FAR BELOW THE 180MMHG REQUIRED OF THE VALVE. THE CAUSE OF THE INTRODUCER DAMAGE - MECHANICAL WAS A DAMAGED VALVE DUE TO BLEEDING FROM THE SIDE OF THE INTRODUCER HUB AND LEAKING OF THE INTRODUCER REPRODUCING UNDER 180MMHG OF PRESSURE. ACCESS SITE BLEEDING - MAJOR: AFTER THERE WAS SIGNIFICANT BLEEDING FROM PERFORATION POINTS DOWN THE SIDES OF THE HUB. NEW AXILLARY SHEATH PLACED, ISSUE RESOLVED. THE AXILLARY SHEATH WAS RETURNED AND INSPECTED. VISUAL INSPECTION REVEALED DRIED BLOOD AT THE HEMOSTATIC VALVE OF THE SHEATH WHERE THE TWO HALVES WOULD PEEL AWAY. THE SHEATH WAS LEAK TESTED WHERE FLUID WOULD FLOW OUT OF THE VALVE AT 105.6MMHG, FAR BELOW THE 180MMHG REQUIRED OF THE VALVE. THE CAUSE OF THE ACCESS SITE BLEEDING - MAJOR WAS A DAMAGED VALVE DUE TO BLEEDING FROM THE SIDE OF THE INTRODUCER HUB AND LEAKING OF THE INTRODUCER REPRODUCING UNDER 180MMHG OF PRESSURE. LOW PUMP FLOW: AFTER 1 DAY ON SUPPORT, PT'S OUTPUT DROPPED CAUSING SUCTION ALARMS ON P-8. DOWN TO P7 AND FLOW STARTED TO HELP RIGHT TO LEFT FLOW. EF AROUND 20-25%. PT ON VENT. PUMP WAS NOT RETURNED. DATA LOG REVIEW SHOWED TO SEPARATED SETS OF SUCTION EVENTS. BOTH SUCTION EVENTS RESOLVED BY REDUCING THE P-LEVEL FROM P-8 TO P-7, INDICATING THE SUCTION EVENTS WERE TRIGGERED BY A LACK OF BLOOD VOLUME IN THE LV. THE CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY PATIENT CONDITION RELATED DUE TO THE SUCTION EVENTS RESOLVING WHEN THE P-LEVEL WAS REDUCED. H10 NARRATIVE ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-28519 STATED THE DEVICE WAS NOT GOING TO BE RECEIVED, HOWEVER THE DEVICE HAD BEEN RECEIVED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22044293&pc=NIV | 1028232-2025-02433 | 2025-04-30 04:00 | Death | BIOTRONIK AG, BUELACH, SWITZERLAND | 2025-05-19 04:00 | NIV | PK PAPYRUS US 3.0/20 | Patient Device Interaction Problem | Hypoxia; Irregular Pulse; Vascular Dissection | H170004 |  | 1 | Event Description: PK PAPYRUS COVERED STENT SYSTEMS WERE SELECTED FOR TREATMENT OF A TYPE III CS/IV PERFORATION (15 MM LENGTH) OCCURRED DURING POST-DILATATION OF A DES. AFTER THE FIRST PK PAPYRUS WAS PLACED, THE PERFORATION WAS NOT COMPLETELY SEALED (REPORTED SEPARATELY), SO A SECOND PK PAPYRUS WAS DEPLOYED (3.0/20, COMPLAINT DEVICE). A POSSIBLE EDGE DISSECTION AT THE EDGE OF THE SECOND STENT SUSPECTED, SO AN ONYX FRONTIER WAS DEPLOYED. AFTERWARDS, AN IMPELLA WAS PLACED AND THE PATIENT LEFT CATH LAB TO ICU. PATIENT CONTINUED TO HAVE SEVERE HEMODYNAMIC COMPROMISE DESPITE INCREASING MECHANICAL SUPPORT. FAMILY ULTIMATELY DECIDED TO WITHDRAW CARE. Manufacturer Narrative: THE AFFECTED PK PAPYRUS STENT SYSTEM WAS NOT RETURNED TO BIOTRONIK AND COULD THEREFORE NOT BE SUBJECTED TO A TECHNICAL INVESTIGATION. IMAGES OF THE CASE SUCH AS ANGIOGRAPHIES COULD ALSO NOT BE OBTAINED. THE PROVIDED CLINICAL INFORMATION (PK PAPYRUS DEVICE REGISTRATION FORM, CATHLAB REPORT) AND THE PRODUCTION DOCUMENTATION WERE REVIEWED TO ESTABLISH WHETHER A DEVICE DEFICIENCY CONTRIBUTED TO THE EVENT. REVIEW OF THE PRODUCTION DOCUMENTATION CONFIRMED THAT THE DEVICE WAS MANUFACTURED ACCORDING TO SPECIFICATIONS AND FULFILLED ALL THE REQUIREMENTS OF IN-PROCESS AND FINAL INSPECTION. BASED ON THE PROVIDED DOCUMENTATION AND THE CONDUCTED REVIEW, NO MANUFACTURING RELATED ROOT CAUSE COULD BE IDENTIFIED. THE TREATING PHYSICIAN RATED THE OUTCOME AS NO DEVICE- BUT PROCEDURE-RELATED COMPLICATION. PK PAPYRUS IS INDICATED FOR THE TREATMENT OF ACUTE PERFORATIONS OF NATIVE CORONARY ARTERIES AND CORONARY BYPASS GRAFTS IN VESSELS 2.5 TO 5.0 MM IN DIAMETER. AS SUCH, THIS DEVICE IS TYPICALLY USED IN EMERGENT RESPONSE TO A LIFE-THREATENING ADVERSE PROCEDURAL EVENT UNRELATED TO USE OF THE PK PAPYRUS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22044301&pc=NIV | 1028232-2025-02437 | 2025-04-30 04:00 | Death | BIOTRONIK AG, BUELACH, SWITZERLAND | 2025-05-19 04:00 | NIV | PK PAPYRUS US 3.0/15 | Lack of Effect | Cardiac Tamponade | H170004 |  | 1 | Event Description: PK PAPYRUS COVERED STENT SYSTEMS WERE SELECTED FOR TREATMENT OF A TYPE III CS/IV PERFORATION (15 MM LENGTH) OCCURRED DURING POST-DILATATION OF A DES. AFTER THE FIRST PK PAPYRUS (COMPLAINT DEVICE) WAS PLACED, THE PERFORATION WAS NOT COMPLETELY SEALED SO A SECOND PK PAPYRUS WAS DEPLOYED. A POSSIBLE EDGE DISSECTION AT THE EDGE OF THE SECOND STENT SUSPECTED (REPORTED SEPARATELY), SO AN ONYX FRONTIER WAS DEPLOYED. AFTERWARDS, AN IMPELLA WAS PLACED AND THE PATIENT LEFT CATH LAB TO ICU. PATIENT CONTINUED TO HAVE SEVERE HEMODYNAMIC COMPROMISE DESPITE INCREASING MECHANICAL SUPPORT. FAMILY ULTIMATELY DECIDED TO WITHDRAW CARE. ADDITIONAL PATIENT CODE E0504. Manufacturer Narrative: THE AFFECTED PK PAPYRUS STENT SYSTEM WAS NOT RETURNED TO BIOTRONIK AND COULD THEREFORE NOT BE SUBJECTED TO A TECHNICAL INVESTIGATION. IMAGES OF THE CASE SUCH AS ANGIOGRAPHIES COULD ALSO NOT BE OBTAINED. THE PROVIDED CLINICAL INFORMATION (PK PAPYRUS DEVICE REGISTRATION FORM, CATHLAB REPORT) AND THE PRODUCTION DOCUMENTATION WERE REVIEWED TO ESTABLISH WHETHER A DEVICE DEFICIENCY CONTRIBUTED TO THE EVENT. REVIEW OF THE PRODUCTION DOCUMENTATION CONFIRMED THAT THE DEVICE WAS MANUFACTURED ACCORDING TO SPECIFICATIONS AND FULFILLED ALL THE REQUIREMENTS OF IN-PROCESS AND FINAL INSPECTION. BASED ON THE PROVIDED DOCUMENTATION AND THE CONDUCTED REVIEW, NO MANUFACTURING RELATED ROOT CAUSE COULD BE IDENTIFIED. THE TREATING PHYSICIAN RATED THE OUTCOME AS NO DEVICE- BUT PROCEDURE-RELATED COMPLICATION. IT SHOULD BE NOTED THAT THE IFU ADVISES THE USER TO SELECT THE STENT SIZE TO MATCH THE DIAMETER OF THE VESSEL TO ACHIEVE A FINAL STENT DIAMETER TO VESSEL RATIO OF 1:1. PK PAPYRUS IS INDICATED FOR THE TREATMENT OF ACUTE PERFORATIONS OF NATIVE CORONARY ARTERIES AND CORONARY BYPASS GRAFTS IN VESSELS 2.5 TO 5.0 MM IN DIAMETER. AS SUCH, THIS DEVICE IS TYPICALLY USED IN EMERGENT RESPONSE TO A LIFE-THREATENING ADVERSE PROCEDURAL EVENT UNRELATED TO USE OF THE PK PAPYRUS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22044306&pc=NPT | 2025587-2025-03757 | 2025-05-06 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-05-19 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 10 MONTHS FOLLOWING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC AORTIC VALVE SUDDEN CARDIAC DEATH OCCURRED. THE PHYSICIAN INDICATED THE DEATH WAS POSSIBLY RELATED TO THE IMPLANT PROCEDURE AND VALVE.Event Description: ADDITIONAL INFORMATION INDICATED THE OFFICIAL CAUSE OF DEATH WAS NOT KNOWN. THE VALVE REMAINED IMPLANTED.Event Description: IT WAS REPORTED THAT APPROXIMATELY 10 MONTHS FOLLOWING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC AORTIC VALVE SUDDEN CARDIAC DEATH OCCURRED. THE PHYSICIAN INDICATED THE DEATH WAS POSSIBLY RELATED TO THE IMPLANT PROCEDURE AND VALVE. ADDITIONAL INFORMATION INDICATED THE OFFICIAL CAUSE OF DEATH WAS NOT KNOWN. THE VALVE REMAINED IMPLANTED. ADDITIONAL INFORMATION WAS RECEIVED TO CLARIFY NEITHER THE MEDTRONIC VALVE NOR THE PROCEDURE PERFORMED TO IMPLANT THE VALVE WERE RELATED TO THE PATIENT'S DEATH. Manufacturer Narrative: UPDATED DATA: D4 - UDI H6 - ANNEX B. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A THIRD PARAGRAPH WAS ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22044335&pc=DSQ | 2916596-2025-03063 | 2019-07-01 04:00 | Death | THORATEC CORPORATION | 2025-05-19 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemorrhage/Bleeding; Unspecified Infection; Thrombosis/Thrombus | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED ¿EVALUATION OF SAFETY CRITERIA FOR ENOXAPARIN BRIDGING IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES IN AN OUTPATIENT CARE SETTING¿ THAT THE HEARTMATE 3 MAY BE ATTRIBUTED TO BLEEDING, INFECTION, CARDIAC ARRHYTHMIA, PUMP THROMBOSIS, AND DEATH. THE STUDY SOUGHT TO EVALUATE BOTH THE SHORT-TERM ADVERSE EVENTS AND LONG-TERM OUTCOMES OF ENOXAPARIN BRIDGING AT A MAJOR TRANSPLANT CENTER, FOLLOWING THE IMPLEMENTATION OF BRIDGING SAFETY CRITERIA. A TOTAL OF 85 PATIENTS BETWEEN JUL2019 AND MAY2022 WERE INCLUDED IN THE STUDY, WITH 51 PATIENTS BRIDGED WITH ENOXAPARIN AND 34 NON-BRIDGED. OF THE PATIENTS BRIDGED WITH ENOXAPARIN, 53% WERE SUPPORTED BY HEARTMATE III (N=27), 35% BY HVAD (N=18), AND 12% BY HEARTMATE II (N=6). THE PRIMARY ENDPOINT WAS THE OCCURRENCE OF MAJOR BLEEDING/THROMBOSIS EVENTS WITHIN 30DAYS OF BRIDGING. THE SECONDARY ENDPOINT WAS FREEDOM FROM EVENTS 30 DAYS AFTER ENOXAPARIN INITIATION AND OVERALL PATIENT SURVIVABILITY UNTIL THE LAST FOLLOW-UP. WITHIN 30 DAYS, NO MAJOR BLEEDING/THROMBOTIC EVENTS WERE NOTED. AFTER 30 DAYS, THE MAJOR BLEEDING RATE WAS 5.8% IN BRIDGED VS. 11.8% IN NON-BRIDGED PATIENTS. THE INCIDENCE OF MAJOR BLEEDING IN THE BRIDGED GROUP INCLUDED ONE CASE OF GASTROINTESTINAL (GI) BLEEDING IN AN HVAD PATIENT AND ONE NOSE/DENTAL BLEEDING IN EACH HMII AND HMIII PATIENTS. TWO HVAD PATIENTS DEVELOPED STROKES: ONE HEMORRHAGIC AND ONE ISCHEMIC/EMBOLIC. STATISTICALLY, THERE WERE NO DIFFERENCES IN BOTH GROUPS WITH 3.9% IN BRIDGED AND 2.9% IN THE NON-BRIDGED GROUP. OVERALL, 3-YEAR SURVIVAL WAS 78% IN THE BRIDGED VS. 49% IN NON-BRIDGED PATIENTS. ANALYSIS OF NON-BRIDGED PATIENTS REVEALED MORTALITIES PRIMARILY DUE TO MAJOR INFECTION (33%), CARDIAC ARRHYTHMIA (33%) AND THE REMAINING 33% WERE COMPLICATIONS DURING EARLY POST-LVAD IMPLANT, PUMP THROMBOSIS, AND PATIENT WITHDRAWAL OF CARE. THE STUDY CONCLUDED THAT WHEN ELIGIBILITY CRITERIA FOR SAFE BRIDGING WERE APPLIED, THE USE OF ENOXAPARIN BRIDGING WAS ASSOCIATED WITH NO MAJOR BLEEDING/THROMBOTIC EVENTS DURING BRIDGING AND IMPROVED 3-YEAR SURVIVAL IN LEFT VENTRICULAR ASSIST DEVICE PATIENTS. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. CEDARS SINAI COMPREHENSIVE TRANSPLANT CENTER, LOS ANGELES, CALIFORNIA, USA ; SMIDT HEART INSTITUTE, CEDARS SINAI MEDICAL CENTER, LOS ANGELES, CALIFORNIA, USA ; DEPARTMENT OF PHARMACY, CEDARS SINAI MEDICAL CENTER, LOS ANGELES, CALIFORNIA, USA ; DEPARTMENT OF PATHOLOGY, CEDARS SINAI MEDICAL CENTER, LOS ANGELES, CALIFORNIA, USA. LAM, L. D., CZER, L. S., RUNYAN, C., OTAROLA, I., JANG, J., LAU, J., GAU, M., HERNANDEZ, K., NGO, T., AGUILLON, M., HUIE, N., CHEN, J. W., COLE, R., MORIGUCHI, J., & VOLOD, O. (2024). EVALUATION OF SAFETY CRITERIA FOR ENOXAPARIN BRIDGING IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES IN AN OUTPATIENT CARE SETTING. ARTIFICIAL ORGANS. HTTPS://DOI.ORG/10.1111/AOR.14900. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE HEARTMATE 3 DEVICE SERIAL NUMBERS AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE HEARTMATE 3 LVAS IFU, DOCUMENT REVISION D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA, INFECTION, DEVICE THROMBOSIS, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION. MODIFICATIONS. SECTION 6 OF THE IFU (UNDER "ONGOING PATIENT ASSESSMENT AND CARE") LISTS INFECTION, ARRHYTHMIA, AND THROMBOEMBOLISM AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. FURTHERMORE, SEVERAL SECTIONS OF THE HM3 IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22044497&pc=KRH | 9612164-2025-02471 | 2020-11-30 05:00 | Death | MEDTRONIC MEXICO | 2025-05-19 04:00 | KRH | SIMUFORM | Insufficient Information | Insufficient Information | K201449 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT 28 DAYS POST IMPLANT OF THIS SIMUFORM RING, THE PATIENT DIED. THE CAUSE OF DEATH WAS REPORTED AS A BRAIN ANEURYSM. RELATEDNESS OF THE DEATH TO THE SIMUFORM RING COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PHYSICIAN DOES NOT BELIEVE THAT MEDTRONIC PRODUCT CAUSED OR CONTRIBUTED TO THIS DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5 AND D4 UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22044705&pc=LMF | 1810189-2025-00018 |  | Death | PFIZER, INC. (DEVICE) | 2025-05-19 04:00 | LMF | GELFOAM | Off-Label Use | Multiple Organ Failure | 18-286 |  | 1 | Event Description: EVENT VERBATIM [PREFERRED TERM] INDICATION: GASTRIC VARICES [OFF LABEL USE], INDICATION: GASTRIC VARICES [DEVICE USE ISSUE], ONE PATIENT SUCCUMBED, WITH MODS IN THE COIL + GF GROUP [MULTIPLE ORGAN DYSFUNCTION SYNDROME], , NARRATIVE: THIS IS A LITERATURE REPORT FOR THE FOLLOWING LITERATURE SOURCE(S): "EFFICACY AND SAFETY OF ENDOSCOPIC ULTRASOUND GUIDED COIL EMBOLIZATION WITH ABSORBABLE GELATIN SPONGE VERSUS COIL EMBOLIZATION WITH CYANOACRYLATE GLUE IN GASTRIC VARICES: A RANDOMIZED COMPARATIVE STUDY", GASTROINTESTINAL ENDOSCOPY, 2025; VOL:101(5 SUPPLE), PGS:S561, DOI:10.1016/J.GIE.2025.03.949. A PATIENT (AGE AND GENDER NOT PROVIDED) RECEIVED ABSORBABLE GELATIN (GELFOAM), FOR GASTRIC VARICES. THE PATIENT'S RELEVANT MEDICAL HISTORY WAS NOT REPORTED. THE PATIENT TOOK CONCOMITANT MEDICATIONS. THE FOLLOWING INFORMATION WAS REPORTED: OFF LABEL USE (DEATH, MEDICALLY SIGNIFICANT), DEVICE USE ISSUE (DEATH, MEDICALLY SIGNIFICANT), OUTCOME "FATAL" AND ALL DESCRIBED AS "INDICATION: GASTRIC VARICES"; MULTIPLE ORGAN DYSFUNCTION SYNDROME (DEATH, MEDICALLY SIGNIFICANT), OUTCOME "FATAL", DESCRIBED AS "ONE PATIENT SUCCUMBED, WITH MODS IN THE COIL + GF GROUP". THE PATIENT UNDERWENT THE FOLLOWING LABORATORY TESTS AND PROCEDURES: ENDOSCOPIC ULTRASOUND: UNKNOWN RESULTS; ULTRASOUND DOPPLER: UNKNOWN RESULTS. THE PATIENT DATE OF DEATH WAS UNKNOWN. REPORTED CAUSE OF DEATH: "INDICATION: GASTRIC VARICES", "MULTIPLE ORGAN DYSFUNCTION SYNDROME". IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. CLINICAL COURSE: THIS STUDY AIMS TO COMPARE THE EFFICACY AND SAFETY OF ENDOSCOPIC ULTRASOUND (EUS) GUIDED COIL EMBOLIZATION WITH GELATIN SPONGE (GELFOAM) VERSUS COIL EMBOLIZATION WITH CYA GLUE INJECTION FOR MANAGING PATIENTS WITH GASTRIC VARICES. PATIENTS DIAGNOSED WITH GASTRIC VARICES (GOV2/IGV1) WERE INCLUDED, WITH 70-PATIENTS RANDOMIZED IN 1:1 RATIO INTO TWO GROUPS: EUS-GUIDED COIL EMBOLIZATION WITH ABSORBABLE GELATIN SPONGE (COIL + GF; N = 35) OR EUS-GUIDED COIL EMBOLIZATION WITH CYA (COIL + CYA; N = 35). FOLLOWUP EUS-DOPPLER WAS SCHEDULED AT THE 1 ST, 3 RD, 6 TH AND 12 TH MONTHS POST-INITIAL PROCEDURE TO EVALUATE GASTRIC VARICES REAPPEARANCE, NEED FOR REINTERVENTION, SURVIVAL, AND ADVERSE EVENTS. ONE PATIENT FROM EACH GROUP SUCCUMBED, WITH MODS IN THE COIL + GF GROUP AND SUSPECTED NON-TARGET EMBOLIC EVENT IN THE COIL + CYA GROUP AT 6 WEEKS. CAUSALITY FOR "INDICATION: GASTRIC VARICES" AND "ONE PATIENT SUCCUMBED, WITH MODS IN THE COIL + GF GROUP" WAS DETERMINED ASSOCIATED TO ABSORBABLE GELATIN (MALFUNCTION)., COMMENT: THE EVENT MULTIPLE ORGAN DYSFUNCTION SYNDROME IS ASSESSED AS SERIOUS, ASSOCIATED WITH THE PRODUCT ABSORBABLE GELATIN (GELFOAM) IN THE CONTEXT OF OFF LABEL USE, AND UNLISTED IN THE CDS OF THE PFIZER SUSPECT PRODUCT ABSORBABLE GELATIN. THIS CASE WILL BE REASSESSED WHEN ADDITIONAL INFORMATION IS RECEIVED.Event Description: EVENT [PREFERRED TERM] INDICATION: GASTRIC VARICES [OFF LABEL USE], INDICATION: GASTRIC VARICES [DEVICE USE ISSUE], ONE PATIENT SUCCUMBED, WITH MODS IN THE COIL + GF GROUP [MULTIPLE ORGAN DYSFUNCTION SYNDROME], NARRATIVE: THIS IS A LITERATURE REPORT FOR THE FOLLOWING LITERATURE SOURCE(S): "EFFICACY AND SAFETY OF ENDOSCOPIC ULTRASOUND GUIDED COIL EMBOLIZATION WITH ABSORBABLE GELATIN SPONGE VERSUS COIL EMBOLIZATION WITH CYANOACRYLATE GLUE IN GASTRIC VARICES: A RANDOMIZED COMPARATIVE STUDY", GASTROINTESTINAL ENDOSCOPY, 2025; VOL:101(5 SUPPLE), PGS: S561, DOI: 10.1016/J.GIE.2025.03.949. A PATIENT (AGE AND GENDER NOT PROVIDED) RECEIVED ABSORBABLE GELATIN (GELFOAM), FOR GASTRIC VARICES. THE PATIENT'S RELEVANT MEDICAL HISTORY WAS NOT REPORTED. THE PATIENT TOOK CONCOMITANT MEDICATIONS. THE FOLLOWING INFORMATION WAS REPORTED: OFF LABEL USE (DEATH, MEDICALLY SIGNIFICANT), DEVICE USE ISSUE (DEATH, MEDICALLY SIGNIFICANT), OUTCOME "FATAL" AND ALL DESCRIBED AS "INDICATION: GASTRIC VARICES"; MULTIPLE ORGAN DYSFUNCTION SYNDROME (DEATH, MEDICALLY SIGNIFICANT), OUTCOME "FATAL", DESCRIBED AS "ONE PATIENT SUCCUMBED, WITH MODS IN THE COIL + GF GROUP". THE PATIENT UNDERWENT THE FOLLOWING LABORATORY TESTS AND PROCEDURES: ENDOSCOPIC ULTRASOUND: UNKNOWN RESULTS; ULTRASOUND DOPPLER: UNKNOWN RESULTS. THE PATIENT DATE OF DEATH WAS UNKNOWN. REPORTED CAUSE OF DEATH: "INDICATION: GASTRIC VARICES", "MULTIPLE ORGAN DYSFUNCTION SYNDROME". IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. CLINICAL COURSE: THIS STUDY AIMS TO COMPARE THE EFFICACY AND SAFETY OF ENDOSCOPIC ULTRASOUND (EUS) GUIDED COIL EMBOLIZATION WITH GELATIN SPONGE (GELFOAM) VERSUS COIL EMBOLIZATION WITH CYA GLUE INJECTION FOR MANAGING PATIENTS WITH GASTRIC VARICES. PATIENTS DIAGNOSED WITH GASTRIC VARICES (GOV2/IGV1) WERE INCLUDED, WITH 70-PATIENTS RANDOMIZED IN 1:1 RATIO INTO TWO GROUPS: EUS-GUIDED COIL EMBOLIZATION WITH ABSORBABLE GELATIN SPONGE (COIL + GF; N = 35) OR EUS-GUIDED COIL EMBOLIZATION WITH CYA (COIL + CYA; N = 35). FOLLOWUP EUS-DOPPLER WAS SCHEDULED AT THE 1 ST, 3 RD, 6 TH AND 12 TH MONTHS POST-INITIAL PROCEDURE TO EVALUATE GASTRIC VARICES REAPPEARANCE, NEED FOR REINTERVENTION, SURVIVAL, AND ADVERSE EVENTS. ONE PATIENT FROM EACH GROUP SUCCUMBED, WITH MODS IN THE COIL + GF GROUP AND SUSPECTED NON-TARGET EMBOLIC EVENT IN THE COIL + CYA GROUP AT 6 WEEKS. CAUSALITY FOR "INDICATION: GASTRIC VARICES" AND "ONE PATIENT SUCCUMBED, WITH MODS IN THE COIL + GF GROUP" WAS DETERMINED ASSOCIATED TO ABSORBABLE GELATIN (MALFUNCTION). FOLLOW-UP (14MAY2025): THIS IS A LITERATURE REPORT FOR THE FOLLOWING LITERATURE SOURCE ENTITLED EFFICACY AND SAFETY OF ENDOSCOPIC ULTRASOUND GUIDED COIL EMBOLIZATION WITH ABSORBABLE GELATIN SPONGE VERSUS COIL EMBOLIZATION WITH CYANOACRYLATE GLUE IN GASTRIC VARICES: A RANDOMIZED COMPARATIVE STUDY, GASTROINTESTINAL ENDOSCOPY, 2025, VOL 101(5 SUPPLE), PGS S561, DOI: 10.1016/J.GIE.2025.03.949. THIS IS A FOLLOW-UP REPORT BASED ON THE RECEIPT OF THE PUBLICATION; THE CASE HAS BEEN UPDATED TO INCLUDE ADDITIONAL INFORMATION IDENTIFIED IN THE PUBLICATION. UPDATED INFORMATION INCLUDED: REPORTER INFORMATION., COMMENT: THE EVENT MULTIPLE ORGAN DYSFUNCTION SYNDROME IS ASSESSED AS SERIOUS, ASSOCIATED WITH THE PRODUCT ABSORBABLE GELATIN (GELFOAM) IN THE CONTEXT OF OFF LABEL USE, AND UNLISTED IN THE CDS OF THE PFIZER SUSPECT PRODUCT ABSORBABLE GELATIN. THIS CASE WILL BE REASSESSED WHEN ADDITIONAL INFORMATION IS RECEIVED.Event Description: EVENT VERBATIM [PREFERRED TERM] INDICATION: GASTRIC VARICES [OFF LABEL USE], INDICATION: GASTRIC VARICES [DEVICE USE ISSUE], ONE PATIENT SUCCUMBED, WITH MODS IN THE COIL + GF GROUP [MULTIPLE ORGAN DYSFUNCTION SYNDROME], , NARRATIVE: THIS IS A LITERATURE REPORT FROM PRODUCT QUALITY GROUP FOR THE FOLLOWING LITERATURE SOURCE(S): "EFFICACY AND SAFETY OF ENDOSCOPIC ULTRASOUND GUIDED COIL EMBOLIZATION WITH ABSORBABLE GELATIN SPONGE VERSUS COIL EMBOLIZATION WITH CYANOACRYLATE GLUE IN GASTRIC VARICES: A RANDOMIZED COMPARATIVE STUDY", GASTROINTESTINAL ENDOSCOPY, 2025; VOL:101(5 SUPPLE), PGS:S561, DOI:10.1016/J.GIE.2025.03.949. A PATIENT (AGE AND GENDER NOT PROVIDED) RECEIVED ABSORBABLE GELATIN (GELFOAM), FOR GASTRIC VARICES. THE PATIENT'S RELEVANT MEDICAL HISTORY WAS NOT REPORTED. THE PATIENT TOOK CONCOMITANT MEDICATIONS. THE FOLLOWING INFORMATION WAS REPORTED: OFF LABEL USE (DEATH, MEDICALLY SIGNIFICANT), DEVICE USE ISSUE (DEATH, MEDICALLY SIGNIFICANT), OUTCOME "FATAL" AND ALL DESCRIBED AS "INDICATION: GASTRIC VARICES"; MULTIPLE ORGAN DYSFUNCTION SYNDROME (DEATH, MEDICALLY SIGNIFICANT), OUTCOME "FATAL", DESCRIBED AS "ONE PATIENT SUCCUMBED, WITH MODS IN THE COIL + GF GROUP". THE PATIENT UNDERWENT THE FOLLOWING LABORATORY TESTS AND PROCEDURES: ENDOSCOPIC ULTRASOUND: UNKNOWN RESULTS; ULTRASOUND DOPPLER: UNKNOWN RESULTS. THE PATIENT DATE OF DEATH WAS UNKNOWN. REPORTED CAUSE OF DEATH: "INDICATION: GASTRIC VARICES", "MULTIPLE ORGAN DYSFUNCTION SYNDROME". IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. CLINICAL COURSE: THIS STUDY AIMS TO COMPARE THE EFFICACY AND SAFETY OF ENDOSCOPIC ULTRASOUND (EUS) GUIDED COIL EMBOLIZATION WITH GELATIN SPONGE (GELFOAM) VERSUS COIL EMBOLIZATION WITH CYA GLUE INJECTION FOR MANAGING PATIENTS WITH GASTRIC VARICES. PATIENTS DIAGNOSED WITH GASTRIC VARICES (GOV2/IGV1) WERE INCLUDED, WITH 70-PATIENTS RANDOMIZED IN 1:1 RATIO INTO TWO GROUPS: EUS-GUIDED COIL EMBOLIZATION WITH ABSORBABLE GELATIN SPONGE (COIL + GF; N = 35) OR EUS-GUIDED COIL EMBOLIZATION WITH CYA (COIL + CYA; N = 35). FOLLOWUP EUS-DOPPLER WAS SCHEDULED AT THE 1 ST, 3 RD, 6 TH AND 12 TH MONTHS POST-INITIAL PROCEDURE TO EVALUATE GASTRIC VARICES REAPPEARANCE, NEED FOR REINTERVENTION, SURVIVAL, AND ADVERSE EVENTS. ONE PATIENT FROM EACH GROUP SUCCUMBED, WITH MODS IN THE COIL + GF GROUP AND SUSPECTED NON-TARGET EMBOLIC EVENT IN THE COIL + CYA GROUP AT 6 WEEKS. CAUSALITY FOR "INDICATION: GASTRIC VARICES" AND "ONE PATIENT SUCCUMBED, WITH MODS IN THE COIL + GF GROUP" WAS DETERMINED ASSOCIATED TO ABSORBABLE GELATIN (MALFUNCTION). PRODUCT QUALITY GROUP PROVIDED INVESTIGATIONAL RESULTS ON 04JUN2025 FOR ABSORBABLE GELATIN: INVESTIGATION SUMMARY AND CONCLUSION: THE COMPLAINT FOR 'ADVERSE EVENT MD/MDCP' FOR AN UNKNOWN BATCH OF GELFOAM ABSORBABLE MATERIAL WAS INVESTIGATED. THE INVESTIGATION INCLUDED REVIEWING APRRS FOR THE PRODUCT AND A MEDICAL DEVICE TREND REVIEW. THE FINAL SCOPE WAS DETERMINED TO BE ALL BATCHES OF GELFOAM MANUFACTURED BY THE SITE WITHIN THE 36 MONTHS (THE EXPIRY INTERVAL OF THE PRODUCT) PRIOR TO THE RECEIPT DATE OF THE COMPLAINT. A COMPLAINT SAMPLE WAS NOT RETURNED. NO RELATED QUALITY ISSUES WERE IDENTIFIED DURING THE INVESTIGATION. NO ROOT CAUSE OR CAPA WERE IDENTIFIED AS THE COMPLAINT WAS NOT CONFIRMED. NO RELATED QUALITY DEFECTS WERE IDENTIFIED THROUGH INVESTIGATION. THERE IS NO IMPACT ON MEDICAL DEVICE QUALITY, REGULATORY, VALIDATION, OR STABILITY.ADDITIONAL PRODUCT QUALITY INFORMATION: "INVESTIGATION SUMMARY AND CONCLUSION: THE COMPLAINT FOR 'ADVERSE EVENT MD/MDCP' FOR AN UNKNOWN BATCH OF GELFOAM ABSORBABLE MATERIAL WAS INVESTIGATED. THE INVESTIGATION INCLUDED REVIEWING APRRS FOR THE PRODUCT AND A MEDICAL DEVICE TREND REVIEW. THE FINAL SCOPE WAS DETERMINED TO BE ALL BATCHES OF GELFOAM MANUFACTURED BY THE SITE WITHIN THE 36 MONTHS (THE EXPIRY INTERVAL OF THE PRODUCT) PRIOR TO THE RECEIPT DATE OF THE COMPLAINT. A COMPLAINT SAMPLE WAS NOT RETURNED. NO RELATED QUALITY ISSUES WERE IDENTIFIED DURING THE INVESTIGATION. NO ROOT CAUSE OR CAPA WERE IDENTIFIED AS THE COMPLAINT WAS NOT CONFIRMED. NO RELATED QUALITY DEFECTS WERE IDENTIFIED THROUGH INVESTIGATION. THERE IS NO IMPACT ON MEDICAL DEVICE QUALITY, REGULATORY, VALIDATION, OR STABILITY." FOLLOW-UP (14MAY2025): THIS IS A LITERATURE REPORT FOR THE FOLLOWING LITERATURE SOURCE ENTITLED EFFICACY AND SAFETY OF ENDOSCOPIC ULTRASOUND GUIDED COIL EMBOLIZATION WITH ABSORBABLE GELATIN SPONGE VERSUS COIL EMBOLIZATION WITH CYANOACRYLATE GLUE IN GASTRIC VARICES: A RANDOMIZED COMPARATIVE STUDY, GASTROINTESTINAL ENDOSCOPY, 2025, VOL 101(5 SUPPLE), PGS S561, DOI: 10.1016/J.GIE.2025.03.949. THIS IS A FOLLOW-UP REPORT BASED ON THE RECEIPT OF THE PUBLICATION; THE CASE HAS BEEN UPDATED TO INCLUDE ADDITIONAL INFORMATION IDENTIFIED IN THE PUBLICATION. UPDATED INFORMATION INCLUDED: REPORTER INFORMATION. FOLLOW-UP (04JUN2025): THIS IS A FOLLOW-UP REPORT FROM PRODUCT QUALITY GROUP PROVIDING INVESTIGATION RESULTS., COMMENT: THE EVENT MULTIPLE ORGAN DYSFUNCTION SYNDROME IS ASSESSED AS SERIOUS, ASSOCIATED WITH THE PRODUCT ABSORBABLE GELATIN (GELFOAM) IN THE CONTEXT OF OFF LABEL USE, AND UNLISTED IN THE CDS OF THE PFIZER SUSPECT PRODUCT ABSORBABLE GELATIN. THIS CASE WILL BE REASSESSED WHEN ADDITIONAL INFORMATION IS RECEIVED. Manufacturer Narrative: INVESTIGATION SUMMARY AND CONCLUSION: THE COMPLAINT FOR 'ADVERSE EVENT MD/MDCP' FOR AN UNKNOWN BATCH OF GELFOAM ABSORBABLE MATERIAL WAS INVESTIGATED. THE INVESTIGATION INCLUDED REVIEWING APRRS FOR THE PRODUCT AND A MEDICAL DEVICE TREND REVIEW. THE FINAL SCOPE WAS DETERMINED TO BE ALL BATCHES OF GELFOAM MANUFACTURED BY THE SITE WITHIN THE 36 MONTHS (THE EXPIRY INTERVAL OF THE PRODUCT) PRIOR TO THE RECEIPT DATE OF THE COMPLAINT. A COMPLAINT SAMPLE WAS NOT RETURNED. NO RELATED QUALITY ISSUES WERE IDENTIFIED DURING THE INVESTIGATION. NO ROOT CAUSE OR CAPA WERE IDENTIFIED AS THE COMPLAINT WAS NOT CONFIRMED. NO RELATED QUALITY DEFECTS WERE IDENTIFIED THROUGH INVESTIGATION. THERE IS NO IMPACT ON MEDICAL DEVICE QUALITY, REGULATORY, VALIDATION, OR STABILITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22045152&pc=DSQ | 22045152 | 2025-05-10 04:00 | Death | BERLIN HEART GMBH | 2025-05-19 04:00 | DSQ | EXCOR | Crack; Fracture; Material Split, Cut or Torn; Air/Gas in Device | Cardiac Arrest; Hemorrhage/Bleeding; Obstruction/Occlusion; Convulsion/Seizure |  |  | 1 | Event Description: THE PATIENT WAS ON A BERLIN HEART VENTRICULAR ASSIST DEVICE (VAD.) THE NURSE NOTICED BLEEDING FROM THE INSERTION SITE AT THE INFLOW CANNULA. THE PATIENT PROVIDERS WERE CALLED TO THE BEDSIDE. AT THAT TIME, AIR WAS NOTED IN THE VAD. AN ATTEMPT WAS MADE TO TRANSITION THE PATIENT TO AN EXTRA-CORPOREAL LIFE SUPPORT (ECLS) SYSTEM WHICH WAS UNSUCCESSFUL. IT WAS NOTICED THAT INFLOW CANNULA HAD A TEAR IN THE TUBING. OPERATIVE NOTES: DATE OF OPERATION: [REDACTED DATE]. PREOPERATIVE DIAGNOSIS: DILATED CARDIOMYOPATHY, STATUS POST BERLIN LEFT VENTRICULAR ASSIST DEVICE (LVAD), LISTED FOR HEART TRANSPLANTATION STATUS 1A, BLEEDING, CARDIAC ARREST WITH CPR IN PROGRESS. POSTOPERATIVE DIAGNOSIS: DILATED CARDIOMYOPATHY, STATUS POST BERLIN LEFT VENTRICULAR ASSIST DEVICE (LVAD), LISTED FOR HEART TRANSPLANTATION STATUS 1A, BLEEDING, CARDIAC ARREST. OPERATION: EMERGENT CANNULATION FOR ARTERIOVENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). THIS PATIENT IS A [REDACTED AGE] WITH A DIAGNOSIS OF DILATED CARDIOMYOPATHY, WHO IS STATUS POST PLACEMENT OF A BERLIN LEFT VENTRICULAR ASSIST DEVICE ON [REDACTED DATE]. PATIENT HAS BEEN ON A BERLIN LVAD FOR 291 DAYS. WE WERE EMERGENTLY CALLED TO THE BEDSIDE SECONDARY TO ACUTE DECOMPENSATION FROM BLEEDING EMANATING FROM THE BERLIN INFLOW CANNULA. THIS WAS CONTROLLED BY MANUAL COMPRESSION BY REPORT, AND TRANSFUSION PROTOCOL HAD BEEN INSTITUTED. PATIENT WAS THEN NOTED TO HAVE SEIZURES AND WAS SUBSEQUENTLY INTUBATED. THE TEAM THEN NOTED AIR IN THE BERLIN PUMP, AND THE TEAM PROCEEDED TO CLAMP THE INFLOW CANNULA WHILE THE PATIENT WAS CONTINUING TO BE RESUSCITATED. CPR WAS THEN INSTITUTED, WHICH INCLUDED CHEST COMPRESSIONS AS WELL AS INSTITUTION OF OUR CODE PROTOCOL. THE PATIENT ALSO REQUIRED REINTUBATION. AT THIS POINT, WE ARRIVED, AND GIVEN THE SITUATION, OUR PLAN WAS TO EMERGENTLY GO ON ARTERIOVENOUS ECMO VIA RIGHT NECK CANNULATION. WE DISCUSSED THE SITUATION WITH OUR TEAM INCLUDING CARDIOLOGISTS AND INTENSIVE CARE TEAM AND OUR ONLY OPTION WAS TO ON ECMO. OF NOTE, THE PATIENT'S RIGHT INTERNAL JUGULAR VEIN WAS OCCLUDED; HOWEVER, WE STILL ATTEMPTED TO PLACE CANNULAS GIVEN THE ACUTE NATURE OF THE SITUATION. WE DID COMMUNICATE WITH THE FAMILY THE GRAVITY OF THE SITUATION AND EXPLAINED THAT WE WOULD ATTEMPT TO GO ON ECMO. PATIENT WAS ENROUTE DURING THE EVENT, AND WE MAINTAINED COMMUNICATION AND UPDATES WITH THE FAMILY. FINDINGS: OCCLUDED RIGHT INTERNAL JUGULAR VEIN. PERFORATION/CRACK IN THE INFLOW CANNULA (1-2MM) OF BERLIN LVAD, NEAR THE CONNECTION TO THE VELOUR OF THE CANNULA. PROCEDURE: THE RIGHT NECK WAS THEN PREPPED AND DRAPED WHILE CHEST COMPRESSIONS WERE ONGOING. AN INCISION WAS MADE IN THE RIGHT NECK, AND THE RIGHT INTERNAL CAROTID ARTERY WAS DISSECTED FREE AND ENCIRCLED. OF NOTE, WE COULD REALLY NOT IDENTIFY THE RIGHT INTERNAL JUGULAR VEIN; HOWEVER, THERE WAS A COLLATERAL THAT WAS IDENTIFIED AND WAS QUITE SMALL. WE PROCEEDED TO HEPARINIZE, AND WE CANNULATED THE RIGHT CAROTID ARTERY WITH A 10-FRENCH ARTERIAL CANNULA. THIS WAS DE-AIRED AND CONNECTED TO THE CIRCUIT. WE THEN TRANSECTED THIS SMALL INTERNAL JUGULAR VEIN AND PLACED A 10-FRENCH VENOUS CANNULA; HOWEVER, WE DID MEET RESISTANCE. THIS WAS CONNECTED TO OUR VENOUS CIRCUIT DESPITE NO SIGNIFICANT BACK-BLEEDING. THE CANNULA WAS FILLED WITH SALINE AND CONNECTED, AND WE PROCEEDED TO ATTEMPT TO GO ON ARTERIOVENOUS ECMO. GIVEN THE OCCLUSION OF THE INTERNAL JUGULAR VEIN, WE WERE UNABLE TO ACHIEVE ANY FLOW ON ECMO. AT THIS POINT, WE DECIDED TO DISCONTINUE OUR RESUSCITATIVE EFFORTS, AND THE PATIENT WAS PRONOUNCED AT 2151. WE THEN PROCEEDED TO INSPECT THE INFLOW CANNULA AND NOTED THAT THERE WAS A FRACTURE IN THE INFLOW CANNULA IMMEDIATELY ADJACENT TO THE VELOUR. THERE WAS NO EVIDENCE OF ANY OTHER FRACTURES OR DISRUPTION. THE OUTFLOW CANNULA WAS INTACT WITHOUT ANY ABNORMALITIES. MANUFACTURER RESPONSE FOR VENTRICULAR ASSIST DEVICE (VAD) DRIVER, IKUS STATIONARY DRIVING UNIT FOR EXCOR PARACORPOREAL VAD SYSTEM (PER SITE REPORTER). BERLIN HEART WILL PERFORM AN INVESTIGATION. MANUFACTURER RESPONSE FOR APEX CANNULA, EXCOR APEX CANNULA FOR SMALL CHILDREN (PER SITE REPORTER), BERLIN HEART WILL INVESTIGATE. MANUFACTURER RESPONSE FOR APEX CANNULA, EXCOR APEX CANNULAE (PER SITE REPORTER), BERLIN HEART WILL PERFORM AN INVESTIGATION. MANUFACTURER RESPONSE FOR VAD BLOOD PUMP, EXCOR BLOOD PUMP WITH PU VALVES (PER SITE REPORTER), BERLIN HEART WILL PERFORM AN INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22045308&pc=NGV | 2135147-2025-02629 | 2025-04-24 04:00 | Death | ABBOTT MEDICAL | 2025-05-19 04:00 | NGV | AMPLATZER AMULET | Patient-Device Incompatibility | Perforation; Pericardial Effusion | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS CHOSEN FOR A PROCEDURE USING A 12F TORQVUE 45X45 DELIVERY SHEATH. DURING PROCEDURE, AFTER THREE PARTIAL RECAPTURES IT WAS DETERMINED THAT A LARGER DEVICE WAS NEEDED FOR BETTER TISSUE OPPOSITION, AS COLOR FLOW AROUND THE DEVICE AND MINIMAL COMPRESSION WAS SEEN WITH THE 25MM AMULET. THERE WAS NO ADDITIONAL PRODUCT EXPERIENCE OTHER THAN THE MIS-SIZING OCCURRED ON THE 25MM AMULET AND IT WAS REMOVED FROM THE PATIENT PRIOR TO RELEASE FROM THE DELIVERY CABLE. A NEW 28MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS CHOSEN AND PLACED WITH A SINGLE POSITION ATTEMPT WITH A 14F TORQVUE 45X45 DELIVERY SHEATH. THE DEVICE POSITION MET CLOSE CRITERIA, ANGIOGRAPHY AND COLOR FLOW DOPPLER SHOWED NO LEAKS AROUND THE DEVICE, AND A 30 SECOND TENSION TEST WAS PERFORMED WITH NO MOVEMENT NOTED ON ECHOCARDIOGRAM (ECHO) OR FLUOROSCOPY. THE DEVICE WAS RELEASED AND THE CASE WAS FINISHED. HEPARIN WAS REVERSED WITH 50MG OF PROTAMINE AT THE CONCLUSION OF THE PROCEDURE. THE PATIENT WAS RECOVERED IN THE HOSPITAL FOR SEVERAL HOURS AND ECHOCARDIOGRAM PRIOR TO DISCHARGE SHOWED NO ABNORMAL FINDINGS. ON (B)(6) 2025, IT WAS REPORTED THAT PATIENT HAD DIED AT HOME THE NIGHT HE WAS DISCHARGED FROM THE HOSPITAL. THE CAUSE OF DEATH WAS DEVICE-RELATED PULMONARY ARTERY EROSION AND PERICARDIAL EFFUSION. THE AUTOPSY REPORT DISCLOSED AS 400ML OF BLOOD WAS FOUND IN THE PERICARDIAL SPACE. THE LEFT ATRIAL APPENDAGE WAS FOUND TO BE INTACT. NO DEVICE MATERIAL OR STABILIZING WIRES WERE OBSERVED OUTSIDE OF THE LEFT ATRIAL APPENDAGE. THE PULMONARY ARTERY DEMONSTRATED AN EROSION SITE OF 15MMX10MM ON THE EXTERNAL WALL OF THE PULMONARY ARTERY, AND A 2MM DEFECT ON THE INTERNAL WALL OF THE PULMONARY ARTERY. THE POST-MORTEM REPORT SPECULATED CAUSE OF EROSION FROM AN INTERACTION BETWEEN THE LEFT ATRIAL APPENDAGE AND THE PULMONARY ARTERY. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT OF PATIENT-DEVICE INCOMPATIBILITY (IMPLANT-RELATED TISSUE DAMAGE), PERICARDIAL EFFUSION AND PATIENT DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THERE ARE MULTIPLE RISK FACTORS FOR PERICARDIAL EFFUSION AFTER A LEFT ATRIAL APPENDAGE CLOSURE PROCEDURE, INCLUDING ANTICOAGULATION STATUS OF THE PATIENT, MANEUVERING OF THE GUIDEWIRE OR SHEATHS WITHIN THE HEART, PATIENT ANATOMY, MULTIPLE RECAPTURE OF DEVICES, AND THE TRANS-SEPTAL PUNCTURE. FIELD INDICATED THAT THREE PARTIAL RECAPTURES WERE PERFORMED BEFORE EXCHANGE WITH LARGER DEVICE. FIELD ALSO NOTED THAT CAUSE OF DEATH WAS DEVICE-RELATED PULMONARY ARTERY EROSION AND PERICARDIAL EFFUSION. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED PATIENT-DEVICE INCOMPATIBILITY COULD NOT BE DETERMINED. THE REPORTED PERICARDIAL EFFUSION AND DEATH COULD BE RELATED TO PROCEDURE. HOWEVER, THIS CANNOT BE CONFIRMED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. FURTHERMORE, THIS COMPLAINT WAS REVIEWED BY ABBOTT. THE REVIEWER STATED THAT, THE PRESENCE OF AN INTACT LAA WITH NO EVIDENCE OF AN INJURY OR PROTRUSION OF ANCHORS THROUGH THE LAA MAKE IT DIFFICULT TO DETERMINE THE EXACT MECHANISM FOR EROSION. NONETHELESS, THE CAUSE OF DEATH IS PROCEDURE RELATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22045327&pc=NPT | 2135147-2025-02630 | 2025-04-25 04:00 | Death | ABBOTT MEDICAL | 2025-05-19 04:00 | NPT | NAVITOR VISION | Central Regurgitation | Heart Block; Heart Failure/Congestive Heart Failure; Aortic Valve Insufficiency/ Regurgitation | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 35MM NAVITOR VISIONS VALVE WAS CHOSEN FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM. THEY SCREENED THE VALVE PRIOR TO ENTERING BODY. THE VALVE WAS PREDILATED WITH A 20MM NON-ABBOTT BALLOON, CROSSED EASILY AND DELIVERED WITHOUT ISSUE. THE AORTA WAS HORIZONTAL AND THE FIRST IMPLANT DEPTH WAS FRACTION TOO HIGH. A PARTIAL RECAPTURE WAS DONE, BUT 2ND ATTEMPT WAS TOO LOW. THEY DECIDED TO FULLY RECAPTURE AND RE-DEPLOYMENT. THE VALVE WOULD NOT RECAPTURE DUE TO KINK AND POSSIBLE INTUSUSCEPTION OF DELIVERY CAPSULE. THE VALVE COULD NOT DEPLOYED AS THEY HAD COME ACROSS ATRIOVENTRICULAR (AV). THE VALVE WAS APPROXIMATELY 70% RECAPTURED BUT KINK PREVENTED CAPSULE COMING BACK. THE VALVE WAS DRAGGED ACROSS TRANSVERSE ARCH TO DESCENDING AORTA, SNARED NOSE CONE AND VALVE TO ATTEMPT TO STRAIGHTEN OUT CAPSULE, BUT SNARING WAS FAILED. THEY DECIDED TO DEPLOY VALVE IN PROXIMAL DESCENDING AORTA AND IT WAS SUCCESSFUL. THEN THEY RE-SNARED CAPSULE AND SUCCESSFULLY STRAIGHTENED OUT CAPSULE TO ALLOW RETRIEVAL VIA LEFT FEMORAL ARTERY (LFA). THE LFA WAS CLOSED WITH TWO NON-ABBOTT DEVICES. UNFORTUNATELY, DUE TO A COMBINATION OF AORTIC REGURGITATION AND LEFT BUNDLE BRANCH BLOCK (LBBB), PATIENT WENT INTO LOW OUTPUT HEART FAILURE AND DIED 24 HOURS LATER. THE PATIENT WAS TOO FRAIL FOR RE-INTERVENTION. THE PHYSICIAN MENTIONED THE CAUSE OF DEATH WAS HEART FAILURE.Event Description: N/A. Manufacturer Narrative: AN EVENT OF CENTRAL REGURGITATION AND PATIENT DEATH WAS REPORTED. INFORMATION FROM INDICATED THAT THE PATIENT HAD HORIZONTAL AORTA AND A PARTIAL RECAPTURE WAS DONE AS THE VALVE WAS FRACTIONED TOO HIGH, BUT SECOND ATTEMPT WAS TOO LOW. FOLLOWING THIS, A RE-SHEATH OF THE VALVE WAS ATTEMPTED BUT WAS UNSUCCESSFUL WITH DAMAGE TO THE CAPSULE NOTED. THE VALVE COULD NOT BE RE-SHEATHED AND WAS DEPLOYED IN PROXIMAL DESCENDING AORTA. THE DELIVERY SYSTEM WAS SUBSEQUENTLY REMOVED FROM THE BODY AND THE FEMORAL ARTERIAL WAS CLOSED. UNFORTUNATELY DUE TO A COMBINATION OF AORTIC REGURGITATION AND LEFT BUNDLE BRANCH BLOCK, PATIENT WENT INTO LOW OUTPUT HEART FAILURE AND EXPIRED THE FOLLOWING DAY. THE CAUSE OF DEATH WAS REPORTED AS HEART FAILURE. FIELD INDICATED THAT THE PATIENT WAS TOO FRAIL FOR RE-INTERVENTION. THE DEVICE REMAINS IMPLANTED AND WILL NOT RETURN TO ABBOTT FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON CINE REVIEW OF THE ANGIOGRAPHIC IMAGING REVEALED SEVERELY TORTUOUS ILIOFEMORAL ARTERIES, A TORTUOUS THORACIC AORTA, AND A HORIZONTAL ASCENDING AORTA. BASED ON THE AVAILABLE INFORMATION AND MEDICAL REVIEW, THE DIFFICULTY IN RE-SHEATHING THE VALVE APPEARS TO BE DUE TO A COMBINATION OF COMPLEX ARTERIAL ANATOMY AND MULTIPLE RE-SHEATHING ATTEMPTS IN A PATIENT WITH SIGNIFICANT COMORBIDITIES, INCLUDING LEFT VENTRICULAR IMPAIRMENT AND CHRONIC KIDNEY DISEASE. THE REPORTED AORTIC REGURGITATION, LEFT BUNDLE BRANCH BLOCK, HEART FAILURE AND PATIENT DEATH APPEARS TO BE A CASCADING EFFECT. THE EXACT CAUSE COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22045340&pc=DSQ | 2916596-2025-02925 | 2024-12-25 05:00 | Death | THORATEC CORPORATION | 2025-05-19 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SYSTEMIC INFECTION. THIS WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. THE SOURCE OF THE INFECTION WAS NOT SHARED. THERE WERE NO CLINICAL SYMPTOMS/CULTURES IDENTIFIED. THE INFECTION WAS TREATED WITH STANDARD INFECTION TREATMENT USED BY THE CUSTOMER. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22045564&pc=DSQ | 2916596-2025-02928 | 2024-07-25 04:00 | Death | THORATEC CORPORATION | 2025-05-19 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SYSTEMIC INFECTION. THIS WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT RETURN. THERE WERE NO CLINICAL SYMPTOMS/CULTURES IDENTIFIED. THE INFECTION WAS TREATED WITH STANDARD INFECTION TREATMENT USED BY THIS HOSPITAL. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22045583&pc=NVN | 2649622-2025-13541 | 2025-04-24 04:00 | Death | MPRI | 2025-05-19 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY TWO DAYS AFTER THE IMPLANT OF THE LEADS. ADDITIONAL INFORMATION OBTAINED SUGGESTED THAT THE CAUSE OF DEATH COULD NOT BE IDENTIFIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22045603&pc=DTB | 2182208-2025-02441 | 2025-04-24 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-05-19 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY TWO DAYS AFTER THE IMPLANT OF THE LEADS. ADDITIONAL INFORMATION OBTAINED SUGGESTED THAT THE CAUSE OF DEATH COULD NOT BE IDENTIFIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22045674&pc=BZD | 2518422-2025-045052 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-19 04:00 | BZD | REP DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Hypersensitivity/Allergic reaction; Cancer; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION IN REFERENCE TO THE "REP DREAMSTATION AUTO CPAP" DEVICE ALLEGING DEATH OF THE PATIENT, EYE IRRITATION, NOSE IRRITATION, HYPERSENSITIVITY AND CANCER. THERE WAS A REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22046123&pc=DSQ | 2916596-2025-02968 | 2025-04-23 04:00 | Death | THORATEC CORPORATION | 2025-05-19 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22046172&pc=NGV | 2135147-2025-02638 | 2010-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-05-19 04:00 | NGV | AMPLATZER AMULET | Migration or Expulsion of Device; Patient Device Interaction Problem | Insufficient Information | P200049 |  | 1 | Event Description: THE ARTICLE, "SHORT- AND LONG-TERM OUTCOMES OF PERCUTANEOUS LEFT ATRIAL APPENDAGE OCCLUSION IN CANCER PATIENTS", WAS REVIEWED. THE ARTICLE PRESENTED AN OBSERVATIONAL AND RETROSPECTIVE STUDY TO COMPARE THE SAFETY AND EFFECTIVENESS IN SHORT- AND LONG-TERM OUTCOMES AMONG PATIENTS WITH AND WITHOUT CANCER WHO UNDERWENT LEFT ATRIAL APPENDAGE OCCLUSION (LAAO). DEVICES INCLUDED IN THE STUDY WERE E AMPLATZER CARDIAC PLUG¿ (ABBOTT®), AMULET¿ (ABBOTT®), WATCHMAN¿ (BOSTON SCIENTIFIC®), AND LAMBRE¿ (LIFETECH SCIENTIFIC®). THE ARTICLE CONCLUDED THAT LAAO PROCEDURE WAS SAFE AND EFFECTIVE IN BOTH POPULATIONS. CANCER PATIENTS EXPERIENCED HIGHER RATES OF ALL-CAUSE MORTALITY, WITH NO DIFFERENCES IN STROKE AND BLEEDING OUTCOMES BETWEEN GROUPS. [THE PRIMARY AND CORRESPONDING AUTHOR WAS RODRIGO EST´EVEZ-LOUREIRO, DEPARTMENT OF CARDIOLOGY. UNIVERSITY HOSPITAL ALVARO CUNQUEIRO, VIGO, SPAIN, WITH CORRESPONDING EMAIL: DEPARTMENT OF CARDIOLOGY. UNIVERSITY HOSPITAL ALVARO CUNQUEIRO, VIGO, SPAIN, WITH CORRESPONDING EMAIL: RODRIGO.ESTEVEZ.LOUREIRO@SERGAS.ES]. THE TIME FRAME OF THE STUDY WAS FROM APRIL 2010 TO DECEMBER 2023. A TOTAL OF 361 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 186 (51.5%) RECEIVED AN ABBOTT DEVICE. IN THE CANCER PATIENTS GROUP, THE AVERAGE AGE WAS 76 YEARS AND THE MAJORITY GENDER WAS MALE. IN THE NON-CANCER PATIENTS GROUP, THE AVERAGE AGE WAS 75 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, HEART FAILURE, HYPERTENSION, DIABETES, STROKE, TRANSIENT ISCHEMIC ATTACK, THROMBOEMBOLISM, VASCULAR DISEASE, ABNORMAL RENAL FUNCTION, ABNORMAL LIVER FUNCTION, BLEEDING (GASTROINTESTINAL, INTRACRANIAL, GENITOURINARY, PULMONARY), LABILE INR, ALCOHOL USE, CORONARY ARTERY DISEASE, PRIOR CORONARY ARTERY BYPASS GRAFT, VALVULAR HEART DISEASE, CARDIOMYOPATHY, OBSTRUCTIVE SLEEP APNEA, CHRONIC LUNG DISEASE.Event Description: N/A Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: SHORT- AND LONG-TERM OUTCOMES OF PERCUTANEOUS LEFT ATRIAL APPENDAGE OCCLUSION IN CANCER PATIENTS. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF AMPLATZER AMULET WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING ATRIAL FIBRILLATION, HEART FAILURE, HYPERTENSION, DIABETES, STROKE, TRANSIENT ISCHEMIC ATTACK, THROMBOEMBOLISM, VASCULAR DISEASE, ABNORMAL RENAL FUNCTION, ABNORMAL LIVER FUNCTION, BLEEDING (GASTROINTESTINAL, INTRACRANIAL, GENITOURINARY, PULMONARY), LABILE INR, ALCOHOL USE, CORONARY ARTERY DISEASE, PRIOR CORONARY ARTERY BYPASS GRAFT, VALVULAR HEART DISEASE, CARDIOMYOPATHY, OBSTRUCTIVE SLEEP APNEA, CHRONIC LUNG DISEASE. COMPLICATIONS REPORTED INCLUDED DEATH, HOSPITALIZATION, CARDIAC TAMPONADE, STROKE, BLEEDING, HEART FAILURE, EPISTAXIS, GASTROINTESTINAL BLEEDING, INTRACRANIAL BLEEDING, THROMBUS, DEVICE MIGRATION, PERI-DEVICE LEAK. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DEATH DATE IS ESTIMATED. B3: EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: SHORT- AND LONG-TERM OUTCOMES OF PERCUTANEOUS LEFT ATRIAL APPENDAGE OCCLUSION IN CANCER PATIENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22046176&pc=DSQ | 2916596-2025-02980 | 2025-04-25 04:00 | Death | THORATEC CORPORATION | 2025-05-19 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH A SOCIAL MEDIA POST THAT THE PATIENT PASSED AWAY. FURTHER INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. ATTEMPTS TO REACH THE CUSTOMER FOR A PRIVATE MESSAGE WITH ADDITIONAL INFORMATION HAD BEEN MADE. NO FURTHER INFORMATION WAS PROVIDED BY THE CUSTOMER. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. NO PRODUCT WAS RETURNED FOR EVALUATION. A REVIEW OF THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED AS THE SERIAL NUMBER OF THE DEVICE WAS NOT COMMUNICATED/IDENTIFIED. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A: PATIENT INFORMATION HAS NOT YET BEEN PROVIDED. SECTION B3: DATE OF EVENT WAS ESTIMATED. SECTION B2 DATE OF DEATH WAS NOT PROVIDED. SECTION D4: DEVICE SERIAL WAS NOT PROVIDED, SO THE EXPIRATION DATE CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22046375&pc=NPT | 2015691-2025-03948 | 2025-03-10 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-19 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem |  | P140031 |  | 1 | Event Description: AS REPORTED TO EDWARDS PATIENT SUPPORT, APPROXIMATELY 25 DAYS POST TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WITH A 29 MM SAPIEN 3 ULTRA RESILIA VALVE, THE PATIENT PASSED AWAY. IT WAS REPORTED THAT THE PATIENT HAD BEEN DISCHARGED HOME ON THE DAY AFTER THE TAVR PROCEDURE. APPROXIMATELY 2 DAYS LATER, THE PATIENT HAD A CARDIAC ARREST. THE PATIENT WAS ABLE TO BE RESUSCITATED AFTER CHEST COMPRESSIONS AND WAS TAKEN BACK TO THE IMPLANTING HOSPITAL. THE PATIENT REMAINED IN THE INTENSIVE CARE UNIT (ICU) BUT WAS NOT IN GOOD CONDITION AND WAS UNABLE TO RECOVER. PER THE DEATH CERTIFICATE, THE CAUSES OF DEATH WERE LISTED AS CARDIAC ARREST, PULMONARY EMBOLISM, ACUTE KIDNEY FAILURE AND THORACIC AORTIC DISSECTION. NO ALLEGATIONS WERE MADE REGARDING THE EDWARDS DEVICE. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR CORRECTION AND INCLUDES ADDITIONAL INFORMATION FROM THE PROVIDED MEDICAL RECORDS. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN CORRECTED/UPDATED: A2 (AGE AT TIME OF THE EVENT), B7 (OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS), H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS), H10 (RELATED REPORT NUMBERS) AND H11 (ADDITIONAL NARRATIVE/DATA). THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. PLEASE SEE MANUFACTURER REPORT NUMBER 2015691-2025-04685 FOR DETAILS OF THE OTHER EVENT. ADDITIONAL INFORMATION WAS RECEIVED VIA MEDICAL RECORDS REVEALING THE PATIENT UNDERWENT A SUCCESSFUL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WITH A 29 MM SAPIEN 3 ULTRA RESILIA VALVE AND THERE WAS NO PARAVALVULAR LEAK (PVL). IN ADDITION, THERE WERE NO DEVICE MALFUNCTIONS NOR PROCEDURAL ADVERSE EVENTS LISTED. THE PATIENT WAS DISCHARGED APPROXIMATELY 1 DAY POST TAVR PROCEDURE. SUBSEQUENTLY, APPROXIMATELY 3 DAYS POST DISCHARGE, THE PATIENT WAS READMITTED TO THE HOSPITAL FOR SHOCK SECONDARY TO A TYPE B AORTIC DISSECTION WITH CARDIAC ARREST. COMPLICATIONS OF THIS ADMISSION INCLUDED STERNAL FRACTURES FROM CARDIOPULMONARY RESUSCITATION (CPR), ACUTE RENAL FAILURE WITH CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT), HEPARIN-INDUCED THROMBOCYTOPENIA (HIT) WITH PULMONARY EMBOLISM (PE) AND DEEP VEIN THROMBOSIS (DVT) AND ACUTE HYPOXIC RESPIRATORY FAILURE ALL SECONDARY TO SHOCK. THE PATIENT WAS PLACED ON PALLIATIVE CARE AND PASSED AWAY APPROXIMATELY 25 DAYS POST TAVR PROCEDURE. THERE WAS NO ALLEGATION OR INDICATION THE TAVR PROCEDURE, NOR AN EDWARDS DEVICE CAUSED OR CONTRIBUTED TO THE EVENT. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE RELIES ON CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ALTHOUGH A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AT THIS TIME, THE EVENT MAY BE DUE TO THE MECHANISMS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22046407&pc=QZI | 2124215-2025-32068 | 2025-04-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-05-19 04:00 | QZI | FARAWAVE | Insufficient Information | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED A PATIENT DEATH OCCURRED FOLLOWING AN ATRIAL FIBRILLATION PULSE FIELD ABLATION PROCEDURE INVOLVING A FARAWAVE CATHETER. PRIOR TO DISCHARGE, THE PATIENT EXPIRED. NO DEVICE ISSUES OR MALFUNCTIONS WERE REPORTED. THE RELATIONSHIP OF THE DEATH TO THE PROCEDURE WAS UNKNOWN. NO FURTHER DETAILS PROVIDED. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS. Manufacturer Narrative: EXACT PROCEDURE AND DATE OF DEATH WERE UNKNOWN. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22046637&pc=OZD | 1220648-2025-28539 | 2025-04-30 04:00 | Death | ABIOMED, INC. | 2025-05-19 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS ON IMPELLA CP SUPPORT AND EXPERIENCED HEAVY BLEEDING. THE IMPELLA CP WAS WEANED AND EXPLANTED. THE PATIENT'S CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. ABIOMED'S MEDICAL SAFETY OFFICER REVIEWED THE EVENT AND DETERMINED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME. ADDITIONAL INFORMATION WAS RECEIVED. THE PATIENT CAME INTO THE CLINIC SPONTANEOUSLY BREATHING. AFTER FIRST EXAMINATION, IT WAS DECIDED TO DO A CARDIAC CATHETERIZATION. A PRONOUNCED OCCLUSION OF THE LEFT MAIN STEM WAS NOTED. THE PATIENT PRESENTED WITH AN EJECTION FRACTION OF 25% AND A SYSTOLIC BLOOD PRESSURE WELL BELOW 90. IT WAS DECIDED TO INTERVENE UNDER IMPELLA SUPPORT. DURING THE APPLICATION, THE PATIENT WAS IN NEED OF RESUSCITATION FOR ATRIAL FIBRILLATION. AFTER ABOUT TWO MINUTES, THE PATIENT HAD A RETURN OF SPONTANEOUS CIRCULATION. INTERVENTION WENT WELL UNDER IMPELLA PROTECTION. STENTING OF THE LEFT MAIN WAS SUCCESSFUL, AFTER WHICH THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT. THE PATIENT WAS VERY UNSTABLE AND DEVELOPED RESPIRATORY INSUFFICIENCY AND HAD TO BE INTUBATED AND VENTILATED. THE PATIENT HAD AN INCREASING NEED FOR VASOPRESSORS/INOTROPES. DIFFUSE BLEEDING WAS NOTED AS THERE WAS A DROP IN HEMOGLOBIN DUE TO LIVER FAILURE. THERE WAS ALSO BLEEDING FROM THE INSERTION SITE OF THE IMPELLA (DESPITE BEST PRACTICE) SO THE DECISION WAS MADE TO EXPLANT THE IMPELLA CP. THE PATIENT DIED IN THE COURSE OF THE AFTERNOON FROM SEVERITY OF THE SHOCK. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER. ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22046930&pc=NPT | 2015691-2025-03952 | 2025-04-29 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-19 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Endocarditis | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES AFFILIATE IN THE UNITED KINGDOM, A PATIENT WHO UNDERWENT A TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WITH A 26 MM SAPIEN 3 ULTRA RESILIA VALVE SUBSEQUENTLY DEVELOPED ENDOCARDITIS AND PASSED AWAY. THE ONSET DATE OF THE ENDOCARDITIS RELATIVE TO THE TAVR PROCEDURE IS UNKNOWN AT THIS TIME. THE DATE OF DEATH IS ALSO UNKNOWN AT THIS TIME. NO ADDITIONAL DETAILS HAVE BEEN PROVIDED. Manufacturer Narrative: THIS IS ONE OF THREE MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. PLEASE SEE MANUFACTURER REPORT NUMBERS 2015691-2025-01027 AND 2015691-2025-01031 FOR DETAILS OF THE OTHER EVENTS. PER THE INSTRUCTIONS FOR USE (IFU), INFECTIONS SUCH AS SEPTICEMIA AND ENDOCARDITIS ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. ENDOCARDITIS IS AN INFECTION OF A NATIVE OR PROSTHETIC VALVE, IS TREATED WITH ANTIBIOTICS, AND MAY REQUIRE VALVE REPLACEMENT IF ANTIBIOTIC THERAPY IS NOT EFFECTIVE. CAUSES OF PROSTHETIC VALVE ENDOCARDITIS ARE WELL DOCUMENTED IN THE LITERATURE AND ARE TYPICALLY CLASSIFIED AS EARLY (LESS THAN 60 DAYS) OR LATE (GREATER THAN 60 DAYS). THE TWO CATEGORIES SHOW DEFINITE DIFFERENCES IN CLINICAL FEATURES, MICROBIAL PATTERNS, AND MORTALITY. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, EARLY PROSTHETIC VALVE ENDOCARDITIS IS USUALLY CAUSED BY PERIOPERATIVE BACTERIAL CONTAMINATION OF THE VALVE. IN EARLY CASES OF PROSTHETIC ENDOCARDITIS, SUBSEQUENT INFECTIONS MAY BE RELATED TO CONTAMINATION AT THE TIME OF SURGERY (POOR STERILE TECHNIQUE). ADDITIONAL CAUSES OF EARLY PROSTHETIC VALVE ENDOCARDITIS (PVE) CAN OCCUR FROM OTHER SOURCES OF INFECTION AND NOT FROM THE VALVE (E.G., DENTAL TREATMENT OR SURGICAL PROCEDURES INVOLVING THE UPPER RESPIRATORY TRACT, URINARY TRACT INFECTIONS (UTIS), PNEUMONIA). ADDITIONALLY, A NON-CONFORMANCE IN THE STERILITY OR PACKAGING PROCESSES OF THE VALVE MAY ALSO MANIFEST IN THE EARLY POST-OPERATIVE PERIOD CAUSING EARLY PVE. HOWEVER, THERE ARE SEVERAL TYPES OF BACTERIA THAT WOULD NOT SURVIVE THE EDWARDS STERILIZATION PROCESS. EDWARDS LIFESCIENCES PRODUCES AND PROVIDES STERILE TISSUE BIOPROSTHESIS TO ITS CUSTOMERS BY FOLLOWING CAREFULLY DESIGNED ROBUST STERILIZATION PROCESSES. THESE MANUFACTURING PROCESSES HAVE BEEN VALIDATED AND DEMONSTRATED TO CONSISTENTLY PROVIDE A SIGNIFICANT SAFETY FACTOR FROM WHICH MICROORGANISMS COULD NOT SURVIVE. MICROBIOLOGY AND PROCESS MONITORING IS ROUTINELY REVIEWED WITHIN QUALITY SYSTEMS TO MAINTAIN STERILITY CONTROL. VALIDATED TESTING HAS DEMONSTRATED THAT MICROORGANISMS COULD NOT SURVIVE EDWARDS' MULTI-STAGE PROCESSING WITH ENHANCED STERILANT OR HEATED GLUTARALDEHYDE TERMINAL STERILANT SOLUTION. THESE MULTIPLE, REDUNDANT MANUFACTURING CONTROLS ENSURE THE STERILITY OF EDWARDS' VALVES AS PROVIDED TO CUSTOMERS. THEREFORE, THE PROBABILITY OF ENDOCARDITIS RELATED TO EDWARDS' BIOPROSTHESIS IS REMOTE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ALTHOUGH A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AT THIS TIME, THE EVENT MAY BE DUE TO THE MECHANISMS DESCRIBED ABOVE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22047078&pc=NPT | 2015691-2025-03956 | 2025-05-01 04:00 | Death | EDWARDS LIFESCIENCES | 2025-05-19 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Fluid/Blood Leak; Malposition of Device | Insufficient Information | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING IMPLANT OF A 29MM SAPIEN 3 ULTRA RESILIA (S3UR) AND ALTERRA IN THE AORTIC POSITION IN AN OFF-LABEL CASE VIA TRANSEPTAL (THROUGH THE FONTAN), THE ALTERRA SUCCESSFULLY LANDED IN THE AORTIC POSITION. THE TEAM SUCCESSFULLY DELIVERED THE S3UR INTO THE ALTERRA, HOWEVER THE WHOLE SYSTEM SEEMED TO MIGRATE LV AND THERE WAS SEVERE AORTIC INSUFFICIENCY (AI). A SECOND 29MM S3UR VALVE WAS PREPPED AND DELIVERED INTO THE FIRST S3UR AND ALTERRA. THE PATIENT WENT INTO PEA AND WAS NEVER ABLE TO RECOVER. ADDITIONAL INFORMATION RECEIVED NOTED THAT THE SEVERE AI WAS A PARASTENT LEAK. IT SEEMED PRESENT ON THE ANGIO POST VALVE DEPLOYMENT. THERE WAS NO VALVE DYSFUNCTION OF THE FIRST S3UR. THE SECOND S3UR WAS ONLY PLACED TO FIX THE AI THROUGH THE OPEN STRUTS OF THE ALTERRA AND THE FIRST S3UR.Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING IMPLANT OF A 29MM SAPIEN 3 ULTRA RESILIA (S3UR) AND ALTERRA IN THE AORTIC POSITION IN AN OFF-LABEL CASE VIA TRANSEPTAL (THROUGH THE FONTAN), THE ALTERRA SUCCESSFULLY DEPLOYED IN THE AORTIC POSITION. THE TEAM SUCCESSFULLY DELIVERED THE S3UR INTO THE ALTERRA, HOWEVER THE WHOLE SYSTEM SEEMED TO MIGRATE LV AND THERE WAS SEVERE AORTIC INSUFFICIENCY (AI). A SECOND 29MM S3UR VALVE WAS PREPPED AND DELIVERED INTO THE FIRST S3UR AND ALTERRA. THE PATIENT WENT INTO PEA AND WAS NEVER ABLE TO RECOVER. ADDITIONAL INFORMATION RECEIVED NOTED THAT THE SEVERE AI WAS A PARASTENT LEAK. IT SEEMED PRESENT ON THE ANGIO POST VALVE DEPLOYMENT. THERE WAS NO VALVE DYSFUNCTION OF THE FIRST S3UR. THE SECOND S3UR WAS ONLY PLACED TO FIX THE AI THROUGH THE OPEN STRUTS OF THE ALTERRA AND THE FIRST S3UR. ACCORDING TO THE MEDICAL RECORD, PATIENT WITH AN EF (27%) AND SEVERE AORTIC VALVE REGURGITATION WITH NO AORTIC ANNULAR CALCIUM UNDERWENT TREATMENT WITH AN ALTERRA PRESTENT AND A 29 MM SAPIEN 3. THE PROCEDURE WAS PERFORMED AS SALVAGE TAVR VALVE GIVEN THAT THE PATIENT HAD NO OTHER TREATMENT OPTIONS WITH SURGERY FOR VALVE REPLACEMENT AND HAD BEEN TURNED DOWN FOR TRANSPLANT EXTREME RISK. ALSO, THE ANNULUS WAS TOO LARGE FOR A JENA VALVE. A SECOND 29 MM +1 CC SAPIEN 3 VALVE WAS REQUIRED AND DEPLOYED DURING ACLS GIVEN THE VENTRICULAR APPEARANCE OF THE ALTERRA STENT AND FIRST SAPIEN VALVE. AN 'A-V RAIL SYSTEM' WAS ESTABLISHED THROUGH WHICH SUBSEQUENT EQUIPMENT WOULD BE ADVANCED THOUGH A SEPTOSTOMY. THE ALTERRA PRESTENT WAS PLACED IN THE AORTIC POSITION WITH NO ISSUES. THE 29 MM +1 CC VOLUME EDWARDS SAPIEN VALVE WITH AS PLACED WITHIN THE ALTERRA WHILE RAPIDLY PACING THE HEART (180 BPM). HOWEVER, THE PATIENT BECAME SIGNIFICANTLY HYPOTENSIVE AND RECOVERED WITH SOME CHEST COMPRESSIONS AND EPINEPHRINE BOLUSES. TEE SHOWED THAT THE VALVE AND DOCKING SYSTEM WERE EITHER DEPLOYED LOW (BELOW THE NATIVE AORTIC VALVE LEAFLETS) OR LIKELY MIGRATED MORE VENTRICULAR DURING THE RELEASE OF THE ALTERRA DOCKING STENT (THOUGH THIS WAS NOT APPRECIATED ON THE FLUORO). AT THIS POINT THE PATIENT BECAME HYPOTENSIVE AGAIN AND ACLS WAS INITIATED, MEDICATIONS WERE PROVIDED AND CHEST COMPRESSIONS PERFORMED. SIMULTANEOUSLY, THE SEPTOSTOMY WAS CLOSED WITH AN AMPLATZ 10MM SEPTAL OCCLUDE. ACLS WAS CONTINUED THROUGHOUT, A BRIEF INTERRUPTION OF CHEST COMPRESSIONS NOTED PULSELESS ELECTRICAL ACTIVITY (PEA). DURING THIS TIME AN AORTIC ROOT ANGIOGRAM WAS PERFORMED AND SHOWED A DEEPLY SEATED VALVE AND DOCKING SYSTEM. A PIGTAIL CATHETER WAS ADVANCED ACROSS THE RECENTLY DEPLOYED TAVR VALVE, AND A SECOND 29 MM +1 CC VOLUME EDWARDS SAPIEN VALVE WAS ADVANCED AND POSITIONED MORE PROXIMALLY WHERE IT WAS ULTIMATELY DEPLOYED WITHOUT COMPLICATION. THE PATIENT REMAINED IN PEA, ACLS WAS CONTINUED WITH USE OF IV MEDS WITH EPINEPHRINE AND ONGOING CHEST COMPRESSIONS. LAB WORK DEMONSTRATED APPROPRIATE SATURATIONS OF 97% AND NO EVIDENCE OF ACUTE BLOOD LOSS WITH STABLE HEMOGLOBIN AND HEMATOCRIT. CHEST COMPRESSIONS CONTINUED WHILE WE REVIEWED AND TREATED POTENTIAL REVERSIBLE CAUSES. ONGOING ASSESSMENT WITH TEE DEMONSTRATED CARDIAC STANDSTILL. NO READILY REVERSIBLE CAUSES WERE REVEALED. AFTER APPROXIMATELY 30 MINUTES OF ONGOING ACLS, THE PATIENT UNFORTUNATELY EXPIRED. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR COMPLETION TO THE ENGINEERING EVALUATION. THE PRODUCT WAS NOT RETURNED; THEREFORE, A NO PRODUCT RETURN INVESTIGATION WAS COMPLETED. AS A DEVICE WAS NOT RETURNED, VISUAL INSPECTION, FUNCTIONAL TESTING, AND DIMENSIONAL TESTING WERE UNABLE TO BE DONE. NO IMAGERY WAS PROVIDED. AS NO DEVICE WAS RETURNED AND THERE IS NO EVIDENCE TO SUPPORT A MANUFACTURING/DESIGN DEFECT POTENTIALLY CONTRIBUTED TO THE COMPLAINT, A MANUFACTURING MITIGATION REVIEW IS NOT REQUIRED. IT WAS NOTED THAT ALTERRA AND S3UR WERE IMPLANTED IN AORTIC VALVE POSITION VIA TRANSEPTAL APPROACH FOR THIS CASE WITH COMMANDER DELIVERY SYSTEM. ALTERRA IS INDICATED FOR USE IN PULMONARY POSITION PER IFU. THEREFORE, THIS CASE WAS AN OFF-LABEL OPERATION. AVAILABLE IFU/TRAINING MATERIALS ARE REVIEWED FOR INSTRUCTIONS THAT MAY BE RELEVANT TO THE COMPLAINT EVENT; HOWEVER, INSTRUCTIONS MAY NOT DESCRIBE THE ENTIRE PROCEDURE AS DESCRIBED PER THE CER. THE ALTERRA ADAPTIVE PRESTENT SYSTEM IFU WAS REVIEWED. PRECAUTIONS INCLUDE, 'CORRECT SIZING OF THE PRESTENT INTO THE RVOT IS ESSENTIAL TO MINIMIZE RISKS SUCH AS PARAVALVULAR LEAK, MIGRATION, EMBOLIZATION, AND/OR RVOT RUPTURE'. POTENTIAL ADVERSE EVENTS NOTE, 'POTENTIAL RISKS THAT MAY OR MAY NOT REQUIRE INTERVENTION ASSOCIATED WITH THE PRESTENT, DELIVERY SYSTEM AND/OR ACCESSORIES INCLUDE, BUT MAY NOT BE LIMITED TO, THE FOLLOWING: DEVICE MIGRATION OR MALPOSITION'. NO EVIDENCE OF PRODUCT NON-CONFORMANCES OR LABELING/IFU INADEQUACIES WERE IDENTIFIED IN THE EVALUATION.' THE COMPLAINTS FOR MALPOSITIONED PRESTENT AND DEPLOYED PRESTENT EXHIBITS PARA-STENT LEAK WERE CONFIRMED BASED ON PROVIDED MEDICAL RECORDS. AS THE DEVICE WAS NOT RETURNED, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL EXAMINATION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THEREFORE, THE PRESENCE OF A MANUFACTURING NON-CONFORMANCE WAS UNABLE TO BE DETERMINED. HOWEVER, REVIEW OF THE DHR AND LOT HISTORY DID NOT PROVIDE ANY INDICATION THAT A MANUFACTURING NON-CONFORMANCE WOULD HAVE CONTRIBUTED TO THE COMPLAINT EVENT. A REVIEW OF IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. FURTHERMORE, NO ABNORMALITIES WERE REPORTED DURING DEVICE UNPACKING OR PREPARATION. AS REPORTED, 'DURING IMPLANT OF A 29MM SAPIEN 3 ULTRA RESILIA (S3UR) AND ALTERRA IN THE AORTIC POSITION IN AN OFF-LABEL CASE VIA TRANSEPTAL (THROUGH THE FONTAN), THE ALTERRA SUCCESSFULLY LANDED IN THE AORTIC VALVE. THE TEAM SUCCESSFULLY DELIVERED THE S3UR INTO THE ALTERRA, HOWEVER THE WHOLE SYSTEM SEEMED TO MIGRATE LV AND THERE WAS SEVERE AORTIC INSUFFICIENCY (AI). ADDITIONAL INFORMATION RECEIVED NOTED THAT THE SEVERE AI WAS A PARASTENT LEAK. IT SEEMED PRESENT ON THE ANGIO POST VALVE DEPLOYMENT. THERE WAS NO VALVE DYSFUNCTION OF THE FIRST S3UR.' BASED ON THE MEDICAL RECORDS, THE PATIENT APPEARED TO HAVE LARGE ANNULUS WITH THE ABSENCE OF CALCIUM. IN THIS CASE, PATIENT FACTORS AS SUCH MAY HAVE CAUSED A CHALLENGING PATHWAY FOR THE DEVICE PLACEMENT WHICH HAS RESULTED IN THE MALPOSITION OF ALTERRA AND THV TO A LOWER POSITION (VENTRICULAR) DURING THE PROCEDURE. ADDITIONALLY, THE ALTERRA IS NOT INDICATED FOR USE IN THE AORTIC POSITION. IT IS POSSIBLE THAT A COMBINATION OF PATIENT FACTORS AND THE USE OF THE DEVICE IN A NON-INDICATED LOCATION MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CHALLENGING ANATOMY) AND OFF-LABEL USE (NON-INDICATED LOCATION) MAY HAVE CONTRIBUTED TO THE COMPLAINT EVENT. ADDITIONALLY, THE REPORTED PRESTENT MIGRATION/MALPOSITION COULD POTENTIALLY FURTHER LEAD TO PARA-STENT LEAK AS REPORTED DUE TO PRESTENT MIGRATION AS IT MAY COMPROMISE THE INTENDED FUNCTION OF A PRESTENT TO PROVIDE A PROPER SEAL AGAINST HEART TISSUE. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PROCEDURAL FACTORS (PRESTENT MALPOSITION) AND OFF-LABEL USE (NON-INDICATED LOCATION) MAY HAVE CONTRIBUTED TO THE COMPLAINT EVENT. PER THE ASSESSMENT, IT HAS BEEN DETERMINED THAT NO CAPA, SCAR, OR PRA IS REQUIRED AS THE ESTABLISHED TRIGGERS HAVE NOT BEEN MET. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. Manufacturer Narrative: PER ADDITIONAL CLINICAL REVIEW OF THE CASE BY CLINICAL PRODUCT QUALITY MANAGER, UPDATE TO B4, G3, G6, AND H2. CORRECTION TO B2. ADDED ADDITIONAL CODE TO H6 HEALTH EFFECT IMPACT CODE. INVESTIGATION IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO MEDICAL RECORDS BEING RECEIVED. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED OR CORRECTED: B4, B5, B7, G3, G6, H2, H6, H11. THE INVESTIGATION IS ONGOING. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22047104&pc=OZD | 1220648-2025-28543 | 2025-04-28 04:00 | Death | ABIOMED, INC | 2025-05-19 04:00 | OZD | IMPELLA CP | Malposition of Device; Detachment of Device or Device Component | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP HAD OOZING AT THE ACCESS SITE DESPITE ANGLE MATCHING AND A PURSE STRING SUTURE. A FEMSTOP WAS PLACED AND BLOOD PRODUCTS WERE TRANSFUSED TO OBTAIN HEMOSTASIS. ADDITIONALLY, THE PUMP WAS REPOSITIONED DUE TO ALARMS AND THE ANTICONTAMINATION SLEEVE WAS TORN. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING, PRODUCT DAMAGE ( SLEEVE DAMAGE), AND POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. AS PER THE CLINICAL INFORMATION PROVIDED, THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY ANTICOAGULATION MANAGEMENT RELATED DUE TO HIGH PATIENT ACTIVATED CLOTTING TIME (ACT) VALUES. POSITIONING ISSUE: CLINICAL STATES THAT THERE WERE "IMPELLA IN VENTRICLE" ALARMS OVERNIGHT AND THE PUMP WAS REPOSITIONED WHICH HELPED TROUBLESHOOT THE ALARMS. PUMP WAS NOT RETURNED. THE REASON FOR THE PUMP POSITION CHANGE IS NOT CLEAR. THEREFORE, THE ROOT CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY USE RELATED AS POSITION ISSUE RESOLVED WITH TROUBLESHOOTING DURING SUPPORT. PRODUCT DAMAGE (SLEEVE): CLINICAL STATES THAT THE STERILE SLEEVE WAS TORN ON (B)(6) 2025 BUT NO OTHER SUPPORTING INFORMATION IS PROVIDED. PUMP WAS NOT RETURNED. THEREFORE, THE ROOT CAUSE OF THE PRODUCT DAMAGE - SLEEVE ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED FOR INVESTIGATION AND INSUFFICIENT CLINICAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22047132&pc=DYB | 2015691-2025-03958 | 2025-04-24 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-05-19 04:00 | DYB | EDWARDS ESHEATH+ INTRODUCER SET | Physical Resistance/Sticking | Ischemia; Perforation; Perforation of Vessels; Stenosis; Thromboembolism | K200258 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES JAPAN ASSOCIATE, DURING A LEFT TRANSFEMORAL TAVR PROCEDURE WITH A 23 MM SAPIEN 3 ULTRA RESILIA, INABILITY OF THE DELIVERY SYSTEM TO ADVANCE THROUGH THE SHEATH, AND INJURY IN THE ACCESS VESSEL OCCURRED. INITIALLY, THE TAVR PROCEDURE WAS PERFORMED VIA LEFT TRANSFEMORAL APPROACH. THE 14 FR ESHEATH+ WAS INSERTED WITHOUT RESISTANCE. THE 23 MM DELIVERY SYSTEM WAS INSERTED INTO THE SHEATH AND ADVANCED, BUT THE DELIVERY SYSTEM WAS UNABLE TO PASS THROUGH THE SHEATH DESPITE SEVERAL ATTEMPTS. THE PHYSICIAN DECIDED TO CONVERT TO A LEFT SUBCLAVIAN ARTERY APPROACH. A SECOND KIT WAS OPENED, AND A 23 MM SAPIEN 3 ULTRA RESILIA WAS IMPLANTED WITHOUT ANY ISSUES. THERE WAS VASCULAR RUPTURE IN THE LEFT GROIN WHERE THE SHEATH INITIAL INSERTION WAS ATTEMPTED, AND SURGICAL REPAIR WAS PERFORMED. ANGIOGRAPHY REVEALED THROMBOEMBOLISM IN THE VESSEL DISTAL TO THE KNEE, SO THROMBECTOMY WAS ADDITIONALLY PERFORMED, SUCCESSFULLY ACHIEVING RECANALIZATION. ON THE WEEKEND OF TAVR PROCEDURE (2 OR 3 DAYS LATER), THROMBECTOMY WAS PERFORMED AGAIN, BUT THE PATIENT PASSED AWAY. PER MEDICAL OPINION, THE POOR ELASTIC VESSELS RESULTING FROM DIALYSIS USER MAY HAVE CONTRIBUTED INABILITY OF THE DELIVERY SYSTEM TO ADVANCE THROUGH THE SHEATH. NO FURTHER INFORMATION WAS AVAILABLE.Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES JAPAN ASSOCIATE AND SUPPLEMENTED BY INFORMATION FROM THE JAPANESE TAVI REGISTRY, THE PATIENT UNDERWENT A LEFT TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE USING A 23 MM SAPIEN 3 ULTRA RESILIA VALVE. THE INITIAL APPROACH INVOLVED INSERTION OF A 14 FR ESHEATH+ WITHOUT RESISTANCE. HOWEVER, THE 23 MM DELIVERY SYSTEM WAS UNABLE TO ADVANCE THROUGH THE SHEATH DESPITE SEVERAL ATTEMPTS, PROMPTING THE PHYSICIAN TO SWITCH TO A LEFT SUBCLAVIAN APPROACH. A SECOND DEVICE KIT WAS OPENED, AND THE VALVE WAS SUCCESSFULLY IMPLANTED VIA THE SUBCLAVIAN ARTERY WITHOUT FURTHER TECHNICAL DIFFICULTIES. DURING THE ATTEMPTED TRANSFEMORAL ACCESS, A VASCULAR RUPTURE OCCURRED IN THE LEFT COMMON FEMORAL ARTERY (CFA), WHICH REQUIRED SURGICAL REPAIR. POST-DEPLOYMENT ANGIOGRAPHY REVEALED THROMBOEMBOLISM IN THE VESSEL DISTAL TO THE KNEE, FOR WHICH ENDOVASCULAR TREATMENT (EVT), INCLUDING THROMBECTOMY, WAS PERFORMED WITH SUCCESSFUL RECANALIZATION. ON POSTOPERATIVE DAY (POD) 2, THE PATIENT DEVELOPED ACUTE OCCLUSION OF THE LEFT LOWER LIMB ARTERY AND UNDERWENT EMERGENCY THROMBECTOMY AND EVT. HOWEVER, LIMB ISCHEMIA PERSISTED. ON POD 3, CONTRAST-ENHANCED COMPUTED TOMOGRAPHY SHOWED STENOSIS OF THE LEFT SUPERFICIAL FEMORAL ARTERY (SFA) AND CFA. EMERGENCY EVT WAS AGAIN PERFORMED WHILE THE PATIENT WAS ON CONTINUOUS HEMODIAFILTRATION. ALTHOUGH PARTIAL IMPROVEMENT IN BLOOD FLOW WAS OBSERVED, THE PATIENT DEVELOPED WORSENING ACIDOSIS AND REFRACTORY HYPOTENSION, ULTIMATELY RESULTING IN DEATH ON POD 3. THE CAUSE OF DEATH WAS REPORTED AS ACUTE LOWER LIMB ARTERIAL OCCLUSION. PER THE ATTENDING MEDICAL TEAM, THE PATIENT'S LONG-TERM DIALYSIS LIKELY CONTRIBUTED TO VASCULAR FRAGILITY, WHICH MAY HAVE BEEN A FACTOR IN BOTH THE DIFFICULTY ADVANCING THE DELIVERY SYSTEM AND THE SUBSEQUENT VASCULAR COMPLICATIONS. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED, FOLLOWING COMPLETION OF THE ENGINEERING EVALUATION. B4, G3, G6, AND H2 HAVE BEEN UPDATED. H6: HEALTH EFFECT - CLINICAL, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS, AND H11 HAVE BEEN CORRECTED. THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT AND ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. IMAGERY OF THE PATIENT'S ANATOMY WAS PROVIDED AND REVEALED THE PRESENCE OF CALCIFICATION AND TORTUOSITY IN THE PATIENT'S LEFT ACCESS VESSEL. PER THE INSTRUCTIONS FOR USE (IFU), POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE INCLUDE THROMBUS FORMATION, PLAQUE DISLODGMENT, AND EMBOLIZATION THAT MAY RESULT IN MYOCARDIAL INFARCTION, STROKE, DISTAL PERIPHERAL OCCLUSION, AND/OR DEATH. THROMBOEMBOLISM IS THE OBSTRUCTION OF A BLOOD VESSEL BY A BLOOD CLOT THAT HAS BECOME DISLODGED FROM ANOTHER SITE IN THE CIRCULATION. IT IS THE NATURAL TENDENCY OF THE BODY TO FORM A CLOT ON FOREIGN OBJECTS IN THE VASCULAR SPACE. THESE PATIENTS ARE ANTICOAGULATED FOR THE PROCEDURE AND INTERVENTIONAL BEST PRACTICES MANDATE METICULOUS WIPING AND FLUSHING OF THE DEVICES TO PREVENT AND/OR REMOVE THE CLOT. DISPLACEMENT OF CALCIUM DEPOSITS OR ATHEROMA WITH EMBOLIZATION OF DEBRIS CAN ALSO TRIGGER THE FORMATION OF A THROMBUS WHICH CAN THEN EMBOLIZE AND CAUSE BLOCKAGE IN A SMALLER VESSEL. CALCIFICATION AND ATHEROMA (CELLULAR DEBRIS, INFLAMMATORY CELLS, AND CHOLESTEROL) CAN ACCUMULATE ALONG THE WALLS OF THE VASCULATURE AND WITHIN CARDIAC STRUCTURES AND CAN VARY IN SIZE. THERE MAY BE CASES WHERE A PATIENT HAS A PROTRUDING ATHEROMA OR CALCIFIED PLAQUE PRIOR TO THE PROCEDURE. IT IS ALSO POSSIBLE FOR ONE OF THE INTERVENTIONAL DEVICES USED DURING THE THV PROCEDURE TO DISRUPT THE MATERIAL, RESULTING IN MOBILE DEBRIS. ANNULAR STRUCTURE TISSUE OR CALCIFICATION CAN ALSO BE DISRUPTED DURING SHEATH INSERTION, BALLOON VALVULOPLASTY, AND DEPLOYMENT OF THE PROSTHETIC VALVE. THE DEVICE TRAINING MANUALS AND IFU INSTRUCT THE OPERATOR TO ADMINISTER HEPARIN AND MAINTAIN THE ACT AT = 250 SEC. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THE REPORTED THROMBOEMBOLISM. THE EXACT CAUSE OF THE REPORTED EVENT CANNOT BE DETERMINED. HOWEVER, IN ADDITION TO THE MECHANISMS DESCRIBED ABOVE, ANOTHER VASCULAR COMPLICATION (RUPTURE IN THE LEFT COMMON FEMORAL ARTERY) MAY HAVE CONTRIBUTED TO THE EVENT. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH AND POTENTIAL VALVE FRAME DAMAGE USING THE S3U/S3UR VALVE CONFIGURATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, UNDERSIZED VESSELS, AND/OR STEEP INSERTION ANGLE. IN ADDITION, VALVE FRAME AND/OR SHEATH DAMAGE CAN BE A RESULT OF INCREASED PUSH FORCE AND ANY EXCESSIVE DEVICE MANIPULATION OR SHEATH-VALVE INTERACTION. IN THIS CASE, THE INABILITY TO ADVANCE THROUGH THE SHEATH WAS UNABLE TO BE CONFIRMED DUE TO UNAVAILABILITY OF RETURNED DEVICE/APPLICABLE IMAGERY. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION, TORTUOSITY) MAY HAVE CONTRIBUTED TO THE REPORTED EVENT, AS CONFIRMED VIA IMAGERY OF THE PATIENT'S ANATOMY. CALCIFICATION CAN REDUCE THE VESSEL LUMEN DIAMETER AND MAY INCREASE RESTRICTION LEADING TO RESISTANCE. CALCIFICATION CAN ALSO RESULT IN THE CREATION OF SUB-OPTIMAL ANGLES DURING DELIVERY SYSTEM INSERTION THAT MAY LEAD TO RESISTANCE. SCRATCHES OBSERVED ON THE SHEATH SHAFT CAN BE INDICATIVE OF THE PRESENCE OF CALCIFIED NODULES WITHIN THE ACCESS VESSEL. TORTUOUS PATIENT ANATOMY CAN CREATE SUB-OPTIMAL ANGLES THAT CAN LEAD TO NON-COAXIAL ALIGNMENT BETWEEN THE DELIVERY SYSTEM WITH CRIMPED VALVE AND SHEATH INNER LUMEN. KINKS OR CURVATURE ALONG THE SHEATH SHAFT CAN BE INDICATIVE OF THE PRESENCE OF VESSEL TORTUOSITY. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AN INVESTIGATION HAS BEEN OPENED TO DETERMINE THE ROOT CAUSE AND IMPLEMENT ANY NECESSARY CORRECTIVE ACTIONS. Manufacturer Narrative: THE INVESTIGATION FOR THIS EVENT IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22047422&pc=CAW | 3013095415-2025-00443 | 2025-04-16 04:00 | Death | INVACARE FLORIDA | 2025-05-19 04:00 | CAW | INVACARE® PLATINUM® OXYGEN CONCENTRATOR | Insufficient Information | Insufficient Information | K160630 |  | 1 | Event Description: THE DEVICE'S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THE FOLLOWING EVENT: "CONCENTRATOR STOPPED WORKING, WAS ALARMING, TECH ARRIVED WITHIN 2HRS AND PATIENT HAD DIED. FAMILY MEMBER WOULD NOT LET THE TECH TAKE THE CONCENTRATOR. TECH WAS ABLE TO CHECK THE CONCENTRATOR, HE MENTIONED WHEN HE TOOK THE TUBING OFF THE CONCENTRATOR WOULD RUN BUT WHEN HE PUT THE TUBING ON IT WOULD SHUT DOWN AND START ALARMING. THERE WAS NO HUMIDIFIER BOTTLE ATTACHED TO THE CONCENTRATOR." THE REPORTED EVENT INVOLVED AN INVACARE® PLATINUM® OXYGEN CONCENTRATOR. NO FURTHER INFORMATION ABOUT THE PATIENT WAS PROVIDED. Manufacturer Narrative: H6: VENTEC HAS DETERMINED THAT THIS DEVICE WAS PLACED INTO COMMERCIAL DISTRIBUTION BACK IN AUGUST OF 2014, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE ¿NONE.¿ VENTEC SPOKE TO THE INITIAL REPORTER (DISTRIBUTOR) WHO ADVISED THAT THEY HAVE THE TUBING AND THE O2 CONCENTRATOR IN THEIR POSSESSION. THE DISTRIBUTOR FURTHER ADVISED THAT THEY WOULD NOT BE SENDING IT TO VENTEC (REACT HEALTH) FOR AN EVALUATION. THE DISTRIBUTOR STATED THAT IF THAT CHANGES IN THE FUTURE, THEY WOULD CONTACT VENTEC AND REQUEST A RETURN AUTHORIZATION (RMA). THE DEVICE WAS NOT RETURNED TO VENTEC FOR EVALUATION. VENTEC HAS DETERMINED THAT THERE IS INSUFFICIENT INFORMATION TO DETERMINE IF THE DEVICE USE CONTRIBUTED TO THE PATIENT'S OUTCOME. THE CAUSE OF THE REPORTED ISSUE COULD NOT BE DETERMINED. IF VENTEC RECEIVES ANY NEW INFORMATION REGARDING THIS EVENT, A FOLLOW-UP REPORT SHALL BE SUBMITTED AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22047563&pc=FKX | 0008030665-2025-01126 | 2025-04-23 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-05-19 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181108 |  | 1 | Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS CUSTOMER SERVICE THIS PD PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED WHILE CONNECTED TO THE CYCLER. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW-UP WITH THE PATIENT¿S PDRN, IT WAS REPORTED THIS PATIENT EXPIRED AT HOME ON (B)(6) 2025 DUE TO AN UNREPORTED CAUSE. IT WAS AFFIRMED THAT THE PATIENT WAS CONNECTED TO THEIR LIBERTY SELECT CYCLER AT THE TIME OF DEATH. THE PATIENT WAS FOUND BY FAMILY TO BE UNRESPONSIVE AND PULSELESS ON THE MORNING OF (B)(6) 2025. EMERGENCY SERVICES WERE ACTIVATED BUT THE PATIENT WAS NOT TRANSPORTED AND PRONOUNCED DECEASED IN HOME. THE CAUSE OF DEATH WAS NOT REPORTED TO THE OUTPATIENT CLINIC; HOWEVER, THE PATIENT HAD A RECENT DECLINE IN HEALTH INVOLVING ADVANCED DEMENTIA THAT PREEXISTED HIS START ON PD THERAPY IN (B)(6) 2025. IT WAS CONFIRMED, THOUGH THE CAUSE OF DEATH WAS NOT REPORTED TO THE CLINIC, THE PATIENT¿S DEATH WAS DETERMINED TO BE UNRELATED TO PD THERAPY AND NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. THE CAUSE OF THIS PATIENT¿S DEATH CANNOT BE DETERMINED; HOWEVER, THERE WAS NO INDICATION THE PATIENT¿S DEATH WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22047674&pc=NIK | 2017865-2025-64064 | 2025-02-18 05:00 | Death | ABBOTT | 2025-05-19 04:00 | NIK | QUADRA ASSURA ICD | Inappropriate/Inadequate Shock/Stimulation | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE INAPPROPRIATE HIGH-VOLTAGE (HV) THERAPY WAS DELIVERED BY THE DEVICE. TECHNICAL SUPPORT WAS CONTACTED, BUT NO INTERVENTION HAS BEEN PERFORMED AT THIS TIME. THE PATIENT WAS STABLE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEVICE WAS PERFORMING AS PROGRAMMED. ADDITIONALLY, THE PATIENT PASSED AWAY DUE TO HEART FAILURE. NO ABBOTT PRODUCTS WERE ALLEGED TO HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22047690&pc=DSQ | 2916596-2025-03180 | 2025-04-29 04:00 | Death | THORATEC CORPORATION | 2025-05-19 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT PASSED AWAY, AND THE CAUSE OF DEATH WAS NOT KNOWN. DUE TO HOSPITAL POLICY, NO FURTHER INFORMATION WAS PROVIDED. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. IT WAS UNKNOWN WHETHER THE PUMP WAS EXPLANTED OR RETURNED. TO DATE, THE LVAS HAS NOT RETURNED TO ABBOTT FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22047697&pc=BZD | 2518422-2025-045062 | 2024-02-01 05:00 | Death | RESPIRONICS, INC. | 2025-05-19 04:00 | BZD | REMSTAR PRO C-FLEX+ | Degraded | No Clinical Signs, Symptoms or Conditions | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALLEGED DEVICE TURNS OFF RANDOMLY. THE PATIENT WAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE CUSTOMER ALLEGED DEVICE TURNS OFF RANDOMLY. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. AFTER MULTIPLE ATTEMPTS, THE CUSTOMER CONFIRMED THAT THE DEVICE DID NOT CONTRIBUTE TO THE PATIENT'S DEATH. THERE WAS NO PATIENT HARM OR INJURY ASSOCIATED WITH THIS REPORT AND NO INCREASED RISK TO THE INTENDED USER. BASED ON THE INFORMATION AVAILABLE, IT IS A NON-REPORTABLE COMPLAINT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22047809&pc=OZD | 1220648-2025-28546 | 2025-04-22 04:00 | Death | ABIOMED, INC. | 2025-05-19 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Failure to Pump; Pumping Stopped |  | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN CARDIOGENIC SHOCK FROM AN ACUTE MYOCARDIAL INFARCTION WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND WOULD EXPERIENCE PUMP FAILURE DUE TO AN ACCIDENT WITH THE AUTOMATED IMPELLA CONTROLLER. THE PATIENT WOULD EVENTUALLY EXPIRE. THE IMPELLA CP WAS SUCCESSFULLY PLACED, AND PERCUTANEOUS CORONARY INTERVENTION WAS COMPLETED VIA A DRUG-ELUTING STENT TO THE LEFT ANTERIOR DESCENDING ARTERY. THAT EVENING AN ECHOCARDIOGRAM WAS COMPLETED WITH PUMP AT "GOOD" POSITION, 3.3CM FROM AORTIC VALVE. THE NEXT MORNING DISCUSSION WAS MADE WITH CARDIOLOGIST AND CARDIOVASCULAR AND THORACIC SURGEON, AND THE PLAN WAS NOTED TO ESCALATE TO 5.5. THE REQUEST WAS MADE AND ACCEPTED TO TRANSFER THE PATIENT TO ANOTHER HOSPITAL LATER IN THE DAY. THE TRANSFER WAS COMPLETED BY FLIGHT CREW. THE AIC WAS ATTACHED TO SIDE OF STRETCHER AND WHEN ENTERING THE ELEVATOR, THE BLUE PORTION OF THE IMPELLA PLUG HIT THE ELEVATOR DOOR AND BROKE OFF. PUMP FAILURE OCCURRED; SUPPORT LEVEL DIMINISHED TO P-0. THE PATIENT WAS IN CRITICAL CONDITION. CONSEQUENTLY, THE AIC WAS EXCHANGED AND THE IMPELLA CP DEVICE PUMP 1 WAS EXPLANTED AND REPLACED WITH ANOTHER IMPELLA CP DEVICE PUMP 2 FOR CONTINUATION OF THERAPY. ON THE REPLACEMENT IMPELLA CP DEVICE PUMP 2, THE PATIENT WOULD DEVELOP LIMB ISCHEMIA UNDERGOING A FEM-FEM BYPASS AND EVENTUALLY EXPIRE THREE DAYS THEREAFTER. CARE WAS WITHDRAWN. Manufacturer Narrative: MEDICAL SAFETY REVIEWED THE EVENT AND NOTED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE AUTOMATED IMPELLA CONTROLLER (AIC) AS AN ASSOCIATED FACTOR TO THE OUTCOME (PATIENT'S DEMISE). THE PUMP STOP/FAILURE FOR THE IMPELLA CP PUMP 1 IS A SERIOUS INJURY AS WELL AS IMPELLA CP PUMP 2 (REPLACEMENT PUMP) WHERE THE PATIENT DEVELOPED LIMB ISCHEMIA. LIMB ISCHEMIA CAN LEAD TO MORE COMPLICATIONS DOWN THE LINE, BUT THE IMMEDIATE ISSUE IS THE AIC PUMP FAILURE (AND THIS IS WHAT LED TO OR CONTRIBUTED MORE TO THE DEMISE OUTCOME). THE AIC IS ONE OF THREE DEVICES ASSOCIATED WITH THE EVENT. NOTE THE FOLLOWING: REFER TO INITIAL MEDICAL DEVICE REPORT FOR THE IMPELLA CP PUMP 1 SERIAL NUMBER (B)(6) WITH DATE OF EVENT 22-APR-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). REFER TO MANUFACTURING REPORT NUMBER 1220648-2025-28324 FOR THE IMPELLA CP PUMP 2 (REPLACEMENT PUMP) REPORT. THE AUTOMATED IMPELLA CONTROLLER DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22047952&pc=KDJ | 1416980-2025-02921 | 2025-04-24 04:00 | Death | VANTIVE US HEALTHCARE LLC | 2025-05-19 04:00 | KDJ | NI | Adverse Event Without Identified Device or Use Problem | Peritonitis | NI |  | 1 | Event Description: A PERITONEAL DIALYSIS (PD) PATIENT EXPERIENCED PERITONITIS MANIFESTED BY CLOUDY PD EFFLUENT. THE CAUSE OF PERITONITIS WAS UNKNOWN. THE SAME DAY AS EVENT ONSET, THE PATIENT WAS HOSPITALIZED AND TREATED WITH VANCOMYCIN INJECTION (1GM, EVERY THREE DAYS, INTRAPERITONEAL, DISCONTINUED) FOR PERITONITIS. THE PATIENT WAS DISCHARGED SIX DAYS AFTER HOSPITAL ADMISSION. AT THE TIME OF THIS REPORT, THE PATIENT WAS RECOVERING FROM PERITONITIS. PD THERAPY WAS ONGOING. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: B2, B5, B6, H1, H6. B5: UPON FOLLOW UP, THE PATIENT WAS TREATED WITH MEROPENEM (1GM, ONCE DAILY, INTRAPERITONEAL, DISCONTINUED). THE VANCOMYCIN WAS DISCONTINUED AFTER 4 DAYS. IT WAS REPORTED THAT THE PATIENT DIED DUE TO THE SECOND EPISODE OF PERITONITIS AND WEAKNESS. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. THE PD THERAPY WAS ONGOING TILL DEATH. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED AND THE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22048556&pc=PNJ | 9612164-2025-02363 | 2025-04-22 04:00 | Death | MEDTRONIC IRELAND | 2025-05-19 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Dyspnea; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED POST OPERATIVELY. IT WAS FURTHER NOTED THAT THE PATIENT EXPERIENCED CARDIAC ARREST, CARDIAC PERFORATION PERICARDIAL EFFUSION CARDIAC TAMPONADE, SHORTNESS OF BREATH AND PULSELESS ELECTRICAL ACTIVITY.Event Description: IT WAS FURTHER REPORTED THAT THE AUTOPSY REPORT NOTED THAT THE TINES OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) WERE OUTSIDE THE HEART. Manufacturer Narrative: CORRECTED: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MC2AVR1] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22048756&pc=BZD | 2518422-2025-105862 | 2025-05-16 04:00 | Death | RESPIRONICS, INC. | 2025-05-19 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED TO THE MANUFACTURER AND INSPECTED. NO ERRORS WERE FOUND; THEREFORE, NO COMPONENTS WERE REPLACED TO CORRECT A MALFUNCTION. THE DEVICE PASSED ALL TESTS AND WAS RETURNED TO STOCK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22048936&pc=MNS | 2518422-2025-105865 | 2025-05-16 04:00 | Death | RESPIRONICS, INC. | 2025-05-19 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22049035&pc=LYJ | 1644487-2025-10149 | 2025-03-07 05:00 | Death | HOUSTON | 2025-05-19 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DIED DUE TO CARDIAC ISSUES. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22049444&pc=NAY | 2955842-2025-21034 | 2024-12-07 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-05-19 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Thrombocytopenia | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN UNEVENTFUL PANCREATODUODENECTOMY PROCEDURE, THE PATIENT DEVELOPED A COAGULATION DISORDER AND EXPIRED. THE CUSTOMER REPORTED THAT SOME MINOR, ANTICIPATED BLEEDING OCCURRED DURING THE PROCEDURE. THERE WERE NO INTRAOPERATIVE COMPLICATIONS AND NO DA VINCI PRODUCT MALFUNCTIONS OR ISSUES CONTRIBUTED TO THE PATIENT¿S DEATH. IT WAS REPORTED THAT ON POST-OPERATIVE DAY 2, THE PATIENT EXHIBITED HYPERACTIVE DELIRIUM WITH PSYCHOMOTOR AGITATION THAT REQUIRED ADMINISTRATION OF ANTIPSYCHOTIC MEDICATIONS. ON POST-OPERATIVE DAY 4, THE PATIENT DEVELOPED THROMBOCYTOPENIA, EVIDENCED BY A PLATELET COUNT OF 90,000/MM3, ALONG WITH DIFFUSE ABDOMINAL AND "RESIDUAL PANCREATIC PARENCHYMA" BLEEDING. THE SURGEON BELIEVES THE COAGULATION DISORDER DEVELOPED AS A SECONDARY EFFECT OF THE ANTIPSYCHOTIC MEDICATIONS. THE PATIENT SUBSEQUENTLY EXPERIENCED CARDIAC ARREST, AND AN EMERGENT EXPLORATORY LAPAROTOMY WAS PERFORMED. THE SEQUENCE OF EVENTS LEADING TO THE PATIENT DEATH ON POST-OPERATIVE DAY 6 IS UNKNOWN. AN AUTOPSY WAS NOT PERFORMED. THE CAUSE OF DEATH RECORDED ON THE DEATH CERTIFICATE WAS HYPOVOLEMIC SHOCK.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO RELEVANT SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE, PERMANENT CAUTERY HOOK, CADIERE FORCEPS, MARYLAND BIPOLAR FORCEPS, LARGE NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCE'S DOCUMENTED FOR THE DEVICES THAT WOULD BE RELATED TO THIS EVENT. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FROM HYPOVOLEMIC SHOCK SEVERAL DAYS FOLLOWING A PANCREATICODUODENECTOMY. POTENTIAL CONTRIBUTING FACTORS REPORTED INCLUDE A COAGULATION DISORDER WHICH WAS ATTRIBUTED TO MEDICATIONS. THE PATIENT DEVELOPED DIFFUSE ABDOMINAL BLEEDING. HOW THIS WAS DISCOVERED WAS NOT PROVIDED. NO OTHER POTENTIAL SOURCES WERE PROVIDED. THE PATIENT WENT INTO CARDIAC ARREST AND AN EMERGENT EXPLORATION WAS PERFORMED. THE DETAILS OF THE FINDINGS WERE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. Manufacturer Narrative: SECTION B5 ADDITIONAL INFORMATION: PRIOR TO THE PATIENT'S CARDIAC ARREST, BLOOD WAS NOTED IN THE SURGICAL DRAIN; THERE WAS NO BLOOD IN THE DRAIN EARLIER IN THE DAY. AN EMERGENCY EXPLORATORY LAPAROTOMY WAS PERFORMED, REVEALING A HEMOPERITONEUM. THE BLEEDING SOURCE WAS IDENTIFIED AT THE JUNCTION BETWEEN THE PANCREATIC HEAD AND BODY, SPECIFICALLY AT THE SECTIONED RAW TRANSITION AREA OF THE "RESIDUAL PANCREATIC PARENCHYMA". THIS HEMORRHAGE WAS ATTRIBUTED TO THROMBOCYTOPENIA. A REVIEW OF THE ADDITIONAL EVENT INFORMATION PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT DIED FOLLOWING A LATE BLEEDING EVENT SEVERAL DAYS AFTER A PANCREATICODUODENECTOMY. THE LOCATION OF THE BLEED WAS NOTED TO BE THE RAW SURFACE OF THE PANCREATIC TRANSECTION. THE CAUSE OF THE BLEED IS UNCLEAR ALTHOUGH IT WAS NOTED THE PATIENT HAD BECOME THROMBOCYTOPENIC AND MAY HAVE DEVELOPED COAGULOPATHY. THE REASON FOR THE COAGULOPATHY IS UNCLEAR. NO ALLEGATION IS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22041352&pc=BZD | 2518422-2025-045014 | 2024-07-26 04:00 | Death | RESPIRONICS, INC. | 2025-05-18 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY AND WIFE REPORTED THAT SHE RECEIVED A LETTER STATING TO RETURN THE RECALLED DEVICE. THERE WAS A REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND VISIBLE FOAM PARTICLES WERE NOT OBSERVED. ADDITIONALLY, THE PATIENT¿S COMPLAINT WAS NOT CONFIRMED, AND THE DEVICE WAS SCRAPPED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: IN PREVIOUS REPORT, THE MANUFACTURER CAPTURED INCORRECT INFORMATION IN SECTION B, IN SECTION D AND IN SECTION H. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. IN SECTION B: ADVERSE EVENT/PRODUCT PROBLEM AND OUTCOMES ATTRIBUTED TO AE HAS BEEN CORRECTED. IN SECTION D: PRODUCT UDI HAS BEEN UPDATED. IN SECTION H: COMPONENT CODE GRID HAS BEEN CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22041095&pc=BZD | 2518422-2025-105811 | 2025-04-09 04:00 | Death | RESPIRONICS, INC. | 2025-05-17 04:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A NOTIFICATION RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.