## Sheet1
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22191438&pc=BZD | 2518422-2025-106553 | 2025-06-06 04:00 | Death | RESPIRONICS, INC. | 2025-06-12 04:00 | BZD | REMSTAR PRO C-FLEX+ | No Apparent Adverse Event | Insufficient Information | K091319 |  | 1 | Event Description: THE REPORTER, WHO IS THE RELATIVE OF THE PATIENT, REPORTED THE PATIENT EXPIRED. THE REPORTER STATED THE PATIENT PREVIOUSLY USED; HOWEVER, THE PATIENT WAS NOT USING THE DEVICE AT THE TIME OF DEATH. THE PATIENT WAS SWITCHED TO A DIFFERENT (UNSPECIFIED) DEVICE. THE REPORTER HAD TO MAKE ANOTHER CALL AND WAS UNABLE TO ANSWER ANY FURTHER QUESTIONS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22191855&pc=OZP | 2032227-2025-189979 | 2025-05-13 04:00 | Death | MEDTRONIC MINIMED | 2025-06-12 04:00 | OZP | UNKNOWN NGP | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia |  |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS DIABETES. NO CONTRIBUTING FACTORS WERE IDENTIFIED FOR OTHER RELEVANT HISTORY, INCLUDING PREEXISTING MEDICAL CONDITIONS. THE PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: UNK_NGP, MMT-332A, UNOMEDICAL. PRODUCT RETURN WAS REQUIRED FOR UNK_NGP. PRODUCT RETURN WAS REQUIRED FOR MMT-332A. PRODUCT RETURN WAS REQUIRED FOR MMT-378. Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. SELECT PATIENT INFORMATION CANNOT BE PROVIDED DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22191996&pc=FRN | 2016493-2025-86528 | 2025-05-12 04:00 | Death | CAREFUSION SD | 2025-06-12 04:00 | FRN | ALARIS SYSTEM | Use of Device Problem; Inaccurate Delivery | Low Blood Pressure/ Hypotension | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER REQUESTED ALOG REVIEW TO KNOW WHAT WAS PROGRAMMED ON PATIENT ON (B)(6) 2025. IT WAS INITIALLY REPORTED THERE WAS PATIENT INVOLVEMENT BUT HARM WAS UNKNOWN. THROUGH DUE DILIGENCE ATTEMPTS, ADDITIONAL INFORMATION WAS RECEIVED. PATIENT WAS TAKEN FOR EMERGENCY SURGERY TO REPAIR A TRANSECTED POPLITEAL ARTERY AFTER SUFFERING A MECHANICAL FALL. POST-OPERATIVELY ON RETURN TO THE ICU, THE SURGEON ORDERED A FLUID BOLUS OF LACTATED RINGERS SOLUTION. THE CLINICIAN PROGRAMMED THE PUMP TO 999 ML/HOUR. UPON HEARING THE ALARM FOR AIR IN LINE, IT WAS DISCOVERED THAT THE CLINICIAN HAD PROGRAMMED THAT RATE OF 999ML/HR ON THE CHANNEL THAT HAD A FENTANYL DRIP RUNNING AT 1ML/HOUR. THE PATIENT RECEIVED THE ENTIRE BAG OF FENTANYL 2,500 MCG/100ML OVER APPROXIMATELY 20-30 MINUTES. NARCAN 2MG WAS ADMINISTERED IMMEDIATELY AFTER THE FENTANYL ADMINISTRATION. THE PATIENT WAS ALREADY ON A NEO SYNEPHRINE DRIP TO MAINTAIN BLOOD PRESSURE, THIS WAS TITRATED UP TO A MAX RATE OF 300 MCG/MINUTE. LEVOPHED DRIP WAS ADDED APPROXIMATELY 6 HOURS AFTER THE FENTANYL ADMINISTRATION. OVER THE NEXT 6 HOURS THE PATIENT¿S BLOOD PRESSURE RANGED FROM 83/46 MMHG TO 133/53MMHG. MEAN ARTERIAL PRESSURE RANGED FROM 34 MMHG TO 129 MMHG. CUSTOMER STATES THEY ARE UNABLE TO DETERMINE HOW MUCH OF THE BLOOD PRESSURE CHANGES WERE DUE TO THE FENTANYL VS POST-OPERATIVE HYPOTENSION AND KNOWN SIGNIFICANT INTRA-OP BLOOD LOSS DUE TO COAGULOPATHY, AND PRE-OP BLOOD LOSS FROM HEMATOMA. THE PATIENT CONTINUED TO EXPERIENCE POST-OP HYPOTENSION FROM LARGE VOLUME BLOOD LOSS, TREATED WITH BLOOD TRANSFUSIONS, FLUID BOLUSES AND IV PRESSORS. PATIENT ULTIMATELY SUFFERED MULTI-ORGAN FAILURE AND EXPIRED. ADDITIONAL MEDICATIONS INFUSING AT THE TIME INCLUDED: FENTANYL, 2500MCG IN 100ML NACL 0.9%, RUNNING AT 1MCG/HR (OR 1ML/HR) NEO-SYNEPHRINE, 50,000 MCG IN 250ML NACL 0.9%, RUNNING AT 60-300 MCG/MIN.Event Description: IT WAS REPORTED THAT THE CUSTOMER REQUESTED A LOG REVIEW TO KNOW WHAT WAS PROGRAMMED ON PATIENT ON (B)(6) 2025. IT WAS INITIALLY REPORTED THERE WAS PATIENT INVOLVEMENT, BUT OUTCOME WAS UNKNOWN. THROUGH DUE DILIGENCE ATTEMPTS, ADDITIONAL INFORMATION WAS RECEIVED. PATIENT WAS TAKEN FOR EMERGENCY SURGERY TO REPAIR A TRANSECTED POPLITEAL ARTERY AFTER SUFFERING A MECHANICAL FALL. POST-OPERATIVELY ON RETURN TO THE ICU, THE SURGEON ORDERED A FLUID BOLUS OF LACTATED RINGERS SOLUTION. THE CLINICIAN PROGRAMMED THE PUMP TO 999 ML/HOUR. UPON HEARING THE ALARM FOR AIR IN LINE, IT WAS DISCOVERED THAT THE CLINICIAN HAD PROGRAMMED THAT RATE OF 999ML/HR. ON THE CHANNEL THAT HAD A FENTANYL DRIP RUNNING AT 1ML/HOUR. THE PATIENT RECEIVED THE ENTIRE BAG OF FENTANYL 2,500 MCG/100ML OVER APPROXIMATELY 20-30 MINUTES. NARCAN 2MG WAS ADMINISTERED IMMEDIATELY AFTER THE FENTANYL ADMINISTRATION. THE PATIENT WAS ALREADY ON A NEO SYNEPHRINE DRIP TO MAINTAIN BLOOD PRESSURE; THIS WAS TITRATED UP TO A MAX RATE OF 300 MCG/MINUTE. LEVOPHED DRIP WAS ADDED APPROXIMATELY 6 HOURS AFTER THE FENTANYL ADMINISTRATION. OVER THE NEXT 6 HOURS THE PATIENT¿S BLOOD PRESSURE RANGED FROM 83/46 MMHG TO 133/53MMHG. MEAN ARTERIAL PRESSURE RANGED FROM 34 MMHG TO 129 MMHG. CUSTOMER STATES THEY ARE UNABLE TO DETERMINE HOW MUCH OF THE BLOOD PRESSURE CHANGES WERE DUE TO THE FENTANYL VS POST-OPERATIVE HYPOTENSION AND KNOWN SIGNIFICANT INTRA-OP BLOOD LOSS DUE TO COAGULOPATHY, AND PRE-OP BLOOD LOSS FROM HEMATOMA. THE PATIENT CONTINUED TO EXPERIENCE POST-OP HYPOTENSION FROM LARGE VOLUME BLOOD LOSS, TREATED WITH BLOOD TRANSFUSIONS, FLUID BOLUSES AND IV PRESSORS. PATIENT ULTIMATELY SUFFERED MULTI-ORGAN FAILURE AND EXPIRED. ADDITIONAL MEDICATIONS INFUSING AT THE TIME INCLUDED: FENTANYL, 2500MCG IN 100ML NACL 0.9%, RUNNING AT 1MCG/HR (OR 1ML/HR). NEO-SYNEPHRINE, 50,000 MCG IN 250ML NACL 0.9%, RUNNING AT 60-300 MCG/MIN. Manufacturer Narrative: OMIT: C20 - NO FINDINGS AVAILABLE, D15 - CAUSE NOT ESTABLISHED. CORRECTION: DEVICE MANUFACTURE DATE. ADDITIONAL INFORMATION: IMDRF ANNEX A, B, C, D, G CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. INVESTIGATION SUMMARY: THE REPORTED ISSUE OF A PROGRAMMING ISSUE INVOLVING FENTANYL COULD NOT BE DEFINITIVELY CONFIRMED OR REPLICATED. LOG ANALYSIS COULD NOT CONFIRM A USER PROGRAMMING ERROR RELATED TO THE FENTANYL INFUSION, AS THE INFUSION WAS PROGRAMMED UNDER THE GUARDRAILS IV FLUIDS PROFILE. THIS PROFILE DOES NOT ASSOCIATE A SPECIFIC DRUG ID, MAKING IT IMPOSSIBLE TO VERIFY WHETHER FENTANYL WAS THE ACTUAL DRUG ADMINISTERED AT THE 999 ML/H RATE. HOWEVER, THE ANALYSIS DID CONFIRM THAT THE SUSPECT PUMP MODULE WAS PROGRAMMED AT A RATE OF 999 ML/H, AND THAT AN AIR-IN-LINE ALARM OCCURRED, CONSISTENT WITH THE CLINICIAN¿S REPORT LABORATORY TESTING, AS WELL AS EXTERNAL AND INTERNAL INSPECTION OF THE SUSPECT DEVICES COULD NOT BE PERFORMED, AS THEY WERE NOT RETURNED FOR INVESTIGATION. THE PC UNIT (PCU) AND PUMP MODULE WERE NOT RETURNED, DESPITE MULTIPLE ATTEMPTS TO RETRIEVE THEM. THE FACILITY CONFIRMED THAT THE SUSPECT DEVICES AND ASSOCIATED COMPONENTS WOULD NOT BE RELEASED TO BD. HOWEVER, THE FACILITY PROVIDED THE LOGS FROM THE PUMP MODULES AND PCU, ALONG WITH A DATA SET, TO HELP DETERMINE THE DETAILS OF THE REPORTED COMPLAINT. A REVIEW OF THE PUMP MODULE AND PCU ERROR LOGS REVEALED THAT NO ERRORS WERE RECORDED DURING THE EVENT OR IN CONNECTION WITH THE REPORTED INCIDENT. ADDITIONALLY, LOG ANALYSIS CONFIRMED THAT PUMP MODULE SN (B)(6) WAS ADDED TO THE ALARIS CONFIGURATION BUT WAS NOT IN USE ON THE DATE OF THE REPORTED EVENT. THE EVENT DATE WAS REPORTED AS (B)(6) 2025, ALTHOUGH NO SPECIFIC TIME WAS PROVIDED. BASED ON THE ANALYSIS IT WAS CONCLUDED THAT THE TIME OF THE EVENT WAS 8:39 PM AS THE AIR-IN-LINE WAS RECORDED ON THE SUSPECT PUMP MODULE S/N: (B)(6). THIS ALARM FOLLOWED THE PROGRAMMING OF AN INFUSION AT 999 ML/H, WHICH ALIGNS WITH THE CLINICIAN¿S REPORT OF A FENTANYL INFUSION ERROR. SHORTLY AFTER, THE SUSPECT PUMP MODULE WAS REPROGRAMMED TO INFUSE MIDAZOLAM 100 MG/100 ML (DRUG ID (B)(4)) AT A DOSE OF 2 MCG/H, AS ALSO REPORTED BY THE CLINICIAN. HOWEVER, BECAUSE THE EARLIER INFUSION WAS PROGRAMMED UNDER THE GUARDRAILS IV FLUIDS PROFILE, IT IS NOT POSSIBLE TO CONFIRM THE DRUG ID ASSOCIATED WITH THE 999 ML/H RATE. LOG ANALYSIS CONFIRMED THE INFUSION OF PHENYLEPHRINE (DRUG ID (B)(4)) REPORTED BY THE FACILITY AS NEO-SYNEPHRINE AND MIDAZOLAM (DRUG ID (B)(4)). OF THESE, ONLY PHENYLEPHRINE (NEO-SYNEPHRINE) WAS REPORTED BY THE FACILITY. THE OTHER DRUGS IDENTIFIED IN THE LOGS DID NOT MATCH THOSE REPORTED BY THE FACILITY, WHICH INCLUDED LACTATED RINGERS, FENTANYL, AND NARCAN. THE SUSPECT PUMP MODULE (SN (B)(6)) WAS ADDED TO THE PCU AT 8:24 PM ON (B)(6) 2025. IT WAS PROGRAMMED NINE MINUTES LATER TO INFUSE MAINTENANCE IV FLUIDS USING THE GUARDRAILS IV FLUIDS PROFILE, AT A RATE OF 999 ML/H AND A VOLUME TO BE INFUSED (VTBI) OF 500 ML. THE LOG REVIEW CONFIRMED THE PROGRAMMED INFUSION AT 999 ML/H AND THE AIR-IN-LINE ALARM ON THE SUSPECT PUMP MODULE. HOWEVER, THE ALARM OCCURRED ONLY BETWEEN 8:34 PM AND 8:39 PM IT IS NOT POSSIBLE TO DETERMINE WHAT THE ACTUAL VOLUME IS WITHIN AN IV CONTAINER AT THE START AND ENDING OF AN INFUSION FROM A REVIEW OF A PCU EVENT LOG. THIS IS NOT A FUNCTION OF THE EVENT LOG; IT ONLY CAN REFLECT THE INPUTTED DRUG AND PROGRAMMING INFORMATION IT RECEIVES EITHER MANUALLY OR REMOTELY WITH RESPECT TO VOLUME TO BE DELIVERED. PROPER OPERATION OF THE BD ALARIS¿ SYSTEM REQUIRES FAMILIARITY WITH ITS FEATURES, SETUP, PROGRAMMING, IV SETS, AND ACCESSORIES. BEFORE USE, READ ALL INSTRUCTIONS, INCLUDING THOSE FOR ANY ATTACHED MODULES. THE BD ALARIS¿ PCU IS THE CORE OF THE BD ALARIS¿ SYSTEM, PROVIDING A COMMON USER INTERFACE FOR PROGRAMMING INFUSIONS. HOWEVER, THE BD ALARIS¿ SYSTEM USER MANUAL EMPHASIZES VERIFYING INFUSION PARAMETERS BEFORE SELECTING THE START KEY. THE ADMINISTRATION SET WAS REQUESTED BUT WAS NOT RETURNED; THEREFORE, IT COULD NOT BE INSPECTED OR TESTED. THE DEVICE WAS IN USE FOR TREATMENT PURPOSES AS INTENDED PER 21 CFR 820.198(D)(2). ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED PROGRAMMING ERROR DURING THE FENTANYL INFUSION COULD NOT BE DEFINITIVELY CONFIRMED. HOWEVER, A REVIEW OF THE LOGS DID CONFIRM AN INFUSION RATE OF 999 ML/HR, AS REPORTED BY THE FACILITY ON THE DAY OF THE EVENT, HOWEVER IT IS NOT KNOWN IF THE IV BAG CONTAINING FENTANYL WAS CONNECTED TO THIS DEVICE. THIS RATE WAS REPORTED TO HAVE BEEN PROGRAMMED FOR FENTANYL INFUSION AND COINCIDED WITH AN AIR-IN-LINE ALARM. INSPECTION AND LABORATORY TESTING OF THE DEVICES COULD NOT BE PERFORMED BECAUSE THE DEVICES WERE NOT RETURNED FOR INVESTIGATION. Manufacturer Narrative: THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22191999&pc=FRN | 2016493-2025-86527 | 2025-05-12 04:00 | Death | CAREFUSION SD | 2025-06-12 04:00 | FRN | ALARIS SYSTEM | Use of Device Problem; Inaccurate Delivery | Low Blood Pressure/ Hypotension | K133532 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER REQUESTED ALOG REVIEW TO KNOW WHAT WAS PROGRAMMED ON PATIENT ON (B)(6) 2025. IT WAS INITIALLY REPORTED THERE WAS PATIENT INVOLVEMENT BUT HARM WAS UNKNOWN. THROUGH DUE DILIGENCE ATTEMPTS, ADDITIONAL INFORMATION WAS RECEIVED. PATIENT WAS TAKEN FOR EMERGENCY SURGERY TO REPAIR A TRANSECTED POPLITEAL ARTERY AFTER SUFFERING A MECHANICAL FALL. POST-OPERATIVELY ON RETURN TO THE ICU, THE SURGEON ORDERED A FLUID BOLUS OF LACTATED RINGERS SOLUTION. THE CLINICIAN PROGRAMMED THE PUMP TO 999 ML/HOUR. UPON HEARING THE ALARM FOR AIR IN LINE, IT WAS DISCOVERED THAT THE CLINICIAN HAD PROGRAMMED THAT RATE OF 999ML/HR ON THE CHANNEL THAT HAD A FENTANYL DRIP RUNNING AT 1ML/HOUR. THE PATIENT RECEIVED THE ENTIRE BAG OF FENTANYL 2,500 MCG/100ML OVER APPROXIMATELY 20-30 MINUTES. NARCAN 2MG WAS ADMINISTERED IMMEDIATELY AFTER THE FENTANYL ADMINISTRATION. THE PATIENT WAS ALREADY ON A NEO SYNEPHRINE DRIP TO MAINTAIN BLOOD PRESSURE, THIS WAS TITRATED UP TO A MAX RATE OF 300 MCG/MINUTE. LEVOPHED DRIP WAS ADDED APPROXIMATELY 6 HOURS AFTER THE FENTANYL ADMINISTRATION. OVER THE NEXT 6 HOURS THE PATIENT¿S BLOOD PRESSURE RANGED FROM 83/46 MMHG TO 133/53MMHG. MEAN ARTERIAL PRESSURE RANGED FROM 34 MMHG TO 129 MMHG. CUSTOMER STATES THEY ARE UNABLE TO DETERMINE HOW MUCH OF THE BLOOD PRESSURE CHANGES WERE DUE TO THE FENTANYL VS POST-OPERATIVE HYPOTENSION AND KNOWN SIGNIFICANT INTRA-OP BLOOD LOSS DUE TO COAGULOPATHY, AND PRE-OP BLOOD LOSS FROM HEMATOMA. THE PATIENT CONTINUED TO EXPERIENCE POST-OP HYPOTENSION FROM LARGE VOLUME BLOOD LOSS, TREATED WITH BLOOD TRANSFUSIONS, FLUID BOLUSES AND IV PRESSORS. PATIENT ULTIMATELY SUFFERED MULTI-ORGAN FAILURE AND EXPIRED. ADDITIONAL MEDICATIONS INFUSING AT THE TIME INCLUDED: FENTANYL, 2500MCG IN 100ML NACL 0.9%, RUNNING AT 1MCG/HR (OR 1ML/HR) NEO-SYNEPHRINE, 50,000 MCG IN 250ML NACL 0.9%, RUNNING AT 60-300 MCG/MIN.Event Description: IT WAS REPORTED THAT THE CUSTOMER REQUESTED A LOG REVIEW TO KNOW WHAT WAS PROGRAMMED ON PATIENT ON (B)(6) 2025. IT WAS INITIALLY REPORTED THERE WAS PATIENT INVOLVEMENT, BUT OUTCOME WAS UNKNOWN. THROUGH DUE DILIGENCE ATTEMPTS, ADDITIONAL INFORMATION WAS RECEIVED. PATIENT WAS TAKEN FOR EMERGENCY SURGERY TO REPAIR A TRANSECTED POPLITEAL ARTERY AFTER SUFFERING A MECHANICAL FALL. POST-OPERATIVELY ON RETURN TO THE ICU, THE SURGEON ORDERED A FLUID BOLUS OF LACTATED RINGERS SOLUTION. THE CLINICIAN PROGRAMMED THE PUMP TO 999 ML/HOUR. UPON HEARING THE ALARM FOR AIR IN LINE, IT WAS DISCOVERED THAT THE CLINICIAN HAD PROGRAMMED THAT RATE OF 999ML/HR. ON THE CHANNEL THAT HAD A FENTANYL DRIP RUNNING AT 1ML/HOUR. THE PATIENT RECEIVED THE ENTIRE BAG OF FENTANYL 2,500 MCG/100ML OVER APPROXIMATELY 20-30 MINUTES. NARCAN 2MG WAS ADMINISTERED IMMEDIATELY AFTER THE FENTANYL ADMINISTRATION. THE PATIENT WAS ALREADY ON A NEO SYNEPHRINE DRIP TO MAINTAIN BLOOD PRESSURE, THIS WAS TITRATED UP TO A MAX RATE OF 300 MCG/MINUTE. LEVOPHED DRIP WAS ADDED APPROXIMATELY 6 HOURS AFTER THE FENTANYL ADMINISTRATION. OVER THE NEXT 6 HOURS THE PATIENT¿S BLOOD PRESSURE RANGED FROM 83/46 MMHG TO 133/53MMHG. MEAN ARTERIAL PRESSURE RANGED FROM 34 MMHG TO 129 MMHG. CUSTOMER STATES THEY ARE UNABLE TO DETERMINE HOW MUCH OF THE BLOOD PRESSURE CHANGES WERE DUE TO THE FENTANYL VS POST-OPERATIVE HYPOTENSION AND KNOWN SIGNIFICANT INTRA-OP BLOOD LOSS DUE TO COAGULOPATHY, AND PRE-OP BLOOD LOSS FROM HEMATOMA. THE PATIENT CONTINUED TO EXPERIENCE POST-OP HYPOTENSION FROM LARGE VOLUME BLOOD LOSS, TREATED WITH BLOOD TRANSFUSIONS, FLUID BOLUSES AND IV PRESSORS. PATIENT ULTIMATELY SUFFERED MULTI-ORGAN FAILURE AND EXPIRED. ADDITIONAL MEDICATIONS INFUSING AT THE TIME INCLUDED: FENTANYL, 2500MCG IN 100ML NACL 0.9%, RUNNING AT 1MCG/HR (OR 1ML/HR). NEO-SYNEPHRINE, 50,000 MCG IN 250ML NACL 0.9%, RUNNING AT 60-300 MCG/MIN. Manufacturer Narrative: THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. Manufacturer Narrative: OMIT : C20 - NO FINDINGS AVAILABLE, D15 - CAUSE NOT ESTABLISHED. ADDITIONAL INFORMATION: IMDRF ANNEX A, B, C, D, G CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. INVESTIGATION SUMMARY: THE REPORTED ISSUE OF A PROGRAMMING ISSUE INVOLVING FENTANYL COULD NOT BE DEFINITIVELY CONFIRMED OR REPLICATED. LOG ANALYSIS COULD NOT CONFIRM A USER PROGRAMMING ERROR RELATED TO THE FENTANYL INFUSION, AS THE INFUSION WAS PROGRAMMED UNDER THE GUARDRAILS IV FLUIDS PROFILE. THIS PROFILE DOES NOT ASSOCIATE A SPECIFIC DRUG ID, MAKING IT IMPOSSIBLE TO VERIFY WHETHER FENTANYL WAS THE ACTUAL DRUG ADMINISTERED AT THE 999 ML/H RATE. HOWEVER, THE ANALYSIS DID CONFIRM THAT THE SUSPECT PUMP MODULE WAS PROGRAMMED AT A RATE OF 999 ML/H, AND THAT AN AIR-IN-LINE ALARM OCCURRED, CONSISTENT WITH THE CLINICIAN¿S REPORT LABORATORY TESTING, AS WELL AS EXTERNAL AND INTERNAL INSPECTION OF THE SUSPECT DEVICES COULD NOT BE PERFORMED, AS THEY WERE NOT RETURNED FOR INVESTIGATION. THE PC UNIT (PCU) AND PUMP MODULE WERE NOT RETURNED, DESPITE MULTIPLE ATTEMPTS TO RETRIEVE THEM. THE FACILITY CONFIRMED THAT THE SUSPECT DEVICES AND ASSOCIATED COMPONENTS WOULD NOT BE RELEASED TO BD. HOWEVER, THE FACILITY PROVIDED THE LOGS FROM THE PUMP MODULES AND PCU, ALONG WITH A DATA SET, TO HELP DETERMINE THE DETAILS OF THE REPORTED COMPLAINT. A REVIEW OF THE PUMP MODULE AND PCU ERROR LOGS REVEALED THAT NO ERRORS WERE RECORDED DURING THE EVENT OR IN CONNECTION WITH THE REPORTED INCIDENT. ADDITIONALLY, LOG ANALYSIS CONFIRMED THAT PUMP MODULE SN (B)(6) WAS ADDED TO THE ALARIS CONFIGURATION BUT WAS NOT IN USE ON THE DATE OF THE REPORTED EVENT. THE EVENT DATE WAS REPORTED AS 12MAY2025, ALTHOUGH NO SPECIFIC TIME WAS PROVIDED. BASED ON THE ANALYSIS IT WAS CONCLUDED THAT THE TIME OF THE EVENT WAS 8:39 PM AS THE AIR-IN-LINE WAS RECORDED ON THE SUSPECT PUMP MODULE S/N:(B)(6). THIS ALARM FOLLOWED THE PROGRAMMING OF AN INFUSION AT 999 ML/H, WHICH ALIGNS WITH THE CLINICIAN¿S REPORT OF A FENTANYL INFUSION ERROR. SHORTLY AFTER, THE SUSPECT PUMP MODULE WAS REPROGRAMMED TO INFUSE MIDAZOLAM 100 MG/100 ML (DRUG ID 259) AT A DOSE OF 2 MCG/H, AS ALSO REPORTED BY THE CLINICIAN. HOWEVER, BECAUSE THE EARLIER INFUSION WAS PROGRAMMED UNDER THE GUARDRAILS IV FLUIDS PROFILE, IT IS NOT POSSIBLE TO CONFIRM THE DRUG ID ASSOCIATED WITH THE 999 ML/H RATE. LOG ANALYSIS CONFIRMED THE INFUSION OF PHENYLEPHRINE (DRUG ID 299) REPORTED BY THE FACILITY AS NEO-SYNEPHRINE AND MIDAZOLAM (DRUG ID 259). OF THESE, ONLY PHENYLEPHRINE (NEO-SYNEPHRINE) WAS REPORTED BY THE FACILITY. THE OTHER DRUGS IDENTIFIED IN THE LOGS DID NOT MATCH THOSE REPORTED BY THE FACILITY, WHICH INCLUDED LACTATED RINGERS, FENTANYL, AND NARCAN. THE SUSPECT PUMP MODULE (SN (B)(6)) WAS ADDED TO THE PCU AT 8:24 PM ON 12 MAY 2025. IT WAS PROGRAMMED NINE MINUTES LATER TO INFUSE MAINTENANCE IV FLUIDS USING THE GUARDRAILS IV FLUIDS PROFILE, AT A RATE OF 999 ML/H AND A VOLUME TO BE INFUSED (VTBI) OF 500 ML. THE LOG REVIEW CONFIRMED THE PROGRAMMED INFUSION AT 999 ML/H AND THE AIR-IN-LINE ALARM ON THE SUSPECT PUMP MODULE. HOWEVER, THE ALARM OCCURRED ONLY BETWEEN 8:34 PM AND 8:39 PM IT IS NOT POSSIBLE TO DETERMINE WHAT THE ACTUAL VOLUME IS WITHIN AN IV CONTAINER AT THE START AND ENDING OF AN INFUSION FROM A REVIEW OF A PCU EVENT LOG. THIS IS NOT A FUNCTION OF THE EVENT LOG; IT ONLY CAN REFLECT THE INPUTTED DRUG AND PROGRAMMING INFORMATION IT RECEIVES EITHER MANUALLY OR REMOTELY WITH RESPECT TO VOLUME TO BE DELIVERED. PROPER OPERATION OF THE BD ALARIS¿ SYSTEM REQUIRES FAMILIARITY WITH ITS FEATURES, SETUP, PROGRAMMING, IV SETS, AND ACCESSORIES. BEFORE USE, READ ALL INSTRUCTIONS, INCLUDING THOSE FOR ANY ATTACHED MODULES. THE BD ALARIS¿ PCU IS THE CORE OF THE BD ALARIS¿ SYSTEM, PROVIDING A COMMON USER INTERFACE FOR PROGRAMMING INFUSIONS. HOWEVER, THE BD ALARIS¿ SYSTEM USER MANUAL EMPHASIZES VERIFYING INFUSION PARAMETERS BEFORE SELECTING THE START KEY. THE ADMINISTRATION SET WAS REQUESTED BUT WAS NOT RETURNED; THEREFORE, IT COULD NOT BE INSPECTED OR TESTED. THE DEVICE WAS IN USE FOR TREATMENT PURPOSES AS INTENDED PER 21 CFR 820.198(D)(2). ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED PROGRAMMING ERROR DURING THE FENTANYL INFUSION COULD NOT BE DEFINITIVELY CONFIRMED. HOWEVER, A REVIEW OF THE LOGS DID CONFIRM AN INFUSION RATE OF 999 ML/HR, AS REPORTED BY THE FACILITY ON THE DAY OF THE EVENT, HOWEVER IT IS NOT KNOWN IF THE IV BAG CONTAINING FENTANYL WAS CONNECTED TO THIS DEVICE. THIS RATE WAS REPORTED TO HAVE BEEN PROGRAMMED FOR FENTANYL INFUSION AND COINCIDED WITH AN AIR-IN-LINE ALARM. INSPECTION AND LABORATORY TESTING OF THE DEVICES COULD NOT BE PERFORMED BECAUSE THE DEVICES WERE NOT RETURNED FOR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22192151&pc=OYC | 2032227-2025-190073 | 2024-04-25 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-12 04:00 | OYC | PARADIGM REAL-TIME REVEL INSULIN INFUSION PUMP MMT-723NAS | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P150019 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY AT THE HOSPITAL ON (B)(6) 2024. THE CUSTOMER WAS ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS DIABETES AND RESPIRATORY FAILURE. THE CUSTOMER¿S BLOOD GLUCOSE WAS UNKNOWN. THE EVENT INVOLVED PRODUCT(S) UNK_SENSOR, UNOMEDICAL, MMT-332A, MMT-723NAS. THE INSULIN PUMP WAS NOT WORN AT THE TIME OF PASSING. SENSOR WAS NOT WORN AT TIME OF PASSING. THE INSULIN PUMP WAS LAST WORN ON (B)(6) 2024, PRIOR TO PASSING. NO PRODUCT RETURN IS REQUIRED FOR UNK_SENSOR, UNOMEDICAL, MMT-332A, MMT-723NAS. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22192305&pc=BZD | 2518422-2025-046818 | 2023-07-14 04:00 | Death | RESPIRONICS, INC. | 2025-06-12 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH DUE TO ALZHEIMER'S. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, NO SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULD CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22192365&pc=BZD | 2518422-2025-046835 | 2024-07-17 04:00 | Death | RESPIRONICS, INC. | 2025-06-12 04:00 | BZD | DREAMSTATION AUTO CPAP | Contamination; Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE REPORTER HAS ALLEGED HER HUSBAND PASSED AWAY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING INVESTIGATION, THE MANUFACTURER OBSERVED AN UNKNOWN LIGHT CONTAMINATION AT THE AIR INLET AND THE BOTTOM OF THE BLOWER BOX. LIGHT BLACK SPECKS IN THE BOTTOM ENCLOSURE. UNKNOWN LIGHT WHITE DUST-LIKE CONTAMINATION WAS ON TOP AND BOTTOM OF THE BLOWER MOTOR AND THE BLOWER MOTOR SEAL. MISSING AIR INLET FILTER. DURING INVESTIGATION THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER SUPPLY AND POWER CORD TO THE DEVICE AND VERIFIED AIRFLOW AND CONFIRMED THAT THE DEVICE POWERS ON PROVIDES THERAPY. THE MANUFACTURER WAS NOT ABLE TO CONFIRM THE COMPLAINT. NO SPECIFIC COMPLAINT GIVEN. THE MANUFACTURER OBSERVED DUST CONTAMINATION IN THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22192367&pc=NIK | 9614453-2025-02264 | 2025-04-30 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-06-12 04:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Failure to Deliver Shock/Stimulation | Myocardial Infarction | P010031 |  | 1 | Event Description: IT WAS REPORTED BY THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) PATIENT'S SPOUSE THAT THE PATIENT EXPIRED TWO WEEKS POST-IMPLANT AND THE DAY AFTER LEAVING THE HOSPITAL. IT WAS REPORTED THAT THE DEVICE DID NOT WORK, DID NOT DO ANYTHING, AND DID NOT SHOCK THE PATIENT. THOUGH THE SPOUSE WAS REPORTEDLY INFORMED THAT THE PATIENT HAD A HEART ATTACK AND DIED OF NATURAL CAUSES, THEY BELIEVED THE DEATH WAS DEVICE-RELATED. ADDITIONAL INFORMATION WAS REQUESTED BUT NOT OBTAINED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22192666&pc=MIH | 3002808486-2025-00131 | 2020-11-20 05:00 | Death | WILLIAM COOK EUROPE | 2025-06-12 04:00 | MIH | ZENITH ALPHA¿ THORACIC ENDOVASCULAR GRAFT | Patient Device Interaction Problem | Hemorrhage/Bleeding; Shock |  |  | 1 | Event Description: DESCRIPTION OF EVENT ACCORDING TO STUDY (B)(6): ZENITH ALPHA THORACIC POST MARKET RETROSPECTIVE STUDY): SYNOPSIS: THE PATIENT EXPERIENCED HEMORRHAGE/HEMORRHAGIC SHOCK 18 DAYS POST-PROCEDURE AND DIED 27 DAYS POST-PROCEDURE. ON (B)(6) 2020, THIS 79-YEAR-OLD MALE PATIENT WAS ROUTINELY TREATED FOR CHRONIC THORACIC AORTIC DISSECTION TYPE B WITH PLACEMENT OF A PROXIMAL ZTA-PT-42-38-225 AND A DISTAL ZTA-PT-38-34-217, STARTING AT ZONE 2. ON (B)(6) 2020, THE PATIENT EXPERIENCED HEMORRHAGE/HEMORRHAGIC SHOCK. THE EVENT WAS TREATED WITH ANTIBIOTICS, INOTROPIC CARDIAC SUPPORT, AND OXYGEN. THE EVENT LED TO DEATH. THE SITE NOTED THE RELATIONSHIP OF THE EVENT TO THE STUDY DEVICE, PROCEDURE, AND TREATED AORTIC DISEASE AS ¿RETROSPECTIVE EVENT, UNABLE TO DETERMINE. PATIENT OUTCOME: ON (B)(6) 2020, THE PATIENT DIED. THE CAUSE OF DEATH WAS HEMORRHAGIC SHOCK. DESCRIPTION OF CIRCUMSTANCES INCLUDES ¿PCR COVID 19 POSITIVE.¿ RELATIONSHIP ASSESSMENT OF DEATH TO STUDY DEVICE AND PROCEDURE WAS NOT PROVIDED.Event Description: NO ADDITIONAL INFORMATION REGARDING THE PATIENT AND/OR EVENT HAS BEEN RECEIVED SINCE THE PREVIOUS MEDWATCH REPORT WAS SENT. Manufacturer Narrative: MANUFACTURERS REF# (B)(4). BLANK FIELDS ON THIS FORM INDICATE THE INFORMATION IS UNKNOWN OR UNAVAILABLE. D4) CATALOG #: ZTA-PT-42-38-225 AND ZTA-PT-38-34-217. D4) LOT #: E3943633 AND E3953730. G4) SIMILAR TO DEVICE MARKETED UNDER 510(K)/PMA: P140016. INVESTIGATION IS STILL IN PROGRESS. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED. Manufacturer Narrative: MANUFACTURERS REF#: (B)(4). BASED ON RESPONSE RECEIVED TO REQUESTED INFORMATION FROM THE SITE IT IS CONCLUDED THAT THERE ARE NO INDICATIONS THAT THE ZTA DEVICES ARE IMPLICATED IN THE EVENT OR THAT THE PROCEDURE IS RELATED TO THE DEATH AND THEREBY THE EVENT IS NO LONGER CONSIDERED REPORTABLE. COOK MEDICAL WILL CONTINUE TO MONITOR FOR SIMILAR EVENTS. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22192704&pc=DSQ | 2916596-2025-03680 | 2025-03-12 04:00 | Death | THORATEC CORPORATION | 2025-06-12 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED DUE TO PATIENT PRIVACY LAWS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22192710&pc=DSQ | 2916596-2025-03570 | 2025-05-09 04:00 | Death | THORATEC CORPORATION | 2025-06-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Mechanical Problem; Pumping Stopped; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Head Injury; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS INVOLVED IN A CAR CRASH ON (B)(6) 2025. THE DRIVELINE AND ONE OF THE BATTERY LEADS WERE CUT BY THE MEDICS AT THE LOCATION OF THE CRASH. IT WAS UNCLEAR WHETHER THE PATIENT'S DRIVELINE AND WHITE SYSTEM CONTROLLER LEAD WERE CUT BY THE PATIENT OR IF EMERGENCY MEDICAL SERVICES (EMS) CUT THE EQUIPMENT WHILE CUTTING OFF THEIR CLOTHES. AN AUTOPSY WAS REQUESTED BY THE FAMILY AND THE CORONER REQUESTED LOG FILE DATA FROM THE PUMP OR THE CONTROLLER TO DETERMINE WHEN THE DRIVELINE WAS CUT AND WHEN THE PUMP TURNED OFF.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT HAD A HEAD INJURY. THE PATIENT ARRIVED AT THE EMERGENCY ROOM IN PULSELESS ELECTRICAL ACTIVITY (PEA) AFTER RECEIVING CARDIOPULMONARY RESUSCITATION FOR 45 MINUTES. AFTER ONGOING RESUSCITATION ATTEMPTS, TIME OF DEATH WAS CALLED. IT WAS UNKNOWN AS AN AUTOPSY REPORT WAS NOT PROVIDED FROM THE CORONER. THE DEVICE OPERATED AS EXPECTED BEFORE THE DRIVELINE AND POWER CABLE WERE CUT PER THE DIAGNOSTICS REPORT. Manufacturer Narrative: D6B - DATE OF EXPLANT - CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: EVALUATION OF HEARTMATE 3, LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), DID NOT REVEAL ANY FUNCTIONAL ISSUES. A SPECIFIC CAUSE FOR THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. (B)(6) WAS RETURNED ASSEMBLED WITH THE PUMP CABLE CONNECTED TO THE RETURNED SECTION OF THE MODULAR CABLE. THE SEALED OUTFLOW GRAFT WAS RETURNED ATTACHED TO THE PUMP COVER OUTLET PORT WITH THE OUTFLOW GRAFT BEND RELIEF PROPERLY ENGAGED TO THE GRAFT HARDWARE. THE APICAL CUFF WAS RETURNED WITH THE CUFF LOCK FULLY ENGAGED. UPON DISASSEMBLY OF THE RETURNED PUMP, VISUAL INSPECTION OF THE BLOOD CONTACTING SURFACES DID NOT REVEAL ANY ADHERED DEPOSITIONS OR THROMBUS FORMATIONS. THE RETRIEVED LVAD (LEFT VENTRICULAR ASSIST DEVICE) LOG FILES FROM THE RETURNED DEVICE SHOWED THE PUMP OPERATING AS INTENDED UNTIL THE DRIVELINE WAS CUT. (B)(6) WAS CLEANED, REBUILT, AND FUNCTIONALLY TESTED ON A MOCK CIRCULATORY LOOP. THE DEVICE OPERATED AS INTENDED AND IN ACCORDANCE WITH MANUFACTURING SPECIFICATIONS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE IFU AND PATIENT HANDBOOK CONTAIN INFORMATION ON HOW TO CLEAN AND CARE FOR THE DRIVELINE. SECTION 6 OF THE IFU AND SECTION 4 OF THE PATIENT HANDBOOK INSTRUCT THE USER TO CHECK THE DRIVELINE DAILY FOR SIGNS OF DAMAGE, SUCH AS CUTS, HOLES, OR TEARS. THE PATIENT HANDBOOK ALSO STATES, ¿CALL YOUR HOSPITAL CONTACT RIGHT AWAY IF THE DRIVELINE IS DAMAGED (OR MIGHT BE DAMAGED)." SECTION 7 OF THE IFU AND SECTION 5 OF THE PATIENT HANDBOOK PROVIDE ADDITIONAL INSTRUCTIONS REGARDING DRIVELINE CARE IN SUB-SECTIONS ENTITLED, "WHAT NOT TO DO: DRIVELINE AND CABLES". NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22192753&pc=OZD | 1220648-2025-29238 | 2025-05-18 04:00 | Death | ABIOMED, INC. | 2025-06-12 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 FOR MECHANICAL CIRCULATORY SUPPORT. DURING SUPPORT, AN ISCHEMIC STROKE WAS NOTED. THE PATIENT WAS NOT WAKING UP FROM SEDATION AFTER SURGERY. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S DEMISE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22193063&pc=BZD | 2518422-2025-046853 | 2025-04-24 04:00 | Death | RESPIRONICS, INC. | 2025-06-12 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND IS ALSO REQUESTING FOR A RETURN LABEL SENT TO THE MAILING ADDRESS DOCUMENTED ABOVE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION. THERE WAS 0 ERROR CODE FOUND. THE MANUFACTURER SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22193382&pc=CAH | MW5171363 | 2024-11-16 05:00 | Death | CYTIVA PALL FILETER / PALL NEWQUAY | 2025-06-12 04:00 | CAH | PALL ULTIPOR 55 BREATHING CIRCUIT FILTER | Device Difficult to Setup or Prepare; Tidal Volume Fluctuations; Improper or Incorrect Procedure or Method; Obstruction of Flow | Brain Injury; Unspecified Respiratory Problem |  |  | 1 | Event Description: ON (B)(6) 2024, AN ADVERSE EVENT OCCURRED DURING A TRANSPLANT SURGERY INVOLVING THE USE OF PALL ULTIPOR 55 BREATHING CIRCUIT FILTER WITH AN AEROGEN SOLO NEBULIZER. THE PATIENT, A 57-YEAR-OLD MALE, EXPERIENCED PROLONGED RESPIRATORY CHALLENGES DUE TO AN INCORRECT SETUP OF THE CIRCUIT. WHICH LIKELY CAUSED THE FILTER TO CLOG. THE FILTER AEROGEN SOLO NEBULIZER WAS INTENDED TO DELIVER INHALED EPOPROSTENOL AFTER THE PATIENT WAS LIBERATED FROM CARDIOPULMONARY BYPASS. THE PATIENT WAS TAKEN OFF BYPASS AND CONNECTED TO A VENTILATOR FOR THE DELIVERY OF INHALED EPOPROSTENOL. HOWEVER, WITHIN TWO HOURS, THE PATIENT'S TIDAL VOLUMES DECREASED, INDICATING DIFFICULTY IN VENTILATION AND OXYGENATION, LIKELY DUE TO THE CLOGGED FILTER OBSTRUCTING AIRFLOW AND EPOPROSTENOL DELIVERY FROM THE AEROGEN SOLO NEBULIZER. THIS LED TO AN ANOXIC BRAIN INJURY AND ULTIMATELY THE PATIENT'S DEATH, WHICH OCCURRED ON (B)(6) 2024. THE ADVERSE EVENT WAS IDENTIFIED AS POTENTIALLY LINKED TO THE SETUP OF THE FILTER WITH THE AEROGEN SOLO NEBULIZER AFTER DISCOVERING THAT THE FILTER ON THE VENTILATOR HAD CLOGGED, WHICH OBSTRUCTED THE DELIVERY OF OXYGEN AND EPOPROSTENOL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22193786&pc=MKJ | 0003015876-2025-01203 | 2021-01-06 05:00 | Death | PHYSIO-CONTROL, INC. - 3015876 | 2025-06-12 04:00 | MKJ | LIFEPAK(R) CR2 DEFIBRILLATOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P170018 |  | 1 | Event Description: A SOCIAL MEDIA POST INVOLVING LPCR2 WAS DISCOVERED BY A STRYKER EMPLOYEE. THE VIDEO ILLUSTRATED THE USER'S INITIAL INABILITY TO EFFECTIVELY DEPLOY THE DEVICE DURING A PATIENT EVENT. THE PATIENT ASSOCIATED WITH THE REPORTED EVENT DID NOT SURVIVE. Manufacturer Narrative: STRYKER REQUESTED ADDITIONAL INFORMATION ON THE PATIENT AND EVENT. STRYKER REQUESTS PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK. STRYKER CONTINUES TO INVESTIGATE THE REPORTED ISSUE AND WILL SUBMIT A SUPPLEMENTAL REPORT ON THIS EVENT TO THE FDA AS PROVIDED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22193985&pc=DSQ | 2916596-2025-03880 | 2025-06-04 04:00 | Death | THORATEC CORPORATION | 2025-06-12 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WILL NOT BE EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. NO ADDITIONAL INFORMATION WOULD BE PROVIDED DUE TO PATIENT PRIVACY LAWS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22194131&pc=OZO | 2032227-2025-190343 | 2025-04-14 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-12 04:00 | OZO | 630G INSULIN PUMP MMT-1715KL 630G | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Lymphoma; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER WAS PASSED AWAY DUE TO HEART ATTACK, LYMPHOMA AND DIABETES. THE TREATMENT AND THE BLOOD VALUE INVOLVED IN THE EVENT IS UNKNOWN AND THE CUSTOMER DATE OF PASSING WAS (B)(6) 2025. THE EVENT INVOLVED PRODUCT(S) UNOMEDICAL, UNK_SENSOR, MMT-1715KL, UNK_RESERVOIR. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING AND OBSERVED THAT THE PUMP WAS WORN AT THE TIME OF PASSING. IT WAS UNKNOWN WHETHER THE CUSTOMER HAD BEEN USING THE INSULIN PUMP WITHIN 48 HOURS OF THE REPORTED EVENT AND UNKNOWN WHETHER THE AUTO MODE/SMARTGUARD FEATURE WAS ACTIVE AT THE TIME OF THE EVENT. NO FURTHER PATIENT COMPLICATIONS WERE REPORTED. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR UNK_SENSOR. NO PRODUCT RETURN IS REQUIRED FOR MMT-1715KL. NO PRODUCT RETURN IS REQUIRED FOR UNK_RESERVOIR.Event Description: UPDATED SUMMARY: THE EVENT INVOLVED PRODUCT(S) UNOMEDICAL, MMT-1715KL, MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR MMT-1715KL. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: ADDITIONAL INFORMATION HAS BEEN RECEIVED WHICH WAS NOT INCLUDED IN THE INITIAL REPORT. THE UPDATED INFORMATION HAS BEEN PROVIDED IN BELOW SECTIONS WITH THIS FOLLOW-UP REPORT. 1. SECTION B5 - UPDATED THE SUMMARY. 2. SECTION D10 - UPDATED CONCOMITANT PRODUCTS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22195014&pc=DSQ | 3007042319-2025-01266 | 2025-06-08 04:00 | Death | HEARTWARE, INC. | 2025-06-12 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE (VAD) IMPLANT EXPERIENCED WORSENING HEART FAILURE AND WAS FAILING TO THRIVE. THE PATIENT WAS PLACED IN HOSPICE CARE AS AN END-OF-LIFE DECISION AND SUBSEQUENTLY DIED IN HOSPICE. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINS IMPLANTED. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT WAS PLACED ON HOSPICE CARE AND SUBSEQUENTLY EXPIRED. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSE OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, WORSENING HEART FAILURE AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22195037&pc=LPB | 2029046-2025-01921 |  | Death | BIOSENSE WEBSTER INC | 2025-06-12 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Stenosis; Thrombosis/Thrombus | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PVC ABLATION PROCEDURE WITH A THERMOCOOL SMARTTOUCH SF CATHETER AND EXPERIENCED CARDIAC ARREST DUE TO THROMBUS IN THE COMMON TRUNK, STENTING OF THE LEFT CORONARY ARTERY TREATED WITH STENTING AND TRANSFERRED TO ANOTHER HOSPITAL. THE PATIENT DIED OF HEMORRHAGIC SHOCK. MINUTES AFTER THE PATIENT WAS SENT BACK TO HIS ROOM, CARDIAC ARREST OCCURRED BECAUSE OF A THROMBUS IN THE COMMON TRUNK AND NEEDED STENTING IN THE LEFT CORONARY. HOWEVER, RESTENOSIS WAS NOTED OVER THE STENT AND HAD ANOTHER ARREST AND NEEDED REANIMATION FOR ABOUT 2 HOURS BEFORE BEING TRANSFERRED TO ANOTHER HOSPITAL FOR SURGERY. THE PVC PROCEDURE WAS ON THE (LVOT) LEFT VENTRICLE OUTFLOW TRACT. THE PROCEDURE LASTED FOR 1 HOUR AND 30 MINUTES WITH 6 ABLATIONS. THE PVCS DISAPPEARED AFTER THE SECOND ABLATION. WAITED 10 MINUTES AFTER THE LAST ABLATION. THE RF ABLATION TIME WAS LESS THAN 2 MINUTES. ADDITIONAL INFORMATION WAS RECEIVED ON 20-MAY-2025. AFTER TRANSFERRED TO THE OTHER HOSPITAL, THE PATIENT DIED OF HEMORRHAGIC SHOCK. MULTIPLE ATTEMPTS HAVE BEEN MADE TO OBTAIN CLARIFICATION OF THIS COMPLAINT. HOWEVER, NO FURTHER INFORMATION HAS BEEN MADE AVAILABLE. Manufacturer Narrative: D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THE PICTURE INVESTIGATION DETAILS. A PICTURE WAS RECEIVED FOR EVALUATION FOLLOWING JOHNSON & JOHNSON MEDTECH'S PROCEDURES. ACCORDING TO PICTURES PROVIDED BY THE CUSTOMER, PROCEDURE SETTINGS WERE OBSERVED ON THE CARTO 3 SCREEN. HOWEVER, NONE OF THE REPORTED ADVERSE EVENTS CAN BE CONFIRMED BASED ONLY ON THE IMAGES PROVIDED. THE LOT NUMBER IS NOT AVAILABLE; THEREFORE, THE MANUFACTURING RECORD EVALUATION COULD NOT BE PERFORMED. THE CUSTOMER COMPLAINT WAS NOT CONFIRMED BASED ON THE PICTURE RECEIVED. THE DEVICE HAS NOT BEEN RETURNED FOR ANALYSIS AND A ROOT CAUSE IS UNABLE TO BE ASSIGNED BASED ON THE CURRENT EVIDENCE. IF THE DEVICE IS RECEIVED IN THE FUTURE, THE PRODUCT INVESTIGATION WILL BE PERFORMED AND UPDATED ACCORDINGLY, AND ANY ACTION WILL BE TAKEN IF NECESSARY. H6. INVESTIGATION FINDINGS CODE OF ¿APPROPRIATE TERM/CODE NOT AVAILABLE¿ REPRESENTS PHOTO/VIDEO ANALYSIS. MANUFACTURER¿S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED ON 08-JUL-2025. THE PHYSICIAN¿S OPINION ON THE CAUSE OF THIS ADVERSE EVENT WAS THAT BY GOING RETRO AORTIC, THEY COULD HAVE DESTABILIZED AN EXISTING ATHEROSCLEROTIC PLAQUE CREATING A CHAIN REACTION LEADING TO THROMBUS. THE INTERVENTION PROVIDED WAS STENT PROCEDURE IN THE LEFT CORONARY. HOWEVER, THE VESSEL RE-STENOSED OVER THE STENT AND NEEDED REANIMATION FOR ABOUT 2 HOURS. THEY PROVIDED CARDIAC MASSAGE AND PUT HIM ON ECMO BEFORE TRANSFERRING HIM TO (B)(6) HOSPITAL. DETAILS ON EVENT: THEY HAD A PVC PROCEDURE ON THE LEFT VENTRICLE LVOT. THE PROCEDURE LASTED FOR 1:30H. THEY DID 6 ABLATIONS. THE PVCS DISAPPEARED AFTER THE SECOND ABLATION, THEN WAITED 10 MINUTES AFTER THE LAST ABLATION. THE RADIO FREQUENCY (RF) ABLATION TIME WAS LESS THAN 2 MINUTES. ABOUT 10 MINUTES AFTER THE PATIENT WAS SENT BACK TO HIS ROOM, HE HAD CARDIAC ARREST BECAUSE OF A THROMBUS IN THE COMMON TRUNK. THE PATIENT HAD A SUCCESSFUL STENTING PROCEDURE IN THE LEFT CORONARY. BUT ABOUT 30 MINUTES AFTER HE WAS SENT BACK TO HIS ROOM AFTER THE STENTING PROCEDURE, HE RE-STENOSED OVER THE STENT. HE HAD ANOTHER CARDIAC ARREST AND NEEDED REANIMATION FOR ABOUT SEVERAL HOURS. THEY PROVIDED CARDIAC MASSAGE AND PUT HIM ON ECMO BEFORE TRANSFERRING HIM TO (B)(6). IN ADDITION, PROVIDED PATIENT INFORMATION AND THEREFORE, PROCESSED THE FOLLOWING FIELDS: A2. PATIENT AGE AT THE TIME OF EVENT, A2. AGE UNIT, A3A. SEX, A5. ETHNICITY. ALSO PROVIDED THE PRODUCT INFORMATION FOR A CONCOMITANT GENERATOR. THEREFORE, PROCESSED THE D10. CONCOMITANT MEDICAL PRODUCTS AND THERAPY DATES SECTION. IN THE 3500A INITIAL IT WAS REPORTED UNDER ADDITIONAL INFORMATION RECEIVED ON 20-MAY-2025 THAT AFTER THE PATIENT WAS TRANSFERRED TO THE OTHER HOSPITAL, THE PATIENT DIED OF HEMORRHAGIC SHOCK. DURING AN INTERNAL REVIEW ON 02-AUG-2025, NOTED A CORRECTION TO THE 3500A INITIAL AS THE B2. DATE OF DEATH FIELD WAS NOT PROCESSED. THEREFORE, PROCESSED AS (B)(6) 2025. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22195215&pc=MHX | 8030229-2025-04980 | 2025-05-16 04:00 | Death | NIHON KOHDEN CORPORATION | 2025-06-12 04:00 | MHX | CNS-2101 CENTRAL MONITOR | No Audible Alarm; Protective Measures Problem | No Clinical Signs, Symptoms or Conditions | K223567 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT THE HEART RATE (HR) FOR A PATIENT DROPPED TO 30BPM; HOWEVER, THERE WERE NO ALARMS AT ALL AT THE CENTRAL NURSE'S STATION (CNS).Event Description: THE CUSTOMER REPORTED THAT THE HEART RATE (HR) FOR A PATIENT DROPPED TO 30BPM; HOWEVER, THERE WERE NO ALARMS AT ALL AT THE CENTRAL NURSE'S STATION (CNS). Manufacturer Narrative: DETAILS OF COMPLAINT: THE CUSTOMER REPORTED THAT THE HEART RATE (HR) FOR A PATIENT DROPPED TO 30BPM BETWEEN 1:00PM AND 1:45PM; HOWEVER, THERE WERE NO ALARMS AT ALL AT THE CENTRAL NURSE'S STATION (CNS). WHEN THEY CHECKED THE CNS, IT SHOWED THAT THE ALARMS WERE SUSPENDED. THE PATIENT EXPERIENCED BRADYCARDIA FOR APPROXIMATELY 1-2 MINUTES, WITH THE HEART RATE DROPPING INTO THE 30S. IT WAS DURING THIS TIME THEY DISCOVERED THE ALARMS WERE TURNED OFF AT THE BEDSIDE. THE PATIENT UNFORTUNATELY EXPIRED. THE CUSTOMER WANTS THE LOGS REVIEWED TO SEE IF ONE OF THE NURSES SUSPENDED OR CHANGED THE ALARMS DURING THIS TIME FRAME OR IF SOMETHING ELSE CAUSED THE ALARMS TO BE SUSPENDED. INVESTIGATION SUMMARY: THE CUSTOMER PROVIDED THE EVENT LOGS. THE LOGS WERE REVIEWED FROM BOTH THE CNS AND THE G5 BEDSIDE MONITOR (CU-152RA SN: (B)(6)). ANALYSIS REVEALED THAT ALARMS WERE SUSPENDED DUE TO THE COMFORT CARE MODE BEING ENABLED ON THE G5 AT 11:54:52 ON (B)(6) 2025. THE ROOT CAUSE IS THAT COMFORT CARE WAS SELECTED TO BE ON WHICH SUSPENDS ALARMS AUTOMATICALLY. THE FOLLOWING FIELD(S) CONTAIN NO INFORMATION (NI), AS ATTEMPTS TO OBTAIN THE INFORMATION WERE MADE, BUT NOT PROVIDED: D10 ATTEMPT # 1: 06/02/2025 EMAILED THE CUSTOMER VIA MICROSOFT OUTLOOK FOR ALL INFORMATION IN THE NI LIST ABOVE: NO REPLY WAS RECEIVED. ATTEMPT # 2: 06/04/2025 EMAILED THE CUSTOMER VIA MICROSOFT OUTLOOK FOR ALL INFORMATION IN THE NI LIST ABOVE: NO REPLY WAS RECEIVED. ATTEMPT # 3: 06/11/2025 EMAILED THE CUSTOMER VIA MICROSOFT OUTLOOK FOR ALL INFORMATION IN THE NI LIST ABOVE: NO REPLY WAS RECEIVED. ADDITIONAL INFORMATION: B4 DATE OF THIS REPORT. G3 DATE RECEIVED BY MANUFACTURER. G6 TYPE OF REPORT. H2 IF FOLLOW UP, WHAT TYPE? H11 ADDITIONAL MANUFACTURER NARRATIVE. Manufacturer Narrative: THE CUSTOMER REPORTED THAT THE HEART RATE (HR) FOR A PATIENT DROPPED TO 30BPM BETWEEN 1:00PM AND 1:45PM; HOWEVER, THERE WERE NO ALARMS AT ALL AT THE CENTRAL NURSE'S STATION (CNS). WHEN THEY CHECKED THE CNS, IT SHOWED THAT THE ALARMS WERE SUSPENDED. THE PATIENT EXPERIENCED BRADYCARDIA FOR APPROXIMATELY 1-2 MINUTES, WITH THE HEART RATE DROPPING INTO THE 30S. IT WAS DURING THIS TIME THEY DISCOVERED THE ALARMS WERE TURNED OFF AT THE BEDSIDE. THE PATIENT UNFORTUNATELY EXPIRED. THE CUSTOMER WANTS THE LOGS REVIEWED TO SEE IF ONE OF THE NURSES SUSPENDED OR CHANGED THE ALARMS DURING THIS TIME FRAME OR IF SOMETHING ELSE CAUSED THE ALARMS TO BE SUSPENDED. NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. THE FOLLOWING FIELD(S) CONTAIN NO INFORMATION (NI), AS ATTEMPTS TO OBTAIN THE INFORMATION WERE MADE, BUT NOT PROVIDED:
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22195236&pc=DSQ | 2916596-2025-03750 | 2025-05-21 04:00 | Death | THORATEC CORPORATION | 2025-06-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT WAS ON COMFORT CARE THEN PASSED AWAY AFTER LEFT VENTRICULAR ASSIST DEVICE WAS TURNED OFF. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER: (B)(6) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22195371&pc=OZD | 1220648-2025-29256 | 2025-05-20 04:00 | Death | ABIOMED, INC. | 2025-06-12 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT PRESENTED IMPLANTED WITH AN IMPELLA CP AND WITH A MOTTLED RIGHT LEG. THE PATIENT WAS TAKEN TO THE CATH LAB TO BE ESCALATED TO AN IMPELLA 5.5. AFTER AN EXTERNAL FEM-FEM BYPASS DISTAL PULSES WERE PRESENT. THE PATIENT EXPRESSED THEY DID NOT WANT EXTREME MEASURES TAKEN AND CARE WAS WITHDRAWN. THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22196130&pc=NPT | 2015691-2025-04795 | 2025-05-23 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-12 04:00 | NPT | COMMANDER DELIVERY SYSTEM | Fluid/Blood Leak; Difficult to Remove; Adverse Event Without Identified Device or Use Problem | Thromboembolism; No Clinical Signs, Symptoms or Conditions | P140031 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN FRANCE. AS REPORTED, THIS WAS A CASE OF AN IMPLANT OF A 26MM SAPIEN 3 ULTRA IN THE AORTIC POSITION BY TRANSAXILLARY APPROACH. DURING INFLATION IT WAS NOTED THAT THE BALLOON COULD NOT BE INFLATED. THEREFORE, THE VALVE DEPLOYMENT COULD NOT BE PERFORMED. IT WAS ALSO OBSERVED THAT THERE WAS SOME BLOOD IN THE SYRINGE. IT WAS DECIDED TO REMOVE THE VALVE THROUGH THE TRANSAXILLARY ACCESS BUT IT WAS NOT POSSIBLE. IT WAS DECIDED TO MANAGE THE REMOVAL THROUGH THE TRANSFEMORAL APPROACH BUT, THE PATIENT HAD THROMBOSIS IN THE FEMORAL. DURING THE WITHDRAWAL A THROMBUS EMBOLIZED. FINALLY, THE VALVE WAS REMOVED AND A SECOND ONE WAS IMPLANTED. THE PROCEDURE WAS SUCCESSFUL AND THE PATIENT WAS NOTED TO BE STABLE. FEW DAYS AFTER THE PROCEDURE THE PATIENT, UNFORTUNATELY PASSED AWAY DUE TO THE THROMBOEMBOLISM. Manufacturer Narrative: THE INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE COMPLETED UPON COMPLETION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR EVALUATION AS IT WAS DISCARDED. THEREFORE, A NO PRODUCT RETURN ENGINEERING EVALUATION WAS PERFORMED. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE COMPLAINTS OF DELIVERY SYSTEM LEAKAGE AND WITHDRAW DIFFICULTY WERE UNABLE TO BE CONFIRMED AS NEITHER THE COMPLAINT DEVICE NOR APPLICABLE PROCEDURAL IMAGERY WERE PROVIDED FOR INVESTIGATION. DUE TO THE UNAVAILABILITY OF THE DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL, FUNCTIONAL, OR DIMENSIONAL ANALYSIS; THEREFORE, A MANUFACTURING NON-CONFORMANCE WAS UNABLE TO BE DETERMINED. A REVIEW OF DHR AND LOT HISTORY DID NOT PROVIDE ANY INDICATION THAT A MANUFACTURING NON-CONFORMANCE WOULD HAVE CONTRIBUTED TO THE COMPLAINT. A REVIEW OF THE IFU AND TRAINING MANUALS REVEALED NO DEFICIENCIES. FURTHERMORE, NO ABNORMALITIES WERE REPORTED DURING DEVICE UNPACKING OR PREPARATION. AS REPORTED, ''26MM SAPIEN 3 ULTRA IN AORTIC POSITION BY TRANSAXILLARY APPROACH. DURING INFLATION IT WAS NOTED THAT THE BALLOON COULD NOT BE INFLATED, SO THE VALVE DEPLOYMENT COULD NOT BE DONE. IT WAS ALSO OBSERVED SOME BLOOD IN THE SYRINGE. IT WAS DECIDED TO REMOVE THE VALVE THROUGH THE TRANSAXILLARY ACCESS BUT IT WAS NOT POSSIBLE. IT WAS DECIDED TO MANAGE DE REMOVAL THROUGH THE TRANSFEMORAL APPROACH BUT, THE PATIENT HAD THROMBOSIS IN THE FEMORAL AND DURING THE WITHDRAWAL A THROMBUS EMBOLIZED''. FOR THE COMPLAINT OF DELIVERY SYSTEM LEAKAGE, AS THE DELIVERY SYSTEMS ARE 100% TESTED FOR LEAKAGE, AND AS THIS ISSUE WAS UNDETECTED DURING PREPARATION, IT IS LIKELY THAT DAMAGE TO THE DELIVERY SYSTEM OCCURRED DURING THE PROCEDURE. IT IS POSSIBLE THAT DAMAGE TO THE BALLOON DURING THE PROCEDURE OCCURRED AS THERE WAS OBSERVED BLOOD IN THE INFLATION DEVICE. HOWEVER, DUE TO THE LIMITED INFORMATION, A DEFINITE ROOT CAUSE IS UNABLE TO BE DETERMINED AT THIS TIME. NO MANUFACTURING NON-CONFORMANCES WERE IDENTIFIED DURING EVALUATION. NO IFU/LABELING/TRAINING MANUAL INADEQUACIES WERE IDENTIFIED. THEREFORE, NO CORRECTIVE OR PREVENTIVE ACTIONS ARE REQUIRED AT THIS TIME. FOR THE COMPLAINTS OF WITHDRAW DIFFICULTY, NO INFORMATION WAS PROVIDED ON PATIENT'S VESSELS CHARACTERISTICS (CALCIFICATION, TORTUOSITY, MINIMUM LUMINAL DIAMETER). THE PRESENCE OF CALCIFICATION, TORTUOSITY, AND UNDERSIZED VESSEL CAN CREATE CHALLENGING PATHWAY DURING THE WITHDRAWAL OF THE DELIVERY SYSTEM WITH CRIMPED THV. ADDITIONALLY, BECAUSE THE DELIVERY SYSTEM WAS WITHDRAWN AFTER VALVE ALIGNMENT, THE OUTER DIAMETER OF THE CRIMPED VALVE ON THE INFLATION BALLOON IS LARGER THAN THAT OF THE CRIMPED VALVE ON THE CRIMP BALLOON, WHICH CAN FURTHER CONTRIBUTE TO THE WITHDRAWAL DIFFICULTY. TRAINING MANUAL INDICATED THAT ''CRIMPED THV ALIGNED ON BALLOON IS LARGER THAN CRIMPED THV OFF BALLOON. TAKE CARE IF DECIDING TO RETRIEVE'', AND ''DO NOT FORCE THE THV INTO THE SHEATH. IF RESISTANCE IS FELT, THE THV MAY BE CAUGHT ON THE SHEATH TIP''. DUE TO THE LIMITED INFORMATION, A DEFINITE ROOT CAUSE IS UNABLE TO BE DETERMINED; HOWEVER, AVAILABLE INFORMATION INDICATES THAT PROCEDURAL FACTORS (WITHDRAWAL OF CRIMPED VALVE) MAY HAVE CONTRIBUTED TO THE REPORTED EVENT AND SUBSEQUENT THROMBUS EMBOLIZATION AND DEATH. NO MANUFACTURING NON-CONFORMANCES WERE IDENTIFIED DURING EVALUATION. NO IFU/LABELING/TRAINING MANUAL INADEQUACIES WERE IDENTIFIED. THEREFORE, NO CORRECTIVE OR PREVENTIVE ACTIONS ARE REQUIRED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22196637&pc=MVK | 3008642652-2025-05776 | 2025-04-12 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-12 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO NON-LIFEVEST DEFIBRILLATIONS. THE DEVICE WAS STARTED UP AT 19:46:36 ON (B)(6) 2025. AT 22:56:45, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT. THE DEVICE PROPERLY DETECTED VF. CPR/MOTION ARTIFACT PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. AT 23:00:12, THE PATIENT RECEIVED THE FIRST NON-LIFEVEST DEFIBRILLATION. RHYTHM AT THE TIME OF THE FIRST NON-LIFEVEST DEFIBRILLATION WAS VF WITH CPR MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 23:04:20, THE PATIENT RECEIVED THE SECOND NON-LIFEVEST DEFIBRILLATION. RHYTHM AT THE TIME OF THE SECOND NON-LIFEVEST DEFIBRILLATION WAS VF WITH CPR MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 23:11:16 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22196785&pc=DXY | 2182208-2025-02859 | 2018-12-26 05:00 | Death | MEDTRONIC, INC. | 2025-06-12 04:00 | DXY | MEDTRONIC ICD | Inappropriate/Inadequate Shock/Stimulation | Shock from Patient Lead(s); Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS DECEASED. THE PATIENT REPRESENTATIVE REPORTED THAT THEIR FATHER PASSED AWAY WITHOUT THE FAMILY KNOWING BECAUSE THE MONITOR FAILED TO NOTIFY THE FAMILY IN THE HOUSE. IT WAS REPORTED THAT THE "MACHINE DIDN'T EVEN REGISTER THE OCCURRENCE." THE PATIENT REPRESENTATIVE CALLED THE PHYSICIAN'S OFFICE WHO STATED THAT THEY DID NOT RECEIVE ANY INDICATION THAT THERE WAS A PROBLEM OCCURRING. IT WAS ALSO NOTED THAT AT FIRST, THE DEVICE WAS "TOO SENSITIVE" AND WOULD SHOCK THEIR FATHER WITHOUT A REASON. THE PHYSICIAN HAD TO "RECALIBRATE" THE DEVICE. THE DISPOSITION OF THE DEVICE IS UNKNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22196986&pc=DQY | 2124215-2025-38817 | 2025-05-21 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-12 04:00 | DQY | WATCHMAN TRUSTEER? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiogenic Shock; Cardiac Perforation; Pericardial Effusion | K240018 |  | 1 | Event Description: REPORTED VIA STUDY. IT WAS REPORTED THAT THE PATIENT DIED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN POSITIONED THE WAS IN THE LAA OF THE PATIENT AND THEN INSERTED THE WDS. THE CLOSURE DEVICE WAS DEPLOYED IN THE PATIENT BUT THE PHYSICIAN WAS UNABLE TO POSITION THE DEVICE PROPERLY IN THE LAA. THE WDS WAS PULLED BACK INTO THE RIGHT ATRIUM AND THE PHYSICIAN PERFORMED A SECOND TRANSSEPTAL PUNCTURE. SHORTLY, AFTER THIS WAS PERFORMED A PERICARDIAL EFFUSION WITH CARDIAC TAMPONADE WAS NOTED. THE PATIENTS BLOOD PRESSURE DROPPED WHICH REQUIRED VASOPRESSOR SUPPORT AND THE PATIENT RECEIVED A BLOOD TRANSFUSION. THE PROCEDURE WAS ENDED, AND THE PATIENT WAS BROUGHT TO SURGERY WHERE THEY UNDERWENT A PERICARDIAL WINDOW PROCEDURE WHICH WAS CONVERTED A FULL STERNOTOMY. THE PHYSICIAN USED SUTURE TO CLOSE A PERFORATION IN THE LAA OF THE PATIENT AND THE LAA WAS CLIPPED. THE PATIENT REMAINED IN THE HOSPITAL TO RECOVER FROM THIS EVENT. WHILE IN THE HOSPITAL THE PATIENT BECAME SEPTIC AND EXPERIENCED CARDIOGENIC SHOCK. EIGHT DAYS POST PROCEDURE THE PATIENT HAD AN INCREASED DEMAND FOR PRESSOR SUPPORT. NINE DAYS POST PROCEDURE THE PATIENT WAS EXPERIENCING PERSISTENT LACTIC ACIDOSIS, ACUTE KIDNEY INJURY RESPIRATORY FAILURE AND THEN MULTIPLE SYSTEM ORGAN FAILURES BEFORE THE PATIENT ULTIMATELY DIED.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN POSITIONED THE WAS IN THE LAA OF THE PATIENT AND THEN INSERTED THE WDS. THE CLOSURE DEVICE WAS DEPLOYED IN THE PATIENT BUT THE PHYSICIAN WAS UNABLE TO POSITION THE DEVICE PROPERLY IN THE LAA. THE WDS WAS PULLED BACK INTO THE RIGHT ATRIUM AND THE PHYSICIAN PERFORMED A SECOND TRANSSEPTAL PUNCTURE. SHORTLY, AFTER THIS WAS PERFORMED A PERICARDIAL EFFUSION WITH CARDIAC TAMPONADE WAS NOTED. THE PATIENTS BLOOD PRESSURE DROPPED WHICH REQUIRED VASOPRESSOR SUPPORT AND THE PATIENT RECEIVED A BLOOD TRANSFUSION. THE PROCEDURE WAS ENDED, AND THE PATIENT WAS BROUGHT TO SURGERY WHERE THEY UNDERWENT A PERICARDIAL WINDOW PROCEDURE WHICH WAS CONVERTED A FULL STERNOTOMY. THE PHYSICIAN USED SUTURE TO CLOSE A PERFORATION IN THE LAA OF THE PATIENT AND THE LAA WAS CLIPPED. THE PATIENT REMAINED IN THE HOSPITAL TO RECOVER FROM THIS EVENT. WHILE IN THE HOSPITAL THE PATIENT BECAME SEPTIC AND EXPERIENCED CARDIOGENIC SHOCK. EIGHT DAYS POST PROCEDURE THE PATIENT HAD AN INCREASED DEMAND FOR PRESSOR SUPPORT. NINE DAYS POST PROCEDURE THE PATIENT WAS EXPERIENCING PERSISTENT LACTIC ACIDOSIS, ACUTE KIDNEY INJURY RESPIRATORY FAILURE AND THEN MULTIPLE SYSTEM ORGAN FAILURES BEFORE THE PATIENT ULTIMATELY DIED. THIS EVENT WAS ALSO REPORTED UNDER REPORT NUMBER 2124215-2025-34314. Manufacturer Narrative: H2 CORRECTION: THIS EVENT WAS ALSO REPORTED UNDER REPORT NUMBER 2124215-2025-34314. THIS COMPLAINT WILL BE CLOSED AS A DUPLICATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22197015&pc=DSQ | 2916596-2025-03764 | 2024-10-27 04:00 | Death | THORATEC CORPORATION | 2025-06-12 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO STROKE. THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22197053&pc=DSQ | 2916596-2025-03608 | 2018-01-23 05:00 | Death | THORATEC CORPORATION | 2025-06-12 04:00 | DSQ | HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM, JP | Pumping Stopped; Gas/Air Leak; Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Low Blood Pressure/ Hypotension | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ATRIOVENTRICULAR FIX AND TORE THE AORTA SO HARD TO DO EMERGENCY HOMOGRAFT INSERTION. LOTS OF AIR WENT INTO THE PUMP. A PUMP STOPPAGE OCCURRED AT THE TIME OF THE OPERATION. DURING THE ANALYSIS OF THE LOG FILE, PUMP SOPS WERE OBSERVED WHILE THE PUMP WAS AT 6000RPM WHICH DID NOT SEEM TO BE AN ISSUE. AFTER THE SPEED WAS BROUGHT UP TO 8000RPM, THE PUMP WAS RUNNING AS EXPECTED AND THE SET SPEED WAS AT 8800 RPM. IT WAS NOTED THAT THE SYSTEM CONTROLLER HAD 28 DAYS ACCUMULATED ON THE TIMER. THE PATIENT HAD A NEUROLOGICAL DYSFUNCTION, RIGHT HEMISPHERE SUBDURAL HEMATOMAS. COMPUTED TOMOGRAPHY WAS PERFORMED, AND STROKE WAS NOTED. THE PATIENT WAS ON BLOOD THINNER TREATMENT, ASPIRIN AT THE TIME OF THE EVENT. AN EMERGENCY SURGICAL TREATMENT WAS PERFORMED. THE ANESTHESIA CAUSED A DROP IN BLOOD PRESSURE, AND A CONTINUOUS INFUSION OF BOSMIN WAS STARTED. THE PATIENT'S BLOOD PRESSURE DID NOT IMPROVE AFTER SURGERY AND THERE WAS NO IMPROVEMENT AFTER TREATMENT WITH PITRESSIN AND NOA; HOWEVER, THE PATIENT'S CONDITION WORSENED AND ULTIMATELY PASSED AWAY DUE TO THE NEUROLOGIC DISORDER. THE EVENT DID NOT HAVE CAUSAL RELATIONSHIP WITH THE DEVICE. IT WAS UNKNOWN IF THE DEVICE FUNCTIONED PROPERLY AT THE TIME OF DEATH. THE DEVICE WAS EXPLANTED. THE PATIENT'S OPERATION AND PUMP STOP WAS PREVIOUSLY REPORTED UNDER MFR 2916596-2018-00508. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22197212&pc=DSQ | 2916596-2025-03798 | 2024-11-25 05:00 | Death | THORATEC CORPORATION | 2025-06-12 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SPONTANEOUS INTRACRANIAL BLEEDING (ICB). THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE PATIENT SUFFERED FROM A SPONTANEOUS ICB WHICH WAS REFRACTORY AND RELATED TO THE OUTCOME AT LEAST. NO ANTICOAGULATION ISSUES OR DEVICE RELATED ISSUES WERE MENTIONED PRIOR TO THE OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22197242&pc=DSQ | 2916596-2025-03821 | 2025-05-31 04:00 | Death | THORATEC CORPORATION | 2025-06-12 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22197307&pc=LWQ | 3019892983-2025-01012 | 2024-09-16 04:00 | Death | CORCYM S.R.L. | 2025-06-12 04:00 | LWQ | CARBOMEDICS TOP HAT MECHANICAL HEART VALVE | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: MANUFACTURER WAS INFORMED OF THIS EVENT THROUGH DEVICE TRACKING AUDIT CONDUCTED. BASED ON THE INFORMATION AVAILABLE, ON (B)(6) 2021, A CARBOMEDICS TOP-HAT AORTIC HEART VALVE, S5-025 WAS IMPLANTED IN A PATIENT. AS REPORTED, THE PATIENT DIED ON (B)(6) 2024. NO OTHER INFORMATION WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22197327&pc=LWQ | 3005687633-2025-00117 | 2024-09-16 04:00 | Death | CORCYM S.R.L. | 2025-06-12 04:00 | LWQ | CARBOMEDICS TOP HAT MECHANICAL HEART VALVE | Insufficient Information | Insufficient Information | P900060 |  | 1 | Event Description: MANUFACTURER WAS INFORMED OF THIS EVENT THROUGH DEVICE TRACKING AUDIT CONDUCTED. BASED ON THE INFORMATION AVAILABLE, ON (B)(6) 2021, A CARBOMEDICS TOP-HAT AORTIC HEART VALVE, S5-025 WAS IMPLANTED IN A PATIENT. AS REPORTED, THE PATIENT DIED (B)(6) 2024. NO OTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: THE MANUFACTURING AND MATERIAL RECORDS FOR THE DEVICE INVOLVED IN THIS EVENT, AS THEY PERTAIN TO THE REPORTED ISSUE, WERE RETRIEVED AND REVIEWED BY THE MANUFACTURER¿S QUALITY ENGINEERING. THE RESULTS CONFIRMED THAT THIS PROSTHESIS SATISFIED ALL REQUIRED MATERIAL, VISUAL, AND PERFORMANCE STANDARDS AT THE TIME OF MANUFACTURE AND RELEASE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEFINITIVE CAUSE OF THE REPORTED EVENT CANNOT BE ESTABLISHED AT THIS TIME. HOWEVER, FROM THE DOCUMENT REVIEW PERFORMED, NO MANUFACTURING DEFICIENCIES WERE NOTED. SHOULD FURTHER INFORMATION BE RECEIVED IN THE FUTURE, A FOLLOW UP REPORT WILL BE PROVIDED. Manufacturer Narrative: MANUFACTURER WILL SUBMIT FOLLOW-UP REPORT UPON AVAILABILITY OF NEW, RELEVANT DETAILS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22197411&pc=LWP | 2017865-2025-87106 | 2024-11-16 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-12 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | High Blood Pressure/ Hypertension | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIOVASCULAR DISEASES AND HEART BLOCK. THE PACEMAKER WAS EXPLANTED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE OR LEAD RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO PATIENT EXPIRED. THE DEVICE WAS RECEIVED WITH A BATTERY LEVEL WITHIN THE NORMAL RANGE OF OPERATION. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. AFTER ALL ELECTRICAL TESTS PERFORMED, INCLUDING THERMAL AND MECHANICAL STRESS TESTING THE DEVICE EXHIBITS NORMAL CHARACTERISTICS. LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY. NO ANOMALIES WERE FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22197421&pc=NVN | 2017865-2025-87108 | 2024-11-16 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-12 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | High Blood Pressure/ Hypertension | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIOVASCULAR DISEASES AND HEART BLOCK. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE OR LEAD RELATED. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22197422&pc=NVN | 2017865-2025-87109 | 2024-11-16 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-12 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | High Blood Pressure/ Hypertension | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIOVASCULAR DISEASES AND HEART BLOCK. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE OR LEAD RELATED. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22197618&pc=MQV | 2249852-2025-00010 | 2024-11-20 05:00 | Death | COLLAGEN MATRIX, INC. | 2025-06-12 04:00 | MQV | BIA-GEN | Adverse Event Without Identified Device or Use Problem | Swelling/ Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K182074 |  | 1 | Event Description: COMPLAINT SUBMITTED THROUGH CLINICAL DATA COLLECTION. LIMITED INFORMATION PROVIDED. PATIENT WAS INVOLVED IN A CLINICAL STUDY FOR PSEUDOCLAUDICATION SYNDROME. INITIAL SURGERY WAS PERFORMED ON (B)(6) 2024 UTILIZING THE BIA-GEN BIOACTIVE MOLDABLE BONE GRAFT MATRIX. THE PRODUCT WAS PLACED IN VERTEBRAE LOCATIONS L2-L3-L4-L5. THERE WERE NO OTHER PRODUCTS USED IN CONJUNCTION WITH BIA-GEN DURING THE INITIAL SURGERY AND NO COMPLICATIONS NOTED. THE CLINICIAN NOTED THE PRDOUCT HAD HANDLED AND PERFORMED EXCELLENT. ON (B)(6) 2024, THE PATIENT WENT TO THE EMERGENCY DEPARTMENT FOR EMESIS, LEFT FOOT SWELLING AND EPIGASTRIC DISCOMFORT. PATIENT THEN BECAME PALE, UNRESPONSIVE, AND EXPIRED. CLINICIAN BELIEVES CAUSE OF DEATH TO BE CARDIAC ARREST, HYPOXIC ARREST, OR A MISSED MYOCARDIAL INFRACTION. CLINICIAN STATED THE SERIOUS ADVERSE EVENT IS UNRELATED TO THE BIA-GEN PRODUCT. NO FURTHER INFORMATION SURROUNDING PATIENT INFORMATION AND EVENT WILL BE PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22198148&pc=MVK | 3008642652-2025-05786 | 2025-05-26 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-12 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE APPROPRIATE TREATMENTS AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 06:50:08 ON (B)(6) 2025. AT 07:51:14, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS TACHYCARDIA @ 100 BPM WITH PVC¿S. THE RHYTHM THEN DEGRADED TO VF. AT 07:51:50, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM WITH PAC¿S. AT 08:24:45, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF. AT 08:25:17, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM. AT 08:25:56, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM WITH PAC¿S. AT 08:26:35, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 30 BPM WITH PAC¿S. THE ELECTRODE BELT WAS DISCONNECTED AT 09:06:45 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22198362&pc=BZD | 2518422-2025-106577 | 2024-02-28 05:00 | Death | RESPIRONICS, INC. | 2025-06-12 04:00 | BZD | PHILIPS CPAP DEVICE | Degraded | Respiratory Failure; Swollen Lymph Nodes/Glands; Heart Failure/Congestive Heart Failure; Cough |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT EXPERIENCED A HACKING COUGH, ENLARGED LYMPH NODES IN THE TRACHEA AND EVIDENCE OF BLACK POWDER IN TUBING. THE PATIENT PASSED AWAY AFTER SUFFERING FROM CARDIOPULMONARY FAILURE. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT EXPERIENCED A HACKING COUGH, ENLARGED LYMPH NODES IN THE TRACHEA AND EVIDENCE OF BLACK POWDER IN TUBING. THE PATIENT PASSED AWAY AFTER SUFFERING FROM CARDIOPULMONARY FAILURE. THE PREVIOUS REPORT INCORRECTLY STATED THE REPORTER COUNTRY WAS "UNITED STATES" THIS HAS BEEN UPDATED IN BOX E TO "ITALY". Manufacturer Narrative: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT EXPERIENCED A HACKING COUGH, ENLARGED LYMPH NODES IN THE TRACHEA AND EVIDENCE OF BLACK POWDER IN TUBING. THE PATIENT PASSED AWAY AFTER SUFFERING FROM CARDIOPULMONARY FAILURE. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED, BECAUSE THE CASE WAS REPORTED FROM LEGAL PROCEEDINGS WITH NO SN AND NO INFORMATION IS AVAILABLE IF DEVICE CAN BE RETURNED FOR INVESTIGATION AT THIS MOMENT IN TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22198928&pc=LJE | 2124215-2025-38055 | 2025-03-03 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-12 04:00 | LJE | PERCUFLEX URINARY DIVERSION STENT SET | Break; Difficult to Remove | Sepsis; Foreign Body In Patient | K830803 |  | 1 | Event Description: IT WAS REPORTED THAT THE FRACTURED SINGLE J STENT WAS UNABLE TO BE REMOVED AT THE TIME OF STENT CHANGE RESULTING IN RETAINED FRAGMENT WITH SUBSEQUENT SEPSIS, AND DEATH.Event Description: IT WAS REPORTED THAT THE FRACTURED SINGLE J STENT WAS UNABLE TO BE REMOVED AT THE TIME OF STENT CHANGE RESULTING IN RETAINED FRAGMENT WITH SUBSEQUENT SEPSIS, AND DEATH. Manufacturer Narrative: BLOCK H6: IMDRF DEVICE CODE A0401 CAPTURES THE REPORTABLE EVENT OF STENT SHAFT BREAK. IMDRF PATIENT CODE E0306 CAPTURES THE REPORTABLE EVENT OF SEPSIS. IMDRF PATIENT CODE E2008 CAPTURES THE REPORTABLE EVENT OF FOREIGN BODY IN PATIENT. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF DEATH. Manufacturer Narrative: BLOCK H6: IMDRF DEVICE CODE A0401 CAPTURES THE REPORTABLE EVENT OF STENT SHAFT BREAK. IMDRF PATIENT CODE E0306 CAPTURES THE REPORTABLE EVENT OF SEPSIS. IMDRF PATIENT CODE E2008 CAPTURES THE REPORTABLE EVENT OF FOREIGN BODY IN PATIENT. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF DEATH. BLOCK H11: THIS REPORT IS REPORT IS BEING SUBMITTED TO CORRECT THE INIT RPTR ALSO SENT REP TO FDA QUESTION (E4) SECTION E HAS BEEN CHANGE TO YES. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UDI AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22199504&pc=CBK | 2023050-2025-00029 | 2025-05-17 04:00 | Death | NEWPORT MEDICAL INSTRUMENT, INC. | 2025-06-12 04:00 | CBK | HT70 VENTILATOR | No Flow | Respiratory Failure | K111146 |  | 1 | Event Description: IT WAS REPORTED THAT WHILE IN USE ON A CRITICALLY ILL PATIENT, THE HT70 VENTILATOR GENERATED AN ALARM, THE SCREEN WENT COMPLETELY DARK AND THE VENTILATION STOPPED. THE PATIENT HAD NO SPONTANEOUS BREATHING, WAS UNDER A DO NOT RESUSCITATE (DNR) ORDER, AND PASSED AWAY. Manufacturer Narrative: SECTION A PATIENT INFORMATION AND SECTION E INITIAL REPORTER CANNOT BE PROVIDED DUE TO JAPANESE PRIVACY REGULATIONS. SECTION E. JAPAN MEDTRONIC PERSONNEL REPORTED EVENT TO MANUFACTURER ON BEHALF OF THE CUSTOMER. H3: MEDTRONIC HAS NOT RECEIVED THE SUSPECT DEVICE/COMPONENT FROM THE CUSTOMER FOR EVALUATION NOR HAS THE DEVICE BEEN EVALUATED BY THE MEDTRONIC SERVICE ENGINEER. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22181956&pc=LNH | 3018423337-2025-00004 | 2025-05-27 04:00 | Death | FUJIFILM CORPORATION | 2025-06-11 04:00 | LNH | ECHELON OVAL MRI SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K221619 |  | 1 | Event Description: ON MAY 29TH, 2025, FUJIFILM CORPORATION WAS INFORMED OF AN EVENT INVOLVING ECHELON OVAL MRI SYSTEM. PATIENT HAS DIED WHILE BEING SCANNED FOR ABDOMEN ON ECHELON OVAL MRI SYSTEM. Manufacturer Narrative: ON JUNE 25, 2025, A REPORT WAS RECEIVED FROM THE IMPORTER STATING THAT THERE WERE NO ISSUES WITH THE IMAGE QUALITY OF THE DEVICE. WE DETERMINED THAT THE DEVICE WAS OPERATING NORMALLY BECAUSE THERE WERE NO DEVICE ERROR LOGS AND THERE WERE NO IMAGE QUALITY ISSUES. Manufacturer Narrative: THE INVESTIGATION IS CURRENTLY ONGOING. THE INFORMATION CURRENTLY AVAILABLE IS AS FOLLOWS. 1. WE HAVE RECEIVED REPORTS FROM MEDICAL TECHNICIANS THAT THE PATIENT HAD MANY MEDICAL PROBLEMS. 2.THERE WERE NO DEVICE ERROR LOGS. 3. WE HAVE RECEIVED REPORTS THAT THE DEVICE WAS OPERATING NORMALLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22182341&pc=PNJ | 9612164-2025-02843 | 2023-08-25 04:00 | Death | MEDTRONIC IRELAND | 2025-06-11 04:00 | PNJ | MICRA | Electrical /Electronic Property Problem; Use of Device Problem | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) POOR NUMBER WERE NOTED AFTER THE FIRST DEPLOYMENT. THERE WAS DIFFICULTY REALIGNING THE CUP AND CONE, THEN THE DEPLOYMENT BUTTON WOULD NOT SLIDE FORWARD. THE LEADLESS IPG WAS RECAPTURED. UNDER FLUOROSCOPY THERE APPEARED TO BE A SPACE BETWEEN THE CUP AND CONE. DEPLOYMENT WAS DONE AGAIN WITH SIMILAR POOR TESTING R ESULTS. RECAPTURE WAS DONE WITH SOME LIMITED RESISTANCE. IT WAS ALSO NOTED THAT THERE WAS A CLOT ISSUE RELATED TO THE LEADLESS IPG AND DELIVERY SYSTEM. THE LEADLESS IPG WAS NOT USED AND WAS REPLACED. THE REPLACEMENT LEADLESS IPG ALSO EXHIBITED POOR ALIGNMENT BETWEEN THE CUP AND CONE. THERE WAS RESISTANCE NOTED WITH THE DEPLOYMENT BUTTON. THE LEADLESS IPG WAS SUCCESSFULLY IMPLANTED. IT WAS FURTHER NOTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22182356&pc=KRH | 9612164-2025-02844 | 2024-09-18 04:00 | Death | MEDTRONIC MEXICO | 2025-06-11 04:00 | KRH | TRI-AD BAND | Insufficient Information | Insufficient Information | K181132 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT AFTER AN UNKNOWN DURATION POST IMPLANT OF THIS 28MM TRI-AD BAND, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE TRI-AD BAND COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: UPDATED DATA: A.3A: SEX: A.3B: GENDER: A.4: WEIGHT IN LBS: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22182460&pc=KRH | 9612164-2025-02848 | 2024-11-03 04:00 | Death | MEDTRONIC MEXICO | 2025-06-11 04:00 | KRH | TRI-AD BAND | Insufficient Information | Insufficient Information | K181132 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT AFTER AN UNKNOWN DURATION POST IMPLANT OF THIS 28MM TRI-AD BAND, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE TRI-AD BAND COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED THAT 3 YEARS POST IMPLANT OF THIS 28 MM TRI-AD BAND, THE PATIENT DIED. Manufacturer Narrative: SECTIONS UPDATED: B5, G6, H2. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: A.3A: SEX: A.3B: GENDER: A.4: PATIENT WEIGHT B.2: DATE OF DEATH: B.3: DATE OF EVENT: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22183017&pc=BZD | 2518422-2025-106535 | 2025-06-05 04:00 | Death | RESPIRONICS, INC. | 2025-06-11 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED NOTIFICATION THAT THE SALES REPRESENTATIVE WAS INFORMED BY THE DEALER OF A PATIENT DEATH CONFIRMED ON (B)(6) 2025, WHILE THE PATIENT WAS USING THE DEVICE. AS THE DEVICE WAS IN USE AT THE TIME, IT HAS BEEN SEIZED BY THE POLICE. REPORTEDLY, THE PRESCRIBING PHYSICIAN HAS BEEN INTERVIEWED BY AUTHORITIES. AT PRESENT, THERE IS NO INDICATION OF A CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE PATIENT¿S DEATH. NO PERSONALLY IDENTIFIABLE PATIENT INFORMATION HAS BEEN DISCLOSED. THE DURATION OF DEVICE USE BY THE PATIENT REMAINS UNKNOWN. THE MANDATORY QUESTION, ¿DID THE ALARM SOUND AT THE TIME OF THE EVENT?¿ WAS MARKED AS ¿UNKNOWN.¿ ACCORDING TO THE DEALER, THE POLICE HAVE INDICATED THAT THE DEVICE IS OPERATIONAL. HOWEVER, THE EXACT CONDITION OF THE DEVICE AT THE TIME OF THE INCIDENT HAS NOT BEEN CLEARLY ESTABLISHED. NO DEVICE MALFUNCTION, OPERATIONAL STOP, OR ERROR DISPLAY HAS BEEN CONFIRMED. AS THE DEVICE IS CURRENTLY IN POLICE CUSTODY, IT IS UNAVAILABLE FOR FURTHER INVESTIGATION. NO MEDICAL INTERVENTION RELATED TO THE EVENT HAS BEEN REPORTED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED NOTIFICATION THAT THE SALES REPRESENTATIVE WAS INFORMED BY THE DEALER OF A PATIENT DEATH CONFIRMED ON (B)(6) 2025, WHILE THE PATIENT WAS USING THE DEVICE. AS THE DEVICE WAS IN USE AT THE TIME, IT HAD BEEN SEIZED BY THE POLICE. REPORTEDLY, THE PRESCRIBING PHYSICIAN HAD BEEN INTERVIEWED BY AUTHORITIES. AT PRESENT, THERE IS NO INDICATION OF A CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE PATIENT¿S DEATH. NO PERSONALLY IDENTIFIABLE PATIENT INFORMATION HAD BEEN DISCLOSED. THE DURATION OF DEVICE USE BY THE PATIENT REMAINS UNKNOWN. THE MANDATORY QUESTION, ¿DID THE ALARM SOUND AT THE TIME OF THE EVENT?¿ WAS MARKED AS ¿UNKNOWN.¿ ACCORDING TO THE DEALER, THE POLICE HAVE INDICATED THAT THE DEVICE IS OPERATIONAL. HOWEVER, THE EXACT CONDITION OF THE DEVICE AT THE TIME OF THE INCIDENT HAD NOT BEEN CLEARLY ESTABLISHED. NO DEVICE MALFUNCTION, OPERATIONAL STOP, OR ERROR DISPLAY HAD BEEN CONFIRMED. AS THE DEVICE IS CURRENTLY IN POLICE CUSTODY, IT WAS UNAVAILABLE FOR FURTHER INVESTIGATION. NO MEDICAL INTERVENTION RELATED TO THE EVENT HAD BEEN REPORTED. THE DEVICE WAS RETURNED FOR INVESTIGATION. AFTER CONDUCTING FUNCTIONAL CHECKS AND INTERNAL INSPECTIONS, NO ABNORMALITIES WERE FOUND IN THE EQUIPMENT ITSELF. ODORS RELATED TO EVERYDAY LIFE WERE CONFIRMED. UPON RECEIVING A REQUEST FROM THE CUSTOMER TO CANCEL THE REPAIR, THE REPAIR HAD BEEN HALTED, AND THE EQUIPMENT WILL BE RETURNED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22183024&pc=OZD | 1220648-2025-29186 | 2025-05-17 04:00 | Death | ABIOMED, INC. | 2025-06-11 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA 5.5 WAS PLACED ON A PATIENT AND WOULD BE WORKED UP FOR VENTRICULAR ASSIST DEVICE. INITIALLY DURING IMPELLA 5.5 SUPPORT IT WAS NOTED THAT THE PATIENT WAS DOING WELL AND HEMODYNAMICALLY STABLE. HOWEVER, SEVERAL DAYS AFTER SUPPORT STARTED THE PATIENT BEGAN TO HAVE A GASTROINTESTINAL BLEED AND HEPARIN WAS BEING HELD. ADDITIONALLY, THE PATIENT WAS CLOSTRIDIOIDES DIFFICILE POSITIVE. IT WAS NOTED THAT AFTER THAT, THE PATIENT CONTINUE TO HAVE BLEEDS WITH DARK STOOLS AND THE PATIENT RECEIVED SEVERAL UNITS OF BLOOD. ULTIMATELY, IT WAS NOTED THAT THE PATIENT BEGAN HAVING ABDOMEN PAIN RESULTING IN A COMPUTED TOMOGRAPHY THAT REVEALED FREE AIR IN ABDOMEN. GENERAL SURGERY WAS CONSULTED. HOWEVER, THE PATIENT BEGAN TO DECOMPENSATE QUICKLY. IT WAS BELIEVED THAT THE PATIENT HAD ISCHEMIC BOWEL AND SURGERY WAS DEEMED NOT TO BE AN OPTION. THE PATIENT CONTINUED TO DECOMPENSATE AND EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22183228&pc=BZD | 2518422-2025-106536 | 2025-06-10 04:00 | Death | RESPIRONICS, INC. | 2025-06-11 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22183329&pc=OZD | 1220648-2025-29191 | 2025-05-17 04:00 | Death | ABIOMED, INC. | 2025-06-11 04:00 | OZD | IMPELLA CP | Infusion or Flow Problem | Tachycardia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT ON IMPELLA CP SUPPORT HAD VENTRICULAR TACHYCARDIA/FIBRILLATION REQUIRING ONE SHOCK. AN ECHOCARDIOGRAM WAS DONE FOR PLACEMENT AFTER SHOCK AND NO ADJUSTMENTS WAS REQUIRED. FLOW SATURATION WAS ALSO NOTED THAT SAME DAY. TISSUE PLASMINOGEN ACTIVATOR WAS INITIATED. ADDITIONAL INFORMATION WAS PROVIDED THAT THE PATIENT CONTINUED TO DETERIORATE AND REQUIRED MORE PRESSOR/INOTROPE SUPPORT. AS A RESULT, THE FAMILY DECIDED TO WITHDRAW CARE. THE PATIENT EXPIRED SHORTLY AFTER. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22183518&pc=OZD | 1220648-2025-29192 | 2025-06-03 04:00 | Death | ABIOMED, INC. | 2025-06-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: INFORMATION WAS OBTAINED FROM A LITERATURE REPORT (B)(4) REPORTING THAT A PATIENT PRESENTED WITH ACUTE MYOCARDIAL INFARCTION. THE PATIENT WAS MANAGED WITH ECPELLA BECAUSE OF HEMODYNAMIC COMPROMISE. HOWEVER, A WEEK LATER, THERE WAS A SUDDEN INCREASE IN VENOUS OXYGEN SATURATION. TRANSTHORACIC ECHOCARDIOGRAPHY REVEALED VENTRICULAR SEPTAL RUPTURE, AND FREE WALL RUPTURE. INTRA-VENTRICULAR THROMBUS WAS ALSO OBSERVED DESPITE STANDARD ANTICOAGULATION THERAPY. EVEN WITH DOUBLE CARDIAC RUPTURE, ECPELLA COULD FACILITATE LEFT VENTRICULAR UNLOADING AND SUSTAIN HEMODYNAMICS. HOWEVER, BECAUSE OF THE RISK OF DEVICE FAILURE DUE TO THROMBUS ASPIRATION INTO THE LMPELLA CP, THE PATIENT UNDERWENT REPAIR SURGERY. POSTOPERATIVELY, THE PATIENT WAS TEMPORARILY WEANED OFF OF ECPELLA, AND HIS HEMODYNAMICS DETERIORATED AGAIN, AND EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: A.1, A.4, AND A.5 ARE UNKNOWN. D.4 CATALOG NUMBER, SERIAL NUMBER, LOT NUMBER, EXPIRATION DATE AND UNIQUE IDENTIFIER (UDI) # ARE UNKNOWN. D.6A AND D.6B IMPLANT AND EXPLANT DATES ARE UNKNOWN. E.1 NAME PREFIX/TITLE, FIRST NAME AND LAST NAME ARE UNKNOWN. H.4 DEVICE MANUFACTURE DATE IS UNKNOWN. THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: WARNINGS & CAUTIONS: CAUTIONS. ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION. SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22183537&pc=PNJ | 9612164-2025-02859 | 2025-05-21 04:00 | Death | MEDTRONIC IRELAND | 2025-06-11 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO IMPLANT A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), AN ISSUE OCCURRED DURING MULTIPLE DEPLOYMENTS. DESPITE THE DEPLOYMENT BUTTON BEING FULLY PULLED, DEPLOYMENT OF THE IPG WAS NOT POSSIBLE. PULLING THE DELIVERY SYSTEM DID NOT ENABLE DEPLOYMENT, BUT RETRACTION WAS STILL POSSIBLE. WHEN THE DELIVERY SYSTEM WAS REMOVED FROM THE BODY, DEPLOYMENT BECAME POSSIBLE. ANOTHER ATTEMPT TO DEPLOY THE IPG INSIDE THE BODY FAILED. THE IPG SYSTEM WAS ATTEMPTED/NOT USED AND WAS REPLACED WITH A SECOND LEADLESS IPG FOR PLACEMENT. IT WAS REPORTED THAT POST THE IMPLANT PROCEDURE, THE PATIENT EXPERIENCED CARDIAC PERFORATION, CARDIAC TAMPONADE, PERICARDIAL EFFUSION, BLOOD PRESSURE DROP AND FLUCTUATING BLOOD PRESSURE. PERICARDIOCENTESIS DRAINAGE WAS PERFORMED. AFTERWARDS, PERICARDIAL EFFUSION WAS NOT OBSERVED, AND IT WAS DETERMINED THAT THERE WERE NO PROBLEMS WITH THE ECHOCARDIOGRAM, AND THE DRAIN WAS REMOVED APPROXIMATELY TWO DAYS LATER. SUBSEQUENTLY, THE BLOOD PRESSURE DID NOT INCREASE SUFFICIENTLY, BUT DIALYSIS WAS PERFORMED USING CONTINUOUS HEMODIAFILTRATION (CHDF) INSTEAD OF NORMAL DIALYSIS. IT WAS REPORTED THAT THE PATIENT DIED APPROXIMATELY FOUR DAYS POST IMPLANT OF THE LEADLESS IPG. CAUSE OF DEATH WAS PROVIDED AS NON-CARDIAC DEATH, HISTORY OF SEVERE AORTIC VALVE STENOSIS AND INSUFFICIENT BLOOD PRESSURE, WHICH PREVENTED CONVENTIONAL DIALYSIS, AND CHRONIC RENAL FAILURE. THE SECOND IPG REMAINS IN THE PATIENT. Manufacturer Narrative: CORRECTION H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6)). D9: NO. MC2AVR1 LEADLESS IPG; EXPLANTED (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22183553&pc=PNJ | 9612164-2025-02860 | 2025-05-21 04:00 | Death | MEDTRONIC IRELAND | 2025-06-11 04:00 | PNJ | MICRA¿ AV2 | Activation, Positioning or Separation Problem; Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE TO IMPLANT A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), AN ISSUE OCCURRED DURING MULTIPLE DEPLOYMENTS. DESPITE THE DEPLOYMENT BUTTON BEING FULLY PULLED, DEPLOYMENT OF THE IPG WAS NOT POSSIBLE. PULLING THE DELIVERY SYSTEM DID NOT ENABLE DEPLOYMENT, BUT RETRACTION WAS STILL POSSIBLE. WHEN THE DELIVERY SYSTEM WAS REMOVED FROM THE BODY, DEPLOYMENT BECAME POSSIBLE. ANOTHER ATTEMPT TO DEPLOY THE IPG INSIDE THE BODY FAILED. THE IPG SYSTEM WAS ATTEMPTED/NOT USED AND WAS REPLACED WITH A SECOND LEADLESS IPG FOR PLACEMENT. IT WAS REPORTED THAT POST THE IMPLANT PROCEDURE, THE PATIENT EXPERIENCED CARDIAC PERFORATION, CARDIAC TAMPONADE, PERICARDIAL EFFUSION, BLOOD PRESSURE DROP AND FLUCTUATING BLOOD PRESSURE. PERICARDIOCENTESIS DRAINAGE WAS PERFORMED. AFTERWARDS, PERICARDIAL EFFUSION WAS NOT OBSERVED, AND IT WAS DETERMINED THAT THERE WERE NO PROBLEMS WITH THE ECHOCARDIOGRAM, AND THE DRAIN WAS REMOVED APPROXIMATELY TWO DAYS LATER. SUBSEQUENTLY, THE BLOOD PRESSURE DID NOT INCREASE SUFFICIENTLY, BUT DIALYSIS WAS PERFORMED USING CONTINUOUS HEMODIAFILTRATION (CHDF) INSTEAD OF NORMAL DIALYSIS. IT WAS REPORTED THAT THE PATIENT DIED APPROXIMATELY FOUR DAYS POST IMPLANT OF THE LEADLESS IPG. CAUSE OF DEATH WAS PROVIDED AS NON-CARDIAC DEATH, HISTORY OF SEVERE AORTIC VALVE STENOSIS AND INSUFFICIENT BLOOD PRESSURE, WHICH PREVENTED CONVENTIONAL DIALYSIS, AND CHRONIC RENAL FAILURE. THE SECOND IPG REMAINS IN THE PATIENT. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1] (SERIAL: (B)(6). D9: <(><<)>NO> MC2AVR1 LEADLESS IPG, IMPLANT DATE (B)(6) 2025 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22184084&pc=OZD | 1220648-2025-29196 | 2025-05-17 04:00 | Death | ABIOMED, INC. | 2025-06-11 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA 5.5 WAS PLACED ON A PATIENT THAT HAD PLACEMENT SUCTION OCCUR WHEN THE PUMP WAS DECREASED TO P-6 LEVEL. THE PATIENT CONTINUED TO HAVE INTERMITTENT SUCTION EVENTS. THE IMPELLA 5.5 WAS REPOSITIONED, BUT SUCTION WAS NOT RESOLVED. IT IS BELIEVED THAT THE RIGHT SIDE OF THE HEART WAS FAILING. IT WAS FURTHER REPORTED THAT CARE WAS WITHDREW AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: COMPLAINT DEVICE NOT RETURNED FOR ANALYZING. DATA LOG REVIEWED: NO MC SPIKE OUTSIDE OF P-LEVEL CHANGE OBSERVED. MANY SUCTIONS OCCURRED DURING ENTIRE IMPELLA SUPPORT WITH IMPELLA POSITION UNKNOWN ALARMS. PLACEMENT SIGNAL SLIGHTLY TRENDING DOWN AT END OF SUPPORT. CLINICAL STATED: PATIENT HAD RV DYSFUNCTION AND VSD ISSUES, CONTINUES TO HAVE INTERMITTENT SUCTION EVENTS, MOST LIKELY RELATED TO RIGHT-SIDED FAILURE AND NEGATIVE VOLUME STATUS. REPEATED EVENTS AFTER A FEW DAYS OF SUPPORT DUE TO RIGHT SIDED DYSFUNCTION. PUMP REPOSITIONED DID NOT RESOLVE ISSUE. BLOOD VOLUME NOT GIVEN. NO CANNULA KINKS. NO LV THROMBUS OBSERVED. LOW PUMP FLOW: THE CAUSE OF LOW PUMP FLOW ISSUE MOST LIKELY WAS PATIENT CONDITION RELATED DUE TO RV DYSFUNCTION RIGHT SIDED FAILURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22184252&pc=DXF | 2124215-2025-38318 | 2025-05-21 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-06-11 04:00 | DXF | VERSACROSS CONNECT LAAC ACCESS SOLUTION | Difficult to Advance | Low Blood Pressure/ Hypotension; Cardiac Perforation; Multiple Organ Failure; Pericardial Effusion | K150709 |  | 1 | Event Description: IT WAS REPORTED A PERICARDIAL EFFUSION (PE) WAS NOTED. THE PROCEDURE WAS CANCELLED. THE PATIENT PASSED AWAY. DURING A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A VERSACROSS CONNECT LAAC WAS SELECTED FOR USE. THE APPENDAGE WAS HEAVILY PECTINATED AND SHALLOW. THE PHYSICIAN HAD DIFFICULTY PERFORMING TRANSSEPTAL PUNCTURE (TSP) DUE TO PATIENT HAVING MULTIPLE PREVIOUS ABLATIONS AND A SCARRED SEPTUM. TWO TSPS WERE PERFORMED. AFTER THE SECOND TSP, PRIOR TO ATTEMPTING ANOTHER WATCHMAN DEVICE, THE BLOOD PRESSURE DROPPED AND AN EFFUSION WAS NOTED ON TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). THERE WAS A PERICARDIAL EFFUSION NOTED AROUND THE RIGHT VENTRICLE AND APPENDAGE/TRANSVERSE SINUS. A PERFORATION WAS IN THE DISTAL END OF THE APPENDAGE. THE PATIENT WAS THEN SENT TO CARDIAC SURGERY AND HAD A STERNOTOMY. THE PATIENT WAS STILL INTUBATED AT THE HOSPITAL, BUT EXPECTED TO FULLY RECOVER. THE PATIENT DEVELOPED LACTIC ACIDOSIS TWO DAYS AFTER THE STERNOTOMY, AND HAD MULTIPLE ORGAN FAILURE. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (RETAINED). IT WAS REPORTED THAT THE VERSACROSS RF WIRE WAS TENTING LESS THAN IDEALLY, AND STILL CLEAR. THERE WERE NO MALFUNCTIONS REPORTER AND IN THE PHYSICIAN'S OPINION, THERE IS NO CONTRIBUTION FROM VERSACROSS DEVICES TO THE PATIENT COMPLICATIONS. ACT WAS THERAPEUTIC. A PERICARDIOCENTESIS WAS NOT DONE, SINCE THE PERFORATION WAS NOT IN A SPOT THAT COULD BE ACCESSED TO DO SO. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22184254&pc=NGV | 2124215-2025-37954 | 2025-05-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-11 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: REPORTED VIA CLINICAL STUDY. IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. THE PATIENT WAS ENROLLED IN A CLINICAL STUDY AND A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED ON (B)(6) 2024. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) WAS USED AND A 27MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS IMPLANTED WITH A COMPLETE LAA SEAL AND DEPLOYED DEVICE DIAMETER OF 24 MM. THE PATIENT WAS DISCHARGED THE SAME DAY ON A MEDICATION REGIME OF ASPIRIN AND CLOPIDOGREL. ON (B)(6) 2024, THE PATIENT PRESENTED FOR THE PROTOCOL SCHEDULED 45 DAYS FOLLOW UP VISIT. TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) REVEALED A COMPLETE SEAL, NO THROMBUS, NO PERICARDIAL EFFUSION, AND NO ATRIAL SEPTAL SHUNT. LEFT VENTRICULAR EJECTION FRACTION WAS 55%. ON AN UNKNOWN DATE POST INDEX PROCEDURE, THE PATIENT WAS REPORTED TO HAVE DIED. NO ADDITIONAL INFORMATION ON THE CAUSE OF DEATH WAS PROVIDED.Event Description: REPORTED VIA CLINICAL STUDY. IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. THE PATIENT WAS ENROLLED IN A CLINICAL STUDY AND A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED ON (B)(6) 2024. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) WAS USED AND A 27MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS IMPLANTED WITH A COMPLETE LAA SEAL AND DEPLOYED DEVICE DIAMETER OF 24 MM. THE PATIENT WAS DISCHARGED THE SAME DAY ON A MEDICATION REGIME OF ASPIRIN AND CLOPIDOGREL. ON (B)(6) 2024, THE PATIENT PRESENTED FOR THE PROTOCOL SCHEDULED 45 DAYS FOLLOW UP VISIT. TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) REVEALED A COMPLETE SEAL, NO THROMBUS, NO PERICARDIAL EFFUSION, AND NO ATRIAL SEPTAL SHUNT. LEFT VENTRICULAR EJECTION FRACTION WAS 55%. ON AN UNKNOWN DATE POST INDEX PROCEDURE, THE PATIENT WAS REPORTED TO HAVE DIED. NO ADDITIONAL INFORMATION ON THE CAUSE OF DEATH WAS PROVIDED. IT WAS FURTHER REPORTED THAT ON (B)(6) 2025, 330 DAYS POST INDEX PROCEDURE, THE PATIENT EXPIRED, AS REPORTED PER END OF STUDY FORM. THE CAUSE OF DEATH REMAINS UNKNOWN. Manufacturer Narrative: B2 DATE OF DEATH: AWARE DATE WAS USED AS DEATH DATE AS ACTUAL DEATH DATE WAS NOT KNOWN/PROVIDED. B3 DATE OF EVENT: AWARE DATE WAS USED AS EVENT DATE AS ACTUAL EVENT DATE WAS NOT KNOWN/PROVIDED. Manufacturer Narrative: B2 DATE OF DEATH ADDED. B3 DATE OF EVENT : EVENT DATE UPDATED. B5 DESCRIBE EVENT OR PROBLEM: UPDATED WITH ADDITIONAL INFORMATION RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22184859&pc=FTL | MW5171291 | 2019-10-23 04:00 | Death | BARD DOVAL / DAVOL INC., SUB. C.R. BARD, INC. | 2025-06-11 04:00 | FTL | HERNIA MESH (POLYPROPYLENE) | Defective Device; Contamination /Decontamination Problem; Unintended Movement; Patient Device Interaction Problem | Wound Dehiscence; Fistula; Neuropathy; Internal Organ Perforation; Pain; Sepsis; Urinary Tract Infection; Distress; Impaired Healing; Coma; Quadriplegia |  |  | 1 | Event Description: MY NAME IS (B)(6) THE SON OF (B)(6) WHO DIED AS A RESULT OF DEFECTIVE HERNIA MESH PRODUCTS, SPECIFICALLY BARD DOVAL AND COVIDIEN AS WELL AS A THIRD PRODUCT I BELIEVE. I WAS MY FATHER'S CARE PROVIDER FOR THE LAST THREE YEARS OF HIS LIFE AS HIS CONDITION BECAME SEVERE AND DEBILITATING. IN 2019 HE WENT IN FOR EMERGENCY SURGERY DUE TO A LIFE-THREATENING BOWEL BLOCKAGE THAT BECAME SEPTIC. IT TOOK 11 DAYS FOR HIM TO WAKE UP AND HE NEVER WALKED AGAIN FOR THE LAST TWO YEARS OF HIS LIFE. WHEN I MET WITH THE SURGEON DR. (B)(6) HE STATED "THIS WAS ONE OF THE MOST DIFFICULT SURGERIES HE EVER PERFORMED". BEAR IN MIND DR. (B)(6) HAD BEEN PERFORMING SURGERIES FOR 35 YEARS TO PUT THAT STATEMENT IN PERSPECTIVE. THE PRODUCTS USED IN MY FATHER'S HERNIA SURGERIES DESTROYED HIS INTERNAL ORGANS TO THE POINT THEY HAD TO BE REMOVED AND RESECTED. HIS ORGANS WERE DEAD, DYING, PERFORATED, AND SO ON. EVERYTHING I SPEAK ABOUT IN THIS EMAIL IS FULLY DOCUMENTED IN HIS MEDICAL RECORDS. I PERSONALLY SPENT 3 YEARS READING THROUGH EVERY PAGE AND GOING OVER THEM WITH A FINE TOOTH COMB. I SOLD MY HOME AND LEFT MY CAREER AS A MORTGAGE COMPLIANCE OFFICER TO LIVE WITH MY FATHER THE LAST 2-2.5 YEARS OF HIS LIFE TO TAKE CARE OF HIM 24/7 AS MY MOTHER, HIS WIFE PASSED AWAY IN 2011. AFTER HIS LAST HERNIA MESH REVISION SURGERY IN 2019 DUE TO A BOWEL BLOCKAGE HE NEVER RECOVERED. HE SUFFERED EXTREME NEUROPATHY, FUNCTIONAL QUADRIPLEGIA, LARGE ABDOMINAL WOUNDS WITH FISTULAS THAT I HAD TO PACK EVERY DAY, NUMEROUS UTI'S FROM HIS CATHETER, ETC... IN THE WEEKS FOLLOWING HIS SURGERY HIS SURGEON HAD TO PULL MESH THROUGH HIS FISTULAS DURING AN OFFICE VISIT DUE TO THE EXTREME PAIN HE WAS IN. I SPENT TWO YEARS LISTENING TO HIM CRYING OUT IN PAIN AND SOBBING ON A DAILY BASIS, AND UNLESS YOU'VE WATCHED A PARENT SUFFER LIKE THAT YOU HAVE NO IDEA WHAT IT DOES TO YOU. HIS WOUNDS NEVER HEALED DUE TO THE MESH INFECTION. I HAD TO CALL 911 OVER 20 TIMES IN THE LAST YEAR OF HIS LIFE AND HE HAD TO ENDURE THE PAIN OF TRAVELING TO THE HOSPITAL AND THE PAIN WAS EXCRUCIATING. HIS QUALITY OF LIFE WAS GONE COMPLETELY. MY FATHER DIED A VERY SLOW, AGONIZING, DEPRESSIVE, DEMORALIZING, TORTUOUS DEATH, TO PUT IT MILDLY. THESE ARE FACTS. TRUST ME, I HAD THE TRAUMATIZING TASK OF BEING THERE 24/7 THROUGHOUT THE ENTIRE 2 PLUS YEAR ORDEAL. MAKE NO MISTAKE HE SUFFERED FROM EVERY HERNIA MESH COMPLICATION YOU CAN FIND ON GOOGLE PLUS THE EXTREME NEUROPATHY, FUNCTION QUADRIPLEGIA, AND DEATH. HE COULD NOT WALK, FEED HIMSELF, BATHE, USE THE TOILET, SCRATCH HIS NOSE, PET HIS DOG, OR ANYTHING ELSE FOR THAT MATTER. NOBODY SHOULD EVER HAVE TO GO THROUGH THE HELL MY FATHER, MYSELF, AND MY FAMILY WENT THROUGH DUE TO THESE MEDICAL PRODUCTS. IT WAS A TRAGEDY OF EPIC PROPORTIONS FOR ALL OF US TO SAY THE VERY LEAST. REFERENCE REPORTS MW5171292, MW5171293.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22184888&pc=DSQ | 2916596-2025-03682 | 2025-04-01 04:00 | Death | THORATEC CORPORATION | 2025-06-11 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Shock; Vasodilatation | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿THE DOWNWARD SPIRAL OF HEMOCHROMATOSIS¿ THAT THE HEARTMATE 3 IS ASSOCIATED WITH VASOPLEGIA, REFRACTORY SHOCK, AND DEATH. THE ARTICLE DESCRIBED A CASE STUDY INVOLVING A 48-YEAR-OLD MALE WITH HEREDITARY HEMOCHROMATOSIS (HOMOZYGOUS C282Y) DIAGNOSED 20 YEARS AGO WHICH REQUIRED FREQUENT PHLEBOTOMIES, BUT LOST TO THERAPY FOLLOW UP. THEY PRESENTED WITH VOLUME OVERLOAD AND RIGHT SIDED VISUAL DEFICITS SECONDARY TO AN ACUTE EMBOLIC STROKE. ECHOCARDIOGRAM SHOWED BIVENTRICULAR HEART FAILURE WITH DILATED LEFT VENTRICULAR CAVITY (LV END DIASTOLIC DIAMETER 6.2CM) AND LEFT VENTRICULAR EJECTION FRACTION (LVEF) OF 10-15% WITH GLOBAL HYPOKINESIS AND A LARGE LV THROMBUS. CARDIAC MRI CONFIRMED LVEF OF 15% FROM MYOCARDIAL IRON DEPOSITION. RIGHT HEART CATHETERIZATION FINDINGS SHOWED RIGHT ATRIAL PRESSURE OF 10, PULMONARY ARTERY PRESSURE OF 30/20 (23), PULMONARY CAPILLARY WEDGE PRESSURE OF 19, AND CARDIAC INDEX OF 2.1. THEY HAD A NOTABLE FERRITIN OF 5133. DUE TO CARDIOGENIC SHOCK FROM HEMOCHROMATOSIS INDUCED CARDIOMYOPATHY, THEY WERE STARTED ON INOTROPES WITH DIURESIS. PATIENT DID NOT GET PHLEBOTOMY OR IRON CHELATION DUE TO HIS HEMODYNAMIC INSTABILITY. THEY RAPIDLY DETERIORATED REQUIRING INTRA-AORTIC BALLOON PUMP. THEY WERE EXCLUDED FROM HEART TRANSPLANTATION DUE TO SEVERE MALNUTRITION AND UNCONTROLLED DIABETES (A1C 10.2%). LIVER ECHOTEXTURE BY MRI SHOWED MODERATE INVOLVEMENT WITH INTACT SYNTHETIC FUNCTION. PATIENT HAD A HM3 LEFT VENTRICULAR ASSIST DEVICE SUPPORT WITH TEMPORARY RIGHT VENTRICULAR ASSIST DEVICE IN AN ATTEMPT TO BRIDGE TO POTENTIAL HEART-LIVER TRANSPLANTATION. UNFORTUNATELY, THEIR POSTOPERATIVE COURSE WAS COMPLICATED BY PROFOUND VASOPLEGIA LEADING TO REFRACTORY SHOCK DESPITE AGGRESSIVE THERAPIES. THE PATIENT SUCCUMBED TO THEIR PROFOUND CRITICAL ILLNESS. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBERS WERE DOCUMENTED AS UNKNOWN. EVENT DATE IS ESTIMATED TO BE APRIL 1, 2025 AS THIS POSTER WAS PUBLISHED ON THAT DATE. TIME OF DEVICE IMPLANTATION WAS UNKNOWN ASCENSION ST. VINCENT HOSPITAL, INDIANAPOLIS, IN, USA ISHOLA, F. O., ARMAN, M. E., LI, A., GARCIA-CORTES, R. S., SUDHARSHAN, S., CHAUDHRY, S.-P. S., RAVICHANDRAN, A. K., MORRIS, K., & WILLIAMS, A. J. (2025). THE DOWNWARD SPIRAL OF HEMOCHROMATOSIS. JOURNAL OF THE AMERICAN COLLEGE OF CARDIOLOGY, 85(12), 3899. HTTPS://DOI.ORG/10.1016/S0735-1097(25)04383-9 THE REPORTABLE AWARE DATE IS THE DATE THE SJM NOTIFIER COMPLETED READING THE ARTICLE AND ENTERED IN THE COMPLAINT DATABASE (PER SECTION 6.3.4 OF(B)(4)). NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION B1 CORRECTION. SECTION B2 CORRECTION. SECTION E2 CORRECTION. SECTION E3 CORRECTION. SECTION G2 CORRECTION. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT AND PATIENT INFORMATION; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22184889&pc=FTL | MW5171292 | 2019-10-23 04:00 | Death | COVIDIEN | 2025-06-11 04:00 | FTL | HERNIA MESH (POLYPROPYLENE) | Use of Device Problem; Therapeutic or Diagnostic Output Failure | Fistula; Failure of Implant; Neuropathy; Internal Organ Perforation; Pain; Urinary Tract Infection; Depression; Impaired Healing; Obstruction/Occlusion; Post Operative Wound Infection; Quadriplegia |  |  | 1 | Event Description: MY NAME IS (B)(6) THE SON OF (B)(6) WHO DIED AS A RESULT OF DEFECTIVE HERNIA MESH PRODUCTS, SPECIFICALLY BARD DOVAL AND COVIDIEN AS WELL AS A THIRD PRODUCT I BELIEVE. I WAS MY FATHER'S CARE PROVIDER FOR THE LAST THREE YEARS OF HIS LIFE AS HIS CONDITION BECAME SEVERE AND DEBILITATING. IN 2019 HE WENT IN FOR EMERGENCY SURGERY DUE TO A LIFE-THREATENING BOWEL BLOCKAGE THAT BECAME SEPTIC. IT TOOK 11 DAYS FOR HIM TO WAKE UP AND HE NEVER WALKED AGAIN FOR THE LAST TWO YEARS OF HIS LIFE. WHEN I MET WITH THE SURGEON DR. (B)(6) HE STATED "THIS WAS ONE OF THE MOST DIFFICULT SURGERIES HE EVER PERFORMED". BEAR IN MIND DR. (B)(6) HAD BEEN PERFORMING SURGERIES FOR 35 YEARS TO PUT THAT STATEMENT IN PERSPECTIVE. THE PRODUCTS USED IN MY FATHER'S HERNIA SURGERIES DESTROYED HIS INTERNAL ORGANS TO THE POINT THEY HAD TO BE REMOVED AND RESECTED. HIS ORGANS WERE DEAD, DYING, PERFORATED, AND SO ON. EVERYTHING I SPEAK ABOUT IN THIS EMAIL IS FULLY DOCUMENTED IN HIS MEDICAL RECORDS. I PERSONALLY SPENT 3 YEARS READING THROUGH EVERY PAGE AND GOING OVER THEM WITH A FINE TOOTH COMB. I SOLD MY HOME AND LEFT MY CAREER AS A MORTGAGE COMPLIANCE OFFICER TO LIVE WITH MY FATHER THE LAST 2-2.5 YEARS OF HIS LIFE TO TAKE CARE OF HIM 24/7 AS MY MOTHER, HIS WIFE PASSED AWAY IN 2011. AFTER HIS LAST HERNIA MESH REVISION SURGERY IN 2019 DUE TO A BOWEL BLOCKAGE HE NEVER RECOVERED. HE SUFFERED EXTREME NEUROPATHY, FUNCTIONAL QUADRIPLEGIA, LARGE ABDOMINAL WOUNDS WITH FISTULAS THAT I HAD TO PACK EVERY DAY, NUMEROUS UTI'S FROM HIS CATHETER, ETC... IN THE WEEKS FOLLOWING HIS SURGERY HIS SURGEON HAD TO PULL MESH THROUGH HIS FISTULAS DURING AN OFFICE VISIT DUE TO THE EXTREME PAIN HE WAS IN. I SPENT TWO YEARS LISTENING TO HIM CRYING OUT IN PAIN AND SOBBING ON A DAILY BASIS, AND UNLESS YOU'VE WATCHED A PARENT SUFFER LIKE THAT YOU HAVE NO IDEA WHAT IT DOES TO YOU. HIS WOUNDS NEVER HEALED DUE TO THE MESH INFECTION. I HAD TO CALL 911 OVER 20 TIMES IN THE LAST YEAR OF HIS LIFE AND HE HAD TO ENDURE THE PAIN OF TRAVELING TO THE HOSPITAL AND THE PAIN WAS EXCRUCIATING. HIS QUALITY OF LIFE WAS GONE COMPLETELY. MY FATHER DIED A VERY SLOW, AGONIZING, DEPRESSIVE, DEMORALIZING, TORTUOUS DEATH, TO PUT IT MILDLY. THESE ARE FACTS. TRUST ME, I HAD THE TRAUMATIZING TASK OF BEING THERE 24/7 THROUGHOUT THE ENTIRE 2 PLUS YEAR ORDEAL. MAKE NO MISTAKE HE SUFFERED FROM EVERY HERNIA MESH COMPLICATION YOU CAN FIND ON GOOGLE PLUS THE EXTREME NEUROPATHY, FUNCTION QUADRIPLEGIA, AND DEATH. HE COULD NOT WALK, FEED HIMSELF, BATHE, USE THE TOILET, SCRATCH HIS NOSE, PET HIS DOG, OR ANYTHING ELSE FOR THAT MATTER. NOBODY SHOULD EVER HAVE TO GO THROUGH THE HELL MY FATHER, MYSELF, AND MY FAMILY WENT THROUGH DUE TO THESE MEDICAL PRODUCTS. IT WAS A TRAGEDY OF EPIC PROPORTIONS FOR ALL OF US TO SAY THE VERY LEAST. REFERENCE REPORTS MW5171291, MW5171293.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22184899&pc=FTL | MW5171293 | 2019-10-23 04:00 | Death | MEDICAL CORPORATION | 2025-06-11 04:00 | FTL | HERNIA MESH (POLYPROPYLENE) | Contamination; Defective Device; Unintended Movement; Patient Device Interaction Problem | Wound Dehiscence; Cardiac Arrest; Fistula; Neuropathy; Internal Organ Perforation; Pain; Sepsis; Urinary Tract Infection; Distress; Impaired Healing; Coma; Quadriplegia |  |  | 1 | Event Description: MY NAME IS (B)(6) THE SON OF (B)(6) WHO DIED AS A RESULT OF DEFECTIVE HERNIA MESH PRODUCTS, SPECIFICALLY BARD DOVAL AND COVIDIEN AS WELL AS A THIRD PRODUCT I BELIEVE. I WAS MY FATHER'S CARE PROVIDER FOR THE LAST THREE YEARS OF HIS LIFE AS HIS CONDITION BECAME SEVERE AND DEBILITATING. IN 2019 HE WENT IN FOR EMERGENCY SURGERY DUE TO A LIFE-THREATENING BOWEL BLOCKAGE THAT BECAME SEPTIC. IT TOOK 11 DAYS FOR HIM TO WAKE UP AND HE NEVER WALKED AGAIN FOR THE LAST TWO YEARS OF HIS LIFE. WHEN I MET WITH THE SURGEON DR. (B)(6) HE STATED "THIS WAS ONE OF THE MOST DIFFICULT SURGERIES HE EVER PERFORMED". BEAR IN MIND DR. (B)(6) HAD BEEN PERFORMING SURGERIES FOR 35 YEARS TO PUT THAT STATEMENT IN PERSPECTIVE. THE PRODUCTS USED IN MY FATHER'S HERNIA SURGERIES DESTROYED HIS INTERNAL ORGANS TO THE POINT THEY HAD TO BE REMOVED AND RESECTED. HIS ORGANS WERE DEAD, DYING, PERFORATED, AND SO ON. EVERYTHING I SPEAK ABOUT IN THIS EMAIL IS FULLY DOCUMENTED IN HIS MEDICAL RECORDS. I PERSONALLY SPENT 3 YEARS READING THROUGH EVERY PAGE AND GOING OVER THEM WITH A FINE-TOOTH COMB. I SOLD MY HOME AND LEFT MY CAREER AS A MORTGAGE COMPLIANCE OFFICER TO LIVE WITH MY FATHER THE LAST 2-2.5 YEARS OF HIS LIFE TO TAKE CARE OF HIM 24/7 AS MY MOTHER, HIS WIFE PASSED AWAY IN 2011. AFTER HIS LAST HERNIA MESH REVISION SURGERY IN 2019 DUE TO A BOWEL BLOCKAGE HE NEVER RECOVERED. HE SUFFERED EXTREME NEUROPATHY, FUNCTIONAL QUADRIPLEGIA, LARGE ABDOMINAL WOUNDS WITH FISTULAS THAT I HAD TO PACK EVERY DAY, NUMEROUS UTI'S FROM HIS CATHETER, ETC... IN THE WEEKS FOLLOWING HIS SURGERY HIS SURGEON HAD TO PULL MESH THROUGH HIS FISTULAS DURING AN OFFICE VISIT DUE TO THE EXTREME PAIN HE WAS IN. I SPENT TWO YEARS LISTENING TO HIM CRYING OUT IN PAIN AND SOBBING ON A DAILY BASIS, AND UNLESS YOU'VE WATCHED A PARENT SUFFER LIKE THAT YOU HAVE NO IDEA WHAT IT DOES TO YOU. HIS WOUNDS NEVER HEALED DUE TO THE MESH INFECTION. I HAD TO CALL 911 OVER 20 TIMES IN THE LAST YEAR OF HIS LIFE AND HE HAD TO ENDURE THE PAIN OF TRAVELING TO THE HOSPITAL AND THE PAIN WAS EXCRUCIATING. HIS QUALITY OF LIFE WAS GONE COMPLETELY. MY FATHER DIED A VERY SLOW, AGONIZING, DEPRESSIVE, DEMORALIZING, TORTUOUS DEATH, TO PUT IT MILDLY. THESE ARE FACTS. TRUST ME, I HAD THE TRAUMATIZING TASK OF BEING THERE 24/7 THROUGHOUT THE ENTIRE 2 PLUS YEAR ORDEAL. MAKE NO MISTAKE HE SUFFERED FROM EVERY HERNIA MESH COMPLICATION YOU CAN FIND ON GOOGLE PLUS THE EXTREME NEUROPATHY, FUNCTION QUADRIPLEGIA, AND DEATH. HE COULD NOT WALK, FEED HIMSELF, BATHE, USE THE TOILET, SCRATCH HIS NOSE, PET HIS DOG, OR ANYTHING ELSE FOR THAT MATTER. NOBODY SHOULD EVER HAVE TO GO THROUGH THE HELL MY FATHER, MYSELF, AND MY FAMILY WENT THROUGH DUE TO THESE MEDICAL PRODUCTS. IT WAS A TRAGEDY OF EPIC PROPORTIONS FOR ALL OF US TO SAY THE VERY LEAST. REFERENCE REPORTS MW5171291, MW5171292.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22184984&pc=NVY | 2649622-2025-15812 | 2025-05-19 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-06-11 04:00 | NVY | SPRINT QUATTRO SECURE MRI SURESCAN | Under-Sensing | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED INTERMITTENT UNDERSENSING DURING A VENTRICULAR TACHYCARDIA (VT) EVENT. THE RV LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT APPROXIMATELY A WEEK LATER, THE PATIENT UNDERWENT A VENTRICULAR TACHYCARDIA (VT) ABLATION. APPROXIMATELY TWO WEEKS LATER, THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANTED: (B)(6) 2013. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22185041&pc=NGV | 2124215-2025-38228 | 2025-05-29 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-11 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED, AND A 27MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS IMPLANTED. THE PATIENT WAS DISCHARGED. A FEW DAYS OR WEEKS AFTER THE WATCHMAN PROCEDURE, THE PATIENT DIED. THE PHYSICIAN WAS UNSURE THE EXACT DATE OF DEATH OR CAUSE OF DEATH. THE PHYSICIAN DID NOT BELIEVE THE CAUSE OF DEATH WAS RELATED TO THE LEFT ATRIAL APPENDAGE CLOSURE (LAAC), BECAUSE THE PATIENT WAS IN NEED OF TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) AND HAD MANY OTHER COMORBIDITIES AND CONDITIONS. ACCORDING TO THE PHYSICIAN THE AUTOPSY IS NOT AVAILABLE. THERE WAS NO FURTHER INFORMATION PROVIDED. Manufacturer Narrative: B2: AWARE DATE WAS USED AS DATE OF DEATH AS ACTUAL DATE OF DEATH WAS NOT KNOWN. B3: AWARE DATE WAS USED AS EVENT DATE AS ACTUAL EVENT DATE WAS NOT KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22185412&pc=GDW | 3005075853-2025-04434 | 2025-05-16 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-06-11 04:00 | GDW | ETS FLEX ARTICNG LNR CUTR 45MM | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Cardiac Arrest; Hemorrhage/Bleeding | K020779 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN VIDEO LAPAROSCOPIC PROCEDURE ABOUT THE PATIENT'S DEATH, WHICH OCCURRED IN THE EARLY HOURS OF THE SAME DAY. TO DATE, THE CAUSE OF DEATH HAS NOT BEEN CONFIRMED. A FAST ULTRASOUND EXAM WAS PERFORMED, WHICH IDENTIFIED FREE FLUID IN THE SUPRAPUBIC REGION. IT WAS REPORTED THAT THE DEATH OCCURRED IN THE ICU. AFTER THE THIRD AND FINAL STAPLING STAGE TO COMPLETE THE NEW STOMACH, USING THE LOAD, IT WAS NECESSARY TO REINFORCE THE STAPLE LINE DUE TO OVERLAPPING TISSUE. IT WAS OBSERVED THAT THE STAPLE LINE APPEARED TO BE "FLAT". DURING THE METHYLENE BLUE TEST, A LEAK WAS IDENTIFIED IN THE ANASTOMOSIS REGION, SPECIFICALLY IN THE SUTURE PERFORMED BY THE DOCTOR, A NEW SUTURE WAS PERFORMED AND THE SURGERY WAS COMPLETED. ACCORDING TO INFORMATION, THE PATIENT PRESENTED AN INCIDENT IN THE ICU, AND DIED AFTER TWO RESPIRATORY ARRESTS, A FAST ULTRASOUND WAS PERFORMED, SHOWING FREE FLUID IN THE SUPRAPUBIC REGION. Manufacturer Narrative: (B)(4). DATE SENT: 6/11/2025. D4: BATCH # UNK. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: 1) WHAT WERE THE INDICATIONS FOR THE PROCEDURE? 1. THIS IS A PATIENT WHO WAS SEEN BY A MULTIDISCIPLINARY TEAM (SUS), WHERE THE PROCEDURE WAS AUTHORIZED DUE TO HER BMI. 2) WHAT IS THE SURGEON¿S EXPERIENCE WITH THE DEVICE? HE HAS PERFORMED MORE THAN 200 SURGERIES USING THE ATS. 3) DOES THE SURGEON TYPICALLY USE THE ATS FOR THIS PROCEDURE? IF HE USED THIS STAPLER FOR BARIATRIC SURGERIES, SUS, CONTRACT. 4) IF HE USED THIS STAPLER FOR BARIATRIC SURGERIES, SUS, CONTRACT. I DIDN'T UNDERSTAND THE QUESTION. FIRING? NO. 5) PLEASE PROVIDE MOR INFORMATION ON " STAPLE LINE APPEARED TO BE "FLAT". WHAT IS MEANT BY "IT WAS NECESSARY TO REINFORCE THE STAPLE LINE DUE TO OVERLAPPING TISSUE"? THE STAPLE LINE, ACCORDING TO THE SURGEON, HAD BLEEDING POINTS, AND THE STAPLE LINE NEEDED TO BE REINFORCED DUE TO THE BLEEDING. 6) PLEASE PROVIDE A TIME LINE OF EVENTS. WHAT MEDICAL OR SURGICAL INTERVENTION WAS ATTEMPTED PRIOR TO THE PATIENT'S DEATH? BEFORE THE INCIDENT, A PRODUCT COMPLAINT WAS FILED DUE TO (B)(6) CASE. AND THE SURGEON REQUESTED THAT THE ATS BE REPLACED WITH ECHELON. 7) WHAT WAS THE COD? WE WERE NOT NOTIFIED. 8) DOES THE SURGEON BELIEVE THE DIFFICULTIES WITH THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH? IF SO, WHY? THE SURGEON HAS NO DIFFICULTY WITH THE ATS. 9) IS THERE AN AUTOPSY REPORT AVAILABLE? NO. 10) IS THE SURGEON INTERESTED IN SPEAKING WITH ETHICON'S MEDICAL AND ENGINEERING TEAM? NOT AT THIS TIME. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22185615&pc=DSQ | 3007042319-2025-01257 | 2025-06-07 04:00 | Death | HEARTWARE, INC. | 2025-06-11 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Disconnection; Failure to Pump; Pumping Stopped; Power Problem | Fever; Nausea; Vomiting | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A VENTRICULAR ASSIST DEVICE (VAD) PATIENT ARRIVED TO THE EMERGENCY DEPARTMENT WITH FEVER AND ACUTE NAUSEA AND VOMITING FOR SEVERAL DAYS. THERE WAS A VAD STOP DUE TO A DOUBLE DISCONNECT AND THE CONTROLLERS HAD MULTIPLE LOSSES OF POWER. THE EMERGENCY DEPARTMENT WAS UNABLE TO RESTART THE PUMP DESPITE SEVERAL CONTROLLER EXCHANGES, INCLUDING THE USE OF CONTROLLERS WITH MODIFIED SOFTWARE FOLLOWING THE DOUBLE POWER DISCONNECT. THE PATIENT SUBSEQUENTLY DIED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR DEVICE EVALUATION AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: ONE (1) VENTRICULAR ASSIST DEVICE (VAD) ((B)(6) AND TWO (2) CONTROLLERS ((B)(6)) WERE NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. REVIEW OF THE AVAILABLE AUTOLOGS REPORT ASSOCIATED WITH (B)(6) REVEALED FIVE (5) VAD DISCONNECT ALARMS LOGGED SINCE (B)(6) 2025, INDICATING A PHYSICAL DISCONNECTION OF THE DRIVELINE FROM THE CONTROLLER. ADDITIONALLY, FIVE (5) VAD STOPPED ALARMS WERE LOGGED SINCE (B)(6) 2025 AND SEVEN (7) FAILED MOTOR START ATTEMPTS WERE LOGGED ON (B)(6) 2025 BETWEEN 20:25:43 AND 20:55:09, INDICATING THAT THE PUMP FAILED TO RESTART AFTER MULTIPLE ATTEMPTS. FURTHER REVIEW OF THE AVAILABLE AUTOLOGS REPORT REVEALED FIVE (5) CONTROLLER POWER-UP EVENTS WERE LOGGED SINCE (B)(6) 2025, INDICATING LOSSES OF POWER TO THE CONTROLLER. REVIEW OF THE AVAILABLE AUTOLOGS REPORT ASSOCIATED WITH (B)(6) REVEALED TWO (2) VAD DISCONNECT ALARMS LOGGED ON (B)(6) 2025 AT 20:43:38 AND 20:46:03, INDICATING A PHYSICAL DISCONNECTION OF THE DRIVELINE FROM THE CONTROLLER, AND TWO (2) VAD STOPPED ALARMS LOGGED SINCE (B)(6) 2025, INDICATING THAT THE PUMP FAILED TO RESTART AFTER MULTIPLE ATTEMPTS. ADDITIONALLY, TWO (2) CONTROLLER POWER-UP EVENTS WERE LOGGED ON (B)(6) 2025 AT 20:44:45 AND AT 20:47:04, INDICATING LOSSES OF POWER TO THE CONTROLLER. PRIOR TO THE SECOND LOSS OF POWER, A POWER ADAPTER WAS CONNECTED TO POWER PORT 1 AND (B)(6) WAS CONNECTED TO POWER PORT 2. THE CONTROLLER WAS WITHOUT POWER FOR FIFTY-NINE (59) SECONDS AND SEVENTEEN (17) SECONDS, RESPECTIVELY. OF NOTE, RAW LOG FILES ASSOCIATED WITH (B)(6) WERE NOT AVAILABLE FOR ANALYSIS. (B)(6) ARE LOADED WITH THE ALTERNATE SOFTWARE FOR THE PUMP START ALGORITHM. AS A RESULT, THE REPORTED EVENTS WERE CONFIRMED. THE MOST LIKELY ROOT CAUSE OF THE VAD STOPPED ALARM CAN BE ATTRIBUTED TO A FAILURE OF THE PUMP TO RESTART AFTER SEVERAL ATTEMPTS. (B)(6) WAS IN SCOPE OF FCA CVG-21-Q3-21 (NON-SUBGROUP). CAPA PR00532915 IS INVESTIGATING PUMP FAILURES TO RESTART. BASED ON AN INVESTIGATION CONDUCTED UNDER CAPA PR00532915, THE MOST LIKELY ROOT CAUSE OF THE FAILURE TO RESTART EVENT MAY BE ATTRIBUTED TO OUTER SHROUD CONTACT THAT CREATED MORE FRICTION AT THE HOUSING TO IMPELLER INTERFACE. THE MOST LIKELY ROOT CAUSE OF THE LOSS OF POWER EVENT CAN BE ATTRIBUTED TO A DISCONNECTION OF BOTH POWER SOURCES AND/OR TO AN INTERMITTENT DISCONNECTION ON ONE OR BOTH POWER SOURCES. CAPA PR00551638 INVESTIGATED CONTROLLER LOSSES OF POWER. EVEN THOUGH THIS CAPA IS NOW CLOSED, (B)(6) FALL IN SCOPE OF THIS CAPA. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. ADDITIONAL PRODUCTS: CONTROLLER 2.0 (B)(6) H3: YES CONTROLLER 2.0 (B)(6) H3: YES INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM CONTROLLER 2.0, D4: MODEL #: 1420 / CATALOG #: 1420 / EXPIRATION DATE: UNKNOWN / SERIAL OR LOT#: (B)(6), D9: NO, H3: NO, H4: MFG DATE: UNKNOWN, H5: NO, H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. D1: HEARTWARE VENTRICULAR ASSIST SYSTEM CONTROLLER 2.0 D4: MODEL #: 1420 / CATALOG #: 1420 / EXPIRATION DATE: UNKNOWN / SERIAL OR LOT#: (B)(6), D9: NO, H3: NO, H4: MFG DATE: UNKNOWN, H5: NO, H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22185858&pc=OZD | 1220648-2025-29213 | 2025-05-31 04:00 | Death | ABIOMED, INC. | 2025-06-11 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Septic Shock | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT THE PATIENT DEVELOPED SEPTIC SHOCK BUT THE TIMING OF ONSET AND TREATMENT ARE UNKNOWN. IN ADDITION, SINCE MANY CATHETERS WERE INSERTED IN ADDITION TO IMPELLA, SUCH AS EXTRACORPOREAL MEMBRANE OXYGENATION, SOLUBLE GUANYLATE CYCLASE, AND CONTROLLED VENTILATION, IT WAS UNCLEAR WHETHER THERE WAS A DIRECT CAUSAL RELATIONSHIP BETWEEN SEPSIS SHOCK AND IMPELLA. THE DOCTOR REPORTED THAT THERE WERE NO PROBLEMS WITH THE IMPELLA STERILIZED SLEEVE OR SKIN PROBLEMS AT THE IMPELLA INSERTION SITE. IT WAS ALSO REPORTED THAT MULTIPLE ORGAN FAILURE PROGRESSED DUE TO SEPSIS SHOCK AND THE PATIENT DIED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22186166&pc=NIQ | 2124215-2025-34652 | 2025-05-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-11 04:00 | NIQ | PROMUS PREMIER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Chest Pain |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH CORONARY ARTERY DISEASE AND WAS UNDERGOING PERCUTANEOUS CORONARY INTERVENTION. A 38 X 3.00 PROMUS PREMIER DRUG-ELUTING STENT WAS ADVANCED AND DEPLOYED IN THE LEFT CIRCUMFLEX ARTERY WITHOUT ANY ISSUES. AFTER THE PROCEDURE, THE PATIENT WAS MOVED TO A CARDIAC CARE UNIT. AFTER A FEW HOURS, THE PATIENT EXPERIENCED CHEST PAIN AND WENT INTO CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS IMMEDIATELY PERFORMED BUT THE PATIENT DIED. THE OFFICIAL CAUSE OF DEATH WAS CARDIAC ARREST, AND AN AUTOPSY WAS NOT PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22186172&pc=MCX | 2124215-2025-37743 | 2025-05-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-11 04:00 | MCX | ROTAPRO? | Entrapment of Device; Output Problem; Noise, Audible | Perforation | P900056 |  | 1 | Event Description: IT WAS REPORTED THAT DEVICE BECAME STUCK, AND PERFORATION OCCURRED. THE TARGET LESION WAS LOCATED IN THE CALCIFIED RIGHT CORONARY ARTERY. A 1.50 MM ROTAPRO WAS SELECTED FOR USE. DURING THE PROCEDURE, AFTER THE SECOND TO THIRD PASS, THE BURR BECAME STUCK. THE DEVICE WAS SUCCESSFULLY REMOVED. AFTER REMOVAL, THE PATIENT WAS PRESENTED WITH A MINIMAL PERFORATION VIA ANGIOGRAPHY AND PATIENT PRESSURE ALSO DROPPED. AN ECHOCARDIOGRAM WAS PERFORMED SHOWING NO INFUSION AT THE TIME. THE PHYSICIAN THEN PLACED A COVERED STENT TO COVER THE PERFORATION. THE PATIENT WAS ADMITTED FOR FURTHER OBSERVATION AND WAS EXPECTED TO FULLY RECOVER. THE PROCEDURE WAS COMPLETED USING ORIGINAL METHOD OR DEVICE.Event Description: IT WAS REPORTED THAT DEVICE BECAME STUCK, AND PERFORATION OCCURRED. THE TARGET LESION WAS LOCATED IN THE CALCIFIED RIGHT CORONARY ARTERY. A 1.50 MM ROTAPRO WAS SELECTED FOR USE. DURING THE PROCEDURE, AFTER THE SECOND TO THIRD PASS, THE BURR BECAME STUCK. THE DEVICE WAS SUCCESSFULLY REMOVED. AFTER REMOVAL, THE PATIENT WAS PRESENTED WITH A MINIMAL PERFORATION VIA ANGIOGRAPHY AND PATIENT PRESSURE ALSO DROPPED. AN ECHOCARDIOGRAM WAS PERFORMED SHOWING NO INFUSION AT THE TIME. THE PHYSICIAN THEN PLACED A COVERED STENT TO COVER THE PERFORATION. THE PATIENT WAS ADMITTED FOR FURTHER OBSERVATION AND WAS EXPECTED TO FULLY RECOVER. THE PROCEDURE WAS COMPLETED USING ORIGINAL METHOD OR DEVICE. IT WAS FURTHER REPORTED THAT THE TARGET LESION WAS MILDLY TORTUOUS AND SEVERELY CALCIFIED. FURTHERMORE, THE PATIENT WAS TRANSFERRED TO THE ICU FOR FURTHER CARE, HOWEVER THE PATIENT DIED.Event Description: IT WAS REPORTED THAT DEVICE BECAME STUCK, AND PERFORATION OCCURRED. THE TARGET LESION WAS LOCATED IN THE CALCIFIED RIGHT CORONARY ARTERY. A 1.50 MM ROTAPRO WAS SELECTED FOR USE. DURING THE PROCEDURE, AFTER THE SECOND TO THIRD PASS, THE BURR BECAME STUCK. THE DEVICE WAS SUCCESSFULLY REMOVED. AFTER REMOVAL, THE PATIENT WAS PRESENTED WITH A MINIMAL PERFORATION VIA ANGIOGRAPHY AND PATIENT PRESSURE ALSO DROPPED. AN ECHOCARDIOGRAM WAS PERFORMED SHOWING NO INFUSION AT THE TIME. THE PHYSICIAN THEN PLACED A COVERED STENT TO COVER THE PERFORATION. THE PATIENT WAS ADMITTED FOR FURTHER OBSERVATION AND WAS EXPECTED TO FULLY RECOVER. THE PROCEDURE WAS COMPLETED USING ORIGINAL METHOD OR DEVICE. IT WAS FURTHER REPORTED THAT THE TARGET LESION WAS MILDLY TORTUOUS AND SEVERELY CALCIFIED. FURTHERMORE, THE PATIENT WAS TRANSFERRED TO THE ICU FOR FURTHER CARE, HOWEVER THE PATIENT DIED. Manufacturer Narrative: H6 - DEVICE CODES: UPDATED. B2 - OUTCOMES ATTRIB TO ADV EVENT: UPDATED. D4 - LOT NUMBER: UPDATED. G1 - MANUFACTURING SITE: UPDATED. H1 - TYPE OF REPORTABLE EVENT: UPDATED. D4. LOT NUMBER: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. GOOD FAITH EFFORT ATTEMPT WAS MADE TO TRY AND RETRIEVE THE PRODUCT DETAILS. Manufacturer Narrative: B2 - OUTCOMES ATTRIB TO ADV EVENT: UPDATED. D4 - LOT NUMBER: UPDATED. G1 - MANUFACTURING SITE: UPDATED. H1 - TYPE OF REPORTABLE EVENT: UPDATED. D4. LOT NUMBER: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. GOOD FAITH EFFORT ATTEMPT WAS MADE TO TRY AND RETRIEVE THE PRODUCT DETAILS. Manufacturer Narrative: D4. LOT NUMBER: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. GOOD FAITH EFFORT ATTEMPT WAS MADE TO TRY AND RETRIEVE THE PRODUCT DETAILS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22186182&pc=GAM | 2210968-2025-06622 | 2024-02-01 05:00 | Death | ETHICON INC. | 2025-06-11 04:00 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Adverse Event Without Identified Device or Use Problem | Shock; Multiple Organ Failure | K022269 |  | 1 | Event Description: FLAP MANAGEMENT FOLLOWING ORTHOTOPIC HEART TRANSPLANTATION: A SINGLE INSTITUTION'S REVIEW OF 66 STERNAL WOUND COMPLICATIONS. THE AIM OF THIS STUDY IS TO FOCUSES ON THE INDICATIONS, OPERATIVE CHARACTERISTICS, AND OUTCOMES IN THE SUBSET OF PATIENTS WHO HAVE UNDERGONE HEART TRANSPLANTATION. BETWEEN (B)(6) 1996 AND (B)(6) 2023, A TOTAL OF 66 PATIENTS WHO UNDERWENT OHT (ORTHOTOPIC HEART TRANSPLANTATION) WHILE USING #2 VICRYL OR POLYSORB SUTURES. REPORTED COMPLICATIONS ARE N=4; PARTIAL DEHISCENCE TREATMENT: NOT MENTIONED. N=4; SUBPECTORAL COLLECTION TREATMENT: NOT MENTIONED. N=3; SEROMA TREATMENT: NOT MENTIONED. N=2; SKIN-EDGE NECROSIS TREATMENT: NOT MENTIONED. N=1; COMPLETE DEHISCENCE TREATMENT: NOT MENTIONED. N=1; PAIN TREATMENT: INCREASED NARCOTIC USE. N=10; UNSPECIFIED COMPLICATION TREATMENT: REVISIONARY SURGERY . N=4; RECURRENT INFECTION TREATMENT: N=?; SECOND STERNAL WOUND SURGERY. N=3; HEMATOMA TREATMENT: N=?; SECOND STERNAL WOUND SURGERY. N=?; WOUND DEHISCENCE TREATMENT: N=?; SECOND STERNAL WOUND SURGERY. N=2; DEATH DUE TO STERNAL WOUNDS INVOLVING MULTIORGAN FAILURE AFTER SHOCK TREATMENT: NOT MENTIONED. IN CONCLUSION, THIS REVIEW OF SWC IN HEART TRANSPLANT RECIPIENTS OVER 28 YEARS AT OUR INSTITUTION REVEALS NUMEROUS INSIGHTS INTO THE MANAGEMENT AND OUTCOMES OF THIS COMPLEX PATIENT POPU LATION. THERE IS A HEIGHTENED RISK OF STERNOTOMY WOUND COMPLICATIONS IN THIS GROUP, LIKELY DUE TO IMMUNOSUPPRES SION AND POOR BASELINE HEALTH STATUS. OUR STRATEGY FOCUSES ON IMMEDIATE AND AGGRESSIVE ANTIBIOTIC THERAPY, THOROUGH DEBRIDEMENT, AND DEFINITIVE CLOSURE WITH BILATERAL PECTOR ALIS MYOCUTANEOUS ADVANCEMENT FLAPS. THIS APPROACH HAS YIELDED COMPLICATION AND MORTALITY RATES THAT ARE COMPAR ABLE TO THE GENERAL CARDIAC SURGERY POPULATION, AS WELL AS EXCELLENT FUNCTIONAL AND ESTHETIC RESULTS. SINCE OUR 1995 SERIES, A MORE INDOLENT PRESENTATION OF STERNAL WOUND IN FECTIONS HAS BEEN NOTED. CONSEQUENTLY, WE ADVOCATE FOR EARLY INTERVENTION AT THE INITIAL CLINICAL SIGNS OF INFECTION, EVEN IN CASES WHERE WOUND CULTURES ARE NOT YET POSITIVE. A HIGH INDEX OF SUSPICION AND RAPID, DEFINITIVE TREATMENT IS IMPORTANT IN AVOIDING POSSIBLE LIFE-THREATENING CON SEQUENCES. FUTURE TREATMENT PROTOCOLS CAN MAKE USE OF THESE FINDINGS TO BETTER ASSESS PATIENT PROGNOSIS AND OPTI MIZE OPERATIVE MANAGEMENT. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DOES THE SURGEON BELIEVE THAT ANY OF THE ETHICON PRODUCTS INVOLVED CAUSED AND/OR CONTRIBUTED TO THE POST-OPERATIVE COMPLICATIONS/DEATH DESCRIBED IN THE ARTICLE? WHICH SPECIFIC ETHICON PRODUCTS HAVE BEEN USED DURING THE PROCEDURES (PRODUCT CODE, LOT NUMBER)? DOES THE SURGEON BELIEVE THERE WAS ANY DEFICIENCY WITH ANY OF THE ETHICON PRODUCTS USED IN THIS PROCEDURE? IF SO, PLEASE PROVIDE DETAILS. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE A COMPLAINT REFERENCE NUMBER. PATIENT DEMOGRAPHICS? HAS AN AUTOPSY BEEN PERFORMED FOR DEATH CASE? IF SO, CAN THE RESULTS BE SHARED? WOULD THE JOURNAL AUTHOR / PHYSICIAN LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. H6 COMPONENT CODE: G07002 ¿ DEVICE NOT RETURNED. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. CITATION: J PLAST RECONSTR AESTHET SURG. 2024 MAR;90:266-272. DOI: 10.1016/J.BJPS.2024.02.025. EPUB 2024 FEB 7. PMID: 38401198. HTTPS://DOI.ORG/10.1016/J.BJPS.2024.02.025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22186184&pc=GAM | 2210968-2025-06624 | 2024-02-01 05:00 | Death | ETHICON INC. | 2025-06-11 04:00 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Adverse Event Without Identified Device or Use Problem | Shock; Multiple Organ Failure | K022269 |  | 1 | Event Description: FLAP MANAGEMENT FOLLOWING ORTHOTOPIC HEART TRANSPLANTATION: A SINGLE INSTITUTION'S REVIEW OF 66 STERNAL WOUND COMPLICATIONS. THE AIM OF THIS STUDY IS TO FOCUSES ON THE INDICATIONS, OPERATIVE CHARACTERISTICS, AND OUTCOMES IN THE SUBSET OF PATIENTS WHO HAVE UNDERGONE HEART TRANSPLANTATION. BETWEEN MAY 1996 AND MAY 2023, A TOTAL OF 66 PATIENTS WHO UNDERWENT OHT (ORTHOTOPIC HEART TRANSPLANTATION) WHILE USING #2 VICRYL OR POLYSORB SUTURES. REPORTED COMPLICATIONS ARE N=4; PARTIAL DEHISCENCE TREATMENT: NOT MENTIONED N=4; SUBPECTORAL COLLECTION TREATMENT: NOT MENTIONED N=3; SEROMA TREATMENT: NOT MENTIONED N=2; SKIN-EDGE NECROSIS TREATMENT: NOT MENTIONED N=1; COMPLETE DEHISCENCE TREATMENT: NOT MENTIONED N=1; PAIN TREATMENT: INCREASED NARCOTIC USE N=10; UNSPECIFIED COMPLICATION TREATMENT: REVISIONARY SURGERY N=4; RECURRENT INFECTION TREATMENT: N=?; SECOND STERNAL WOUND SURGERY N=3; HEMATOMA TREATMENT: N=?; SECOND STERNAL WOUND SURGERY N=?; WOUND DEHISCENCE TREATMENT: N=?; SECOND STERNAL WOUND SURGERY N=2; DEATH DUE TO STERNAL WOUNDS INVOLVING MULTIORGAN FAILURE AFTER SHOCK TREATMENT: NOT MENTIONED IN CONCLUSION, THIS REVIEW OF SWC IN HEART TRANSPLANT RECIPIENTS OVER 28 YEARS AT OUR INSTITUTION REVEALS NUMEROUS INSIGHTS INTO THE MANAGEMENT AND OUTCOMES OF THIS COMPLEX PATIENT POPULATION. THERE IS A HEIGHTENED RISK OF STERNOTOMY WOUND COMPLICATIONS IN THIS GROUP, LIKELY DUE TO IMMUNOSUPPRES SION AND POOR BASELINE HEALTH STATUS. OUR STRATEGY FOCUSES ON IMMEDIATE AND AGGRESSIVE ANTIBIOTIC THERAPY, THOROUGH DEBRIDEMENT, AND DEFINITIVE CLOSURE WITH BILATERAL PECTOR ALIS MYOCUTANEOUS ADVANCEMENT FLAPS. THIS APPROACH HAS YIELDED COMPLICATION AND MORTALITY RATES THAT ARE COMPAR ABLE TO THE GENERAL CARDIAC SURGERY POPULATION, AS WELL AS EXCELLENT FUNCTIONAL AND ESTHETIC RESULTS. SINCE OUR 1995 SERIES, A MORE INDOLENT PRESENTATION OF STERNAL WOUND IN FECTIONS HAS BEEN NOTED. CONSEQUENTLY, WE ADVOCATE FOR EARLY INTERVENTION AT THE INITIAL CLINICAL SIGNS OF INFECTION, EVEN IN CASES WHERE WOUND CULTURES ARE NOT YET POSITIVE. A HIGH INDEX OF SUSPICION AND RAPID, DEFINITIVE TREATMENT IS IMPORTANT IN AVOIDING POSSIBLE LIFE-THREATENING CON SEQUENCES. FUTURE TREATMENT PROTOCOLS CAN MAKE USE OF THESE FINDINGS TO BETTER ASSESS PATIENT PROGNOSIS AND OPTI MIZE OPERATIVE MANAGEMENT. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DOES THE SURGEON BELIEVE THAT ANY OF THE ETHICON PRODUCTS INVOLVED CAUSED AND/OR CONTRIBUTED TO THE POST-OPERATIVE COMPLICATIONS/DEATH DESCRIBED IN THE ARTICLE? WHICH SPECIFIC ETHICON PRODUCTS HAVE BEEN USED DURING THE PROCEDURES (PRODUCT CODE, LOT NUMBER)? DOES THE SURGEON BELIEVE THERE WAS ANY DEFICIENCY WITH ANY OF THE ETHICON PRODUCTS USED IN THIS PROCEDURE? IF SO, PLEASE PROVIDE DETAILS. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE A COMPLAINT REFERENCE NUMBER. PATIENT DEMOGRAPHICS? HAS AN AUTOPSY BEEN PERFORMED FOR DEATH CASE? IF SO, CAN THE RESULTS BE SHARED? WOULD THE JOURNAL AUTHOR / PHYSICIAN LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. H6 COMPONENT CODE: G07002 ¿ DEVICE NOT RETURNED D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. CITATION: J PLAST RECONSTR AESTHET SURG. 2024 MAR;90:266-272. DOI: 10.1016/J.BJPS.2024.02.025. EPUB 2024 FEB 7. PMID: 38401198. HTTPS://DOI.ORG/10.1016/J.BJPS.2024.02.025
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22186222&pc=NGV | 2135147-2025-03093 | 2025-05-17 04:00 | Death | ABBOTT MEDICAL | 2025-06-11 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Angina; Cardiac Arrest; Non specific EKG/ECG Changes; Low Blood Pressure/ Hypotension; Cardiac Tamponade; Pericardial Effusion; Movement Disorder | P200049 |  | 1 | Event Description: IT WAS REPORTED ON (B)(6) 2025 AN 18MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS SELECTED FOR IMPLANT USING A 14F AMPLATZER TORQVUE 45X45 DELIVERY SHEATH. THE PATIENT HAD A PRE-EXISTING TRACE EFFUSION AT THE START OF THE PROCEDURE. TRANSSEPTAL PUNCTURE WAS INFERIOR AND MID. THE PATIENT HAD A HYPERCONTRACTILE APPENDAGE. THE PATIENT'S LEFT ATRIAL PRESSURE WAS 13 MMHG. CLOSE CRITERIA WERE MET AND THE OCCLUDER WAS IMPLANTED. THE TRIVIAL EFFUSION WAS UNCHANGED. ON (B)(6) 2025 THE PATIENT RETURNED TO THE HOSPITAL WITH DIZZINESS AND CHEST PAIN AND BECAME HYPOTENSIVE. A SURFACE ECHOCARDIOGRAM REVEALED A MODERATE, CIRCUMFERENTIAL PERICARDIAL EFFUSION WITH A RIGHT VENTRICLE CHAMBER COLLAPSE. THE EFFUSION DEVELOPED INTO A CARDIAC TAMPONADE. THE PATIENT DECOMPENSATED WITH ST CHANGES AND BECAME HEMODYNAMICALLY UNSTABLE. THE PATIENT WENT INTO CARDIAC ARREST AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. A PERICARDIOCENTESIS WAS PERFORMED WHILE THE PATIENT WAS CODING AND 150CC OF LIQUID WAS REMOVED. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE CAUSE OF DEATH WAS THE PERICARDIAL EFFUSION. THE USER BELIEVED THE PERICARDIAL EFFUSION WAS DUE TO THE OCCLUDER. Manufacturer Narrative: AN EVENT OF DIZZINESS, CHEST PAIN, HYPOTENSION, CIRCUMFERENTIAL PERICARDIAL EFFUSION AND CARDIAC TAMPONADE WERE REPORTED TWO DAYS POST AMULET DEVICE IMPLANT. WHILE IN THE HOSPITAL THE PATIENT DECOMPENSATED FURTHER AND WENT INTO CARDIAC ARREST REQUIRING CARDIOPULMONARY RESUSCITATION (CPR). A PERICARDIOCENTESIS WAS PERFORMED WHILE THE PATIENT WAS CODING. HOWEVER, THE PATIENT EXPIRED AND THE CAUSE OF DEATH WAS REPORTED AS PERICARDIAL EFFUSION. THE DEVICE REMAINED IMPLANTED AND WILL NOT RETURN TO ABBOTT FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE CINE REVIEW PERFORMED AT ABBOTT, THE FINAL AMULET DEVICE POSITION APPEARED TO BE DEEP INTO THE LAA. A DEEP AMULET IMPLANT INTO THE LAA MAY HAVE INCREASED THE RISK FOR DEVELOPING A PERICARDIAL EFFUSION. BASED ON MEDICAL REVIEW, THE PATIENT DEATH APPEARS TO BE RELATED TO POST-PROCEDURE BLEEDING INTO PERICARDIAL SPACE WITH DEVELOPMENT OF CARDIAC TAMPONADE. THE CAUSE FOR THE BLEEDING COULD NOT BE DETERMINED. IT IS POSSIBLE THAT BLEEDING MAY BE RESULT OF MICRO-PERFORATION OF THE LAA AND/OR INJURY TO THE PULMONARY ARTERY. HOWEVER, THIS COULD NOT BE CONFIRMED. A PERICARDIAL EFFUSION IS A KNOWN POTENTIAL ADVERSE EVENT FOLLOWING LAAO. THE REPORTED PATIENT EFFECTS OF ANGINA, MOVEMENT DISORDER, PERICARDIAL EFFUSION, CARDIAC TAMPONADE, HYPOTENSION, CARDIAC ARREST AND DEATH APPEARS TO BE A CASCADING EFFECT. THE REPORTED HOSPITALIZATION AND UNEXPECTED MEDICAL INTERVENTION WERE RESULT OF CASE-SPECIFIC CIRCUMSTANCE AS PERICARDIOCENTESIS WAS PERFORMED FOR PERFORATION AND RESUSCITATION WAS REQUIRED DUE TO CARDIAC ARREST. THERE WERE NO COMPLAINTS ASSOCIATED WITH ANY OTHER DEVICES FROM THE LOT. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22186489&pc=DYB | 3007284313-2025-04027 | 2025-01-07 05:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-11 04:00 | DYB | GORE® DRYSEAL FLEX INTRODUCER SHEATH | Insufficient Information | Perforation of Vessels | K160254 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, A PATIENT WAS BEING TREATED USING GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL (CTAG) USING AN INTRODUCER SHEATH THAT IS LIKELY A GORE® DRYSEAL FLEX INTRODUCER SHEATH. AFTER THE DEVICE WAS FULLY DEPLOYED, THE PHYSICIAN ATTEMPTED TO REMOVED THE GREY ANGULATION HANDLE, BUT WAS UNABLE TO DO SO. TWO STRINGS WERE STILL ATTACHED, AND WHEN TRYING TO PULL THE HANDLE THE DEVICE BEGAN ANGULATING ON THE INNER CURVE. THE PHYSICIAN THEN OPENED THE HATCH. LINES 3 AND 4 WERE STILL AVAILABLE IN THE HATCH. THE PHYSICIAN ATTEMPTED TO PULL LINE 4, BUT THE DEVICE CONTINUED TO ANGULATE. WHEN THE PHYSICIAN TRIED TO REMOVE LINE 3, THERE WAS RESISTANCE, AND THE PHYSICIAN WAS UNABLE TO REMOVE LINE 3. THE PHYSICIAN THEN BALLOONED THE STENT GRAFT. THIS RELEASED THE DEVICE FROM THE STENT GRAFT. IMMEDIATELY FOLLOWING THE BALLOONING, THE BLOOD PRESSURE IN THE PATIENT DROPPED. THE PHYSICIAN PERFORMED A RETROGRADE ANGIOGRAM, AND DISCOVERED THAT THE RIGHT ILIAC ARTERY WAS TRANSECTED, AND COMPLETED DISCONNECTED. THE PHYSICIAN CONVERTED TO OPEN SURGERY TO REPAIR THE ILIAC, BUT THE PATIENT DIED. ACCORDING TO THE VOLUNTARY MEDWATCH SUBMITTED BY (B)(6) ON (B)(6) 2025, THE ILIAC ARTERY WAS TRANSECTED BY THE GRAFT. AT THE TIME OF THE EVENT, BOTH THE FSA AND A PRODUCT SPECIALIST REPORTED THAT THE SURGEON HAD TOLD THEM THE RUPTURED ILIAC ARTERY WAS UNRELATED TO THE IMPLANTED DEVICE, AND IN A FOLLOW-UP CONVERSATION ON (B)(6) 2025, THE SURGEON REITERATED THAT HE DID NOT BELIEVE THE STENT WAS THE CAUSE OF THE RUPTURE. ACCORDING TO (B)(6), THEIR OFFICE DID NOT DIRECTLY EXAMINE THE PATIENT BODY. Manufacturer Narrative: H.6. TYPE OF INVESTIGATION CODE B17: THE DEVICE WAS REQUESTED FROM THE FACILITY, BUT WAS NOT MADE AVAILABLE. H.6. TYPE OF INVESTIGATION CODE B22: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE CONDUCTED BECAUSE THE SERIAL/LOT NUMBER REMAINS UNKNOWN. H.6. INVESTIGATION CONCLUSIONS CODE D15: THERE IS INSUFFICIENT INFORMATION AVAILABLE FOR GORE TO REASONABLY DRAW CONCLUSIONS RELATED TO ASPECTS OF THE EVENT, THEREFORE CONCLUSION CODE ¿D15: CAUSE NOT ESTABLISHED¿ IS BEING USED. INSUFFICIENT INFORMATION MAY INCLUDE LIMITED OR MISSING RELEVANT MEDICAL RECORDS, INVOLVEMENT OF MULTIPLE IMPLANTED DEVICES (INCLUDING NON-GORE DEVICES) IN THE FIELD OF TREATMENT, PATIENT NON-COMPLIANCE, AND/OR A GENERAL LACK OF AVAILABLE DETAIL OR SPECIFICITY RELATED TO AN ADVERSE EVENT AND/OR DEVICE. Manufacturer Narrative: THE PRODUCT MALFUNCTION REFERENCED IN THE EVENT DESCRIPTION WAS PREVIOUSLY REPORTED ON REPORT 3013164176-2025-02370. AT THE TIME OF THAT REPORT, THE INFORMATION FROM THE PHYSICIAN INDICATED THAT THE GORE® DEVICES PLAYED NO ROLE IN THE RUPTURE OF THE ILIAC THAT LED TO THE PATIENT'S DEATH. ON MAY 15, 2025, WE BECAME AWARE OF VOLUNTARY MEDWATCH MW5170196, ALLEGING THAT GORE® TAG® CONFORMABLE THORACIC STENT GRAFT CAUSED THE TRANSECTED ILIAC ARTERY. BASED ON FURTHER CONVERSATIONS WITH THE PHYSICIAN, THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT TRAVERSED THAT SECTION OF THE VASCULATURE WITHIN A GORE® DRYSEAL FLEX INTRODUCER SHEATH. THERE IS NO REASON TO SUSPECT THAT THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT CAUSED OR CONTRIBUTED TO THE RUPTURE. ALTHOUGH THERE IS NO DIRECT ALLEGATION AGAINST THE GORE® DRYSEAL FLEX INTRODUCER SHEATH, WE ARE REPORTING ON THE DEVICE OUT OF AN ABUNDANCE OF CAUTION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22186522&pc=PYX | 1220648-2025-29217 | 2025-04-02 04:00 | Death | ABIOMED, INC. | 2025-06-11 04:00 | PYX | IMPELLA RP | Failure to Pump | No Clinical Signs, Symptoms or Conditions | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE STAFF WAS UNABLE TO ADEQUATELY CALIBRATE / ZERO THE SIGNAL OF AN IMPLANTED IMPELLA RP PUMP. ROUGHLY SIX DAYS INTO SUPPORT, POSITIONING ISSUES WERE ENCOUNTERED FOLLOWING SWAN GANZ CATHETER REMOVAL AND THE PUMP STOPPED UNEXPECTEDLY MULTIPLE TIMES. THE PUMP WAS REPLACED IN RESPONSE TO THE NOTED FAILURE. FURTHER DETAILS STATE THAT THE PATIENT WAS UNSTABLE AND CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE PUMP WAS RETURNED FOR INVESTIGATION. NO PHYSICAL DAMAGE WAS OBSERVED ON THE RETURNED PRODUCT. THE PUMP PASSED ALL ELECTRICAL RESISTANCE TESTS ON THE MOTOR CABLE AND THE DP SENSOR. THE PUMP WAS FURTHER CONNECTED TO THE AIC, AND THE IMPELLA STOPPED; A CONTROLLER FAILURE ALARM WAS OBSERVED DURING THE START OF THE PUMP. A SIGNIFICANT AMOUNT OF BIOMATERIAL WAS OBSERVED ON THE IMPELLER. HIGH PURGE PRESSURE WAS REPRODUCED DURING THE TEST. FURTHER TEARDOWN OF THE INFLOW CAGE WAS PERFORMED TO CLEAR THE BIOMATERIAL. AFTER CLEARING THE BIOMATERIAL, THE PUMP WAS ABLE TO RUN IN A BUCKET OF WATER WITH NO PURGE-RELATED ISSUES AND ALSO CONFIRMED THAT THE PURGE GAP APPEARED NORMAL ON THE PUMP. PUMP STOP: THE DATA LOG SHOWS SEVERAL PUMP STOP ALARMS WITH MOTOR CURRENT SPIKES AT THE END OF THE CASE RUN ON DAY 7. PURGE PRESSURE WAS SLIGHTLY ELEVATED DURING THE PUMP STOP EVENT. THE CAUSE OF THE PUMP STOP ISSUE WAS BIOMATERIAL INGESTION DURING THE CASE RUN, BASED ON RETURNED PRODUCT INVESTIGATION AND DATA LOG REVIEW. PLACEMENT SIGNAL ISSUE: THE LT LOG SHOWS NOTIFICATIONS OF COMPLETING THE ZEROING SENSOR EVENT. ADDITIONALLY, THE IMC LOG WAS REVIEWED AND NOTED THAT THE ZEROING OF THE DP SENSOR WAS PER SPECIFICATION AND THAT THE CORRECT OFFSET WAS MAINTAINED UNTIL THE END OF THE CASE RUN. THE IMC LOG CONFIRMS THAT THE DOCTOR PRESSED THE P-0 PUMP START INPUT THREE TIMES AFTER ZEROING THE SENSOR. THE CAUSE OF THE PLACEMENT SIGNAL ISSUE WAS NOT DETERMINED, AS THE DATA LOG SUGGESTS NO ISSUES IN CALCULATING THE OFFSET DURING THE ZEROING PROCESS, AND SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED TO CONFIRM ANY USER-RELATED ASPECTS. THE CAUSE OF THE POSITIONING ISSUE WAS NOT DETERMINED SINCE SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED. THE CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS NOT DETERMINED SINCE SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN RECEIVED BUT HAS NOT YET BEEN EVALUATED. WHEN THE DEVICE INVESTIGATION HAS BEEN COMPLETED OR ANY NEW INFORMATION HAS BEEN RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: TROUBLESHOOTING THE PURGE SYSTEM ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA RP WITH SMARTASSIST SYSTEM CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22186779&pc=OZD | 1220648-2025-29220 | 2025-05-21 04:00 | Death | ABIOMED, INC. | 2025-06-11 04:00 | OZD | IMPELLA 5.5 | Failure to Advance | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD AN IMPELLA 5.5 PLACED TO SUPPORT AND REPLACE A PRIOR IMPELLA 5.5 THAT MALFUNCTIONED ON THE FIRST CONSOLE. A SECOND PUMP FAILED TO DELIVER DESPITE ALL EFFORTS AND WAS REMOVED. ADDITIONALLY, BLEEDING WAS OBSERVED DURING THE PROCEDURE, AND THE PATIENT RECEIVED 20 UNITS OF FRESH FROZEN PLASMA, 19 UNITS OF RED BLOOD CELLS, 2 UNITS OF PLATELETS, AND 3 UNITS OF CRYOPRECIPITATE. THE PATIENT REMAINED NON-PULSATILE WITH COMPROMISED HEMODYNAMICS AND CONTINUED TO REQUIRE ONGOING TRANSFUSIONS OF BLOOD PRODUCTS. ULTIMATELY, THE DECISION WAS MADE BY THE FAMILY TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA 5.5 WITH SMARTASSIST SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22187076&pc=NPT | 2025587-2025-04184 | 2025-06-04 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-11 04:00 | NPT | EVOLUT FX DCS | Difficult to Insert; Failure to Advance; Material Protrusion/Extrusion; Patient Device Interaction Problem | Thrombosis/Thrombus | P130021 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT HAD STENOSIS THROUGHOUT THE FEMORAL AND ILIAC STENTS WHICH HAD BEEN TREATED WITH COATED STENTS. WHILE ADVANCING THE DELIVERY CATHETER SYSTEM (DCS), THIS CAUSED BOTH THE STENTS TO MIGRATE OVER THE DCS AND GUIDEWIRE. THE AORTIC STENT DETACHED AND BECAME STUCK IN THE ILIAC ARTERY STENT. THE DCS BECAME DAMAGED AND THE SYSTEM NEEDED TO BE REPLACED. THROMBOSIS AT THE LEVEL OF THE ABDOMINAL AORTA WAS REPORTED AS THE STENT BECAME COMPACTED. THE PATIENT UNDERWENT SURGERY.Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT HAD STENOSIS THROUGHOUT THE FEMORAL AND ILIAC STENTS WHICH HAD BEEN TREATED WITH COATED STENTS. WHILE ADVANCING THE DELIVERY CATHETER SYSTEM (DCS), THIS CAUSED BOTH THE STENTS TO MIGRATE OVER THE DCS AND GUIDEWIRE. THE AORTIC STENT DETACHED AND BECAME STUCK IN THE ILIAC ARTERY STENT. THE DCS BECAME DAMAGED AND THE SYSTEM NEEDED TO BE REPLACED. THROMBOSIS AT THE LEVEL OF THE ABDOMINAL AORTA WAS REPORTED AS THE STENT BECAME COMPACTED. THE PATIENT UNDERWENT SURGERY. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PROTRUSION WAS NOTED AFTER INSERTION INTO THE PATIENT. THE MINIMUM DIAMETER OF THE ACCESS VESSEL WAS 4.2/3.5 MM BUT WAS PREVIOUSLY TREATED WITH A STENT. A 14 FRENCH MEDTRONIC (SENTRANT) EXTERNAL INTRODUCER SHEATH WAS USED FOR THE PROCEDURE. THERE WAS DIFFICULTY INSERTING THE DCS AND IT WOULD NOT PASS THE ABDOMINAL AORTA. A PROCEDURAL DELAY OCCURRED AS A RESULT OF THE PROTRUSION. SURGERY WAS PERFORMED TO REMOVE THE THROMBOSED AORTIC STENT. A NEW VALVE WAS IMPLANTED. THE PATIENT DIED AFTER THE THIRD SURGERY.Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT HAD STENOSIS THROUGHOUT THE FEMORAL AND ILIAC STENTS WHICH HAD BEEN TREATED WITH COATED STENTS. WHILE ADVANCING THE DELIVERY CATHETER SYSTEM (DCS), THIS CAUSED BOTH THE STENTS TO MIGRATE OVER THE DCS AND GUIDEWIRE. THE AORTIC STENT DETACHED AND BECAME STUCK IN THE ILIAC ARTERY STENT. THE DCS BECAME DAMAGED AND THE SYSTEM NEEDED TO BE REPLACED. THROMBOSIS AT THE LEVEL OF THE ABDOMINAL AORTA WAS REPORTED AS THE STENT BECAME COMPACTED. THE PATIENT UNDERWENT SURGERY. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PROTRUSION WAS NOTED AFTER INSERTION INTO THE PATIENT. THE MINIMUM DIAMETER OF THE ACCESS VESSEL WAS 4.2/3.5 MM BUT WAS PREVIOUSLY TREATED WITH A STENT. A 14 FRENCH MEDTRONIC (SENTRANT) EXTERNAL INTRODUCER SHEATH WAS USED FOR THE PROCEDURE. THERE WAS DIFFICULTY INSERTING THE DCS AND IT WOULD NOT PASS THE ABDOMINAL AORTA. A PROCEDURAL DELAY OCCURRED AS A RESULT OF THE PROTRUSION. SURGERY WAS PERFORMED TO REMOVE THE THROMBOSED AORTIC STENT. A NEW VALVE WAS IMPLANTED. THE PATIENT DIED AFTER THE THIRD SURGERY. ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEATH OCCURRED ONE DAY FOLLOWING THE VALVE IMPLANT PROCEDURE. THE CAUSE OF DEATH WAS NOT KNOWN, BUT THE PHYSICIAN ATTRIBUTED THE DEATH TO THE FIRST DCS BUT NOT THE IMPLANTED VALVE OR SECOND DCS.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE VALVE IMPLANT, THREE SURGERIES WERE PERFORMED. THE FIRST TO BRIDGE THE FEMORAL ARTERIES, THE SECOND FOR THE ABDOMINAL THROMBOSIS AND THE THIRD FAILED WHEN THEY TRIED TO REMOVE THE BLOOD CLOTS. NON-MEDTRONIC STENT (BEGRAFT) WAS USED DURING THE PROCEDURE. DURING THE VALVE ADVANCEMENT, THE PROTRUSION OCCURRED WITH THE DELIVERY CATHETER SYSTEM (DCS) CAPSULE AS IT INTERACTED WITH THE ILIAC STENTS AND SUBSEQUENT AORTIC STENT. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: EVOLUTFX-26 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: ON (B)(6) 2025; EXPLANT DATE: ON (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B2, B5, D4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW: THREE STATIC IMAGES WERE PROVIDED FOR REVIEW. THE PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW THEREFORE IT IS NOT POSSIBLE TO CONFIRM MINIMUM VESSEL DIAMETERS. IT WAS REPORTED THAT THE PATIENT¿S FEMORAL AND ILIAC ARTERIES HAD STENOSIS AND STENTS WERE IMPLANTED. IMAGES SHOW THAT THE DELIVERY CATHETER SYSTEM (DCS) INTERACTED WITH THE PERIPHERAL STENTS CAUSING DISLODGEMENT OF THE STENTS AND DAMAGE OF THE DCS. IT WAS REPORTED THAT THE PROCEDURE WAS ABORTED, AND THE PATIENT UNDERWENT SURGERY. UPDATED: B2, B5, H1, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22187354&pc=OZD | 1220648-2025-29226 | 2025-05-19 04:00 | Death | ABIOMED, INC. | 2025-06-11 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT ENDURED DIFFUSE BLEEDING FROM THEIR AXILLARY POCKET IMMEDIATELY FOLLOWING THE IMPELLA IMPLANT. ANESTHESIA WAS UNABLE TO OXYGENATE THE PATIENT WITH FRANK RED BLOOD OUT OF THE ET TUBE AND THE DIFFUSE BLEEDING. WHILE ATTEMPTING TO CANNULATE, THE PATIENT CODED AND PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMART ASSIST SYSTEM INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22187582&pc=OZD | 1220648-2025-29229 | 2025-05-23 04:00 | Death | ABIOMED, INC. | 2025-06-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE STAFF WAS UNABLE TO OBTAIN PULSES AFTER THE REMOVAL OF THE PEEL AWAY SHEATH. THE PATIENT CODED AND THE PUMP WAS TURNED TO P2. HOWEVER, THE PATIENT LATER EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22187832&pc=OZD | 1220648-2025-29230 | 2025-05-19 04:00 | Death | ABIOMED, INC. | 2025-06-11 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT DEVELOPED HEMOLYSIS. DUE TO THE PATIENT'S WORSENING UNDERLYING CONDITION, THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA CP WITH SMARTASSIST SYSTEM INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22188051&pc=NVZ | 1028232-2025-02821 | 2025-05-05 04:00 | Death | BIOTRONIK SE & CO. KG | 2025-06-11 04:00 | NVZ | ENITRA 8 DR-T | Adverse Event Without Identified Device or Use Problem | Syncope/Fainting | P950037 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE STAY IN HOSPITAL THE PATIENT SUFFERED SYNCOPE AND WAS RESUSCITATED, BUT WITHOUT SUCCESS. THE PATIENT DIED. HOSPITAL STAY WAS NOT FOR CARDIOLOGICAL REASONS. Manufacturer Narrative: THE PACEMAKER AND THE PROXIMAL FRAGMENTS OF THE LEADS WERE RETURNED AND SUBJECTED TO AN EXTENSIVE ANALYSIS. PRIOR TO THE ANALYSIS OF THE DEVICES, THE QUALITY DOCUMENTS ACCOMPANYING THE MANUFACTURING PROCESS FOR THESE DEVICES WERE RE-INVESTIGATED. ALL PRODUCTION STEPS WERE PERFORMED ACCORDINGLY, AND IN PARTICULAR THE FINAL ACCEPTANCE TESTS PROVED THE DEVICES FUNCTIONS TO BE AS SPECIFIED. ENITRA 8 DR-T: UPON RECEIPT, THE PACEMAKER WAS INTERROGATED AND THE MEMORY CONTENT WAS ANALYZED INDICATING NO ANOMALIES. THE BATTERY STATUS WAS FOUND TO BE LARGER THAN 85%. THE ABILITY OF THE DEVICE TO DELIVER THERAPIES WAS VERIFIED. THE ANTI-BRADYCARDIA PACING PULSES PROVED TO BE NORMAL AND IN AMPLITUDE AND FREQUENCY AS PROGRAMMED. THERE WAS NO INDICATION OF A DEVICE MALFUNCTION. AN ADDITIONAL LONG-TERM PACING TEST WAS INITIATED. DURING THE TEST, EACH PACING PULSE WAS RECORDED. THE EVALUATION OF THESE PACING PULSES DOCUMENTED REGULAR DEVICE BEHAVIOR. THERE WAS NO INTERMITTENT OR PERMANENT LOSS OF OUTPUT. IN CONCLUSION, THE DEVICE IS FULLY FUNCTIONAL. THE ANALYSIS DID NOT SHOW ANY DEVIATIONS FROM THE TECHNICAL SPECIFICATIONS. THE ANALYSIS DID NOT REVEAL ANY SIGN OF A MATERIAL OR MANUFACTURING PROBLEM. SOLIA S 53 AND S 60: UPON RECEIPT, BOTH LEADS WERE FOUND CUT APPROX. 10 CM DISTAL TO THE IS-1 CONNECTOR PIN. ONLY THE PROXIMAL FRAGMENTS OF THE LEADS WERE AVAILABLE FOR ANALYSIS. IN THE COURSE OF THE VISUAL INSPECTION, NO DEVIATIONS WERE NOTED, WHICH CAN BE CONSIDERED TO BE THE ROOT CAUSE OF THE CLINICAL OBSERVATION. THE LEAD FRAGMENTS PROVED TO BE WITHOUT FAULT THROUGHOUT ITS INSPECTION. IN CONCLUSION, THE ANALYSIS DID NOT REVEAL ANY SIGN OF A MATERIAL OR MANUFACTURING PROBLEM OF THE AVAILABLE LEAD FRAGMENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22188352&pc=FKX | 0008030665-2025-01323 | 2025-05-20 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-06-11 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K222318 |  | 1 | Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED TO FRESENIUS CUSTOMER SUPPORT THAT A PATIENT PASSED AWAY ON THE CYCLER. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING HOURS OF (B)(6) 2025. THE PATIENT WAS DISCOVERED BY A FAMILY MEMBER WHEN THEY ENTERED THE BEDROOM TO ASSIST THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA INDICATED THE LIBERTY SELECT CYCLER COMPLETED THERAPY, AND THERE WERE NO TREATMENT RELATED CONCERNS. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING AND TRANSPORTED THE PATIENT DIRECTLY TO THE MORGUE (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO HER ADVANCED AGE AND EXTENSIVE CARDIAC HISTORY. PER THE PDRN, THE EVENTS WERE UNRELATED TO ANY FRESENIUS PRODUCT(S) AND/OR DEVICE(S) DEFICIENCY OR MALFUNCTION. IT IS UNKNOWN IF THE DEVICE HAD BEEN SHIPPED BACK TO THE MANUFACTURER.Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED TO FRESENIUS CUSTOMER SUPPORT THAT A PATIENT PASSED AWAY ON THE CYCLER. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING HOURS OF (B)(6) 2025. THE PATIENT WAS DISCOVERED BY A FAMILY MEMBER WHEN THEY ENTERED THE BEDROOM TO ASSIST THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA INDICATED THE LIBERTY SELECT CYCLER COMPLETED THERAPY, AND THERE WERE NO TREATMENT RELATED CONCERNS. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING AND TRANSPORTED THE PATIENT DIRECTLY TO THE MORGUE (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO HER ADVANCED AGE AND EXTENSIVE CARDIAC HISTORY. PER THE PDRN, THE EVENTS WERE UNRELATED TO ANY FRESENIUS PRODUCT(S) AND/OR DEVICE(S) DEFICIENCY OR MALFUNCTION. IT IS UNKNOWN IF THE DEVICE HAD BEEN SHIPPED BACK TO THE MANUFACTURER. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, G4, H3. PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER WITH NO PHYSICAL DAMAGE NOTED. ALTHOUGH THERE WAS NO EVIDENCE OF DRIED FLUID PRESENT WITHIN THE CASSETTE COMPARTMENT, THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED THE CASSETTE MEMBRANE. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THE CYCLER UNDERWENT AND PASSED A SYSTEM AIR LEAK TEST, VALVE ACTUATION TEST, AND TEACH PUMPS TEST. THERE WERE NO VISUAL DISCREPANCIES ENCOUNTERED DURING THE INTERNAL INSPECTION. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED. Manufacturer Narrative: CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND SUBSEQUENT DEATH, AS THE PATIENT WAS UNDERGOING CCPD THERAPY WHEN SHE EXPIRED. THE PDRN REPORTED THE PRIMARY CAUSE OF DEATH WAS A CARDIAC ARREST, SECONDARY TO THE PATIENT¿S EXTENSIVE CARDIAC HISTORY AND ADVANCED AGE. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22188465&pc=GDW | 3005075853-2025-04429 | 2025-06-03 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-06-11 04:00 | GDW | CIRCULAR MECH STD 33MM | Failure to Fire | Necrosis; Shock; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K983536 |  | 1 | Event Description: IT WAS REPORTED THAT POST OP TO AN UNKNOWN PROCEDURE AT THE TIME OF COLORECTAL ANASTOMOSIS WITH THE CIRCULAR STAPLER, THE EQUIPMENT FAILED, RESULTING IN APPROXIMATELY 50% OF OPEN INTESTINAL LIGHT. THIS FAILURE SERIOUSLY COMPROMISED PATIENT SAFETY AND PROCEDURE INTEGRITY. THE DIRECT CONSEQUENCES OF THIS FAILURE FOR THE PATIENT WERE: SIGNIFICANT INCREASE IN SURGICAL TIME; INCREASED CHANCE OF POSTOPERATIVE INFECTION; DECOMPENSATION OF PREEXISTING COMORBIDITIES; NEED FOR A NEW SURGICAL PROCEDURE FOR CORRECTION, WITH ALL RISKS ASSOCIATED WITH A REINTERVENTION. THE SURGERY WAS DELAYED. IT WAS NECESSARY TO PERFORM THE ANASTOMOSIS MANUALLY, AND AN ILEOSTOMY WAS REQUIRED. Manufacturer Narrative: (B)(4). DATE SENT: 8/5/2025. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: IS THERE AN AUTOPSY REPORT AVAILABLE? NO, THERE ISN¿T. THE PATIENT DIED DUE TO PNEUMONIA AT THE END. IS THE SURGEON INTERESTED IN SPEAKING WITH ETHICON MEDICAL AND ENGINEERING STAFF? IF SO, PLEASE PROVIDE 3 DATES/TIME THE SURGEON WOULD BE ABLE TO MEET. I GUESS I PROVIDED ALL THE INFORMATION ABOUT THE EVENT, SO THERE¿S NO NEED FOR ONE CALL. Manufacturer Narrative: (B)(4) DATE SENT: 7/7/2025 H6: HEALTH EFFECT ¿ CLINICAL CODE GASTROINTESTINAL SYSTEM (E10) ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: WHY WAS THE NEED TO DIVERT REQUIRED? PATIENT UNDERWENT SURGERY FOR PERITONITIS DUE TO PERFORATED DIVERTICULITIS, WITH SIGNS OF HEMODYNAMIC INSTABILITY. THE SECOND SURGERY WAS PERFORMED TO TAKE DOWN THE COLOSTOMY, AND INTESTINAL TRANSIT RECONSTRUCTION. WHAT WERE THE INDICATIONS FOR SURGERY? COLOSTOMY TAKEDOWN. WHAT SURGICAL PROCEDURE WAS PERFORMED? LAPAROTOMY, AND ANASTOMOSIS BETWEEN TRANSVERSES COLON AND RECTUM. DID THE PATIENT RECEIVE ANY PREOPERATIVE CHEMOTHERAPY OR RADIATION? NO WERE THERE ANY ISSUES EXPERIENCED WITH THE DEVICE IN THE INITIAL SURGICAL PROCEDURE? NO WHAT HEALTHCARE PROFESSIONAL FIRED THE DEVICE AND WHAT IS HIS/HER EXPERIENCE WITH THE DEVICE? ME IN PERSON. 18 YEARS OF SURGICAL PRACTICE. WHERE IN THE GREEN GAP SETTING SCALE WAS THE INDICATOR LOCATED PRIOR TO FIRING (LOW-B, MIDDLE-B, OR HIGH-B)? MIDDLE DID THE HEALTHCARE PROFESSIONAL WAIT 15 SECONDS AFTER CLOSING THE DEVICE AND THEN RETIGHTEN PRIOR TO FIRING? YES WAS THE DEVICE DIFFICULT TO CLOSE? NO WAS THE DEVICE DIFFICULT TO FIRE? NO HOW MAY COUNTER-CLOCKWISE REVOLUTIONS OF THE ADJUSTING KNOB WERE USED TO OPEN THE DEVICE? TWO OR THREE DID THE HEALTHCARE PROFESSIONAL RECEIVE AUDIBLE & TACTILE FEEDBACK WHEN FIRING THE DEVICE? YES WERE THE DONUTS INSPECTED? IF SO, PLEASE DESCRIBE. YES, BOTH OF THEM WITH TOTAL CIRCUMFERENCE. WAS A COMPLETE TRANSECTION OF THE WHITE BREAKAWAY WASHER VISUALLY CONFIRMED? YES WERE THERE ANY ISSUES NOTED WITH STAPLE FORMATION AT ANY POINT THROUGHOUT THE CARE OF THE PATIENT? IF SO, PLEASE DESCRIBE THE SHAPE AND LOCATION. AFTER THAT, I SAW THAT 50% OF THE CIRCUMFERENCE WAS OPEN. THERE WAS STAPLES ONLY AT RECTUM SIDE AT THIS POINT. WHAT IS THE CURRENT STATUS OF THE PATIENT? INSTABLE, TRACHEOSTOMIZED, WITH NECROSIS OF FEET AND HANDS DUE TO PROLONGED SHOCK. UNDERGOING HEMODIALYSIS AND WITH PROBABLE NEUROLOGICAL SEQUELAE. DOES THE SURGEON BELIEVE THE POST-OPERATIVE COMPLICATIONS WERE RELATED TO AN ALLEGED DEFICIENCY OF THE DEVICE OR WERE THERE OTHER CONTRIBUTING FACTORS TO INCLUDE PATIENT TISSUE CONDITION? BOTH. THE PATIENT HAS SEQUELAE OF A STROKE, CHRONIC DISEASES, AND THE FACT THAT THE ANASTOMOSIS FAILED IN A LOW RECTAL STUMP LED US TO CHOOSE TO PERFORM A DIVERSION ILEOSTOMY. AS A CONSEQUENCE, THE PATIENT HAD DIFFICULTIES IN THE POSTOPERATIVE PERIOD. THE PATIENT HAD POST-SURGICAL COMPLICATIONS AND IS EVEN HOSPITALIZED. Manufacturer Narrative: (B)(4). DATE SENT 6/11/2025. D4 BATCH # UNK. H6: HEALTH EFFECT ¿ CLINICAL CODE GASTROINTESTINAL SYSTEM (E10). AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHY WAS THE NEED TO DIVERT REQUIRED? WHAT WERE THE INDICATIONS FOR SURGERY? WHAT SURGICAL PROCEDURE WAS PERFORMED? DID THE PATIENT RECEIVE ANY PREOPERATIVE CHEMOTHERAPY OR RADIATION? WERE THERE ANY ISSUES EXPERIENCED WITH THE DEVICE IN THE INITIAL SURGICAL PROCEDURE? WHAT HEALTHCARE PROFESSIONAL FIRED THE DEVICE AND WHAT IS HIS/HER EXPERIENCE WITH THE DEVICE? WHERE IN THE GREEN GAP SETTING SCALE WAS THE INDICATOR LOCATED PRIOR TO FIRING (LOW-B, MIDDLE-B, OR HIGH-B)? DID THE HEALTHCARE PROFESSIONAL WAIT 15 SECONDS AFTER CLOSING THE DEVICE AND THEN RETIGHTEN PRIOR TO FIRING? WAS THE DEVICE DIFFICULT TO CLOSE? WAS THE DEVICE DIFFICULT TO FIRE? HOW MAY COUNTER-CLOCKWISE REVOLUTIONS OF THE ADJUSTING KNOB WERE USED TO OPEN THE DEVICE? DID THE HEALTHCARE PROFESSIONAL RECEIVE AUDIBLE & TACTILE FEEDBACK WHEN FIRING THE DEVICE? WERE THE DONUTS INSPECTED? IF SO, PLEASE DESCRIBE. WAS A COMPLETE TRANSECTION OF THE WHITE BREAKAWAY WASHER VISUALLY CONFIRMED? WERE THERE ANY ISSUES NOTED WITH STAPLE FORMATION AT ANY POINT THROUGHOUT THE CARE OF THE PATIENT? IF SO, PLEASE DESCRIBE THE SHAPE AND LOCATION. WHAT IS THE CURRENT STATUS OF THE PATIENT? DOES THE SURGEON BELIEVE THE POST-OPERATIVE COMPLICATIONS WERE RELATED TO AN ALLEGED DEFICIENCY OF THE DEVICE OR WERE THERE OTHER CONTRIBUTING FACTORS TO INCLUDE PATIENT TISSUE CONDITION? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22188679&pc=QEW | 3005168196-2025-00197 | 2024-11-10 05:00 | Death | PENUMBRA, INC. | 2025-06-11 04:00 | QEW | INDIGO SYSTEM ASPIRATION CATHETER 12 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest |  |  | 1 | Event Description: THE PATIENT UNDERWENT A THROMBECTOMY PROCEDURE FOR PULMONARY EMBOLISM USING AN INDIGO ASPIRATION SYSTEM. IT WAS REPORTED THAT THE PATIENT WAS ADMITTED IN CRITICAL CONDITION AFTER EXPERIENCING CARDIAC ARREST REQUIRING RESUSCITATION PRIOR TO TREATMENT. THE PATIENT WAS TREATED WITH THE INDIGO ASPIRATION SYSTEM CAVT TUBING. THE PATIENT WAS THEN TRANSFERRED TO THE RECOVERY ROOM. THE PATIENT EXPERIENCED CIRCULATORY AND RESPIRATORY DECOMPENSATION FOLLOWED BY SUDDEN CARDIAC ARREST DUE TO ASYSTOLE. THE PATIENT WAS INTUBATED, TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU), AND EXPIRED A FEW HOURS LATER. THE PATIENT'S DEATH WAS REPORTED TO BE A SERIOUS ADVERSE EVENT WITH A POSSIBLE RELATIONSHIP TO THE INDIGO ASPIRATION SYSTEM AND THE INDEX PROCEDURE.Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED IN CRITICAL CONDITION AFTER EXPERIENCING CARDIAC ARREST AND REQUIRING RESUSCITATION. THE PATIENT HAD DEEP LOSS OF CONSCIOUSNESS, WIDE AND NON-REACTIVE PUPILS, AND EXTREME CIRCULATORY AND RESPIRATORY FAILURE. IN LABORATORY TESTS TAKEN PRIOR TO TREATMENT, THERE WAS INTENSIFIED METABOLIC AND RESPIRATORY ACIDOSIS AND EVIDENCE OF LIVER INJURY. THE PATIENT THEN UNDERWENT A SUCCESSFUL THROMBECTOMY PROCEDURE TO TREAT A PULMONARY EMBOLISM IN THE RIGHT AND LEFT PULMONARY ARTERIES AND ITS BRANCHING LOBAR ARTERIES USING AN INDIGO SYSTEM ASPIRATION CATHETER 12 (CAT12), A LIGHTNING ASPIRATION TUBING (LIGHTNING), AN INDIGO SYSTEM SEPARATOR 12 (SEP12), AND A PENUMBRA ENGINE (ENGINE). DURING THE PROCEDURE, SIGNIFICANT AMOUNTS OF EMBOLIC MATERIAL WERE REMOVED; HOWEVER, TOWARDS THE END OF THE SURGERY THE PATIENT REQUIRED CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) OXYGEN THERAPY TO MAINTAIN ATRIAL OXYGEN SATURATION AT GREATER THAN 80%, AND CATECHOLAMINES INFUSION. AFTER THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARDIAC CASE UNIT. THE PATIENT THEN EXPERIENCED CIRCULATORY AND RESPIRATORY DECOMPENSATION FOLLOWED BY SUDDEN CARDIAC ARREST DUE TO ASYSTOLE. AFTER AN APPROXIMATELY HOUR-LONG RESUSCITATION WITH A LUCAS DEVICE AND ADMINISTRATION OF ALTEPLASE, ADRENALINE, AND NORADRENALINE, A RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED. THE INTENSIVE LIFE SUPPORT THAT STARTED DURING RESUSCITATION WAS CONTINUED. DESPITE THE ADMINISTERED INTENSIVE CARE, THE PATIENT'S CONDITION KEPT DETERIORATING GRADUALLY. THE PATIENT EXPIRED A FEW HOURS LATER. THE PATIENT'S DEATH WAS REPORTED TO BE A SERIOUS ADVERSE EVENT WITH A POSSIBLE RELATIONSHIP TO THE CAT12 AND THE INDEX PROCEDURE. Manufacturer Narrative: THE PRODUCT WAS NOT RETURNED FOR EVALUATION. THEREFORE, A PHYSICAL DEVICE INVESTIGATION WAS UNABLE TO BE PERFORMED. FROM THE INFORMATION PROVIDED, THERE IS NO INDICATION THAT THERE WAS ANY DEVICE MALFUNCTION, NONCONFORMANCE, OR MISUSE THAT CONTRIBUTED TO THE REPORTED EVENT. THE PRODUCT LOT NUMBER WAS NOT PROVIDED; THEREFORE, THE MANUFACTURING RECORDS COULD NOT BE REVIEWED. Manufacturer Narrative: PLEASE NOTE THAT THE FOLLOWING SECTIONS ARE BEING UPDATED BASED ON ADDITIONAL INFORMATION PROVIDED BY THE PENUMBRA CLINICAL TEAM: 1. SECTION A. BOX 2. AGE AT TIME OF EVENT/AGE UNIT, 2. SECTION A. BOX 3A. SEX, 3. SECTION B. BOX 2. OUTCOMES ATTRIBUTED TO ADVERSE EVENT, 4. SECTION B. BOX 5. DESCRIBE EVENT OR PROBLEM, 5. SECTION D. BOX 1. BRAND NAME, 6. SECTION D. BOX 4. CATALOG #, UNIQUE IDENTIFIER, 7. SECTION H. BOX 6. HEALTH EFFECT IMPACT CODE 2.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22188806&pc=MVK | 3008642652-2025-05739 | 2025-04-11 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-11 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED A NON-LIFEVEST DEFIBRILLATION. THE DEVICE WAS STARTED UP AT 11:12:34 ON (B)(6) 2025. AT 22:33:42 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VT FROM 140-160 BPM WITH CPR/MOTION ARTIFACT DEGRADING TO VF WITH VARYING RATES AND AMPLITUDES. THE LIFEVEST PROPERLY DETECTED VT/VF. CPR/MOTION ARTIFACT AND VF WITH VARYING RATES/AMPLITUDES PREVENTED THE LIFEVEST FROM DELIVERING A TREATMENT TO THE PATIENT. AT 22:33:45, THE PATIENT RECEIVED THE NON-LIFEVEST DEFIBRILLATION. RHYTHM AT THE TIME OF THE NON-LIFEVEST DEFIBRILLATION WAS VF WITH VARYING RATES/AMPLITUDES AND MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS BRADYCARDIA @ 40 BPM WITH CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 22:42:17 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22189313&pc=PRL | 2017233-2025-06296 | 2025-05-20 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-11 04:00 | PRL | GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS | Difficult to Remove | Perforation of Vessels | P160021 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT UNDERWENT TREATMENT FOR AORTOILIAC OCCLUSIVE DISEASE IN THE DISTAL AORTA, BILATERAL COMMON AND EXTERNAL ILIAC ARTERIES. GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) DEVICES WERE IMPLANTED IN THE EXTERNAL ILIAC ARTERY. THE PHYSICIAN BEGAN WITH FEMORAL ENDARTERECTOMY OF THE COMMON FEMORAL ARTERIES. AFTER COMPLETION OF THE ENDARTERECTOMY, BILATERAL ACCESS WAS GAINED AND CORDIS 8 FR BRITE TIP® INTRODUCER SHEATHS WERE INSERTED. THE LESIONS WERE CROSSED; HOWEVER, THE PHYSICIAN ENCOUNTERED A SUBINTIMAL POSITION AT THE DISTAL AORTA. INITIALLY, A COOK MEDICAL BEBACK CROSSING CATHETER WAS ATTEMPTED TO BE USED. A PHILIPS PIONEER PLUS IVUS-GUIDED CATHETER WAS SELECTED FOR RE-ENTRY INTO THE DISTAL AORTA. TWO 8X79 VBX DEVICES WERE DEPLOYED BILATERALLY INTO THE DISTAL AORTA AND COMMON ILIAC ARTERIES USING A KISSING STENT TECHNIQUE. TO EXCLUDE THE AORTIC DISEASE, THE VBX DEVICES WERE SLIGHTLY RAISED IN THE BIFURCATION. TO PRESERVE THE HYPOGASTRIC ARTERY, AN 8X60 BOSTON SCIENTIFIC INNOVA¿ VASCULAR SELF-EXPANDING STENT SYSTEM (INNOVA STENT) WAS DEPLOYED INTO THE 8X79 VBX DEVICE THAT WAS IMPLANTED ON THE RIGHT SIDE. THE INNOVA STENT WAS OVERLAPPED WITH AN 8X59 VBX DEVICE THAT EXTENDED INTO THE EXTERNAL ILIAC ARTERY. THE 8X59 VBX BALLOON WAS INFLATED AGAIN TO BALLOON THE OVERLAP OF THE INNOVA STENT AND VBX DEVICE. AS REPORTED, THERE WERE NO ISSUES WITH BALLOON INFLATION. HOWEVER, WHEN ATTEMPTING TO WITHDRAW THE BALLOON CATHETER, THE PHYSICIAN NOTICED THE BALLOON CATHETER WAS STUCK TO THE GUIDEWIRE, PREVENTING ADVANCEMENT OR RETRACTION. SEVERAL ATTEMPTS WERE MADE TO DISENGAGE THE CATHETER BY PULLING NEGATIVE ON THE BALLOON, BUT ALL WERE UNSUCCESSFUL. IT IS UNKNOWN WHY THE BALLOON CATHETER BECAME STUCK ON THE GUIDEWIRE. THE BALLOON CATHETER, GUIDEWIRE, AND SHEATH WERE ALL REMOVED AS ONE FROM THE PATIENT. AFTER REMOVAL OF THE BALLOON CATHETER, GUIDEWIRE, AND SHEATH, THE PATIENT STARTED BLEEDING HEAVILY AND EXPERIENCED A DROP IN BLOOD PRESSURE, AND SUBSEQUENTLY CODED. A COOK MEDICAL CODA® BALLOON CATHETER WAS INSERTED FROM THE CONTRALATERAL SIDE TO CONTROL THE BLEEDING. AN 8 X 15 GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE WAS SUCCESSFULLY DEPLOYED ACROSS THE ILIAC ARTERY PERFORATION. CPR WAS PERFORMED FOR A SIGNIFICANT DURATION, BUT THE PATIENT EXPIRED. THE PATIENT'S DROP IN BLOOD PRESSURE WAS CAUSED BY AN ARTERIAL RUPTURE AND INTERNAL BLEEDING. HOWEVER, THE EXACT CAUSE OF DEATH REMAINS UNKNOWN. THE PHYSICIAN STATED THERE WAS SPECULATION THAT THE CATHETER, WHERE THE BALLOON WAS MOUNTED, MAY HAVE OVALIZED AND CAUSED THE CATHETER TO GET STUCK ON THE WIRE. THIS COULD NOT BE CONFIRMED AS THE BALLOON CATHETER AND GUIDEWIRE WERE IMMEDIATELY THROWN AWAY. Manufacturer Narrative: SECTION H6 UPDATED TO REFLECT COMPLETION OF INVESTIGATION. THE REPORTED COMPLAINT COULD NOT BE INDEPENDENTLY CONFIRMED BECAUSE NO PRODUCT NOR IMAGES ENABLING DIRECT ASSESSMENT OF PRODUCT PERFORMANCE WERE RETURNED FOR EVALUATION. THEREFORE, AN INDEPENDENT ASSESSMENT OF PRODUCT PERFORMANCE COULD NOT BE CONDUCTED. AS REPORTED, THE PATIENT'S DROP IN BLOOD PRESSURE WAS CAUSED BY AN ARTERIAL RUPTURE AND INTERNAL BLEEDING. IT IS UNKNOWN WHETHER THIS PATIENT HARM OCCURRED AS A RESULT OF DILATING AT THE OVERLAP OF THE VBX DEVICE AND THE BARE METAL STENT OR DURING REMOVAL OF THE DELIVERY SYSTEM, GUIDEWIRE, AND INTRODUCER SHEATH. IT WAS REPORTED THAT THE EXACT CAUSE OF DEATH WAS UNKNOWN. THE CAUSE FOR THE REPORTED COMPLAINT COULD NOT BE ESTABLISHED WITH THE AVAILABLE INFORMATION. FURTHER INFORMATION REGARDING THIS EVENT WAS REQUESTED BY GORE, BUT NO FURTHER INFORMATION HAS BEEN REPORTED, THEREFORE THIS INVESTIGATION IS CONSIDERED COMPLETE. Manufacturer Narrative: CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. H6 CODE C19: A REVIEW OF THE MANUFACTURING RECORDS INDICATED THE DEVICE MET PRE-RELEASE SPECIFICATIONS. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22189480&pc=QFG | 3013756811-2025-101194 | 2025-03-28 04:00 | Death | TANDEM DIABETES CARE | 2025-06-11 04:00 | QFG | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Diabetic Ketoacidosis | K203234 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. CAUSE OF DEATH WAS DUE TO DIABETIC KETOACIDOSIS (DKA) WITH A BLOOD GLUCOSE (BG) LEVEL OF 702 MG/DL. ACCORDING TO THE HEALTHCARE PROVIDER (HCP), THE PUMP INDICATED ELEVATED CONTINUOUS GLUCOSE MONITOR SENSOR READINGS STARTING ON (B)(6); CAUSE WAS NOT KNOWN. THE HCP ALSO REPORTED THAT THE CUSTOMER PERFORMED AN INFUSION SET CHANGE AND DELIVERED CORRECTION BOLUSES VIA THE PUMP. AN INVESTIGATION PERFORMED BY FORENSIC MEDICINE IN SWEDEN FOUND NO UNDERLYING CAUSE FOR THE DKA, AND NO INFECTION OR POISONING WAS FOUND. NO ADDITIONAL INFORMATION COULD BE CONFIRMED AT THE TIME OF THIS REPORT, AS PUMP DATA IS NOT AVAILABLE FOR REVIEW. Manufacturer Narrative: THE DEVICE IS EXPECTED TO BE RETURNED; HOWEVER, THE DEVICE HAS NOT YET BEEN RECEIVED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF THE DEVICE IS RECEIVED AND EVALUATION IS PERFORMED. Manufacturer Narrative: D2A D2A
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22189564&pc=HAW | 1723170-2025-02288 | 2025-01-14 05:00 | Death | MEDTRONIC NAVIGATION, INC | 2025-06-11 04:00 | HAW | MEDTRONIC NAVIGATION | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage |  |  | 1 | Event Description: CITATION: DOI: 10.1002/ACN3.70000. SUMMARY: OBJECTIVE: A SUBSTANTIAL PART OF CENTRAL NERVOUS SYSTEM (CNS) DISORDERS REMAINS UNEXPLAINED, DESPITE VARIOUS NEW AND MINIMALLY INVASIVE DIAGNOSTIC TECHNIQUES. WITHIN THIS RAPIDLY DEVELOPING DIAGNOSTIC FIELD, THE PRECISE ROLE OF BRAIN BIOPSY IS UNKNOWN. WE AIMED TO STUDY THE CLINICAL IMPACT AND SAFETY OF BRAIN BIOPSIES IN UNEXPLAINED CNS DISORDERS. METHODS: IN THIS RETROSPECTIVE COHORT STUDY, WE INCLUDED ALL ADULT PATIENTS WHO WERE REFERRED FOR A DIAGNOSTIC WORK-UP TO OUR ACADEMIC CENTER WITH NEURO-INFLAMMATORY, NEURO-ONCOLOGICAL, AND NEURO-INFECTIOUS EXPERTISE AND UNDERWENT A BRAIN BIOPSY BETWEEN (B)(6) 2010 AND (B)(6) 2023. TYPICAL CASES OF CNS NEOPLASMS AND INFECTIONS WERE NOT ANALYZED. BRAIN BIOPSIES WERE EVALUATED WITH RESPECT TO DIAGNOSTIC AND THERAPEUTIC IMPACT AND COMPLICATION RISK. RESULTS: BRAIN BIOPSY WAS PERFORMED IN 587 PATIENTS. NINETY-FOUR PATIENTS WITH A CNS DISORDER OF UNKNOWN CAUSE, WITH 107 BIOPSIES, WERE ANALYZED (44% FEMALE, MEDIAN AGE 58 YEARS). POSTOPERATIVE DIAGNOSES INCLUDED BRAIN TUMORS/LYMPHOMAS (37/94, 39%), INFLAMMATORY DISORDERS (11/94, 12%), INFECTIONS (8/94, 9%), AUTOIMMUNE ENCEPHALITIS (8/94, 9%), AND PRIMARY ANGIITIS OF THE CNS (4/94, 4%). DIAGNOSTIC YIELD OF BRAIN BIOPSY WAS 62%, INCREASING UP TO 72% AFTER REPEAT BIOPSIES, AS 10 ADDITIONAL PATIENTS WERE DIAGNOSED WITH A BRAIN TUMOR. IN 77% OF PATIENTS, BRAIN BIOPSY CHANGED THE TREATMENT STRATEGY. SYMPTOMATIC INTRACRANIAL HEMORRHAGE OCCURRED IN 4 OF 107 BRAIN BIOPSIES (4%). INTERPRETATION: IN A SELECTED POPULATION OF PATIENTS WITH UNEXPLAINED CNS DISORDERS, CLINICAL IMPACT OF BRAIN BIOPSIES IS HIGH, WHILE BEING RELATIVELY SAFE. A MULTIDISCIPLINARY TEAM APPROACH IS FUNDAMENTAL IN ESTABLISHING OPTIMAL INDICATION FOR BRAIN BIOPSY AND SUBSEQUENT TREATMENT DECISIONS. REPORTED EVENT(S): THIS WAS A RETROSPECTIVE STUDY OF 94 PATIENTS, 41 FEMALE AND 53 MALE. THE AVERAGE PATIENT AGE AT THE TIME OF BIOPSY WAS 58 YEARS. POST-OPERATIVE DEATH WITHIN 30 DAYS OCCURRED IN 3 OF 107 BRAIN BIOPSIES, ALTHOUGH DEATH WAS CONSIDERED CLEARLY BIOPSY-RELATED IN 1 BRAIN BIOPSY, CONCERNING A PATIENT WITH A MASSIVE INTRACRANIAL HEMORRHAGE 5 DAYS POSTOPERATIVELY. IN ONE CASE, DEATH WAS CONSIDERED NOT BIOPSY-RELATED, IN VIEW OF THE LONG DURATION BETWEEN SURGICAL PROCEDURE AND DEATH (29 DAYS), WHEREAS A DIRECT RELATION BETWEEN BRAIN BIOPSY AND DEATH COULD NOT BE RULED OUT IN THE REMAINING PATIENT.Event Description: ADDITIONAL INFORMATION WAS RECEIVED. IT WAS REPORTED THAT NO DEVICES NOR SOFTWARE OF MEDTRONIC WERE USED FOR THESE BIOPSIES. Manufacturer Narrative: G2: THIS EVENT OCCURRED IN THE NETHERLANDS, SEE SECTION E. H3, H6: NO PRODUCTS WERE RETURNED TO MEDTRONIC FOR ANALYSIS. CODES B17, C20, AND D15 ARE APPLICABLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H2, B5: EVENT DETAILS HAVE BEEN UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22189608&pc=NPT | 2135147-2025-03098 | 2025-05-19 04:00 | Death | ABBOTT MEDICAL | 2025-06-11 04:00 | NPT | NAVITOR VISION | Perivalvular Leak | No Clinical Signs, Symptoms or Conditions | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 27MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM. THE PATIENT'S ANNULAR DIMENSIONS WERE PERIMETER 76.3, DIAMETER- 20.6 X 28.1, AREA- 440.3. PRE-DILATATION WAS PERFORMED WITH A 22MM TRUE BALLOON. THE NAVITOR WAS SUCCESSFULLY IMPLANTED INTO THE PATIENT AT A FINAL DEPTH OF 3MM. HOWEVER, IMMEDIATE ANGIOGRAM SHOWED THE PATIENT HAD MODERATE LEAK DUE TO NOT FULLY EXPANDING. THEREFORE, THE VALVE WAS POST-DILATED TO HELP MITIGATE THE LEAK. POST-DILATION BALLOON CAUSED A RUPTURE AND ON (B)(6) 2025, THE PATIENT PASSED AWAY. Manufacturer Narrative: AN EVENT OF PARAVALVULAR LEAK AND PATIENT DEATH WERE REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED AND WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. IT WAS INDICATED THAT VALVE WAS IMPLANTED SUCCESSFULLY AT 3 MM DEPTH. FIELD ALSO INDICATED THAT THE POST-DILATATION WAS PERFORMED. THE CAUSE OF REPORTED EVENT OF PARAVALVULAR LEAK COULD NOT BE DETERMINED. THE CAUSE OF PATIENT DEATH WAS RELATED PROCEDURAL CONDITIONS (POST-DILATATION). THE REPORTED UNEXPECTED MEDICAL INTERVENTION IS RESULTS OF CASE-SPECIFIC CIRCUMSTANCES, AS THE POST-IMPLANT VALVULOPLASTY WAS PERFORMED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. FROM MEDICAL REVIEW: THE CAUSE OF DEATH WAS BELIEVED TO BE ANNULAR RUPTURE SECONDARY TO POST-DILATION. IT IS UNKNOWN AT THIS TIME IF THE PATIENT UNDERWENT SURGERY TO REPAIR THE ANNULAR RUPTURE. Manufacturer Narrative: THE DEVICE WILL NOT BE RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22190919&pc=DRA | 2135147-2025-03105 | 2025-05-20 04:00 | Death | ABBOTT MEDICAL | 2025-06-11 04:00 | DRA | MITRACLIP | Difficult to Insert | Fistula; Perforation of Vessels | K190167 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT MIXED MITRAL REGURGITATION (MR) WITH A GRADE OF 3-4. A STEERABLE GUIDE CATHETER (SGC) WAS PREPARED PER THE INSTRUCTIONS FOR USE (IFU) AND INSERTED INTO THE LEFT ATRIUM (LA). HOWEVER, DIFFICULTIES OBTAINING GROIN ACCESS OCCURRED, AND MULTIPLE ATTEMPTS TO ACCESS THE FEMORAL VEIN WERE MADE, AND A STRUCTURE WAS VISIBLE, ATTACHED TO THE WIRE. IT WAS NOTED THAT IT WAS ONLY VISIBLE FOR A MOMENT. IN THE PHYSICIAN¿S OPINION, THE STRUCTURE WAS A PIECE OF TISSUE ORIGINATING FROM THE VENOUS VASCULAR SYSTEM AND THE SGC DID NOT CAUSE OR CONTRIBUTE TO THE TISSUE DAMAGE. ADDITIONAL HEPARIN WAS ADMINISTERED. IT WAS NOTED THAT THE PHYSICIAN DID NOT WANT TO ASPIRATE. THE SGC, INCLUDING THE WIRE, WAS REMOVED FROM THE PATIENT. THE DEVICES WERE EXAMINED, AND A SMALL TISSUE STRUCTURE WAS OBSERVED. THE PROCEDURE WAS CONTINUED. ONE CLIP WAS THEN IMPLANTED, REDUCING MR TO A GRADE OF 1-2. IT WAS NOTED THAT THE CLIP WAS STABLE ON BOTH LEAFLETS. ON (B)(6) 2025, THE PATIENT WAS DIAGNOSED WITH ARTERIOVENOUS (AV) FISTULA. IN THE PHYSICIAN'S OPINION, THE SGC DID NOT CAUSE OR CONTRIBUTE TO THE AV FISTULA. ON (B)(6) 2025, DURING THE TREATMENT OF THE COMPLICATION, THE PATIENT REQUIRED RESUSCITATION AND PASSED AWAY. IN THE PHYSICIAN'S OPINION, THE SGC DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT DEATH. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO SIMILAR COMPLAINTS REPORTED FROM THIS LOT. BASED ON THE INFORMATION PROVIDED, THE REPORTED DIFFICULT TO INSERT INTO ANATOMY CANNOT BE DETERMINED. ADDITIONALLY, THE REPORTED PATIENT EFFECT OF PERFORATION OF VESSELS, FISTULA, AND DEATH ARE LISTED IN THE MITRACLIP SYSTEM INSTRUCTIONS FOR USE AND ARE KNOWN POSSIBLE COMPLICATION ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED PERFORATION OF VESSELS, FISTULA, AND DEATH COULD NOT BE DETERMINED. THE REPORTED HOSPITALIZATION, SURGICAL INTERVENTION, MEDICATION REQUIRED, AND UNEXPECTED MEDICAL INTERVENTION WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: THE DEVICE WILL NOT BE RETURNED FOR EVALUATION AS THE DEVICE WAS REPORTEDLY DISCARDED. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22191148&pc=NRY | 3005168196-2025-00182 | 2025-05-10 04:00 | Death | PENUMBRA, INC. | 2025-06-11 04:00 | NRY | PENUMBRA SYSTEM RED 72 REPERFUSION CATHETER | Fracture | Foreign Body Embolism | K211654 |  | 1 | Event Description: THE PATIENT WAS UNDERGOING A THROMBECTOMY PROCEDURE TO TREAT A LARGE VESSEL OCCLUSION (LVO) IN THE RIGHT INTERNAL CAROTID ARTERY (ICA) USING A PENUMBRA SYSTEM RED 72 REPERFUSION CATHETER (RED72) AND A BENCHMARK BMX96 ACCESS SYSTEM (BMX96). DURING THE PROCEDURE, AFTER PERFORMING A DIRECT ASPIRATION FIRST-PASS TECHNIQUE (ADAPT) FOR APPROXIMATELY NINETY SECONDS, THE PHYSICIAN RETRACTED THE RED72 UNDER ASPIRATION AND NOTICED THAT A DISTAL SEGMENT OF THE RED72 FRACTURED WITHIN THE RIGHT ICA. THE PHYSICIAN DECIDED NOT TO REMOVE THE FRACTURED SEGMENT OF THE RED72 AND CONTINUED THE PROCEDURE. THE PROCEDURE WAS COMPLETED USING ANOTHER RED72, THE SAME BMX96, AND A NON-PENUMBRA MICROCATHETER WITH A THROMBOLYSIS IN CEREBRAL INFARCTION (TICI) SCORE OF 2A+. THE PHYSICIAN THEN STARTED A DOUBLE ANTI-AGGREGATION THERAPY. IT WAS REPORTED THAT THE CAROTID BRANCH THAT PREVIOUSLY HAD SLOWER FLOW, HAD INCREASED FLOW DUE TO THE IMPLANTED PART OF THE RED72. THE PART OF THE RED72 WAS NOT REMOVED DUE TO THE PATIENT'S DELICATE MEDICAL CONDITION. IT WAS ALSO REPORTED THAT THE BRANCH OF THE RIGHT ICA WHERE THE PART OF THE RED72 REMAINED, LATER BECAME RE-OCCLUDED. THE PATIENT WAS ALSO TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) FOR RECOVERY. HOWEVER, THE PATIENT LATER PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: EVALUATION OF THE RETURNED RED72 CONFIRMED THAT THE CATHETER WAS FRACTURED. EVALUATION REVEALED APPROXIMATELY 8 CM OF STRETCHING PROXIMAL TO THE FRACTURED LOCATION. THIS IS CONSISTENT WITH THE DEVICE BEING RETRACTED AGAINST RESISTANCE PRIOR TO FRACTURE. IF THE RED72 IS RETRACTED AGAINST RESISTANCE, DAMAGE SUCH AS STRETCHING AND SUBSEQUENT FRACTURE MAY OCCUR. IT WAS REPORTED THAT THE PATIENT'S ANATOMY WAS TORTUOUS. THIS MAY HAVE CONTRIBUTED TO THE RESISTANCE EXPERIENCED DURING THE PROCEDURE. THE ROOT CAUSE COULD NOT BE DETERMINED. FURTHER EVALUATION REVEALED OVALIZATION AND KINKS ON THE CATHETER SHAFT. THIS DAMAGE WAS LIKELY INCIDENTAL TO THE REPORTED COMPLAINT AND THE ROOT CAUSE COULD NOT BE DETERMINED. PENUMBRA CATHETERS ARE INSPECTED DURING IN-PROCESS INSPECTION AND DURING QUALITY INSPECTION AFTER MANUFACTURING. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS. Manufacturer Narrative: THIS DEVICE IS AVAILABLE FOR RETURN. A FOLLOW UP MDR WILL BE SUBMITTED UPON COMPLETION OF THE DEVICE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22173707&pc=OZO | 2032227-2025-188397 | 2023-05-05 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-10 04:00 | OZO | 630G INSULIN PUMP MMT-1715KL 630G | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Heart Failure/Congestive Heart Failure | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY AT HOME ON (B)(6), 2023. THE CUSTOMER WAS NOT ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS A HEART CONDITION AND COMPLICATIONS OF DIABETES. CUSTOMER BLOOD GLUCOSE WAS UNKNOWN AT THE TIME OF THE INCIDENT. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1715KL, UNOMEDICAL, AND MMT-UNKNOWN. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING, CUSTOMER HAS WORN INSULIN PUMP AT THE TIME OF PASSING. CUSTOMER WAS USING THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT AND THE CUSTOMER WAS USING SENSORS. THERE IS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. NO PRODUCT RETURN IS REQUIRED FOR MMT-UNKNOWN. NO PRODUCT RETURN IS REQUIRED FOR MMT-1715KL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22174564&pc=BZD | 2518422-2025-046664 | 2023-06-15 04:00 | Death | RESPIRONICS, INC. | 2025-06-10 04:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Degraded | Cancer; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED/RECERTIFIED DREAM STATION CPAP WITH AN ALLEGATION METASTATIC ESOPHAGEAL CARCINOMA. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22174728&pc=GEI | 1717344-2025-00779 | 2024-07-16 04:00 | Death | COVIDIEN MFG DC BOULDER | 2025-06-10 04:00 | GEI | VALLEYLAB | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Hypovolemic Shock | K151649 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN OPEN TOTAL ABDOMINAL HYSTERECTOMY AND BILATERAL SALPINGO-OOPHORECTOMY, THE PATIENT GOT INTO HYPOVOLEMIC SHOCK THE FOLLOWING MORNING. THE BLEEDING VESSELS WERE IDENTIFIED WHEN THE PATIENT WAS TAKEN BACK TO THEATRE AS AN EMERGENCY. A MASSIVE TRANSFUSION PROTOCOL (MTP) WAS INITIATED, RESUSCITATED AND GAVE FLUIDS AND THEN TAKEN BY THE AMBULANCE CREW TO TRANSFER TO A DIFFERENT HOSPITAL WHERE THEY OPENED THE PATIENT UP AND PERFORMED HEMOSTASIS BUT THE PATIENT SUBSEQUENTLY DIED OF HYPOVOLEMIC SHOCK. AN INQUEST WAS MADE BY THE CORONER, AND IT WAS IDENTIFIED THAT THE PATIENT EXPERIENCED A DELAYED INTRA-ABDOMINAL BLEED FROM 2 OF THE 9 BLOOD VESSELS THAT WERE SEALED WITH THE LIGASURE HAD FAILED. Manufacturer Narrative: D10 CONCOMITANT PRODUCTS: UNKNOWN LIGASUR, UNKNOWN LIGASURE INSTRUMENT (LOT#: UNKNOWN). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22174788&pc=DYB | 2015691-2025-04719 | 2025-04-11 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-10 04:00 | DYB | EDWARDS ESHEATH PLUS INTRODUCER SHEATH | Physical Resistance/Sticking | Perforation of Vessels | K200258 |  | 1 | Event Description: AS REPORTED BY OUR GERMAN AFFILIATES, DURING IMPLANT OF A 23MM SAPIEN 3 ULTRA (S3U) VALVE IN AORTIC POSITION BY RIGHT TRANSFEMORAL APPROACH. DURING THE INSERTION OF THE COMMANDER WITH THE VALVE THROUGH THE ESHEATH+, THE VALVE COULD NOT ADVANCE. THE SYSTEM WAS REMOVED AND DISCARDED WITH NO PATIENT INJURY REPORTED AT THIS POINT. A NEW 23MM S3U VALVE WAS PREPARED AND A 16FR ESHEATH+ WAS USED THIS TIME. THE NEW 16FR ESHEATH+ WAS EASILY INSERTED BUT IT WAS ALSO DIFFICULT TO ADVANCE THE VALVE. AFTER A CONSIDERABLE EFFORT AND FORCE, THE VALVE WAS ABLE TO EXIT THE SHEATH AND THE VALVE WAS SUCCESSFULLY DEPLOYED. AFTER THE REMOVAL OF THE ESHEATH+, DESPITE THE USE OF TWO PROSTYLE SUTURES DEVICES AND 8F ANGIOSEAL, THE VESSEL WAS NOT TIGHTLY CLOSED. A NON-EDWARDS CROSSOVER SHEATH WAS INSERTED FROM THE LEFT TO THE RIGHT VESSEL. THEN, ANGIOGRAPHY REVEALED THAT THE RIGHT COMMON FEMORAL WAS PERFORATED AND A NON-EDWARDS BALLOON WAS ADVANCE TO CONTROL THE BLEEDING. THE PATIENT WAS THEN MOVED TO VASCULAR SURGERY AND THE VESSEL WAS RESECTED AND A LONG 6MM/ 4CM LONG PROSTHETIC INTERPOSITION WAS IMPLANTED. AFTER REPEATED ANGIOGRAPHY, THE EXTERNAL ILIAC ARTERY WAS SUB TOTALLY OCCLUDED BY A DISSECTED DUCT, AND A STENT WAS IMPLANTED. THE FINAL ANGIOGRAPHY SHOWED GOOD VISUALIZATION OF ALL VESSELS. THE PATIENT WAS MOVED TO ICU FOR MONITORING BUT PASSED AWAY DURING THE NIGHT DUE TO HEMORRHAGIC HYPOVOLEMIC SHOCK AFTER MASSIVE BLEEDING FROM THE COMMON FEMORAL ARTERY. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO THE INVESTIGATION BEING COMPLETED. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: B4, G3, G6, H2, H6, H11. THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. REVIEW OF PROVIDED 3MENSIO REVEALED THE FOLLOWING OBSERVATIONS: PATIENT'S RIGHT ACCESS VESSEL HAD PRESENCE OF TORTUOSITY, CALCIFICATION, AND UNDERSIZED DIMENSIONS. MINIMAL LUMINAL DIAMETER FOR 16F ESHEATH IS ' 6.0MM. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH AND POTENTIAL VALVE FRAME DAMAGE USING THE S3U/S3UR VALVE CONFIGURATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, UNDERSIZED VESSELS, AND/OR STEEP INSERTION ANGLE. IN ADDITION, VALVE FRAME AND/OR SHEATH DAMAGE CAN BE A RESULT OF INCREASED PUSH FORCE AND ANY EXCESSIVE DEVICE MANIPULATION OR SHEATH-VALVE INTERACTION. THE COMPLAINT FOR DIFFICULTY ADVANCING THROUGH SHEATH WAS CONFIRMED EMPIRICALLY BASED ON MEDICAL RECORD. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION, TORTUOSITY, UNDERSIZED VESSELS) MAY HAVE CONTRIBUTED TO THE REPORTED EVENT, AS CONFIRMED VIA 3MENSIO. CALCIFICATION CAN REDUCE THE VESSEL LUMEN DIAMETER AND MAY INCREASE RESTRICTION LEADING TO RESISTANCE. CALCIFICATION CAN ALSO RESULT IN THE CREATION OF SUB-OPTIMAL ANGLES DURING DELIVERY SYSTEM INSERTION THAT MAY LEAD TO RESISTANCE. SCRATCHES OBSERVED ON THE SHEATH SHAFT CAN BE INDICATIVE OF THE PRESENCE OF CALCIFIED NODULES WITHIN THE ACCESS VESSEL. TORTUOUS PATIENT ANATOMY CAN CREATE SUB-OPTIMAL ANGLES THAT CAN LEAD TO NON-COAXIAL ALIGNMENT BETWEEN THE DELIVERY SYSTEM WITH CRIMPED VALVE AND SHEATH INNER LUMEN. KINKS OR CURVATURE ALONG THE SHEATH SHAFT CAN BE INDICATIVE OF THE PRESENCE OF VESSEL TORTUOSITY. UNDERSIZED VESSELS (E.G. DUE TO ANATOMICAL VARIATION, PRE-EXISTING STENT OR GRAFT, SCAR TISSUE, OR FIBROSIS) CAN CREATE A RESTRICTED PATHWAY OR CONSTRAINED CONDITION RESULTING IN DIFFICULTY DURING SHEATH EXPANSION AND INCREASED RESISTANCE DURING THE ADVANCEMENT OF THE DELIVERY SYSTEM. THE CAUSE OF THE RIGHT COMMON FEMORAL PERFORATION, WITH SUBSEQUENT BLEEDING AND DEATH IS UNKNOWN. HOWEVER, MAYBE RELATED TO THE TWO PROSTYLE SUTURES DEVICES USED FOR VESSEL CLOSURE OR RELATED TO THE RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR AN UPDATE TO THE ENGINEERING EVALUATION. THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. REVIEW OF PROVIDED 3MENSIO REVEALED THE FOLLOWING OBSERVATIONS: PATIENT'S RIGHT ACCESS VESSEL HAD PRESENCE OF TORTUOSITY, CALCIFICATION, AND UNDERSIZED DIMENSIONS. MINIMAL LUMINAL DIAMETER FOR 16F ESHEATH IS ' 6.0MM. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATION'S INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH AND POTENTIAL VALVE FRAME DAMAGE USING THE S3U/S3UR VALVE CONFIGURATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCE'S. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, UNDERSIZED VESSELS, AND/OR STEEP INSERTION ANGLE. IN ADDITION, VALVE FRAME AND/OR SHEATH DAMAGE CAN BE A RESULT OF INCREASED PUSH FORCE AND ANY EXCESSIVE DEVICE MANIPULATION OR SHEATH-VALVE INTERACTION. THE COMPLAINT FOR DIFFICULTY ADVANCING THROUGH SHEATH WAS CONFIRMED EMPIRICALLY BASED ON MEDICAL RECORD. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION, TORTUOSITY, UNDERSIZED VESSELS) MAY HAVE CONTRIBUTED TO THE REPORTED EVENT, AS CONFIRMED VIA 3MENSIO. CALCIFICATION CAN REDUCE THE VESSEL LUMEN DIAMETER AND MAY INCREASE RESTRICTION LEADING TO RESISTANCE. CALCIFICATION CAN ALSO RESULT IN THE CREATION OF SUB-OPTIMAL ANGLES DURING DELIVERY SYSTEM INSERTION THAT MAY LEAD TO RESISTANCE. TORTUOUS PATIENT ANATOMY CAN CREATE SUB-OPTIMAL ANGLES THAT CAN LEAD TO NON-COAXIAL ALIGNMENT BETWEEN THE DELIVERY SYSTEM WITH CRIMPED VALVE AND SHEATH INNER LUMEN. UNDERSIZED VESSELS (E.G. DUE TO ANATOMICAL VARIATION, PRE-EXISTING STENT OR GRAFT, SCAR TISSUE, OR FIBROSIS) CAN CREATE A RESTRICTED PATHWAY OR CONSTRAINED CONDITION RESULTING IN DIFFICULTY DURING SHEATH EXPANSION AND INCREASED RESISTANCE DURING THE ADVANCEMENT OF THE DELIVERY SYSTEM. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO A FOLLOW UP RESPONSE BEING RECEIVED. THIS IS ONE OF TWO MANUFACTURERS BEING SUBMITTED FOR THIS CASE. THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22174971&pc=OZP | 2032227-2025-188866 | 2024-11-29 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-10 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Appropriate Term/Code Not Available | Hyperglycemia; Pneumonia; Sepsis; Diabetic Ketoacidosis; Ischemic Heart Disease; Heart Failure/Congestive Heart Failure | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER HAD PASSED AWAY, AND THE CAUSE OF DEATH WAS CONGESTIVE HEART FAILURE, ISCHEMIC HEART DISEASE, AND DIABETES. THE CUSTOMER REPORTED A BLOOD GLUCOSE VALUE OF 800 MG/DL. THE CUSTOMER REPORTED THE SEPSIS, DIABETIC KETOACIDOSIS, HYPERGLYCEMIA, AND DOUBLE PNEUMONIA. THE EVENT INVOLVED PRODUCTS MMT-397A, MMT-1884, AND MMT-397A. TROUBLESHOOTING WAS PERFORMED AND FOUND THAT THE DOCUMENT OF ANY HEALTH ISSUES OR ILLNESS THAT MAY HAVE CONTRIBUTED OR LED UP TO PASSING WAS THE HEART DISEASE. THE CUSTOMER WAS DECEASED DUE TO THE INSULIN PUMP WAS NOT WORKING AND SO THAT THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP. NO PRODUCT RETURN IS REQUIRED FOR MMT-397A, MMT-1884, AND MMT-397A.Event Description: UPDATED SUMMARY: THE EVENT IS NOT-REPORTABLE AS PUMP WAS NOT USED AT THE TIME OF REPORTED EVENT AND INSULIN PUMP WAS ALREADY REPORTED UNDER REGULATORY REPORT 2032227-2025-187641. SO, EVENT IS A NOT-REPORTABLE SCENARIO. Manufacturer Narrative: ADDITIONAL REVIEW OF THE EVENT AND/OR ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DID NOT AND DOES NOT MEET THE REPORTING REQUIREMENTS STIPULATED IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22175011&pc=CAW | 3009096682-2025-00005 | 2025-05-20 04:00 | Death | QINGDAO KINGON MEDICAL SCIENCE AND TECHNOLOGY CO. | 2025-06-10 04:00 | CAW | REACT POC2000 | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: REACT HEALTH RECEIVED A COMPLAINT FROM A DURABLE MEDICAL EQUIPMENT PROVIDER OF A PATIENT PASSING AWAY WHILE USING THE DEVICE. THE DEVICE WAS RETURNED TO REACT HEALTH AND EVALUATED ON (B)(6) 2025, WHERE IT WAS DETERMINED TO BE OPERATING ACCORDING TO THE MANUFACTURER'S SPECIFICATIONS. THE MANUFACTURER HAS NOT TAKEN POSSESSION OF THE DEVICE AS OF THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22175226&pc=PNJ | 9612164-2025-02836 | 2025-05-16 04:00 | Death | MEDTRONIC IRELAND | 2025-06-10 04:00 | PNJ | MICRA¿ VR2 | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Cardiac Perforation | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) DELIVERY SYSTEM (DS) THE PATIENT EXPERIENCED CARDIAC TAMPONADE / PERFORATION AND A PERICARDIOCENTESIS WAS PERFORMED. THE LEADLESS IPG DS WAS ATTEMPTED NOT USED. THE PATIENT PASSED AWAY. Manufacturer Narrative: CORRECTION H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: [MICRA] PRODUCT ID: [MC2VR01-DELSYS] (SERIAL: [(B)(6)]). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22175615&pc=QZK | 2017233-2025-06291 | 2025-05-15 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-10 04:00 | QZK | GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P230023 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT USING A GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS, GORE® EXCLUDER® AAA ENDOPROSTHESIS, GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) IN ZONE 5 WHERE A GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) DEVICES WERE USED AS BRANCH DEVICES IN THE VISCERAL ARTERIES. IT WAS REPORTED PRIOR TO THE PROCEDURE; THE PATIENT'S LEFT RENAL ARTERY WAS ALREADY OCCLUDED. DURING THE PROCEDURE, THERE WAS A DISTAL PERFORATION IN THE RIGHT RENAL DUE TO A 0.035" ROSEN GUIDEWIRE. THE PERFORATION WAS RESOLVED BY PLACEMENT OF A NON-GORE CORONARY COVERED STENT. THE PHYSICIAN HAD RIGHT SIDED ACCESS DUE TO THE LEFT SUBCLAVIAN BEING STENOTIC AND ALREADY HAVING AN UNKNOWN STENT IN PLACE FROM A PREVIOUS SURGERY. THE PHYSICIAN IMPLANTED ALL DEVICES WITHOUT ANY ISSUES. TWO VBX DEVICES WERE IMPLANTED IN THE SMA, THREE VBX DEVICES WERE IMPLANTED IN THE RIGHT RENAL, AND TWO VBX DEVICES WERE IMPLANTED IN THE CELIAC. THE PATIENT TOLERATED THE PROCEDURE. ON (B)(6) 2025, IT WAS REPORTED THE PATIENT HAD A STROKE AND EXPIRED. THE PHYSICIAN STATED DISPLACEMENT OF ATHEROMA (PLAQUE) IN THE ARCH, MOVED TO THE PATIENT'S HEAD AND LEGS CAUSING THE STROKE. Manufacturer Narrative: ACCORDING TO THE GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE NEUROLOGIC DAMAGE, LOCAL OR SYSTEMIC (E.G., STROKE) AND DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: ADDED G3/G4, H1/H2, H6. CORRECTED B3.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22175639&pc=DQY | 3007284313-2025-04023 | 2025-05-15 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-10 04:00 | DQY | GORE® TRI LUMEN CATHETER | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | K180919 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT USING A GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS, GORE® EXCLUDER® AAA ENDOPROSTHESIS, GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) IN ZONE 5 WHERE A GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) DEVICES WERE USED AS BRANCH DEVICES IN THE VISCERAL ARTERIES. IT WAS REPORTED PRIOR TO THE PROCEDURE; THE PATIENT'S LEFT RENAL ARTERY WAS ALREADY OCCLUDED. DURING THE PROCEDURE, THERE WAS A DISTAL PERFORATION IN THE RIGHT RENAL DUE TO A 0.035" ROSEN GUIDEWIRE. THE PERFORATION WAS RESOLVED BY PLACEMENT OF A NON-GORE CORONARY COVERED STENT. THE PHYSICIAN HAD RIGHT SIDED ACCESS DUE TO THE LEFT SUBCLAVIAN BEING STENOTIC AND ALREADY HAVING AN UNKNOWN STENT IN PLACE FROM A PREVIOUS SURGERY. THE PHYSICIAN IMPLANTED ALL DEVICES WITHOUT ANY ISSUES. TWO VBX DEVICES WERE IMPLANTED IN THE SMA, THREE VBX DEVICES WERE IMPLANTED IN THE RIGHT RENAL, AND TWO VBX DEVICES WERE IMPLANTED IN THE CELIAC. THE PATIENT TOLERATED THE PROCEDURE. ON (B)(6) 2025, IT WAS REPORTED THE PATIENT HAD A STROKE AND EXPIRED. THE PHYSICIAN STATED DISPLACEMENT OF ATHEROMA (PLAQUE) IN THE ARCH, MOVED TO THE PATIENT'S HEAD AND LEGS CAUSING THE STROKE. Manufacturer Narrative: CORRECTION: B3. Manufacturer Narrative: ACCORDING TO THE GORE® TRI LUMEN CATHETER INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE STROKE AND DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22175668&pc=OZD | 1220648-2025-29151 | 2025-05-15 04:00 | Death | ABIOMED, INC. | 2025-06-10 04:00 | OZD | IMPELLA CP | Restricted Flow rate | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PUMP WAS PLACED, AND FLOWS WERE LOW. THE 2ND CP WAS PLACED AS REPLACEMENT, AND ADDITIONALLY THE TEAM MADE THE DECISION TO DISCUSS RF FLEX PLACEMENT TO ASSIST WITH THE VENTRICULAR FAILURE. THE FAMILY MADE THE CHOICE TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22175867&pc=NIQ | 9612164-2025-02837 | 2025-05-23 04:00 | Death | MEDTRONIC IRELAND | 2025-06-10 04:00 | NIQ | ONYX FRONTIER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Ventricular Fibrillation; Thrombosis/Thrombus | P160043 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT PERCUTANEOUS CORONARY INTERVENTION (PCI) INVOLVING THE IMPLANTATION OF ONE ONYX FRONTIER CORONARY DRUG-ELUTING STENT AND SUBSEQUENTLY DIED. DURING THE PROCEDURE, IN ADDITION TO THE ONYX FRONTIER DEVICE THAT WAS DEPLOYED,AN ASSIST DEVICE WAS USED AND INTRAVASCULAR LITHOTRIPSY (IVL) WAS ALSO PERFORMED. THE LESION BEING TREATED WAS NON-TORTUOUS AND NON-CALCIFIED WITH 90% STENOSIS LOCATED IN THE LEFT ANTERIOR DESCENDING ARTERY (CASS#8). THE DEVICE WAS INSPECTED PRIOR TO USE AND NEGATIVE PREP WAS PERFORMED WITH NO ISSUES. THE DEVICE DID NOT PASS THROUGH A PREVIOUSLY-DEPLOYED STENT. POST-STENT PLACEMENT IMAGING SHOWED NO ABNORMALITIES AND THE PROCEDURE WAS COMPLETED. THE PATIENT'S ACTIVATED CLOTTING TIME (ACT) WAS 120 OR LESS. APPROXIMATELY FIFTEEN TO TWENTY MINUTES AFTER COMPLETION OF THE PROCEDURE AND EXITING THE CATHETERISATION LABORATORY, THE PATIENT DEVELOPED VENTRICULAR FIBRILLATION LEADING TO CARDIAC ARREST AND DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S CONDITION AT THE START OF THE PROCEDURE WAS DESCRIBED AS NORMAL. DURING THE PROCEDURE, IN ADDITION TO THE ONYX FRONTIER DEVICE THAT WAS DEPLOYED, AN ASSIST INTRA-AORTIC BALLOON PUMP (IABP) DEVICE WAS USED, AND INTRAVASCULAR LITHOTRIPSY (IVL) WAS ALSO PERFORMED. THE IABP ASSIST DEVICE WAS USED BEFORE STATING THE PCI AS IT WAS A WIDE-RANGE TRIAD LESION AND HIGH-RISK PCI, SO THE ASSIST DEVICE WAS USED FOR RISK PREVENTION. IT WAS CLARIFIED THAT IVL WAS NOT USED DURING THE PROCEDURE. POST-STENT PLACEMENT VESSEL FLOW WAS CHECKED BY CONTRAST RADIOGRAPHY WHICH SHOWED NO ABNORMALITIES, AND THE PROCEDURE WAS COMPLETED. THERE WAS NO EVIDENCE OF RESTENOSIS. A THROMBUS WAS ADHERED TO THE GUIDE CATHETER REMOVED FROM THE BODY. THE PHYSICIAN DID NOT ASSESS THAT THE DEATH EVENT WAS DIRECTLY RELATED TO THE USE OF THE ONYX FRONTIER STENT. LOT NUMBER PROVIDED. CORRECTION: THE DIRECT CAUSE OF DEATH IS UNKNOWN BUT RELATED TO THE LAD PCI THERAPY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: ANNEX D CODE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22175868&pc=DYB | 3007284313-2025-04024 | 2025-05-15 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-10 04:00 | DYB | GORE® DRYSEAL FLEX INTRODUCER SHEATH | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | K160254 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT USING A GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS, GORE® EXCLUDER® AAA ENDOPROSTHESIS, GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) IN ZONE 5 WHERE A GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) DEVICES WERE USED AS BRANCH DEVICES IN THE VISCERAL ARTERIES. IT WAS REPORTED PRIOR TO THE PROCEDURE; THE PATIENT'S LEFT RENAL ARTERY WAS ALREADY OCCLUDED. DURING THE PROCEDURE, THERE WAS A DISTAL PERFORATION IN THE RIGHT RENAL DUE TO A 0.035" ROSEN GUIDEWIRE. THE PERFORATION WAS RESOLVED BY PLACEMENT OF A NON-GORE CORONARY COVERED STENT. THE PHYSICIAN HAD RIGHT SIDED ACCESS DUE TO THE LEFT SUBCLAVIAN BEING STENOTIC AND ALREADY HAVING AN UNKNOWN STENT IN PLACE FROM A PREVIOUS SURGERY. THE PHYSICIAN IMPLANTED ALL DEVICES WITHOUT ANY ISSUES. TWO VBX DEVICES WERE IMPLANTED IN THE SMA, THREE VBX DEVICES WERE IMPLANTED IN THE RIGHT RENAL, AND TWO VBX DEVICES WERE IMPLANTED IN THE CELIAC. THE PATIENT TOLERATED THE PROCEDURE. ON (B)(6) 2025, IT WAS REPORTED THE PATIENT HAD A STROKE AND EXPIRED. THE PHYSICIAN STATED DISPLACEMENT OF ATHEROMA (PLAQUE) IN THE ARCH, MOVED TO THE PATIENT'S HEAD AND LEGS CAUSING THE STROKE. Manufacturer Narrative: H6: B22 AND C20 ¿ PRODUCT HISTORY REVIEW COULD NOT BE PERFORMED AS THE SERIAL# IS NOT AVAILABLE. ACCORDING TO THE GORE® DRYSEAL FLEX INTRODUCER SHEATH INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: ADDED. CORRECTION: B3.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22176207&pc=DSQ | 2916596-2025-03591 | 2025-05-12 04:00 | Death | THORATEC CORPORATION | 2025-06-10 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-ORGAN FAILURE. THE OUTCOME WAS RELATED TO A REFRACTORY MULTI-ORGAN FAILURE IN THE COURSE, BUT NOT CONSIDERED DEVICE OR THERAPY RELATED. NO DEVICE RELATED ISSUES WERE MENTIONED PRIOR OUTCOME WHICH COULD LEAD TO THE MULTIORGAN FAILURE. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22176288&pc=OZD | 1220648-2025-29156 | 2025-05-16 04:00 | Death | ABIOMED, INC. | 2025-06-10 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD A PATIENT ON IMPELLA CP ESCALATED TO THE 5.5 PUMP FOR CARE ESCALATION. IT WAS REPORTED THAT WHILE ON IMPELLA 5.5 A DIFFUSE HEAD BLEED WAS NOTED ON A COMPUTED TOMOGRAPHY SCAN. ANTICOAGULATION WAS STOPPED. CARE WAS WITHDRAWN TWO DAYS LATER AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22176381&pc=DSQ | 2916596-2025-03465 | 2025-04-28 04:00 | Death | THORATEC CORPORATION | 2025-06-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Gastrointestinal Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT WAS ADMITTED ON (B)(6) 2025 WITH GASTROINTESTINAL BLEEDING. THEY CONTINUED TO BLEED THROUGHOUT THE HOSPITAL ADMISSION. THEY WERE TRANSFERRED BACK DOWN TO THE INTENSIVE CARE UNIT FOR HYPOTENSION AND ONGOING BLEEDING ON (B)(6) 2025. THIS WAS STATUS POST COILS WITH INTERVENTIONAL RADIOLOGY (IVR). THEY CONTINUED TO REQUIRE TRANSFUSION WITH MASSIVE AMOUNTS OF GASTROINTESTINAL BLEEDING. THEY WERE NOT DEEMED A SUITABLE CANDIDATE FOR ANY MORE INVASIVE PROCEDURES AND BLEEDING WAS ONGOING DESPITE MASS TRANSFUSION. THE PATIENT'S FAMILY DECIDED TO STOP TREATMENT AND MOVE TO COMFORT MEASURES. PRESSORS WERE STOPPED AND THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS DISCONNECTED. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF GASTROINTESTINAL (GI) BLEEDING COULD NOT BE CONCLUSIVELY DETERMINED. BLEEDING IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM}. GI BLEEDING HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22176451&pc=DSQ | 2916596-2025-03472 | 2022-03-08 05:00 | Death | THORATEC CORPORATION | 2025-06-10 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Renal Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT HAD RENAL FAILURE THAT LED UP TO THE DEATH. THE PATIENT WAS DISCHARGED ON HOSPICE AND WISHED TO NO LONGER PROCEED WITH OUTPATIENT HEMODIALYSIS. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE RENAL DYSFUNCTION WAS NOT DETERMINED TO BE RELATED TO OR CAUSED BY THE DEVICE OR DEVICE THERAPY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RENAL DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A4: PATIENT WEIGHT WAS NOT PROVIDED. SECTION E1: REPORTER EMAIL WAS NOT AVAILABLE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22176707&pc=NPT | 2015691-2025-04722 | 2025-05-19 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-10 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM A FIELD CLINICAL SPECIALIST THAT A PATIENT UNDERWENT TMVR IN MITRAL ANNULAR CALCIFICATION (MAC) WITH LAMPOON. AS REPORTED, THE 29MM S3UR WAS PREPPED PER IFU AND SUCCESSFULLY DEPLOYED IN MAC. AFTER DEPLOYMENT, LARGE PERICARDIAL EFFUSION WAS NOTED. THE PERICARDIAL EFFUSION WAS DRAINED. THERE WAS NO DIFFICULTY TRACKING OR WITHDRAWING THE DELIVERY SYSTEM THROUGH THE PATIENT'S ANATOMY. THE PATIENT WAS ALIVE AND TRANSPORTED TO ICU AFTER THE END OF THE CASE. THE PATIENT WAS SINCE BROUGHT BACK TO THE CATH LAB FOR REOCCURRING PERICARDIAL EFFUSION. THE TEAM WAS UNABLE TO STOP THE BLEEDING AND THE PATIENT PASSED AWAY. THE PATIENT WAS NOT A CANDIDATE FOR OPEN HEART SURGERY. THE INITIAL REPORTER DID NOT ALLEGE THAT ANY OF THE EW DEVICES WERE DEFICIENT IN SOME WAY. THE CAUSE OF THE EFFUSION WAS THOUGHT TO BE A POSSIBLE WIRE PERF IN THE LV. HOWEVER, IT WAS UNCLEAR WHETHER A RUPTURE OR WIRE PERF. THE DEGREE OF TORTUOSITY WAS MODERATE AND THE DEGREE OF ANNULAR CALCIFICATION WAS SEVERE. Manufacturer Narrative: THE DEVICE WAS USED IN OFF-LABEL IMPLANTATION IN THE NATIVE MITRAL POSITION. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO NATIVE MITRAL PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THAT PATIENT RELATED FACTORS (TMVR IN MITRAL ANNULAR CALCIFICATION (MAC)) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: SUPPLEMENTAL REPORT SUBMITTED TO INCLUDE ADDITIONAL INFORMATION RECEIVED THROUGH FOLLOW-UP. UPDATED A2 AND B7. THE CONCLUSION REMAINS THE SAME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22176822&pc=DSQ | 2916596-2025-03460 | 2025-05-17 04:00 | Death | THORATEC CORPORATION | 2025-06-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT LEFT THE HOSPITAL AGAINST MEDICAL ADVICE ON (B)(6) 2025. THE PATIENT'S FAMILY MEMBER CALLED ON (B)(6) 2025 AND INFORMED THE OFFICE THAT THE PATIENT WAS TO GO TO HOSPICE CARE. THE PATIENT PASSED AWAY AT HOME ON (B)(6) 2025 SURROUNDED BY FAMILY. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22176827&pc=MHX | 9610816-2025-000475 | 2025-05-09 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-06-10 04:00 | MHX | INTELLIVUE PATIENT MONITOR MX750 | Defective Alarm | Respiratory Failure; Insufficient Information | K210906 |  | 1 | Event Description: IT WAS REPORTED THE MONITOR DID NOT ALARM FOR SPO2 AND THE PATIENT PASSED AWAY. IN ADDITION, IT WAS NOTED THERE WERE NO ERRORS ON THE MONITOR OR CENTRAL STATION. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. E1: REPORTING INSTITUTION PHONE #: (B)(6). Manufacturer Narrative: A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) REVIEWED PROVIDED ALARM LOGS AND AUDIT DATA FROM THE PHILIPS PATIENT INFORMATION CENTER IX (PIC IX) THAT WAS CONNECTED WITH THE MX750 MONITOR. THE PSE FOUND A LOT OF SPO2 GENERATED ALARMS IN THE LOGS FROM THE PIC IX. THE PSE NOTED THAT THERE WERE DESATURATION (DESAT) ALARMS AROUND 21:35, AND THEN AT 21:36:17 THE SPO2 ALARMS WERE SHUT OFF (FROM ON TO OFF). FOLLOWING THE TIMELINE, THE SPO2 WAS ALREADY BETWEEN 86 AND 80%, AND RELEASED BEFORE BEING SHUT OFF DESAT ALARMS (RED ALARM DESAT 82 < 84). THAT WAS CHANGED AT 22:10:29 SPO2; ALARMS FROM OFF TO ON, WHEN THE SPO2 WAS ALREADY TOO LOW BETWEEN 27% AND 23%, GENERATING AT 22:11:44 AND 22:12:37 DESAT 0<84 ALARMS (MEANING VERY CRITICAL LOW LEVEL OF SPO2); AFTER THAT, THE TIMELINE SHOWS LOWEST LEVEL SPO2 AND THE MONITOR GENERATED A LOT OF SPO2, NO PULSE, AND SPO2, AND "NOISY SIGNAL" ALARMS, AS WELL AS SPO2 SHUT FROM ON TO OFF, FOLLOWED BY AN SPO2 ALARMS FROM ON TO OFF. FURTHER INFORMATION WAS PROVIDED BY A PHILIPS REMOTE SERVICE ENGINEER (RSE) INDICATING THAT SPO2 ALARMS WERE TURNED OFF, AND PATIENT CARE GIVEN TO SOMEONE ELSE FOR BREAK RELIEF. THE PATIENT'S RESPIRATORY HEALTH DECLINED, AND ATTENTION DRAWN WHEN A RED ALARM FOR ARTERIAL BLOOD PRESURE (ABP) OCCURRED. THE PATIENT WAS DO NOT ATTEMPT RESUSCITATION (DNAR), BUT PATIENT WAS STABLE WAITING FOR A WARD BED. DUE TO LACK OF ALARM, HELP WAS NOT SOUGHT AT TIME OF INITIAL RESPIRATORY DECLINE. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS THE USERS TURNING OFF THE SPO2 ALARMS. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22176828&pc=NVY | 2649622-2025-15654 | 2025-05-09 04:00 | Death | MPRI | 2025-06-10 04:00 | NVY | EPSILA EV¿ MRI SURESCAN¿ | Signal Artifact/Noise; Over-Sensing; Defibrillation/Stimulation Problem; Inappropriate/Inadequate Shock/Stimulation; Under-Sensing | Ventricular Fibrillation; Brain Injury | P220012 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT RECEIVED INAPPROPRIATE HIGH-VOLTAGE THERAPY DUE TO P-WAVE OVERSENSING (PWOS) AND NOISE ON THE EXTRAVASCULAR (EV) LEAD DESPITE PRIOR REPROGRAMMING, AND UNDERSENSING WAS NOTED ON STORED ELECTROGRAMS (EGM) POST-DETECTION. ADDITIONAL PROGRAMMING OPTIONS WERE PROVIDED, HOWEVER THE PATIENT DID NOT WISH TO BE SEEN AND ONLY WANTED ALERT TONES TURNED OFF. IT WAS LATER REPORTED THAT DUE TO ADDITIONAL OVERSENSING OF NOISE WHILE THE PATIENT WAS AT THE GYM, AN INAPPROPRIATE SHOCK WAS DELIVERED WHICH INDUCED VENTRICULAR FIBRILLATION (VF). TWO ADDITIONAL HIGH-VOLTAGE SHOCKS WERE DELIVERED WHICH WERE UNSUCCESSFUL AT TERMINATING THE RHYTHM, AND SUBSEQUENT UNDERSENSING WAS OBSERVED WHICH ALLOWED THE RHYTHM TO BE DEVICE-CLASSIFIED AS TERMINATED WHILE THE RHYTHM CONTINUED WITHOUT ADDITIONAL THERAPY. THE PATIENT WAS DIAGNOSED WITH A NOXIOUS BRAIN INJURY AND WAS INTUBATED AND HOSPITALIZED. LIFE SUPPORT WAS WITHDRAWN AT THE REQUEST OF THE PATIENT'S REPRESENTATIVES, AND THE PATIENT ULTIMATELY PASSED AWAY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED OVERSENSING. ANALYSIS OF THE DEVICE MEMORY INDICATED RIGHT VENTRICULAR UNDERSENSING. ANALYSIS OF THE DEVICE MEMORY OBSERVED NOISE ON THE ELECTROGRAM WAVEFORMS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22176843&pc=NPT | 2025587-2025-04149 | 2025-03-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-10 04:00 | NPT | EVOLUT FX DCS | Device Dislodged or Dislocated; Insufficient Information; Patient Device Interaction Problem |  | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY WAS PERFORMED AS STANDARD FOR THE IMPLANTING PHYSICIAN. FOLLOWING DEPLOYMENT, THE FIRST VALVE DISLODGED. A SECOND VALVE WAS LOADED INTOA DELIVERY CATHETER SYSTEM; HOWEVER, THE PATIENT BEGAN TO DECOMPENSATE BEFORE DELIVERY COULD BEGIN. AN ASCENDING AORTIC DISSECTION WAS IDENTIFIED. INTERVENTIONAL RADIOLOGY AND VASCULAR SURGERY WERE CALLED. AN ATTEMPT WAS MADE TO SNARE THE DISLODGED VALVE AND PLACE A STENT. DESPITE THESE EFFORTS, THE PATIENT DIED. THE MEDTRONIC DEVICE WAS NOTED TO HAVE CONTRIBUTED TO THE ADVERSE OUTCOME. OTHER CONTRIBUTING FACTORS INCLUDED THE PATIENT'S DILATED ASCENDING AORTA. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN; THE DELIVERY CATHETER SYSTEM DID NOT CONTRIBUTE TO THE PATIENT'S DEATH. THE PATIENT'S UNDERLYING CONDITION OF SEVERE AORTIC STENOSIS WHICH REQUIRED THE TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED THAT THE SECOND VALVE WAS NEVER DEPLOYED, HOWEVER, THE SECOND SYSTEM WAS ATTEMPTED IN THE PATIENT.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE AORTIC DISSECTION OCCURRED WHEN THE SECOND SYSTEM WAS ADVANCING OVER THE THE AORTIC ARCH, MOVING THE FIRST VALVE THAT HAD DISLODGED. PER THE PHYSICIAN, THE SECOND DCS DID NOT CAUSE OR CONTRIBUTE TO THE DISSECTION. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID {EVFXPLUS-29}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {0195-HERAT VALVES}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22176860&pc=NVN | 2182208-2025-02808 | 2025-01-01 05:00 | Death | MEDTRONIC, INC. | 2025-06-10 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction; Heart Failure/Congestive Heart Failure | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE FEASIBILITY AND SAFETY OF LEFT BUNDLE BRANCH AREA PACING (LBBAP) IN PATIENTS WITH A HISTORY OF OPEN-HEART SURGERY (OHS). PATIENTS WERE STUDIED IN GROUPS OF PRIOR OHS AND NON-OHS. THE AUTHORS DESCRIBED PATIENT DEATHS IN BOTH GROUPS; THE CAUSE OF DEATH WAS END-STAGE CHRONIC HEART FAILURE IN ONE PATIENT IN THE OHS GROUP. THREE DEATHS OCCURRED IN THE NON-OHS GROUP: ONE DUE TO END-STAGE HEART FAILURE, ONE DUE TO SUSPECTED ACUTE MYOCARDIAL INFARCTION, AND ONE UNKNOWN CAUSE. THERE IS NO ALLEGATION OF LEAD-DEATH RELATEDNESS INDICATED IN THE ARTICLE. THERE WAS ONE LEAD DISLODGEMENT IN THE NON-OHS GROUP WHICH REQUIRED REIMPLANTATION. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: COMPARATIVE ANALYSIS OF LEFT BUNDLE BRANCH AREA PACING IN PATIENTS WITH AND WITHOUT A HISTORY OF OPEN-HEART SURGERY. JOURNAL OF ARRHYTHMIA. 2025; 41:E70010. DOI: 10.1002/JOA3.70010. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22177072&pc=DSQ | 2916596-2025-03484 | 2025-05-15 04:00 | Death | THORATEC CORPORATION | 2025-06-10 04:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO END STAGE HEART FAILURE. THE PATIENT WAS ON HOSPICE. ADDITIONAL DETAILS LEADING UP TO DEATH WERE UNKNOWN. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. NO DEVICE WOULD BE RETURNED. Manufacturer Narrative: SECTION D4: EXPIRATION DATE AND MANUFACTURE DATE (H4) WILL BE PROVIDED UPON INVESTIGATION. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION A2: PATIENT AGE CORRECTED. SECTION G2: REPORT SOURCE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22177134&pc=NKM | 2135147-2025-03069 | 2025-05-15 04:00 | Death | ABBOTT MEDICAL | 2025-06-10 04:00 | NKM | MITRACLIP G4 CLIP DELIVERY SYSTEM (MDR) | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A PATIENT PRESENTED WITH GRADE 3-4 DEGENERATIVE MITRAL REGURGITATION (MR) AND A POSTERIOR 2 LEAFLET FLAIL FOR A MITRACLIP PROCEDURE. TWO MITRACLIPS WERE IMPLANTED SUCCESSFULLY. THE FINAL MR WAS GRADE 1-2. THERE WAS NO CLINICALLY SIGNIFICANT DELAY REPORTED. FOUR HOURS AFTER THE PROCEDURE, THE PATIENT PRESENTED WITH THE INABILITY TO MOVE THE LEFT ARM AND A STROKE CODE WAS CALLED. A COMPUTERIZED TOMOGRAPHY SCAN (CT) AND AN MAGNETIC RESONANCE IMAGING (MRI) WERE NEGATIVE FOR STROKE. PER THE PHYSICIAN, IT WAS THEN CONSIDERED TO BE A RESULT OF THE ANESTHESIA ACCESS DURING THE PROCEDURE AND A STROKE WAS NO LONGER THE DIAGNOSIS. HOWEVER, THE PATIENT'S CONDITION CONTINUED TO WORSEN AND HIS LEFT LEG WAS NO LONGER MOBILE EITHER. THE PHYSICIAN AND STAFF DO NOT KNOW THE ORIGIN OF THE POTENTIAL EMBOLI-DRIVEN STROKE, BUT THEY DO NOT BELIEVE IT WAS RELATED TO ANY OF THE ABBOTT DEVICES USED DURING THE PROCEDURE. ON (B)(6) 2025, THE PATIENT'S STATUS DECLINED AND WAS PRONOUNCED DEAD. THE CAUSE OF DEATH WAS A BRAIN HEMORRHAGE. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS NOT DUE TO THE MITRACLIP DEVICES. THE CAUSE FOR THE STROKE IS INCONCLUSIVE.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A PATIENT PRESENTED WITH GRADE 3-4 DEGENERATIVE MITRAL REGURGITATION (MR) AND A POSTERIOR 2 LEAFLET FLAIL FOR A MITRACLIP PROCEDURE. TWO MITRACLIPS WERE IMPLANTED SUCCESSFULLY. THE FINAL MR WAS GRADE 1-2. THERE WAS NO CLINICALLY SIGNIFICANT DELAY REPORTED. FOUR HOURS AFTER THE PROCEDURE THE PATIENT PRESENTED WITH THE INABILITY TO MOVE THE LEFT ARM AND A STROKE CODE WAS CALLED. A COMPUTERIZED TOMOGRAPHY SCAN (CT) AND AN MAGNETIC RESONANCE IMAGING (MRI) WERE NEGATIVE FOR STROKE. PER THE PHYSICIAN, IT WAS THEN CONSIDERED TO BE A RESULT OF THE ANESTHESIA ACCESS DURING THE PROCEDURE AND A STROKE WAS NO LONGER THE DIAGNOSIS. HOWEVER, THE PATIENT'S CONDITION CONTINUED TO WORSEN AND HIS LEFT LEG WAS NO LONGER MOBILE EITHER. THE PHYSICIAN AND STAFF DO NOT KNOW THE ORIGIN OF THE POTENTIAL EMBOLI-DRIVEN STROKE, BUT THEY DO NOT BELIEVE IT WAS RELATED TO ANY OF THE ABBOTT DEVICES USED DURING THE PROCEDURE. ON (B)(6) 2025, THE PATIENT'S STATUS DECLINED AND WAS PRONOUNCED DEAD. THE CAUSE OF DEATH WAS A BRAIN HEMORRHAGE. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS NOT DUE TO THE MITRACLIP DEVICES. THE CAUSE FOR THE STROKE IS INCONCLUSIVE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THE REPORTED CEREBROVASCULAR ACCIDENT OR DEATH. BASED ON AVAILABLE INFORMATION, THE REPORTED CEREBROVASCULAR ACCIDENT APPEARS TO BE RELATED TO THE ANESTHESIA ACCESS DURING THE PROCEDURE. THE REPORTED DEATH APPEARS TO BE RELATED TO A BRAIN HEMORRHAGE UNRELATED TO THE MITRACLIP DEVICES. THE REPORTED PATIENT EFFECTS OF CEREBROVASCULAR ACCIDENT AND DEATH, AS LISTED IN THE MITRACLIP SYSTEM INSTRUCTIONS FOR USE, ARE KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED HOSPITALIZATION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22177140&pc=NKM | 2135147-2025-03070 | 2025-05-15 04:00 | Death | ABBOTT MEDICAL | 2025-06-10 04:00 | NKM | MITRACLIP G4 CLIP DELIVERY SYSTEM (MDR) | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A PATIENT PRESENTED WITH GRADE 3-4 DEGENERATIVE MITRAL REGURGITATION (MR) AND A POSTERIOR 2 LEAFLET FLAIL FOR A MITRACLIP PROCEDURE. TWO MITRACLIPS WERE IMPLANTED SUCCESSFULLY. THE FINAL MR WAS GRADE 1-2. THERE WAS NO CLINICALLY SIGNIFICANT DELAY REPORTED. FOUR HOURS AFTER THE PROCEDURE THE PATIENT PRESENTED WITH THE INABILITY TO MOVE THE LEFT ARM AND A STROKE CODE WAS CALLED. A COMPUTERIZED TOMOGRAPHY SCAN (CT) AND AN MAGNETIC RESONANCE IMAGING (MRI) WERE NEGATIVE FOR STROKE. PER THE PHYSICIAN, IT WAS THEN CONSIDERED TO BE A RESULT OF THE ANESTHESIA ACCESS DURING THE PROCEDURE AND A STROKE WAS NO LONGER THE DIAGNOSIS. HOWEVER, THE PATIENT'S CONDITION CONTINUED TO WORSEN AND HIS LEFT LEG WAS NO LONGER MOBILE EITHER. THE PHYSICIAN AND STAFF DO NOT KNOW THE ORIGIN OF THE POTENTIAL EMBOLI-DRIVEN STROKE, BUT THEY DO NOT BELIEVE IT WAS RELATED TO ANY OF THE ABBOTT DEVICES USED DURING THE PROCEDURE. ON (B)(6) 2025, THE PATIENT'S STATUS DECLINED AND WAS PRONOUNCED DEAD. THE CAUSE OF DEATH WAS A BRAIN HEMORRHAGE. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS NOT DUE TO THE MITRACLIP DEVICES. THE CAUSE FOR THE STROKE IS INCONCLUSIVE.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A PATIENT PRESENTED WITH GRADE 3-4 DEGENERATIVE MITRAL REGURGITATION (MR) AND A POSTERIOR 2 LEAFLET FLAIL FOR A MITRACLIP PROCEDURE. TWO MITRACLIPS WERE IMPLANTED SUCCESSFULLY. THE FINAL MR WAS GRADE 1-2. THERE WAS NO CLINICALLY SIGNIFICANT DELAY REPORTED. FOUR HOURS AFTER THE PROCEDURE THE PATIENT PRESENTED WITH THE INABILITY TO MOVE THE LEFT ARM AND A STROKE CODE WAS CALLED. A COMPUTERIZED TOMOGRAPHY SCAN (CT) AND AN MAGNETIC RESONANCE IMAGING (MRI) WERE NEGATIVE FOR STROKE. PER THE PHYSICIAN, IT WAS THEN CONSIDERED TO BE A RESULT OF THE ANESTHESIA ACCESS DURING THE PROCEDURE AND A STROKE WAS NO LONGER THE DIAGNOSIS. HOWEVER, THE PATIENT'S CONDITION CONTINUED TO WORSEN AND HIS LEFT LEG WAS NO LONGER MOBILE EITHER. THE PHYSICIAN AND STAFF DO NOT KNOW THE ORIGIN OF THE POTENTIAL EMBOLI-DRIVEN STROKE, BUT THEY DO NOT BELIEVE IT WAS RELATED TO ANY OF THE ABBOTT DEVICES USED DURING THE PROCEDURE. ON (B)(6) 2025, THE PATIENT'S STATUS DECLINED AND WAS PRONOUNCED DEAD. THE CAUSE OF DEATH WAS A BRAIN HEMMEOAGE. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS NOT DUE TO THE MITRACLIP DEVICES. THE CAUSE FOR THE STROKE IS INCONCLUSIVE. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THE REPORTED CEREBROVASCULAR ACCIDENT OR DEATH. BASED ON AVAILABLE INFORMATION, THE REPORTED CEREBROVASCULAR ACCIDENT APPEARS TO BE RELATED TO THE ANESTHESIA ACCESS DURING THE PROCEDURE. THE REPORTED DEATH APPEARS TO BE RELATED TO A BRAIN HEMORRHAGE UNRELATED TO THE MITRACLIP DEVICES. THE REPORTED PATIENT EFFECTS OF CEREBROVASCULAR ACCIDENT AND DEATH, AS LISTED IN THE MITRACLIP SYSTEM INSTRUCTIONS FOR USE, ARE KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED HOSPITALIZATION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22177259&pc=DSQ | 2916596-2025-03500 | 2025-05-16 04:00 | Death | THORATEC CORPORATION | 2025-06-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Gastrointestinal Hemorrhage; Liver Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD BEEN ADMITTED TO THE HOSPITAL, PRIOR TO BEING DISCHARGED ON HOSPICE, WITH GASTROINTESTINAL (GI) BLEED AND LIVER FAILURE. THE PATIENT PASSED AWAY ON HOSPICE ON (B)(6) 2025. THE VENTRICULAR ASSIST DEVICE (VAD) HAD NO PART IN THE CAUSE OF THE DEATH.Event Description: ADDITIONAL INFORMATION RECEIVED STATED THAT THE PATIENT HAD LONG STANDING DRIVELINE INFECTION WITH MULTIPLE DEBRIDEMENTS AND WAS ON CHRONIC INTRAVENOUS ANTIBIOTICS. MOREOVER, THE PATIENT HAD RETURNED HOME AFTER BEING PUT ON HOSPICE AND PASSED AWAY QUIETLY WITH FAMILY AT BEDSIDE. HOME HEALTH WITHDREW LEFT VENTRICULAR ASSIST DEVICE (LVAD) SUPPORT AND REPORTED TIME OF DEATH ON (B)(6) 2025 AT 10:18 PM. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING, HEPATIC DYSFUNCTION, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22177631&pc=LWS | 2649622-2025-15668 | 2025-05-18 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-06-10 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Under-Sensing | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING BEFORE DETECTION. THE RV LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY THE FOLLOWING DAY. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS IN A VENTRICULAR TACHYCARDIA (VT) STORM PRIOR TO THEIR DEATH. THE CAUSE OF DEATH WAS DETERMINED TO BE CARDIOGENIC SHOCK AND ORGAN FAILURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: DTPC2QQ CRT-D, IMPLANTED: (B)(6) 2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22177802&pc=NPT | 2015691-2025-04733 | 2025-05-20 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-10 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Device Dislodged or Dislocated | Foreign Body In Patient | P140031 |  | 1 | Event Description: AS REPORTED BY OUR EDWARDS LIFESCIENCES AFFILIATE IN GERMANY, REGARDING A 26MM SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE (THV) WAS IMPLANTED IN THE AORTIC POSITION VIA A RIGHT TRANSFEMORAL APPROACH IN A PATIENT WITH A PRE-EXISTING NON-EDWARDS BIOPROSTHETIC VALVE, DURING THE PROCEDURE, COMPLICATIONS AROSE WHEN THE CRIMPED THV BECAME LODGED BETWEEN THE LEAFLETS AND THE UPPER STENT-FRAME RING OF THE NON-EDWARDS DURING POSITIONING, PRIOR TO ANY ATTEMPT AT DEPLOYMENT. SEVERAL MANEUVERS WERE ATTEMPTED AND AT THE END, 1 ML WAS INJECTED INTO THE COMMANDER DELIVERY SYSTEM IN AN EFFORT TO SLIGHTLY INFLATE THE BALLOON AND FACILITATE RELEASE OF THE THV. DURING THE PULL-BACK MANEUVER TO DISENGAGE THE THV, THE VALVE APPEARED PARTIALLY DEFORMED, RAISING CONCERNS ABOUT POTENTIAL DAMAGE. SUBSEQUENTLY, THE THV DISLODGED FROM THE COMMANDER END, AS CONFIRMED BY FLUOROSCOPIC IMAGING. DUE TO THE VALVE DISLODGEMENT, SURGICAL INTERVENTION WAS REQUIRED TO REMOVE THE DEVICE. APPROXIMATELY 15 MINUTES AFTER THE PROCEDURE CONCLUDED, THE PATIENT WENT INTO SHOCK, LIKELY DUE TO PROLONGED LOWER BODY ISCHEMIA DURING VASCULAR ACCESS AND PERFUSION INTERRUPTION. DESPITE RESUSCITATIVE EFFORTS, THE PATIENT COULD NOT RECOVER FROM THE SHOCK AND PASSED AWAY SHORTLY THEREAFTER. Manufacturer Narrative: THIS IS ONE OF TWO REPORTS BEING SUBMITTED FOR THIS CASE. THE INVESTIGATION IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR ADDITIONAL INFORMATION FROM A PRODUCT INVESTIGATION. THE COMPLAINT FOR THV DISLODGED OFF BALLOON AND SUBSEQUENT PATIENT DEATH WAS CONFIRMED BASED ON THE EVALUATION OF THE PROVIDED IMAGERY. REVIEW OF THE DHR AND LOT HISTORY DID NOT PROVIDE ANY INDICATION THAT A MANUFACTURING NON-CONFORMANCE WOULD HAVE CONTRIBUTED TO THE COMPLAINT EVENT. A REVIEW OF IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. FURTHERMORE, NO ABNORMALITIES WERE REPORTED DURING DEVICE UNPACKING OR PREPARATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. PER REVIEW, NO EVIDENCE OF PRODUCT NON-CONFORMANCE'S OR LABELING/IFU INADEQUACIES WERE IDENTIFIED IN THE EVALUATION. AS PER DESCRIPTION EVENT, ''1 ML WAS INJECTED INTO THE COMMANDER DELIVERY SYSTEM IN AN EFFORT TO SLIGHTLY INFLATE THE BALLOON AND FACILITATE RELEASE OF THE THV''. PROVIDED IMAGERY SHOWS THE THV FRAME PARTIALLY AND ASYMMETRICALLY EXPANDED AND FLARED AT INFLOW SIDE. ALSO, VALVE IS OBSERVED DISLODGED FROM DELIVERY SYSTEM IN THE ABDOMINAL AORTA. NO DELIVERY SYSTEM NOR GUIDEWIRE IS VISIBLE ACROSS THE THV. IN THIS CASE, THE BALLOON WAS SLIGHTLY INFLATED (1ML) TO ATTEMPT TO WITHDRAW THE ENTIRE SYSTEM FROM THE PRE-EXISTING IMPLANTED VALVE. DUE TO INCREASE OF VALVE PROFILE, WHEN THE COMMANDER DELIVERY SYSTEM WAS ATTEMPTED TO WITHDRAWAL, THE PARTIALLY EXPANDED VALVE MAY HAVE DISENGAGED FROM THE SYSTEM, LEADING TO DISLODGE INTO THE VASCULATURE, AS REPORTED. ADDITIONALLY, DEVICE MANIPULATION APPLIED TO OVERCOME THE RESISTANCE FELT DURING WITHDRAWAL CAN ALSO CONTRIBUTE TO THE THV DISLODGMENT. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PROCEDURAL FACTORS (WITHDRAWAL OF PARTIALLY EXPANDED VALVE, EXCESSIVE DEVICE MANIPULATION) LIKELY CONTRIBUTED TO THE EVENT. HOWEVER A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22177867&pc=MKJ | 1220908-2025-01991 | 2025-05-17 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-06-10 04:00 | MKJ | R SERIES | Failure to Power Up | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE WOULD NOT POWER UP. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED.Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT AN 83-YEAR-OLD FEMALE PATIENT, THE DEVICE WOULD NOT POWER UP. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED.Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A 71-YEAR-OLD FEMALE PATIENT, THE DEVICE WOULD NOT POWER UP. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: JUSTIFICATION FOR NO UDI, THIS DEVICE WAS MANUFACTURED BEFORE UDI REQUIREMENTS. ZOLL MEDICAL CORPORATION HAS RECEIVED THE PRODUCT AND WILL BE PROVIDING A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS COMPLETED. Manufacturer Narrative: THIS SUPPLEMENTAL MEDWATCH REPORT IS REPORTING THE EVALUATION OF THE DEVICE AND CORRECTING INFORMATION SUBMITTED ON THE INITIAL MEDWATCH REPORT. PLEASE REFERENCE SECTIONS A2, A3A, B1, AND B5 UPDATED. THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. THE CUSTOMER'S REPORT WAS CONFIRMED AND ATTRIBUTED TO A LOOSE FLEX CABLE CONNECTION AT THE CONNECTOR ON THE CONTROL BOARD. DURING THE INVESTIGATION OF THE DEVICE TO DETERMINE THE ROOT CAUSE, THE TECHNICIAN OBSERVED DAMAGE TO THE FRONT PANEL OVERMOLD, THE SHOCK BUTTON LED PADS WERE LIFTED FROM THE CONTROL BOARD, AND A DAMAGED PACE/DEFIB (PD) ENGINE. THE DAMAGE IS NOT CONSISTENT WITH WHAT WE WOULD EXPECT TO BE TYPICAL HANDLING OR PREVENTIVE MAINTENANCE ACTIVITY. THE OBSERVED DAMAGE IS CONSISTENT WITH IMPACT DAMAGE. THE R SERIES OPERATOR'S GUIDE (PN: 9650-0912-01) (REV: YA) DETAILS ROUTINE PROCEDURES FOR THE R SERIES DEVICE INCLUDING DAILY VISUAL INSPECTIONS. IT STATES, "ENSURE THAT THE UNIT IS CLEAN (WITH NO FLUID SPILLS) AND FREE OF VISIBLE DAMAGE." ADDITIONALLY, IT STATES, "ALWAYS INSPECT THE UNIT FOR DAMAGE IF IT HAS BEEN DROPPED." THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS FOR REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: THIS SUPPLEMENTAL MEDWATCH REPORT IS CORRECTING INFORMATION SUBMITTED ON THE INITIAL MEDWATCH REPORT. PLEASE REFERENCE SECTIONS A2, A4, A5A, A5B, B5, AND H10. THIS CORRECTION IS IN RESPONSE TO A USER SUBMITTED MEDWATCH REPORT. THE END USER HAS CONFIRMED THE CORRECT PATIENT DEMOGRAPHICS FOR THIS REPORTED CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22178182&pc=OZD | 1220648-2025-29173 | 2025-05-16 04:00 | Death | ABIOMED, INC. | 2025-06-10 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT HAD LIMB ISCHEMIA. BILATERAL UPPER AND LOWER EXTREMITIES WERE NOTED TO BE EXTREMELY MOTTLED. ORDER FOR BILATERAL ARTERIAL DUPLEXES OF UPPER AND LOWER EXTREMITIES PLACED. THE PATIENT CONTINUED TO DECLINE, AND THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST SYSTEM INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22178354&pc=DSQ | 2916596-2025-03435 | 2025-05-15 04:00 | Death | THORATEC CORPORATION | 2025-06-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Chest Pain; Cardiovascular Insufficiency | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED ON (B)(6) 2025 AROUND 23:30 WITH CHEST PAIN. THE PATIENT BECAME MORE SYMPTOMATIC AND HAD AN ¿ASYSTOLIC EVENT¿ AROUND 02:00 ON (B)(6) 2025. THE PATIENT WAS TAKEN TO THE CATHETERIZATION LAB FOR ASSESSMENT AND TREATMENT AROUND 04:00. DURING CATHETERIZATION, CODE BLUE WAS CALLED AT 04:45. THE PATIENT HAD A CARDIAC ARREST. THEY WERE CANNULATED FOR EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WITH A START TIME OF 05:50. LOW FLOW ALARMS WERE CONSTANT AFTER 04:15. EXTENDED SILENCE WAS PLACED ON THE PATIENT PER CARDIOTHORACIC AND VASCULAR SURGERY (CTVS). LOG FILES WERE SENT FOR REVIEW. THE FILE CAPTURED LOW FLOW EVENTS ON (B)(6) 2025. THE FILE ALSO SHOWED POWER DECREASING WITH PULSATILITY INDEX (PI) BECOMING MORE STATIC. THIS COMBINATION OF TRENDING COULD BE CAUSED BY SOME TYPE OF OBSTRUCTION. THE PATIENT PASSED AWAY DUE TO CARDIOVASCULAR COLLAPSE ON (B)(6) 2025. THE WERE LOW FLOWS NOTED PRIOR TO CARDIOPULMONARY RESUSCITATION (CPR) INITIATION. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: SECTION G6 CORRECTION. SECTION H6 CORRECTION: HEALTH EFFECT - IMPACT CODE 4642 - ADDITIONAL DEVICE REQUIRED ADDED. SECTION H6 CORRECTION: MEDICAL DEVICE PROBLEM CODE 2993 - ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM ADDED. MANUFACTURER'S INVESTIGATION CONCLUSION: REVIEW OF THE SUBMITTED LOG FILE CONFIRMED THE REPORTED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THE EVENT COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2025 THROUGH (B)(6) 2025. LOW FLOW HAZARD ALARMS WERE CAPTURED ON (B)(6) 2025. NO OTHER NOTABLE EVENTS OR ALARMS ASSOCIATED WITH THE PUMP WERE CAPTURED, AND THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE SET SPEED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. ADDITIONALLY, IT ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS AND EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, PROVIDES INFORMATION ON ALL SYSTEM ALARM CONDITIONS AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22178361&pc=DSQ | 2916596-2025-03473 | 2025-02-08 05:00 | Death | THORATEC CORPORATION | 2025-06-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED BY THE CUSTOMER. ADDITIONAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: E3: OCCUPATION: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22178611&pc=DSQ | 2916596-2025-03678 | 2025-05-17 04:00 | Death | THORATEC CORPORATION | 2025-06-10 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO PULSELESS ELECTRICAL ACTIVITY (PEA) ARREST. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: THE PEA ETIOLOGY WAS UNKNOWN, AND IT WAS AN UNEXPECTED DEATH. THE DEVICE WAS OPERATING NORMALLY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22178619&pc=OZD | 1220648-2025-29176 | 2025-05-16 04:00 | Death | ABIOMED, INC. | 2025-06-10 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN POSTCARDIOTOMY CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP AND RP FLEX DEVICES FOR MECHANICAL CIRCULATORY BIPELLA SUPPORT AND DEVELOPED HEMOLYSIS REQUIRING INTERVENTION. THE PATIENT WOULD SUBSEQUENTLY EXPIRE. THE PATIENT CAME OUT OF THE OPERATING ROOM POST CORONARY ARTERY BYPASS GRAFT SURGERY ON TWO TO THREE VASOPRESSORS/INOTROPES. OVER THE NEXT TWO HOURS, THE PATIENT CONTINUED TO DECOMPENSATE REQUIRING INCREASING INOTROPIC SUPPORT WITH WORSENING BLOOD GASES. THE DECISION MADE TO BRING PATIENT TO CATHETERIZATION LAB FOR IMPELLA PLACEMENT IN THE SETTING OF WORSENING POST-CARDIOTOMY CARDIOGENIC SHOCK. THE IMPELLA CP DEVICE WAS IMPLANTED AND FOLLOWED BY THE IMPELLA RP FLEX APPROXIMATELY 45 MINUTES LATER PLACED IN THE SETTING OF CARDIOPULMONARY RESUSCITATION. THE PATIENT WAS STABILIZED AND TRANSPORTED TO UNIT IN CRITICAL CONDITION FOLLOWING 30MIN OF ADVANCED CARDIAC LIFE SUPPORT. HOWEVER, POST IMPELLAS IMPLANT THE PATIENT DEVELOPED HEMOLYSIS. A BEDSIDE ECHOCARDIOGRAM WAS PERFORMED TO TRY TO DETERMINE SOURCE OF HEMOLYSIS. THE IMPELLA CP WAS NOTED TO BE FREE FLOATING MID LEFT VENTRICLE AND FREE OF ALL STRUCTURES. THE PATIENT RECEIVED MULTIPLE BLOOD PRODUCTS: FRESH FROZEN PLASMA, AND RED BLOOD CELLS. HOWEVER APPROXIMATELY SIX HOURS LATER, THE PATIENT EXPIRED. Manufacturer Narrative: COMPLAINT DEVICE NOT RETURNED FOR ANALYZING. DATA LOG REVIEWED: NO MC SPIKE OBSERVED OUTSIDE OF P-LEVEL CHANGES. PS WAS OBSERVED TO BE TRENDING DOWN. PLACEMENT SIGNAL ALARMS OBSERVED FOLLOWED BY IMPELLA POSITION IN VENTRICLE/ AORTA AT P-9 IN THE FIRST HOUR OF SUPPORT AFTER INSERTION. PURGE PRESSURE SPIKE OBSERVED FOLLOWED BY MANY SUCTIONS AT P-3. DROP IN FLOW OBSERVED AT -P4 TO THE END OF CP SUPPORT CLINICAL STATED: BEDSIDE ECHO PERFORMED TO TRY TO DETERMINE SOURCE OF HEMOLYSIS. IMPELLA FREE FLOATING MID LV AND FREE OF ALL STRUCTURES. PATIENT RECEIVING MULTIPLE BLOOD PRODUCTS, FFP, PLASMA, RBCS. MULTIPLE AMPS BICARBONATE. RED URINE NOTED. CONCERN FOR DIC. GOOD POSITION CONFIRMED. PATIENT WAS SEVERELY ACIDOTIC AND WAS CHASING BLEEDING. HEMOLYSIS WAS NOT FROM THE PUMP. HEMOLYSIS: THE CAUSE OF THE HEMOLYSIS ISSUE MOST LIKELY WAS DUE PATIENT CONDITION SINCE PATIENT WAS SEVERELY ACIDOTIC AND WAS CHASING BLEEDING. LOW PUMP FLOW: FAILURE MODE LOW PUMP FLOW ADDED BASED ON THE OBSERVATION FROM COMPLAINT PUMP LOG REVIEW. THE CAUSE OF THE LOW PUMP FLOW ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED SINCE PATIENT WAS SEVERELY ACIDOTIC, CARDIAC OUTPUT WAS LESS THAN THE MINIMUM AND UNABLE TO ACHIEVE ADEQUATE BLOOD VOLUME TO THE LEFT SIDE TO ACHIEVE ADEQUATE FLOW. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED AND NOTED THE PATIENT WAS SEVERELY ACIDOTIC AND WAS CHASING BLEEDING. THE LOCAL STAFF AND TEAM UNDERSTOOD THE PATIENT WOULD PROBABLY NOT SURVIVE. MINIMUM CARDIAC OUTPUT PATIENT NEEDED WAS AT LEAST 4.1 L/MIN AND WAS UNABLE TO BE ACHIEVED; UNABLE TO GET THE VOLUME TO THE LEFT SIDE TO ACHIEVE ADEQUATE FLOWS AND PATIENT¿S LACTATE CONTINUED TO RISE. MEDICAL SAFETY OFFICER REVIEWED THE EVENT AND DETERMINED THE PATIENT ON BIPELLA SUPPORT DEVELOPED HEMOLYSIS. ATTEMPT MADE TO DETERMINE THE SOURCE FOR THE HEMOLYSIS; NO FURTHER INFORMATION NOTED. THE PATIENT CRITICALLY ILL, EXPIRES AFTER APPROXIMATELY EIGHT HOURS AFTER START OF IMPELLAS MECHANICAL CIRCULATORY SUPPORT. WITH THE INABILITY TO DISSOCIATE, THE DEMISE IS TO BE FILED AGAINST THE IMPELLA CP (WHICH WAS THE FIRST IMPLANT) AND IMPELLA IMPELLA RP FLEX IS A SERIOUS INJURY WHICH IS REPORTED UNDER MANUFACTURING REPORT NUMBER 1220648-2025-29109. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22178680&pc=DSQ | 2916596-2025-03691 | 2025-05-18 04:00 | Death | THORATEC CORPORATION | 2025-06-10 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO REFRACTORY MULTI-ORGAN FAILURE (MOF). THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. REFRACTORY AND PROGRESSIVE MOF WAS RELATED TO THE OUTCOME. NO DEVICE RELATED ISSUES LED TO MOF. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS RUNNING AS EXPECTED. NO DEVICE ISSUES WERE MENTIONED PRIOR TO THE OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22178932&pc=OZD | 1220648-2025-29177 | 2025-05-11 04:00 | Death | ABIOMED, INC. | 2025-06-10 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Hemorrhage/Bleeding; Thrombocytopenia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT ON IMPELLA 5.5 SUPPORT EXPERIENCED SUCTION EVENTS. TWO UNITS OF PACKED RED BLOOD CELLS WERE GIVEN. THE NEXT DAY HEPARIN WAS ON HOLD DUE TO HIGH PARTIAL THROMBOPLASTIN TIME AND LOW XA. THE PATIENT WAS BROUGHT TO THE OPERATING ROOM THE FOLLOWING DAY FOR A WASHOUT AND CLOSURE. RED BLOOD CELLS AND PLATELETS WERE GIVEN AND LABS WERE REDRAWN. AN ENDOSCOPY WAS PERFORMED THE NEXT DAY FOR A GASTROINTESTINAL BLEED. BLEEDING CONTINUED FOR DAYS REQUIRING PLATELETS AND RED BLOOD CELLS AT EVERY SHIFT. THE PATIENT REMAINED CRITICALLY ILL THROUGHOUT SUPPORT. IT WAS NOTED THAT IT WAS EVIDENT THAT THE PATIENT HAD MULTISYSTEM ORGAN FAILURE AND CONTINUED TO HAVE RECTAL BLEEDING. HEPARIN WAS STOPPED. ULTIMATELY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22179185&pc=OZD | 1220648-2025-29179 | 2025-05-16 04:00 | Death | ABIOMED, INC. | 2025-06-10 04:00 | OZD | IMPELLA CP | Positioning Failure | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, IT WAS NOTED THAT THE PATIENT IS HEMODYNAMICALLY STABLE AND LACTIC WAS NORMALIZED. IT WAS DECIDED TO START A WEAN PROCESS AND SEE IF THE PATIENT WOULD NEED TO BE TRANSFERRED FOR ESCALATION OF CARE. AFTER WEANING, THE PATIENT REQUIRED THE NEED FOR INCREASED VASOPRESSORS. IN ADDITION, THERE WERE CONTINUOUS ALARMS AND DROPPED TO P-2. THE PATIENT WAS HAVING FREQUENT RUNS OF V-FIB. ULTIMATELY THE PATIENT CODED AND EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: WARNINGS, CAUTIONS AND PRECAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ UNDERSTANDING AND MANAGING IMPELLA CATHETER POSITION ALARMS ¿IF THE IMPELLA CATHETER IS COMPLETELY OUT OF THE VENTRICLE, DO NOT ATTEMPT TO REPOSITION THE CATHETER ACROSS THE VALVE WITHOUT A GUIDEWIRE.¿ CORRECT IMPELLA CATHETER POSITION ¿FOR OPTIMAL POSITIONING OF THE IMPELLA CATHETER, THE INLET AREA OF THE CATHETER SHOULD BE 3.5 CM BELOW THE AORTIC VALVE ANNULUS AND WELL AWAY FROM PAPILLARY MUSCLE AND SUBANNULAR STRUCTURES. THE OUTLET AREA SHOULD BE WELL ABOVE THE AORTIC VALVE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22179271&pc=CBK | 1220908-2025-02121 | 2025-05-16 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-06-10 04:00 | CBK | MAIN ASSEMBLY, VENTILATOR, COMMERCIAL, ZVENT, BASI | Failure of Device to Self-Test | Insufficient Information | K162832 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE DISPLAYED A "SELF CHECK" ERROR MESSAGE. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. HOWEVER, THE DEVICE LOG WAS PROVIDED FOR REVIEW. THE REPORTED EVENT OCCURRED IN (B)(6) 2025, BUT NO ACTIVITY HAS BEEN LOGGED SINCE DECEMBER 2024. BASED ON THE EVIDENCE, THE DEVICE APPEARS NOT TO HAVE BEEN USED DURING THE REPORTED EVENT. THIS CLAIM WILL BE CLOSED AS NO SAMPLE RETURNED. ANALYSIS FOR REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: JUSTIFICATION FOR NO UDI: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22179284&pc=OZD | 1220648-2025-29180 | 2025-05-13 04:00 | Death | ABIOMED, INC. | 2025-06-10 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Arrhythmia; Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT (UNKNOWN ETHNICITY) WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS INITIALLY IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND DEVELOPED HEMOLYSIS RESULTING IN DEVICE EXPLANT WITH REPLACEMENT TO AN IMPELLA 5.5. DURING THE IMPELLA 5.5 DEVICE SUPPORT, THE PATIENT EXPERIENCED LOW PUMP FLOW, TACHYCARDIA, RETURN OF HEMOLYSIS AND WOULD SUBSEQUENTLY EXPIRE. THE PATIENT PRESENTED WITH SHORTNESS OF BREATH AND DIAGNOSTICS REVEALED AN EJECTION FRACTION OF 15% WITH MULTI-VESSEL DISEASE. PERCUTANEOUS CORONARY INTERVENTION (PCI) WAS PERFORMED WITH TWO STENTS TO THE LEFT ANTERIOR DESCENDING ARTERY AND ONE STENT TO THE RAMUS ARTERY WHICH OCCURRED APPROXIMATELY FIVE DAYS PRIOR TO IMPELLA CP DEVICE IMPLANT FOR HEMODYNAMIC SUPPORT. THE IMPELLA CP WAS SUCCESSFULLY PLACED, AND AT THIS TIME CORONARIES WERE ASSESSED, AND IT WAS REVEALED THE RECENTLY PLACED RAMUS ARTERY STENT THROMBOSED. CORONARY ANGIOPLASTY WAS PERFORMED WITH GOOD RETURN VASCULARIZATION. AFTER THIS PCI, COMPANION AND PEEL-AWAY SHEATHS WERE REMOVED, REPOSITIONING SHEATH PLACED, ANGLE MATCHED WITH NO OOZING AT THE SITE. THE PATIENT WAS SENT TO THE UNIT ON IMPELLA MCS. APPROXIMATELY 16 HOURS AFTER START OF THE IMPELLA CP MCS, THE PATIENT DEVELOPED HEMOLYSIS MANAGED WITH DECREASE IN PERFORMANCE LEVEL AND ESCALATION TO AN IMPELLA 5.5 WHICH WAS SUCCESSFULLY PLACED. THE PATIENT SURVIVED THE EXPLANT WITH REPLACEMENT AND WAS NOTED TO BE IN STABLE CONDITION. THE PATIENT'S URINE OUTPUT INCREASED AND WAS CLEARING; NEXT DAY, URINE WAS CLEAR/YELLOW. APPROXIMATELY, TWO DAYS THEREAFTER, DURING SHOCK ROUNDS, THE PHYSICIAN ATTEMPTED TO REPOSITION THE IMPELLA 5.5 AS AN ATTEMPT TO INCREASE THE PERFORMANCE LEVEL FROM P-4. THE LEVEL WAS INCREASED TO P-5 AND THERE WAS SUCTION. THEREFORE, THE PHYSICIAN STATED THAT THE IMPELLA 5.5 WAS OPTIMIZED AS MUCH AS POSSIBLE WITH P-4. TWO DAYS LATER, THE PATIENT WAS REPORTEDLY HAVING AN INCREASINGLY TACHYCARDIAC EVENT, AND THE FOLLOWING DAY HEMOLYSIS WAS NOTED: PINK TINGED URINE WITH SUCTION AND INCREASED PERFORMANCE LEVEL OVERNIGHT. PERFORMANCE LEVEL WAS DECREASED TO MANAGE THE HEMOLYSIS, AND AN ECHOCARDIOGRAM WAS PERFORMED (RESULTS OF THE ECHOCARDIOGRAM WERE UNNOTED). THE PATIENT CONTINUED SUPPORT; FOUR DAYS THEREAFTER, IT WAS NOTED THE PATIENT BEGAN TO DECOMPENSATE THROUGHOUT THE DAY AT A LOWERED PERFORMANCE LEVEL OF P-2. THE PATIENT WAS PLACED ON BILEVEL POSITIVE AIRWAY PRESSURE OVERNIGHT. VENOUS OXYGENATION SATURATION WAS DOWN TO 26% AND TACHYCARDIA WORSENED. THREE DAYS LATER, THE PATIENT WENT INTO CARDIAC ARREST AND WAS SUBSEQUENTLY INTUBATED AND STARTED ON HIGH DOSE PRESSORS POST RETURN OF SPONTANEOUS CIRCULATION. OVERNIGHT, THE PATIENT EXPERIENCED ADDITIONAL CARDIAC EVENTS REQUIRING DEFIBRILLATION. THE PATIENT WAS NOTED NOT AN EXTRACORPOREAL MEMBRANE OXYGENATION CANDIDATE; URINE OUTPUT WAS DOWN 10-15 CC/HOUR ON LASIX DRIP. CONTINUOUS RENAL REPLACEMENT THERAPY WAS NOTED AS A POSSIBLE NEED, BUT THE PATIENT WOULD SUBSEQUENTLY EXPIRE. Manufacturer Narrative: MEDICAL SAFETY OFFICER REVIEWED THE EVENT AND DETERMINED THE HEMOLYSIS OCCURRING WITH THE IMPELLA CP CLEARED AFTER ESCALATION AND IS A SERIOUS INJURY. HOWEVER, IMPELLA 5.5-RELATED EVENTS MAY HAVE ADDED TO AN ALREADY COMPLEX SITUATION AND THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA 5.5 DEVICE USED AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME. THE IMPELLA 5.5 DEVICE IS ONE OF TWO DEVICES ASSOCIATED WITH THE EVENT. REFER TO INITIAL MEDICAL DEVICE REPORT FOR IMPELLA CP SERIAL NUMBER (B)(6) WITH DATE OF EVENT 11-MAY-2025 AND PATIENT IDENTIFIER ENDING IN (B)(6). THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS - ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE - ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS - ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION - ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22179565&pc=NKM | 2135147-2025-03076 | 2025-05-16 04:00 | Death | ABBOTT MEDICAL | 2025-06-10 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "IMPACT OF INSTITUTIONAL MONTHLY VOLUME OF TRANSCATHETER EDGE-TO-EDGE REPAIR PROCEDURES FOR SIGNIFICANT MITRAL REGURGITATION: EVIDENCE FROM THE GIOTTO-VAT STUDY" WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE IMPACT OF INSTITUTIONAL MONTHLY VOLUME ON OUTCOMES OF TEER TO REMEDY SIGNIFICANT MITRAL REGURGITATION. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT HIGH INSTITUTIONAL MONTHLY VOLUME OF TEER MITRAL VALVE REPAIR APPEARS TO CORRELATE WITH AN IMPROVED PROCEDURAL SUCCESS RATE AND SHORTER HOSPITALIZATIONS. SIMILARLY FAVORABLE RESULTS WERE FOUND FOR LONG-TERM RATES OF CARDIAC DEATH OR HOSPITALIZATION FOR HEART FAILURE. THESE FINDINGS INFORM ON THE IMPORTANCE OF OPERATOR EXPERIENCE AND CENTER EXPERTISE IN ACHIEVING STATE-OF-THE-ART RESULTS WITH TEER, WHILE CONFIRMING THE USEFULNESS OF THE PROCTORING APPROACH WHEN NAÏVE CENTERS BEGIN A TEER PROGRAM. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS NICOLA CORCIONE AT SAPIENZA UNIVERSITY OF ROME INSTITUTION WITH CORRESPONDING EMAIL NICOLACORCIONE72@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS NOT SPECIFIED. A TOTAL OF 2213 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 2213 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, CORONARY ARTERY DISEASE, PRIOR MYOCARDIAL INFARCTION, PRIOR CEREBROVASCULAR EVENT. (B)(6), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED BLEEDING, HEART FAILURE, PROLONGED HOSPITALIZATION, SURGICAL INTERVENTION, STROKE, ATRIAL FIBRILLATION. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22179694&pc=NVN | 2182208-2025-02818 | 2025-03-20 04:00 | Death | MEDTRONIC, INC. | 2025-06-10 04:00 | NVN | SELECTSECURE MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P030036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE CLINICAL OUTCOMES OF HIS BUNDLE PACING (HBP) VERSUS RIGHT VENTRICULAR PACING (RVP) IN PATIENTS WHO DEVELOP CONDUCTION DISTURBANCES FOLLOWING TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR). THE AUTHORS DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE OR CARDIOVASCULAR MORTALITY. THERE WERE PATIENTS WHO EXPERIENCED OTHER MAJOR ADVERSE CARDIOVASCULAR EVENTS (MACE) WHICH INCLUDED HEART FAILURE HOSPITALIZATIONS, ATRIAL FIBRILLATION (AF), AND LEADS WHICH EXHIBITED DISLODGMENTS AND WERE REVISED. THE STATUS OF THE LEADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS FEMALE/74 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CLINICAL OUTCOMES OF HIS BUNDLE PACING VS. RIGHT VENTRICULAR PACING IN PATIENTS WITH CONDUCTION DISTURBANCES FOLLOWING TRANSCATHETER AORTIC VALVE REPLACEMENT. BMC CARDIOVASCULAR DISORDERS. 2025. 25:202. DOI: 10.1186/S12872-025-04643-6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22179702&pc=GAM | 2210968-2025-06581 | 2024-12-10 05:00 | Death | ETHICON INC. | 2025-06-10 04:00 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Post Operative Wound Infection | K022269 |  | 1 | Event Description: SURGICAL OUTCOME AFTER AUTOLOGOUS BONE CHIPS REPLACEMENT IN DEPRESSED SKULL FRACTURES: A SINGLE CENTER EXPERIENCE. THE AIM OF THIS STUDY IS TO EVALUATE DIFFERENT PATTERNS OF OUTCOME INCLUDING (SAFETY, EFFICACY, COMPLICATIONS, PATIENTS¿ SATISFACTION AND NEED FOR LATE CP) OF USING AUTOLOGOUS BONE CHIPS FOR MANAGEMENT OF DEPRESSED SKULL FRACTURES (DSFS). BETWEEN MARCH 2023 AND NOVEMBER 2023, A TOTAL OF 68 TRAUMATIC BRAIN INJURY (TBI) PATIENTS WITH DSFS WERE OPERATED UPON IN OUR NEUROSURGERY DEPARTMENT. BASED ON OUR ELIGIBILITY CRITERIA, 59 PATIENTS FULFILLED OUR INCLUSION CRITERIA AND WERE ACTIVELY INCLUDED. AMONG OUR INCLUDED CASES, (89.8%) HAD MILD TBI AND THE REMAINING (10.2%) HAD MODERATE TBI. THE MEAN AGE WAS (16.505±12.42) YEARS WITH RANGE FROM 6 MONTHS TO 59 YEARS. THE VAST MAJORITY OF PATIENTS 41 CASES (69.5%) WERE =20 YEARS AND 18 PATIENTS (30.5%) WERE >20 YEARS. MALES CONSTITUTED THE MAIN BULK OF STUDY POPULATION (78%) AND FEMALE REPRESENTED ONLY (22%). RTA WAS THE MOST FREQUENT MOOD OF TRAUMA (35.6%) FOLLOWED BY ASSAULT (23.7%), DIRECT TRAUMA (23.47%) THEN: FFH (16.9%). VICRYL 0 (ETHICON) WAS USED TO MAKE A NET THROUGH EDGES OF THE GALEA THEN BONE FRAGMENTS WERE REPLACED AND FIXED USING THIS NET. REPORTED COMPLICATIONS ARE : VICRYL 0 (ETHICON) -SEIZURES (N=3.4%) TREATMENT: NOT REPORTED -CSF LEAKAGE (N=3.4%) TREATMENT : NOT REPORTED -WOUND INFECTION (N=1) TREATMENT : WE STARTED EMPIRIC BROAD-SPECTRUM ANTIBIOTIC REGIMEN WITH VANCOMYCIN (15 MG/KG IV EVERY 12 H) PLUS A THIRD A GENERATION CEPHALOSPORIN (CEFTRIAXONE 2 G IV DAILY) AND SHOWED GOOD RESPONSE TO THE ANTIBIOTIC TREATMENT VICRYL 0 (ETHICON) DEATH (N=1); THIS PATIENT HAD WOUND INFECTION AND DEVELOPED OSTEOMYELITIS TREATMENT: THE PATIENT HAD WOUND INFECTION WHICH WAS TREATED WITH EMPIRIC BROAD-SPECTRUM ANTIBIOTIC REGIMEN WITH VANCOMYCIN (15 MG/KG IV EVERY 12 H) PLUS A THIRD A GENERATION CEPHALOSPORIN (CEFTRIAXONE 2 G IV DAILY); DEVELOPED OSTEOMYELITIS AND UNDERWENT RE-SURGERY FOR WOUND DEBRIDEMENT AND REMOVAL OF BONE FRAGMENTS, LATER THIS CASE WAS EXPIRED. IN CONCLUSION, USING AUTOLOGOUS DEPRESSED BONE CHIPS FOR SKULL DEFECT RECONSTRUCTION CAN BE A SAFE AND FEASIBLE SURGICAL TECHNIQUE FOR TBI PATIENTS SUFFERING DSFS WITH GOOD AESTHETIC RESULTS, HIGH PATIENT SATISFACTION, DECREASED NEED FOR LATER CRANIOPLASTY AND CONSEQUENTLY LOW OVERALL MANAGEMENT COST. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DOES THE SURGEON BELIEVE THAT ANY OF THE ETHICON PRODUCTS INVOLVED CAUSED AND/OR CONTRIBUTED TO THE POST-OPERATIVE COMPLICATIONS DESCRIBED IN THE ARTICLE? WHICH SPECIFIC ETHICON PRODUCTS HAVE BEEN USED DURING THE PROCEDURES (PRODUCT CODE, LOT NUMBER)? DOES THE SURGEON BELIEVE THERE WAS ANY DEFICIENCY WITH ANY OF THE ETHICON PRODUCTS USED IN THIS PROCEDURE? IF SO, PLEASE PROVIDE DETAILS. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE A COMPLAINT REFERENCE NUMBER. PATIENT DEMOGRAPHICS? WOULD THE JOURNAL AUTHOR / PHYSICIAN LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. H6 COMPONENT CODE: G07002 ¿ DEVICE NOT RETURNED. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. CITATION: NEUROSURG REV. 2024 DEC 10;47(1):898. DOI: 10.1007/S10143-024-03128-Y. PMID: 39656311; PMCID: PMC11632004. HTTPS://DOI.ORG/10.1007/S10143-024-03128-Y.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22179986&pc=NVN | 2017865-2025-85904 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-10 04:00 | NVN | TENDRIL MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS CONGESTIVE HEART FAILURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22180018&pc=NIK | 2017865-2025-85906 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-10 04:00 | NIK | QUARTET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS CONGESTIVE HEART FAILURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22180022&pc=NIK | 2017865-2025-85903 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-10 04:00 | NIK | QUADRA ALLURE MP RF CRT-P | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS CONGESTIVE HEART FAILURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22180045&pc=NVN | 2017865-2025-85908 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-10 04:00 | NVN | TENDRIL MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS CONGESTIVE HEART FAILURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22180253&pc=MSD | 1282497-2025-00402 | 2023-11-22 05:00 | Death | COVIDIEN LLC | 2025-06-10 04:00 | MSD | DIALYSIS UNKNOWN | Product Quality Problem | Cardiac Tamponade |  |  | 1 | Event Description: PLI 10: ACCORDING TO THE LITERATURE, A 25-YEAR-OLD FEMALE WITH A HISTORY OF MULTIPLE DIALYSIS ACCESS FAILURES PRESENTED FOR AN EXCHANGE OF A TUNNELED CENTRAL VENOUS CATHETER (CVC) VIA THE LEFT INTERNAL JUGULAR VEIN. THE PATIENT MEDICAL HISTORY INCLUDED HYPERTENSION, CONGESTIVE HEART FAILURE WITH A REDUCED EJECTION FRACTION OF 25%, AND END-STAGE KIDNEY DISEASE (ESKD). IN AUGUST 2022, PATIENT HAD BEEN ADMITTED FOR THE CREATION OF A NEW RIGHT UPPER ARM ARTERIOVENOUS GRAFT (AVG) AND EXCHANGE OF HER LEFT INTERNAL JUGULAR (IJ) CVC DUE TO CATHETER DYSFUNCTION ATTRIBUTED TO RETRACTION. DURING THE PROCEDURE, A SINGLE STIFF GLIDE WIRE (0.035", 180 CM; TERUMO, NJ) WAS ADVANCED THROUGH ONE PORT OF THE EXISTING CATHETER INTO THE INFERIOR VENA CAVA (IVC). THE OLD CVC WAS THEN REMOVED ALONG THE PREVIOUS TUNNEL. A 23 CM PRECISION CATHETER WAS INSERTED OVER THE GUIDEWIRE USING THE PROVIDED STIFFENERS AND ADVANCED TOWARD THE CENTRAL CIRCULATION. HOWEVER, SIGNIFICANT RESISTANCE WAS ENCOUNTERED AT THE JUNCTION OF THE SUPERIOR VENA CAVA (SVC) AND RIGHT ATRIUM. DURING THE PROCEDURE THE WIRE WAS RETRACTED INTO WHAT WAS BELIEVED TO BE THE RIGHT VENTRICLE. SEVERAL ATTEMPTS TO REDIRECT THE GUIDEWIRE INTO THE IVC (INFERIOR VENA CAVA) AT THIS POINT, SVC STENOSIS WAS SUSPECTED. THE PATIENT EXPERIENCED ACUTE CARDIOPULMONARY DECOMPENSATION AND RAPIDLY DEVELOPING TO CARDIOPULMONARY ARREST. PERICARDIAL TAMPONADE WAS SUSPECTED, PROMPTING THE PLACEMENT OF A PERICARDIAL DRAINAGE CATHETER, WHICH YIELDED 100 ML OF BLOOD. DESPITE THESE INTERVENTIONS, THE PATIENT EXPIRED WITHIN ONE HOUR. POSTMORTEM AUTOPSY SHOWED A 2 MM TRANSMURAL, INTRAPERICARDIAL DEFECT IN THE SVC WALL, CONSISTENT WITH VASCULAR PERFORATION. ADDITIONALLY, A SEVERE, ORGANIZED STENOSIS OF THE SVC JUST ABOVE THE SVC¿RIGHT ATRIAL JUNCTION WAS NOTED. FATAL SUPERIOR VENA CAVA RUPTURE COMPLICATING DIALYSIS CATHETER EXCHANGE-THE JOURNAL OF VASCULAR ACCESS. Manufacturer Narrative: ADDITIONAL INFORMATION: A4, G3, MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: AHMED E ALI, ALIAN AL-BALAS, PAUL V BENSON AND AMMAR ALMEHMI FATAL SUPERIOR VENA CAVA RUPTURE COMPLICATING DIALYSIS CATHETER EXCHANGE THE JOURNAL OF VASCULAR ACCESS ARTICLE REUSE GUIDELINES: SAGEPUB.COM/JOURNALS-PERMISSIONS DOI: 10.1177/11297298231219288. JOURNALS.SAGEPUB.COM/HOME/JVA MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22180322&pc=LWS | 9614453-2025-02245 | 2025-05-09 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-06-10 04:00 | LWS | AURORA EV-ICD¿ MRI SURESCAN¿ | Insufficient Information | Bradycardia; Discomfort | P220012 |  | 1 | Event Description: IT WAS REPORTED THAT AN EXTRAVASCULAR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (EV-ICD) PATIENT EXPERIENCED BRADYCARDIA. THE PATIENT COMPLAINED OF DISCOMFORT WITH PACING, SO THE PACING FEATURE WAS ULTIMATELY LEFT PROGRAMMED OFF. ADDITIONAL INFORMATION RECEIVED INDICATED THE PATIENT EXPIRED TWO DAYS LATER DUE TO METABOLIC CIRCUMSTANCES. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22180374&pc=NPT | 2015691-2025-04744 | 2025-05-16 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-10 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Hematoma; Rupture | P140031 |  | 1 | Event Description: EDWARDS RECEIVED NOTIFICATION FROM OUR AFFILIATE IN POLAND. AS REPORTED, THIS WAS A CASE OF AN IMPLANT OF A 26MM SAPIEN 3 ULTRA IN AORTIC POSITION BY RIGHT TRANSFEMORAL APPROACH. THE PATIENT PRESENTED A TRICUSPID VALVE WITH LARGE CALCIFICATIONS OF THE LEAFLETS. WHEN THE DELIVERY SYSTEM WITH THE VALVE WAS CROSSING THE NATIVE ANNULUS, THE VALVE SLID OFF THE BALLOON AND JUMP INTO THE ANNULUS, PROBABLY DUE TO THE ANNULAR CALCIFICATION. THE VALVE WAS WITHDRAWN WITH THE DELIVERY SYSTEM INTO THE ABDOMINAL AORTA AND IT WAS ATTEMPTED TO RE-INSERT IT IN THE ESHEATH+ TO REMOVE IT BUT IT WAS NOT POSSIBLE. AFTER SEVERAL ATTEMPTS, THE VALVE COMPLETELY SLIPPED OFF THE BALLOON SO, IT WAS DECIDED TO REMOVE THE DELIVERY SYSTEM, LEAVING THE VALVE DISLODGED, AND AN 18F BALLOON WAS PASSED THROUGH IT IN AN ATTEMPT TO RETRACT THE VALVE INTO THE ESHEATH+, WHICH HAD BEEN TORN DURING IMPLANTATION. AS THE VALVE COULD NOT BE REMOVED, THE ESHEATH+ WAS REMOVED AND A NON-EDWARDS 18F SHEATH WAS INSERTED. ANOTHER ATTEMPT TO RETRIEVE THE VALVE WAS DONE BUT IT WAS UNSUCCESSFUL. THE 18F BALLOON WAS EXCHANGED FOR A HIGH-PRESSURE 6F BALLOON, AND THE VALVE WOULD BE PULLED TOGETHER WITH THE SHEATH DOWN TO THE ILIAC ARTERY. IN THIS LOCATION, THE VALVE WOULD THEN BE SURGICALLY REMOVED. A WHOLE NEW SYSTEM WAS PREPARED, AND THE SECOND VALVE WAS SUCCESSFULLY DEPLOYED. A HEMATOMA WAS OBSERVED, BUT NO ACTIONS WERE TAKEN, AND THE PATIENT WAS NOTED TO BE STABLE. THE DAY AFTER PROCEDURE THE PATIENT FELT UNWELL AND A BLEEDING WAS DISCOVERED FROM AN AORTIC ROOT RUPTURE. AN OPEN CHEST SURGERY WAS PERFORMED TO TREAT THE RUPTURE AND IT WAS PERFORMED A SURGICAL AORTIC VALVE REPLACEMENT. THE PATIENT WAS NOTED TO BE IN STABLE CONDITIONS AFTER THE PROCEDURE. HOWEVER, HE PATIENT PASSED AWAY TWO DAYS AFTER THE PROCEDURE. Manufacturer Narrative: SUPPLEMENTAL REPORT SUBMITTED TO CORRECT H10. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THAT PATIENT RELATED FACTORS (LARGE CALCIFICATIONS OF THE NATIVE LEAFLETS) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22180535&pc=NPT | 2015691-2025-04745 | 2025-05-14 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-10 04:00 | NPT | SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS ECUADOR AFFILIATE, DURING A TRANSFEMORAL TAVR PROCEDURE WITH A 29MM SAPIEN 3 TRANSCATHETER HEART VALVE, THE VALVE WAS ADVANCED FOR DEPLOYMENT. THE VALVE EXPANSION WAS CONFIGURED AT -1CC TO ACHIEVE A CONSERVATIVE 6% OVERSIZING CONSIDERING THE PATIENT'S ANATOMY. THE PATIENT REMAINED HEMODYNAMICALLY STABLE POST-IMPLANTATION, BUT SUBSEQUENTLY DEVELOPED HYPOTENSION, WHICH RESPONDED TO VASOPRESSOR SUPPORT. VASOPRESSORS WERE LATER DISCONTINUED AS THE PATIENT'S BLOOD PRESSURE STABILIZED WITHOUT PHARMACOLOGIC SUPPORT. FOLLOW-UP IMAGING SUGGESTED A POSSIBLE ANNULAR RUPTURE; HOWEVER, SUBSEQUENT IMAGING DID NOT SHOW THE RUPTURE. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT WITHOUT VASOPRESSORS. A TRANSTHORACIC ECHOCARDIOGRAM CONFIRMED THAT THE VALVE WAS FUNCTIONING APPROPRIATELY. THE PATIENT EXPERIENCED CARDIAC ARREST 36 HOURS LATER AND PASSED AWAY IN INTENSIVE CARE, UNDER A DO-NOT-RESUSCITATE ORDER SIGNED BY FAMILY MEMBERS IN ACCORDANCE WITH THE PATIENT'S WISHES. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIAC ARREST IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BALLOON VALVULOPLASTY, THE USE OF LOCAL AND/OR GENERAL ANESTHESIA, BIOPROSTHETIC HEART VALVES, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. CARDIAC ARREST OCCURS WHEN A MAJOR CARDIOVASCULAR, RESPIRATORY, OR METABOLIC EVENT RESULTS IN THE INABILITY OF CARDIAC MUSCLE TO GENERATE SUFFICIENT FORCE IN RESPONSE TO ELECTRICAL DEPOLARIZATION. IT MAY BE CAUSED BY A PROFOUND CARDIOVASCULAR INSULT. EXAMPLES INCLUDE SEVERE PROLONGED HYPOXIA OR ACIDOSIS, EXTREME HYPOVOLEMIA, FLOW-RESTRICTING PULMONARY EMBOLUS, CARDIAC TAMPONADE, THROMBOSIS (CORONARY OR PULMONARY). THESE EVENTS MAY BE RELATED TO THE EDWARDS DEVICE IF IT IS ASSOCIATED WITH CARDIAC TAMPONADE, ANNULAR RUPTURE, OR OTHER EDWARDS'S DEVICE RELATED BLEED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ALTHOUGH THE EXACT CAUSE AND SOURCE OF THE CARDIAC ARREST CANNOT BE DETERMINED, IT IS POSSIBLE THE MECHANISMS OF THE TAVR PROCEDURE DESCRIBED ABOVE OR PATIENT FACTORS NOT PROVIDED MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22180542&pc=BZD | 2518422-2025-106527 | 2025-03-05 05:00 | Death | RESPIRONICS, INC. | 2025-06-10 04:00 | BZD | DREAMSTATION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER MAILING SENT OUT DUE TO A DEVICE RECALL. THE MANUFACTURER WAS INFORMED BY A FAMILY MEMBER THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22180829&pc=NPT | 2135147-2025-03083 | 2017-10-01 04:00 | Death | ABBOTT MEDICAL | 2025-06-10 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P190023 |  | 1 | Event Description: THE ARTICLE, "HEMODYNAMIC PERFORMANCE OF A SELF-EXPANDING TRANSCATHETER AORTIC VALVE WITH AN INTRA-ANNULAR LEAFLET POSITION IN PATIENTS WITH A SMALL AORTIC ANNULUS", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE STUDY TO ASSESS THE SHORT-TERM HEMODYNAMIC PERFORMANCE OF THE SELF-EXPANDING PORTICO/NAVITOR TRANSCATHETER AORTIC VALVES (TAVS), WHICH HAS AN INTRA-ANNULAR LEAFLET POSITION, IN PATIENTS WITH SMALL AORTIC ANATOMIES. DEVICES INCLUDED IN THE STUDY WERE NAVITOR AND PORTICO. THE ARTICLE CONCLUDED THAT SELF-EXPANDING TRANSCATHETER AORTIC VALVES WITH AN INTRA-ANNULAR LEAFLET POSITION ARE ASSOCIATED WITH FAVORABLE HEMODYNAMIC PERFORMANCE IN PATIENTS WITH A SMALL AORTIC ANNULUS. [THE PRIMARY AND CORRESPONDING AUTHOR WAS MATJAZ BUNC, DEPARTMENT OF CARDIOLOGY, UNIVERSITY MEDICAL CENTRE LJUBLJANA, 1000 LJUBLJANA, SLOVENIA, WITH CORRESPONDING EMAIL: MBUNCEK@YAHOO.COM] THE TIME FRAME OF THE STUDY WAS FROM OCTOBER 2017 TO AUGUST 2024. A TOTAL OF 37 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 81.6 YEARS AND THE MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED SMALL AORTIC ANNULI. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF HEMODYNAMIC PERFORMANCE OF A SELF-EXPANDING TRANSCATHETER AORTIC VALVE WITH AN INTRA-ANNULAR LEAFLET POSITION IN PATIENTS WITH A SMALL AORTIC ANNULUS WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH CO-MORBIDITIE INCLUDING SMALL AORTIC ANNULI. SOME OF THE COMPLICATIONS REPORTED WERE CARDIAC ARREST, AORTIC VALVE INSUFFICIENCY/ REGURGITATION, CONGESTIVE HEART FAILURE, OBSTRUCTION/OCCLUSIONAND DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B2: DEATH DATE IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: HEMODYNAMIC PERFORMANCE OF A SELF-EXPANDING TRANSCATHETER AORTIC VALVE WITH AN INTRA-ANNULAR LEAFLET POSITION IN PATIENTS WITH A SMALL AORTIC ANNULUS. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DEATH DATE IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: HEMODYNAMIC PERFORMANCE OF A SELF-EXPANDING TRANSCATHETER AORTIC VALVE WITH AN INTRA-ANNULAR LEAFLET POSITION IN PATIENTS WITH A SMALL AORTIC ANNULUS
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22181598&pc=NAY | 2955842-2025-24520 | 2025-05-08 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-06-10 04:00 | NAY | DAVINCI XI | Insufficient Information | Hemorrhage/Bleeding; Asystole | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI ASSISTED LIVER CYSTECTOMY, THE PATIENT EXPERIENCED A SMALL AMOUNT OF BLEEDING FROM THE HEPATIC VEIN SHORTLY BEFORE THE END OF THE PROCEDURE. THE OPERATING ROOM NURSE REPORTED THAT THE PROCEDURE WAS CONVERTED TO OPEN TO ACHIEVE HEMOSTASIS. CONTROLLING THE BLEEDING WAS DIFFICULT, BUT WAS ACHIEVED. THE PATIENT THEN EXPERIENCED ASYSTOLE AND EXPIRED. THE ORIGIN OF THE ASYSTOLE WAS UNCLEAR. THE VENA CAVA WAS REPORTED TO BE WELL-FILLED WITH GOOD CIRCULATION OF THE INTESTINES AND LIVER IMMEDIATELY BEFORE THE ASYSTOLE OCCURRED. THE REPORTED DIFFERENTIAL DIAGNOSES WERE PULMONARY EMBOLISM OR VAGAL STIMULATION AFTER REPOSITIONING OF THE ROCHARD HOOK RETRACTOR. AUTOPSY RESULTS ARE NOT AVAILABLE. THE SURGEON BELIEVES THAT THE DA VINCI SYSTEM HAD NO IMPACT ON THE PATIENT¿S DEATH.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: SECTION B5 ADDITIONAL INFORMATION: THE MOBILIZATION OF THE TUMOR WAS ALMOST COMPLETED WHEN THE BLEEDING BEGAN. A SYNCHROSEAL INSTRUMENT AND UNSPECIFIED BIPOLAR INSTRUMENT WERE BEING USED DURING THIS TASK. THE CUSTOMER SAID THERE WAS NO MALFUNCTION OF A DA VINCI PRODUCT DURING THIS EVENT. THE BLEEDING WAS COMING FROM THE LIVER VEIN, DISTAL TO THE TUMOR. CENTRALLY, THE VEIN WAS ADHERED TO THE TUMOR. THE DECISION TO CONVERT TO OPEN WAS MORE DUE TO THE VEIN BEING ADHERED TO THE TUMOR IN THE CENTRAL PART THAN THE BLEEDING. THE RESECTION OF THE TUMOR AND HEMOSTASIS WERE BOTH COMPLETED WHILE OPEN. THE BLOOD SUPPLY TO THE LIVER AND INTESTINES WERE WELL MAINTAINED AT ALL TIMES. THE VENA CAVA WAS WELL FILLED AT ALL TIMES, SO THE RETURN FLOW TO THE HEART WAS MORE VIGOROUS THAN INSUFFICIENT. THE CUSTOMER STATED THAT THE ORIGIN OF THE ASYSTOLE REMAINS UNEXPLAINED. A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE FIVE SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. Manufacturer Narrative: THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 0 DEGREE ENDOSCOPE PLUS, CADIERE FORCEPS, FORCE BIPOLAR, MONOPOLAR CURVED SCISSORS, SYNCHROSEAL. A SITE HISTORY REVIEW FOUND NO RELEVANT COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE INSTRUMENTS USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. REVIEW OF THE ENERGY LOGS FOR THE SYNCHROSEAL INSTRUMENT FOUND 10 COAGULATION, 45 SEAL AND 9 TRANSECT EVENTS WITH NO ASSOCIATED ERRORS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT DIED DURING A LIVER RESECTION REPORTEDLY FOR A LIVER CYST. THERE WAS SOME BLEEDING DURING THE CASE BUT THAT BLEEDING WAS REPORTEDLY CONTROLLED. ASYSTOLE DEVELOPED FOR REASONS THAT ARE UNCLEAR. THERE IS NO ALLEGATION AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION A2 PATIENT AGE: THE PATIENT'S REPORTED YEAR OF BIRTH WAS 1988. ESTIMATED DATE OF BIRTH USED IS (B)(6) 1988.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22165276&pc=DTB | 2649622-2025-15411 | 2025-04-05 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-06-09 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Low impedance | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED APPROXIMATELY TWO MONTHS POST IMPLANT THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED LOW IMPEDANCE. THE RA LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS DECEASED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT'S DEATH WAS NOT LINKED TO THE PERFORMANCE ISSUE OF THE IPG SYSTEM AND THE PATIENT WAS MADE COMFORT MEASURES. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22165634&pc=OZD | 1220648-2025-29106 | 2025-05-13 04:00 | Death | ABIOMED, INC. | 2025-06-09 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT WAS ON IMPELLA CP SUPPORT HAD NO PULSES TO THE LOWER EXTREMITY. DUPLEX STUDY WAS NEGATIVE FOR FLOW. A DISTAL PERFUSION CATHETER WAS PLACED. HOWEVER, THERE WAS STILL NO FLOW. ADDITIONALLY, UPON INTUBATION, THE PATIENT BEGAN TO DROP FLOWS, HAVING SUCTION AT EVERY P LEVEL, REQUIRING REDUCTION TO P3. THE NEXT DAY, EXTRACORPOREAL MEMBRANE OXYGENATION CIRCUIT WITH CHATTER AND IMPELLA WITH SUCTION WAS NOTED. ALBUMIN WAS ADMINISTERED AND HYPOTENSION, SUCTION, AND CHATTER IMPROVED. IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: LIMB ISCHEMIA - IRREVERSIBLE: CLINICAL DETAILS STATE THAT THERE WERE NO PULSES TO LOWER EXTREMITY AS OF (B)(6) 2025. DUPLEX STUDY NEGATIVE FOR FLOW. DPC PLACED HOWEVER STILL NO FLOW. LIMB ISCHEMIA CAN OCCUR DURING IMPELLA SUPPORT. WHEN MAKING A DETERMINATION AS TO THE CAUSE OF THE LIMB ISCHEMIA, THE FOLLOWING CLINICAL INFORMATION IS NECESSARY: ¿ PATIENTS PRE-MORBID CONDITION AND PERI-PROCEDURAL CONDITION. ¿ ANY CONCOMITANT MEDICATION. ¿ VASCULAR SIZE OR VARIATIONS THEREOF. ¿ DETAILED DESCRIPTION OF THE SYMPTOMS EXPERIENCED BY THE PATIENT. ¿ PHYSICAL EXAMINATION FINDINGS, INCLUDING PULSE ASSESSMENT AND VISUAL EXAMINATION OF THE AFFECTED LIMB. ¿ DIAGNOSTIC IMAGING RESULTS, SUCH AS DOPPLER ULTRASOUND, ANGIOGRAPHY, OR MAGNETIC RESONANCE ANGIOGRAPHY. ¿ LABORATORY TEST RESULTS, INCLUDING BLOOD TESTS FOR MARKERS OF INFLAMMATION OR CLOTTING DISORDERS. ¿ ANY RELEVANT INFORMATION ABOUT RISK FACTORS FOR PERIPHERAL ARTERIAL DISEASE, SUCH AS SMOKING, DIABETES, OR HYPERTENSION. ¿ PATIENT'S RESPONSE TO PREVIOUS TREATMENTS OR INTERVENTIONS FOR VASCULAR CONDITIONS WITH A RETURNED IMPELLA, THE MAX OD COULD BE ASSESSED TO ENSURE IT IS WITHIN SPECIFICATION. THE PRODUCT COULD ALSO BE EVALUATED FOR ANY BURRS OR SHARP EDGES. ADDITIONALLY, THE CLINICAL INFORMATION ABOVE WOULD NEED TO BE CONSIDERED IN THE CONTEXT OF THE RETURNED PRODUCT. TO DETERMINE THE TRUE ROOT CAUSE OF LIMB ISCHEMIA, THE CLINICAL INFORMATION MENTIONED ABOVE WOULD NEED TO BE ASSESSED IN CONJUNCTION WITH EVIDENCE FROM THE RETURNED DEVICE. AS ALL THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS NOT DETERMINED. LOW PUMP FLOW: CLINICAL DETAILS STATE THAT ON (B)(6) 2025 UPON ARRIVING TO ICU AND BEING INTUBATED, PATIENT BEGAN TO DROP FLOWS, HAVING SUCTION AT EVERY P LEVEL, REQUIRING REDUCTION TO P3. ECHO AT BEDSIDE, RV AKINETIC, LV WITH AN EF OF APPROX 7%. ON (B)(6) 2025, ECMO CIRCUIT HAD CHATTER AND IMPELLA HAD SUCTION. ALBUMIN ADMINISTERED AND HYPOTENSION, SUCTION, AND CHATTER IMPROVED. TTE AT BEDSIDE SHOWING UNDERFILLED LV, ADEQUATE IMPELLA POSITIONING. DATA LOG REVIEW SHOWED A SET OF SUCTION ALARMS ON (B)(6) 2025 WITH VARYING FLOWS WITH THE SEVERAL CHANGES IN P-LEVEL. THERE WAS ALSO A GENERAL PS TREND DOWNWARD ON THIS DAY. THE NEXT DAY, SUCTION ALARMS ALIGNED WITH DROPS IN PLACEMENT SIGNAL. PRODUCT WAS NOT RETURNED. THE ROOT CAUSE OF THE LOW FLOW AND SUCTION WAS MOST LIKELY PATIENT CONDITION SINCE CLINICAL DETAILS STATE THE PATIENT HAD AN UNDERFILLED LV AND DATA LOG REVIEW SHOWS PATIENT DETERIORATION TRENDS IN PS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22165854&pc=MIH | 9680654-2025-00027 | 2025-01-27 05:00 | Death | WILLIAM A. COOK AUSTRALIA, PTY LTD | 2025-06-09 04:00 | MIH | ZENITH FENESTRATED AAA ENDOVASCULAR GRAFT PROXIMAL BODY | Migration |  | P020018 |  | 1 | Event Description: THIS (B)(6): PATIENT EXPIRED.TYPE 3C ENDOLEAK WAS DISCOVERED ON IMAGING REVIEW FOR THE RELATED (B)(6). SEE (B)(6) FOR PATIENT EXPIRED.TYPE 3B (TEARS) IN GRAFT MATERIAL (ORIGINALLY REPORTED TO BE GRAFT SEPARATION). Manufacturer Narrative: THE GRAFT REMAINED IN SITU AND THEREFORE WAS NOT RETURNED FOR EVALUATION. HOWEVER, IMAGES WERE PROVIDED AND REVIEWED BY CLINICAL PERSONNEL, AND THE FOLLOWING WAS DETERMINED "THIS CASE SUPPOSEDLY INVOLVES A ZFEN (P AND D) THAT WAS IMPLANTED IN 2022, WITH WHAT APPEARS TO BE PREEXISTING ALPHA THORACIC GRAFTS IN THE THORACIC AORTA. THE IMAGING WE HAVE IS FROM 2023 AND 2024, BUT ON BOTH OF THOSE FOLLOW-UP CT SCANS, THERE IS NO SIGN OF A BIFURCATED COMPONENT LIKE A ZFEN-D. BUT THE SITE REPORTS SEPARATION OF THE ZFEN-P AND THE ZFEN-D. I CAN¿T SEE A ZFEN-D THERE AT ALL, AND THE LARGE ENDOLEAKS ARE HIGHER THAN THAT. THE PATIENT APPARENTLY HAD A SIGNIFICANT FALL IN (B)(6) 2024, REQUIRING HOSPITALISATION, AND SUBSEQUENTLY DIED WITH LARGE ENDOLEAKS FROM THE ENDOGRAFT. THE CT SCAN FROM (B)(6) 2024 SHOWS AT LEAST THREE LARGE TYPE 3 ENDOLEAKS: 1. LARGE ENDOLEAK FROM THE REGION OF THE TOP OF THE OVERLAP BETWEEN THE FENESTRATED GRAFT AND THE OUTER ALPHA THORACIC GRAFT. THIS MAY BE A FABRIC TEAR DUE TO THE BARBS OF THE FENESTRATED GRAFT HAVING TORN THE FABRIC OF THE ALPHA GRAFT ¿ JUST MY SPECULATION BUT WOULD EXPLAIN THIS LARGE ENDOLEAK. IT IS HARD TO SEE AND CONFIRM FROM THE CT SCAN THAT THERE WERE INDEED FIXATION BARBS ON THE FENESTRATED GRAFT. 2. SLIGHTLY LESS VOLUME ENDOLEAK FROM THE REGION OF THE 2ND STENT DOWN THE FENESTRATED GRAFT. 3. SIGNIFICANT ENDOLEAK FROM THE SMA FENESTRATION. ENDOLEAK 1 WOULD BE A TYPE 3B (FABRIC TEAR). ENDOLEAK 2 ALSO PROBABLY A TYPE 3B. ENDOLEAK 3 WOULD BE A TYPE 3C. GIVEN THAT I CAN¿T SEE ANY OBVIOUS GRAFT SEPARATION OF COMPONENTS APART FROM THE LOOSENING OF THE SMA STENT IN THE SMA FENESTRATION, I THINK THE MOST LIKELY EXPLANATION FOR THIS UNFORTUNATE CASE IS THAT THE PATIENT¿S FALL MAY HAVE CAUSED DOWNWARD (DISTAL) JOLTING MOVEMENT OF THE FENESTRATED GRAFT THAT IN TURN CAUSED FABRIC TEARS IN THE ALPHA THORACIC GRAFT FABRIC, AND THE TYPE 3B ENDOLEAK(S). SIMILARLY, THE SMA STENT MAY HAVE BEEN JOLTED LOOSE DUE TO THE MOVEMENT OF THE MAIN FENESTRATED GRAFT BODY, LEADING TO THE TYPE 3C ENDOLEAK FROM THE SMA FENESTRATION. I CAN¿T SEE ANY CONVINCING EVIDENCE OF GRAFT COMPONENT SEPARATION APART FROM THE LOOSE SMA STENT, AND I CAN¿T SEE ANY EVIDENCE OF A BIFURCATED GRAFT COMPONENT EG ZFEN-D. SUMMARY: LIKELY GRAFT FABRIC TEARS DUE TO PATIENT FALL, LEADING TO MULTIPLE LARGE ENDOLEAKS, AND PATIENT DEATH. NOT ANY INTRINSIC FAULT OR FAILURE OF ANY OF THE DEVICES". REVIEW OF THE DEVICE HISTORY RECORD FOUND THE WORK ORDER FOR LOT AC1100373 APPEARED COMPLETE AND THE QUALITY CONTROL INSPECTION WAS VERIFIED TO ENSURE THAT THE DEVICE PASSED INSPECTION. THERE WERE NO TEMPORARY DEVIATIONS OR NON-CONFORMANCES IN PLACE AT THE TIME OF MANUFACTURE. REVIEW OF SPECIFICATIONS FOUND THAT UPON REVIEWING THE PHOTOS TAKEN OF THE COMPLAINT DEVICE DURING MANUFACTURE IT APPEARS THAT THE DEVICE WAS MANUFACTURED AS PER SPECIFICATIONS. REVIEW OF INSTRUCTIONS FOR USE (IFU) AND LABELLING FOUND IT TO CONTAIN APPROPRIATE WARNINGS, PRECAUTIONS, AND INSTRUCTIONS TO THE USER. THERE IS NO EVIDENCE TO SUGGEST THAT THE USER DID NOT FOLLOW THE INSTRUCTIONS FOR USE. A REVIEW OF THE MANUFACTURING RECORDS AND SPECIFICATIONS DID NOT REVEAL ANY DISCREPANCIES THAT COULD HAVE CONTRIBUTED TO THE REPORTED ISSUE. THE INVESTIGATION CONCLUDED THAT THE COMPLAINT DEVICE WAS MANUFACTURED TO SPECIFICATION. ALTHOUGH A DEFINITIVE ROOT CAUSE WAS NOT ABLE TO BE DETERMINED, BASED ON THE MEDICAL DIRECTOR'S ASSESSMENT, THE MOST LIKELY ROOT CAUSE OF THE SIGNIFICANT ENDOLEAK FROM THE SMA FENESTRATION IS DUE TO A DOWNWARD (DISTAL) JOLTING MOVEMENT OF THE FENESTRATED GRAFT AS A RESULT OF THE PATIENT'S FALL, WHICH LEAD TO MULTIPLE LARGE ENDOLEAKS, AND PATIENT DEATH. AFTER CONSIDERING THIS EVENT THE RISK ASSOCIATED WITH THE USE OF THIS DEVICE IS STILL DEEMED ADEQUATE. THE APPROPRIATE PERSONNEL WILL CONTINUE TO MONITOR FOR SIMILAR COMPLAINTS. SHOULD ADDITIONAL INFORMATION BE RECEIVED AT ANY TIME IN THE FUTURE AN ADDITIONAL REPORT MAY BE SUBMITTED AT THAT TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22165913&pc=PYX | 1220648-2025-29107 | 2025-05-15 04:00 | Death | ABIOMED, INC. | 2025-06-09 04:00 | PYX | IMPELLA RP FLEX | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA RP FLEX FOR MECHANICAL CIRCULATORY SUPPORT. WHEN ATTEMPTING TO GAIN ACCESS FOR THE IMPELLA RP FLEX IMPLANT, THE STAFF ACCIDENTALLY HIT THE FEMORAL ARTERY TWICE WITH A 18 GAUGE NEEDLE. ACTIVATED CLOTTING TIME WAS 280 AT THAT TIME. ONCE ACCESS WAS OBTAINED, THE IMPELLA RP FLEX WAS PLACED WITHOUT ISSUE. HOWEVER, WHEN READY TO LEAVE THE CATHETERIZATION LAB, THE PATIENT'S BLOOD PRESSURE BEGAN TO DECREASE, AND A HEMATOMA WAS DISCOVERED VIA COMPUTED TOMOGRAPHY. THE PATIENT RECEIVED FOUR UNITS OF PACKED RED BLOOD CELLS/PLASMA/PLATELETS AND MANUAL PRESSURE WAS HELD FOR TWO HOURS TO MANAGE THE CONDITION. THE PATIENTS BLOOD PRESSURE STABILIZED. THE PATIENT PASSED AWAY AFTER SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA RP FLEX WITH SMARTASSIST INSTRUCTIONS FOR USE & CLINICAL REFERENCE MANUAL SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP FLEX SMARTASSIST SYSTEM CATHETER. USE OF THE REPOSITIONING SHEATH AND THE 23 FR PEEL-AWAY INTRODUCER ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE INTERNAL JUGULAR VEIN. CLOSE AND DRESS THE WOUND.¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ACCESS SITE BLEEDING: THE ROOT CAUSE OF ACCESS SITE BLEEDING IS DETERMINED TO BE USE ISSUE BECAUSE THE BLEEDING IS CAUSED FORM THE 2 ARTERIAL STICKS THAT HAPPENED DURING THE PROCEDURE AND MANUAL PRESSURE IS HELD FOR 2 HOURS TO CONTROL THE BLEEDING. VASCULAR DAMAGE- PERIPHERAL VESSEL: THE ROOT CAUSE OF VASCULAR DAMAGE IS DETERMINED TO BE USE ISSUE BECAUSE THE DOCTOR HIT THE FEMORAL ARTERY 2 TIMES WITH THE 18G NEEDLE DURING THE INSERTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22167162&pc=OWB | 3003768277-2025-005188 | 2025-05-09 04:00 | Death | PHILIPS MEDICAL SYSTEMS NEDERLAND B.V. | 2025-06-09 04:00 | OWB | AZURION | Radiation Output Failure | Insufficient Information | K163715 |  | 1 | Event Description: IT WAS REPORTED TO PHILIPS THAT DURING A STEMI (ST-ELEVATION MYOCARDIAL INFARCTION) PROCEDURE, IMAGING BECAME UNAVAILABLE. THE PATIENT WAS SUBSEQUENTLY MOVED TO ANOTHER ROOM. THE REPORT INDICATED THAT THE PATIENT PASSED AWAY; HOWEVER, DETAILED INFORMATION REGARDING THE DEATH HAS NOT BEEN PROVIDED TO DATE. AN INVESTIGATION INTO THIS COMPLAINT HAS BEEN INITIATED BY PHILIPS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22167312&pc=OZD | 1220648-2025-29118 | 2025-02-18 05:00 | Death | ABIOMED, INC. | 2025-06-09 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Hemorrhagic Stroke | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WENT INTO HEMORRHAGIC SHOCK AND PASSED AWAY DURING IMPELLA CP SUPPORT. IT WAS NOTED THAT PRIOR TO IMPELLA CP IMPLANT, A RETROPERITONEAL BLEED HAD DEVELOPED AT THE TIME OF EXTRACORPOREAL MEMBRANE OXYGENATION INSERTION. DESPITE THE ONGOING CONDITION, THE PATIENT WAS GIVEN ANTICOAGULANTS FOLLOWING IMPELLA CP IMPLANT LEADING TO THE HEMORRHAGIC SHOCK. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22167339&pc=NGV | 2135147-2025-03041 | 2025-05-15 04:00 | Death | ABBOTT MEDICAL | 2025-06-09 04:00 | NGV | AMPLATZER AMULET | Migration or Expulsion of Device | Cardiac Arrest | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS CHOSEN FOR A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE USING A 12F AMPLATZER AMULET DELIVERY SHEATH. THE SIZING OF THE DEVICE WAS DETERMINED BY COMPUTED TOMOGRAPHY (CT) PRE-OPERATIVE, CONFIRMED WITH TRANSTHORACIC ECHOCARDIOGRAM (TEE) DURING IMPLANT. THE LEFT ATRIAL PRESSURE WAS ABOVE 12MMHG. THE PHYSICIAN TRIED TO IMPLANT THE AMULET DEEPER, BUT IT WAS NOT POSSIBLE TO STAY IN THE ANTERO-SUPERIOR LOB WITH A GOOD AXIS. THE AMULET WAS IMPLANTED BUT THERE WAS A SMALL CONTACT BETWEEN DISC AND LOB AND THE DEVICE WAS WELL COMPRESSED, WITH ROMAN HELMET SHAPE. AFTER THE PROCEDURE, THE PATIENT SUFFERED A CARDIAC ARREST AND RESUSCITATED. A TTE PERFORMED AND SHOWED THE AMULET WAS EMBOLIZED IN THE LEFT VENTRICLE. THERE WAS NO INTERVENTION PERFORMED DUE TO THE AGE OF THE PATIENT. THE PATIENT WAS REPORTED IN THE INTENSIVE CARE UNIT AND PASSED AWAY. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: IT WAS REPORTED THAT THE PATIENT SUFFERED A CARDIAC ARREST AND RESUSCITATED AFTER THE AMULET IMPLANT PROCEDURE. A TTE SHOWED THE AMULET HAD EMBOLIZED IN THE LEFT VENTRICLE. THERE WAS NO INTERVENTION PERFORMED DUE TO THE AGE OF THE PATIENT. REPORTEDLY, THE PATIENT PASSED AWAY IN THE INTENSIVE CARE UNIT. THE DEVICE REMAINED IMPLANTED AND WILL NOT RETURN TO ABBOTT FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. STILL ECHOCARDIOGRAM IMAGES AND FLURO IMAGES PROVIDED BY THE FIELD WERE REVIEWED. THERE WERE REPORTED DIFFICULTIES WITH LOBE ANCHORING AND A LARGE FLARED LEFT ATRIAL APPENDAGE (LAA) OSTIUM WAS NOTED ON CT IMAGING. BASED ON THE INFORMATION AVAILABLE, THE EXACT CAUSE OF THE REPORTED EVENT COULD NOT BE CONCLUSIVELY DETERMINED. IT IS POSSIBLE THAT THE OPERATIONAL DIFFICULTIES (DIFFICULTIES WITH LOBE ANCHORING) MAY HAVE CONTRIBUTED TO THE DEVICE EMBOLIZATION. THE REPORTED CARDIAC ARREST AND PATIENT DEATH APPEARS TO BE A CASCADING EFFECT. THERE WERE NO COMPLAINTS ASSOCIATED WITH ANY OTHER DEVICES FROM THE LOT. BASED ON THE INFORMATION RECEIVED, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING, DESIGN, OR MANUFACTURING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22167359&pc=GDW | 3005075853-2025-04293 | 2025-04-17 04:00 | Death | ETHICON ENDO-SURGERY, LLC. | 2025-06-09 04:00 | GDW | 75MM SELECTABLE NEW TLC | Patient Device Interaction Problem | Insufficient Information | K092577 |  | 1 | Event Description: IT WAS REPORTED THAT POST OPERATION TO A COLOSTOMY CLOSING PROCEDURE THE PATIENT DIED 5 DAYS AFTER USING THE STAPLER AND TWO CHARGES. IN CONVERSATION WITH THE SURGEON, HE INFORMED ME THAT HE WAS A VERY SERIOUS PATIENT, WITH SEVERAL COMORBIDITIES AND CHRONIC RENAL. HOWEVER, THE SAME WAS REPROACHED APPROXIMATELY 5 DAYS AFTER SURGERY, AND IT WAS NOTICED THAT IN THE STAPLING LINE THERE WAS A DESCENT. Manufacturer Narrative: (B)(4), DATE SENT 6/9/2025, D4 BATCH # UNK. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: HOW MANY NTLC¿S WERE USED? HOW MANY RELOADS WERE USED (WHAT WAS THE PRODUCT CODE)? WHAT WERE THE INDICATIONS FOR SURGERY? WHAT SURGICAL PROCEDURE WAS PERFORMED? WHAT COLOR SETTING WAS SELECTED ON THE DEVICE AND WHAT WAS THE SEQUENCE OF THE FIRINGS? WHAT ANATOMICAL STRUCTURES WAS THE DEVICE FIRED ON? WHAT DEVICE WAS USED TO CREATE THE COMMON CHANNEL? WHAT WAS USED TO CLOSE THE COMMON OPENING? WERE THERE ANY ISSUES EXPERIENCED WITH THE DEVICE IN THE INITIAL SURGICAL PROCEDURE? WAS THE DEVICE DIFFICULT TO FIRE? HOW WERE THE POST-OPERATIVE ISSUES IDENTIFIED? WERE THERE ANY ISSUES NOTED WITH STAPLE FORMATION AT ANY POINT THROUGHOUT THE CARE OF THE PATIENT? IF SO, PLEASE DESCRIBE. DOES THE SURGEON BELIEVE THAT THE ALLEGED DIFFICULTIES WITH THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENTS DEATH? IS THERE AN AUTOPSY REPORT AVAILABLE? IS THE SURGEON INTERESTED IN SPEAKING WITH ETHICON MEDICAL AND ENGINEERING STAFF? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22167364&pc=DYB | 9612164-2025-02808 |  | Death | MEDTRONIC IRELAND | 2025-06-09 04:00 | DYB | MICRA | Adverse Event Without Identified Device or Use Problem | Endocarditis | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IT NOTED THAT THE PATIENT EXPERIENCED ENDOCARDITIS AND VEGETATION. THE PATIENT IS A PARTICIPANT IN A POST APPROVAL CLINICAL SURVEILLANCE PRODUCT SURVEILLANCE REGISTRY.Event Description: IT WAS FURTHER NOTED THAT THE DEATH IS NOT RELATED TO THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) OR PROCEDURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22167371&pc=OZD | 1220648-2025-29119 | 2025-05-13 04:00 | Death | ABIOMED, INC. | 2025-06-09 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels; Implant Pain | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, ON THE DAY AFTER IMPLANT THE PATIENT'S LIMB BECAME NUMB/PARESTHESIA DESPITE INTACT PULSES. LIMB WAS NOT ISCHEMIC. THE TEAM OPENED THE AXILLARY INCISION BEDSIDE AND REMOVED MATERIAL THAT MAY HAVE BEEN CAUSING THE COMPRESSION. THE PUMP REMAINED ON FOR SUPPORT FOR APPROXIMATELY TWO WEEKS WHEN THE PATIENT EXPIRED AFTER SEPSIS FROM HERNIA SURGERY AND CARE/COMFORT MEASURES. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED PAIN AT INSERTION SITE AND VASCULAR DAMAGE PERIPHERAL VESSEL ISSUE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AND THE INSUFFICIENT CLINICAL DETAILS PROVIDED WERE NOT SUFFICIENT TO DETERMINE A ROOT CAUSE. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING: SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22167373&pc=PNJ | 9612164-2025-02807 |  | Death | MEDTRONIC IRELAND | 2025-06-09 04:00 | PNJ | MICRA AV | Adverse Event Without Identified Device or Use Problem | Endocarditis | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IT NOTED THAT THE PATIENT EXPERIENCED ENDOCARDITIS AND VEGETATION. THE PATIENT IS A PARTICIPANT IN A POST APPROVAL CLINICAL SURVEILLANCE PRODUCT SURVEILLANCE REGISTRY.Event Description: IT WAS FURTHER NOTED THAT THE DEATH IS NOT RELATED TO THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) OR PROCEDURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID: [MC1AVR1-DELSYS] (SERIAL: [(B)(6)]). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22167580&pc=OHR | 1417592-2025-00281 | 2025-05-15 04:00 | Death | MEDLINE INDUSTRIES LP | 2025-06-09 04:00 | OHR | 1LYR TRY 16FR 10ML 100% SIL UM SNAP | Appropriate Term/Code Not Available | Unspecified Infection; Urinary Tract Infection |  |  | 1 | Event Description: STOOL INTRUSION UNDER SEAL OF CATHETER TO DRAINAGE BAG CONNECTION. Manufacturer Narrative: IT WAS REPORTED THAT A PATIENT WITH AN EXISTING CAUTI WITH FOLEY CATHETER HAD "STOOL INTRUSION" IN THE CATHETER. THE PATIENT REPORTEDLY HAD "SUDDEN DECLINE IN PATIENT STATUS AND SEVERE SEPSIS, REQUIRED PRESSORS AND ULTIMATELY PASSED AWAY". IT WAS REPORTED THAT THE PATIENT "HAD MULTIPLE MEDICAL ISSUES, AND SHE CAN'T SAY FOR SURE THAT THE PATIENT'S DEATH WAS DUE TO CAUTI". IN AN ABUNDANCE OF CAUTION, THIS MEDWATCH IS BEING FILED. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22167686&pc=NGV | 2124215-2025-37307 | 2025-05-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-09 04:00 | NGV | WATCHMAN FLX LEFT ATRIAL APPENDAGE CLOSURE DEVICE WITH DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Pericardial Effusion | P130013 |  | 1 | Event Description: IT WAS REPORTED CARDIAC TAMPONADE AND DEATH OCCURRED. A PULSE FIELD ABLATION (PFA) PROCEDURE WAS PERFORMED WITH A FARAWAVE ABLATION CATHETER WITHOUT ANY COMPLICATIONS. SUBSEQUENTLY THE SAME DAY, A CONCOMITANT NON-BOSTON SCIENTIFIC (BSC) TRANSCATHETER MITRAL VALVE REPAIR (TMVR) WAS IMMEDIATELY FOLLOWED BY A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE USING TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) AND FLUOROSCOPY IMAGING INTRAPROCEDURE. A SMALL PERICARDIAL EFFUSION (PE) WAS NOTED PRIOR TO THE PROCEDURE ON IMAGING, WITH LOCATION AND SIZE UNKNOWN. DURING THE LAA CLOSURE PROCEDURE, A WATCHMAN FLX CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WAS INSERTED, AND A PE IN AN UNKNOWN LOCATION OCCURRED DURING THE PROCEDURE THAT LED TO CARDIAC TAMPONADE THAT WAS PRESUMABLY DUE TO POSITIONING MANEUVERS OF THE WDS. A PERICARDIOCENTESIS WAS PERFORMED. THE WATCHMAN CLOSURE DEVICE WAS IMPLANTED DESPITE THE PE, WITH THE PHYSICIANS HOPING IT WOULD SOLVE THE EVENT. THE PATIENT RECEIVED A BLOOD TRANSFUSION. THE PATIENT DIED THE SAME DAY OF THE INDEX PROCEDURE. THE OFFICIAL CAUSE OF DEATH WAS CARDIAC TAMPONADE. IT WAS NOTED BY THE PHYSICIANS THAT THE INTRAPROCEDURE PE OCCURRED AFTER THE REMOVAL OF THE NON-BSC TMVR DELIVERY SYSTEM, DURING OR AFTER THE INSERTION OF THE WDS. Manufacturer Narrative: D4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UDI AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22167961&pc=DSQ | 2916596-2025-03447 | 2021-10-22 04:00 | Death | THORATEC CORPORATION | 2025-06-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Dyspnea; Lethargy; Heart Failure/Congestive Heart Failure; Swelling/ Edema | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CARDIAC ARREST RELATED TO RIGHT SIDED HEART FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED.Event Description: RIGHT HEART FAILURE EXISTED PRE - LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT, BUT THE SITE COULD NOT FULLY IDENTIFY. A DEVICE RELATED ISSUE DID NOT CAUSE OR CONTRIBUTE TO THE RIGHT HEART FAILURE. THE PATIENT HAD SHORTNESS OF BREATH, LETHARGY AND SWELLING. THE PATIENT HAD MEDICAL OPTIMIZATION AND MECHANICAL VENTILATOR. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE¿) STATES THAT PATIENTS MAY DEVELOP RIGHT VENTRICULAR FAILURE DURING OR SHORTLY AFTER IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THIS SECTION (UNDER ¿CAUTION!¿) ALSO EXPLAINS THAT RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22168051&pc=MKJ | 1220908-2025-02053 | 2025-05-18 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-06-09 04:00 | MKJ | X SERIES | Unexpected Shutdown | Cardiac Arrest | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO PACE A PATIENT DURING A BRADYCARDIA EVENT (AGE & GENDER UNKNOWN), THE DEVICE INAPPROPRIATELY SHUT DOWN. AS THE PATIENT WAS BEING LOADED INTO THE BACK OF THE AMBULANCE, THE PATIENT WENT INTO CARDIAC ARREST. THE CLINICIAN PERFORMED RESUSCITATION EFFORTS TO CONTINUE TO TREAT THE PATIENT. THE COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: ZOLL MEDICAL CORPORATION EVALUATED THE DEVICE, AND THE CUSTOMER'S REPORT WAS NOT REPLICATED OR CONFIRMED. THE DEVICE WAS PUT THROUGH EXTENSIVE TESTING, WHICH INCLUDED BENCH HANDLING AND FUNCTIONAL TESTING WITHOUT DUPLICATING THE CUSTOMER'S REPORT. AN INTERNAL INSPECTION OF THE POWER CABLES AND CONNECTIONS WITHIN THE DEVICE FOUND NO DISCREPANCIES. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22168150&pc=OZD | 1220648-2025-29121 | 2025-05-15 04:00 | Death | ABIOMED, INC. | 2025-06-09 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT AND HAD A CEREBRAL HEMORRHAGE AND GRADUALLY LOST THE ABILITY TO MAINTAIN THEIR VITAL SIGNS. THE DOCTOR DID BELIEVE THE IMPELLA COULD HAVE BEEN A FACTOR. THE PATIENT EXPIRED DUE TO THE CEREBRAL HEMORRHAGE. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22168585&pc=OZD | 1220648-2025-29125 | 2025-05-12 04:00 | Death | ABIOMED, INC. | 2025-06-09 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED 5.5 PUMP SUPPORT ONGOING FOR A PATIENT ADMITTED IN WITH CARDIOGENIC SHOCK. THE PUMP HAD PURGE FLOW/PRESSURE ISSUES THAT WERE REMEDIED BY THE ADDITION OF TISSUE PLASMINOGEN ACTIVATOR (TPA) LYSIS TO THE PURGE FLUID. PUMP REMAINED ON TO DATE DESPITE MALFUNCTION. PATIENT LATER EXPIRED WHEN CARE WAS WITHDRAWN AFTER 11 DAYS OF SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF HIGH PURGE PRESSURE HAS BEEN COMPLETED. THE PUMP WAS RETURNED FOR INVESTIGATION. NO PHYSICAL DAMAGE WAS OBSERVED ON THE RETURNED PUMP. A SIGNIFICANT AMOUNT OF BIOMATERIAL WAS NOTED ON THE OUTFLOW CAGE AND IMPELLER. THE BIOMATERIAL WAS FLUSHED OUT DURING THE START OF THE TEST RUN, AND THE PUMP WAS RUN ACCORDING TO SPECIFICATIONS IN A BUCKET OF WATER. THERE WERE NO ISSUES NOTED WITH THE PURGE PARAMETERS DURING THE TEST RUN. THE DATA LOG SHOWS A SLOW RISE IN PURGE PRESSURE FROM 500 MMHG TO 700 MMHG OVER A 20-HOUR PERIOD DURING THE CASE ON THE 12TH. NO PURGE-RELATED ALARMS WERE TRIGGERED DURING THIS TIME. PURGE FLOW WAS SEEN RESOLVED WITH GRADUAL DECREASE TREND. AS PER THE CLINICAL, LYSIS PROTOCOL APPLIED DURING PURGE PRESSURE RISE EVENT. THE CAUSE OF THE PURGE PRESSURE ISSUE WAS MOST LIKELY BIOMATERIAL BUILDUP, BASED ON DATA LOG REVIEW AND RETURNED PRODUCT INVESTIGATION. Manufacturer Narrative: THE IMPELLA DEVICE HAS BEEN RECEIVED BUT HAS NOT YET BEEN EVALUATED. WHEN THE DEVICE INVESTIGATION HAS BEEN COMPLETED OR ANY NEW INFORMATION HAS BEEN RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE PATIENT'S OUTCOME WAS REVIEWED BY ABIOMED'S MEDICAL SAFETY OFFICER (MSO) AND IT WAS DETERMINED THAT THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMARTASSIST STATES THE FOLLOWING: SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22168783&pc=DSQ | 2916596-2025-03388 | 2025-05-10 04:00 | Death | THORATEC CORPORATION | 2025-06-09 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS/MULTI-ORGAN FAILURE. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT BE EXPLANTED.Event Description: IT WAS REPORTED THAT THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, SEPSIS, AND DEATH. ALTHOUGH ONLY SEPSIS WAS REPORTED BY THE ACCOUNT, SEVERAL SECTIONS OF THE HEARTMATE 3 LVAS IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22168875&pc=DSQ | 2916596-2025-03399 | 2025-04-18 04:00 | Death | THORATEC CORPORATION | 2025-06-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE WAS UNKNOWN.Event Description: IT WAS REPORTED THAT SUPPORT WAS WITHDRAWN WHILE THE PATIENT WAS ON HOSPICE. THE DEATH WAS EXPECTED AND PEACEFUL. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. THE DEVICE WOULD NOT BE RETURNED AS IT WAS SENT WITH THE PATIENT HOME AND THEN TAKEN TO A NEARBY VENTRICULAR ASSIST DEVICE (VAD) CLINIC. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22168919&pc=DSQ | 2916596-2025-03784 | 2024-11-02 04:00 | Death | THORATEC CORPORATION | 2025-06-09 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. NO FURTHER INFORMATION WAS KNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22169060&pc=OZD | 1220648-2025-29128 | 2025-05-14 04:00 | Death | ABIOMED, INC. | 2025-06-09 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT EXPERIENCED RUNS VENTRICULAR TACHYCARDIA REQUIRING INTERVENTION AND WOULD SUBSEQUENTLY EXPIRE. THE PATIENT WAS ADMITTED WITH PLAN FOR SCHEDULE PERCUTANEOUS CORONARY INTERVENTION (PCI) AND DECOMPENSATED THROUGHOUT THE DAY. THE DECISION WAS MADE TO PLACE AN IMPELLA CP AFTER THE PULMONARY CAPILLARY WEDGE PRESSURE OF 30 MMHG. CARDIAC INDEX OF 1.31 L/MIN/M2 WORSENING LACTATE AND PH. DURING INSERTION THE PATIENT PULSELESS ELECTRICAL ACTIVITY (PEA) ARRESTED. COMPRESSIONS INITIATED AND THE PATIENT WAS INTUBATED. PERSISTENT VENTRICULAR TACHYCARDIA/VENTRICULAR FIBRILLATION WAS EXPERIENCED REQUIRING DEFIBRILLATION ALONG WITH ADVANCE CARDIOVASCULAR LIFE SUPPORT MEDICATION. THE DECISION WAS MADE TO FURTHER ESCALATED TO EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). THE PATIENT WAS EVENTUALLY STABILIZED AND RETURNED TO THE UNIT ON SUPPORTS. LATER THAT EVENING, THE PATIENT EXPERIENCED INTERMITTENT VENTRICULAR TACHYCARDIA. TREATING WITH MEDS AND DEFIBRILLATION. THE PATIENT CONTINUED ON SUPPORT WITH UNNOTED COMPLICATIONS. ON DAY SEVEN OF THE SUPPORT, THE PATIENT WAS NOTED TO HAVE EXPIRED. CARE WAS WITHDRAWN. Manufacturer Narrative: ALTHOUGH THE DECISION WAS MADE TO DISCONTINUE IMPELLA MECHANICAL CIRCULATORY SUPPORT, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR TO THE OVERALL OUTCOME (GIVEN THE ARRHYTHMIC EVENTS). THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. HOWEVER, AN INVESTIGATION WAS COMPLETED AND NOTED THE FOLLOWING: CARDIAC ARRYTHMIA CAN BE REPORTED DURING IMPELLA SUPPORT AND TO MAKE A DETERMINATION AS TO THE CAUSE OF THE CARDIAC ARRYTHMIA, THE FOLLOWING CLINICAL INFORMATION MUST BE CONSIDERED: ¿ PATIENT'S MEDICAL HISTORY, INCLUDING ANY PRE-EXISTING CARDIAC CONDITIONS, PREVIOUS EPISODES OF ARRHYTHMIAS, OR STRUCTURAL HEART DISEASE ¿ TIMING OF THE ARRYTHMIA EVENT IN RELATION TO IMPELLA SUPPORT (DURING OR POST-IMPLANT) ¿ ELECTROCARDIOGRAM (EKG) RECORDINGS OF THE ARRHYTHMIA EPISODES ¿ EVALUATION OF ANY UNDERLYING ELECTRICAL DISTURBANCES, SUCH AS QT PROLONGATION OR ELECTROLYTE IMBALANCES ¿ ASSESSMENT OF HEMODYNAMIC INSTABILITY DURING THE ARRHYTHMIA, INCLUDING BLOOD PRESSURE AND CARDIAC OUTPUT MEASUREMENTS ¿ PRESENCE OF OTHER POTENTIALLY CONTRIBUTORY FACTORS, SUCH AS MEDICATION CHANGES, ISCHEMIA, OR ELECTROLYTE ABNORMALITIES ¿ RESPONSE TO ANY ANTIARRHYTHMIC TREATMENTS OR INTERVENTIONS DURING THE EVENT ¿ ANY DOCUMENTED DEVICE-RELATED ISSUES OR COMPLICATIONS DURING IMPELLA SUPPORT ¿ EVALUATION OF POTENTIAL TRIGGERS OR PRECIPITATING FACTORS, SUCH AS CATHETER MANIPULATION, REPERFUSION INJURY, OR INCREASED MYOCARDIAL OXYGEN DEMAND ¿ REVIEW OF ANY CONCOMITANT MEDICATIONS THAT MAY INFLUENCE CARDIAC ELECTROPHYSIOLOGY WITH A RETURNED IMPELLA, THE DEVICE CAN BE INSPECTED FOR ANY BURRS OR ROUGH EDGES THAT MAY CONTACT THE HEART WALL. THEREFORE, TO DETERMINE THE TRUE ROOT CAUSE OF THE VENTRICULAR TACHYCARDIA, THE CLINICAL INFORMATION MENTIONED ABOVE WOULD NEED TO BE ASSESSED IN CONJUNCTION WITH EVIDENCE FROM THE RETURNED DEVICE. IN CONCLUSION, AS THE NECESSARY CLINICAL INFORMATION WAS NOT PROVIDED AND THE DEVICE WAS NOT RETURNED, THE ROOT CAUSE OF THE VENTRICULAR TACHYCARDIA WAS NOT DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22169665&pc=LWP | 2017865-2025-85753 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-09 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED DUE TO UNKNOWN CAUSES. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO PATIENT DEATH. THE DEVICE WAS RECEIVED WITH A BATTERY LEVEL WITHIN THE NORMAL RANGE OF OPERATION. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. AFTER ALL ELECTRICAL TESTS PERFORMED, INCLUDING THERMAL AND MECHANICAL STRESS TESTING THE DEVICE EXHIBITS NORMAL CHARACTERISTICS. LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY. NO ANOMALIES WERE FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22169708&pc=NVN | 2017865-2025-85754 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-09 04:00 | NVN | TENDRIL ST | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED DUE TO UNKNOWN CAUSES. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGES. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES. ELECTRICAL SHORTING WAS FOUND BETWEEN CONDUCTORS AT THE CONNECTOR REGION CONSISTENT WITH PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22169732&pc=NVN | 2017865-2025-85755 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-09 04:00 | NVN | TENDRIL ST | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED DUE TO UNKNOWN CAUSES. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22169929&pc=OZD | 1220648-2025-29132 | 2025-05-17 04:00 | Death | ABIOMED, INC. | 2025-06-09 04:00 | OZD | IMPELLA 5.5 | Patient-Device Incompatibility | Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA 5.5 WAS TO BE IMPLANTED IN A PATIENT ADMITTED IN WITH MULTIPLE CARDIAC AND SYSTEMIC COMORBIDITIES. THE PUMP PLACEMENT/DELIVERY WAS COMPLICATED BY A WIRE INTERACTION. THE WIRE SNAGGED AND WAS DIFFICULT TO REMOVE BUT SUPPORT WAS PLACED. LATER, DURING PUMP SUPPORT, THE TEAM OBSERVED SIGNS OF HEMOLYSIS IN ELEVATED FREE HEMOGLOBIN. THE PUMP SUPPORTED THE PATIENT FOR 10 DAYS.Event Description: THE DECISION WAS MADE TO WITHDRAW CARE AND THE PATIENT EXPIRED. IT IS UNKNOWN IF THE USE OF THE IMPELLA WAS RELATED TO PATIENT DEATH, THUS THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA PUMP AND GUIDEWIRE WERE RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA 5.5 WITH SMARTASSIST: SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES AND HEMOLYSIS HAS BEEN COMPLETED. THROMBUS FAILURE MODE ADDED FOR BIOMATERIAL OBSERVED ON IMPELLER. THE RETURNED PRODUCT WAS INSPECTED AND BIOMATERIAL WAS OBSERVED BENEATH IMPELLER BLADES AND IN THE CANNULA. THE PUMP PASSED THE LIGHT CONTINUITY TEST AND THE 5S PARAMETERS WERE WITHIN NORMAL RANGE. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE THROMBUS WAS NOT DETERMINED DELIVERY ISSUES: CLINICAL MENTIONED GUIDEWIRE MET RESISTANCE DURING REMOVAL AND WAS SNAGGED. THE RETURNED GUIDEWIRE WAS INSPECTED AND DAMAGE WAS CONFIRMED. THE ROOT CAUSE OF THE DELIVERY ISSUES WAS MOST LIKELY USE RELATED AS EVIDENCED BY THE DAMAGE ON THE RETURNED PRODUCT AND THE CLINICAL MENTION OF RESISTANCE MET. HEMOLYSIS: BIOMATERIAL WAS OBSERVED WRAPPED AROUND THE IMPELLER BLADES AND IN THE CANNULA. CLINICAL REPORTED HEMOLYSIS ON (B)(6). THE DATA LOGS DO NOT SHOW ANY MOTOR CURRENT SPIKES BEFORE REPORTED HEMOLYSIS. THERE WERE NO FREQUENT OR CONTINUOUS SUCTION ALARMS OR POSITION RELATED ALARMS. CLINICAL DETAILS ALSO MENTION PUMP POSITION WAS CONFIRMED VIA IMAGING. THE ROOT CAUSE OF THE HEMOLYSIS WAS MOST LIKELY BIOMATERIAL AS EVIDENCED BY THE BIOMATERIAL WRAPPED AROUND THE IMPELLER ON THE RETURNED PRODUCT B2 DEATH AND DATE OF DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29132. B5 PATIENT OUTCOMES WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29132. H1 TYPE OF REPORT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29132. H6 HEALTH EFFECT - IMPACT CODE 1802 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29132.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22169936&pc=LWR | 9612164-2025-02816 | 2025-03-26 04:00 | Death | MEDTRONIC MEXICO | 2025-06-09 04:00 | LWR | AVALUS AORTIC TISSUE VALVE | Insufficient Information | Pulmonary Edema; Insufficient Information | P170006 |  | 1 | Event Description: IT WAS REPORTED THAT POST IMPLANT OF THIS 23MM AORTIC BIOPROSTHETIC VALVE, THE PATIENT DIED. IT WAS STATED THAT THE PATIENT COULD HAVE DIED FROM HEART FAILURE. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMEDEvent Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE OF DEATH WAS REPORTED AS PULMONARY EDEMA. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5 AND H6 UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22169959&pc=MHX | 9610816-2025-000472 | 2025-05-14 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-06-09 04:00 | MHX | PATIENT INFORMATION CENTER IX | No Audible Alarm | Insufficient Information | K211900 |  | 1 | Event Description: IT WAS REPORTED NO ALARMS WERE GENERATED ON THE NURSE PHONES AND THE PATIENT PASSED AWAY. IT WAS INDICATED RED ALARMS WOULD BE DISPLAYED BRIEFLY ON THE NURSE PHONE BUT IT APPEARS THERE WAS NO AUDIBLE ALARM. Manufacturer Narrative: E1 REPORTING INSTITUTION PHONE#: (B)(6). IT WAS ALSO INDICATED THE DEVICE (ASCOM PHONE) WAS NOT USABLE AND A REPLACEMENT WAS REQUESTED. THE CUSTOMER ALLEGED THAT AN ASYSTOLE ALARM WAS MISSED. THE PATIENT UNDERWENT RESUSCITATION MEASURES, HOWEVER THE PATIENT PASSED AWAY. A PHILIPS PROJECT MANAGER AND APPLICATION SPECIALIST WERE CONTACTED. IT WAS NOTED THAT THE CONFIGURATION IS SET FOR DELAYED PHONE NOTIFICATION, HOWEVER THE ALLEGATION INDICATES THAT THE MESSAGE WAS RECEIVED (VERY BRIEFLY) AND DISAPPEARED. A PHILIPS REMOTE SERVICE ENGINEER (RSE) SPOKE WITH THE CUSTOMER AND NOTED THAT THE EVENT TOOK PLACE ON (B)(6) 2025 AT APPROXIMATELY 19:50 HOURS. THE ONSITE TECHNICIAN TRIGGERED A RED ALARM AND THIS WAS RECEIVED AT THE END DEVICE (MYCO 3 SMARTPHONE) WITH CAREASSIST 4.4.0. THE TEST WAS SUCCESSFUL. THE SOFTWARE WAS PROACTIVELY UPDATED TO VERSION 4.3.7. A PHILIPS FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED AND THE PRELIMINARY INVESTIGATION AND FOUND THAT STARTING AT 19:55:07, AN ALARM WAS GENERATED AT THE PATIENT INFORMATION CENTER IX (PIC IX) INDICATING THAT THERE WAS NO MOBILE DEVICE AVAILABLE FOR ALERTING FOR THE AFFECTED BED LOCATION. FROM 19:55:15 AND 19:55:18, RED ALARMS WERE GENERATED ON THE MX40 MONITOR AND THE PICIX. THESE WERE ACKNOWLEDGED AT THE PICIX AT 19:55:28 AND 19:55:29, WHEREUPON THE ALARM WAS TERMINATED AT THE PICIX. ACCORDING TO THE CAREGIVER, THE NECESSARY ACTIONS WERE THEN IMMEDIATELY INITIATED FOR THE PATIENT. THE EARLY INVESTIGATION INDICATES THAT THE MOBILE PHONE WAS NOT IN USE. AS THIS IS CONTRADICTORY TO THE ORIGINAL ALLEGATION, QUALITY PERSONNEL ARE WORKING WITH THE FIELD ENGINEER TO CLARIFY THE EVENT. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22170350&pc=OUT | 2029214-2025-01337 | 2025-05-01 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-06-09 04:00 | OUT | PIPELINE | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P100018 |  | 1 | Event Description: PANDO, A., HA, C. J., THIBAULT, D., KHANDELWAL, P., NAIR, A., SINGLA, A., & SUN, H. (2025). FLOW DIVERTER ASSISTED EMBOLIZATION OF RUPTURED ANEURYSMS IS ASSOCIATED WITH INCREASED HEMORRHAGIC COMPLICATIONS: PROGNOSTIC FACTORS AND OUTCOMES IN NEUROENDOVASCULAR TREATMENT OF SUBARACHNOID HEMORRHAGES. WORLD NEUROSURGERY. HTTPS://DOI.ORG/10.1016/J.WNEU.2025.124061. MEDTRONIC REVIEW OF THE LITERATURE ARTICLE FOUND A RETROSPECTIVE REVIEW OF THE 2016 TO 2021 NATIONAL INPATIENT SAMPLE (NIS) DATABASE WHICH IDENTIFIED 62,567 PATIENTS WITH NON-TRAUMATIC SUBARACHNOID HEMORRHAGE (SAH). OF THESE PATIENTS, 409 UNDERWENT TREATMENT WITH PIPELINE EMBOLIZATION DEVICE FLOW DIVERTER STENT IMPLANTATION. 6,834 PATIENTS UNDERWENT COIL EMBOLIZATION (MANUFACTURER, BRAND NOT SPECIFIED). THE REMAINDER OF THE PATIENTS UNDERWENT NON-ENDOVASCULAR MANAGEMENT. IT SHOULD BE NOTED THAT THIS TREATMENT IS OFF-LABEL FOR THE PIPELINE DEVICE. NO DEVICE MALFUNCTION WAS REPORTED IN THE ARTICLE, 43 PATIENTS IN THE PIPELINE TREATMENT GROUP HAD AN OUTCOME OF MORTALITY. CAUSE OF DEATH AND/OR RELATIONSHIP TO THE PIPELINE OR THE PROCEDURE WAS NOT SPECIFIED IN THE ARTICLE. Manufacturer Narrative: A2. REPORTED PATIENT AGE (55 YEARS) IS THE MEAN AGE FOR ALL PATIENTS IN THE PIPELINE TREATMENT GROUP IN THE ARTICLE. A3A. REPORTED PATIENT SEX (FEMALE) REPRESENTS THE MAJORITY OF PATIENTS INCLUDED IN THE PIPELINE TREATMENT GROUP, AND THE STUDY OVERALL. B3. REPORTED EVENT DATE IS THE DATE THE ARTICLE WAS ACCEPTED FOR PUBLICATION. SECTION E: THE INITIAL REPORTER INFORMATION AND ADDRESS PERTAINS TO THE COMMUNICATING AUTHOR. THE EVENTS REPORTED IN THE ARTICLE WERE FROM THE NATIONAL INPATIENT SAMPLE (NIS) WHICH INVOLVED MANY FACILITIES ACROSS THE UNITED STATES. THIS REPORT IS SUBMITTED REPRESENTATIVE OF PATIENTS WHO HAD OUTCOME OF MORTALITY REPORTED IN THE ARTICLE. SERIOUS INJURIES THAT DID NOT NECESSARILY END IN DEATH WILL BE REPORTED SEPARATELY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22170394&pc=MVK | 3008642652-2025-05629 | 2025-03-10 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-09 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 18:50:55 ON (B)(6) 2025. AT 20:11:03 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. AT 20:11:40, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH TACTILE ARTIFACT. AT 20:14:05, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY. AT 20:15:40, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT ELECTRODE LEAD FALLOFF. THE ELECTRODE BELT WAS DISCONNECTED AT 20:54:36 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22170476&pc=MVK | 3008642652-2025-05631 | 2025-05-16 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-09 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 08:58:57 ON (B)(6) 2025. AT 11:49:55, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 11:51:37, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 11:53:26 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22170829&pc=DSQ | 2916596-2025-03409 | 2025-05-11 04:00 | Death | THORATEC CORPORATION | 2025-06-09 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE WAS NOT PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. ADDITIONAL INFORMATION WAS UNKNOWN.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE OPERATED AS EXPECTED. IT WAS COMMUNICATED THAT THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22170979&pc=NIQ | 2024168-2025-05945 | 2024-06-10 04:00 | Death | ABBOTT VASCULAR INC. | 2025-06-09 04:00 | NIQ | XIENCE SKYPOINT¿ | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Vascular Dissection | P110019 |  | 1 | Event Description: PATIENT ID: (B)(6). ON (B)(6) 2024 PERCUTANEOUS CORONARY INTERVENTION (PCI) WAS PERFORMED TO TREAT LESIONS IN THE LEFT MAIN TRUNK (LMT) AND LEFT ANTERIOR DESCENDING (LAD) ARTERIES WITH MILD TORTUOSITY AND MODERATE CALCIFICATION. PRIOR TO STENT IMPLANTATION, THE LESIONS WERE TREATED WITH INTRAVASCULAR LITHOTRIPSY (IVL) AND PRE-DILATED WITH A NON-COMPLIANT BALLOON. THEN, A 2.25X18MM, 3X33MM AND 3.5X28MM XIENCE SKYPOINT STENTS WERE IMPLANTED FOLLOWED BY POST DILATATION. AT SOME POINT DURING THE PROCEDURE A CORONARY ARTERY DISSECTION OCCURRED; HOWEVER WAS COVERED WITH STENTS. THEN, ON (B)(6) 2024 THE PATIENT EXPERIENCED DIFFICULTY WALKING AND WAS HOSPITALIZED. AN MRI REVEALED ISCHEMIC STROKE. ON (B)(6) 2024 THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL, WHERE ON (B)(6) 2024 AN ABLATION PROCEDURE WAS PERFORMED DURING WHICH REPORTED BLEEDING OCCURRED IN THE RIGHT FEMORAL PUNCTURE SITE. THE PATIENT WAS DISCHARGED A FEW DAYS LATER. DURING THE 1 YEAR FOLLOW-UP CALL IT WAS REPORTED THAT ON (B)(6) 2025 THE PATIENT DIED DURING HOSPITALIZATION. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT(S), AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED; HOWEVER, THE SUBSEQUENT TREATMENT APPEARS TO BE RELATED TO THE OPERATIONAL CONTEXT OF THE PROCEDURE. THE REPORTED PATIENT EFFECT(S) OF DEATH, STROKE AND DISSECTION ARE LISTED IN THE XIENCE SKYPOINT EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22171147&pc=LWS | 2649622-2025-15537 | 2024-09-23 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-06-09 04:00 | LWS | SPRINT QUATTRO SECURE MRI SURESCAN | Impedance Problem; High Capture Threshold | No Clinical Signs, Symptoms or Conditions | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED ALERTS DUE TO OUT OF RANGE (OOR) PACING IMPEDANCE, DEFIBRILLATION IMPEDANCE, AND SUPERIOR VENA CAVA (SVC) IMPEDANCE MEASUREMENTS. IT WAS ALSO NOTED THAT THE RV LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA). IT WAS ALSO NOTED THAT THE RIGHT ATRIAL (RA) LEAD TRIGGERED AN ALERT DUE TO OUT OF RANGE (OOR) PACING IMPEDANCE. THE RV LEAD AND RA LEAD WERE EXPLANTED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THE PATIENT PASSED AWAY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. THE ANALYST NOTED ON THE PROXIMAL PORTION OF THE WAS RETURNED WITH THE IS-1 LEG MEASURING 3.8 CM LONG, THE RV-DF1 LEG MEASURING 3.5 CM LONG, AND THE SVC-DF1 LEG MEASURING 3 CM LONG. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD, IMPLANTED: (B)(6) 2012 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22171160&pc=DTB | 2649622-2025-15539 | 2024-09-23 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-06-09 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Impedance Problem | No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED ALERTS DUE TO OUT OF RANGE (OOR) PACING IMPEDANCE, DEFIBRILLATION IMPEDANCE, AND SUPERIOR VENA CAVA (SVC) IMPEDANCE MEASUREMENTS. IT WAS ALSO NOTED THAT THE RV LEAD TRIGGERED A LEAD INTEGRITY ALERT (LIA). IT WAS ALSO NOTED THAT THE RIGHT ATRIAL (RA) LEAD TRIGGERED AN ALERT DUE TO OUT OF RANGE (OOR) PACING IMPEDANCE. THE RV LEAD AND RA LEAD WERE EXPLANTED. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: 694765 LEAD IMPLANTED: (B)(6) 2010. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. THE ANALYST NOTED A PROXIMAL SEGMENT OF THE LEAD 3.4 CM LONG WAS RETURNED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22171544&pc=MFA | 3007284006-2025-00113 | 2025-06-02 04:00 | Death | SPECTRANETICS CORPORATION | 2025-06-09 04:00 | MFA | GLIDELIGHT LASER SHEATH | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels | P960042 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE TWO RIGHT ATRIAL (RA) AND TWO LEFT VENTRICULAR (LV) LEADS DUE TO FRACTURE. SPECTRANETICS LLD LEAD LOCKING DEVICES (LLDS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. WHILE OPENING OF THE POCKET, ONE OF THE RA LEADS STARTED TO COME APART, BUT DID NOT CUT TO MAINTAIN ITS STRUCTURE. BEGINNING WITH A SPECTRANETICS 14F GLIDELIGHT LASER SHEATH, THE TWO LV LEADS, AND ONE RA LEAD WERE REMOVED SUCCESSFULLY. DURING EXTRACTION OF THE LAST RA LEAD, WHEN THE GLIDELIGHT WAS BEING USED AS A MECHANICAL SHEATH, IT WAS NOTICED THAT THE TIP WAS BENT. HOWEVER, EXTRACTION CONTINUED BUT THE LEAD BROKE, AND THE LLD, ALONG WITH A PORTION OF THE LEAD WERE REMOVED. THE GLIDELIGHT WAS REMOVED FROM THE BODY, AND THE BLOOD PRESSURE DROPPED. RESCUE EFFORTS BEGAN, INCLUDING RESCUE BALLOON, STERNOTOMY, AND CELL SAVER. AFTER SOME TIME, A PERFORATION IN THE SUBCLAVIAN VEIN WAS DISCOVERED, WHICH WAS ORIGINALLY THOUGHT TO HAVE BEEN IN THE SUPERIOR VENA CAVA (SVC). THE PATIENT WAS TRANSFERRED TO THE OPERATING ROOM, REPAIR WAS COMPLETED, AND THE PATIENT STABILIZED; UNFORTUNATELY, DID NOT SURVIVE THE PROCEDURE. THIS REPORT CAPTURES THE GLIDELIGHT IN USE WHEN THE SUBCLAVIAN PERFORATION OCCURRED, REQUIRING INTERVENTION, BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF THE GLIDELIGHT DEVICE IN USE DURING THE PROCEDURE. Manufacturer Narrative: A2) PATIENT DATE OF BIRTH - NOT AVAILABLE FROM FACILITY. A3B) PATIENT GENDER - NOT AVAILABLE FROM FACILITY. D4) DEVICE SERIAL NUMBER - NOT AVAILABLE FROM FACILITY. H3) THE GLIDELIGHT WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) PERFORATION OF VESSELS AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE GLIDELIGHT DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22171580&pc=BZD | 2518422-2025-106485 | 2024-12-01 05:00 | Death | RESPIRONICS, INC. | 2025-06-09 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A USER OF A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE REPORTS THE PATIENT WAS NOT USING THE DEVICE AT THE TIME OF PASSING. THE PATIENT HAS NOT USED THE DEVICE FOR A VERY LONG TIME. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE IS NO ALLEGATION THAT THE DEVICE CAUSED THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22172117&pc=LWP | 3006705815-2025-04242 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-09 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS CONGESTIVE HEART FAILURE. Manufacturer Narrative: CORRECTION: SECTION D - MANUFACTURER INFORMATION WAS CORRECTED. THE MANUFACTURER INFORMATION SHOULD HAVE BEEN "ST. JUDE MEDICAL, INC.(CRM-SYLMAR)" INSTEAD OF "ST. JUDE MEDICAL - NEUROMODULATION (PUERTO RICO, LLC)" WHICH WAS SELECTED IN ERROR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22172135&pc=NVN | 2017865-2025-85778 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-09 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS CONGESTIVE HEART FAILURE. Manufacturer Narrative: AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22172191&pc=NVN | 2017865-2025-85780 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-09 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS CONGESTIVE HEART FAILURE. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR THE CUT END WHICH IS CONSISTENT WITH PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22172381&pc=PNJ | 2017865-2025-85748 | 2025-05-09 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-09 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension | P150035 |  | 1 | Event Description: VOLUNTARY MEDWATCH RECEIVED STATES THE PATIENT PRESENTED FOR IMPLANT PROCEDURE. DURING SURGERY, THE PATIENT'S BLOOD PRESSURE BEGAN TO DROP. OPEN CHEST SURGERY AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS PERFORMED. THE PATIENT DECEASED. Manufacturer Narrative: D4 - PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER CANNOT BE PROVIDED AS A PRODUCT IDENTIFIER COULD NOT BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22172387&pc=PNJ | 2017865-2025-85751 | 2025-05-09 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-09 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension | P150035 |  | 1 | Event Description: VOLUNTARY MEDWATCH RECEIVED STATES THE PATIENT PRESENTED FOR IMPLANT PROCEDURE. DURING SURGERY, THE PATIENT'S BLOOD PRESSURE BEGAN TO DROP. OPEN CHEST SURGERY AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS PERFORMED. THE PATIENT DECEASED. Manufacturer Narrative: D4 - PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER CANNOT BE PROVIDED AS A PRODUCT IDENTIFIER COULD NOT BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22164401&pc=BZD | 2518422-2025-106450 | 2025-06-04 04:00 | Death | RESPIRONICS, INC. | 2025-06-08 04:00 | BZD | DREAMSTATION | Degraded | Cancer; Unspecified Gastrointestinal Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING TROUBLE SWALLOWING, ESOPHAGEAL CANCER, AND DEATH. THE PATIENT HAD REQUIRED CHEMOTHERAPY AND WAS ON FEEDING TUBES. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22162538&pc=QFG | 3013756811-2025-107457 | 2025-05-18 04:00 | Death | TANDEM DIABETES CARE | 2025-06-07 04:00 | QFG | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | Adverse Event Without Identified Device or Use Problem | Hypoglycemia | K232380 |  | 1 | Event Description: IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. PER THE CONTACT, THE CUSTOMER WAS WEARING THE PUMP AT THE TIME OF DEATH. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS A LOW BLOOD GLUCOSE LEVEL. NO ADDITIONAL INFORMATION WAS PROVIDED. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. Manufacturer Narrative: THE DEVICE IS EXPECTED TO BE RETURNED; HOWEVER, THE DEVICE HAS NOT YET BEEN RECEIVED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF THE DEVICE IS RECEIVED AND EVALUATION IS PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22162622&pc=DSQ | 2916596-2025-03553 | 2024-11-14 05:00 | Death | THORATEC CORPORATION | 2025-06-07 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN UNKNOWN ETIOLOGY. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE DEATH WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. IT WAS COMMUNICATED THAT DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WILL NOT COMMUNICATE ANY FURTHER PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22162817&pc=OAE | 2029046-2025-01849 | 2025-05-15 04:00 | Death | BIOSENSE WEBSTER INC | 2025-06-07 04:00 | OAE | QDOT MICRO¿ CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Ventricular Fibrillation; Stenosis | P210027 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN IDIOPATHIC VENTRICULAR TACHYCARDIA / PVC (PREMATURE VENTRICULAR CONTRACTION) ABLATION AND THE PATIENT EXPERIENCED LEFT CORONARY ARTERY OCCLUSION / VENTRICULAR FIBRILLATION THAT REQUIRED DEFIBRILLATION, CPR (CARDIOPULMONARY RESUSCITATION), IMPELLA PLACEMENT AND REVASCULARIZATION OF CORONARY ARTERY. THE PATIENT DIED. DURING THE CASE A LEFT CORONARY ARTERY OCCLUSION WAS NOTICED AS THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (V-FIB). THE PATIENT WAS SHOCKED MULTIPLE TIMES TO RESTORE A RHYTHM, AND THE CERTIFIED REGISTERED NURSE ANESTHETIST (CRNA) WAS LOOKING INTO ONE OF THE PATIENT OXYGENATION WAVE FORMS AND NOTED POOR OXYGENATION LEVELS. THE INTERVENTIONAL CARDIOLOGIST ON CALL WAS CALLED AND THEY IMAGED THE CORONARY ARTERIES AND FOUND THE LEFT MAIN WAS OCCLUDED. DEFIBRILLATION AND CPR WERE PERFORMED TO THE PATIENT DURING THIS TIME. THEY WERE ABLE TO OPEN AND RESTORE THE CORONARY ARTERY AND REGAIN BLOOD FLOW. THE CALLER NOTED THAT ANOTHER PHYSICIAN CAME IN AND PLACED AN IMPELLA HEART PUMP FOR VENTRICULAR ASSISTANCE. THE PATIENT IS IN STABLE CONDITION AND WAS TRANSFERRED TO CCU (CARDIAC CARE UNIT) FOR OBSERVATION. THE PATIENT DIED. PATIENT'S RELEVANT HISTORY / PRE-EXISTING CONDITION INCLUDES: UNIFOCAL PVC RESULTING IN VF (VENTRICULAR FIBRILLATION) AND SUBSEQUENT SHOCK FROM ICD (IMPLANTABLE CARDIOVERTER DEFIBRILLATOR). THE DETAILS ABOUT THE DEATH INCLUDE: ACCESS WAS OBTAINED USING ULTRASOUND GUIDANCE. WE USED ICE (INTRACARDIAC ECHOCARDIOGRAPHY) TO TAKE CONTOURS OF THE LV (LEFT VENTRICLE), PAPILLARY MUSCLES, AORTA, AND ALL CORONARY CUSPS (INCLUDING THE LEFT MAIN). HE POSITIONED HIS CATHETER IN THE LEFT CORONARY CUSP, BESIDE THE LEFT MAIN AND PROCEEDED TO COME ON ABLATION. WE BURNED FOR 9 MINS AND 12 SECONDS. THE FORCE WAS IN WITHIN NORMAL PARAMETERS. DURING THE LAST ABLATION, THE PATIENT WENT INTO V-FIB. LIFE SAVING MEASURES WERE PERFORMED INCLUDING CPR (CARDIOPULMONARY RESUSCITATION), IMPELLA PLACEMENT, AND PCI (PERCUTANEOUS CORONARY INTERVENTION) OF THE LAD (LEFT ANTERIOR DESCENDING ARTERY). THE PATIENT WAS TRANSPORTED TO CCU, HE CODED SEVERAL MORE TIMES BEFORE THE FAMILY CHANGED HIS STATUS TO DNR (DO NOT RESUSCITATE). SADLY, HE PASSED A FEW MINUTES LATER. Manufacturer Narrative: ON 19-JUN-2025, THE BWI PRODUCT ANALYSIS LAB RECEIVED THE DEVICE FOR EVALUATION. THE PRODUCT INVESTIGATION WAS SUBSEQUENTLY COMPLETED. IT WAS REPORTED THAT A PATIENT UNDERWENT AN IDIOPATHIC VENTRICULAR TACHYCARDIA / PVC (PREMATURE VENTRICULAR CONTRACTION) ABLATION AND THE PATIENT EXPERIENCED LEFT CORONARY ARTERY OCCLUSION / VENTRICULAR FIBRILLATION THAT REQUIRED DEFIBRILLATION, CPR (CARDIOPULMONARY RESUSCITATION), IMPELLA PLACEMENT AND REVASCULARIZATION OF CORONARY ARTERY. THE PATIENT DIED. DEVICE EVALUATION DETAILS: THE PRODUCT WAS RETURNED TO JOHNSON & JOHNSON MEDTECH (J&J MEDTECH) FOR EVALUATION. A VISUAL INSPECTION AND REVISION OF ALL FEATURES WERE PERFORMED FOLLOWING J&J MEDTECH PROCEDURES. VISUAL ANALYSIS REVEALED NO DAMAGE OR ANOMALIES ON THE DEVICE. THE DEVICE FEATURES WERE REVIEWED, AND NO ISSUES WERE OBSERVED DURING THE PRODUCT INVESTIGATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER 31543001L, AND NO INTERNAL ACTIONS RELATED TO THE REPORTED COMPLAINT WERE IDENTIFIED. NO MALFUNCTION WAS OBSERVED DURING THE PRODUCT ANALYSIS. THE ROOT CAUSE OF THE ADVERSE EVENT REMAINS UNKNOWN. IN ADDITION, NO DEVICE MALFUNCTION WAS REPORTED, THERE MAY HAVE BEEN OTHER CIRCUMSTANCES OR ISSUES THAT OCCURRED DURING THE USE OF THE DEVICE THAT COULD NOT BE REPLICATED DURING THE ANALYSIS. THE INSTRUCTION FOR USE (IFU) STATES THAT CAREFUL CATHETER MANIPULATION MUST BE PERFORMED TO AVOID CARDIAC DAMAGE, PERFORATION, OR TAMPONADE. ADDITIONALLY, THE LOT NUMBER BECAME AVAILABLE DUE TO THE DEVICE'S RETURN. AS A RESULT, THE D4 PRIMARY UDI NUMBER WAS UPDATED TO (B)(4). AS PART OF JOHNSON & JOHNSON MEDTECH¿S QUALITY PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC. OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22163372&pc=BZD | 2518422-2025-106445 | 2025-05-29 04:00 | Death | RESPIRONICS, INC. | 2025-06-07 04:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER HAS BEEN NOTIFIED THAT A PATIENT HAS PASSED AWAY. HOWEVER, THERE'S NO INDICATION THAT THE DEVICE WAS THE CAUSE OF DEATH. THE PATIENT'S DAUGHTER WAS GIVEN INSTRUCTIONS ON HOW TO RETURN THE RECALLED DEVICE AND ASKED ABOUT THE TIMING OF HER CHECK. ALTHOUGH THERE'S NO EVIDENCE TO SUGGEST THE DEVICE CAUSED THE PATIENT'S DEATH, MORE DETAILS ARE NEEDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED ABOUT THE REPORTED DEATH. IF WE RECEIVE ANY NEW INFORMATION, WE'LL FILE A FOLLOW-UP REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22163670&pc=OZP | 2032227-2025-187526 | 2025-05-05 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-07 04:00 | OZP | PUMP MMT-1884L MM780G BLE 3.0 MG/ML | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Myocardial Infarction | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2025, AND THE CAUSE OF DEATH WAS HEART ATTACK/STROKE. THE EVENT INVOLVED PRODUCTS MMT-7040A, MMT-1884L, MMT-242A, AND MMT-332A. TROUBLESHOOTING WAS PERFORMED, AND ANY HEALTH ISSUES OR ILLNESSES THAT MAY HAVE CONTRIBUTED OR LED UP TO PASSING WAS DIABETES AND SEVERAL STROKES. THE PUMP WAS WORN AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-7040A, MMT-242A, AND MMT-332A. MMT-1884L WAS REQUESTED, AND THE DEVICE WILL BE RETURNED. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22163941&pc=EOQ | 2955842-2025-24205 | 2025-05-13 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-06-07 04:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Low Oxygen Saturation | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT WENT INTO CARDIAC ARREST AND EXPIRED. WHILE BEING TRANSPORTED TO THE PROCEDURE ROOM, THE PATIENT EXPERIENCED OXYGEN DESATURATION. UPON INDUCTION OF ANESTHESIA, THE PATIENT DECOMPENSATED, EXHIBITING LOW HEART RATE, HYPOTENSION, AND FURTHER OXYGEN DESATURATION. THE ANESTHESIOLOGIST ATTEMPTED TO STABILIZE THE PATIENT WHILE THE PHYSICIAN INITIATED ION REGISTRATION; HOWEVER, THE PATIENT WENT INTO CARDIAC ARREST AND THE PROCEDURE WAS ABORTED. THE ION BIOPSY HAD NOT STARTED AT THIS POINT. ADVANCED CARDIAC LIFE SUPPORT (ACLS) PROTOCOL WAS ADMINISTERED BUT THE PATIENT EXPIRED. THE PHYSICIAN STATED THAT THE CARDIAC ARREST AND THE PATIENT'S OUTCOME WERE NOT RELATED TO THE ION SYSTEM, BUT RATHER TO THE PATIENT'S SIGNIFICANT UNDERLYING COMORBIDITIES, POTENTIALLY EXACERBATED BY ANESTHESIA, AND THAT A SIMILAR OUTCOME MIGHT HAVE OCCURRED WITH ANOTHER PROCEDURE MODALITY. THERE WAS NO DEVICE MALFUNCTION ASSOCIATED WITH THE EVENT. Manufacturer Narrative: SYSTEM LOGS ARE NOT AVAILABLE FOR REVIEW. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT WITH SIGNIFICANT COMORBIDITIES INCLUDING COPD WITH CHRONIC HYPOXIC RESPIRATORY FAILURE ON CHRONIC OXYGEN, CORONARY ARTERY DISEASE WITH PRIOR MI POST STENT, CHF, ATRIAL FLUTTER AND B-CELL LYMPHOMA UNDERWENT AN ION LUNG BIOPSY. DURING TRANSPORT TO THE PROCEDURE SUITE OXYGEN DESATURATION WAS NOTED. DURING INDUCTION WITH ANESTHESIA HYPOTENSION, BRADYCARDIA, AND OXYGEN DESATURATION WAS NOTED WITH CONTINUED DETERIORATION DESPITE ATTEMPTS TO STABILIZE BY ANESTHESIA DURING THE REGISTRATION PROCESS WHICH WAS ABORTED PRIOR TO NAVIGATION. NO BIOPSIES WERE PERFORMED. THE PATIENT ULTIMATELY FURTHER DECOMPENSATED INTO CARDIAC ARREST AND DIED DESPITE RESUSCITATIVE EFFORTS FOR 15 MINUTES INCLUDING ACLS. BASED ON THE AVAILABLE DATA THE EVENTS WERE PROCEDURE/ANESTHESIA RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22163945&pc=EOQ | 2955842-2025-24301 | 2025-05-14 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-06-07 04:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Insufficient Information | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE THE PATIENT DIED. THE BIOPSIED LESION WAS LOCATED IN THE RIGHT UPPER LOBE. AFTER UNDOCKING THE ION SYSTEM THE PATIENT CODED. CARDIOPULMONARY RESUSCITATION WAS INITIATED BUT WAS UNSUCCESSFUL. THERE WAS NO REPORT OF ANY ISSUES DURING THE PROCEDURE, AND NO MALFUNCTION OF AN ION SYSTEM, INSTRUMENT, OR ACCESSORY OCCURRED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE PHYSICIAN, NO RESPONSE HAS BEEN RECEIVED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. ADDITIONAL INFORMATION CAN BE FOUND IN THE FOLLOWING FIELDS: H11 AND B7. Manufacturer Narrative: SECTION B5: ADDITIONAL INFORMATION - THE ENDOLUMINAL SALES REPRESENTATIVE RECEIVED ADDITIONAL INFORMATION THAT THE LUNG LESION SIZE WAS UNDER 3 CM AND THERE WERE NO ISSUES WITH THE ION SYSTEM DURING THE PROCEDURE. THE PATIENT HAD A KNOWN CARDIAC HISTORY, WAS INFORMED OF THE RISKS PRIOR TO THE ION PROCEDURE, AND OPTED TO PROCEED WITH THE LUNG BIOPSY. CARDIOPULMONARY RESUSCITATION WAS INITIATED FOR THE CARDIAC ARREST AND WAS PERFORMED FOR 30 MINUTES. THE PATIENT MAY HAVE DIED FROM A CARDIAC RELATED EVENT. NO OTHER INFORMATION WAS AVAILABLE. A RE-REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) UPON RECEIPT OF ADDITIONAL INFORMATION CONCLUDED THAT THE PATIENT HAD A PRIOR UNSPECIFIED CARDIOVASCULAR HISTORY BUT DECIDED TO UNDERGO AN ION LUNG BIOPSY AFTER DISCUSSING THE POTENTIAL RISK INVOLVED. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE AND THE ION SYSTEM WAS UNDOCKED. THE PATIENT THEN CODED AND EFFORTS AT RESUSCITATION INCLUDING CPR FOR 30 MINUTES WERE UNSUCCESSFUL. THERE WAS NO REPORTED MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. BASED ON THE AVAILABLE DATA THE EVENT WAS POSSIBLY PROCEDURE RELATED IN THE SETTING OF KNOWN PRIOR CARDIOVASCULAR HISTORY. THERE IS NO ALLEGATION THAT A MALFUNCTION OF THE ION SYSTEM, INSTRUMENT, OR ACCESSORY OCCURRED. GIVEN THE AVAILABLE DATA THE EVENT WAS LIKELY DUE TO THE PROCEDURE UNDER GENERAL ANESTHESIA AND NOT ASSOCIATED WITH THE ION SYSTEM, INSTRUMENTS, OR ACCESSORIES. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%), 5 ARRHYTHMIAS (0.02%) AND 1 MYOCARDIAL INFARCTION (0.004%). ONE PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%) AND NO ASSOCIATED EVENTS OF MYOCARDIAL INFARCTION OR ARRHYTHMIAS. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% INCLUDING A RATE OF 0.02% OF ANY ARRHYTHMIA AND 1 DEATH. ANOTHER CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO CARDIAC EVENTS. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. Manufacturer Narrative: SYSTEM LOGS ARE NOT AVAILABLE FOR REVIEW. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES THAT WOULD BE RELATED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT UNDERWENT AN ION LUNG BIOPSY. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE AND THE ION SYSTEM WAS UNDOCKED. THE PATIENT THEN CODED AND EFFORTS AT RESUSCITATION WERE UNSUCCESSFUL. EFFORTS AT OBTAINING FURTHER DATA HAVE BEEN UNSUCCESSFUL. BASED ON THE AVAILABLE DATA THE EVENTS WERE POSSIBLY PROCEDURE RELATED. THERE WAS NO REPORTED MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. SECTION A2 PATIENT AGE: THE PATIENT'S AGE WAS REPORTED AS IN THEIR "EARLY 80'S" YRS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22163984&pc=LYJ | 1644487-2025-10234 | 2023-05-01 04:00 | Death | HOUSTON | 2025-06-07 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. PER THE PHYSICIAN, THE PATIENT DIED SUDDENLY; EXACT CAUSE OF DEATH WAS UNKNOWN. THE PHYSICIAN NOTED THAT AT THE LAST RECORDED INTERROGATION, NO PROBLEMS WERE SEEN. NO ADDITIONAL RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22156107&pc=OZO | 2032227-2025-186889 | 2025-02-26 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-06 04:00 | OZO | 630G INSULIN PUMP MMT-1715K 630G BLACK MG | Adverse Event Without Identified Device or Use Problem | Hypoglycemia; Respiratory Failure; Cancer | P150001 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY AT THE HOSPITAL ON (B)(6) 2025. THE CUSTOMER WAS ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS HYPOGLYCEMIA, RESPIRATORY FAILURE AND CANCER WAS HOSPITALIZED. THE CUSTOMER¿S BLOOD GLUCOSE WAS 22 MG/DL. THE EVENT INVOLVED PRODUCT(S) MMT-1715K. THE CUSTOMER WAS NOT WEARING THE INSULIN PUMP AT THE TIME OF DEATH. THE CUSTOMER WAS NOT USING SENSORS. THE INSULIN PUMP WAS LAST WORN ON (B)(6) 2025, PRIOR TO PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-1715K. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22156764&pc=QSK | 9612515-2025-00056 | 2025-05-13 04:00 | Death | VASCUTEK LTD. | 2025-06-06 04:00 | QSK | THORAFLEX HYBRID | Fluid/Blood Leak; Device Stenosis | Hemorrhage/Bleeding; Stenosis | P210006 |  | 1 | Event Description: EVENT DATE: 13 MAY 2025. IMPLANT DATE: (B)(6) 2025. EVENT REPORTED FROM (B)(6). THORAFLEX ANTE-FLOW DEVICE IMPLANTED IN A 51-YEAR-OLD MALE PATIENT WITH CHRONIC TYPE A AORTIC DISSECTION (CTAD) AND CAD W/ BYPASS AT TIME OF ARCH REPAIR. AFTER THE PROCEDURE THE DEVICE APPEARED NARROW USING TEE (UNDERSTOOD TO BE TRANSOPHAGEAL ECHOCARDIOGRAM) AT THE ANASTOMOSIS OF THE THORAFLEX COLLAR. IVUS (UNDERSTOOD TO BE INTRAVASCULAR ULTRASOUND) WAS USED AND AN AORTIC BALLOON WAS INSERTED THROUGH THE PERFUSION ARM OF THE THORAFLEX DEVICE, THE NARROWING WAS BALLOONED AND UNDER IVUS VISUALIZATION IT WAS EXPANDED. HOWEVER, IT DID NOT STAY OPEN. THE SURGEON DECIDED TO BEGIN CLOSING AND WOULD PLAN ON COMING BACK AND EXPANDING WITH A BX (UNDERSTOOD TO BE BALLOON- EXPANDABLE) STENT. HOWEVER, THE PATIENT WAS BLEEDING FROM AN INDETERMINATE SOURCE AND THEY REMAINED IN THE ROOM TRYING TO IDENTIFY THE LOCATION. THEY DECIDED TO STOP AND TRANSPORT THE PATIENT TO RECOVER WHERE HE PASSED. IT'S UNCERTAIN IF THE NARROWING WAS A TECHNICAL OR DEVICE RELATED COMPLICATION.Event Description: EVENT DATE : (B)(6) 2025. IMPLANT DATE (B)(6) 2025. EVENT REPORTED FROM (B)(6). THORAFLEX ANTE-FLOW DEVICE IMPLANTED IN A 51 YEAR OLD MALE PATIENT WITH CHRONIC TYPE A AORTIC DISSECTION (CTAD) AND CAD W/ BYPASS AT TIME OF ARCH REPAIR. AFTER THE PROCEDURE THE DEVICE APPEARED NARROW USING TEE( UNDERSTOOD TO BE TRANSOPHAGEAL ECHOCARDIOGRAM) AT THE ANASTOMOSIS OF THE THORAFLEX COLLAR. IVUS ( UNDERSTOOD TO BE INTRAVASCULAR ULTRASOUND) WAS USED AND AN AORTIC BALLOON WAS INSERTED THROUGH THE PERFUSION ARM OF THE THORAFLEX DEVICE, THE NARROWING WAS BALLOONED AND UNDER IVUS VISUALIZATION IT WAS EXPANDED. HOWEVER, IT DID NOT STAY OPEN. THE SURGEON DECIDED TO BEGIN CLOSING AND WOULD PLAN ON COMING BACK AND EXPANDING WITH A BX (UNDERSTOOD TO BE BALLOON- EXPANDABLE ) STENT. HOWEVER, THE PATIENT WAS BLEEDING FROM AN INDETERMINATE SOURCE AND THEY REMAINED IN THE ROOM TRYING TO IDENTIFY THE LOCATION. THEY DECIDED TO STOP AND TRANSPORT THE PATIENT TO RECOVER WHERE HE PASSED. IT'S UNCERTAIN IF THE NARROWING WAS A TECHNICAL OR DEVICE RELATED COMPLICATION. THIS REPORT IS BEING SUBMITTED AS FOLLOW UP #1 FOR MANUFACTURING REPORT NUMBER 9612515-2025-00056 TO PROVIDE EVENT CLOSURE INFORMATION FOR (B)(4). Manufacturer Narrative: MANUFACTURER NARRATIVE: CLINICAL CODE: 1888 - HAEMORRHAGE/ BLEEDING - BLOOD LEAKAGE FORM INDETERMINATE SOURCE . 2263 - STENOSIS - THORAFLEX HYBRID DEVICE APPEARED NARROW USING TEE AT THE ANASTOMOSIS OF THE THORAFLEX COLLAR. IMPACT CODE: 1802 - DEATH - PATIENT PASSED AWAY (B)(6) 2025 NO AUTOPSY WAS PERFORMED. DEVICE PROBLEM CODE : 4066 - DEVICE STENOSIS - AN AORTIC BALLOON WAS INSERTED THROUGH THE PERFUSION ARM OF THE THORAFLEX DEVICE, THE NARROWING WAS BALLOONED AND UNDER IVUS (UNDERSTOOD TO BE ) VISUALIZATION IT WAS EXPANDED. HOWEVER, IT DID NOT STAY OPEN. 1250 -FLUID /BLOOD LEAK - THE PATIENT WAS BLEEDING FROM AN INDETERMINATE SOURCE AND THEY REMAINED IN THE ROOM TRYING TO IDENTIFY THE LOCATION. THEY DECIDED TO STOP AND TRANSPORT THE PATIENT TO RECOVER WHERE HE PASSED. THE LOCATION OF THE BLOOD LEAKAGE WAS NOT DETERMINED. COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4110 - TREND ANALYSIS: A 4-YEAR REVIEW OF SIMILAR COMPLAINTS THORAFLEX HYBRID/ GRAFT DILATION JAN 21 - MAY 25 SHOWED THIS IS THE FIRST EVENT REPORTED. NO TREND REQUIRING ACTION HAS BEEN IDENTIFIED. 4111 - COMMUNICATION INTERVIEW: ADDITIONAL INFORMATION - THE PATIENT'S AORTIC ANATOMY SHOWED NO SIGNIFICANT CURVATURE OR OTHER NOTABLE ANATOMICAL FEATURES WITHIN THE DISTAL LANDING ZONE. THE TRUE LUMEN DIAMETER AT THE COLLAR ANASTOMOSIS SITE WAS ESTIMATED AT 15MM, WHILE THE OVERALL AORTIC DIAMETER WAS 26-29MM IN THE DISTAL LANDING ZONE AND 27MM AT THE COLLAR. A PLANNED OVERSIZE OF 10-15% WAS IMPLEMENTED DURING THE PROCEDURE. HOWEVER, THE THORAFLEX HYBRID DEVICE WAS NOT INTENDED TO REMODEL THE TRUE LUMEN. THE VESSEL WALL EXHIBITED NO SIGNS OF CALCIFICATION AND WAS NOT OTHERWISE COMPROMISED. THE LOCATION OF LEAKAGE WAS NOT DETERMINED. ADDITIONALLY, AN AUTOPSY WAS NOT PERFORMED ON THIS PATIENT. THE DEVICE REMAINED IMPLANTED, RENDERING IT UNAVAILABLE FOR ADDITIONAL INVESTIGATION. SCANS FOR REVIEW WERE PROVIDED. 3331 - ANALYSIS OF PRODUCTION RECORDS: FULL BATCH REVIEW PERFORMED FROM BASE FABRIC TO FINISHED PRODUCT WHICH CONFIRMED THE BATCH WAS MANUFACTURED TO SPECIFICATION. 4117 - DEVICE NOT RETURNED: DEVICE REMAINED IMPLANTED. 4112 - INFORMATION PROVIDED BY 3RD PARTY - CANS WERE PROVIDED FOR REVIEW. INVESTIGATION FINDINGS: 213 - NO DEVICE PROBLEM FOUND - FULL BATCH REVIEW FROM BASE FABRIC TO FINISHED PRODUCT WAS PERFORMED WHICH CONFIRMED THE DEVICE WAS MANUFACTURED TO SPECIFICATION, ADDITIONAL INFORMATION RECEIVED FROM THE SITE STATED THE DEVICE DEPLOYED AND FUNCTIONED AS EXPECTED WITHOUT ANY APPARENT ABNORMALITIES. PRE OP SCANS WERE REVIEWED ON 23RD JUN 2025, THE PRE OP SCANS DID NOT DEMONSTRATE ANY CHALLENGES THAT COULD RESULT IN NARROWING OF THE DEVICE OR THE LEAKAGE. INVESTIGATION CONCLUSION: 4310 - CAUSE TRACED TO NON - DEVICE RELATED FACTORS - ROOT CAUSE WAS DEEMED TO BE PROCEDURE / TECHNICAL RELATED. - LOCATION OF THE BLOOD LEAK WAS NOT FOUND/ STOPPED. Manufacturer Narrative: MANUFACTURER NARRATIVE: CLINICAL CODE: 1888 - HAEMORRHAGE/ BLEEDING - BLOOD LEAKAGE FORM INDETERMINATE SOURCE. 2263 - STENOSIS - THORAFLEX HYBRID DEVICE APPEARED NARROW USING TEE AT THE ANASTOMOSIS OF THE THORAFLEX COLLAR. IMPACT CODE: 1802 - DEATH - PATIENT PASSED AWAY 13 MAY 2025 NO AUTOPSY WAS PERFORMED. DEVICE PROBLEM CODE: 4066 - DEVICE STENOSIS - AN AORTIC BALLOON WAS INSERTED THROUGH THE PERFUSION ARM OF THE THORAFLEX DEVICE, THE NARROWING WAS BALLOONED AND UNDER IVUS (UNDERSTOOD TO BE) VISUALIZATION IT WAS EXPANDED. HOWEVER, IT DID NOT STAY OPEN. 1250 -FLUID /BLOOD LEAK - THE PATIENT WAS BLEEDING FROM AN INDETERMINATE SOURCE AND THEY REMAINED IN THE ROOM TRYING TO IDENTIFY THE LOCATION. THEY DECIDED TO STOP AND TRANSPORT THE PATIENT TO RECOVER WHERE HE PASSED. THE LOCATION OF THE BLOOD LEAKAGE WAS NOT DETERMINED. COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. TYPE OF INVESTIGATION: 4110 - TREND ANALYSIS: A 4-YEAR REVIEW OF SIMILAR COMPLAINTS THORAFLEX HYBRID/ GRAFT DILATION JAN 2021 - MAY 2025 SHOWED THIS IS THE FIRST EVENT REPORTED. NO TREND REQUIRING ACTION HAS BEEN IDENTIFIED. 4111 - COMMUNICATION INTERVIEW: ADDITIONAL INFORMATION - THE PATIENT'S AORTIC ANATOMY SHOWED NO SIGNIFICANT CURVATURE OR OTHER NOTABLE ANATOMICAL FEATURES WITHIN THE DISTAL LANDING ZONE. THE TRUE LUMEN DIAMETER AT THE COLLAR ANASTOMOSIS SITE WAS ESTIMATED AT 15MM, WHILE THE OVERALL AORTIC DIAMETER WAS 26-29MM IN THE DISTAL LANDING ZONE AND 27MM AT THE COLLAR. A PLANNED OVERSIZE OF 10-15% WAS IMPLEMENTED DURING THE PROCEDURE. HOWEVER, THE THORAFLEX HYBRID DEVICE WAS NOT INTENDED TO REMODEL THE TRUE LUMEN. THE VESSEL WALL EXHIBITED NO SIGNS OF CALCIFICATION AND WAS NOT OTHERWISE COMPROMISED. THE LOCATION OF LEAKAGE WAS NOT DETERMINED. ADDITIONALLY, AN AUTOPSY WAS NOT PERFORMED ON THIS PATIENT. THE DEVICE REMAINED IMPLANTED, RENDERING IT UNAVAILABLE FOR ADDITIONAL INVESTIGATION. SCANS HAVE BEEN REQUESTED. 3331 - ANALYSIS OF PRODUCTION RECORDS: FULL BATCH REVIEW PERFORMED FROM BASE FABRIC TO FINISHED PRODUCT WHICH CONFIRMED THE BATCH WAS MANUFACTURED TO SPECIFICATION. 4117 - DEVICE NOT RETURNED: DEVICE REMAINED IMPLANTED. INVESTIGATION FINDINGS: 3233 - RESULTS PENDING COMPLETION OF INVESTIGATION. INVESTIGATION CONCLUSION: 11 - CONCLUSION NOT YET AVAILABLE AS INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22156969&pc=OZP | 2032227-2025-187191 | 2024-08-18 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-06 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; High Blood Pressure/ Hypertension; Hypoglycemia; Renal Failure | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. CAUSE OF DEATH WAS FROM NATURAL CAUSES RESULTING FROM KIDNEY FAILURE, HYPERTENSION, AND DIABETES TYPE 2. THE CUSTOMER WAS WEARING THE PUMP AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-242A, MMT-1884, MMT-332A, AND MMT-7040A. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING. NO PRODUCT RETURN WAS REQUIRED FOR MMT-242A. NO PRODUCT RETURN WAS REQUIRED FOR MMT-1884. NO PRODUCT RETURN WAS REQUIRED FOR MMT-332A. NO PRODUCT RETURN WAS REQUIRED FOR MMT-7040A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22157058&pc=OZD | 1220648-2025-29055 | 2025-05-18 04:00 | Death | ABIOMED, INC. | 2025-06-06 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Ischemia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT AN IMPELLA CP WAS PLACED ON A PATIENT WITH CARDIOGENETIC SHOCK. DURING SUPPORT, IT WAS NOTED THAT THE PUMP RAN WELL WITHOUT ANY PROBLEMS. EXTRACORPOREAL MEMBRANE OXYGENATION WAS IN WEANING PROCESS AND VENTING WORKED WELL. NO ESCALATION TO AN IMPELLA 5.5. INITIAL INDICATION WAS DUE TO ISCHEMIA OF THE LEFT LEG. HOWEVER, THE PROBLEM WAS RESOLVED WITH THE HELP OF AN EXTRACORPOREAL FEMOROFEMORAL BYPASS. THE NEXT DAY IT WAS NOTED THAT THE PATIENT WAS IN CRITICAL CONDITION OVERALL. ESCALATION TO AN IMPELLA 5.5 COULD BE AN OPTION. HOWEVER, THEY FIRST WANTED TO CHECK THE PATIENTS NEUROLOGICAL STATUS IN CAUSE OF PROLONGED RESUSCITATION. THE CRANIAL COMPUTED TOMOGRAPHY WAS POSITIVE FOR BRAIN DAMAGE. THE PHYSICIAN WITHDREW CARE IN CAUSE OF DIAGNOSED BRAINDEAD AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22157103&pc=OZD | 1220648-2025-29056 | 2025-05-16 04:00 | Death | ABIOMED, INC. | 2025-06-06 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND SOON THEREAFTER THE PATIENT VENTRICULAR FIBRILLATION ARRESTED AND EXPIRED. THE PATIENT WAS BROUGHT TO CATHETERIZATION LAB, AND INITIAL ANGIOGRAM SHOWED THE LEFT ANTERIOR DESCENDING (LAD) ARTERY WAS TOTALLY OCCLUDED. EPINEPHRINE AND LEVOPHED DRIPS WERE INITIATED, AND THE DECISION WAS MADE TO PLACE AN IMPELLA CP FOR THE REMAINDER OF THE PERCUTANEOUS CORONARY INTERVENTION (PCI) PROCEDURE. THE LAD WAS SUCCESSFULLY OPENED PERCUTANEOUS TRANSLUMINAL CORONARY ANGIOPLASTY AND DRUG-ELUTING STENT. FINAL ANGIOGRAMS SHOWED IMPROVED FLOWS TO THE TREATED AREA. DRIPS WERE INCREASED SLIGHTLY OVER THE COURSE OF THE CASE DUE TO SIGNIFICANT DROPS IN PRESSURES. THE DECISION WAS MADE TO KEEP PATIENT ON IMPELLA MCS POST PCI WHILE IN THE INTENSIVE CARE UNIT. DURING TRANSFER TO THE UNIT, THE PATIENT BECAME RESTLESS AND BEGAN TO SIGNIFICANTLY DECLINE. CODE BLUE WAS CALLED ONCE ON THE FLOOR, AND IT WAS NOTED THAT THE PATIENT WAS IN VENTRICULAR FIBRILLATION ARREST. CARDIOPULMONARY RESUSCITATION AND DEFIBRILLATION WERE INITIATED, AND INTRAVENOUS PUSHES WERE ADMINISTERED OVER THE COURSE OF 20 MINUTES. THE DECISION WAS MADE TO CALL THE CODE AFTER FINAL PULSE CHECK WAS UNSUCCESSFUL. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22157145&pc=OZD | 1220648-2025-29057 | 2025-05-15 04:00 | Death | ABIOMED, INC. | 2025-06-06 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE HEMASHIELD PLATINUM GRAFT STARTED OOZING A FEW MINUTES AFTER IMPELLA IMPLANT. BEFORE THE PERCUTANEOUS CORONARY INTERVENTION, THE SURGEON MEDICATED THE GRAFT WITH GLUE AND TABOTAMP BUT THE COMPLICATION COULD NOT BE MANAGED. IT WAS FURTHER REPORTED THAT THE CORONARY REVASCULARIZATION WAS SUCCESSFULLY COMPLETED, AND THE GRAFT OOZING APPEARED CONTROLLED. DURING THE TRANSFER FROM THE OPERATING TABLE TO THE TRANSPORT BED, THE PATIENT DEVELOPED VENTRICULAR ARRHYTHMIA WITH HEMODYNAMIC COLLAPSE, WHICH WAS RESOLVED AFTER SINUS RHYTHM RESTORATION, ADRENALINE AND FLUID ADMINISTRATION. A FEW MINUTES LATER, A SIMILAR EPISODE OCCURRED BUT THE PATIENT WAS UNRESPONSIVE TO TREATMENT. ECHOCARDIOGRAPHY SHOWED COMPLETE VENTRICULAR DYSFUNCTION AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22157168&pc=MKJ | 1220908-2025-02159 | 2025-05-25 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-06-06 04:00 | MKJ | X SERIES | Unable to Obtain Readings; Incorrect Interpretation of Signal | Insufficient Information | P110022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A (B)(6) MALE PATIENT, IN CARDIAC ARREST, THE DEVICE PROMPTED ATTACH PADS AND FAILED TO ANALYZE. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED; HOWEVER, A DNR WAS PRESENTED AND CPR HALTED. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL UNITED KINGDOM; THE CUSTOMER'S REPORT WAS DUPLICATED AND ATTRIBUTED TO THE CPR ADAPTOR. THE CPR ADAPTOR WAS REPLACED TO RESOLVE THE REPORT. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS FOR REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: JUSTIFICATION FOR NO UDI: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS RECEIVED THE PRODUCT AND WILL BE PROVIDING A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22157332&pc=OZD | 1220648-2025-29061 | 2025-05-15 04:00 | Death | ABIOMED, INC. | 2025-06-06 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE TEAM HAD SUSPICIONS THAT THE AORTIC LEAFLET WAS EITHER ¿IMPINGED OR PERFORATED.¿ THE PHYSICIAN STATED THAT THE PATIENT¿S TISSUE WAS ¿FRIABLE.¿ THE PUMP WAS EXPLANTED DUE TO WITHDRAWING CARE AND THE PATIENT EXPIRED. TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) SHOWED PROBABLE PERFORATION OF THE AORTIC LEAFLET. Manufacturer Narrative: THE INVESTIGATION OF HEART VALVE DAMAGE HAS BEEN COMPLETED. IMPELLA 5.5 WAS RETURNED FOR EVALUATION. UPON VISUAL INSPECTION OF THE RETURNED PRODUCT NO DAMAGE WAS OBSERVED. BIOMATERIAL WAS PRESENT WITHIN THE CANNULA AT THE IMPELLER BLADE. DATA LOGS SHOWED A FEW INSTANCES OF "IMPELLA IN AORTA" ALARMS DURING SUPPORT. NO PERSISTENT POSITIONING ALARMS NOTED DURING SUPPORT. CLINICAL DETAILS NOTED THAT PATIENT HAD FRIABLE TISSUE AND TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) SHOWED PERFORATION OF THE AORTIC LEAFLET AFTER IMPELLA WAS REMOVED. THE ROOT CAUSE OF THE HEART VALVE DAMAGE WAS MOST LIKELY PATIENT CONDITION RELATED AS CLINICAL DETAILS NOTED THAT PATIENT HAD FRIABLE TISSUE AND NO ABNORMALITIES WERE FOUND ON RETURNED PRODUCT. Manufacturer Narrative: THE IMPELLA PUMP WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITES STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: AXILLARY INSERTION OF IMPELLA 5.5 CATHETER ¿USE FLUOROSCOPY FOR PLACEMENT IMPELLA CATHETER PERFORMANCE WILL BE COMPROMISED IF CORRECT PLACEMENT CANNOT BE CONFIRMED. WHILE OTHER IMAGING TECHNIQUES, SUCH AS TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE), CAN HELP CONFIRM THE POSITION OF THE IMPELLA CATHETER AFTER PLACEMENT, TEE DOES NOT ALLOW VISUALIZATION OF THE ENTIRE CATHETER ASSEMBLY AND IS INADEQUATE FOR RELIABLY PLACING THE IMPELLA CATHETER ACROSS THE AORTIC VALVE¿. SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS ¿UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH.¿ IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF IMPELLA CATHETER ¿IF THE INLET AREA OF THE IMPELLA CATHETER IS TOO CLOSE TO OR ENTANGLED IN THE PAPILLARY MUSCLE AND/OR SUBANNULAR STRUCTURES SURROUNDING THE MITRAL VALVE, IT CAN AFFECT MITRAL VALVE FUNCTION AND NEGATIVELY IMPACT CATHETER FLOW.¿ SECTION: UNDERSTANDING AND MANAGING IMPELLA CATHETER POSITION ALARMS. ¿IF THE IMPELLA CATHETER IS COMPLETELY OUT OF THE VENTRICLE, DO NOT ATTEMPT TO REPOSITION THE CATHETER ACROSS THE VALVE WITHOUT A GUIDEWIRE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22157441&pc=PNJ | 9612164-2025-02767 | 2025-05-08 04:00 | Death | MEDTRONIC IRELAND | 2025-06-06 04:00 | PNJ | MICRA¿ AV2 | Use of Device Problem; Device Dislodged or Dislocated | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) BECAME DISLODGED DUE TO THE TETHER CUTTING. THE LEADLESS IPG WAS EXPLANTED. IT WAS FURTHER NOTED THAT THE PATIENT PASSED AWAY THREE DAYS LATER.Event Description: IT WAS FURTHER REPORTED THAT THE DEATH WAS NOT BELIEVED TO BE LINKED TO THE LEADLESS IPG. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [(B)(6)]). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22157487&pc=DSQ | 2916596-2025-03508 | 2025-04-15 04:00 | Death | THORATEC CORPORATION | 2025-06-06 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Loss of consciousness; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND UNRESPONSIVE AT HOME. CARDIOPULMONARY RESUSCITATION (CPR) WAS STARTED AT HOME. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. NO AUTOPSY WAS PERFORMED, AND THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: SECTION H6: HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE PUMP WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22157693&pc=OZD | 1220648-2025-29067 | 2025-05-13 04:00 | Death | ABIOMED, INC. | 2025-06-06 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE PATIENT ON IMPELLA CP SUPPORT HAD BRADYCARDIC ARRHYTHMIA WITH FLAT HEART RHYTHM AND A LINE, BUT PATIENT REMAINED AWAKE AND ALERT. THE PHYSICIAN TOOK THE PATIENT FOR TEMPORARY PACEMAKER PLACEMENT. ADDITIONALLY, THE PATIENT RECEIVED TWO UNITS OF RED BLOOD CELLS (PRBCS) FOR AN UNKNOWN SOURCE OF BLEEDING. PATIENT REQUIRED 2 MORE UNITS OF PRBCS. NO OBVIOUS SIGNS OF ACTIVE BLEEDING BUT POTENTIALLY LEFT GROIN VENOUS SHEATH. IMPELLA SITE DID NOT HAVE ANY OOZING OR HEMATOMA. IT WAS FURTHER REPORTED THAT THE PATIENT HAD VENTRICULAR FIBRILLATION ARRHYTHMIA ON DAY TWO OF IMPELLA SUPPORT AND HEART RATE SUSTAINING IN THE 190S. THE PHYSICIAN SPOKE WITH FAMILY AND DECIDED TO WITHDRAW CARE. THE PATIENT EXPIRED SHORTLY AFTER. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION FOR THE REPORTED VENTRICULAR FIBRILLATION (VF) AND UNKNOWN SOURCE OF BLEEDING HAS BEEN COMPLETED. THE DEVICE NOR SUFFICIENT CLINICAL INFORMATION WAS RETURNED FOR EVALUATION. THEREFORE, THE ROOT CAUSE OF THE BLEEDING NOR THE VF COULD NOT BE DETERMINED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22158020&pc=OZD | 1220648-2025-29070 | 2025-05-13 04:00 | Death | ABIOMED, INC. | 2025-06-06 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT'S RIGHT ARM BECAME MOTTLED DURING IMPELLA 5.5 SUPPORT. VASCULAR SURGERY WAS SET TO MONITOR THE CONDITION, HOWEVER THE PATIENT PASSED AWAY THE FOLLOWING DAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22158208&pc=OZD | 1220648-2025-29072 | 2025-05-26 04:00 | Death | ABIOMED, INC. | 2025-06-06 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED VENTRICULAR PERFORATION LEADING TO MULTIPLE ORGAN FAILURE AND DEATH. THERE WERE NO PROBLEMS WITH THE IMPELLA INSERTION OR OPERATION AND THE ATTENDING PHYSICIAN STATED THE DEATH WAS NOT DUE TO THE IMPELLA. ABIOMED'S MEDICAL SAFETY OFFICER REVIEWED THE EVENT AND DETERMINED THE DEATH CANNOT BE DISSOCIATED AND THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THEIMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: A2, A3, AND A4 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST SYSTEM LISTS VENTRICULAR ARRYTHMIA AS AN ADVERSE EVENT THAT OCCURRED WHILE IN USE WITH THE IMPELLA DEVICE. IT STATES THAT "ONLY A FRACTION OF THESE RATES WERE ATTRIBUTED TO THE IMPELLA DEVICE AND THE EVENTS RESOLVED WITHOUT RESIDUAL EFFECT IN MOST OF THE CASES, UNLESS THE EVENT OF DEATH OCCURRED. OVERALL, THE RESULTS DID NOT SHOW ANY EVIDENCE OF INCREASED MORBIDITY ASSOCIATED WITH THE IMPELLA SUPPORT IN CARDIOMYOPATHY, MYOCARDITIS, AND PERIPARTUM CARDIOMYOPATHY (PPCM) PATIENTS." SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22158396&pc=LJT | 3006260740-2025-03673 | 2025-05-14 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-06-06 04:00 | LJT | X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | K022983 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME AFTER PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY DEVELOPED WITH UNSPECIFIED INFECTION. REPORTEDLY, THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE/WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22158614&pc=NPT | 2135147-2025-02964 | 2013-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-06-06 04:00 | NPT | NAVITOR TRANSCATHETER AORTIC VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P190023 |  | 1 | Event Description: THE ARTICLE "TRANSCATHETER AORTIC VALVE REPLACEMENT IN SMALL AORTIC ANNULI: A PROPENSITY-MATCHED COMPARISON BETWEEN INTRA-ANNULAR SELF-EXPANDING VALVES AND BALLOON-EXPANDABLE VALVES" WAS REVIEWED. THE ARTICLE PRESENTED AN OBSERVATIONAL, RETROSPECTIVE, SINGLE CENTER STUDY TO COMPARE INTRA-ANNULAR SELF-EXPANDING VALVES (SEV) WITH BALLOON-EXPANDABLE VALVES (BEV) IN THE SMALL AORTIC ANNULUS (SAA) POPULATION IN TERMS OF CLINICAL AND ECHOCARDIOGRAPHIC CHARACTERISTICS. DEVICES INCLUDED IN THIS STUDY WERE NAVITOR TRANSCATHETER AORTIC VALVE, PORTICO TRANSCATHETER AORTIC VALVE, SAPIEN 3, SAPIEN 3 ULTRA, AND SAPIEN XT TISSUE HEART VALVES. THE ARTICLE CONCLUDED THAT INTRA-ANNULAR SEV PRESENTED BETTER HEMODYNAMICS COMPARED TO BEV; HOWEVER, SURVIVAL AND INCIDENCE OF STROKE WERE COMPARABLE BETWEEN THE GROUPS AT 2 YEARS. POTENTIAL RISK OF NONSTRUCTURAL VALVE DETERIORATION WITH BEV NEEDED FURTHER INVESTIGATION WITH LONGER FOLLOW-UP. [THE PRIMARY AND CORRESPONDING AUTHOR WAS IBRAHIM SULTAN, UNIVERSITY OF PITTSBURGH, PITTSBURGH, PENNSYLVANIA, UNITED STATES OF AMERICA, WITH CORRESPONDING EMAIL: SULTANI@UPMC.EDU] THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2013 TO DECEMBER 2023. A TOTAL OF 663 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 16% RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 82.0 YEARS, AND THE MAJORITY GENDER WAS FEMALE. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, DIALYSIS, PRIOR PERIPHERAL ARTERY DISEASE, STROKE, CORONARY ARTERY BYPASS GRAFT, PERCUTANEOUS CORONARY INTERVENTION, UNABLE TO WALK, HEART BLOCK, AND HEART FAILURE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRANSCATHETER AORTIC VALVE REPLACEMENT IN SMALL AORTIC ANNULI: A PROPENSITY-MATCHED COMPARISON BETWEEN INTRA-ANNULAR SELF-EXPANDING VALVES AND BALLOON-EXPANDABLE VALVES WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, DIALYSIS, PRIOR PERIPHERAL ARTERY DISEASE, STROKE, CORONARY ARTERY BYPASS GRAFT, PERCUTANEOUS CORONARY INTERVENTION, UNABLE TO WALK, HEART BLOCK, AND HEART FAILURE. SOME OF THE COMPLICATIONS REPORTED WERE AORTIC VALVE STENOSIS AND DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: TRANSCATHETER AORTIC VALVE REPLACEMENT IN SMALL AORTIC ANNULI: A PROPENSITY-MATCHED COMPARISON BETWEEN INTRA-ANNULAR SELF-EXPANDING VALVES AND BALLOON-EXPANDABLE VALVES. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22158754&pc=OZD | 1220648-2025-29075 | 2025-05-11 04:00 | Death | ABIOMED, INC. | 2025-06-06 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AN IMPLANTED IMPELLA 5.5 PUMP BEGAN TO EXPERIENCE HIGH PURGE PRESSURE DURING PATIENT SUPPORT. A LYSIS PROTOCOL WAS SENT TO THE TO THE STAFF IN RESPONSE TO THE DECREASE IN PURGE FLOW. SUPPORT CONTINUED FOR ANOTHER FIVE DAYS, AT WHICH POINT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER. ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22158766&pc=DSQ | 2916596-2025-03482 | 2025-05-16 04:00 | Death | THORATEC CORPORATION | 2025-06-06 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO WITHDRAWAL OF SUPPORT. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22158936&pc=OZD | 1220648-2025-29077 | 2025-05-13 04:00 | Death | ABIOMED, INC. | 2025-06-06 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT LOST ALL DISTAL PULSE DURING IMPELLA CP SUPPORT. DUE TO THE PATIENT'S UNDERLYING CONDITION, THE PATIENT'S FAMILY WITHDREW CARE AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22158948&pc=ESW | 3005099803-2025-02561 | 2025-05-02 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-06 04:00 | ESW | WALLFLEX ESOPHAGEAL | Adverse Event Without Identified Device or Use Problem | Aspiration/Inhalation | K233939 |  | 1 | Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT A WALLFLEX ESOPHAGEAL STENT WAS IMPLANTED IN THE ESOPHAGUS TO TREAT A MALIGNANT OBSTRUCTION DURING A GASTROSCOPY WITH STENT PLACEMENT PROCEDURE PERFORMED ON (B)(6) 2025. THE WALLFLEX ESOPHAGEAL STENT WAS SUCCESSFULLY DEPLOYED. HOWEVER, UNIDENTIFIED LIQUID IN THE PATIENT'S STOMACH WAS OBSERVED. WHEN THE STENT WAS ABLE TO OPEN UP THE OBSTRUCTION IN THE ESOPHAGUS, THE LIQUID CAME UP CAUSING THE PATIENT TO ASPIRATE. EMERGENCY RESUSCITATION WAS ATTEMPTED, BUT THE PATIENT PASSED AWAY. NO FURTHER INFORMATION HAS BEEN OBTAINED DESPITE GOOD FAITH EFFORTS. Manufacturer Narrative: BLOCK D4, H4: THE COMPLAINANT WAS UNABLE TO PROVIDE THE SUSPECT DEVICE UPN AND LOT NUMBER. THEREFORE, THE MANUFACTURE AND EXPIRATION DATES ARE UNKNOWN. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK H6: IMDRF PATIENT CODE E0704 CAPTURES THE REPORTABLE EVENT OF ASPIRATION. IMDRF IMPACT CODE F2306 CAPTURES THE REPORTABLE EVENT OF EMERGENCY RESUSCITATION WAS ATTEMPTED. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF PATIENT PASSED AWAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22159535&pc=DSQ | 2916596-2025-03257 | 2025-05-11 04:00 | Death | THORATEC CORPORATION | 2025-06-06 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Heart Failure/Congestive Heart Failure; Swelling/ Edema | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RENAL FAILURE. THE RENAL FAILURE WAS RELATED TO HEART FAILURE. THE PATIENT HAD ELEVATED CREATININE LEVEL, EDEMA, AND DECREASED URINE OUTPUT. LAB TESTS WERE DONE AND THE PATIENT WAS TREATED WITH DIURETICS. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE REPORTEDLY OPERATED AS EXPECTED. THE PUMP WAS NOT EXPLANTED, AND AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(4) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RENAL DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22159691&pc=MIH | 3011063223-2025-00103 |  | Death | ENDOLOGIX | 2025-06-06 04:00 | MIH | AFX2 | Insufficient Information | Aneurysm | P040002 |  | 1 | Event Description: THE REPORTED INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH REVIEW OF AN INDUSTRY RELEVANT JOURNAL ARTICLE. DOI.ORG/10.1016/J.AVSG.2024.04.027. SAFETY EVALUATION OF UNIBODY ENDOGRAFTS FOR ABDOMINAL AORTIC ANEURYSM REPAIR: A SYSTEMATIC REVIEW AND META-ANALYSIS SAFIA ESSAM, MIRNA HUSSEIN, ABDELRAHMAN MAMDOUH AHMED, LUJAINA AHMED, HAMED GABER, HASSAN EL-MASRY, RODINA MOHAMED ABDELAAL, NOURHAN GALAL, AHMED KASSEM, AND WAEL SHAALAN, ALEXANDRIA, EGYPT BACKGROUND: THE UNIBODY BIFURCATED AORTIC ENDOGRAFT (AFX/AFX2) HAS EMERGED AS A TREATMENT OPTION FOR ABDOMINAL AORTIC ANEURYSMS (AAAS). THIS SYSTEMATIC REVIEW AND META-ANALYSIS AIMED TO EVALUATE THE SAFETY OF THE UNIBODY ENDOGRAFT. METHODS: A LITERATURE SEARCH WAS CONDUCTED IN COCHRANE LIBRARY, SCOPUS, WEB OF SCIENCE, AND PUBMED. STUDIES ASSESSING THE UNIBODY ENDOGRAFT FOR AAA REPAIR BETWEEN 2014 AND 2023 WERE INCLUDED. THE DEFINED PRIMARY OUTCOMES WERE THE INCIDENCES OF TYPE I, II, AND III ENDOLEAKS. THE SECONDARY OUTCOMES WERE ACCESS SITE PROBLEMS, ANEURYSM-RELATED MORTALITY, ANEURYSM RUPTURE, ALL-CAUSE MORTALITY, ANEURYSM SAC GROWTH, LIMB OCCLUSION, STENT GRAFT MIGRATION, AND TECHNICAL SUCCESS RATE. RESULTS: FOURTEEN STUDIES INCLUDING 12 OBSERVATIONAL STUDIES AND 2 RANDOMIZED CONTROLLED TRIALS WERE INCLUDED IN THE SYSTEMATIC REVIEW. THE META-ANALYSIS INCLUDED 10 STUDIES WITH (B)(4) PATIENTS THAT REPORTED THE MEASURED OUTCOMES AND EXCLUDED 4 STUDIES THAT DID NOT. TYPE II ENDOLEAKS HAD THE HIGHEST INCIDENCE OF 12% (95% CONFIDENCE INTERVAL [CI]: 4E20%), FOLLOWED BY TYPE III ENDOLEAKS WITH AN INCIDENCE OF 3% (95% CI: 1E5%). THE INCIDENCE OF TYPE I ENDOLEAKS WAS 1% (95% CI: 0E2%). A SUBGROUP ANALYSIS BY FOLLOW-UP DURATION SHOWED THAT TYPE II ENDOLEAK INCIDENCE WAS HIGHER AFTER 1 TO 2 YEARS OF FOLLOW-UP THAN 3 TO 4 YEARS OF FOLLOW-UP. THE INCIDENCE OF ANEURYSMAL MORTALITY WAS 2% (95% CI: 0E7%), LIMB OCCLUSION WAS 1% (95% CI: 0E1%), STENT GRAFT MIGRATION WAS 1% (95% CI: 0E2%), ANEURYSMAL RUPTURE WAS 6% (95% CI: 2E11%), ACCESS SITE PROBLEMS WERE 7% (95% CI: 2E13%), ANEURYSM SAC GROWTH WAS 2% (95% CI: 0E4%), ALL-CAUSE MORTALITY WAS 21% (95% CI: 4E38%), AND TECHNICAL SUCCESS RATE WAS 100% (95% CI: 98E 100%). CONCLUSIONS: THE UNIBODY ENDOGRAFT IS A SAFE AND MINIMALLY INVASIVE APPROACH FOR AAA REPAIR. HOWEVER, POTENTIAL COMPLICATIONS NECESSITATE CLOSE PATIENT FOLLOW-UP AFTER THE INTERVENTION. NO ADDITIONAL INFORMATION WILL BECOME AVAILABLE REGARDING THIS INITIAL/FINAL REPORT DUE TO THE ABSENCE OF PATIENT IDENTIFIERS. THIS REPORT ADDRESSES PATIENT 3 OF 3. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. THE REPORTED ADVERSE EVENT/INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH AN ONGOING REVIEW OF INDUSTRY RELEVANT JOURNAL ARTICLES WHERE PATIENT RELATED OUTCOMES ARE PRESENTED ANONYMOUSLY, AND PATIENT SPECIFIC DEVICE INFORMATION IS UNAVAILABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE COMPLETED. THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION REMAINS UNKNOWN. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT COULD NOT BE COMPLETED AS THE REPORTED EVENT WAS FOUND DURING PUBLICATION REVIEW WHERE THE PATIENT INFORMATION IS ANONYMOUS. DEVICE, USE, PROCEDURE, AND/OR ANATOMY RELATEDNESS TO THIS ADVERSE EVENT/INCIDENT COULD NOT BE EVALUATED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22160187&pc=OZD | 1220648-2025-29085 | 2024-06-12 04:00 | Death | ABIOMED, INC. | 2025-06-06 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhagic Stroke | P140003 |  | 1 | Event Description: ABIOMED LEGAL ORIGINALLY NOTIFIED OF AN ALLEGATION REGARDING A PATIENT WHO WAS IMPLANTED WITH AN IMPELLA 5.5 FOR MECHANICAL CIRCULATORY SUPPORT AND DURING SUPPORT, THE PATIENT DEVELOPED A HEMATOMA AT THE IMPELLA INSERTION SITE, A MASSIVE SUBDURAL HEMATOMA, AND SUBARACHNOID HEMORRHAGE. NO FURTHER DETAILS SURROUNDING THE EVENT WERE PROVIDED. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION). SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22160201&pc=DSQ | 2916596-2025-03818 | 2025-05-30 04:00 | Death | THORATEC CORPORATION | 2025-06-06 04:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. THE PUMP WOULD NOT BE EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LVAS, SERIAL NUMBER (B)(6) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO HOSPITAL POLICY, NO FURTHER INFORMATION PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22160404&pc=DSQ | 2916596-2025-03556 | 2025-05-05 04:00 | Death | THORATEC CORPORATION | 2025-06-06 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. PATIENT SUFFERED FROM A REFRACTORY, PROGRESSIVE MULTI-ORGAN FAILURE WHICH LED TO DEATH AT THE END. NO DEVICE RELATED ISSUES LED TO MULTI-ORGAN FAILURE. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS RUNNING AS EXPECTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED MULTISYSTEM ORGAN FAILURE (MSOF) AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22160540&pc=LWS | 2649622-2025-15369 | 2025-05-27 04:00 | Death | MPRI | 2025-06-06 04:00 | LWS | SPRINT QUATTRO SECURE MRI SURESCAN | Defibrillation/Stimulation Problem; Under-Sensing | Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. A REVIEW OF CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) DEVICE DATA REVEALED THAT THE PATIENT WAS IN A TRUE VENTRICULAR FIBRILLATION (VF) EPISODE AND IT WAS NOTED THAT THE PATIENT'S RHYTHM BEGIN TO SLOW. THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDER SENSING OF THE VF THAT RESULTED IN INAPPROPRIATE WITHHOLDING OF THERAPY. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED RIGHT VENTRICULAR UNDERSENSING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: 5076-58 LEAD, DTPB2D1 CRT-D, UNKNOWN COMPETITOR IMPLANTED: (B)(6) 2023. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22160614&pc=DSQ | 2916596-2025-03756 | 2024-12-13 05:00 | Death | THORATEC CORPORATION | 2025-06-06 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE CAUSE OF DEATH WAS NOT RELATED TO LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT EXPIRED. ALTHOUGH THE CAUSE OF DEATH WAS REPORTED AS UNKNOWN, THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO FURTHER INFORMATION WAS KNOWN. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22160641&pc=OZD | 1220648-2025-29087 | 2025-05-09 04:00 | Death | ABIOMED, INC. | 2025-06-06 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED A IMPELLA CP TO SUPPORT THE PATIENT WITH MYOCARDITIS. AFTER 7 DAYS THE PATIENT EXPIRED FROM MULTIPLE ORGAN FAILURE AND THE RELATIONSHIP TO IMPELLA IS UNKNOWN. ADDITIONAL INFORMATION PROVIDED FROM THE DOCTOR THAT THE EXTRACORPOREAL MEMBRANE OXYGENATION CIRCUIT CAME LOOSE, CAUSING THE PATIENT'S OVERALL CONDITION TO DETERIORATE, LEADING TO DISSEMINATED INTRAVASCULAR COAGULATION AND SUBSEQUENT DEATH FROM MULTIPLE ORGAN FAILURE. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. ABIOMED'S MEDICAL SAFETY OFFICER REVIEW THE PATIENTS OUTCOME AND DETERMINED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22160777&pc=LWP | 2017865-2025-75639 | 2025-05-13 04:00 | Death | ABBOTT | 2025-06-06 04:00 | LWP | ASSURITY MRI | Premature Discharge of Battery | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT AN ASSESSMENT OF THE DEVICE BATTERY PERFORMED BY TECHNICAL SUPPORT SHOWED PREMATURE BATTERY DEPLETION. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE PATIENT PASSED AWAY; THE ELECTRONIC PERFORMANCE INDICATOR (EPI) WAS RECEIVED AFTER THE PATIENT WAS DECEASED. THERE WAS NO ALLEGATION THAT THE PATIENT'S DEATH WAS RELATED TO THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22160856&pc=MIH | 2247858-2025-00141 | 2025-04-29 04:00 | Death | BOLTON MEDICAL, INC. | 2025-06-06 04:00 | MIH | TREO ABDOMINAL STENT-GRAFT SYSTEM | Patient Device Interaction Problem | Cardiovascular Insufficiency | P190015 |  | 1 | Event Description: "PATIENT (B)(6) DIED ON THE (B)(6) APRIL 2025, THE INCIDENT OCCURRED IN A DIFFERENT HOSPITAL. SITE IS NOT AWARE OF THE EXACT CAUSE OF DEATH YET BUT MORE INFORMATION TO COME. THIS EVENT WAS UNANTICIPATED, IT IS UNDETERMINED SO FAR THAT IT IS RELATED TO THE PROCEDURE, DEVICE AND PRE-EXISTING CONDITIONS." PATIENT OUTCOME: "DEATH ON (B)(6) 2025." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH TWO DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR 2247858-2025-00141, AND DEVICE 2 IS BEING REPORTED UNDER MDR 2247858-2025-00142. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22160857&pc=MIH | 2247858-2025-00142 | 2025-04-29 04:00 | Death | BOLTON MEDICAL, INC. | 2025-06-06 04:00 | MIH | TREO ABDOMINAL STENT-GRAFT SYSTEM | Patient Device Interaction Problem | Cardiovascular Insufficiency | P190015 |  | 1 | Event Description: "PATIENT (B)(6) DIED ON THE (B)(6) 2025, THE INCIDENT OCCURRED IN A DIFFERENT HOSPITAL. SITE IS NOT AWARE OF THE EXACT CAUSE OF DEATH YET BUT MORE INFORMATION TO COME. THIS EVENT WAS UNANTICIPATED, IT IS UNDETERMINED SO FAR THAT IT IS RELATED TO THE PROCEDURE, DEVICE AND PRE-EXISTING CONDITIONS." PATIENT OUTCOME: "DEATH ON (B)(6) 2025." Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH TWO DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR 2247858-2025-00141, AND DEVICE 2 IS BEING REPORTED UNDER MDR 2247858-2025-00142. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22161009&pc=DSQ | 2916596-2025-03542 | 2021-11-29 05:00 | Death | THORATEC CORPORATION | 2025-06-06 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE PATIENT OUTCOME, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. Manufacturer Narrative: SECTION A: HOSPITAL POLICY PROHIBITS THE RELEASE OF THE INFORMATION. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WILL NOT COMMUNICATE ANY FURTHER PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22161245&pc=NVN | 2017865-2025-75642 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-06 04:00 | NVN | ULTIPACE¿ LEAD, 58CM MRI | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE AND INFECTIVE ENDOCARDITIS. IT WAS SUSPECTED THAT THE INFECTION CAME FROM THE TOE AMPUTATION AND THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS RELATED TO DEVICE OR DEVICE RELATED PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22161255&pc=OZD | 1220648-2025-29095 | 2025-05-09 04:00 | Death | ABIOMED, INC. | 2025-06-06 04:00 | OZD | IMPELLA CP | Defective Device | Atrial Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. PRIOR TO IMPELLA PLACEMENT, THE PATIENT WAS BEING TREATED FOR ATRIAL FIBRILLATION RAPID VENTRICULAR RESPONSE. THE PATIENT BECAME HYPOTENSIVE AND LEVO WAS STARTED. IT WAS FURTHER REPORTED THAT THE PATIENT CODED AND HAD AROUND 30 MINUTES OF CARDIOPULMONARY RESUSCITATION BEFORE RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED. THE PATIENT WAS EMERGENTLY TAKEN TO CATH LAB FOR IMPELLA PLACEMENT. DURING PLACEMENT, WHEN PIGTAIL WAS ADVANCED THROUGH THE PEEL AWAY, THE SHEATH HAD A KINK IN IT. A DECISION WAS MADE TO EXCHANGE LONG PEEL AWAY FOR THE SHORT ONE AND IMPELLA WAS PLACED. ADDITIONAL INFORMATION NOTED CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22161432&pc=FGA | 2429304-2025-00173 | 2025-05-05 04:00 | Death | OLYMPUS WINTER & IBE GMBH | 2025-06-06 04:00 | FGA | OP TELESCOPE, 7°, 11 FR. X 220 MM, 7.5 FR. CHANNEL, ANGLED OCULAR | Material Twisted/Bent | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Sepsis; Septic Shock; Ventilator Dependent |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CYSTOSCOPY AND HOLMIUM LASER OR BOTH PENILE AND BULBAR URETHRAL STRICTURES, COMPLEX URETHRAL DILATION, PLACEMENT OF A FOLEY CATHETER, REMOVAL OF THE ARTIFICIAL URINARY SPHINCTER, AND PLACEMENT OF A SUPRAPUBIC CYSTOSTOMY TUBE. AFTER THE PROCEDURE WAS PERFORMED, STAFF FOUND THAT THE 11 FRENCH 7-DEGREE PEDIATRIC CYSTOSCOPE (THE ONLY SCOPE THAT COULD PASS THROUGH PROXIMAL PENILE URETHRAL STRICTURE) WAS SLIGHTLY BENT. PATIENT RECOVERED WELL WITH MINIMAL BLEEDING IN POST ANESTHESIA CARE UNIT (PACU) AND WAS DISCHARGED HOME WITH NO COMPLICATIONS. THE PATIENT LATER REPORTED WAKING UP AT HOME TO FIND BED SATURATED WITH BLOOD. THE PATIENT WENT TO LOCAL EMERGENCY ROOM AND WAS SUBSEQUENTLY TRANSFERRED BACK TO THE FACILITY WITH ANEMIA AND HYPOTENSION. THE PATIENT WAS PLACED ON PRESSORS, VENTILATED, AND TREATED FOR SEPSIS. PATIENT WAS THEN TRANSFERRED TO A HIGHER LEVEL OF CARE FOR PROCEDURES UNABLE TO BE PERFORMED AT THE FACILITY. CUSTOMER NOTES THAT IT IS UNKNOWN WHETHER THE DEVICE IS RELATED TO OUTCOMES. E2 - HEALTH PROFESSIONAL: WAS UNABLE TO BE CHANGED FOR NO BUT SHOULD BE YES. SINCE THE DEVICE DESCRIBED WAS A "11 FRENCH 7-DEGREE PEDIATRIC CYSTOSCOPE", WE SELECTED "(B)(6)" AS A REPRESENTATIVE PRODUCT.Event Description: ADDITIONAL INFORMATION RECEIVED THE PATIENT WAS RE-ADMITTED TO THE FACILITY, RECEIVED BLOOD TRANSFUSIONS, WAS TREATED FOR SEPTIC SHOCK, AND WAS ADMINISTERED THE ANTIBIOTICS ZOSYN AND LINEZOLID. THE PATIENT WAS SUBSEQUENTLY TRANSFERRED FOR SLEDD (SUSTAINED LOW-EFFICIENCY DAILY DIALYSIS) AND HAS SINCE EXPIRED. THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO UPDATE FIELDS (A2, A3, A4, A5, A6, B1, B2, B5, B7, F6, F10) AND ATTACH THE RELATED MEDWATCH REPORT. ADDITIONALLY, THIS SUPPLEMENTAL REPORT IS PROVIDING FURTHER INFORMATION SHARED BY THE CUSTOMER.Event Description: ADDITIONAL INFORMATION WAS RECEIVED FROM THE CUSTOMER INDICATING THAT THE PATIENT WAS RE-ADMITTED TO A DIFFERENT FACILITY FOR TREATMENT OF AN INFECTION. THE PATIENT SUBSEQUENTLY EXPIRED AT THAT FACILITY, AND IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. THE CUSTOMER INDICATED THAT THEY DID NOT CONSIDER THE PATIENT¿S DEATH TO BE RELATED TO THE DEVICE. THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO UPDATE FIELDS (B5 AND F9) AND ADDITIONAL INFORMATION PROVIDED BY THE CUSTOMER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22161484&pc=NPT | 2025587-2025-04104 | 2025-04-01 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-06 04:00 | NPT | MEDTRONIC TRANSCATHETER VALVE | Gradient Increase; Patient Device Interaction Problem |  | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE PREDICTIVE VALUE OF INVASIVE AND ECHOCARDIOGRAPHY-DERIVED MEAN GRADIENT MEASUREMENTS AFTER TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI). THE STUDY POPULATION CONSISTED OF 1 ,227 PATIENTS WHO UNDERWENT TAVI WITH EITHER A SELF-EXPANDING VALVE (SEV) OR BALLOON-EXPANDABLE VALVE (BEV). MEDTRONIC EVOLUT R/PRO/PRO+ VALVE BRANDS WERE USED IN THE SEV GROUP (N = 573), WHILE NON-MEDTRONIC SAPIEN VALVES WERE USED IN THE BEV GROUP (N = 654). THE AUTHORS ULTIMATELY FOUND THAT ONLY INVASIVE-DERIVED MEAN GRADIENTS POST-TAVI WERE ASSOCIATED WITH EARLY, ONE-YEAR, AND TWO-YEAR MORTALITY IN BOTH THE SEV AND BEV GROUPS. SPECIFICALLY, A TOTAL OF 98 DEATHS OCCURRED IN THE SEV GROUP WITHIN TWO YEARS OF TAVI. ADDITIONAL ADVERSE OUTCOMES THAT HAPPENED IN THE SEV GROUP: ELEVATED/HIGH MEAN GRADIENTS, REHOSPITALIZATION FOR HEART FAILURE, AND BIOPROSTHETIC VALVE FAILURE (NO FURTHER DETAILS GIVEN). Manufacturer Narrative: CITATION: MARK M.P. VAN DEN DORPEL, ET AL. PROGNOSTIC VALUE OF INVASIVE VERSUS ECHOCARDIOGRAPHY-DERIVED AORTIC GRADIENT IN PATIENTS UNDERGOING TAVI. EUROINTERVENTION 2025;21:E411-E425. PUBLISHED ONLINE E-EDITION APRIL 2025. DOI: 10.4244/EIJ-D-24-00341. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH: (B)(6) 2025 (MONTH AND YEAR VALID). EARLIEST APPROVED EVOLUT PRODUCT USED FOR PRODUCT CODE AND PMA#. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22161532&pc=DXE | 2124215-2025-35733 | 2025-05-03 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-06 04:00 | DXE | ANGIOJET? SOLENT? OMNI | Adverse Event Without Identified Device or Use Problem | Ischemia; Septic Shock; No Clinical Signs, Symptoms or Conditions | K111182 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE TARGET LESION WAS LOCATED IN THE POPLITEAL ARTERY. THE PROCEDURE WENT WELL, HOWEVER, THE DOCTOR NOTED THAT THE PATIENT MIGHT DEVELOP COMPARTMENT SYNDROME. NO COMPLICATIONS RELATED TO THE DEVICE, OR THE PROCEDURE WERE NOTED; HOWEVER, THE PATIENT SUBSEQUENTLY DECEASED.Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE TARGET LESION WAS LOCATED IN THE POPLITEAL ARTERY. THE PROCEDURE WENT WELL, HOWEVER, THE DOCTOR NOTED THAT THE PATIENT MIGHT DEVELOP COMPARTMENT SYNDROME. POST-PROCEDURE, THE PATIENT EXHIBITED SOME UNKNOWN COMORBIDITIES. NO COMPLICATIONS RELATED TO THE DEVICE, OR THE PROCEDURE WERE NOTED; HOWEVER, THE PATIENT SUBSEQUENTLY DECEASED.Event Description: LOCATIONIT WAS REPORTED THAT THE PATIENT DIED. THE TARGET LESION WAS LOCATED IN THE POPLITEAL ARTERY. THE PROCEDURE WENT WELL, HOWEVER, THE DOCTOR NOTED THAT THE PATIENT MIGHT DEVELOP COMPARTMENT SYNDROME. POST-PROCEDURE, THE PATIENT EXHIBITED SOME UNKNOWN COMORBIDITIES. NO COMPLICATIONS RELATED TO THE DEVICE, OR THE PROCEDURE WERE NOTED; HOWEVER, THE PATIENT SUBSEQUENTLY DECEASED. IT WAS FURTHER REPORTED THAT THE PROCEDURE TOOK PLACE IN THE MID TO LATE AFTERNOON. THE ANGIOJET WAS USED IN THE POPLITEAL ARTERY WHICH WAS COMPLETELY OCCLUDED ALL THE WAY DOWN. THERE WERE NO ISSUES WITH THE ANGIOJET CATHETER. DUE TO THE PATIENT'S LEG BEING COLD FOR A LONG PERIOD OF TIME (DUE TO THE OCCLUSION/LACK OF BLOOD FLOW), IT WOULD HAVE BEEN IDEAL TO GRADUALLY INTRODUCE BLOOD FLOW BACK TO THE ARTERY. HOWEVER, IT WAS NOTED THAT THE ANGIOJET WAS ABLE TO OPEN THE VESSEL VERY QUICKLY WHICH IS WHY IT WAS THOUGHT THAT THE PATIENT MAY EXPERIENCE COMPARTMENT SYNDROME. DURING RECOVERY, THE PATIENTS LEG BEGAN SWELLING. THE PATIENT WAS TAKEN STRAIGHT TO THE OPERATING ROOM, AND A UNILATERAL FASCIOTOMY WAS PERFORMED. POST-PROCEDURE, THE PATIENT SHOWED SIGNS OF SHOCK, WITH ABNORMAL LAB RESULTS. THE PATIENT'S CONDITION DETERIORATED OVERNIGHT, AND THEY WERE DECLARED BRAIN DEAD THE FOLLOWING MORNING. THE PATIENT PASSED AWAY THAT DAY. THE EXACT CAUSE OF DEATH WAS NOT CONFIRMED BUT WAS BELIEVED TO BE SOME FORM OF SHOCK, POSSIBLY SEPTIC SHOCK. Manufacturer Narrative: B5 DESCRIBE EVENT OR PROBLEM: UPDATED. H6 - PATIENT CODES: UPDATED. H6 - IMPACT CODES: UPDATED. Manufacturer Narrative: H6 - PATIENT CODES: UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22161565&pc=OUT | 2029214-2025-01325 | 2025-05-27 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-06-06 04:00 | OUT | PIPELINE VANTAGE WITH SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Paralysis; Ischemia Stroke; Thrombosis/Thrombus | P100018 |  | 1 | Event Description: MEDTRONIC RECEIVED A REPORT THAT A PATIENT EXPERIENCED RIGHT HEMIPLEGIA POST PROCEDURE. THE PATIENT WAS UNDERGOING SURGERY FOR TREATMENT OF A SACCULAR, UNRUPTURED ANEURYSM OF THE LEFT POSTERIOR COMMUNICATING ARTERY (PCOM) WITH A MAX DIAMETER OF 3.2MM AND A 3.5MM NECK DIAMETER. THE LANDING ZONE WAS 2.6MM DISTALLY AND 3.2MM PROXIMALLY. IT WAS NOTED THE PATIENT'S VESSEL TORTUOSITY WAS NORMAL. DAPT (DUAL ANTIPLATELET TREATMENT) WAS ADMINISTERED; ASPIRIN AND BRILLENTA. THE PRU LEVEL WAS 188. THE ANGIOGRAPHIC RESULT POST PROCEDURE SHOWED THE DEVICE WAS PATENT AND ALL VESSELS WERE FILLING NORMALLY. IT WAS REPORTED THAT THE INDEX PROCEDURE WENT WELL ON (B)(6) 2025 AND THE PATIENT WENT TO ROOM POST PROCEDURE. LATER IN THE AFTERNOON THE PATIENT DEVELOPED SYMPTOMS. THE PATIENT WENT FOR CT AND SMALL STROKES WERE OBSERVED. THE PATIENT WAS GIVEN INTEGRILLIN AND WAS DISCHARGED ON (B)(6). THE PATIENT CAME BACK TO THE ER ON (B)(6), DUE TO SYMPTOMS (RIGHT HEMIPLEGIA) AND CTA SHOWED THROMBUS AT NECK OF THE ANEURYSM, BUT THE DEVICE WAS OPEN. THE PATIENT IS NOW ADMITTED AND ON INTEGRILLIN. THE PIPELINE WAS USED FOR AN INDICATION THAT IS APPROVED (ON-LABEL). THE REPORTED DEVICE AND ANY ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU).Event Description: ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE PATIENT PASSED AWAY. THE PATIENT EXPERIENCED STROKE. THE ENTIRE PIPE CLOTTED AND PATIENT HAD A MASSIVE STROKE DESPITE INTEGRILLIN DRIP AND THROMBECTOMY. THE CAUSE OF THE EVENT WAS NOT DETERMINED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT PASSED AWAY IN (B)(6) 2025 AT (B)(6) HOSPITAL. NO AUTOPSY RECORDS WILL BE AVAILABLE. THE DEATH IS THOUGHT TO BE RELATED TO THE MEDTRONIC DEVICE OR THERAPY. THE ENTIRE DEVICE CLOTTED OFF. CODE STROKE WAS CALLED. PATIENT HAD MASSIVE STROKE DESPITE INTEGRILIN DRIP AND THROMBECTOMY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B2: THE DATE OF DEATH WAS REPORTED AS (B)(6) 2025. EXACT DATE NOT REPORTED, MONTH AND YEAR VALID. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22161651&pc=DTZ | 2916596-2025-03320 | 2025-05-08 04:00 | Death | THORATEC CORPORATION | 2025-06-06 04:00 | DTZ | CENTRIMAG ADULT ECMO KIT (US) | Improper or Incorrect Procedure or Method; Gas/Air Leak; No Apparent Adverse Event | Insufficient Information; No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PRE-CONNECTED PACK HAD ENTRAPPED AIR DURING THE PRIMING PROCESS, SEEMINGLY CAUSING THE PUMP TO MALFUNCTION. THE PRE-CONNECTED PACK WAS BEING PREPPED BY LESS COMFORTABLE STAFF THAT SEEMED TO ENTRAP A FAIR AMOUNT OF AIR IN THE CIRCUIT CAUSING THE PUMP TO SHUT DOWN. THIS WAS NOT DEEMED A DEVICE FAILURE BY THE PRACTITIONERS.Event Description: AFTER FURTHER SPEAKING WITH THE CLINICIANS, THEY DID NOT FAULT THE DEVICE AT ALL. THE PATIENT WAS NOTED TO BE CRITICALLY ILL AND HAD MANY UNDERLYING CONDITIONS. THE PHYSICIANS DID NOT WANT TO PURSUE THIS AS A DEVICE FAILURE AND DID NOT WISH TO EXPLORE THIS FURTHER. NO PRODUCT WOULD BE RETURNED. NO LOG FILES WERE AVAILABLE. Manufacturer Narrative: SECTION A - NO PATIENT INFORMATION WAS PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22161816&pc=DSQ | 2916596-2025-03700 | 2025-05-06 04:00 | Death | THORATEC CORPORATION | 2025-06-06 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Encephalopathy; Renal Failure; Cardiogenic Shock; Melena; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS LATER REPORTED THAT THE HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS INSUFFICIENT TO SUPPORT THE PATIENT'S CARDIOGENIC SHOCK AND POST LVAD THE PATIENT ENDED UP BEING CONVERTED TO PERIPHERAL VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (VV-ECMO) WITH LVAD. THE PATIENT EXPERIENCED MELENA, CHOLELITHIASIS OF CRITICAL ILLNESS, HISTOPLASMOSIS, ENTEROCOCCUS, STERNAL INFECTION, AND ENCEPHALOPATHY; ECMO, CONTINUOUS RENAL REPLACEMENT THERAPY, AND NUMEROUS CHEST WASHOUTS WERE PERFORMED. DUE TO THE SEVERITY OF THE PATIENT'S ILLNESS, THE PATIENT WAS NO LONGER DEEMED A TRANSPLANT CANDIDATE AND WAS TRANSITIONED TO COMFORT CARE WHERE THEY PASSED LATER THAT DAY. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE OR PUMP POCKET INFECTION), RENAL FAILURE, OTHER NEUROLOGICAL EVENTS (NOT STROKE-RELATED), AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AND NEUROLOGICAL DYSFUNCTION AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING INFECTION PREVENTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22161935&pc=DSQ | 2916596-2025-03393 | 2025-02-26 05:00 | Death | THORATEC CORPORATION | 2025-06-06 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO UNKNOWN ETIOLOGY. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT BE EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22161983&pc=FGA | 9610773-2025-03851 | 2025-05-05 04:00 | Death | OLYMPUS WINTER & IBE GMBH | 2025-06-06 04:00 | FGA | OP TELESCOPE, 7°, 11 FR. X 220 MM, 7.5 FR. CHANNEL, ANGLED OCULAR | Material Twisted/Bent | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Sepsis; Septic Shock; Ventilator Dependent | K853486 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT CYSTOSCOPY AND HOLMIUM LASER OR BOTH PENILE AND BULBAR URETHRAL STRICTURES, COMPLEX URETHRAL DILATION, PLACEMENT OF A FOLEY CATHETER, REMOVAL OF THE ARTIFICIAL URINARY SPHINCTER, AND PLACEMENT OF A SUPRAPUBIC CYSTOSTOMY TUBE. AFTER THE PROCEDURE WAS PERFORMED, STAFF FOUND THAT THE 11 FRENCH 7 DEGREE PEDIATRIC CYSTOSCOPE (THE ONLY SCOPE THAT COULD PASS THROUGH PROXIMAL PENILE URETHRAL STRICTURE) WAS SLIGHTLY BENT. PATIENT RECOVERED WELL WITH MINIMAL BLEEDING IN POST ANESTHESIA CARE UNIT (PACU) AND WAS DISCHARGED HOME WITH NO COMPLICATIONS. THE PATIENT LATER REPORTED WAKING UP AT HOME TO FIND BED SATURATED WITH BLOOD. THE PATIENT WENT TO LOCAL EMERGENCY ROOM AND WAS SUBSEQUENTLY TRANSFERRED BACK TO THE FACILITY WITH ANEMIA AND HYPOTENSION. THE PATIENT WAS PLACED ON PRESSORS, VENTILATED, AND TREATED FOR SEPSIS. PATIENT WAS THEN TRANSFERRED TO A HIGHER LEVEL OF CARE FOR PROCEDURES UNABLE TO BE PERFORMED AT THE FACILITY. CUSTOMER NOTES THAT IT IS UNKNOWN WHETHER THE DEVICE IS RELATED TO OUTCOMES.Event Description: ADDITIONAL INFORMATION RECEIVED THE PATIENT WAS RE-ADMITTED TO THE FACILITY, RECEIVED BLOOD TRANSFUSIONS, WAS TREATED FOR SEPTIC SHOCK, AND WAS ADMINISTERED THE ANTIBIOTICS ZOSYN AND LINEZOLID. THE PATIENT WAS SUBSEQUENTLY TRANSFERRED FOR SLEDD (SUSTAINED LOW-EFFICIENCY DAILY DIALYSIS) AND HAS SINCE EXPIRED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED FROM THE CUSTOMER INDICATING THAT THE PATIENT WAS RE-ADMITTED TO A DIFFERENT FACILITY FOR TREATMENT OF AN INFECTION. THE PATIENT SUBSEQUENTLY EXPIRED AT THAT FACILITY, AND IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. THE CUSTOMER INDICATED THAT THEY DID NOT CONSIDER THE PATIENT¿S DEATH TO BE RELATED TO THE DEVICE. Manufacturer Narrative: E2 - HEALTH PROFESSIONAL: WAS UNABLE TO BE CHANGED FOR NO BUT SHOULD BE YES. SINCE THE DEVICE DESCRIBED WAS A "11 FRENCH 7-DEGREE PEDIATRIC CYSTOSCOPE", WE SELECTED "(B)(6)" AS A REPRESENTATIVE PRODUCT. THE INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED OR IF ADDITIONAL INFORMATION BECOMES AVAILABLE. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE ADDITIONAL INFORMATION AND THE RESULTS OF THE LEGAL MANUFACTURER'S FINAL INVESTIGATION. UPDATED FIELDS: B5, D9, H4 AND H6 THE DEVICE WAS NOT RETURNED TO OLYMPUS FOR EVALUATION. HOWEVER, THE REPORTED ISSUE OF A BENT SHAFT WAS CONFIRMED BASED ON THE TWO PHOTOS PROVIDED BY THE CUSTOMER, SHOWING VISIBLE BENDING NEAR THE PROXIMAL END. IN ADDITION, THE CUSTOMER PROVIDED A PACKING SLIP WHICH INDICATED THAT THE DEVICE WAS SHIPPED TO A THIRD PARTY AFTER THE REPORTED EVENT OCCURRED. THE PACKING SLIP LISTED INCOMING OBSERVATIONS INCLUDING A BENT SHAFT, PHYSICALLY DAMAGED IMAGE BUNDLE, DAMAGED LIGHT FIBERS AT THE DISTAL END, AND OPTICAL GLUE DAMAGE AT THE OBJECTIVE LENS. THE DEVICE WAS SUBSEQUENTLY REPAIRED BY A THIRD PARTY, INCLUDING IMAGE BUNDLE REPLACEMENT, POLISHING OF LIGHT FIBERS, TUBE REALIGNMENT, AND OBJECTIVE LENS ASSEMBLY REPLACEMENT. A DEVICE HISTORY REVIEW WAS CONDUCTED, AND NO ISSUES WERE IDENTIFIED AND THERE WAS NO SERVICE HISTORY FOR THIS DEVICE. A DEFINITIVE ROOT CAUSE OF THE BENT SHAFT COULD NOT BE IDENTIFIED, AND THE INVESTIGATION FINDINGS DO NOT LEAD TO A CLEAR CONCLUSION ABOUT THE CAUSE OF THE REPORTED ADVERSE EVENT. FURTHER, THE REPORTED ADVERSE EVENTS WERE ASSESSED AS UNLIKELY RELATED TO THE DEVICE. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO UPDATE FIELDS (A2, A3, A4, A5, A6, B1, B2, B5, B7, H1, H6) AND ATTACH THE RELATED MEDWATCH REPORT. ADDITIONALLY, THIS SUPPLEMENTAL REPORT IS PROVIDING FURTHER INFORMATION SHARED BY THE CUSTOMER. IT WAS DETERMINED BY THE OLYMPUS UROLOGY MEDICAL DIRECTOR THAT IT IS UNLIKELY THE OLYMPUS DEVICE CAUSED OR CONTRIBUTED TO THE HEMORRHAGE/BLOOD LOSS/BLEEDING IN THIS CASE. BLEEDING IS A KNOWN PROCEDURAL COMPLICATION AND IS MORE LIKELY RELATED TO THE PATIENT¿S COMPLEX ANATOMY AND COMPLEX UNDERLYING MEDICAL HISTORY. IT WAS ALSO DETERMINED THAT IT IS UNLIKELY THAT THE OLYMPUS DEVICE CAUSED OR CONTRIBUTED TO SEPSIS AS THE PATIENT PRESENTED WITH AN INFECTED OR ERODED ARTIFICIAL URINARY SPHINCTER THAT REQUIRED REMOVAL. THE SEPSIS IN THIS CASE IS MORE LIKELY RELATED TO THE PATIENT¿S PRESENTING DIAGNOSIS AS NOTED ABOVE. BASED ON THE INFORMATION PROVIDED IT IS UNKNOWN IF THE DEATH IN THIS CASE IS A RESULT OF THE BLEEDING OR SEPSIS REPORTED. HOWEVER, BASED UPON THE CURRENT KNOWN INFORMATION IT WAS DETERMINED THE DEATH AS UNLIKELY RELATED TO THE OLYMPUS DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22147378&pc=FRN | 2016493-2025-85046 | 2025-04-13 04:00 | Death | CAREFUSION SD | 2025-06-05 04:00 | FRN | ALARIS SYSTEM | Use of Device Problem; Inaccurate Delivery | Cardiac Arrest; Low Blood Pressure/ Hypotension | K211218 |  | 1 | Event Description: INITIALLY, THE CUSTOMER REQUESTED LOG REVIEW FOR AN UNSPECIFIED ISSUE. AFTER FOLLOW UP BY CLINICAL SPECIALIST, IT WAS REPORTED THAT THERE WAS A PROGRAMMING ERROR WITH A STAT INFUSION OF EPINEPHRINE TO A PATIENT ADMITTED FOR CARDIAC ARREST. THE EPINEPHRINE (0.1MCG/KG/MIN THEN TITRATABLE) WAS INITIALLY PROGRAMMED USING INTEROPERABILITY. THE TOTAL BAG VOLUME AND CONCENTRATION WAS 4MG IN 250ML. AFTER THE INITIAL PROGRAMMING, THE PUMP WAS STOPPED. LATER, THE PUMP WAS RESTARTED WITH MANUAL PROGRAMMING. AFTER THE EPINEPHRINE INFUSION WAS RESTARTED, THE CLINICIAN NOTICED THAT THE PUMP WAS SET TO NON-WEIGHT-BASED DOSING (MCG/MINUTE) HOWEVER THE MEDICATION ORDER AND TITRATIONS WERE INTENDED TO BE ADMINISTERED USING WEIGHT-BASED DOSING. IT WAS REPORTED THAT THE "PATIENT ENDED UP BEING STARTED ON ADDITIONAL MEDICATIONS BECAUSE IT WAS THOUGHT THAT THEY WERE MAXED ON THE CURRENT MEDICATION." FOLLOWING THE EVENT, THE PATIENT EXPERIENCED HYPOTENSION AND WAS STARTED ON DOBUTAMINE. THE PATIENT EXPIRED. IT WAS NOT CLEAR TO WHAT DEGREE THE INFUSION ERROR IMPACTED THE PATIENT ALREADY ADMITTED FOR CARDIAC ARREST.Event Description: INITIALLY, THE CUSTOMER REQUESTED LOG REVIEW FOR AN UNSPECIFIED ISSUE. AFTER FOLLOW UP BY CLINICAL SPECIALIST, IT WAS REPORTED THAT THERE WAS A PROGRAMMING ERROR WITH A STAT INFUSION OF EPINEPHRINE TO A PATIENT ADMITTED FOR CARDIAC ARREST. THE EPINEPHRINE (0.1MCG/KG/MIN THEN TITRATABLE) WAS INITIALLY PROGRAMMED USING INTEROPERABILITY. THE TOTAL BAG VOLUME AND CONCENTRATION WAS 4MG IN 250ML. AFTER THE INITIAL PROGRAMMING, THE PUMP WAS STOPPED. LATER, THE PUMP WAS RESTARTED WITH MANUAL PROGRAMMING. AFTER THE EPINEPHRINE INFUSION WAS RESTARTED, THE CLINICIAN NOTICED THAT THE PUMP WAS SET TO NON-WEIGHT-BASED DOSING (MCG/MINUTE) HOWEVER THE MEDICATION ORDER AND TITRATIONS WERE INTENDED TO BE ADMINISTERED USING WEIGHT-BASED DOSING. IT WAS REPORTED THAT THE "PATIENT ENDED UP BEING STARTED ON ADDITIONAL MEDICATIONS BECAUSE IT WAS THOUGHT THAT THEY WERE MAXED ON THE CURRENT MEDICATION." FOLLOWING THE EVENT, THE PATIENT EXPERIENCED HYPOTENSION AND WAS STARTED ON DOBUTAMINE. THE PATIENT EXPIRED. IT WAS NOT CLEAR TO WHAT DEGREE THE INFUSION ERROR IMPACTED THE PATIENT ALREADY ADMITTED FOR CARDIAC ARREST. Manufacturer Narrative: THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. Manufacturer Narrative: OMIT: C20 - NO FINDINGS AVAILABLE, D15 - CAUSE NOT ESTABLISHED. ADDITIONAL INFORMATION: IMDRF ANNEX A, B, C, D, G CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY: THE OCCURRENCE OF A PROGRAMMING ERROR WAS CONFIRMED THROUGH REVIEW OF THE DEVICE LOGS BUT WAS NOT REPLICATED THROUGH DEVICE TESTING SINCE THEY WERE NOT PROVIDED BY THE FACILITY. ON (B)(6) 2025, AT 1:10 AM, A SAFETY CLAMP OPEN ALARM WAS OBSERVED, FOLLOWED BY A CHECK IV SET ALARM BEFORE THE UNIT WAS CHANNELED OFF. AT 5:16 AM THE CHANNEL WAS SELECTED AND ANOTHER INFUSION FOR AN UNKNOWN DRUG (DRUGID 503, SUSPECTED TO BE EPINEPHRINE) WAS PROGRAMMED MANUALLY AND STARTED AT A CONCENTRATION OF 4 MG / 250 ML (NOT WEIGHT-BASED). THE DOSE WAS 0.1 MCG/MIN, CALCULATING THE RATE AT 0.375 ML/HR AND THE VOLUME TO BE INFUSED (VTBI) WAS 200 ML. BETWEEN 5:26 AM AND 5:33 AM, THE DOSE WAS INCREASED FOUR (4) TIMES IN INCREMENTS OF 0.1 MCG/MINUTE UNTIL IT WAS SETTLED ON 0.5 MCG/MINUTES, CALCULATING A RATE OF 1.875 ML/HR. THE VTBI WAS CHANGED TO 100 ML AT 5:45 AM. AT 11:34 AM, A REMOTE INFUSION (INTEROPERABILITY WORKFLOW) (WEIGHT-BASED) WAS RECEIVED AND STARTED FOR AN UNKNOWN DRUG (DRUGID 642) AT A CONCENTRATION OF 25000 UNITS / 250 ML, A DOSE OF 18 UNITS/HR (RATE OF 8.64 ML/HR) AND A VTBI OF 250 ML. THE PATIENT WEIGHT WAS 48 KG. THE UNIT WAS CHANNELED OFF AT 4:06 PM AND WAS THEN SELECTED AGAIN AT 4:07 PM WHEN A REMOTE INFUSION (WEIGHT-BASED) WAS RECEIVED FOR AN UNKNOWN DRUG (DRUGID 509, SUSPECTED TO BE WEIGHT-BASED EPINEPHRINE) AT A CONCENTRATION OF 4 MG / 250 ML, A DOSE OF 0.0104 MCG/HR, A RATE OF 0 ML/HR AND A VTBI OF 250 ML. HOWEVER, THE INFUSION WAS NOT STARTED. INSTEAD, THE PREVIOUS INFUSION (WEIGHT-BASED) WAS RESTORED AND WAS STARTED ONCE AGAIN AT 4:09 PM. THE INFUSION CONCENTRATION WAS 25000 UNITS / 250 ML, THE DOSE WAS 18 UNITS/HR, THE RATE WAS 8.64 ML/HR AND THE VTBI WAS 250 ML. THE PATIENT WEIGHT WAS 48 KG. AT 6:18 PM, THE DOSE WAS CHANGED TO 15 UNITS/HR (RE-CALCULATING THE RATE TO 7.2 ML/HR) AND A 60-MINUTE DELAY WAS PROGRAMMED. THE INFUSION WAS RESTARTED AT 7:20 PM. THE SYSTEM WAS POWERED OFF ON (B)(6) 2025 AT 12:12 AM. THE TOTAL VOLUME INFUSED (TVI) WAS RECORDED AT 95.843 ML. NO DEVICES WERE RETURNED FOR LABORATORY TESTING; THEREFORE, IT WAS NOT POSSIBLE TO PERFORM AN INTERNAL, EXTERNAL INSPECTION OR LABORATORY TESTING. PROPER OPERATION OF THE BD ALARIS¿ SYSTEM REQUIRES THAT YOU ARE FAMILIAR WITH RELATED FEATURES, SETUP, PROGRAMMING, IV SETS, AND ACCESSORIES. READ ALL INSTRUCTIONS, INCLUDING THOSE FOR ALL ATTACHED MODULE(S) BEFORE USING THE BD ALARIS ¿ SYSTEM. THE DEVICE WAS IN USE FOR TREATMENT PURPOSES AS INTENDED PER 21 CFR 820.198(D)(2). ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED INCIDENT IS ATTRIBUTED TO USER PROGRAMMING ERROR, AS REPORTED BY THE FACILITY AND SUPPORTED BY REVIEW OF THE DEVICE LOGS. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. A REMOTE DRUG ORDER WAS ATTEMPTED ON (B)(6) 2025 AT 6:04 AM BUT WAS NOT COMPLETED, AND INSTEAD THE USER CHANGED THE DOSE FROM THE CURRENTLY INFUSING NON-WEIGHT-BASED ORDER FROM 0.5MCG/MIN TO 0.4MCG/MIN. THE INFUSION WAS STOPPED AT 6:52 AM. AT 11:34 AM A NEW INFUSION OF A DIFFERENT DRUG WITH DIFFERENT DRUG PARAMETERS WAS PROGRAMMED. THAT INFUSION WAS STOPPED AT 4:06 PM. AT 4:07 PM A REMOTE ORDER WAS ATTEMPTED FOR WEIGHT-BASED EPINEPHRINE. THAT PROGRAMING WAS NOT COMPLETED, INSTEAD THE USER WENT BACK AND RESTORED THE PREVIOUS INFUSION WHICH WAS NOT EPINEPHRINE AND HAD A RATE OF 8.64 ML/HOUR. THE INITIAL INFUSION FOR AN UNKNOWN DRUG (SUSPECTED TO BE EPINEPHRINE) WAS PROGRAMMED REMOTELY (INTEROPERABILITY WORKFLOW) AT A CONCENTRATION OF 4 MG / 250 ML AND WAS STARTED ON (B)(6) 2025 AT 1:00 PM (NOT WEIGHT-BASED). THE DOSE WAS 1 MCG/MIN, CALCULATING THE RATE AT 3.75 ML/HR AND THE VOLUME TO BE INFUSED (VTBI) WAS 250 ML. THERE WAS NO ALLEGATION OF A DEVICE MALFUNCTION, AND NO MALFUNCTION WAS IDENTIFIED DURING THE LOG REVIEW.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22147388&pc=FRN | 2016493-2025-85045 | 2025-04-13 04:00 | Death | CAREFUSION SD | 2025-06-05 04:00 | FRN | ALARIS SYSTEM | Use of Device Problem; Inaccurate Delivery | Cardiac Arrest; Low Blood Pressure/ Hypotension | K211218 |  | 1 | Event Description: INITIALLY, THE CUSTOMER REQUESTED LOG REVIEW FOR AN UNSPECIFIED ISSUE. AFTER FOLLOW UP BY CLINICAL SPECIALIST, IT WAS REPORTED THAT THERE WAS A PROGRAMMING ERROR WITH A STAT INFUSION OF EPINEPHRINE TO A PATIENT ADMITTED FOR CARDIAC ARREST. THE EPINEPHRINE (0.1MCG/KG/MIN THEN TITRATABLE) WAS INITIALLY PROGRAMMED USING INTEROPERABILITY. THE TOTAL BAG VOLUME AND CONCENTRATION WAS 4MG IN 250ML. AFTER THE INITIAL PROGRAMMING, THE PUMP WAS STOPPED. LATER, THE PUMP WAS RESTARTED WITH MANUAL PROGRAMMING. AFTER THE EPINEPHRINE INFUSION WAS RESTARTED, THE CLINICIAN NOTICED THAT THE PUMP WAS SET TO NON-WEIGHT-BASED DOSING (MCG/MINUTE) HOWEVER THE MEDICATION ORDER AND TITRATIONS WERE INTENDED TO BE ADMINISTERED USING WEIGHT-BASED DOSING. IT WAS REPORTED THAT THE "PATIENT ENDED UP BEING STARTED ON ADDITIONAL MEDICATIONS BECAUSE IT WAS THOUGHT THAT THEY WERE MAXED ON THE CURRENT MEDICATION." FOLLOWING THE EVENT, THE PATIENT EXPERIENCED HYPOTENSION AND WAS STARTED ON DOBUTAMINE. THE PATIENT EXPIRED. IT WAS NOT CLEAR TO WHAT DEGREE THE INFUSION ERROR IMPACTED THE PATIENT ALREADY ADMITTED FOR CARDIAC ARREST.Event Description: INITIALLY, THE CUSTOMER REQUESTED LOG REVIEW FOR AN UNSPECIFIED ISSUE. AFTER FOLLOW UP BY CLINICAL SPECIALIST, IT WAS REPORTED THAT THERE WAS A PROGRAMMING ERROR WITH A STAT INFUSION OF EPINEPHRINE TO A PATIENT ADMITTED FOR CARDIAC ARREST. THE EPINEPHRINE (0.1MCG/KG/MIN THEN TITRATABLE) WAS INITIALLY PROGRAMMED USING INTEROPERABILITY. THE TOTAL BAG VOLUME AND CONCENTRATION WAS 4MG IN 250ML. AFTER THE INITIAL PROGRAMMING, THE PUMP WAS STOPPED. LATER, THE PUMP WAS RESTARTED WITH MANUAL PROGRAMMING. AFTER THE EPINEPHRINE INFUSION WAS RESTARTED, THE CLINICIAN NOTICED THAT THE PUMP WAS SET TO NON-WEIGHT-BASED DOSING (MCG/MINUTE) HOWEVER THE MEDICATION ORDER AND TITRATIONS WERE INTENDED TO BE ADMINISTERED USING WEIGHT-BASED DOSING. IT WAS REPORTED THAT THE "PATIENT ENDED UP BEING STARTED ON ADDITIONAL MEDICATIONS BECAUSE IT WAS THOUGHT THAT THEY WERE MAXED ON THE CURRENT MEDICATION." FOLLOWING THE EVENT, THE PATIENT EXPERIENCED HYPOTENSION AND WAS STARTED ON DOBUTAMINE. THE PATIENT EXPIRED. IT WAS NOT CLEAR TO WHAT DEGREE THE INFUSION ERROR IMPACTED THE PATIENT ALREADY ADMITTED FOR CARDIAC ARREST. Manufacturer Narrative: THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. Manufacturer Narrative: OMIT: C20 - NO FINDINGS AVAILABLE, D15 - CAUSE NOT ESTABLISHED. ADDITIONAL INFORMATION: IMDRF ANNEX A, B, C, D, G CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. INVESTIGATION SUMMARY: THE OCCURRENCE OF A PROGRAMMING ERROR WAS CONFIRMED THROUGH REVIEW OF THE DEVICE LOGS BUT WAS NOT REPLICATED THROUGH DEVICE TESTING SINCE THEY WERE NOT PROVIDED BY THE FACILITY. ON 14APR2025, AT 1:10 AM, A SAFETY CLAMP OPEN ALARM WAS OBSERVED, FOLLOWED BY A CHECK IV SET ALARM BEFORE THE UNIT WAS CHANNELED OFF. AT 5:16 AM THE CHANNEL WAS SELECTED AND ANOTHER INFUSION FOR AN UNKNOWN DRUG (DRUGID 503, SUSPECTED TO BE EPINEPHRINE) WAS PROGRAMMED MANUALLY AND STARTED AT A CONCENTRATION OF 4 MG / 250 ML (NOT WEIGHT-BASED). THE DOSE WAS 0.1 MCG/MIN, CALCULATING THE RATE AT 0.375 ML/HR AND THE VOLUME TO BE INFUSED (VTBI) WAS 200 ML. BETWEEN 5:26 AM AND 5:33 AM, THE DOSE WAS INCREASED FOUR (4) TIMES IN INCREMENTS OF 0.1 MCG/MINUTE UNTIL IT WAS SETTLED ON 0.5 MCG/MINUTES, CALCULATING A RATE OF 1.875 ML/HR. THE VTBI WAS CHANGED TO 100 ML AT 5:45 AM. AT 11:34 AM, A REMOTE INFUSION (INTEROPERABILITY WORKFLOW) (WEIGHT-BASED) WAS RECEIVED AND STARTED FOR AN UNKNOWN DRUG (DRUGID 642) AT A CONCENTRATION OF 25000 UNITS / 250 ML, A DOSE OF 18 UNITS/HR (RATE OF 8.64 ML/HR) AND A VTBI OF 250 ML. THE PATIENT WEIGHT WAS 48 KG. THE UNIT WAS CHANNELED OFF AT 4:06 PM AND WAS THEN SELECTED AGAIN AT 4:07 PM WHEN A REMOTE INFUSION (WEIGHT-BASED) WAS RECEIVED FOR AN UNKNOWN DRUG (DRUGID 509, SUSPECTED TO BE WEIGHT-BASED EPINEPHRINE) AT A CONCENTRATION OF 4 MG / 250 ML, A DOSE OF 0.0104 MCG/HR, A RATE OF 0 ML/HR AND A VTBI OF 250 ML. HOWEVER, THE INFUSION WAS NOT STARTED. INSTEAD, THE PREVIOUS INFUSION (WEIGHT-BASED) WAS RESTORED AND WAS STARTED ONCE AGAIN AT 4:09 PM. THE INFUSION CONCENTRATION WAS 25000 UNITS / 250 ML, THE DOSE WAS 18 UNITS/HR, THE RATE WAS 8.64 ML/HR AND THE VTBI WAS 250 ML. THE PATIENT WEIGHT WAS 48 KG. AT 6:18 PM, THE DOSE WAS CHANGED TO 15 UNITS/HR (RE-CALCULATING THE RATE TO 7.2 ML/HR) AND A 60-MINUTE DELAY WAS PROGRAMMED. THE INFUSION WAS RESTARTED AT 7:20 PM. THE SYSTEM WAS POWERED OFF ON 15APR2025 AT 12:12 AM. THE TOTAL VOLUME INFUSED (TVI) WAS RECORDED AT 95.843 ML. NO DEVICES WERE RETURNED FOR LABORATORY TESTING; THEREFORE, IT WAS NOT POSSIBLE TO PERFORM AN INTERNAL, EXTERNAL INSPECTION OR LABORATORY TESTING. PROPER OPERATION OF THE BD ALARIS¿ SYSTEM REQUIRES THAT YOU ARE FAMILIAR WITH RELATED FEATURES, SETUP, PROGRAMMING, IV SETS, AND ACCESSORIES. READ ALL INSTRUCTIONS, INCLUDING THOSE FOR ALL ATTACHED MODULE(S) BEFORE USING THE BD ALARIS ¿ SYSTEM. THE DEVICE WAS IN USE FOR TREATMENT PURPOSES AS INTENDED PER 21 CFR 820.198(D)(2). ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED INCIDENT IS ATTRIBUTED TO USER PROGRAMMING ERROR, AS REPORTED BY THE FACILITY AND SUPPORTED BY REVIEW OF THE DEVICE LOGS. A REMOTE DRUG ORDER WAS ATTEMPTED ON (B)(6) 2025 AT 6:04 AM BUT WAS NOT COMPLETED, AND INSTEAD THE USER CHANGED THE DOSE FROM THE CURRENTLY INFUSING NON-WEIGHT-BASED ORDER FROM 0.5MCG/MIN TO 0.4MCG/MIN. THE INFUSION WAS STOPPED AT 6:52 AM. AT 11:34 AM A NEW INFUSION OF A DIFFERENT DRUG WITH DIFFERENT DRUG PARAMETERS WAS PROGRAMMED. THAT INFUSION WAS STOPPED AT 4:06 PM. AT 4:07 PM A REMOTE ORDER WAS ATTEMPTED FOR WEIGHT-BASED EPINEPHRINE. THAT PROGRAMING WAS NOT COMPLETED, INSTEAD THE USER WENT BACK AND RESTORED THE PREVIOUS INFUSION WHICH WAS NOT EPINEPHRINE AND HAD A RATE OF 8.64 ML/HOUR. THE INITIAL INFUSION FOR AN UNKNOWN DRUG (SUSPECTED TO BE EPINEPHRINE) WAS PROGRAMMED REMOTELY (INTEROPERABILITY WORKFLOW) AT A CONCENTRATION OF 4 MG / 250 ML AND WAS STARTED ON (B)(6) 2025 AT 1:00 PM (NOT WEIGHT-BASED). THE DOSE WAS 1 MCG/MIN, CALCULATING THE RATE AT 3.75 ML/HR AND THE VOLUME TO BE INFUSED (VTBI) WAS 250 ML. THERE WAS NO ALLEGATION OF A DEVICE MALFUNCTION, AND NO MALFUNCTION WAS IDENTIFIED DURING THE LOG REVIEW.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22147911&pc=NVZ | 3004209178-2025-09594 | 2025-04-30 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-05 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE FOR A NEW MEDTRONIC IMPLANTABLE PULSE GENERATOR (IPG) REPLACEMENT, THE PATIENT PASSED AWAY, ALTHOUGH NO DEVICE ISSUE WAS IDENTIFIED, THE PATIENT FAMILY MEMBER EMPHASIZED THE PASSING HAPPENED DUE TO THE PROCEDURE. CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID MDT-IPG IMPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22147932&pc=NVZ | 2182208-2025-02739 | 2025-04-30 04:00 | Death | MEDTRONIC, INC. | 2025-06-05 04:00 | NVZ | MEDTRONIC IPG | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE FOR A NEW MEDTRONIC IMPLANTABLE PULSE GENERATOR (IPG) REPLACEMENT, THE PATIENT PASSED AWAY, ALTHOUGH NO DEVICE ISSUE WAS IDENTIFIED, THE PATIENT FAMILY MEMBER EMPHASIZED THE PASSING HAPPENED DUE TO THE PROCEDURE. CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22147945&pc=NVN | 2649622-2025-15158 | 2025-04-30 04:00 | Death | MPRI | 2025-06-05 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE FOR A NEW MEDTRONIC IMPLANTABLE PULSE GENERATOR (IPG) REPLACEMENT, THE PATIENT PASSED AWAY, ALTHOUGH NO DEVICE ISSUE WAS IDENTIFIED, THE PATIENT FAMILY MEMBER EMPHASIZED THE PASSING HAPPENED DUE TO THE PROCEDURE. CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: CONTINUATION OF D10: MDT-IPG IMPLANTED: (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22147980&pc=NVN | 2649622-2025-15160 | 2025-04-30 04:00 | Death | MPRI | 2025-06-05 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE FOR A NEW MEDTRONIC IMPLANTABLE PULSE GENERATOR (IPG) REPLACEMENT, THE PATIENT PASSED AWAY, ALTHOUGH NO DEVICE ISSUE WAS IDENTIFIED, THE PATIENT FAMILY MEMBER EMPHASIZED THE PASSING HAPPENED DUE TO THE PROCEDURE. CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: CONTINUATION OF D10: MDT-IPG IMPLANTED: (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22148392&pc=OZD | 1220648-2025-29010 | 2025-05-31 04:00 | Death | ABIOMED, INC. | 2025-06-05 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Shock | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA CP PUMP WAS PLACED TO SUPPORT A PATIENT WITH ACUTE MYOCARDIAL INFARCTION. AFTER CARDIOPULMONARY RESUSCITATION, AN IMPELLA CP WAS INSERTED TO STABILIZE THE HEMODYNAMICS. BLEEDING AT THE SITE OF CHEST COMPRESSIONS WAS DIFFICULT TO CONTROL. A BLOOD TRANSFUSION WAS PERFORMED. THE BLEEDING LEAD TO HEMORRHAGIC SHOCK AND THE PATIENT DIED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22148636&pc=DSQ | 2916596-2025-03492 | 2025-05-16 04:00 | Death | THORATEC CORPORATION | 2025-06-05 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE WAS NOT PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. IT WAS UNKNOWN IF THE DEVICE WOULD BE RETURNED FOR EVALUATION. Manufacturer Narrative: SECTION A4: PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER: (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22148652&pc=DSQ | 2916596-2025-03496 | 2025-05-19 04:00 | Death | THORATEC CORPORATION | 2025-06-05 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Mechanical Problem; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT WAS WALKING BY THE WATER WITH THEIR FEET IN THE WATER, THE PATIENT STUMBLED AND FELL INTO THE WATER. THE SYSTEM CONTROLLER AND BATTERY CLIPS WERE WET AS A RESULT AND THE SYSTEM CONTROLLER GAVE AN ALARM AFTER THIS EVENT. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED, AND THE DEVICE WOULD NOT RETURN.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE SYSTEM CONTROLLER ALARMED FOR POWER CABLE DISCONNECT, LOW VOLTAGE, AND NO EXTERNAL POWER AFTER THE PATIENT FELL INTO THE WATER. THE DEATH WAS CONSIDERED TO BE DUE TO AN ACCIDENT RELATED TO THE FALL. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND THE REPORTED PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: SECTION B1: TYPE OF REPORT CORRECTED. SECTION H6: MEDICAL DEVICE PROBLEM CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22149010&pc=OZD | 1220648-2025-29017 | 2025-05-12 04:00 | Death | ABIOMED, INC. | 2025-06-05 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Embolism/Embolus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT SUFFERED A CEREBRAL EMBOLISM DURING IMPELLA 5.5 SUPPORT. CARE WAS WITHDRAWN DUE TO THE NOTED CONDITION. THE PATIENT LATER PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22149190&pc=OZD | 1220648-2025-29019 | 2024-06-24 04:00 | Death | ABIOMED, INC. | 2025-06-05 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock | P140003 |  | 1 | Event Description: IT WAS REPORTED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) STUDY THAT A SUBJECT EXPERIENCED CARDIOGENIC SHOCK. LOQI REPORTS: PATIENT ADMITTED ON VENOARTERIAL EXTRACORPOREAL MEMBRANE DUE TO CARDIOGENIC SHOCK COMPLICATED BY A ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION AND UNDERWENT RIGHT GASTRIC ARTERY REVASCULARIZATION WHICH WAS COMPLICATED BY COMPLETE HEART BLOCK. THE PATIENT HAD LMPELLA PLACEMENT FOR CARDIAC SUPPORT. HOSPITAL STAY WAS COMPLICATED BY SUBARACHNOID HEMORRHAGE, SPINAL ISCHEMIA, PNEUMONIA, AND CHRONIC RESPIRATORY FAILURE. ULTIMATELY, THE PATIENT DETERIORATED, AND THE FAMILY OPTED TO PROCEED WITH HOSPICE CARE. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22149350&pc=QFG | 3004464228-2025-24478 | 2024-10-29 04:00 | Death | INSULET CORPORATION | 2025-06-05 04:00 | QFG | OMNIPOD 5 CONTROLLER | Insufficient Information | Diabetic Ketoacidosis | K203768 |  | 1 | Event Description: ON 22MAY25 INSULET RECEIVED AN EMAIL FROM A HEALTHCARE PROFESSIONAL (HCP) NOTIFYING INSULET OF THEIR PATIENT'S DEATH DUE TO DIABETIC KETOACIDOSIS (DKA) AND INQUIRING ABOUT THE PERSONAL DIABETES MANAGER (PDM). INSULET FOLLOWED UP WITH THE HCP AND OBTAINED THE ADDITIONAL BACKGROUND. THE PATIENT HAD COMORBIDITIES THAT INCLUDED RENAL DISEASE AND A PREVIOUS CEREBROVASCULAR ACCIDENT. REPORTEDLY, THE PATIENT WAS AWAITING A RENAL TRANSPLANT. THE PATIENT HAD STRUGGLED WITH MANAGEMENT OF DIABETES PRIOR TO USE OF OMNIPOD 5. APPROXIMATELY A WEEK BEFORE THE PATIENT EXPIRED, SHE WAS EXPERIENCING INCREASED BLOOD GLUCOSE LEVELS. THE PATIENT PASSED AWAY ON (B)(6) 2024. THE PATIENT EXPIRED AT HOME AND WAS WEARING THE DEVICE AT TIME OF DEATH. THE CLINIC HAS POSSESSION OF THE PDM AND WILL MAINTAIN THE DEVICE. INSTRUCTIONS FOR UPLOADING LOGS WERE PROVIDED. THE REQUESTED GLOOKO INFORMATION IS NOT AVAILABLE AS THE PATIENT DID NOT CONSENT TO GLOOKO/OP5 DATA SHARING. THE OP5 WAS BEING USED WITH A DEXCOM G6. DEXCOM HAS BEEN NOTIFIED. Manufacturer Narrative: THE PHYSICAL DEVICE WAS NOT RETURNED. ANALYSIS OF THE AVAILABLE CLOUD DATA FOUND NO EVIDENCE OF ISSUES WITH THE SYSTEM. IT WAS NOTED THAT THE USER DELIVERED THE LARGEST AMOUNTS OF TOTAL INSULIN AND TOTAL BOLUS INSULIN ON THE DAY BEFORE THE DATE OF OCCURRENCE, HOWEVER THERE IS NO EVIDENCE OF A SYSTEM ISSUE CONTRIBUTING TO THIS INSULIN DELIVERY. THE DATA SHOWED THAT THE AUTOMATED ALGORITHM APPROPRIATELY ADAPTED THE AUTOMATED INSULIN DELIVERY BASED ON THE ESTIMATED GLUCOSE VALUES AND BOLUS DELIVERIES. THE DATA SHOWED THAT THE SYSTEM WAS USED PRIMARILY IN AUTOMATED MODE IN THE DAYS LEADING UP TO THE DATE OF OCCURRENCE ((B)(6) 2024). THE DATA SHOWED THAT THE AUTOMATED ALGORITHM WAS ABLE TO APPROPRIATELY ADAPT THE MICROBOLUS AMOUNTS BASED ON THE ESTIMATED GLUCOSE VALUES RECEIVED BY THE POD FROM THE SENSOR. THE DATA SHOWED THAT GLUCOSE VALUES REACHED URGENT HIGH VALUES AT 03:25:39 ON (B)(6) 2024. AT THAT TIME, THE SYSTEM STARTED DELIVERING AT THE MAX AUTOMATED RATE AND CONTINUED TO DELIVER AT THE MAX RATE DUE TO URGENT HIGH BLOOD GLUCOSE LEVELS UNTIL AN AUTOMATED DELIVERY RESTRICTION ALERT WAS GENERATED AT 05:30:40 ON (B)(6) 2024. THE ALERT WAS ACKNOWLEDGED BY THE PATIENT AND THE SYSTEM WAS PUT INTO MANUAL MODE. THE SYSTEM STARTED DELIVERING THE USER'S MANUAL BASAL PROGRAM AS EXPECTED. AN 8.35 U BOLUS WAS DELIVERED SHORTLY AFTER THE SYSTEM MODE CHANGE, RESULTING IN BLOOD GLUCOSE LEVELS DECREASING, REACHING URGENT LOW VALUES AT 07:40:40 ON (B)(6) 2024. BLOOD GLUCOSE LEVELS STARTED INCREASING AGAIN, REACHING URGENT HIGH AT 12:20:41 ON (B)(6) 2024. SENSOR VALUES RECEIVED BY THE POD STAYED URGENT HIGH UNTIL THE SENSOR BECAME INACTIVE AT 22:15:43 ON (B)(6) 2024. NO SENSOR VALUES WERE RECEIVED BY THE POD THROUGH THE LAST DATA POINT AT 11:30:47 ON (B)(6) 2024. THE SYSTEM CONTINUED TO DELIVER THE USER'S MANUAL BASAL PROGRAM DURING THIS PERIOD. MULTIPLE BOLUSES WERE DELIVERED ON (B)(6) 2024, HOWEVER NO BOLUSES WERE DELIVERED ON (B)(6) 2024. NO EVIDENCE OF ISSUES WITH THE SYSTEM WERE OBSERVED IN THE AVAILABLE CLOUD DATA. LOT RELEASE RECORDS WERE REVIEWED AND THE PRODUCT LOT MET ALL ACCEPTANCE CRITERIA. SPECIFICALLY, A POD IS PAIRED TO A PDM AND PUT THROUGH SIMULATED USE TESTING INCLUDING COMMUNICATING WITH THE PDM, DEPLOYMENT, DELIVERING FLUID, OCCLUSION DETECTION, AND FREEDOM FROM HAZARD ALARMS. IF ADDITIONAL INFORMATION IS PROVIDED A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22149394&pc=LWP | 2017865-2025-75527 | 2023-11-29 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-05 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO PATIENT DEATH. ANALYSIS PERFORMED, INCLUDING ELECTRICAL, MECHANICAL, AND ENVIRONMENTAL TESTS INDICATED THAT THE DEVICE EXHIBITED NORMAL DEVICE CHARACTERISTICS WITH NO ANOMALIES. LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22149407&pc=NVN | 2017865-2025-75528 | 2023-11-29 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-05 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22149427&pc=OZD | 1220648-2025-29021 | 2025-05-12 04:00 | Death | ABIOMED, INC. | 2025-06-05 04:00 | OZD | IMPELLA 5.5 | Failure to Advance; Patient-Device Incompatibility | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT THE DECISION FOR IMPELLA 5.5 WAS MADE ON A HIGH-RISK SURGERY PATIENT WITH POSTCARDIOTOMY CARDIOGENIC SHOCK/LOW CARDIAC OUTPUT SYNDROME. IT WAS REPORTED THAT THE LEFT AXILLARY CUTDOWN, FOR LEFT VENTRICLE (LV) ACCESS, INTRA-AORTIC BALLOON PUMP WAS IN PLACE, THE CHEST WAS OPEN, BYPASS FLOWING 2.2 LPM, AL-1 / 035 TO ACCESS LV, HOWEVER THE WIRE WAS SWAPPED FOR V-18. THE IMPELLA 5.5 WOULDN¿T ADVANCE INTO ASCENDING AORTA. AFTER MULTIPLE ATTEMPTS THE 5.5 WAS SWAPPED FOR THE IMPELLA CP HOWEVER THE PATIENT EXPIRED IN THE OPERATING ROOM. ABIOMED'S MEDICAL SAFETY OFFICER REVIEW THE PATIENT'S OUTCOME AND IT WAS DETERMINED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA 5.5 DEVICE AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF DELIVERY ISSUE WAS MOST LIKELY USE ISSUE SINCE NON-ABIOMED PRODUCT WAS USED WHICH IS NOT RECOMMENDED PER INSTRUCTIONS FOR USE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. ABIOMED'S MEDICAL SAFETY OFFICER REVIEW THE PATIENTS OUTCOME AND IT WAS DETERMINED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA 5.5 DEVICE AS AN ASSOCIATED FACTOR TO THE PATIENTS OUTCOME. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22149551&pc=PYX | 1220648-2025-29022 | 2025-05-16 04:00 | Death | ABIOMED, INC. | 2025-06-05 04:00 | PYX | IMPELLA RP FLEX | Material Split, Cut or Torn | Tachycardia; Ventricular Fibrillation | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT IN POSTCARDIOTOMY CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA RP FLEX FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND DURING INSERTION AN ISSUE WITH THE DILATOR WAS ENCOUNTERED. SUBSEQUENTLY, THE PATIENT EXPERIENCED CARDIAC ARRHYTHMIC EVENTS AND EXPIRED. THE PATIENT UNDERWENT SURGICAL UNROOFING OF THE RIGHT CORONARY ARTERY. POST OPEN HEART SURGERY WHILE ON THE UNIT, THE PATIENT WENT INTO ACUTE RIGHT VENTRICULAR FAILURE. THE PATIENT WAS RETURNED TO THE OPERATING ROOM, AND THE DECISION WAS MADE TO PLACE AN IMPELLA RP FLEX. UPON INSERTION INTO THE RIGHT INTERNAL JUGULAR, THE DILATOR OF A 23 FRENCH PEEL AWAY SHEATH SPLIT AT THE TIP. A NEW IMPELLA RP FLEX SHEATH KIT WAS OPENED, AND A NEW DILATOR AND SHEATH WERE PLACED. FOLLOWING SUCCESSFUL PUMP IMPLANT, THE PATIENT ENDURED RUNS OF VENTRICULAR TACHYCARDIA FOR WHICH MULTIPLE SHOCKS WERE REQUIRED. SUPPORT CONTINUED WITH THE CURRENT SETUP FOLLOWING THE INTERVENTION. THE PATIENT REMAINED CRITICALLY ILL WITH VERY POOR PROGNOSIS. THE PATIENT'S CHEST WAS OPEN, MAXED ON MULTIPLE INOTROPES/PRESSORS ON CONTINUOUS RENAL REPLACEMENT THERAPY. APPROXIMATELY THREE DAYS LATER, THE PATIENT DEVELOPED VENTRICULAR FIBRILLATION ARREST IN THE EVENING, WAS SHOCKED, AND CONVERTED TO ASYSTOLE. PACED AT A BACKUP RATE OF 40. ADVANCED CARDIAC LIFE SUPPORT PROTOCOL WAS FOLLOWED MINUS CHEST COMPRESSIONS DUE TO OPEN CHEST STATUS. THE FAMILY WAS NOTIFIED OF CRITICAL CONDITION. ONE HOUR LATER, THE PATIENT VENTRICULAR FIBRILLATION ARRESTED AGAIN AND WAS SHOCKED ONE TIME AT 200 JOULES. BLOOD PRESSURE WAS 67/41. THE FAMILY WAS AT BEDSIDE AND REQUESTED TO STOP ALL MEASURES, AND THE PATIENT SUBSEQUENTLY PASSED AWAY. Manufacturer Narrative: THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR TO THE OVERALL OUTCOME (GIVEN THE ARRHYTHMIC EVENTS REQUIRING ADVANCE CARDIAC LIFE SUPPORT). HOWEVER, IT SHOULD BE NOTED THE PATIENT¿S PROGNOSIS WAS VERY POOR. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER INSTRUCTIONS FOR USE & CLINICAL REFERENCE MANUAL SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ARRHYTHMIA, ATRIAL FIBRILLATION, BLEEDING, CARDIAC TAMPONADE, CARDIOGENIC SHOCK, DEATH, DEVICE MALFUNCTION, HEMOLYSIS, HEPATIC FAILURE, INSERTION SITE INFECTION, PERFORATION, PHLEGMASIA CERULEA DOLENS (A SEVERE FORM OF DEEP VENOUS THROMBOSIS), PULMONARY VALVE INSUFFICIENCY, RESPIRATORY DYSFUNCTION, SEPSIS, THROMBOCYTOPENIA, THROMBOTIC VASCULAR (NON-CENTRAL NERVOUS SYSTEM COMPLICATION, TRICUSPID VALVE INJURY, CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION), VENOUS THROMBOSIS, VENTRICULAR FIBRILLATION AND/OR TACHYCARDIA.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22149680&pc=QBJ | 3004753838-2025-140544 | 2024-05-21 04:00 | Death | Unknown Manufacturer | 2025-06-05 04:00 | QBJ | DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: CLAIMANT¿S ATTORNEY REPORTED TO DEXCOM THAT THE PUMP HARDWARE ALLEGEDLY MALFUNCTIONED AND IT WAS REPORTING THAT THERE WAS ALLEGEDLY A POSSIBLE MALFUNCTION WITH THE DEXCOM SOFTWARE. THE PATIENT WAS REPORTEDLY FOUND PASSED AWAY IN HIS HOME ON APPROXIMATELY (B)(6) 2024. NO PRODUCT OR DATA WAS PROVIDED FOR INVESTIGATION. PROBLEM COULD NOT BE CONFIRMED AND PROBABLE CAUSE CANNOT BE DETERMINED. DEXCOM¿S LEGAL COUNSEL REQUESTED ADDITIONAL REQUESTS ADDITIONAL INFORMATION FROM CLAIMANT¿S ATTORNEY AND NO FURTHER INFORMATION WAS PROVIDED.Event Description: AFTER THE INITIAL MDR, ADDITIONAL INFORMATION WAS ACQUIRED. A REVIEW OF THE DEXCOM PERFORMANCE DATA SHOWED THAT A SENSOR SESSION WAS STARTED ON (B)(6) 2024 AT 2015. THE SENSOR SESSION LASTED APPROXIMATELY 8.5 DAYS. DATA FOR THE FINAL 48 HOURS OF THE SENSOR SESSION INDICATES THAT THE USER WAS WITHIN TARGET RANGE UP UNTIL 2250 ON (B)(6) 2024, WHEN THE ESTIMATED GLUCOSE VALUES (EGVS) WENT ABOVE THE UPPER LIMIT OF USER SET HIGH GLUCOSE THRESHOLD OF 150 MG/DL AND A HIGH GLUCOSE ALERT OCCURRED. THE ALERT SETTING WAS VIBRATE AND THE ALERT REPEATED EVERY 5 MINUTES UNTIL 0220 AM ON (B)(6) 2024 WHEN THE ALERT WAS ACKNOWLEDGED. ON (B)(6) 2024 BETWEEN 0220 AND 1115, THE EGVS CONTINUED TO RISE UNTIL THEY REACHED 400 MG/DL. THE EGVS THEN ROSE ABOVE 400 MG/DL AND REMAINED ABOVE 400 MG/DL UNTIL 0605 ON (B)(6) 2024, WHEN A NO READINGS ALERT OCCURRED. THIS ALERT SETTING WAS ALSO SET TO VIBRATE. AT 0640, THE EGVS RETURNED AND WERE STILL ABOVE 400 MG/DL AND THE HIGH GLUCOSE ALERT OCCURRED AGAIN. NEITHER OF THESE ALERTS WERE ACKNOWLEDGED. ON (B)(6) 2024, AT 1625, THE NO READINGS ALERT OCCURRED AGAIN. THERE WAS NO ACKNOWLEDGEMENT OF THIS ALERT. THE SENSOR SESSION ENDED ON (B)(6) 2024 AT 1640 WHEN THE FINAL EGV OF HIGH WAS RECORDED. THERE WERE NO ISSUES OBSERVED. A PROBABLE CAUSE COULD NOT BE DETERMINED. AT THE TIME OF THE EVENT, THE USER WAS CONNECTED TO AN OMNIPOD 5 INSULIN PUMP, MANUFACTURED BY INSULET. INSULET WAS CONTACTED TO OBTAIN THEIR DATA FOR THE PERIOD OF INTEREST. INSULET¿S DATA WAS REVIEWED, AND IT WAS REPORTED THAT: ¿CLOUD DATA FOR THIS USER FOR THE PERIOD OF (B)(6) 2024 WAS REVIEWED. THE DATA SHOWED THAT THE AUTOMATED ALGORITHM WAS APPROPRIATELY ADAPTING THE AUTOMATED INSULIN DELIVERY BASED ON THE ESTIMATED GLUCOSE VALUES RECEIVED BY THE POD. THE DATA SHOWED THAT THE ESTIMATED GLUCOSE VALUES STARTED INCREASING AND REACHED URGENT HIGH AT 1120 ON (B)(6) 2024. THE SYSTEM RESPONDED APPROPRIATELY TO THE INCREASING AND HIGH ESTIMATED GLUCOSE LEVELS BY DELIVERING THE MAXIMUM AMOUNT OF AUTOMATED INSULIN DETERMINED BY THE USER'S ADAPTIVE BASAL RATE. EXTENDED PERIODS OF THIS MAX INSULIN DELIVERY RESULTED IN AUTOMATED DELIVERY RESTRICTION ALERTS BEING GENERATED AT 1205 AND 1615 ON (B)(6) 2024. AFTER THE AUTOMATED DELIVERY RESTRICTION AT 1615 WAS ACKNOWLEDGED AND THE SYSTEM PUT INTO MANUAL MODE, THE USER KEPT THE SYSTEM IN MANUAL MODE AND THE POD DELIVERED THE USER'S MANUAL BASAL PROGRAM OF 0.8 U/HOUR. MANUAL DELIVERY CONTINUED UNTIL THE LAST DATA POINT AT 1710:40 ON (B)(6) 2024. Manufacturer Narrative: (B)(4) DEXCOM'S LEGAL COUNSEL WAS SERVED WITH A LEGAL ACTION AS A RESULT OF THE PATIENT¿S ALLEGED INJURY. THE REPORTER DID NOT PROVIDE A DESCRIPTION OR INFORMATION REGARDING THE CIRCUMSTANCES OF THE ALLEGED EVENT OR DETAILS OF THE ALLEGED INJURY OTHER THAN WHAT IS ALLEGED IN THE COMPLAINT. DEXCOM HAS PROVIDED TO FDA THE INFORMATION IT CURRENTLY HAS RELATING TO THIS ALLEGED ADVERSE EVENT. DEXCOM'S LEGAL COUNSEL HAS REQUESTED FURTHER INFORMATION FROM THE CLAIMANT¿S ATTORNEY. Manufacturer Narrative: (B)(4). B5 DESCRIBE EVENT OR PROBLEM - ADDITIONAL. H2 ADDITIONAL. H6 TYPE OF INVESTIGATION - ADDITIONAL. H6 INVESTIGATION FINDINGS - ADDITIONAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22150084&pc=JDI | 1818910-2025-09094 | 2013-10-08 04:00 | Death | DEPUY ORTHOPAEDICS INC US | 2025-06-05 04:00 | JDI | UNK HIP ACETABULAR CUP PINNACLE | Adverse Event Without Identified Device or Use Problem | Hypersensitivity/Allergic reaction; Unspecified Infection; Inflammation; Pain; Discomfort |  |  | 1 | Event Description: PATIENT UNDERWENT REVISION SURGERY DUE TO SEVERE PAIN AND DISCOMFORT, INFLAMMATION ON HER LEFT THIGH, BUTTOCKS HIP AND GROIN. THE PINNACLE DEVICE HAD CAUSED A SIGNIFICANT AMOUNT OF METAL DEBRIS OR METALLOIDS TO ACCUMULATE WITHIN THE HIP CAPSULE AND IN THE SOFT TISSUE SURROUNDING THE IMPLANT, PSEUDOTUMOR AND A GREAT DEAL OF INFLAMMATION. MOREOVER, CORROSION, LARGE AMOUNT OF TOXIC COBALT, CHROMIUM AND PARTICLE AND INFECTION WAS ALSO EXPERIENCED. ALL OF THE ABOVEMENTIONED MANIFESTATIONS FROM THE PATIENT WERE THE CONSEQUENCES OF THE IMPLANTATION OF THE PINNACLE DEVICE WHICH RESULTED TO DEATH ON (B)(6) 2023. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803 (AND/OR PART 4, AS APPLICABLE). THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY DEPUY SYNTHES, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. H11 ADDITIONAL MANUFACTURER NARRATIVE: D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. H3, H6: THE DEVICE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE HISTORY REVIEW COULD NOT BE PERFORMED. IF THE LOT/SERIAL NUMBER BECOMES AVAILABLE, THE RECORD WILL BE RE-ASSESSED. NO DEVICE ASSOCIATED WITH THIS REPORT WAS RECEIVED FOR EXAMINATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT/LOT NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS WILL BE CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF THE POST-MARKET SURVEILLANCE. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. Manufacturer Narrative: PRODUCT COMPLAINT # : (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY DEPUY SYNTHES JOINT RECONSTRUCTION, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, DEPUY SYNTHES JOINT RECONSTRUCTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. ADDED: D10 CONCOMITANT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION THAT HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY DEPUY SYNTHES JOINT RECONSTRUCTION, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, DEPUY SYNTHES JOINT RECONSTRUCTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. ADDED: D4 CATALOG. IF INFORMATION THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT IS OBTAINED, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22150106&pc=KWA | 1818910-2025-09095 | 2013-10-08 04:00 | Death | DEPUY ORTHOPAEDICS INC US | 2025-06-05 04:00 | KWA | UNK HIP ACETABULAR LINER METAL PINNACLE | Naturally Worn | Hypersensitivity/Allergic reaction; Unspecified Infection; Inflammation; Pain; Discomfort; Metal Related Pathology |  |  | 1 | Event Description: PATIENT UNDERWENT REVISION SURGERY DUE TO SEVERE PAIN AND DISCOMFORT, INFLAMMATION ON HER LEFT THIGH, BUTTOCKS HIP AND GROIN. THE PINNACLE DEVICE HAD CAUSED A SIGNIFICANT AMOUNT OF METAL DEBRIS OR METALLOIDS TO ACCUMULATE WITHIN THE HIP CAPSULE AND IN THE SOFT TISSUE SURROUNDING THE IMPLANT, PSEUDOTUMOR AND A GREAT DEAL OF INFLAMMATION. MOREOVER, CORROSION, LARGE AMOUNT OF TOXIC COBALT, CHROMIUM AND PARTICLE AND INFECTION WAS ALSO EXPERIENCED. ALL OF THE ABOVE MENTIONED MANIFESTATIONS FROM THE PATIENT WERE THE CONSEQUENCES OF THE IMPLANTATION OF THE PINNACLE DEVICE WHICH RESULTED TO DEATH ON (B)(6) 2023. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803 (AND/OR PART 4, AS APPLICABLE). THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY DEPUY SYNTHES, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. H11 ADDITIONAL MANUFACTURER NARRATIVE: D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. H3, H6: THE DEVICE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE HISTORY REVIEW COULD NOT BE PERFORMED. IF THE LOT/SERIAL NUMBER BECOMES AVAILABLE, THE RECORD WILL BE RE-ASSESSED. NO DEVICE ASSOCIATED WITH THIS REPORT WAS RECEIVED FOR EXAMINATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT/LOT NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS WILL BE CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF THE POST-MARKET SURVEILLANCE. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. Manufacturer Narrative: PRODUCT COMPLAINT # : (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION THAT HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY DEPUY SYNTHES JOINT RECONSTRUCTION, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, DEPUY SYNTHES JOINT RECONSTRUCTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. ADDED: D4 CATALOG. IF INFORMATION THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT IS OBTAINED, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY DEPUY SYNTHES JOINT RECONSTRUCTION, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, DEPUY SYNTHES JOINT RECONSTRUCTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. ADDED: D10 CONCOMITANT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22150113&pc=MKJ | 1220908-2025-02158 | 2025-05-19 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-06-05 04:00 | MKJ | R SERIES | Failure to Deliver Shock/Stimulation | No Clinical Signs, Symptoms or Conditions | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A PATIENT (AGE: 2 WEEKS & GENDER UNKNOWN) THE DEVICE WAS UNABLE TO DISCHARGE. THE COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED.Event Description: IT WAS REPORTED THAT THE DEVICE WAS UNABLE TO DISCHARGE WHILE ATTEMPTING TO DEFIBRILLATE A PEDIATRIC PATIENT (AGE & GENDER UNKNOWN). THE PATIENT EXPIRED; HOWEVER, IT WAS CONFIRMED BY THE REPORTER THAT THE DEVICE MALFUNCTION WAS NOT A CONTRIBUTING FACTOR IN THE PATIENT DEATH.Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A 2-WEEKS-OLD PATIENT (GENDER UNKNOWN), THE DEVICE WAS UNABLE TO DISCHARGE. THE PATIENT EXPIRED; HOWEVER, IT WAS CONFIRMED BY THE REPORTER THAT THE DEVICE MALFUNCTION WAS NOT A CONTRIBUTING FACTOR IN THE PATIENT DEATH. Manufacturer Narrative: THIS SUPPLEMENTAL MEDWATCH REPORT IS REPORTING THE EVALUATION OF THE DEVICE AND CORRECTING INFORMATION SUBMITTED ON THE INITIAL MEDWATCH REPORT. PLEASE REFERENCE SECTION B5 UPDATED. ZOLL MEDICAL CANADA EVALUATED THE DEVICE AND THE CUSTOMER'S REPORT WAS NOT REPLICATED OR CONFIRMED. REVIEW OF LOG CONFIRMED THAT TWO 6J SHOCKS WERE DELIVERED TO THE PATIENT 13 SECONDS APART. THERE WAS NO EVIDENCE THAT THE DEVICE FAILED TO DELIVER A SHOCK OR DISPLAYED A "SHOCK NOT DELIVERED" MESSAGE. THE MESSAGES "DEFIB NOT CHARGED" AND "PRESS CHARGE" APPEARED AFTER EACH SHOCK, WHICH IS EXPECTED WHEN THE SHOCK BUTTON IS PRESSED AGAIN BEFORE RECHARGING. THE CLINICAL LOG WAS NOT AVAILABLE FOR REVIEW. THE RETURNED DEVICE AND CABLES WERE TESTED USING KNOWN GOOD PEDIATRIC PADS AND SHOWED NO ISSUES. ALL TESTS PASSED WITHOUT REPRODUCING THE REPORTED MALFUNCTION. THE DEVICE FUNCTIONED CORRECTLY AND DELIVERED ENERGY AS INTENDED. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE. Manufacturer Narrative: B5: CORRECTION MADE TO PROVIDE MORE DETAILS ON THE REPORTED INCIDENT. D10: CONCOMITANT MEDICAL PRODUCTS AND THERAPY DATES (EXCLUDE TREATMENT OF EVENT): H6: HEALTH EFFECT IMPACT CODE UPDATED. Manufacturer Narrative: JUSTIFICATION FOR NO UDI, THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS RECEIVED THE PRODUCT AND WILL BE PROVIDING A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22150130&pc=LPH | 1818910-2025-09096 | 2013-10-08 04:00 | Death | DEPUY ORTHOPAEDICS INC US | 2025-06-05 04:00 | LPH | UNKNOWN HIP FEMORAL HEAD | Naturally Worn | Hypersensitivity/Allergic reaction; Unspecified Infection; Inflammation; Pain; Discomfort; Metal Related Pathology |  |  | 1 | Event Description: PATIENT UNDERWENT REVISION SURGERY DUE TO SEVERE PAIN AND DISCOMFORT, INFLAMMATION ON HER LEFT THIGH, BUTTOCKS HIP AND GROIN. THE PINNACLE DEVICE HAD CAUSED A SIGNIFICANT AMOUNT OF METAL DEBRIS OR METALLOIDS TO ACCUMULATE WITHIN THE HIP CAPSULE AND IN THE SOFT TISSUE SURROUNDING THE IMPLANT, PSEUDOTUMOR AND A GREAT DEAL OF INFLAMMATION. MOREOVER, CORROSION, LARGE AMOUNT OF TOXIC COBALT, CHROMIUM AND PARTICLE AND INFECTION WAS ALSO EXPERIENCED. ALL OF THE ABOVE MENTIONED MANIFESTATIONS FROM THE PATIENT WERE THE CONSEQUENCES OF THE IMPLANTATION OF THE PINNACLE DEVICE WHICH RESULTED TO DEATH ON (B)(6) 2023. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803 (AND/OR PART 4, AS APPLICABLE). THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY DEPUY SYNTHES, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. H11 ADDITIONAL MANUFACTURER NARRATIVE: D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. H3, H6: THE DEVICE LOT NUMBER IS UNKNOWN, THEREFORE A DEVICE HISTORY REVIEW COULD NOT BE PERFORMED. IF THE LOT/SERIAL NUMBER BECOMES AVAILABLE, THE RECORD WILL BE RE-ASSESSED. NO DEVICE ASSOCIATED WITH THIS REPORT WAS RECEIVED FOR EXAMINATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT/LOT NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS WILL BE CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF THE POST-MARKET SURVEILLANCE. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY DEPUY SYNTHES JOINT RECONSTRUCTION, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, DEPUY SYNTHES JOINT RECONSTRUCTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. ADDED: D10 CONCOMITANT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22150183&pc=LPH | 1818910-2025-09097 | 2013-10-08 04:00 | Death | DEPUY ORTHOPAEDICS INC US | 2025-06-05 04:00 | LPH | UNKNOWN HIP FEMORAL STEM | Degraded | Hypersensitivity/Allergic reaction; Unspecified Infection; Inflammation; Pain; Discomfort; Metal Related Pathology |  |  | 1 | Event Description: PATIENT UNDERWENT REVISION SURGERY DUE TO SEVERE PAIN AND DISCOMFORT, INFLAMMATION ON HER LEFT THIGH, BUTTOCKS HIP AND GROIN. THE PINNACLE DEVICE HAD CAUSED A SIGNIFICANT AMOUNT OF METAL DEBRIS OR METALLOIDS TO ACCUMULATE WITHIN THE HIP CAPSULE AND IN THE SOFT TISSUE SURROUNDING THE IMPLANT, PSEUDOTUMOR AND A GREAT DEAL OF INFLAMMATION. MOREOVER, CORROSION, LARGE AMOUNT OF TOXIC COBALT, CHROMIUM AND PARTICLE AND INFECTION WAS ALSO EXPERIENCED. ALL OF THE ABOVE MENTIONED MANIFESTATIONS FROM THE PATIENT WERE THE CONSEQUENCES OF THE IMPLANTATION OF THE PINNACLE DEVICE WHICH RESULTED TO DEATH ON (B)(6) 2023. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY DEPUY SYNTHES JOINT RECONSTRUCTION, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, DEPUY SYNTHES JOINT RECONSTRUCTION, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. ADDED: D10 CONCOMITANT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. Manufacturer Narrative: PRODUCT COMPLAINT #: (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803 (AND/OR PART 4, AS APPLICABLE). THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY DEPUY SYNTHES, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. H11 ADDITIONAL MANUFACTURER NARRATIVE: D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. H3, H6: THE DEVICE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE HISTORY REVIEW COULD NOT BE PERFORMED. IF THE LOT/SERIAL NUMBER BECOMES AVAILABLE, THE RECORD WILL BE RE-ASSESSED. NO DEVICE ASSOCIATED WITH THIS REPORT WAS RECEIVED FOR EXAMINATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT/LOT NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS WILL BE CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF THE POST-MARKET SURVEILLANCE. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22150375&pc=NVY | 2017865-2025-75539 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-05 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: AS RECEIVED, A PARTIAL CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. THE CUT END OF THIS RETURNED PORTION WAS CONSISTENT WITH PROCEDURAL DAMAGE. DURING ANALYSIS, A FAILURE EVENT WAS OBSERVED, WHICH WAS UNRELATED TO THE REPORTED EVENT. FINAL ANALYSIS FOUND EXTERNAL INSULATION ABRASION AT THE PROXIMAL REGION OF THE LEAD BREACHING THE OPTIM SHEATH. THE CAUSE OF THE EXTERNAL INSULATION ABRASION WAS DUE TO FRICTION TO THE DEVICE CAN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22150414&pc=NVZ | 2017865-2025-75537 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-05 04:00 | NVZ | FORTIFY ASSURA DR ICD, US | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22150417&pc=NVN | 2017865-2025-75538 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-05 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22150581&pc=MGB | 2024168-2025-05778 | 2020-01-01 05:00 | Death | ABBOTT VASCULAR INC. | 2025-06-05 04:00 | MGB | PERCLOSE PROGLIDE¿ | Insufficient Information | Unspecified Heart Problem | P960043 |  | 1 | Event Description: THIS ARTICLE AIMS TO EVALUATE THE EFFICACY AND SAFETY OF BALLOON-EXPANDABLE COVERED STENT IMPLANTATION FOR COMMON FEMORAL ARTERY VASCULAR ACCESS-RELATED COMPLICATIONS ASSOCIATED WITH A LARGE-BORE PUNCTURE. PROGLIDE AND NON-ABBOTT DEVICES WERE USED FOR ACCESS SITE CLOSURES FOR ALL PROCEDURES. IT WAS NOTED THAT IF SIGNS OF ACUTE, PERSISTENT BLEEDING, PSEUDOANEURYSM OR COMPLICATED DISSECTION WERE OBSERVED, A BALLOON EXPANDABLE COVERED STENT IMPLANTATION WAS PERFORMED. A TRANSFUSION WAS ALSO GIVEN TO SOME PATIENTS. A TOTAL OF NINE PATIENTS DIED WITH ONE OF THEM RELATED TO VASCULAR ACCESS SITE COMPLICATIONS. A TOTAL OF 23 PATIENTS WERE ENROLLED IN THE STUDY, WITH A TOTAL OF 12 PATIENTS RECEIVING A PROGLIDE DEVICE. THE STUDY PROCEDURES WERE PERFORMED BETWEEN JANUARY 2020 TO MAY 2023. SPECIFIC PATIENT INFORMATION IS DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED "MIDTERM OUTCOME OF BALLOON-EXPANDABLE COVERED STENTING OF FEMORAL ACCESS SITE COMPLICATIONS." Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED. A REVIEW OF THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) REVIEW WAS NOT PERFORMED BECAUSE A SPECIFIC FAILURE MODE FOR THIS COMPLAINT WAS NOT PROVIDED. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY WAS NOT PERFORMED AS A SIMILARITY COULD NOT BE DETERMINED. THE PATIENT EFFECTS OF DEATH IS LISTED IN THE ELECTRONIC PROGLIDE INSTRUCTIONS FOR USE (IFU) AS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE USE OF THE DEVICE. BASED ON THE CASE INFORMATION, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT OF DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. A DEFINITIVE CAUSE FOR THE UNSPECIFIED DEVICE ISSUE COULD NOT BE DETERMINED. BASED ON THE INFORMATION REPORTED THROUGH THE RESEARCH ARTICLE, A CONCLUSIVE CAUSE FOR THE INSUFFICIENT INFORMATION COULD NOT BE DETERMINED. ADDITIONALLY, A DEFINITIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. B2 - DATE OF DEATH: ESTIMATED. B2 - DATE OF EVENT: ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE IS CAPTURED UNDER SEPARATE MEDWATCH REPORT. LITERATURE ATTACHMENT: ARTICLE TITLE "MIDTERM OUTCOME OF BALLOON-EXPANDABLE COVERED STENTING OF FEMORAL ACCESS SITE COMPLICATIONS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22150625&pc=OZD | 1220648-2025-29028 | 2025-05-12 04:00 | Death | ABIOMED, INC. | 2025-06-05 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD PLACED A IMPELLA CP TO SUPPORT A PATIENT IN SHOCK AND GETTING CARDIOPULMONARY RESUSCITATION. IT WAS REPORTED THAT DURING SUPPORT THE PATIENT HAS BILATERALLY COLD LEGS, WHICH THE REP STATES MAY HAVE BEEN CAUSED BY HIGH CATECHOLAMINES. LATER IT WAS REPORTED THAT THE PATIENT EXPIRED WHILE ON SUPPORT. THE ABIOMED'S MEDICAL SAFETY OFFICER REVIEW THE PATIENT'S OUTCOME AND THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA USED AS AN ASSOCIATED FACTOR TO THE PATIENT OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE ABIOMED'S MEDICAL SAFETY OFFICER REVIEW THE PATIENT'S OUTCOME AND THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA USED AS AN ASSOCIATED FACTOR TO THE PATIENT OUTCOME. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22150960&pc=DSQ | 2916596-2025-03395 | 2025-05-15 04:00 | Death | THORATEC CORPORATION | 2025-06-05 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22150964&pc=DSQ | 2916596-2025-03767 | 2024-01-08 05:00 | Death | THORATEC CORPORATION | 2025-06-05 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22150993&pc=DSQ | 2916596-2025-03762 | 2022-05-18 04:00 | Death | THORATEC CORPORATION | 2025-06-05 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE CAUSE OF DEATH WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22151136&pc=DSQ | 2916596-2025-03361 | 2025-05-11 04:00 | Death | THORATEC CORPORATION | 2025-06-05 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AFTER PALLIATION. THE CAUSE OF DEATH WAS RIGHT VENTRICULAR FAILURE AND CARDIO RENAL SYNDROME WITH A KNOWN RENAL DYSFUNCTION BEFORE IMPLANT. THE RIGHT HEART FAILURE ALSO EXISTED BEFORE IMPLANT AND DEVICE RELATED ISSUES DID NOT CAUSE OR CONTRIBUTE TO IT. THE PATIENT HAD SYMPTOMS OF HIGH CENTRAL VENOUS PRESSURE (CVP), HIGH PULMONARY ARTERY PRESSURE (PAP), AND CONGESTION. THEY WERE TREATED WITH INOTROPES. IT WAS NOTED THAT THE PATIENT NEVER LEFT THE INTENSIVE CARE UNIT (ICU). THE VENTRICULAR ASSIST DEVICE (VAD) WAS TURNED OFF. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED THROUGH THE MANUFACTURING EXECUTION SYSTEM (MES) AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS RIGHT HEART FAILURE, RENAL DYSFUNCTION, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE¿) ALSO OUTLINES INDICATIONS OF RIGHT HEART FAILURE AS WELL AS POSSIBLE TREATMENTS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22151180&pc=OZD | 1220648-2025-29030 | 2025-05-12 04:00 | Death | ABIOMED, INC. | 2025-06-05 04:00 | OZD | IMPELLA CP | Malposition of Device | Cardiac Perforation | P140003 |  | 1 | Event Description: COMPLAINANT IN INDIA REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THERE WAS A LEFT VENTRICLE PERFORATION. THE CLINICAL REP CONFIRMED THAT THE PERFORATION OCCURRED AT THE END OF THE PROCEDURE AND IT WAS CAUSED BY THE IMPELLA. THE PATIENT EXPIRED WITHIN 3 HOURS OF UNIT ADMISSION. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22151447&pc=DSQ | 2916596-2025-03600 | 2025-05-09 04:00 | Death | THORATEC CORPORATION | 2025-06-05 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE CAUSE WAS NOT PROVIDED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEVICE PERFORMED AS INTENDED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE CUSTOMER WILL NOT PROVIDE ANY FURTHER INFORMATION REGARDING THE CAUSE OF DEATH. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(4) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22151601&pc=DSQ | 2916596-2025-03470 | 2025-05-13 04:00 | Death | THORATEC CORPORATION | 2025-06-05 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Muscle Weakness; Tachycardia; Dizziness | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SUSTAINED VENTRICULAR TACHYCARDIA (SVT).Event Description: THE DEVICE WOULD NOT RETURN. IT WAS ADDITIONALLY REPORTED THAT THE PATIENT HAD IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) SHOCKS, WEAKNESS, DIZZINESS, AND TIREDNESS. THE PATIENT HAD A GANGLION BLOCK, ABLATIONS, RADIO FREQUENCY ABLATIONS, MEDICATIONS AND INABILITY TO PREVENT VENTRICULAR TACHYCARDIA. THE ARRHYTHMIA WAS NOT THOUGH TO BE DEVICE OR THERAPY RELATED AND AN ICD WAS IMPLANTED PRIOR TO THE DEVICE. THE SVT WOULD NOT BREAK AND THE PATIENT WISHED FOR "END OF LIFE" SO THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS DEACTIVATED ON (B)(6) 2025.Event Description: VT HAS BEEN PREVIOUSLY REPORTED FOR THIS PATIENT UNDER MFR #:2916596-2025-01450. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED CARDIAC ARRHYTHMIA AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA, AND DEATH THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22151652&pc=OZD | 1220648-2025-29032 | 2023-08-24 04:00 | Death | ABIOMED, INC. | 2025-06-05 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock | P140003 |  | 1 | Event Description: WE RECEIVED A NOTIFICATION VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED WORSENING CARDIOGENIC SHOCK WITH A "UNKNOWN (UNDETERMINED)" RELATIONSHIP TO THE IMPELLA DEVICE USED: ¿PATIENT WAS ADMITTED WITH SEVERE CARDIOMYOPATHY, EJECTION FRACTION OF <20%, CONGESTED HEPATOPATHY. PATIENT IN MULTI-SYSTEM ORGAN FAILURE. PATIENT HAD A BRONCHOSCOPY, ECMO PLACEMENT, HEMODIALYSIS, PATIENT MADE DNR BY FAMILY, TIME OF DEATH 1258 ON (B)(6) 2023.¿. THE PATIENT OUTCOME FATAL. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22151832&pc=OZD | 1220648-2025-29033 | 2025-03-06 05:00 | Death | ABIOMED, INC. | 2025-06-05 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Asystole | P140003 |  | 1 | Event Description: IT WAS REPORTED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY THAT SHORTLY FOLLOWING IMPELLA CP INSERTION, A PATIENT DEVELOPED A COMPLETE HEART BLOCK WITH INTERMITTENT VERY LOW AND WIDE ESCAPE, AND WAS EVENTUALLY ASYSTOLIC. THE PATIENT REQUIRED EMERGENT RIGHT COMMON FEMORAL VEIN TRANSVENOUS PACING (TVP). THIS WAS REMOVED AFTER ITERATIVE DECREMENTAL TESTING DEMONSTRATED PERSISTENCE OF THIS RHYTHM FOLLOWING PERCUTANEOUS CORONARY INTERVENTION (DESPITE IMPELLA REPOSITIONING AND REMOVAL). THE RIGHT INTERNAL JUGULAR WAS ACCESSED AND USED TO DELIVER A SCREW IN TVP IN THE 7 FRENCH MERIT HEATH. A FEW DAYS FOLLOWING EXPLANT, THE PATIENT AGAIN WENT INTO ASYSTOLE WITHOUT CAPTURE OF THEIR PACEMAKER AND WAS AGONALLY BREATHING. CODE BLUE WAS CALLED AND ANESTHESIA INTUBATED FOR AIRWAY PROTECTION. THE CODE LASTED FOR 45 MINUTES WITH MULTIPLE ROUNDS OF CARDIOPULMONARY RESUSCITATION, EPINEPHRINE, CALCIUM, BICARBONATE AND ATTEMPTS TO TRANSCUTANEOUSLY PACE WHICH WAS UNFORTUNATELY UNSUCCESSFUL. THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22151876&pc=OZD | 1220648-2025-29034 | 2025-05-08 04:00 | Death | ABIOMED, INC. | 2025-06-05 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Renal Failure | P140003 |  | 1 | Event Description: WE RECEIVED A NOTIFICATION VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED RENAL FAILURE WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA DEVICE USED: ¿THE PATIENT WAS ADMITTED WITH A HISTORY OF CHRONIC KIDNEY DISEASE TYPE III. THE PATIENT DEVELOPED ACUTE-ON-CHRONIC KIDNEY INSUFFICIENCY DURING INDEX ADMISSION DUE TO POOR RENAL PERFUSION. THE PATIENT WAS STARTED ON CONTINUOUS RENAL REPLACEMENT THERAPY ON (B)(6) 2025. ADMISSION CREATININE 1.90 MG/DL; PEAK CREATININE 4.67 MG/DL ((B)(6) 2025).¿ THE PATIENT OUTCOME NOT RECOVERED/NOT RESOLVED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22152095&pc=LZH | 1423537-2025-00226 | 2025-05-07 04:00 | Death | CARDINAL HEALTH 200, LLC | 2025-06-05 04:00 | LZH | KANGAROO | Insufficient Flow or Under Infusion | Dehydration | K221603 |  | 1 | Event Description: THE VP OF NURSING REACHED OUT TO A CARDINAL HEALTH SALES REP TO INFORM THAT A PATIENT WHO HAD BEEN ON A KANGAROO OMNI PUMP DID NOT APPEAR TO BE RECEIVING THE CORRECT FLUSHING VOLUMES. THE PATIENT WAS STARTED ON (B)(6) 2025 FOR GLUCERNA ENTERAL FEEDINGS. THE ORDERS FOR PUMP DELIVERY WERE TO FEED 40ML/HR AND FLUSH 200ML EVERY 6H. THE DAUGHTER SAID THE PUMP WAS NOT FLUSHING AND THE PATIENT BECAME DEHYDRATED. THE DAUGHTER SAID WHEN SHE WOULD CHANGE THE BAG EVERY DAY SHE SAW THAT THE WATER WAS STILL IN THE BAG BUT DIDN'T CALL TO ASK WHY. THE PATIENT WAS ADMITTED TO ICU ON (B)(6) 2025 FOR DEHYDRATION. AND PASSED AWAY ON (B)(6) 2025. ADDITIONAL INFORMATION RECEIVED: THE PATIENT WAS ADMITTED (B)(6) 2025, FOR ENTERAL THERAPY WITH DYSPHAGIA RELATED TO CVA AND IV ANTIBIOTIC THERAPY FOR POSITIVE BLOOD CULTURES. MEDICAL HISTORY DOCUMENTED AT THAT TIME INCLUDED HYPERLIPIDEMIA, HYPERCHOLESTEROLEMIA, TYPE II DIABETES MELLITUS, OSTEOPOROSIS, DEGENERATIVE LUMBAR SPINE DISEASE, AND SUSPECTED DEMENTIA. Manufacturer Narrative: THE DEVICE WAS RECEIVED MAY 27TH, 2025. THE DEVICE WAS RECEIVED ON SOFTWARE VERSION 1.5.4612. UPON RECEIPT, THE DEVICE WAS DISINFECTED USING ISOPROPANOL ALCOHOL. THE DEVICE WAS SUBJECTED TO A VISUAL INSPECTION TO DETERMINE IF ANY OBVIOUS DAMAGE OR ABNORMALITIES COULD BE DETECTED. THE DEVICE APPEARED TO BE NORMAL WITH NO DAMAGE OR MISSING PARTS DETECTED. DEVICE INTAKE, VISUAL INSPECTION, AND CERTIFICATION WAS COMPLETED. THE DEVICE HAD A STICKER ON THE BACK INDICATING THAT IT HAD UNDERGONE PREVENTATIVE MAINTENANCE AND PASSED TESTING BY (B)(4) ON 3/20/2025. THIS IS SUPPORTED BY THE DEVICE STATING THAT IT HAD LAST PASSED CERTIFICATION 68 DAYS AGO (AS OF 5/27/2025). AS PART OF END OF LINE TESTING, A KANGAROO OMNI PUMP MUST PASS FACTORY CERTIFICATION. THE FACTORY CERTIFICATION FOR THIS SERIAL NUMBER PASSED ON THE FIRST ATTEMPT ON NOVEMBER 12TH, 2024. AFTER THE DEVICE WAS RECEIVED ON MAY 27TH, A SERVICE CERTIFICATION WAS RUN AND PASSED. THE DEVICE WAS POWERED ON AND THE HISTORY FUNCTION WAS ENGAGED. THE HISTORY FEATURE RECORDS AND STORES IN MEMORY 30 DAYS OF DELIVERY VOLUMES. THE HISTORY FEATURE ALLOWS REVIEW IN 1-HOUR INCREMENTS FOR THE LAST 72 HOURS AND REVIEW OF 30 DAYS OF HISTORY IN 1-DAY INCREMENTS. REVIEW OF THIS DEVICE INDICATED NO HISTORY RECORDED UNTIL DAY 24. HISTORY DATA WAS PRESENT FOR DAYS 24 THROUGH 30. THE FEED TOTALIZER HAS A TOTAL FED OF 5601ML AND A TOTAL FLUSHED OF 9ML. THIS INDICATES THAT THE PUMP FED AT LEAST 2,534ML THAT WAS CAPTURED OUTSIDE OF THE HISTORY PERIOD. WHEN THE DEVICE WAS RECEIVED, THE KEEP SETTINGS BUTTON WAS PRESSED TO RECORD HOW THE USER PROGRAMMED THE DEVICE. THE DEVICE WAS PROGRAMMED IN CONTINUOUS MODE WITH A RATE OF 40ML/HR USING A STANDARD FEED/FLUSH SET. THE FLUSH SETTINGS WERE INCOMPLETE WITH ONLY THE FLUSH VOLUME OF 40ML PROGRAMMED AND NO FLUSH INTERVAL PROGRAMMED. FAILURE TO PROGRAM BOTH THE FLUSH VOLUME AND FLUSH INTERVAL WILL RESULT IN THE DEVICE NOT PERFORMING THE FLUSH OPERATION AS THE INPUT WAS NOT FULLY PROGRAMMED. NAVIGATING TO THE SHOW FLUSH PAGES ON EITHER THE READY TO FEED OR FEEDING SCREENS ALSO SHOW THAT THE FLUSH WAS NOT PROPERLY PROGRAMMED. THE KANGAROO OMNI ENTERAL FEEDING PUMP INVOLVED IN COMPLAINT PASSED ALL TESTING AND WAS DETERMINED TO BE FUNCTIONING AS EXPECTED. ANALYSIS OF THE PROGRAM SETTINGS, ALONG WITH REVIEW OF THE HISTORY INDICATE INCOMPLETE PROGRAMMING OF THE FLUSH SETTINGS BY THE USER RESULTED IN THE PUMP NOT FLUSHING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22152273&pc=DSQ | 3007042319-2025-01235 | 2025-05-29 04:00 | Death | HEARTWARE, INC. | 2025-06-05 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A MEDICAL HISTORY OF CARDIOMYOPATHY AND AN ADVANCED STATE OF ILLNESS, WHO HAD A VENTRICULAR ASSIST DEVICE IMPLANTED, PASSED AWAY DUE TO A TERMINAL CLINICAL CONDITION. THE DEATH OCCURRED PEACEFULLY IN A HOSPICE SETTING AND WAS UNRELATED TO THE VENTRICULAR ASSIST DEVICE THERAPY. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. D4: UDI INFORMATION IS UNABLE TO BE OBTAINED AS THE PRODUCT WAS DISTRIBUTED OUTSIDE OF THE UNITED STATES AND D4 UDI IS THEREFORE BLANK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO A TERMINAL CLINICAL CONDITION. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22152336&pc=NGV | 2124215-2025-36967 |  | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-05 04:00 | NGV | WATCHMAN FLX LEFT ATRIAL APPENDAGE CLOSURE DEVICE WITH DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | RWD2400413 | 542 | Event Description: THIS REPORT SUMMARIZES 542 DEATH EVENTS. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. OUT OF THE 542 REPORTED DEATH EVENTS: 11 WERE REPORTED AT DISCHARGE (OR WITHIN 7 DAYS OF IMPLANT, WHICHEVER IS LATER) 50 WERE REPORTED 45 DAYS AFTER IMPLANT 481 WERE REPORTED 1 YEAR POST IMPLANT. MEAN POPULATION AGE AT IMPLANT: 76 YEARS GENDER: FEMALE 4740 (40.8 %), MALE 6878 (59.2%). TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC PERFORMED A RETROSPECTIVE DATA ANALYSIS USING THE AGGREGATE ELECTRONIC HEALTH RECORD DATABASE FROM TRUVETA. THIS ANALYSIS AIMED TO EVALUATE ADVERSE EVENT RATES RELATED TO THE WATCHMAN FLX DEVICE IN REAL-WORLD USE. THE IMPLANT PROCEDURES WERE PERFORMED FROM JANUARY 2015 TO APRIL 2024. WITHIN THIS ANALYSIS, THE TOTAL PATIENT POPULATION SIZE WAS 11 618 INDIVIDUALS. PATIENTS HAD AN AVERAGE FOLLOW-UP DURATION OF 15.4 DAYS BETWEEN IMPLANT AND DATA SNAPSHOT. ADVERSE EVENTS WERE ASSESSED AT DISCHARGE (OR WITHIN 7 DAYS OF IMPLANT, WHICHEVER IS LATER), 45 DAYS, AND 1 YEAR. THESE EVENTS INCLUDE THE OCCURRENCE OF ALL-CAUSE MORTALITY. PATIENT EVENTS WERE IDENTIFIED AS EVENT TERMS WITH RATES. DATA OBTAINED FROM TRUVETA FOR THIS STUDY IS DE-IDENTIFIED BEFORE BEING ACCESSED BY BOSTON SCIENTIFIC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO OUR ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE STUDY DATA DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC.Event Description: THIS REPORT SUMMARIZES 542 DEATH EVENTS. DEVICE RELATIONSHIP TO THE EVENTS REPORTED IS NOT PROVIDED. OUT OF THE 542 REPORTED DEATH EVENTS: 11 WERE REPORTED AT DISCHARGE (OR WITHIN 7 DAYS OF IMPLANT, WHICHEVER IS LATER). 50 WERE REPORTED 45 DAYS AFTER IMPLANT. 481 WERE REPORTED 1 YEAR POST IMPLANT. MEAN POPULATION AGE AT IMPLANT: 76 YEARS. GENDER: FEMALE 4740 (40.8 %), MALE 6878 (59.2%). TYPE OF PROCEDURE: LEFT ATRIAL APPENDAGE CLOSURE. BOSTON SCIENTIFIC PERFORMED A RETROSPECTIVE DATA ANALYSIS USING THE AGGREGATE ELECTRONIC HEALTH RECORD DATABASE FROM TRUVETA. THIS ANALYSIS AIMED TO EVALUATE ADVERSE EVENT RATES RELATED TO THE WATCHMAN FLX DEVICE IN REAL-WORLD USE. THE IMPLANT PROCEDURES WERE PERFORMED FROM JANUARY 2015 TO APRIL 2024. WITHIN THIS ANALYSIS, THE TOTAL PATIENT POPULATION SIZE WAS 11 618 INDIVIDUALS. PATIENTS HAD AN AVERAGE FOLLOW-UP DURATION OF 15.4 DAYS BETWEEN IMPLANT AND DATA SNAPSHOT. ADVERSE EVENTS WERE ASSESSED AT DISCHARGE (OR WITHIN 7 DAYS OF IMPLANT, WHICHEVER IS LATER), 45 DAYS, AND 1 YEAR. THESE EVENTS INCLUDE THE OCCURRENCE OF ALL-CAUSE MORTALITY. PATIENT EVENTS WERE IDENTIFIED AS EVENT TERMS WITH RATES. DATA OBTAINED FROM TRUVETA FOR THIS STUDY IS DE-IDENTIFIED BEFORE BEING ACCESSED BY BOSTON SCIENTIFIC, THUS THERE ARE SIGNIFICANT LIMITATIONS TO OUR ABILITY TO CORRELATE THE DATA TO INFORMATION PREVIOUSLY REPORTED AS A SPONTANEOUS COMPLAINT. THE STUDY DATA DOES NOT PROVIDE A CAUSALITY RELATIONSHIP FOR EACH REPORTED EVENT TO THE DEVICE. NO FURTHER INFORMATION IS AVAILABLE TO BSC.Event Description: Unique Complaint ID Number,Event Date,Date Entered,Manufacturer Aware Date,Brand Name,Generic Name,Model Number,Lot Number,Catalog Number,Serial Number,UDI Number,PMA / 510K Number,Date Returned to Manufacturer,Type of Reportable Event,Event Description,Manufacturer Narrative,Medical History,Patient Age,Patient Gender,Patient Weight,Date Implanted,Date Explanted,Health Effect Clinical Code,Health Effect Impact Code,Device Problem Code,Device Component Code,Investigation Type Code,Investigation Findings Code,Investigation Conclusion Code,Remedial Action Type,Latest Line Item Version
<br>19870445,01/01/2015,06/05/2025,03/18/2025,WATCHMAN FLX,"system, appendage closure, left atrial",Unknown,Unknown,Unknown,NA,Unknown,P130013 ,03/18/2025,D,"It was reported that a death occurred at discharge or within 7 days of implant. <br><br>Type of Procedure: Left Atrial Appendage Closure<br><br>Boston Scientific performed a retrospective data analysis using the aggregate electronic health record database from Truveta. This analysis aimed to evaluate adverse event rates related to the WATCHMAN FLX device in real-world use. The implant procedures were performed from January 2015 to April 2024. Patient events were identified as event terms with rates. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC. <br><br>Patient events were identified as event terms with rates. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC.<br><br>Race: American Indian or Alaska Native 29(0.2 %), Asian 165(1.4 %), Black 401(3.5 %), Other Pacific Islander 5(0.0 %), Other Race 142(1.2 %), Unknown 676(5.8 %), and White 10200(87.8 %).<br>Ethnicity: Hispanic or Latino 162(1.6 %), and Not Hispanic or Latino 10128(98.4 %).","Time Frame of Event Date Collection: Patients in TRUVETA health systems implanted with WATCHMAN FLX device between Jan 2015 and April 2024.<br><br>Summary of Adverse Events: Patients had an average follow-up duration of 15.4 ± 10.8 days between WATCHMAN FLX implant and data snapshot. Adverse events were assessed at discharge (or within 7 days of implant, whichever is later), 45 days, and 1 year. These events include the occurrence of all-cause mortality, any stroke, ischemic stroke, hemorrhagic stroke, systemic embolism, major bleeding, and pericardial effusion requiring intervention.<br><br>Total Number of Patients: 11618<br><br>Device Evaluation: Under the terms and conditions of the study, anonymized data was provided. No products were returned as part of the study. It cannot be determined if these events have been previously reported or if the devices were returned for analysis as part of a previously reported spontaneous complaint. However, the events reported were anticipated in nature as defined by the ""Potential Adverse Events"" list in the FDA-approved Instructions for Use (IFU).<br><br>Contextual Analysis of Study Data:<br><br>Publications of 1-year event rates from similar nationwide, retrospective analyses looking at data from 2016-2018(1) and 2020-2022(2) have reported rates of all-cause death from 8.2%(2) - 8.5%(1), all stroke from 1.5%(2) - 2.1%(1), ischemic stroke from 1.2%(2) - 1.5%(1), hemorrhagic stroke from 0.26%(2) - 0.46%(1), major bleeding from 6.4%(2) - 6.9%(1), and systemic embolism from 0.10%(2) - 0.68%(1). Pericardial effusion was reported at 45-days with a rate of 0.50%.(2)<br><br>Overall, 1-year clinical outcomes after WATCHMAN FLX implantation from Truveta align with expectations derived from published literature. The 1-year event rates found in this analysis for death (5.9%), all stroke (2.0%), hemorrhagic stroke (0.41%), systemic embolism (0.10%), and major bleeding (3.0%) are within or lower than reported rates from analyses of similar nationwide registries of published WATCHMAN literature. The 45-day rate for pericardial effusion requiring intervention (0.31%) as well as the 1-year rate (0.41%) are lower than the published 45-day rate and thus are in-line with expectations. The 1-year ischemic stroke rate (1.7%) is within the 95% confidence interval for the upper bound of the range reported in the literature (1.4% - 1.7%)(1). A similar Japanese nationwide registry published an ischemic event rate of 2.6%(3) for WATCHMAN FLX and single center WATCHMAN studies have reported rates of 2.0%.(4,5) As such, the rate found in this analysis aligns with expectations from published literature. Thus, all rates found in this analysis align with expectations derived from similar published literature.<br><br>References<br>1. Price MJ, Slotwiner D, Du C, et al. Clinical Outcomes at 1 Year Following Transcatheter Left Atrial Appendage Occlusion in the United States. JACC Cardiovasc Interv. Apr 11 2022;15(7):741-750. doi:10.1016/j.jcin.2022.02.009<br>2. Kapadia SR, Yeh RW, Price MJ, et al. Outcomes With the WATCHMAN FLX in Everyday Clinical Practice From the NCDR Left Atrial Appendage Occlusion Registry. Circ Cardiovasc Interv. Sep 2024;17(9):e013750. doi:10.1161/CIRCINTERVENTIONS.123.013750<br>3. Nakashima M, Yamamoto M, Sago M, et al. Comparative Data of Procedural and Midterm Outcomes in Patients Who Underwent Percutaneous Left Atrial Appendage Closure Between the WATCHMAN FLX and WATCHMAN 2.5 Devices?- Insight From the OCEAN-LAAC Registry. Circ J. Jun 25 2024;88(7):1187-1197. doi:10.1253/circj.CJ-24-0062<br>4. Huang WP, Zhang YH, He L, Su X, Yang XW, Guo ZX. Efficacy and Safety of the WATCHMAN Left Atrial Appendage System for Stroke Prevention in Chinese Patients with Nonvalvular Atrial Fibrillation: A Single-center, Prospective, Observational Study. Chin Med J (Engl). Feb 20 2017;130(4):434-438. doi:10.4103/0366-6999.199832<br>5. Hana D, Miller T, Chaker Z, et al. Evaluating Gender-based Differences in Clinical Outcomes for Patients Undergoing Left Atrial Appendage Occlusion: A Single Centre Experience. Curr Probl Cardiol. Mar 2023;48(3):101532. doi:10.1016/j.cpcardiol.2022.101532<br><br>Race: American Indian or Alaska Native 29(0.2 %), Asian 165(1.4 %), Black 401(3.5 %), Other Pacific Islander 5(0.0 %), Other Race 142(1.2 %), Unknown 676(5.8 %), and White 10200(87.8 %).<br>Ethnicity: Hispanic or Latino 162(1.6 %), and Not Hispanic or Latino 10128(98.4 %).<br><br>Detailed product information was not provided to BSC. Because the product is unknown at this time, we are unable to provide the complete Unique Identifier (UDI) # and other specific product information.",,Average 76 years,"Female 4740 (40.8 %)<br>Male 6878 (59.2%)",,,,E2403,F02,A24,G07001,B17,C19,D12,,0
<br>19870445,01/01/2015,06/05/2025,03/18/2025,WATCHMAN FLX,"system, appendage closure, left atrial",Unknown,Unknown,Unknown,NA,Unknown,P130013 ,03/18/2025,D,"It was reported that a death occurred at discharge or within 7 days of implant. <br><br>Type of Procedure: Left Atrial Appendage Closure<br><br>Boston Scientific performed a retrospective data analysis using the aggregate electronic health record database from Truveta. This analysis aimed to evaluate adverse event rates related to the WATCHMAN FLX device in real-world use. The implant procedures were performed from January 2015 to April 2024. Patient events were identified as event terms with rates. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC. <br><br>Patient events were identified as event terms with rates. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC.<br><br>Race: American Indian or Alaska Native 29(0.2 %), Asian 165(1.4 %), Black 401(3.5 %), Other Pacific Islander 5(0.0 %), Other Race 142(1.2 %), Unknown 676(5.8 %), and White 10200(87.8 %).<br>Ethnicity: Hispanic or Latino 162(1.6 %), and Not Hispanic or Latino 10128(98.4 %).","Time Frame of Event Date Collection: Patients in TRUVETA health systems implanted with WATCHMAN FLX device between Jan 2015 and April 2024.<br><br>Summary of Adverse Events: Patients had an average follow-up duration of 15.4 ± 10.8 days between WATCHMAN FLX implant and data snapshot. Adverse events were assessed at discharge (or within 7 days of implant, whichever is later), 45 days, and 1 year. These events include the occurrence of all-cause mortality, any stroke, ischemic stroke, hemorrhagic stroke, systemic embolism, major bleeding, and pericardial effusion requiring intervention.<br><br>Total Number of Patients: 11618<br><br>Device Evaluation: Under the terms and conditions of the study, anonymized data was provided. No products were returned as part of the study. It cannot be determined if these events have been previously reported or if the devices were returned for analysis as part of a previously reported spontaneous complaint. However, the events reported were anticipated in nature as defined by the ""Potential Adverse Events"" list in the FDA-approved Instructions for Use (IFU).<br><br>Contextual Analysis of Study Data:<br><br>Publications of 1-year event rates from similar nationwide, retrospective analyses looking at data from 2016-2018(1) and 2020-2022(2) have reported rates of all-cause death from 8.2%(2) - 8.5%(1), all stroke from 1.5%(2) - 2.1%(1), ischemic stroke from 1.2%(2) - 1.5%(1), hemorrhagic stroke from 0.26%(2) - 0.46%(1), major bleeding from 6.4%(2) - 6.9%(1), and systemic embolism from 0.10%(2) - 0.68%(1). Pericardial effusion was reported at 45-days with a rate of 0.50%.(2)<br><br>Overall, 1-year clinical outcomes after WATCHMAN FLX implantation from Truveta align with expectations derived from published literature. The 1-year event rates found in this analysis for death (5.9%), all stroke (2.0%), hemorrhagic stroke (0.41%), systemic embolism (0.10%), and major bleeding (3.0%) are within or lower than reported rates from analyses of similar nationwide registries of published WATCHMAN literature. The 45-day rate for pericardial effusion requiring intervention (0.31%) as well as the 1-year rate (0.41%) are lower than the published 45-day rate and thus are in-line with expectations. The 1-year ischemic stroke rate (1.7%) is within the 95% confidence interval for the upper bound of the range reported in the literature (1.4% - 1.7%)(1). A similar Japanese nationwide registry published an ischemic event rate of 2.6%(3) for WATCHMAN FLX and single center WATCHMAN studies have reported rates of 2.0%.(4,5) As such, the rate found in this analysis aligns with expectations from published literature. Thus, all rates found in this analysis align with expectations derived from similar published literature.<br><br>References<br>1. Price MJ, Slotwiner D, Du C, et al. Clinical Outcomes at 1 Year Following Transcatheter Left Atrial Appendage Occlusion in the United States. JACC Cardiovasc Interv. Apr 11 2022;15(7):741-750. doi:10.1016/j.jcin.2022.02.009<br>2. Kapadia SR, Yeh RW, Price MJ, et al. Outcomes With the WATCHMAN FLX in Everyday Clinical Practice From the NCDR Left Atrial Appendage Occlusion Registry. Circ Cardiovasc Interv. Sep 2024;17(9):e013750. doi:10.1161/CIRCINTERVENTIONS.123.013750<br>3. Nakashima M, Yamamoto M, Sago M, et al. Comparative Data of Procedural and Midterm Outcomes in Patients Who Underwent Percutaneous Left Atrial Appendage Closure Between the WATCHMAN FLX and WATCHMAN 2.5 Devices?- Insight From the OCEAN-LAAC Registry. Circ J. Jun 25 2024;88(7):1187-1197. doi:10.1253/circj.CJ-24-0062<br>4. Huang WP, Zhang YH, He L, Su X, Yang XW, Guo ZX. Efficacy and Safety of the WATCHMAN Left Atrial Appendage System for Stroke Prevention in Chinese Patients with Nonvalvular Atrial Fibrillation: A Single-center, Prospective, Observational Study. Chin Med J (Engl). Feb 20 2017;130(4):434-438. doi:10.4103/0366-6999.199832<br>5. Hana D, Miller T, Chaker Z, et al. Evaluating Gender-based Differences in Clinical Outcomes for Patients Undergoing Left Atrial Appendage Occlusion: A Single Centre Experience. Curr Probl Cardiol. Mar 2023;48(3):101532. doi:10.1016/j.cpcardiol.2022.101532<br><br>Race: American Indian or Alaska Native 29(0.2 %), Asian 165(1.4 %), Black 401(3.5 %), Other Pacific Islander 5(0.0 %), Other Race 142(1.2 %), Unknown 676(5.8 %), and White 10200(87.8 %).<br>Ethnicity: Hispanic or Latino 162(1.6 %), and Not Hispanic or Latino 10128(98.4 %).<br><br>Detailed product information was not provided to BSC. Because the product is unknown at this time, we are unable to provide the complete Unique Identifier (UDI) # and other specific product information.",,Average 76 years,"Female 4740 (40.8 %)<br>Male 6878 (59.2%)",,,,E2403,F02,A24,G07001,B17,C19,D12,,0
<br>19870445,01/01/2015,06/05/2025,03/18/2025,WATCHMAN FLX,"system, appendage closure, left atrial",Unknown,Unknown,Unknown,NA,Unknown,P130013 ,03/18/2025,D,"It was reported that a death occurred at discharge or within 7 days of implant. <br><br>Type of Procedure: Left Atrial Appendage Closure<br><br>Boston Scientific performed a retrospective data analysis using the aggregate electronic health record database from Truveta. This analysis aimed to evaluate adverse event rates related to the WATCHMAN FLX device in real-world use. The implant procedures were performed from January 2015 to April 2024. Patient events were identified as event terms with rates. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC. <br><br>Patient events were identified as event terms with rates. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC.<br><br>Race: American Indian or Alaska Native 29(0.2 %), Asian 165(1.4 %), Black 401(3.5 %), Other Pacific Islander 5(0.0 %), Other Race 142(1.2 %), Unknown 676(5.8 %), and White 10200(87.8 %).<br>Ethnicity: Hispanic or Latino 162(1.6 %), and Not Hispanic or Latino 10128(98.4 %).","Time Frame of Event Date Collection: Patients in TRUVETA health systems implanted with WATCHMAN FLX device between Jan 2015 and April 2024.<br><br>Summary of Adverse Events: Patients had an average follow-up duration of 15.4 ± 10.8 days between WATCHMAN FLX implant and data snapshot. Adverse events were assessed at discharge (or within 7 days of implant, whichever is later), 45 days, and 1 year. These events include the occurrence of all-cause mortality, any stroke, ischemic stroke, hemorrhagic stroke, systemic embolism, major bleeding, and pericardial effusion requiring intervention.<br><br>Total Number of Patients: 11618<br><br>Device Evaluation: Under the terms and conditions of the study, anonymized data was provided. No products were returned as part of the study. It cannot be determined if these events have been previously reported or if the devices were returned for analysis as part of a previously reported spontaneous complaint. However, the events reported were anticipated in nature as defined by the ""Potential Adverse Events"" list in the FDA-approved Instructions for Use (IFU).<br><br>Contextual Analysis of Study Data:<br><br>Publications of 1-year event rates from similar nationwide, retrospective analyses looking at data from 2016-2018(1) and 2020-2022(2) have reported rates of all-cause death from 8.2%(2) - 8.5%(1), all stroke from 1.5%(2) - 2.1%(1), ischemic stroke from 1.2%(2) - 1.5%(1), hemorrhagic stroke from 0.26%(2) - 0.46%(1), major bleeding from 6.4%(2) - 6.9%(1), and systemic embolism from 0.10%(2) - 0.68%(1). Pericardial effusion was reported at 45-days with a rate of 0.50%.(2)<br><br>Overall, 1-year clinical outcomes after WATCHMAN FLX implantation from Truveta align with expectations derived from published literature. The 1-year event rates found in this analysis for death (5.9%), all stroke (2.0%), hemorrhagic stroke (0.41%), systemic embolism (0.10%), and major bleeding (3.0%) are within or lower than reported rates from analyses of similar nationwide registries of published WATCHMAN literature. The 45-day rate for pericardial effusion requiring intervention (0.31%) as well as the 1-year rate (0.41%) are lower than the published 45-day rate and thus are in-line with expectations. The 1-year ischemic stroke rate (1.7%) is within the 95% confidence interval for the upper bound of the range reported in the literature (1.4% - 1.7%)(1). A similar Japanese nationwide registry published an ischemic event rate of 2.6%(3) for WATCHMAN FLX and single center WATCHMAN studies have reported rates of 2.0%.(4,5) As such, the rate found in this analysis aligns with expectations from published literature. Thus, all rates found in this analysis align with expectations derived from similar published literature.<br><br>References<br>1. Price MJ, Slotwiner D, Du C, et al. Clinical Outcomes at 1 Year Following Transcatheter Left Atrial Appendage Occlusion in the United States. JACC Cardiovasc Interv. Apr 11 2022;15(7):741-750. doi:10.1016/j.jcin.2022.02.009<br>2. Kapadia SR, Yeh RW, Price MJ, et al. Outcomes With the WATCHMAN FLX in Everyday Clinical Practice From the NCDR Left Atrial Appendage Occlusion Registry. Circ Cardiovasc Interv. Sep 2024;17(9):e013750. doi:10.1161/CIRCINTERVENTIONS.123.013750<br>3. Nakashima M, Yamamoto M, Sago M, et al. Comparative Data of Procedural and Midterm Outcomes in Patients Who Underwent Percutaneous Left Atrial Appendage Closure Between the WATCHMAN FLX and WATCHMAN 2.5 Devices?- Insight From the OCEAN-LAAC Registry. Circ J. Jun 25 2024;88(7):1187-1197. doi:10.1253/circj.CJ-24-0062<br>4. Huang WP, Zhang YH, He L, Su X, Yang XW, Guo ZX. Efficacy and Safety of the WATCHMAN Left Atrial Appendage System for Stroke Prevention in Chinese Patients with Nonvalvular Atrial Fibrillation: A Single-center, Prospective, Observational Study. Chin Med J (Engl). Feb 20 2017;130(4):434-438. doi:10.4103/0366-6999.199832<br>5. Hana D, Miller T, Chaker Z, et al. Evaluating Gender-based Differences in Clinical Outcomes for Patients Undergoing Left Atrial Appendage Occlusion: A Single Centre Experience. Curr Probl Cardiol. Mar 2023;48(3):101532. doi:10.1016/j.cpcardiol.2022.101532<br><br>Race: American Indian or Alaska Native 29(0.2 %), Asian 165(1.4 %), Black 401(3.5 %), Other Pacific Islander 5(0.0 %), Other Race 142(1.2 %), Unknown 676(5.8 %), and White 10200(87.8 %).<br>Ethnicity: Hispanic or Latino 162(1.6 %), and Not Hispanic or Latino 10128(98.4 %).<br><br>Detailed product information was not provided to BSC. Because the product is unknown at this time, we are unable to provide the complete Unique Identifier (UDI) # and other specific product information.",,Average 76 years,"Female 4740 (40.8 %)<br>Male 6878 (59.2%)",,,,E2403,F02,A24,G07001,B17,C19,D12,,0
<br>19870445,01/01/2015,06/05/2025,03/18/2025,WATCHMAN FLX,"system, appendage closure, left atrial",Unknown,Unknown,Unknown,NA,Unknown,P130013 ,03/18/2025,D,"It was reported that a death occurred at discharge or within 7 days of implant. <br><br>Type of Procedure: Left Atrial Appendage Closure<br><br>Boston Scientific performed a retrospective data analysis using the aggregate electronic health record database from Truveta. This analysis aimed to evaluate adverse event rates related to the WATCHMAN FLX device in real-world use. The implant procedures were performed from January 2015 to April 2024. Patient events were identified as event terms with rates. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC. <br><br>Patient events were identified as event terms with rates. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC.<br><br>Race: American Indian or Alaska Native 29(0.2 %), Asian 165(1.4 %), Black 401(3.5 %), Other Pacific Islander 5(0.0 %), Other Race 142(1.2 %), Unknown 676(5.8 %), and White 10200(87.8 %).<br>Ethnicity: Hispanic or Latino 162(1.6 %), and Not Hispanic or Latino 10128(98.4 %).","Time Frame of Event Date Collection: Patients in TRUVETA health systems implanted with WATCHMAN FLX device between Jan 2015 and April 2024.<br><br>Summary of Adverse Events: Patients had an average follow-up duration of 15.4 ± 10.8 days between WATCHMAN FLX implant and data snapshot. Adverse events were assessed at discharge (or within 7 days of implant, whichever is later), 45 days, and 1 year. These events include the occurrence of all-cause mortality, any stroke, ischemic stroke, hemorrhagic stroke, systemic embolism, major bleeding, and pericardial effusion requiring intervention.<br><br>Total Number of Patients: 11618<br><br>Device Evaluation: Under the terms and conditions of the study, anonymized data was provided. No products were returned as part of the study. It cannot be determined if these events have been previously reported or if the devices were returned for analysis as part of a previously reported spontaneous complaint. However, the events reported were anticipated in nature as defined by the ""Potential Adverse Events"" list in the FDA-approved Instructions for Use (IFU).<br><br>Contextual Analysis of Study Data:<br><br>Publications of 1-year event rates from similar nationwide, retrospective analyses looking at data from 2016-2018(1) and 2020-2022(2) have reported rates of all-cause death from 8.2%(2) - 8.5%(1), all stroke from 1.5%(2) - 2.1%(1), ischemic stroke from 1.2%(2) - 1.5%(1), hemorrhagic stroke from 0.26%(2) - 0.46%(1), major bleeding from 6.4%(2) - 6.9%(1), and systemic embolism from 0.10%(2) - 0.68%(1). Pericardial effusion was reported at 45-days with a rate of 0.50%.(2)<br><br>Overall, 1-year clinical outcomes after WATCHMAN FLX implantation from Truveta align with expectations derived from published literature. The 1-year event rates found in this analysis for death (5.9%), all stroke (2.0%), hemorrhagic stroke (0.41%), systemic embolism (0.10%), and major bleeding (3.0%) are within or lower than reported rates from analyses of similar nationwide registries of published WATCHMAN literature. The 45-day rate for pericardial effusion requiring intervention (0.31%) as well as the 1-year rate (0.41%) are lower than the published 45-day rate and thus are in-line with expectations. The 1-year ischemic stroke rate (1.7%) is within the 95% confidence interval for the upper bound of the range reported in the literature (1.4% - 1.7%)(1). A similar Japanese nationwide registry published an ischemic event rate of 2.6%(3) for WATCHMAN FLX and single center WATCHMAN studies have reported rates of 2.0%.(4,5) As such, the rate found in this analysis aligns with expectations from published literature. Thus, all rates found in this analysis align with expectations derived from similar published literature.<br><br>References<br>1. Price MJ, Slotwiner D, Du C, et al. Clinical Outcomes at 1 Year Following Transcatheter Left Atrial Appendage Occlusion in the United States. JACC Cardiovasc Interv. Apr 11 2022;15(7):741-750. doi:10.1016/j.jcin.2022.02.009<br>2. Kapadia SR, Yeh RW, Price MJ, et al. Outcomes With the WATCHMAN FLX in Everyday Clinical Practice From the NCDR Left Atrial Appendage Occlusion Registry. Circ Cardiovasc Interv. Sep 2024;17(9):e013750. doi:10.1161/CIRCINTERVENTIONS.123.013750<br>3. Nakashima M, Yamamoto M, Sago M, et al. Comparative Data of Procedural and Midterm Outcomes in Patients Who Underwent Percutaneous Left Atrial Appendage Closure Between the WATCHMAN FLX and WATCHMAN 2.5 Devices?- Insight From the OCEAN-LAAC Registry. Circ J. Jun 25 2024;88(7):1187-1197. doi:10.1253/circj.CJ-24-0062<br>4. Huang WP, Zhang YH, He L, Su X, Yang XW, Guo ZX. Efficacy and Safety of the WATCHMAN Left Atrial Appendage System for Stroke Prevention in Chinese Patients with Nonvalvular Atrial Fibrillation: A Single-center, Prospective, Observational Study. Chin Med J (Engl). Feb 20 2017;130(4):434-438. doi:10.4103/0366-6999.199832<br>5. Hana D, Miller T, Chaker Z, et al. Evaluating Gender-based Differences in Clinical Outcomes for Patients Undergoing Left Atrial Appendage Occlusion: A Single Centre Experience. Curr Probl Cardiol. Mar 2023;48(3):101532. doi:10.1016/j.cpcardiol.2022.101532<br><br>Race: American Indian or Alaska Native 29(0.2 %), Asian 165(1.4 %), Black 401(3.5 %), Other Pacific Islander 5(0.0 %), Other Race 142(1.2 %), Unknown 676(5.8 %), and White 10200(87.8 %).<br>Ethnicity: Hispanic or Latino 162(1.6 %), and Not Hispanic or Latino 10128(98.4 %).<br><br>Detailed product information was not provided to BSC. Because the product is unknown at this time, we are unable to provide the complete Unique Identifier (UDI) # and other specific product information.",,Average 76 years,"Female 4740 (40.8 %)<br>Male 6878 (59.2%)",,,,E2403,F02,A24,G07001,B17,C19,D12,,0
<br>19870445,01/01/2015,06/05/2025,03/18/2025,WATCHMAN FLX,"system, appendage closure, left atrial",Unknown,Unknown,Unknown,NA,Unknown,P130013 ,03/18/2025,D,"It was reported that a death occurred at discharge or within 7 days of implant. <br><br>Type of Procedure: Left Atrial Appendage Closure<br><br>Boston Scientific performed a retrospective data analysis using the aggregate electronic health record database from Truveta. This analysis aimed to evaluate adverse event rates related to the WATCHMAN FLX device in real-world use. The implant procedures were performed from January 2015 to April 2024. Patient events were identified as event terms with rates. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC. <br><br>Patient events were identified as event terms with rates. Data obtained from Truveta for this study is de-identified before being accessed by Boston Scientific, thus there are significant limitations to our ability to correlate the data to information previously reported as a spontaneous complaint. The study data does not provide a causality relationship for each reported event to the device. No further information is available to BSC.<br><br>Race: American Indian or Alaska Native 29(0.2 %), Asian 165(1.4 %), Black 401(3.5 %), Other Pacific Islander 5(0.0 %), Other Race 142(1.2 %), Unknown 676(5.8 %), and White 10200(87.8 %).<br>Ethnicity: Hispanic or Latino 162(1.6 %), and Not Hispanic or Latino 10128(98.4 %).","Time Frame of Event Date Collection: Patients in TRUVETA health systems implanted with WATCHMAN FLX device between Jan 2015 and April 2024.<br><br>Summary of Adverse Events: Patients had an average follow-up duration of 15.4 ± 10.8 days between WATCHMAN FLX implant and data snapshot. Adverse events were assessed at discharge (or within 7 days of implant, whichever is later), 45 days, and 1 year. These events include the occurrence of all-cause mortality, any stroke, ischemic stroke, hemorrhagic stroke, systemic embolism, major bleeding, and pericardial effusion requiring intervention.<br><br>Total Number of Patients: 11618<br><br>Device Evaluation: Under the terms and conditions of the study, anonymized data was provided. No products were returned as part of the study. It cannot be determined if ... Edited for length. See MAUDE Adverse Event Report for complete event narratives.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22152511&pc=DSQ | 2916596-2025-03777 | 2025-05-14 04:00 | Death | THORATEC CORPORATION | 2025-06-05 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-ORGAN FAILURE (MOF). THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE PATIENT SUFFERED FROM A PROGRESSIVE AND REFRACTORY MOF WHICH WAS RELATED TO THE OUTCOME AT LEAST. NO DEVICE RELATED ISSUES LEAD TO MOF, THE PUMP WAS RUNNING AS EXPECTED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22152723&pc=MIH | 3011063223-2025-00102 | 2025-05-04 04:00 | Death | ENDOLOGIX | 2025-06-05 04:00 | MIH | AFX2 | Lack of Effect | Ischemia; Paraplegia; Thrombosis/Thrombus | P040002 |  | 1 | Event Description: THE PATIENT WAS TREATED FOR AN ABDOMINAL AORTIC ANEURYSM (AAA) ON (B)(6) 2025 WITH THE IMPLANT OF AN AFX2 BIFURCATED STENT GRAFT AND AN AFX VELA SUPRARENAL. REPORTEDLY THERE WAS AN AORTIC ARCH ANEURYSM AS WELL AS THE AAA. THE CONDITION OF THE AORTA WAS POOR AND SHAGGY. ALSO, THERE WAS A THROMBUS FROM ABOVE THE RENAL ARTERIES NEAR THE SUPERIOR MESENTERIC ARTERY (SMA). THE AFX2 MAIN BODY AND AFX VELA SUPRARENAL WERE IMPLANTED THROUGH THE AFX SHEATH AS USUAL. HOWEVER, WHEN THE AFX VELA SUPRARENAL WAS ATTEMPTED TO BE PLACED AFTER THE AFX MAIN BODY, AN ANGIOGRAPHY DID NOT SHOW THE RENAL ARTERIES. THEREFORE, CANNULATION TO THE RENAL ARTERIES WAS PERFORMED VIA AN ARM. THE FOLLOWING ANGIOGRAPHY SHOWED THE RENAL ARTERIES, SO THE AFX VELA SUPRARENAL WAS IMPLANTED. SINCE THE JUNCTION OF THE AFX2 MAIN BODY AND THE AFX VELA SUPRARENAL, AND LT-COMMON ILIAC ARTERY (CIA) WERE NARROWED, BALLOON TOUCH-UP WAS PERFORMED. THE FINAL ANGIOGRAPHY SHOWED NO ENDOLEAKS, SO THE PROCEDURE WAS COMPLETED. UNFORTUNATELY, THE PATIENT HAD PARAPLEGIA AFTER THE SURGERY AND DIED FROM MESENTERIC ISCHEMIA AFTERWARDS. REPORTEDLY THE POST-OPERATIVE COMPUTED TOMOGRAPHY SCAN SHOWED A TYPE 1A ENDOLEAK AND A CONTRAST AGENT FLOWING INTO THE ANEURYSM, THUS TURNING OUT THERE HAD BEEN THROMBUS THAT MAY HAVE BEEN INADVERTENTLY CARRIED TO THE PROXIMAL. THE DATE OF DEATH OF THE PATIENT WAS REPORTEDLY CONFIRMED TO BE ON (B)(6) 2025. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. WHERE POSSIBLE, IT IS ENDOLOGIX PRACTICE TO MAKE AT LEAST THREE GOOD FAITH EFFORTS TO RETRIEVE A REPORTED ADVERSE EVENT/INCIDENT-RELATED DEVICE AS WELL AS MEDICAL RECORDS AND MEDICAL IMAGING. AN EVALUATION OF THE MANUFACTURING RECORD WAS COMPLETED. A REVIEW OF THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION SHOWS THE DEVICE TO HAVE BEEN PROPERLY MANUFACTURED AND RELEASED IN ACCORDANCE WITH THE DEVICE MASTER RECORD. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT WAS COMPLETED. AN EXAMINATION OF MEDICAL RECORDS AND/OR MEDICAL IMAGING RECEIVED BY ENDOLOGIX FOUND THE MESENTERIC ISCHEMIA, TYPE IA ENDOLEAK, PARAPLEGIA AND DEATH COMPLAINTS ARE UNCONFIRMED. THIS IS NOT CONSISTENT WITH THE REPORTED ADVERSE EVENT/INCIDENT. THE CLINICAL EVALUATION FOUND REASONABLE EVIDENCE TO SUGGEST THERE WAS A NARROWED LUMINAL DIAMETER OF THE ILIAC ARTERIES AT THE BIFURCATION, AND THE PATIENT UNDERWENT PERCUTANEOUS TRANSLUMINAL ANGIOPLASTY OF THE LEFT COMMON ILIAC ARTERY. WHILE THIS COULD HAVE CONTRIBUTED TO THE REPORTED EVENT, A DEFINITIVE CAUSAL RELATIONSHIP COULD NOT BE ESTABLISHED. THE CLINICAL EVALUATION SHOWS REASONABLE EVIDENCE TO SUGGEST THERE WAS MODERATE PREEXISTING THROMBUS IN THE AORTA. HOWEVER, IT IS UNCLEAR WHETHER THIS CONTRIBUTED TO THE REPORTED EVENT. DEVICE, USER, PROCEDURE OR ANATOMY RELATEDNESS OF COMPLAINT COULD NOT BE DETERMINED. PROCEDURE RELATED HARMS COULD NOT BE DETERMINED. DUE TO THE LACK OF INFORMATION, THE DEATH COULD NOT BE CONFIRMED HOWEVER, THE FINAL PATIENT STATUS WAS REPORTED AS DECEASED ON POSTOPERATIVE DAY TWO. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22152728&pc=OZD | 1220648-2025-29043 | 2025-05-16 04:00 | Death | ABIOMED, INC. | 2025-06-05 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A CP PUMP PLACED TO SUPPORT THE PATIENT FOUND DOWN OUTSIDE WITH VENTRICULAR TACHYCARDIA (VT) AND PLACED ON CP-ECMO SUPPORT. THE PUMP WAS PLACED BUT THE PATIENT STILL EXPIRED AFTER 3 DAYS. PATIENT HAD CEREBRAL EDEMA AND EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22152967&pc=DSQ | 2916596-2025-03518 | 2025-05-11 04:00 | Death | THORATEC CORPORATION | 2025-06-05 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO HOSPITAL PRIVACY POLICY, NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE HEARTMATE 3 LVAS IFU, REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22152986&pc=DSQ | 2916596-2025-03532 | 2025-05-14 04:00 | Death | THORATEC CORPORATION | 2025-06-05 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE WAS NOT PROVIDED/UNKNOWN. IT WAS REPORTED THAT THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. ACCORDING TO VENTRICULAR ASSIST DEVICE (VAD) COORDINATOR, THE VAD WAS OPERATING AS INTENDED, WITH NO DEVICE ISSUES.Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIORGAN SYSTEM FAILURE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED MULTISYSTEM ORGAN FAILURE (MSOF) AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22153001&pc=OZD | 1220648-2025-29045 | 2025-05-16 04:00 | Death | ABIOMED, INC. | 2025-06-05 04:00 | OZD | IMPELLA CP-JP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Cardiac Tamponade | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT EXPERIENCED A CARDIAC TAMPONADE AND DEVELOPED A BLEEDING TENDENCY DURING IMPELLA CP SUPPORT. THE RELATIONSHIP BETWEEN IMPELLA SUPPORT AND THE NOTED CONDITIONS WAS UNKNOWN. AFTER SEVEN DAYS OF SUPPORT, THE PATIENT PASSED AWAY ON SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22153005&pc=DSQ | 2916596-2025-03533 | 2025-05-14 04:00 | Death | THORATEC CORPORATION | 2025-06-05 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS SAID TO BE UNKNOWN. IT WAS REPORTED THAT THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED AND THAT THE VENTRICULAR ASSIST DEVICE WAS OPERATING AS INTENDED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS LATER REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN AS THE PATIENT PASSED AWAY IN THEIR SLEEP. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REVISION, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22153425&pc=MHX | 9610816-2025-000464 | 2025-05-10 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-06-05 04:00 | MHX | INTELLIVUE MULTI-MEASUREMENT MODULE X3 | Wireless Communication Problem | Insufficient Information | K171801 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT PASSED AWAY AFTER LOSS OF MONITORING WHILE TRANSPORTING THE PATIENT. THE CUSTOMER INDICATED WHEN THE PATIENT WAS TRANSPORTED TO OTHER AREAS OF THE HOSPITAL FOR TESTING, THE MONITOR LOSES NETWORK CONNECTIVITY; THEREFORE, THE PATIENT WAS NOT MONITORED AT THE CENTRAL STATION. THE ALARMS WERE PRESENT ON THE X3 DURING TRANSPORT. IT WAS INDICATED THE CUSTOMER WAS NOT ALLEGING THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THE PATIENT WAS BEING MONITORED FOR HEART RATE, SPO2, AND RESPIRATION. THEY WERE DNR AND IN COMFORT CARE AT THE TIME OF THE EVENT, AND THEY WERE IN POOR CONDITION. WHILE BEING TRANSPORTED TO CT, THE PATIENT'S STATUS DECLINED, AND THE NURSES DID NOT SEE THE CHANGE IN CONDITION AS THE X3 MONITOR DROPPED FROM THE CENTRAL STATION. THE RADIOLOGY TECHNICIAN HEARD THE X3 MONITORING BUT DID NOT KNOW WHAT THE MONITOR ALARMS MEANT. NURSES DO NOT TYPICALLY STAY WITH PATIENTS DURING TRANSPORT. NO RESUSCITATION ATTEMPTS WERE MADE DUE TO THE PATIENT'S DNR STATUS. IN ADDITION, IT WAS NOTED THERE ARE NETWORK ACCESS POINTS LOCATED OUTSIDE OF THE EMERGENCY DEPARTMENT TO THE CT LOCATION. Manufacturer Narrative: A PHILIPS FIELD SERVICE ENGINEER (FSE) WENT ONSITE AND COLLECTED THE CLINICAL AUDIT LOGS. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE PHILIPS FIELD SERVICE ENGINEER (FSE) WENT ONSITE AND ASSESSED THE WIRELESS SYSTEM. IT WAS DETERMINED THAT TWO ACCESS POINT (AP) CABLES WERE UNPLUGGED IN THE ER. A COVERAGE AREA VERIFICATION (CAV) FROM (B)(6) 2024 REFLECTS THAT THE WIRELESS SYSTEM WAS HOOKED UP CORRECTLY AND RUNNING AT THE TIME. IT IS UNKNOWN WHY THE ACCESS POINTS WERE LATER DISCONNECTED. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE ISSUE WAS THE DISCONNECTION OF THE ACCESS POINT CABLES. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE FSE PERFORMED A WIRELESS REMEDIATION IN WHICH THE ACCESS POINTS WERE RETURNED TO FACTORY DEFAULT AND RECONFIGURED. WHEN THE FSE WAS ONSITE ON MAY 13, 2025, HE VERIFIED THAT ALL AP¿S WERE ONLINE, AND NO ERRORS WERE DETECTED BY THE UPGRADER TOOL. TO ENSURE PROPER WIRELESS COVERAGE FROM THE INTELLIVUE X3 MONITOR IN THE ER, THE FSE ALSO PERFORMED A CAV CHECK WITH ANOTHER WIRELESS DEVICE (AN MX40).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22153930&pc=GXQ | 2249852-2025-00007 | 2025-04-17 04:00 | Death | COLLAGEN MATRIX, INC. | 2025-06-05 04:00 | GXQ | DURAMATRIX-ONLAY COLLAGEN DURA SUBSTITUTE MEMBRANE | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Low Blood Pressure/ Hypotension; Septic Shock | K061487 |  | 1 | Event Description: LIMITED INFORMATION PROVIDED. CLINICAL DATA COLLECTION SITE REPORTED A "DM/DMO" PRODUCT WAS USED ON A PATIENT FOR A MASS RESECTION ON (B)(6) 2025. NO PRODUCT ITEM NUMBER, PRODUCT LOT NUMBER, OR SPECIFIC PRODUCT NAME WAS PROVIDED. MEDICAL INTERVENTION REQUIRED AND A DECOMPRESSIVE HEMICRANIECTOMY WAS PERFORMED ON (B)(6) 2025. VARIOUS HOSPITAL COMPLICATIONS WERE REPORTED (SEE ATTACHED DOCUMENT FOR ALL REPORTED PATIENT INFORMATION). PATIENT EXPERIENCED ACUTE SEVERE NEUROLOGIC DECLINE ON (B)(6) 2025 AND BECAME HYPOTENSIVE, LIKELY FROM SEPTIC SHOCK. PATIENT WAS TERMINALLY EXTUBATED AND PASSED SHORTLY THEREAFTER ON (B)(6) 2025. IT WAS NOTED THAT "RELATIONSHIP OF THE EVENT" (EVENT OF SEPTIC SHOCK) TO THE INTERVENTION IS UNRELATED. CLINICIAN DOES NOT BELIEVE THERE IS CORRELATION BETWEEN THE "DM/DMO" PRODUCT AND THE SEPTIC SHOCK. NO ADDITIONAL EVENT, PRODUCT, OR PATIENT INFORMATION PROVIDED.Event Description: LIMITED INFORMATION PROVIDED. COMPLAINT WAS SUBMITTED THROUGH CLINICAL DATA COLLECTION. THE CLINICAL DATA COLLECTION SITE REPORTED A DURAMATRIX ONLAY PRODUCT WAS USED ON A PATIENT FOR A MASS RESECTION ON (B)(6) 2025. MEDICAL INTERVENTION REQUIRED AND A DECOMPRESSIVE HEMICRANIOMOTOMY WAS PERFORMED ON (B)(6) 2025. VARIOUS HOSPITAL COMPLICATIONS WERE REPORTED (SEE ATTACHED DOCUMENT FOR ALL REPORTED PATIENT INFORMATION). PATIENT EXPERIENCED ACUTE SEVERE NEUROLOGIC DECLINE ON (B)(6) 2025 AND BECAME HYPOTENSIVE, LIKELY FROM SEPTIC SHOCK. PATIENT WAS TERMINALLY EXTUBATED AND PASSED SHORTLY THEREAFTER ON (B)(6) 2025. IT WAS NOTED THAT "RELATIONSHIP OF THE EVENT" (EVENT OF SEPTIC SHOCK) TO THE INTERVENTION IS UNRELATED. CLINICIAN DOES NOT BELIEVE THERE IS CORRELATION BETWEEN THE PRODUCT AND THE SEPTIC SHOCK/PATIENT DEATH. NO ADDITIONAL EVENT OR PATIENT INFORMATION PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22153994&pc=OZD | 1220648-2025-29052 | 2025-05-16 04:00 | Death | ABIOMED, INC. | 2025-06-05 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD AN IMPELLA CP PUMP PLACED TO SUPPORT A PATIENT ADMITTED IN WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK. THE SUPPORT WAS ONGOING WHEN THE LIMB BECAME ISCHEMIC AND THE TEAM PLACED AN ANTEGRADE SHEATH FOR PERFUSION. THE LIMB DID NOT IMPROVE AND AND THE PATIENT EXPIRED. IT WAS NOTED THAT THERE WERE NO PROBLEMS WITH IMPELLA OPERATION THROUGHOUT MANAGEMENT. HOWEVER, THE LOWER LIMB COMPARTMENT SYNDROME WAS LIKELY CAUSED BY THE IMPELLA. IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS MULTIPLE ORGAN FAILURE DUE TO LOWER LIMB COMPARTMENT SYNDROME. THEREFORE, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA USED AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22154110&pc=MIH | 3011063223-2025-00104 |  | Death | ENDOLOGIX | 2025-06-05 04:00 | MIH | AFX | Insufficient Information | Aneurysm | P040002 |  | 1 | Event Description: THE REPORTED INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH REVIEW OF AN INDUSTRY RELEVANT JOURNAL ARTICLE. DOI.ORG/10.1016/J.AVSG.2024.04.027. SAFETY EVALUATION OF UNIBODY ENDOGRAFTS FOR ABDOMINAL AORTIC ANEURYSM REPAIR: A SYSTEMATIC REVIEW AND META-ANALYSIS SAFIA ESSAM, MIRNA HUSSEIN, ABDELRAHMAN MAMDOUH AHMED, LUJAINA AHMED, HAMED GABER, HASSAN EL-MASRY, RODINA MOHAMED ABDELAAL, NOURHAN GALAL, AHMED KASSEM, AND WAEL SHAALAN, ALEXANDRIA, EGYPT. BACKGROUND: THE UNIBODY BIFURCATED AORTIC ENDOGRAFT (AFX/AFX2) HAS EMERGED AS A TREATMENT OPTION FOR ABDOMINAL AORTIC ANEURYSMS (AAAS). THIS SYSTEMATIC REVIEW AND META-ANALYSIS AIMED TO EVALUATE THE SAFETY OF THE UNIBODY ENDOGRAFT. METHODS: A LITERATURE SEARCH WAS CONDUCTED IN COCHRANE LIBRARY, SCOPUS, WEB OF SCIENCE, AND PUBMED. STUDIES ASSESSING THE UNIBODY ENDOGRAFT FOR AAA REPAIR BETWEEN 2014 AND 2023 WERE INCLUDED. THE DEFINED PRIMARY OUTCOMES WERE THE INCIDENCES OF TYPE I, II, AND III ENDOLEAKS. THE SECONDARY OUTCOMES WERE ACCESS SITE PROBLEMS, ANEURYSM-RELATED MORTALITY, ANEURYSM RUPTURE, ALL-CAUSE MORTALITY, ANEURYSM SAC GROWTH, LIMB OCCLUSION, STENT GRAFT MIGRATION, AND TECHNICAL SUCCESS RATE. RESULTS: FOURTEEN STUDIES INCLUDING 12 OBSERVATIONAL STUDIES AND 2 RANDOMIZED CONTROLLED TRIALS WERE INCLUDED IN THE SYSTEMATIC REVIEW. THE META-ANALYSIS INCLUDED 10 STUDIES WITH 12,690 PATIENTS THAT REPORTED THE MEASURED OUTCOMES AND EXCLUDED 4 STUDIES THAT DID NOT. TYPE II ENDOLEAKS HAD THE HIGHEST INCIDENCE OF 12% (95% CONFIDENCE INTERVAL [CI]: 4E20%), FOLLOWED BY TYPE III ENDOLEAKS WITH AN INCIDENCE OF 3% (95% CI: 1E5%). THE INCIDENCE OF TYPE I ENDOLEAKS WAS 1% (95% CI: 0E2%). A SUBGROUP ANALYSIS BY FOLLOW-UP DURATION SHOWED THAT TYPE II ENDOLEAK INCIDENCE WAS HIGHER AFTER 1 TO 2 YEARS OF FOLLOW-UP THAN 3 TO 4 YEARS OF FOLLOW-UP. THE INCIDENCE OF ANEURYSMAL MORTALITY WAS 2% (95% CI: 0E7%), LIMB OCCLUSION WAS 1% (95% CI: 0E1%), STENT GRAFT MIGRATION WAS 1% (95% CI: 0E2%), ANEURYSMAL RUPTURE WAS 6% (95% CI: 2E11%), ACCESS SITE PROBLEMS WERE 7% (95% CI: 2E13%), ANEURYSM SAC GROWTH WAS 2% (95% CI: 0E4%), ALL-CAUSE MORTALITY WAS 21% (95% CI: 4E38%), AND TECHNICAL SUCCESS RATE WAS 100% (95% CI: 98E 100%). CONCLUSIONS: THE UNIBODY ENDOGRAFT IS A SAFE AND MINIMALLY INVASIVE APPROACH FOR AAA REPAIR. HOWEVER, POTENTIAL COMPLICATIONS NECESSITATE CLOSE PATIENT FOLLOW-UP AFTER THE INTERVENTION. NO ADDITIONAL INFORMATION WILL BECOME AVAILABLE REGARDING THIS INITIAL/FINAL REPORT DUE TO THE ABSENCE OF PATIENT IDENTIFIERS. THIS REPORT ADDRESSES PATIENT 3 OF 3. Manufacturer Narrative: THE REPORTED ADVERSE EVENT/INCIDENT WAS INVESTIGATED IN ALIGNMENT WITH ENDOLOGIX OPERATING PROCEDURES AND WORK INSTRUCTIONS. THE REPORTED ADVERSE EVENT/INCIDENT WAS IDENTIFIED BY ENDOLOGIX THROUGH AN ONGOING REVIEW OF INDUSTRY RELEVANT JOURNAL ARTICLES WHERE PATIENT RELATED OUTCOMES ARE PRESENTED ANONYMOUSLY, AND PATIENT SPECIFIC DEVICE INFORMATION IS UNAVAILABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE COMPLETED. THE PART NUMBER/LOT NUMBER COMBINATION NEEDED FOR DEVICE IDENTIFICATION REMAINS UNKNOWN. AN EVALUATION OF THE DEVICE COULD NOT BE COMPLETED. THE DEVICE WAS NOT RETURNED TO ENDOLOGIX FOR EVALUATION BECAUSE IT REMAINS IMPLANTED. A CLINICAL EVALUATION OF THE ADVERSE EVENT/INCIDENT COULD NOT BE COMPLETED AS THE REPORTED EVENT WAS FOUND DURING PUBLICATION REVIEW WHERE THE PATIENT INFORMATION IS ANONYMOUS. DEVICE, USE, PROCEDURE, AND/OR ANATOMY RELATEDNESS TO THIS ADVERSE EVENT/INCIDENT COULD NOT BE EVALUATED. NO ADDITIONAL INVESTIGATION OF THIS REPORTED ADVERSE EVENT/INCIDENT IS PLANNED. HOWEVER, SHOULD ADDITIONAL INFORMATION RELEVANT TO THE INVESTIGATION OUTCOME BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. ENDOLOGIX WILL CONTINUE TO MONITOR THIS AND SIMILAR ADVERSE EVENTS/INCIDENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22154175&pc=MLV | 2017233-2025-06279 | 2025-04-16 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-05 04:00 | MLV | GORE® CARDIOFORM SEPTAL OCCLUDER | Device Appears to Trigger Rejection | Thrombosis/Thrombus | P050006 |  | 1 | Event Description: IT WAS REPORTED TO GORE THAT THIS PATIENT EXPERIENCED AN EMBOLIC LACUNAR CEREBRAL INFARCTION IN (B)(6) 2024 AND FURTHER EXAMINATION FOUND A PATENT FORAMEN OVALE (PFO). ON (B)(6) 2024, THE DEFECT WAS THEREFORE TREATED WITH A 25 MM GORE® CARDIOFORM SEPTAL OCCLUDER. IN (B)(6) 2024, THE PATENT DEVELOPED ATRIAL FIBRILLATION AND SPONTANEOUSLY CONVERTED TO SINUS RHYTHM. REPORTEDLY, THE PHYSICIAN ATTRIBUTED THE ARRHYTHMIA MOST LIKELY TO THE OCCLUDER DEVICE. FOLLOW-UP TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) IMAGING ON (B)(6) 2024 REVEALED NO DEVICE ISSUES. THE PATIENT PROPHYLACTICALLY RECEIVED ANTICOAGUALANT MEDICATION (APIXABAN) FOLLOWING AN CHA2DS2-VASC ATRIAL FIBRILLATION STROKE RISK ASSESSMENT ANALYSIS WITH A SCORE OF 5 POINTS. HOWEVER, THE PATIENT WAS REPORTED TO HAVE SUFFERED REPEAT TRANSITORY ISCHEMIC ATTACKS IN THE RIGHT MEDIAL VASCULAR AREA DURING (B)(6) 2024. DESPITE NEW ORAL ANTICOAGULANT (NOAC) THERAPY, TEE IMAGING IDENTIFIED THROMBOTIC DEPOSITS IN THE SURROUNDINGS OF THE PFO OCCLUDER, THUS MARCUMAR TREATMENT WAS STARTED IN (B)(6) 2024. ADDITIONAL TEE IMAGING ON (B)(6) 2024 CONFIRMED PERSISTING DEPOSITS ON THE DEVICE MEASURING 8 MM MAXIMUM IN LENGTH AND FLOATING TOWARDS THE LEFT ATRIUM. ON (B)(6) 2024, FOLLOW-UP IMAGING REPORTEDLY SHOWED A CLEARLY RECOGNIZABLE STRUCTURE WITH AN INCREASED MOBILITY MEASURING APPROXIMATELY 4 X 5 MM AND FLOATING TOWARDS THE ATRIOVENTRICULAR VALVES. IT WAS ALSO REPORTED THAT FOLLOW-UP IMAGING PERFORMED ON (B)(6) 2024 NO LONGER DIAGNOSED FLOATING STRUCTURES ON THE OCCLUDER DEVICE. HOWEVER, DURING FOLLOW-UP IMAGING ON (B)(6) 2025, ANOTHER LARGE FLOATING THROMBUS WAS AGAIN SEEN ON THE LEFT DISC MEASURING 11 X 7 MM AND THE DECISION WAS MADE TO REFER THE PATIENT FOR SURGICAL DEVICE REMOVAL. ON (B)(6) 2025, THE DEVICE WAS EXPLANTED. HOWEVER, THE PATIENT SUBSEQUENTLY DEVELOPED MULTI-ORGAN FAILURE AND WAS REPORTED TO HAVE EXPIRED THEREAFTER. Manufacturer Narrative: A1, PATIENT IDENTIFIER (IN CONFIDENCE): A PATIENT IDENTIFIER WAS NOT PROVIDED. DATA PROVIDED IN THIS FIELD REFLECT GORE'S INTERNAL REFERENCE NUMBER FOR THIS CASE. THE REVIEW OF THE MANUFACTURING RECORDS VERIFIED THAT THE LOT INVOLVED IN THIS EVENT MET ALL PRE-RELEASE SPECIFICATIONS. THE DEVICE WAS REPORTED EXPLANTED. THE CURRENT LOCATION OF THE DEVICE IS UNKNOWN. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22154213&pc=NVN | 2017865-2025-75465 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-05 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED DUE TO ACUTE CARDIOPULMONARY ARREST. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22154263&pc=DRA | 2124215-2025-36939 | 2025-05-09 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-05 04:00 | DRA | FARADRIVE | Adverse Event Without Identified Device or Use Problem | Air Embolism; Non specific EKG/ECG Changes; Low Blood Pressure/ Hypotension; Myocardial Infarction; Asystole |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT EXPERIENCED HYPOTENSION, ST SEGMENT ELEVATION, RIGHT CORONARY ARTERY INFARCTION AND BECAME ASYSTOLE. DURING A PULSED FIELD ABLATION (PFA) PROCEDURE TO TREAT A PAROXYSMAL ATRIAL FIBRILLATION A FARADRIVE WAS USED. TRANSEPTAL WAS OBTAINED USING A VERSACROSS CONNECT WITH THE FARADRIVE, PRE-MAPPING WAS DONE AND THEN THE CATHETER WAS REPLACED WITH THE FARAWAVE CATHETER FOR ABLATION. 8 APPLICATIONS WERE GIVEN AND THEN IT WAS NOTICED THAT THE PATIENT'S BLOOD PRESSURE DROPPED SIGNIFICANTLY. INTRACARDIAC ECHOCARDIOGRAPHY (ICE) WAS CHECKED FOR EFFUSION, BUT NOTHING WAS FOUND. IT WAS OBSERVED ON 12-LEAD A ST SEGMENT ELEVATION AND THE PATIENT BECAME ASYSTOLE. COMPRESSIONS WERE GIVEN TO THE PATIENT, AN ANGIOGRAM WAS COMPLETED AND A RIGHT CORONARY ARTERY INFARCT WAS OBSERVED DUE TO BUBBLES. PHYSICIAN COMMENTED THAT PRESSURE BAD CONNECTED TO THE SHEATH WAS EMPTY. THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AND HAD AN ADDITIONAL DEVICE IMPLANTED. THE PATIENT WAS TAKEN TO HAVE A COMPUTED TOMOGRAPHY (CT) SCAN. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22154429&pc=POL | 3012931345-2025-00100 | 2023-12-30 05:00 | Death | STRYKER NEUROVASCULAR-UTAH-SALT LAKE CITY | 2025-06-05 04:00 | POL | UNKNOWN_NEUROVASCULAR_PRODUCT | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K223305 |  | 1 | Event Description: IN DISCOUNT (EVALUATION OF MECHANICAL THROMBECTOMY IN ACUTE ISCHEMIC STROKE RELATED TO A DISTAL ARTERIAL OCCLUSION) STUDY, THERE WAS HEMORRHAGIC TRANSFORMATION WITH WORSENING OF CLINICAL CONDITION IN A CONTEXT OF FAILED RECANALIZATION ON (B)(6) 2023. PATIENT DIED ON (B)(6) 2024. NO ADDITIONAL INFORMATION IS AVAILABLE. THE CAUSE OF DEATH IS UNKNOWN AS THE PATIENT WAS LOST OF VIEW.Event Description: IN DISCOUNT (EVALUATION OF MECHANICAL THROMBECTOMY IN ACUTE ISCHEMIC STROKE RELATED TO A DISTAL ARTERIAL OCCLUSION) STUDY, THERE WAS HEMORRHAGIC TRANSFORMATION WITH WORSENING OF CLINICAL CONDITION IN A CONTEXT OF FAILED RECANALIZATION ON (B)(6) 2023. PATIENT DIED ON (B)(6) 2024. NO ADDITIONAL INFORMATION IS AVAILABLE. THE CAUSE OF DEATH IS UNKNOWN AS THE PATIENT WAS LOST OF VIEW. Manufacturer Narrative: ALTHOUGH THE LOT# IS UNKNOWN THERE ARE CONTROLS IN THE MANUFACTURING PROCESS TO ENSURE THE PRODUCT MET SPECIFICATIONS UPON RELEASE. THE SUBJECT DEVICE IS NOT AVAILABLE; THEREFORE, VISUAL TESTING AS WELL AS FUNCTIONAL TESTING CANNOT BE PERFORMED. THE REPORTED EVENT IS COVERED IN THE DEVICE DIRECTIONS FOR USE (DFU). AS WELL, THE RISK OF THE REPORTED EVENT IS DOCUMENTED IN THE RISK DOCUMENTATION AND THERE ARE CURRENT CONTROLS TO MITIGATE THE RISK OF THE AS REPORTED EVENT. THE REPORTED COMPLAINT COULD NOT BE CONFIRMED, AND IT COULD NOT BE DEFINITIVELY DETERMINED IF THE DEVICE FAILED TO MEET SPECIFICATIONS BECAUSE THE PRODUCT WAS NOT RETURNED. THIS CASE WAS A CLINICAL STUDY "HEMORRHAGIC TRANSFORMATION WITH WORSENING OF CLINICAL CONDITION IN A CONTEXT OF FAILED RECANALIZATION." AS PER THE TREVO NXT IFU, PATIENT DEATH AND INTRACRANIAL HEMORRHAGE ARE DOCUMENTED AS ANTICIPATED PROCEDURAL COMPLICATIONS. AN ASSIGNABLE CAUSE OF ANTICIPATED PROCEDURAL COMPLICATION WILL BE ASSIGNED TO THE REPORTED EVENTS OF 'PATIENT DEATH' AND 'PATIENT INTRACRANIAL HEMORRHAGE', AS A PRODUCT RELATED ROOT CAUSE DOES NOT APPLY AND THE ISSUE IS DUE TO A KNOWN PHYSIOLOGICAL EFFECT OF THE PROCEDURE AND/OR PATIENT CONDITION NOTED WITH THE DIRECTIONS FOR USE, PRODUCT LABELING AND/OR RISK DOCUMENTATION FILES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22154830&pc=EZB | 2124215-2025-36913 | 2022-08-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-05 04:00 | EZB | AMPLATZ SUPER STIFF | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hematoma; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Pleural Effusion; Renal Failure; Prolapse; Multiple Organ Failure; Thrombosis/Thrombus; Asystole; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC VIA AN ARTICLE PUBLISHED IN PEDIATRIC MEDICINE THAT A STUDY WAS CONDUCTED TO REPORT A CASE OF A PEDIATRIC PATIENT WHO UNDERWENT PERCUTANEOUS NEPHROLITHOTOMY (PCNL). DURING THIS PROCEDURE, AN AMPLATZ SUPER STIFF GUIDEWIRE AND AMPLATZ ACCESS SHEATH WERE USED. IT WAS REPORTED THAT THE PROCEDURE WAS CHALLENGING SINCE THE BEGINNING DUE TO THE COMPLEX ANATOMY OF THE PATIENT, AND THE PROCEDURE TOOK LONGER THAN EXPECTED, AS IT LASTED ABOUT 5 HOURS INSTEAD OF THE TYPICAL 30 MINUTES OF PROCEDURE. THE PATIENT DEVELOPED ABDOMINAL COMPARTMENT SYNDROME (ACS) WITHOUT ANY WARNING SIGN. UPON COMPLETION AND REMOVAL OF DRAPES, THE PATIENT WAS TURNED SUPINE, AND IT WAS FOUND THAT THERE WAS A DISTENDED ABDOMEN, RECTAL PROLAPSE AND THE LEGS WERE MOTTLED AND POORLY PERFUSED. AN URGENT ULTRASOUND WAS REQUIRED, AND IT WAS OBSERVED THAT THERE WAS LARGE VOLUME RETROPERITONEAL COLLECTION, WHICH WAS PUSHING THE BLADDER AND BOWELS ANTERIORLY. SEPTATED FLUID WAS ALSO SEEN AROUND THE KIDNEYS, AND MORISON'S PUCH. ADDITIONALLY, THERE WERE BILATERAL PLEURAL EFFUSIONS; THE PATIENT BECAME HYPOTENSIVE AND REQUIRED FLUID RESUSCITATION AND PHENYLEPHRINE BOLUSES. THE PATIENT WAS TRANSFERRED TO THE PEDIATRIC INTENSIVE CARE UNIT (PICU) WHERE TREATMENT CONTINUED AND WHERE THERE WAS ONGOING BLOOD LOSS FROM THE ABDOMEN. FOLLOWING READMISSION TO PICU, IT WAS REVEALED CLOTS IN THE LEFT PARASPINAL REGION/RENAL BED, AND CONSISTENT BLEEDING, MAKING THE EXACT BLEEDING POINTS UNIDENTIFIABLE. THERE WAS A DEVELOPMENT OF SYSTEMIC INFLAMMATORY RESPONSE SYNDROME (SIRS); THE PATIENT ALSO EXPERIENCED MULTIORGAN FAILURE AND ACUTE KIDNEY INJURY. ANTIBIOTIC THERAPY WAS ALSO GIVEN TO THE PATIENT. THE HYPOTENSION CONTINUED DESPITE TREATMENT, AND IT WAS ALSO REVEALED HYPERDYNAMIC LEFT VENTRICULAR FUNCTION WITH MILDLY IMPAIRED RIGHT VENTRICULAR FILLING AND FUNCTION. THERE WAS NO RELENTLESS PROGRESSION OF THE SHOCK DESPITE MAXIMAL MEDICAL THERAPY. THE PATIENT EXPERIENCED INCREASING PERIODS OF BRADYCARDIA AND INTERMITTENT ASYSTOLE WITH WORSENING HYPOTENSION. PRIOR RETURNING TO PICU, FIXED DILATED PUPILS WERE FOUND IN THE PATIENT, THE HAEMODYNAMIC FINDINGS AND NEUROLOGICAL DETERIORATION WERE CONVEYED TO THE FAMILY, AND THE PATIENT ULTIMATELY PASSED AWAY.Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC VIA AN ARTICLE PUBLISHED IN PEDIATRIC MEDICINE THAT A STUDY WAS CONDUCTED TO REPORT A CASE OF A PEDIATRIC PATIENT WHO UNDERWENT PERCUTANEOUS NEPHROLITHOTOMY (PCNL). DURING THIS PROCEDURE, AN AMPLATZ SUPER STIFF GUIDEWIRE AND AMPLATZ ACCESS SHEATH WERE USED. IT WAS REPORTED THAT THE PROCEDURE WAS CHALLENGING SINCE THE BEGINNING DUE TO THE COMPLEX ANATOMY OF THE PATIENT, AND THE PROCEDURE TOOK LONGER THAN EXPECTED, AS IT LASTED ABOUT 5 HOURS INSTEAD OF THE TYPICAL 30 MINUTES OF PROCEDURE. THE PATIENT DEVELOPED ABDOMINAL COMPARTMENT SYNDROME (ACS) WITHOUT ANY WARNING SIGN. UPON COMPLETION AND REMOVAL OF DRAPES, THE PATIENT WAS TURNED SUPINE, AND IT WAS FOUND THAT THERE WAS A DISTENDED ABDOMEN, RECTAL PROLAPSE AND THE LEGS WERE MOTTLED AND POORLY PERFUSED. AN URGENT ULTRASOUND WAS REQUIRED, AND IT WAS OBSERVED THAT THERE WAS LARGE VOLUME RETROPERITONEAL COLLECTION, WHICH WAS PUSHING THE BLADDER AND BOWELS ANTERIORLY. SEPTATED FLUID WAS ALSO SEEN AROUND THE KIDNEYS, AND MOSRISON'S PUCH. ADDITIONALLY, THERE WERE BILATERAL PLEURAL EFFUSIONS; THE PATIENT BECAME HYPOTENSIVE AND REQUIRED FLUID RESUSCITATION AND PHENYLEPHRINE BOLUSES. THE PATIENT WAS TRANSFERRED TO THE PEDIATRIC INTENSIVE CARE UNIT (PICU) WHERE TREATMENT CONTINUED AND WHERE THERE WAS ONGOING BLOOD LOSS FROM THE ABDOMEN. FOLLOWING READMISSION TO PICU, IT WAS REVEALED CLOTS IN THE LEFT PARASPINAL REGION/RENAL BED, AND CONSISTENCE BLEEDING, MAKING THE EXACT BLEEDING POINTS UNIDENTIFIABLE. THERE WAS A DEVELOPMENT OF SYSTEMIC INFLAMMATORY RESPONSE SYNDROME (SIRS); THE PATIENT ALSO EXPERIENCED MULTIORGAN FAILURE AND ACUTE KIDNEY INJURY. ANTIBIOTIC THERAPY WAS ALSO GIVEN TO THE PATIENT. THE HYPOTENSION CONTINUED DESPITE TREATMENT, AND IT WAS ALSO REVEALED HYPERDYNAMIC LEFT VENTRICULAR FUNCTION WITH MILDLY IMPAIRED RIGHT VENTRICULAR FILLING AND FUNCTION. THERE WAS NO RELENTLESS PROGRESSION OF THE SHOCK DESPITE MAXIMAL MEDICAL THERAPY. THE PATIENT EXPERIENCED INCREASING PERIODS OF BRADYCARDIA AND INTERMITTENT ASYSTOLE WITH WORSENING HYPOTENSION. PRIOR RETURNING TO PICU, FIXED DILATED PUPILS WERE FOUND IN THE PATIENT, THE HAEMODYNAMIC FINDINGS AND NEUROLOGICAL DETERIORATION WERE CONVEYED TO THE FAMILY, AND THE PATIENT ULTIMATELY PASSED AWAY. THE PHYSICIAN ASSESSED THAT NEITHER THE SHEATH NOR THE WIRE CONTRIBUTED TO THE PATIENT'S DEATH. THE CAUSE OF DEATH WAS ATTRIBUTED TO THE PATIENT'S UNDERLYING CLINICAL CONDITION AND THE DEVELOPMENT OF ABDOMINAL COMPARTMENT SYNDROME. Manufacturer Narrative: BLOCK B2: DATE OF DEATH AND BLOCK B3: DATE OF EVENT. DATES ARE APPROXIMATED. BLOCK D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK E1: INITIAL REPORTER FACILITY NAME: (B)(6) HOSPITAL. BLOCK G2: LITERATURE SOURCE: ASHRAF, J., DEVALL, M., DOORN, V. D., KUMAR, R. (2022). ABDOMINAL COMPARTMENT SYNDROME AS A COMPLICATION OF PERCUTANEOUS NEPHROLITHOTOMY REQUIRING EXTRA-CORPOREAL LIFE SUPPORT, A CASE REPORT: PEDIATR MED 2022;5:32, HTTPS://DX.DOI.ORG/10.21037/PM-21-94. BLOCK H6: IMDRF PATIENT CODE E2321 CAPTURES THE REPORTABLE EVENT OF LOW BLOOD PRESSURE/HYPOTENSION. IMDRF PATIENT CODE E0506 CAPTURES THE REPORTABLE EVENT OF HEMORRHAGE/BLOOD LOSS/BLEEDING. IMDRF PATIENT CODE E0514 CAPTURES THE REPORTABLE EVENT OF THROMBOSIS/THROMBUS. IMDRF PATIENT CODE E0731 CAPTURES THE REPORTABLE EVENT OF PLEURAL EFFUSION. IMDRF PATIENT CODE E130501 CAPTURES THE REPORTABLE EVENT OF RENAL FAILURE. IMDRF PATIENT CODE E0602 CAPTURES THE REPORTABLE EVENT OF CARDIAC ARREST. IMDRF PATIENT CODE 2342 CAPTURES THE REPORTABLE EVENT OF MULTIPLE ORGAN DYSFUNCTION SYNDROME. IMDRF PATIENT CODE E060101 CAPTURES THE REPORTABLE EVENT OF ASYSTOLE. IMDRF PATIENT CODE E2402 CAPTURES THE REPORTABLE EVENT OF DISTENTION. IMDRF PATIENT CODE E2333 CAPTURES THE REPORTABLE EVENT OF PROLAPSE. IMDRF PATIENT CODE E0519 CAPTURES THE REPORTABLE EVENT OF REDUCED BLOOD FLOW. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF DEATH. IMDRF IMPACT CODE F0801CAPTURES THE REPORTABLE EVENT OF INTENSIVE CARE. IMDRF IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF UNEXPECTED MEDICAL INTERVENTION. IMDRF IMPACT CODE F08 CAPTURES THE REPORTABLE EVENT OF HOSPITALIZATION OR PROLONGED HOSPITALIZATION. IMDRF IMPACT CODE F2303 CAPTURES THE REPORTABLE EVENT OF MEDICATION REQUIRED. Manufacturer Narrative: BLOCK B2: DATE OF DEATH AND BLOCK B3: DATE OF EVENT DATES ARE APPROXIMATED. BLOCK D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK E1: INITIAL REPORTER FACILITY NAME: (B)(6) HOSPITAL. BLOCK G2: LITERATURE SOURCE: ASHRAF, J., DEVALL, M., DOORN, V. D., KUMAR, R. (2022). ABDOMINAL COMPARTMENT SYNDROME AS A COMPLICATION OF PERCUTANEOUS NEPHROLITHOTOMY REQUIRING EXTRA-CORPOREAL LIFE SUPPORT, A CASE REPORT: PEDIATR MED 2022;5:32, HTTPS://DX.DOI.ORG/10.21037/PM-21-94. BLOCK H6: IMDRF PATIENT CODE E2321 CAPTURES THE REPORTABLE EVENT OF LOW BLOOD PRESSURE/HYPOTENSION. IMDRF PATIENT CODE E0506 CAPTURES THE REPORTABLE EVENT OF HEMORRHAGE/BLOOD LOSS/BLEEDING. IMDRF PATIENT CODE E0514 CAPTURES THE REPORTABLE EVENT OF THROMBOSIS/THROMBUS. IMDRF PATIENT CODE E0731 CAPTURES THE REPORTABLE EVENT OF PLEURAL EFFUSION. IMDRF PATIENT CODE E130501 CAPTURES THE REPORTABLE EVENT OF RENAL FAILURE. IMDRF PATIENT CODE E0602 CAPTURES THE REPORTABLE EVENT OF CARDIAC ARREST IMDRF PATIENT CODE 2342CAPTURES THE REPORTABLE EVENT OF MULTIPLE ORGAN DYSFUNCTION SYNDROME. IMDRF PATIENT CODE E060101 CAPTURES THE REPORTABLE EVENT OF ASYSTOLE. IMDRF PATIENT CODE E2402 CAPTURES THE REPORTABLE EVENT OF DISTENTION. IMDRF PATIENT CODE E2333 CAPTURES THE REPORTABLE EVENT OF PROLAPSE. IMDRF PATIENT CODE E0519 CAPTURES THE REPORTABLE EVENT OF REDUCED BLOOD FLOW. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF DEATH. IMDRF IMPACT CODE F0801CAPTURES THE REPORTABLE EVENT OF INTENSIVE CARE. IMDRF IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF UNEXPECTED MEDICAL INTERVENTION. IMDRF IMPACT CODE F08 CAPTURES THE REPORTABLE EVENT OF HOSPITALIZATION OR PROLONGED HOSPITALIZATION. IMDRF IMPACT CODE F2303 CAPTURES THE REPORTABLE EVENT OF MEDICATION REQUIRED. BLOCK H11: ADDITIONAL INFORMATION UPDATED BLOCK B5: DESCRIBE EVENT OR PROBLEM. CORRECTION TO FIELD BLOCK H6: EVALUATION CONCLUSION CODES 1.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22154851&pc=GCC | 2124215-2025-36914 | 2022-08-01 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-05 04:00 | GCC | NEPHROSTOMY SHEATHS AND DILATORS | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hematoma; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Pleural Effusion; Renal Failure; Prolapse; Multiple Organ Failure; Thrombosis/Thrombus; Asystole; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC VIA AN ARTICLE PUBLISHED IN PEDIATRIC MEDICINE THAT A STUDY WAS CONDUCTED TO REPORT A CASE OF A PEDIATRIC PATIENT WHO UNDERWENT PERCUTANEOUS NEPHROLITHOTOMY (PCNL). DURING THIS PROCEDURE, AN AMPLATZ SUPER STIFF GUIDEWIRE AND AMPLATZ ACCESS SHEATH WERE USED. IT WAS REPORTED THAT THE PROCEDURE WAS CHALLENGING SINCE THE BEGINNING DUE TO THE COMPLEX ANATOMY OF THE PATIENT, AND THE PROCEDURE TOOK LONGER THAN EXPECTED, AS IT LASTED ABOUT 5 HOURS INSTEAD OF THE TYPICAL 30 MINUTES OF PROCEDURE. THE PATIENT DEVELOPED ABDOMINAL COMPARTMENT SYNDROME (ACS) WITHOUT ANY WARNING SIGN. UPON COMPLETION AND REMOVAL OF DRAPES, THE PATIENT WAS TURNED SUPINE, AND IT WAS FOUND THAT THERE WAS A DISTENDED ABDOMEN, RECTAL PROLAPSE AND THE LEGS WERE MOTTLED AND POORLY PERFUSED. AN URGENT ULTRASOUND WAS REQUIRED, AND IT WAS OBSERVED THAT THERE WAS LARGE VOLUME RETROPERITONEAL COLLECTION, WHICH WAS PUSHING THE BLADDER AND BOWELS ANTERIORLY. SEPTATED FLUID WAS ALSO SEEN AROUND THE KIDNEYS, AND MORISON'S PUCH. ADDITIONALLY, THERE WERE BILATERAL PLEURAL EFFUSIONS; THE PATIENT BECAME HYPOTENSIVE AND REQUIRED FLUID RESUSCITATION AND PHENYLEPHRINE BOLUSES. THE PATIENT WAS TRANSFERRED TO THE PEDIATRIC INTENSIVE CARE UNIT (PICU) WHERE TREATMENT CONTINUED AND WHERE THERE WAS ONGOING BLOOD LOSS FROM THE ABDOMEN. FOLLOWING READMISSION TO PICU, IT WAS REVEALED CLOTS IN THE LEFT PARASPINAL REGION/RENAL BED, AND CONSISTENCE BLEEDING, MAKING THE EXACT BLEEDING POINTS UNIDENTIFIABLE. THERE WAS A DEVELOPMENT OF SYSTEMIC INFLAMMATORY RESPONSE SYNDROME (SIRS); THE PATIENT ALSO EXPERIENCED MULTIORGAN FAILURE AND ACUTE KIDNEY INJURY. ANTIBIOTIC THERAPY WAS ALSO GIVEN TO THE PATIENT. THE HYPOTENSION CONTINUED DESPITE TREATMENT, AND IT WAS ALSO REVEALED HYPERDYNAMIC LEFT VENTRICULAR FUNCTION WITH MILDLY IMPAIRED RIGHT VENTRICULAR FILLING AND FUNCTION. THERE WAS NO RELENTLESS PROGRESSION OF THE SHOCK DESPITE MAXIMAL MEDICAL THERAPY. THE PATIENT EXPERIENCED INCREASING PERIODS OF BRADYCARDIA AND INTERMITTENT ASYSTOLE WITH WORSENING HYPOTENSION. PRIOR RETURNING TO PICU, FIXED DILATED PUPILS WERE FOUND IN THE PATIENT, THE HAEMODYNAMIC FINDINGS AND NEUROLOGICAL DETERIORATION WERE CONVEYED TO THE FAMILY, AND THE PATIENT ULTIMATELY PASSED AWAY.Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC VIA AN ARTICLE PUBLISHED IN PEDIATRIC MEDICINE THAT A STUDY WAS CONDUCTED TO REPORT A CASE OF A PEDIATRIC PATIENT WHO UNDERWENT PERCUTANEOUS NEPHROLITHOTOMY (PCNL). DURING THIS PROCEDURE, AN AMPLATZ SUPER STIFF GUIDEWIRE AND AMPLATZ ACCESS SHEATH WERE USED. IT WAS REPORTED THAT THE PROCEDURE WAS CHALLENGING SINCE THE BEGINNING DUE TO THE COMPLEX ANATOMY OF THE PATIENT, AND THE PROCEDURE TOOK LONGER THAN EXPECTED, AS IT LASTED ABOUT 5 HOURS INSTEAD OF THE TYPICAL 30 MINUTES OF PROCEDURE. THE PATIENT DEVELOPED ABDOMINAL COMPARTMENT SYNDROME (ACS) WITHOUT ANY WARNING SIGN. UPON COMPLETION AND REMOVAL OF DRAPES, THE PATIENT WAS TURNED SUPINE, AND IT WAS FOUND THAT THERE WAS A DISTENDED ABDOMEN, RECTAL PROLAPSE AND THE LEGS WERE MOTTLED AND POORLY PERFUSED. AN URGENT ULTRASOUND WAS REQUIRED, AND IT WAS OBSERVED THAT THERE WAS LARGE VOLUME RETROPERITONEAL COLLECTION, WHICH WAS PUSHING THE BLADDER AND BOWELS ANTERIORLY. SEPTATED FLUID WAS ALSO SEEN AROUND THE KIDNEYS, AND MORISON'S PUCH. ADDITIONALLY, THERE WERE BILATERAL PLEURAL EFFUSIONS; THE PATIENT BECAME HYPOTENSIVE AND REQUIRED FLUID RESUSCITATION AND PHENYLEPHRINE BOLUSES. THE PATIENT WAS TRANSFERRED TO THE PEDIATRIC INTENSIVE CARE UNIT (PICU) WHERE TREATMENT CONTINUED AND WHERE THERE WAS ONGOING BLOOD LOSS FROM THE ABDOMEN. FOLLOWING READMISSION TO PICU, IT WAS REVEALED CLOTS IN THE LEFT PARASPINAL REGION/RENAL BED, AND CONSISTENCE BLEEDING, MAKING THE EXACT BLEEDING POINTS UNIDENTIFIABLE. THERE WAS A DEVELOPMENT OF SYSTEMIC INFLAMMATORY RESPONSE SYNDROME (SIRS); THE PATIENT ALSO EXPERIENCED MULTIORGAN FAILURE AND ACUTE KIDNEY INJURY. ANTIBIOTIC THERAPY WAS ALSO GIVEN TO THE PATIENT. THE HYPOTENSION CONTINUED DESPITE TREATMENT, AND IT WAS ALSO REVEALED HYPERDYNAMIC LEFT VENTRICULAR FUNCTION WITH MILDLY IMPAIRED RIGHT VENTRICULAR FILLING AND FUNCTION. THERE WAS NO RELENTLESS PROGRESSION OF THE SHOCK DESPITE MAXIMAL MEDICAL THERAPY. THE PATIENT EXPERIENCED INCREASING PERIODS OF BRADYCARDIA AND INTERMITTENT ASYSTOLE WITH WORSENING HYPOTENSION. PRIOR RETURNING TO PICU, FIXED DILATED PUPILS WERE FOUND IN THE PATIENT, THE HAEMODYNAMIC FINDINGS AND NEUROLOGICAL DETERIORATION WERE CONVEYED TO THE FAMILY, AND THE PATIENT ULTIMATELY PASSED AWAY. THE PHYSICIAN ASSESSED THAT NEITHER THE SHEATH NOR THE WIRE CONTRIBUTED TO THE PATIENT'S DEATH. THE CAUSE OF DEATH WAS ATTRIBUTED TO THE PATIENT'S UNDERLYING CLINICAL CONDITION AND THE DEVELOPMENT OF ABDOMINAL COMPARTMENT SYNDROME. Manufacturer Narrative: BLOCK B2: DATE OF DEATH AND BLOCK B3: DATE OF EVENT: DATES ARE APPROXIMATED. BLOCK D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK E1: INITIAL REPORTER FACILITY NAME: (B)(6). BLOCK G2: LITERATURE SOURCE: ASHRAF, J., DEVALL, M., DOORN, V. D., KUMAR, R. (2022). ABDOMINAL COMPARTMENT SYNDROME AS A COMPLICATION OF PERCUTANEOUS NEPHROLITHOTOMY REQUIRING EXTRA-CORPOREAL LIFE SUPPORT, A CASE REPORT: PEDIATR MED 2022;5:32, HTTPS://DX.DOI.ORG/10.21037/PM-21-94 BLOCK H6: IMDRF PATIENT CODE E2321 CAPTURES THE REPORTABLE EVENT OF LOW BLOOD PRESSURE/HYPOTENSION. IMDRF PATIENT CODE E0506 CAPTURES THE REPORTABLE EVENT OF HEMORRHAGE/BLOOD LOSS/BLEEDING. IMDRF PATIENT CODE E0514 CAPTURES THE REPORTABLE EVENT OF THROMBOSIS/THROMBUS. IMDRF PATIENT CODE E0731 CAPTURES THE REPORTABLE EVENT OF PLEURAL EFFUSION. IMDRF PATIENT CODE E130501 CAPTURES THE REPORTABLE EVENT OF RENAL FAILURE. IMDRF PATIENT CODE E0602 CAPTURES THE REPORTABLE EVENT OF CARDIAC ARREST IMDRF PATIENT CODE 2342CAPTURES THE REPORTABLE EVENT OF MULTIPLE ORGAN DYSFUNCTION SYNDROME. IMDRF PATIENT CODE E060101 CAPTURES THE REPORTABLE EVENT OF ASYSTOLE. IMDRF PATIENT CODE E2402 CAPTURES THE REPORTABLE EVENT OF DISTENTION. IMDRF PATIENT CODE E2333 CAPTURES THE REPORTABLE EVENT OF PROLAPSE. IMDRF PATIENT CODE E0519 CAPTURES THE REPORTABLE EVENT OF REDUCED BLOOD FLOW. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF DEATH. IMDRF IMPACT CODE F0801CAPTURES THE REPORTABLE EVENT OF INTENSIVE CARE. IMDRF IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF UNEXPECTED MEDICAL INTERVENTION. IMDRF IMPACT CODE F08 CAPTURES THE REPORTABLE EVENT OF HOSPITALIZATION OR PROLONGED HOSPITALIZATION. IMDRF IMPACT CODE F2303 CAPTURES THE REPORTABLE EVENT OF MEDICATION REQUIRED. Manufacturer Narrative: BLOCK B2: DATE OF DEATH AND BLOCK B3: DATE OF EVENT: DATES ARE APPROXIMATED. BLOCK D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK E1: INITIAL REPORTER FACILITY NAME: (B)(6) HOSPITAL. BLOCK G2: LITERATURE SOURCE: ASHRAF, J., DEVALL, M., DOORN, V. D., KUMAR, R. (2022). ABDOMINAL COMPARTMENT SYNDROME AS A COMPLICATION OF PERCUTANEOUS NEPHROLITHOTOMY REQUIRING EXTRA-CORPOREAL LIFE SUPPORT, A CASE REPORT: PEDIATR MED 2022;5:32, HTTPS://DX.DOI.ORG/10.21037/PM-21-94. BLOCK H6: IMDRF PATIENT CODE E2321 CAPTURES THE REPORTABLE EVENT OF LOW BLOOD PRESSURE/HYPOTENSION. IMDRF PATIENT CODE E0506 CAPTURES THE REPORTABLE EVENT OF HEMORRHAGE/BLOOD LOSS/BLEEDING. IMDRF PATIENT CODE E0514 CAPTURES THE REPORTABLE EVENT OF THROMBOSIS/THROMBUS. IMDRF PATIENT CODE E0731 CAPTURES THE REPORTABLE EVENT OF PLEURAL EFFUSION. IMDRF PATIENT CODE E130501 CAPTURES THE REPORTABLE EVENT OF RENAL FAILURE. IMDRF PATIENT CODE E0602 CAPTURES THE REPORTABLE EVENT OF CARDIAC ARREST IMDRF PATIENT CODE 2342CAPTURES THE REPORTABLE EVENT OF MULTIPLE ORGAN DYSFUNCTION SYNDROME. IMDRF PATIENT CODE E060101 CAPTURES THE REPORTABLE EVENT OF ASYSTOLE. IMDRF PATIENT CODE E2402 CAPTURES THE REPORTABLE EVENT OF DISTENTION. IMDRF PATIENT CODE E2333 CAPTURES THE REPORTABLE EVENT OF PROLAPSE. IMDRF PATIENT CODE E0519 CAPTURES THE REPORTABLE EVENT OF REDUCED BLOOD FLOW. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF DEATH. IMDRF IMPACT CODE F0801CAPTURES THE REPORTABLE EVENT OF INTENSIVE CARE. IMDRF IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF UNEXPECTED MEDICAL INTERVENTION. IMDRF IMPACT CODE F08 CAPTURES THE REPORTABLE EVENT OF HOSPITALIZATION OR PROLONGED HOSPITALIZATION. IMDRF IMPACT CODE F2303 CAPTURES THE REPORTABLE EVENT OF MEDICATION REQUIRED. BLOCK H11: ADDITIONAL INFORMATION UPDATED BLOCK B5: DESCRIBE EVENT OR PROBLEM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22155119&pc=NVN | 2017865-2025-75589 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-05 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED FOR PATIENT DEATH. AS RECEIVED ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22155125&pc=NVN | 2017865-2025-75588 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-05 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22155126&pc=LWP | 2017865-2025-75587 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-05 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO PATIENT EXPIRED. THE DEVICE WAS RECEIVED WITH A BATTERY LEVEL WITHIN THE NORMAL RANGE OF OPERATION. ELECTRICAL, LONGEVITY, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22155132&pc=NPT | 2015691-2025-04610 | 2025-05-13 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-05 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Degraded; Difficult to Open or Close; Device Stenosis; Central Regurgitation | Cardiomyopathy; Myocardial Infarction; Cardiogenic Shock; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES FIELD CLINICAL SPECIALIST, PATIENT HAD PRESENTED WITH A 23MM NON-EW VALVE WITH SEVERE AORTIC INSUFFICIENCY AND WAS AN URGENT CASE. A 23MM SAPIEN 3 ULTRA RESILIA (S3UR) VALVE WAS IMPLANTED INTO THE PREEXISTING NON-EW VALVE AND LANDED 90/10 (AORTIC: VENTRICULAR) WITH NO LEAK PER TTE. BIOPROSTHETIC VALVE FRACTURE WAS NOT PERFORMED. CATHETER GRADIENT MEASURED 3MMHG. THE VALVE WAS UNDER EXPANDED; THE PREEXISTING NON-EW VALVE MEASUREMENT WAS 340MM2. PER THE ACCOUNT, THE PATIENT HAD RISING GRADIENTS IN THE DAYS FOLLOWING THE IMPLANT AND THE VALVE APPEARED TO HAVE "PIN WHEELING" AND CENTRAL LEAK. THE LEAFLETS APPEARED TO BE "TWISTING" AT THEIR COAPTATION POINT PER PHYSICIAN. IMAGERY FROM POST OPERATIVE DAY 3 WAS PROVIDED BY THE SITE AND REVIEWED BY EW PROCTOR/IMAGERY. THE FOLLOWING OBSERVATIONS WERE MADE: CENTRAL REGURGITATION WAS OBSERVED. THE LEAFLETS APPEARED THICKENED, AND SOME SMALL STRANDS APPEARED TO BE ATTACHED TO THE ANTERIOR LEAFLET SIDE WITH INDEPENDENT MOTION. THE LEAFLETS APPEARED TO HAVE LIMITED MOTION, AND THE VALVE WAS PROBABLY STENOTIC. PER THE MEDICAL RECORD WAS RECEIVED, THE PATIENT UNDERWENT TRANSFEMORAL VALVE-IN-VALVE SUCCESSFULLY USING A 23MM S3UR VALVE WITHIN THE EXISTING 23MM NON-EW VALVE. POST-PROCEDURAL ECHOCARDIOGRAPHY SHOWED APPROPRIATE VALVE POSITIONING, NO SIGNIFICANT AORTIC INSUFFICIENCY, AND A MEAN GRADIENT OF 3.6 MMHG. HOWEVER, ELEVATED LEFT VENTRICULAR END-DIASTOLIC PRESSURE (LVEDP) WAS NOTED, CONSISTENT WITH ACUTE DIASTOLIC HEART FAILURE. NO FURTHER DILATION WAS PERFORMED, AND THERE WAS NO EVIDENCE OF SURGICAL VALVE FRACTURE. SUBSEQUENT IMAGING REVEALED PROGRESSIVE DYSFUNCTION OF THE NEW VALVE. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) TWO DAYS POST IMPLANT DEMONSTRATED SEVERE AORTIC STENOSIS (AS), AND A TRANSTHORACIC ECHOCARDIOGRAM (TTE) 3 DAYS POST IMPLANT CONFIRMED ABNORMAL VALVE FUNCTION WITH A MEAN GRADIENT OF 34 MMHG, AN AORTIC VALVE AREA (AVA) BY VELOCITY TIME INTEGRAL (VTI) OF 0.63 CM2, AND MODERATE CENTRAL REGURGITATION. THE PATIENT EXPIRED 7 DAYS POST IMPLANT FOLLOWING TRANSITION TO COMFORT CARE DUE TO POOR PROGNOSIS FROM FAILED VALVE-IN-VALVE TAVR, RESULTING IN CARDIOGENIC SHOCK AND MULTIORGAN FAILURE. Manufacturer Narrative: INVESTIGATION IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR CORRECTION AND ADDITIONAL INFORMATION BASED ON THE ENGINEER EVALUATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: ADDITIONAL CODE ADDED TO H6 COMPONENT CODES, ADDITIONAL CODE ADDED TO H6 DEVICE CODES, ADDITIONAL CODES ADDED TO H6 TYPE OF INVESTIGATION, CORRECTED H6 INVESTIGATION FINDINGS, AND CORRECTED H6 INVESTIGATION CONCLUSIONS. THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. ADDITIONAL ASSESSMENT OF THE FAILURE MODES IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. IMAGERY WAS PROVIDED AND REVIEWED BY (B)(4). THE FOLLOWING OBSERVATIONS WERE MADE: CENTRAL REGURGITATION WAS OBSERVED. THE LEAFLETS APPEARED THICKENED, AND SOME SMALL STRANDS APPEARED TO BE ATTACHED TO THE ANTERIOR LEAFLET SIDE WITH INDEPENDENT MOTION. THE LEAFLETS APPEARED TO HAVE LIMITED MOTION, AND THE VALVE WAS PROBABLY STENOTIC. THE INSTRUCTIONS FOR USE (IFU), AND TRAINING MANUALS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THE SAPIEN 3 ULTRA RESILIA THICKENED VALVE LEAFLETS, STENOSIS, CENTRAL REGURGITATION, LEAFLET MOTION RESTRICTION AND SUBSEQUENT PATIENT DEATH WERE CONFIRMED. DURING THE MANUFACTURING PROCESS, ALL SAPIEN 3 ULTRA RESILIA VALVES ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. HYPOATTENUATED LEAFLET THICKENING (HALT) MAY BE CAUSED BY SEVERAL PATIENT-RELATED FACTORS INCLUDING EARLY VALVE DETERIORATION (SVD) (E.G. STENOSIS/CALCIFICATION/REGURGITATION/INCREASED GRADIENT), NON-STRUCTURAL DYSFUNCTION (E.G. PANNUS), OR THROMBOSIS (FORMATION OF BLOOD CLOTS ON THE VALVE). DUE TO LIMITED INFORMATION, A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AT THIS TIME. REGURGITATION WITH RESTRICTED LEAFLET MOTION DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO A NUMBER OF ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. IN THIS CASE, THICKENED LEAFLET/HALT AND UNKNOWN ATTACHED STRAND WERE OBSERVED FROM THE IMAGERY. THICKENED LEAFLET/HALT AND UNKNOWN ATTACHED STRANDS MAY INTERFERE WITH THE FUNCTIONALITY OF THE DEVICE BY RESTRICTING LEAFLET MOTION AND LEADING TO ABNORMAL COAPTATION WITH REGURGITATION. IN ADDITION, THE REPORTED PINWHEELING OF THE VALVE LEAFLETS SUGGESTS UNDER-EXPANSION OF THE VALVE AT DEPLOYMENT. AN UNDER-EXPANDED VALVE CAN RESULT IN SUBOPTIMAL LEAFLET COAPTATION, LEADING TO RESTRICTED LEAFLET MOTION AND CENTRAL AORTIC REGURGITATION, AS OBSERVED. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (THICKENED LEAFLET, UNKNOWN ATTACHED STRAND) AND/OR PROCEDURAL FACTORS (UNDER-EXPANDED THV) MAY HAVE CONTRIBUTED TO THE CENTRAL REGURGITATION AND LEAFLET MOTION RESTRICTION. HOWEVER, A DEFINITIVE ROOT CAUSE IS UNABLE TO BE DETERMINED. PER THE INSTRUCTION FOR USE (IFU), VALVE STENOSIS IS A POTENTIAL RISK ASSOCIATED WITH BIOPROSTHETIC HEART VALVES. PER THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC), VALVE STENOSIS CAN RESULT FROM A NUMBER OF FACTORS, INCLUDING PANNUS, CALCIFICATION, SUPPORT STRUCTURE DEFORMATION (OUT-OF-ROUND CONFIGURATION), TRAUMA (CARDIA-PULMONARY RESUSCITATION, BLUNT CHEST TRAUMA), ENDOCARDITIS, PROSTHETIC VALVE THROMBOSIS, AND NATIVE LEAFLET PROLAPSE IMPEDING PROSTHETIC LEAFLET MOTION. VALVE STENOSIS MAY RESULT IN SYMPTOMS SUCH AS SHORTNESS OF BREATH AND DECREASED EXERCISE TOLERANCE, WHICH MAY BE ACCOMPANIED BY AN INCREASED GRADIENT ACROSS THE VALVE. THIS COULD BE DUE TO EARLY CALCIFICATION OF THE LEAFLETS, HOST TISSUE OVERGROWTH OR IN RARE CASES, A NON-FUNCTIONING LEAFLET. CALCIFICATION IS A WELL-RECOGNIZED FAILURE MODE OF BIOPROSTHETIC VALVES. THE MECHANISMS FOR BIOPROSTHETIC HEART VALVE TISSUE CALCIFICATION ARE NOT FULLY UNDERSTOOD. MANY FACTORS CAN CONTRIBUTE TO THE ONSET AND PROPAGATION OF CALCIFICATION INCLUDING PATIENT RELATED (E.G. PATIENT AGE, DISEASE STATE, IMMUNE STATUS, AND OTHER CO-MORBIDITIES), PHARMACOLOGICAL, AND INTRINSIC PROPERTIES OF THE VALVE ITSELF. STENOSIS OF AN IMPLANTED VALVE MAY BE A MANIFESTATION OF STRUCTURAL VALVE DETERIORATION (SVD). THIS TERM REFERS TO CHANGES INTRINSIC TO THE VALVE, AND CAN INCLUDE FAILURE MODES SUCH AS WEAR, CALCIFICATION, LEAFLET TEAR, STENT CREEP, LEAFLET DISRUPTION, OR LEAFLET RETRACTION. SVD MAY BE MILD AND NOT REQUIRE ANY INTERVENTION OR IT MAY BE MODERATE TO SEVERE. IT CAN CAUSE THE HEART TO WORK HARDER TO EJECT BLOOD FROM THE VENTRICLE. DEPENDING ON SEVERITY IT COULD BE AN INDICATION FOR VALVE REPLACEMENT OR MEDICAL INTERVENTION. IN THIS CASE, THICKENED LEAFLET AND UNKNOWN ATTACHED STRAND WERE OBSERVED FROM THE IMAGERY. THICKENED LEAFLET/HALT MAY REDUCE THE EOA (EFFECTIVE ORIFICE AREA) OF VALVE, AND ABNORMAL COAPTATION WITH NARROW OPENING (STENOSIS). THE UNKNOWN ATTACHED STRAND COULD AFFECT THE HEMODYNAMIC OF BLOOD FLOW, LEADING TO SUBOPTIMAL LEAFLETS OPENING OR NARROW OPENING (STENOSIS). IN ADDITION, THE IMPLANTED VALVE WITH PINWHEELING APPEARANCE OF THE VALVE LEAFLETS INDICATED VALVE UNDER-EXPANSION, WHICH COULD ALSO LEAD TO THE SUBOPTIMAL VALVE LEAFLETS COAPTATION WITH NARROW OPENING (STENOSIS). AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (THICKENED LEAFLET, UNKNOWN ATTACHED STRAND) AND/OR PROCEDURAL FACTORS (UNDER-EXPANDED THV) MAY HAVE CONTRIBUTED TO THE STENOSIS. HOWEVER, A DEFINITIVE ROOT CAUSE IS UNABLE TO BE DETERMINED. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22155136&pc=OAE | 2029046-2025-01834 | 2025-05-13 04:00 | Death | BIOSENSE WEBSTER INC | 2025-06-05 04:00 | OAE | UNK_QDOT MICRO | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Ischemic Heart Disease | P210027 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A CARDIAC ABLATION WITH A QDOT MICRO CATHETER AND THE PATIENT EXPERIENCED ST ELEVATION AND HEMORRHAGIC STROKE. THE PATIENT DIED. PATIENT HAD A REDO PVI (PULMONARY VEIN ISOLATION) OPERATION UNDER GA (GENERAL ANESTHESIA). CARTO3 SYSTEM WAS USED FOR MAPPING THE LEFT ATRIUM WITH THE OCTARAY. AS THE OCTARAY WAS PULLED OUT AND QDOT WAS BEING INTRODUCED THROUGH THE AGILIS SHEATH, IT WAS NOTICED THAT PATIENT HAD ST ELEVATION. QDOT WAS PULLED OUT AND DOCTOR COMMENTED THAT THERE WAS SOME AIR THAT CAME OUT OF THE SHEATH. THE STAFF AND DOCTOR ATTENDED THE SITUATION AND CALLED FOR ADDITIONAL CARDIOLOGIST SUPPORT. HOWEVER, THE PATIENT STABILIZED, AND NO FURTHER INTERVENTION WAS REQUIRED. AT THIS POINT QDOT WAS REINTRODUCED. THE REQUIRED LINES WERE ABLATED WITH THE COMBINATION OF QMODE PLUS AND QMODE. THE CASE WAS SUCCESSFULLY COMPLETED AND PATIENT LEFT FOR RECOVERY. ABOUT AN HOUR OR SO LATER, THE PATIENT WAS FOUND UNRESPONSIVE AND A NEUROLOGICAL EVENT WAS SUSPECTED AND HE WAS BROUGHT STRAIGHT TO CT (COMPUTED TOMOGRAPHY). ACCORDING TO INFORMATION PROVIDED THE PATIENT WAS FOUND TO HAVE HAD A HEMORRHAGIC STROKE. THE PATIENT WAS MOVED TO ANOTHER HOSPITAL FOR FURTHER TREATMENT AND EVALUATION. THE PHYSICIAN'S OPINION ON THE CAUSE OF THIS ADVERSE EVENT WAS A SUBARRACHONOID HAEMORROID DUE TO INTRACRANIAL ANEURYSM. THE PATIENT REQUIRED EXTENDED HOSPITALIZATION BY ONE WEEK AND RECEIVED A CT AND BRAIN SCAN. HOWEVER, THE OUTCOME WAS DEATH. THE PROCEDURAL ACT (ACTIVATED CLOTTING TIME) WAS 311. THERE WAS NO EVIDENCE OF CHAR DURING THE PROCEDURE. PAST MEDICAL HISTORY INCLUDES MILD CORONARY DISEASE AND HYPERTENSION. Manufacturer Narrative: PER INTERNAL REVIEW ON 20-JUN-2025, IT WAS DETERMINED THAT THIS EVENT IS NOT FDA-REPORTABLE FOR THE FOLLOWING REASONS: ALTHOUGH THE ST ELEVATION IS LIKELY A RESULT OF AIR FROM THE SHEATH, IT IS STILL BEING CONSERVATIVELY REPORTED AS THE PHYSICIAN DID NOT SPECIFICALLY REPORT THE CAUSE OF THE ST ELEVATION AND IT OCCURRED AFTER QDOT WAS REMOVED FROM THE BODY. NO FURTHER INTERVENTION WAS REQUIRED FOR THE ST ELEVATION AND THE PROLONGED HOSPITALIZATION WAS DUE TO THE NEUROLOGICAL EVENT (HEMORRHAGIC STROKE) THAT IS ASSOCIATED WITH THE INTRACRANIAL ANEURYSM. THERE WERE NO MALFUNCTIONS WITH ANY BWI PRODUCTS, THAT WOULD CAUSE THEM TO BE SUSPECTED DEVICES AND BASED ON THE CURRENT INFORMATION, THE FACTORS THAT CONTRIBUTED TO THE DEATH (INTRACRANIAL ANEURYSM) WERE NOT CORRELATED WITH BWI CARDIAC ABLATION DEVICES. THIS EVENT IS NOT FDA REPORTABLE AND THEREFORE NO FURTHER UPDATES WILL BE SENT REGARDING THIS EVENT. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC. OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22155233&pc=NVN | 2017865-2025-75596 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-05 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22155234&pc=LWP | 2017865-2025-75594 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-05 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: DEVICE WAS RETURNED FOR ANALYSIS DUE TO PATIENT DEATH. DEVICE IMAGE WAS SAVED SUCCESSFULLY. ONCE THE DEVICE WAS RECEIVED, IT DISPLAYED NORMAL DEVICE CHARACTERISTICS UPON THE FIRST STAGE OF ANALYSIS. FURTHERMORE, THE DEVICE IMAGE WAS SAVED AND ANALYZED, NO DISCREPANCIES WERE NOTED. AFTER ALL ELECTRICAL TESTS PERFORMED, THE DEVICE CONTINUED EXHIBITING NORMAL DEVICE CHARACTERISTICS. LONGEVITY ASSESSMENT WAS PERFORMED AND DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY. NO OTHER ANOMALIES WERE FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22155235&pc=NVN | 2017865-2025-75595 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-05 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED. THE CAUSE OF DEATH WAS CARDIOPULMONARY ARREST. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTING THE DEATH WAS RELATED TO ANY ABBOTT PRODUCTS OR PROCEDURES. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL, MECHANICAL, AND VISUAL ANALYSIS WAS NORMAL WITH NO ANOMALIES FOUND. Manufacturer Narrative: CORRECTION ¿ B5 - DESCRIBE EVENT OR PROBLEM ¿ INITIAL MDR SUBMITTED ON JUN 5, 2025, WAS MISSING SOME INFORMATION, THIS HAS BEEN UPDATED IN THE REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22155594&pc=LPB | 2029046-2025-01839 | 2025-04-04 04:00 | Death | BIOSENSE WEBSTER INC | 2025-06-05 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Fistula | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION ABLATION PROCEDURE WITH A THERMOCOOL SMARTTOUCH SF AND A SMARTABLATE GENERATOR AND THE PATIENT EXPERIENCED AN ESOPHAGEAL FISTULA THAT RESULTED IN DEATH. THE PHYSICIAN STATED THAT THE PATIENT HAD PASSED FROM COMPLICATIONS OF AN ESOPHAGEAL FISTULA. THE PHYSICIAN STATES THE PATIENT HAD A FOLLOW-UP IN HIS CARDIOLOGY CLINIC, THEN TWO VISITS TO AN OUTSIDE EMERGENCY DEPARTMENT 9 ED). THEN WAS ADMITTED TO AN OUTSIDE HOSPITAL. THE PHYSICIAN STATED THE PATIENT HAD NORMAL LABS AT THE TIME OF FOLLOW UP POST ABLATION. THE DATE OF THE PROCEDURE WAS ON (B)(6) 2025. IT WAS REPORTED THAT THREE ATRIAL FIB PROCEDURES WERE COMPLETED ON (B)(6) WITH THE THERMOCOOL SMARTTOUCH SF ABLATION CATHETER. THE PHYSICIAN HAD ONLY RECENTLY FOUND OUT AND THAT THE DETAILS ARE LIMITED. IT IS UNKNOWN WHEN THE DEATH OCCURRED OR HOW THE ESOPHAGEAL FISTULA WAS DISCOVERED. THE PHYSICIAN STATED THE PATIENT DID NOT RECEIVE AN ESOPHAGOGASTRODUODENOSCOPY (EGD). THE PHYSICIAN STATED THE PATIENT DID RECEIVE A COMPUTED TOMOGRAPHY (CT) SCAN BUT DID NOT KNOW IF CONTRAST WAS USED. THE PHYSICIAN STATED THE CT REPORT DID NOT NOTED ANY EVIDENCE OF ATRIAL ESOPHAGEAL FISTULA. THE PHYSICIAN DID STATE THAT HE REVIEWED HIS PROCEDURE NOTES AND THAT THE ESOPHAGEAL TEMPERATURE DID NOT EXCEED 37.2-37.4 CELSIUS. IT WAS REPORTED THAT WHEN THE PHYSICIAN WENT TO RECALL THE PROCEDURE COMPLETED ON (B)(6) THAT THE PROCEDURE IS MISSING FROM THE EXTERNAL HARD DRIVE AND THE CARTO 3 SYSTEM AND THE LAST DATE OF BACK UP IS ON (B)(6) 2025. THE ACTIVATED CLOTTING TIME (ACT) VALUES IN THE PROCEDURE ARE UNKNOWN, HOWEVER, THE PROCEDURE GOAL WAS 350 AND CHECKED EVERY 10 MINUTES PER HOSPITAL PROTOCOL. SMARTABLATE WAS USED WITH: POWER CONTROL 35W TEMP SETTINGS: WARNING 37*, CUT OFF 40*. IMPEDANCE SETTINGS: MAX 250, MIN 50, SPIKE OFF. VISTAG SURPOINT. FORCE VISUALIZATION FEATURES USED WERE DASHBOARD, VECTOR AND VISITAG. PARAMESTERS FOR STABILITY USED FOR VISITAG MODULE WAS MAXIMUM DISTANCE CHANGE- 3MM, MINIMUM TIME- 3 SECONDS, FORCE OVER TIME - 25%, MINIMUM FORCE - 3 GRAMS. RESPIRATORY ADJUSTMENT WAS USED AND TAG INDEX WAS USED FOR COLOR OPTIONS. ESOPHAGEAL TEMPERATURE MONITORING WAS DONE. BASELINE TEMPERATURE UNKNOWN. THE PHYSICIAN FIRST STATED THE MAXIMUM TEMP REACHED 37.2 PER HIS PROCEDURE REPORT. THE LATER, STATED THE MAXIMUM TEMPERATURE WAS 37.4. NO CRITICAL ERRORS NOTED. ONLY FUNCTIONAL ERRORS OBSERVED. THERE WERE NO ISSUES WITH FLOW RATE CHANGE AT THE START OF ABLATION. THE CATHETER WAS RE-ZEROED PER JNJ PROTOCOL. FREE FLOATING IN THE LEFT ATRIUM. THERE WERE NO REPORTED BWI PRODUCT MALFUNCTIONS. THE "FUNCTIONAL ERRORS" INCLUDED "MAPPING CATHETER OUT OF RANGE" OR "MAPPING CATHETER REQUIRING ZERO." Manufacturer Narrative: DEVICE INVESTIGATION DETAILS: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. HOWEVER, IF THE PRODUCT OR PRODUCT ID NUMBERS ARE RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. NOTE: DATE OF DEATH WAS NOT REPORTED. BEST ESTIMATED DATE IS ON (B)(6) 2025, SAME AS EVENT DATE. NOTE: FOR FIELD D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REF. #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22155599&pc=LPB | 2029046-2025-50001 | 2025-04-04 04:00 | Death | STOCKERT GMBH | 2025-06-05 04:00 | LPB | SMARTABLATE GENERATOR | Adverse Event Without Identified Device or Use Problem | Fistula |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION ABLATION PROCEDURE WITH A THERMOCOOL SMARTTOUCH SF AND A SMARTABLATE GENERATOR AND THE PATIENT EXPERIENCED AN ESOPHAGEAL FISTULA THAT RESULTED IN DEATH. THE PHYSICIAN STATED THAT THE PATIENT HAD PASSED FROM COMPLICATIONS OF AN ESOPHAGEAL FISTULA. THE PHYSICIAN STATES THE PATIENT HAD A FOLLOW-UP IN HIS CARDIOLOGY CLINIC, THEN TWO VISITS TO AN OUTSIDE EMERGENCY DEPARTMENT 9 ED). THEN WAS ADMITTED TO AN OUTSIDE HOSPITAL. THE PHYSICIAN STATED THE PATIENT HAD NORMAL LABS AT THE TIME OF FOLLOW UP POST ABLATION. THE DATE OF THE PROCEDURE WAS (B)(6) 2025. IT WAS REPORTED THAT THREE ATRIAL FIB PROCEDURES WERE COMPLETED ON (B)(6) WITH THE THERMOCOOL SMARTTOUCH SF ABLATION CATHETER. THE PHYSICIAN HAD ONLY RECENTLY FOUND OUT AND THAT THE DETAILS ARE LIMITED. IT IS UNKNOWN WHEN THE DEATH OCCURRED OR HOW THE ESOPHAGEAL FISTULA WAS DISCOVERED. THE PHYSICIAN STATED THE PATIENT DID NOT RECEIVE AN ESOPHAGOGASTRODUODENOSCOPY (EGD). THE PHYSICIAN STATED THE PATIENT DID RECEIVE A COMPUTED TOMOGRAPHY (CT) SCAN BUT DID NOT KNOW IF CONTRAST WAS USED. THE PHYSICIAN STATED THE CT REPORT DID NOT NOTED ANY EVIDENCE OF ATRIAL ESOPHAGEAL FISTULA. THE PHYSICIAN DID STATE THAT HE REVIEWED HIS PROCEDURE NOTES AND THAT THE ESOPHAGEAL TEMPERATURE DID NOT EXCEED 37.2-37.4 CELSIUS. IT WAS REPORTED THAT WHEN THE PHYSICIAN WENT TO RECALL THE PROCEDURE COMPLETED ON (B)(6) THAT THE PROCEDURE IS MISSING FROM THE EXTERNAL HARD DRIVE AND THE CARTO 3 SYSTEM AND THE LAST DATE OF BACK UP IS (B)(6) 2025. THE ACTIVATED CLOTTING TIME (ACT) VALUES IN THE PROCEDURE ARE UNKNOWN, HOWEVER, THE PROCEDURE GOAL WAS 350 AND CHECKED EVERY 10 MINUTES PER HOSPITAL PROTOCOL. SMARTABLATE WAS USED WITH: POWER CONTROL 35W TEMP SETTINGS: WARNING 37*, CUT OFF 40*. IMPEDANCE SETTINGS: MAX 250, MIN 50, SPIKE OFF. VISTAG SURPOINT. FORCE VISUALIZATION FEATURES USED WERE DASHBOARD, VECTOR AND VISITAG. PARAMESTERS FOR STABILITY USED FOR VISITAG MODULE WAS MAXIMUM DISTANCE CHANGE- 3MM, MINIMUM TIME- 3 SECONDS, FORCE OVER TIME - 25%, MINIMUM FORCE - 3 GRAMS. RESPIRATORY ADJUSTMENT WAS USED AND TAG INDEX WAS USED FOR COLOR OPTIONS. ESOPHAGEAL TEMPERATURE MONITORING WAS DONE. BASELINE TEMPERATURE UNKNOWN. THE PHYSICIAN FIRST STATED THE MAXIMUM TEMP REACHED 37.2 PER HIS PROCEDURE REPORT. THE LATER, STATED THE MAXIMUM TEMPERATURE WAS 37.4. NO CRITICAL ERRORS NOTED. ONLY FUNCTIONAL ERRORS OBSERVED. THERE WERE NO ISSUES WITH FLOW RATE CHANGE AT THE START OF ABLATION. THE CATHETER WAS RE-ZEROED PER JNJ PROTOCOL. FREE FLOATING IN THE LEFT ATRIUM. THERE WERE NO REPORTED BWI PRODUCT MALFUNCTIONS. THE "FUNCTIONAL ERRORS" INCLUDED "MAPPING CATHETER OUT OF RANGE" OR "MAPPING CATHETER REQUIRING ZERO." Manufacturer Narrative: THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REF. # (B)(4). Manufacturer Narrative: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION ABLATION PROCEDURE WITH A THERMOCOOL SMARTTOUCH SF AND A SMARTABLATE GENERATOR AND THE PATIENT EXPERIENCED AN ESOPHAGEAL FISTULA THAT RESULTED IN DEATH. THE PHYSICIAN STATED THAT THE PATIENT HAD PASSED FROM COMPLICATIONS OF AN ESOPHAGEAL FISTULA. THE PHYSICIAN STATES THE PATIENT HAD A FOLLOW-UP IN HIS CARDIOLOGY CLINIC, THEN TWO VISITS TO AN OUTSIDE EMERGENCY DEPARTMENT 9 ED). THEN WAS ADMITTED TO AN OUTSIDE HOSPITAL. THE PHYSICIAN STATED THE PATIENT HAD NORMAL LABS AT THE TIME OF FOLLOW UP POST ABLATION. IT IS UNKNOWN WHEN THE DEATH OCCURRED OR HOW THE ESOPHAGEAL FISTULA WAS DISCOVERED. THE PHYSICIAN STATED THE PATIENT DID NOT RECEIVE AN ESOPHAGOGASTRODUODENOSCOPY (EGD). THE PHYSICIAN STATED THE PATIENT DID RECEIVE A COMPUTED TOMOGRAPHY (CT) SCAN BUT DID NOT KNOW IF CONTRAST WAS USED. THE PHYSICIAN STATED THE CT REPORT DID NOT NOTED ANY EVIDENCE OF ATRIAL ESOPHAGEAL FISTULA. DEVICE EVALUATION DETAILS: TECHNICAL SERVICES PERFORMED REMOTE COMMUNICATIONS WITH THE FACILITY WHERE THE DEVICE IS INSTALLED TO UNDERSTAND THE REPORTED EVENT. HOWEVER, SERVICE WAS DECLINED. IT WAS DETERMINED THAT IT WAS NOT CONSIDERED NECESSARY TO PERFORM FURTHER INVESTIGATION ON THIS DEVICE, AS IT PERFORMED AS EXPECTED. THE ACCOUNT DID NOT RETURN THE DEVICE BACK FOR INVESTIGATION. A DEVICE HISTORY RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER, AND NO INTERNAL ACTIONS RELATED TO THE REPORTED CONDITION WERE IDENTIFIED. AS PART OF THE QUALITY PROCESS, THE DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REF. # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22155758&pc=BZD | 2518422-2025-106401 | 2025-05-29 04:00 | Death | RESPIRONICS, INC. | 2025-06-05 04:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE PATIENT'S DAUGHTER INFORMED THE MANUFACTURER OF THE PATIENT'S PASSING. AT THIS TIME, THERE IS NO ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED DEATH. THE DAUGHTER WAS INFORMED THAT THE DEVICE WAS SOLD BY A MEDICAL SUPPLY COMPANY AND STATED THAT SHE WOULD DISPOSE OF IT. SHE THEN THANKED THE MANUFACTURER'S REPRESENTATIVE AND ENDED THE CALL. THE MANUFACTURER'S INVESTIGATION IS CURRENTLY ONGOING, AND ONCE IT IS COMPLETED, A FOLLOW-UP REPORT WILL BE SUBMITTED. ADDITIONALLY, A REPORT WILL BE FILED WITH ALL APPLICABLE REGULATORY AGENCIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22138792&pc=MIH | 3013164176-2025-02544 | 2025-05-13 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-04 04:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Ischemia; Paralysis; Embolism/Embolus | P020004 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: THIS PATIENT WITH ABDOMINAL AORTIC ANEURYSM (AAA) HAD EXTENSIVE THROMBUS WITHIN THE AORTA FROM THE DESCENDING AORTA TO THE RENAL ARTERY BIFURCATION, MAKING IT AN EXTREMELY HIGH-RISK CASE. ALTHOUGH TREATMENT HAD BEEN DECLINED AT ANOTHER HOSPITAL, ENDOVASCULAR AORTIC REPAIR (EVAR) WAS TO BE PERFORMED AT THE REPORTING HOSPITAL CONSIDERING THE RISK OF ANEURYSM RUPTURE. ON (B)(6) 2025, THE PATIENT UNDERWENT AN EVAR FOR AAA USING GORE® EXCLUDER® AAA ENDOPROSTHESIS (EXCLUDER DEVICES). PRIOR TO THE PLACEMENT OF EXCLUDER DEVICES, AN ATTEMPT WAS MADE TO ACCESS THE RIGHT INTERNAL ILIAC ARTERY FOR COIL EMBOLIZATION, BUT IT WAS NOT SUCCEEDED AND TOOK APPROXIMATELY TWO HOURS. GIVEN THE HIGH-RISK NATURE OF THE CASE, EXTREME CAUTION WAS REQUIRED IN MANIPULATING GUIDEWIRES AND CATHETERS, AND ANGIOGRAPHY HAD TO BE MINIMIZED. AS A RESULT, THE PROCEDURE TOOK A TOTAL OF FOUR HOURS TO COMPLETE. THE PROCEDURE WAS SUCCESSFULLY DONE WITHOUT ANY ENDOLEAKS. ON (B)(6) 2025, THE PATIENT DEVELOPED PARAPLEGIA IN THE MORNING. ON (B)(6) 2025, THE PATIENT EXPERIENCED BLOODY STOOLS. ON (B)(6) 2025, IN THE EARLY MORNING, THE PATIENT WENT INTO SHOCK STATE AND SUBSEQUENTLY PASSED AWAY. THE REPORTING PHYSICIAN COMMENTED AS FOLLOWS: EXTREME CARE WAS TAKEN DURING GUIDEWIRE MANIPULATION AND BALLOON INFLATION. ANGIOGRAPHY WAS KEPT TO A MINIMUM, AND ALL POSSIBLE MEASURES WERE TAKEN. HOWEVER, THE THING THAT HAD BEEN FEARED OCCURRED. IT IS BELIEVED THAT THROMBUS DISPERSED DURING THE INDEX PROCEDURE, LEADING TO INTESTINAL ISCHEMIA AND ULTIMATELY RESULTING IN THE PATIENT¿S DEATH. ADDITIONAL PHYSICIAN'S COMMENT: THE ASSOCIATION BETWEEN THE GORE DEVICES AND INTESTINAL ISCHEMIA IS UNCLEAR. FUNDAMENTALLY, THE PATIENT HAD A HIGH-RISK BACKGROUND, AND THE PROCEDURE ITSELF WAS CHALLENGING TO PERFORM. REGARDING THE PARAPLEGIA, THERE WERE NUMEROUS THROMBI FROM THE DESCENDING AORTA TO THE RENAL ARTERY BIFURCATION. IT IS PRESUMED THAT WHEN CONTRAST MEDIUM WAS INJECTED, THE THROMBI WERE DISLODGED AND EMBOLIZED SMALLER ARTERIES SUCH AS THE LUMBAR ARTERIES. Manufacturer Narrative: AS GORE WAS UNABLE TO DETERMINE WHICH DEVICE/DEVICE COMPONENT IS INVOLVED IN THIS REPORTABLE ADVERSE EVENT, THE FOLLOWING ADDITIONAL DEVICE WILL BE IDENTIFIED IN THIS REPORT: CATALOG #RLT261418J, SERIAL #(B)(6), UDI ((B)(4); CATALOG #PLC201400J, SERIAL #(B)(6), UDI (B)(4); AND CATALOG #PLA260300J, SERIAL #(B)(6), UDI (B)(4). H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: ACCORDING TO THE GORE® EXCLUDER® AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION OR ADDITIONAL INTRAOPERATIVE PROCEDURE TIME INCLUDE, BUT ARE NOT LIMITED TO: EMBOLISATION (MICRO AND MACRO) WITH TRANSIENT OR PERMANENT ISCHEMIA. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY HAVE NOT BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22138843&pc=DYB | 3007284313-2025-04012 | 2025-05-13 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-04 04:00 | DYB | GORE® DRYSEAL FLEX INTRODUCER SHEATH | Adverse Event Without Identified Device or Use Problem | Ischemia; Paralysis; Embolism/Embolus | K160254 |  | 1 | Event Description: THE FOLLOWING INFORMATION WAS REPORTED TO GORE: THIS PATIENT WITH ABDOMINAL AORTIC ANEURYSM (AAA) HAD EXTENSIVE THROMBUS WITHIN THE AORTA FROM THE DESCENDING AORTA TO THE RENAL ARTERY BIFURCATION, MAKING IT AN EXTREMELY HIGH-RISK CASE. ALTHOUGH TREATMENT HAD BEEN DECLINED AT ANOTHER HOSPITAL, ENDOVASCULAR AORTIC REPAIR (EVAR) WAS TO BE PERFORMED AT THE REPORTING HOSPITAL CONSIDERING THE RISK OF ANEURYSM RUPTURE. ON (B)(6) 2025, THE PATIENT UNDERWENT AN EVAR FOR AAA USING GORE® EXCLUDER® AAA ENDOPROSTHESIS (EXCLUDER DEVICES). PRIOR TO THE PLACEMENT OF EXCLUDER DEVICES, AN ATTEMPT WAS MADE TO ACCESS THE RIGHT INTERNAL ILIAC ARTERY FOR COIL EMBOLIZATION, BUT IT WAS NOT SUCCEEDED AND TOOK APPROXIMATELY TWO HOURS. GIVEN THE HIGH-RISK NATURE OF THE CASE, EXTREME CAUTION WAS REQUIRED IN MANIPULATING GUIDEWIRES AND CATHETERS, AND ANGIOGRAPHY HAD TO BE MINIMIZED. AS A RESULT, THE PROCEDURE TOOK A TOTAL OF FOUR HOURS TO COMPLETE. THE PROCEDURE WAS SUCCESSFULLY DONE WITHOUT ANY ENDOLEAKS. ON (B)(6) 2025, THE PATIENT DEVELOPED PARAPLEGIA IN THE MORNING. ON (B)(6) 2025, THE PATIENT EXPERIENCED BLOODY STOOLS. ON (B)(6) 2025, IN THE EARLY MORNING, THE PATIENT WENT INTO SHOCK STATE AND SUBSEQUENTLY PASSED AWAY. THE REPORTING PHYSICIAN COMMENTED AS FOLLOWS: EXTREME CARE WAS TAKEN DURING GUIDEWIRE MANIPULATION AND BALLOON INFLATION. ANGIOGRAPHY WAS KEPT TO A MINIMUM, AND ALL POSSIBLE MEASURES WERE TAKEN. HOWEVER, THE THING THAT HAD BEEN FEARED OCCURRED. IT IS BELIEVED THAT THROMBUS DISPERSED DURING THE INDEX PROCEDURE, LEADING TO INTESTINAL ISCHEMIA AND ULTIMATELY RESULTING IN THE PATIENT¿S DEATH. ADDITIONAL PHYSICIAN'S COMMENT: THE ASSOCIATION BETWEEN THE GORE DEVICES AND INTESTINAL ISCHEMIA IS UNCLEAR. FUNDAMENTALLY, THE PATIENT HAD A HIGH-RISK BACKGROUND, AND THE PROCEDURE ITSELF WAS CHALLENGING TO PERFORM. REGARDING THE PARAPLEGIA, THERE WERE NUMEROUS THROMBI FROM THE DESCENDING AORTA TO THE RENAL ARTERY BIFURCATION. IT IS PRESUMED THAT WHEN CONTRAST MEDIUM WAS INJECTED, THE THROMBI WERE DISLODGED AND EMBOLIZED SMALLER ARTERIES SUCH AS THE LUMBAR ARTERIES. Manufacturer Narrative: AS GORE WAS UNABLE TO DETERMINE WHICH DEVICE/DEVICE COMPONENT IS INVOLVED IN THIS REPORTABLE ADVERSE EVENT, THE FOLLOWING ADDITIONAL DEVICE WILL BE IDENTIFIED IN THIS REPORT: CATALOG #DSF1233 WITH UNKNOWN SERIAL # OR UDI. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. H6: CODE D12: ACCORDING TO THE GORE® DRYSEAL FLEX INTRODUCER SHEATH INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION OR ADDITIONAL INTRAOPERATIVE PROCEDURE TIME INCLUDE, BUT ARE NOT LIMITED TO: EMBOLIZATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22138960&pc=OZP | 2032227-2025-185078 | 2023-01-17 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-04 04:00 | OZP | 670G INSULIN PUMP MMT-1780KL | Break | No Clinical Signs, Symptoms or Conditions | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2023 AT HOME. MEDTRONIC LEGAL RECEIVED INFORMATION FROM THE ATTORNEY OF THE FAMILY ON 9TH MAY,2025. MEDTRONIC RECEIVED NOTICE OF A DEVICE/PRODUCT LEGAL HOLD NOTIFICATION, AND THE REPORTER ALLEGES THAT THE USE OF ONE OR MORE MEDTRONIC MINIMED 600 SERIES INSULIN PUMPS WITH A DAMAGED, MISSING, OR BROKEN RETAINER RING CAUSED THE CLAIMANT TO SUFFER PERSONAL INJURIES, AND THE CUSTOMER WAS HOSPITALIZED. THE ALLEGATION DID NOT SPECIFY THE PUMP IDENTIFIER (SERIAL NUMBER) AND THEREFORE ALL 600 SERIES INSULIN PUMPS IN POSSESSION OF THE CLAIMANT, INCLUDING THIS PUMP, ARE CONSIDERED POTENTIALLY WITHIN THE SCOPE OF THE REPORT PENDING CONFIRMATION OF THE AFFECTED SERIAL NUMBER(S). ALL RELATED REPORTS WILL BE UPDATED ACCORDINGLY WHEN AND IF THE AFFECTED SERIAL NUMBER IS IDENTIFIED. THE CAUSE OF DEATH WAS DIABETIC KETOACIDOSIS. IT WAS UNKNOWN WHETHER THE INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-1780KL, MMT-1780KPK, UNOMEDICAL. TROUBLESHOOTING WAS NOT PERFORMED. IT WAS UNKNOWN WHETHER THE CUSTOMER USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT, AND IT WAS UNKNOWN WHETHER THE SMARTGUARD/AUTO MODE OF THE INSULIN PUMP WAS USED OR NOT. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KL. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KPK. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22138983&pc=OZP | 2032227-2025-185080 | 2023-01-17 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-04 04:00 | OZP | PUMP MMT-1780KPK 670G PATHWAY BLACK MG | Break | No Clinical Signs, Symptoms or Conditions | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2023 AT HOME. MEDTRONIC LEGAL RECEIVED INFORMATION FROM THE ATTORNEY OF THE FAMILY ON 9TH MAY, 2025. MEDTRONIC RECEIVED NOTICE OF A DEVICE/PRODUCT LEGAL HOLD NOTIFICATION, AND THE REPORTER ALLEGES THAT THE USE OF ONE OR MORE MEDTRONIC MINIMED 600 SERIES INSULIN PUMPS WITH A DAMAGED, MISSING, OR BROKEN RETAINER RING CAUSED THE CLAIMANT TO SUFFER PERSONAL INJURIES, AND THE CUSTOMER WAS HOSPITALIZED. THE ALLEGATION DID NOT SPECIFY THE PUMP IDENTIFIER (SERIAL NUMBER) AND THEREFORE ALL 600 SERIES INSULIN PUMPS IN POSSESSION OF THE CLAIMANT, INCLUDING THIS PUMP, ARE CONSIDERED POTENTIALLY WITHIN THE SCOPE OF THE REPORT PENDING CONFIRMATION OF THE AFFECTED SERIAL NUMBER(S). ALL RELATED REPORTS WILL BE UPDATED ACCORDINGLY WHEN AND IF THE AFFECTED SERIAL NUMBER IS IDENTIFIED. THE CAUSE OF DEATH WAS DIABETIC KETOACIDOSIS. IT WAS UNKNOWN WHETHER THE INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-1780KL, MMT-1780KPK, UNOMEDICAL. TROUBLESHOOTING WAS NOT PERFORMED. IT WAS UNKNOWN WHETHER THE CUSTOMER USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT, AND IT WAS UNKNOWN WHETHER THE SMARTGUARD/AUTO MODE OF THE INSULIN PUMP WAS USED OR NOT. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KL. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KPK. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22139033&pc=MHX | 9610816-2025-000456 | 2025-01-08 05:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-06-04 04:00 | MHX | INTELLIVUE MX450 PATIENT MONITOR | Device Alarm System | Bradycardia | K130849 |  | 1 | Event Description: IT WAS REPORTED DUE TO A DESATURATION EVENT, THE CUSTOMER REQUESTED ASSISTANCE TO MODIFY THE CRITICALITY OF THE TECHNICAL INOP ALARM FOR 'SPO2 SENSOR DISCONNECTED'. IT WAS INDICATED A RED OR YELLOW ALARM FOR THIS INOP WAS DESIRED AS CAREGIVERS DO NOT PAY ATTENTION TO THE BLUE INOP AND THE PATIENT SEVERELY DESATURATED AND BECAME BRADYCARDIC. A RESUSCITATION WAS UNSUCCESSFUL AND THE VENTILATED PATIENT ULTIMATELY PASSED AWAY. THERE WAS NO ALLEGATION OF MALFUNCTION; HOWEVER, AS THE STAFF DOES NOT PAY ATTENTION TO INOP ALARMS, THE REQUEST FOR A CONFIGURATION CHANGE FOR SPO2 SENSOR DISCONNECTED ALARMS WAS MADE. THE DEVICE WAS IN USE ON A PATIENT. THERE WAS A REPORT OF PATIENT OR USER HARM. Manufacturer Narrative: D4: UDI IS NOT AVAILABLE DUE TO INCOMPLETE SERIAL NUMBER. SERIAL NUMBER WAS REQUESTED, BUT IS UNKNOWN. E1: REPORTING INSTITUTION PHONE # (B)(6). THE FOLLOWING FUNCTIONAL TESTS WERE PERFORMED: THE ISSUE WAS EVALUATED BY A CLINICAL APPLICATION SPECIALIST. IT TURNED OUT THAT THERE WAS NO MALFUNCTION OF THE DEVICE. UPON FURTHER INSPECTION THE CUSTOMER WAS INFORMED THAT A CONFIGURATION FOR "SPO2 SENSOR DISCONNECTED" IN YELLOW OR RED ALARM, INSTEAD OF A BLUE ALARM IS NOT POSSIBLE. THIS CASE WILL BE HANDLED AS AN ENHANCEMENT REQUEST. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS A USER ERROR. THE ENGINEER PROVIDED THEIR ANALYSIS FINDINGS AND CONFIRMED THAT DEVICE TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22139057&pc=OZO | 2032227-2025-185133 | 2022-08-02 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-04 04:00 | OZO | 630G INSULIN PUMP MMT-1715KL 630G | Break | Hyperglycemia | P150001 |  | 1 | Event Description: MEDTRONIC LEGAL RECEIVED INFORMATION FROM THE ATTORNEY OF THE FAMILY ON 9TH MAY,2025. MEDTRONIC RECEIVED NOTICE OF A DEVICE/PRODUCT LEGAL HOLD NOTIFICATION, AND THE REPORTER ALLEGES THAT THE USE OF ONE OR MORE MEDTRONIC MINIMED 600 SERIES INSULIN PUMPS WITH A DAMAGED, MISSING, OR BROKEN RETAINER RING CAUSED THE CLAIMANT TO SUFFER PERSONAL INJURIES, AND THE CUSTOMER WAS HOSPITALIZED. THE ALLEGATION DID NOT SPECIFY THE PUMP IDENTIFIER (SERIAL NUMBER) AND THEREFORE ALL 600 SERIES INSULIN PUMPS IN POSSESSION OF THE CLAIMANT, INCLUDING THIS PUMP, ARE CONSIDERED POTENTIALLY WITHIN THE SCOPE OF THE REPORT PENDING CONFIRMATION OF THE AFFECTED SERIAL NUMBER(S). ALL RELATED REPORTS WILL BE UPDATED ACCORDINGLY WHEN AND IF THE AFFECTED SERIAL NUMBER IS IDENTIFIED. THE CAUSE OF DEATH WAS DIABETIC KETOACIDOSIS. IT WAS UNKNOWN WHETHER THE INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1715KL, MMT-332A, UNOMEDICAL. TROUBLESHOOTING WAS NOT PERFORMED. IT WAS UNKNOWN WHETHER THE CUSTOMER USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT, AND IT WAS UNKNOWN WHETHER THE SMARTGUARD/AUTO MODE OF THE INSULIN PUMP WAS USED OR NOT. NO PRODUCT RETURN IS REQUIRED FOR MMT-1715KL. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22139524&pc=LOD | 0002249697-2025-00587 | 2025-05-12 04:00 | Death | STRYKER ORTHOPAEDICS-MAHWAH | 2025-06-04 04:00 | LOD | SIMPLEX P-JAPANESE TWIN PACK | Patient-Device Incompatibility | Cardiac Arrest |  |  | 1 | Event Description: FOR USE WITH EXETER LONG, THE PLUG WAS INSERTED, AND CEMENT WAS INJECTED USING A CEMENT GUN THREE MINUTES AFTER THE START OF MIXING. WHEN THE STEM WAS INSERTED, A SUDDEN DROP IN BLOOD PRESSURE OF THE PATIENT WAS CONFIRMED. A DEFIBRILLATOR WAS USED, AND CARDIAC MASSAGE WAS CONTINUED FOR ABOUT 60 MINUTES, BUT THE PULSE DID NOT RETURN AND THE PATIENT DIED. Manufacturer Narrative: THE EVENT INVOLVES A DEVICE THAT IS NOT CLEARED FOR SALE IN THE U.S., BUT SIMILAR DEVICE IS COMMERCIALLY AVAILABLE IN THE U.S. REVIEW OF THE DEVICE HISTORY RECORDS INDICATE DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO RELEVANT REPORTED DISCREPANCIES. THERE HAVE BEEN NO SIMILAR EVENTS FOR THE LOT REFERENCED. AN EVALUATION OF THE DEVICE CANNOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE IT WILL BE REPORTED IN A SUPPLEMENTAL REPORT UPON COMPLETION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22139768&pc=DTQ | 8010762-2025-0000239 | 2025-05-27 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-06-04 04:00 | DTQ | CARDIOHELP | Device Handling Problem | Hemorrhage/Bleeding; Asystole | K133598 |  | 1 | Event Description: THE EVENT OCCURRED IN DENMARK DURING TREATMENT. IT WAS REPORTED THAT THE PATIENT ARRIVED IN THE HOSPITAL WITH A LUCAS PUMP PUMPING FOR ONE HOUR AND THE PATIENT HAS A CLOSED LEFT ANTERIOR DESCENDING (LAD). IT WAS HARD TO CANNULATE THE PATIENT AS THE PATIENT WAS BLEEDING ACCORDING TO THE CUSTOMER FROM ¿EVERYWHERE¿. THIS WAS CAUSED DUE TO THE LONG-TERM USE OF THE LUCAS PUMP. THE PATIENT WAS PUT ON ECMO. A CORONARY ARTERIOGRAPHY (KAG) WAS DONE AND THE LAD WAS OPENED. THE CUSTOMER WANTED TO DO A SMALL BYPASS WITH THE CANNULATION TO SECURE THE BLOOD FLOW TO THE LEG AND THEN THE STAFF CLAMPED THE HLS SET. A DEVICE DEFECTIVE ERROR MESSAGE OCCURRED, AS THE CUSTOMER HAS REMOVED THE SET WHILE THE PUMP WAS RUNNING. THE PATIENT WAS HAND CRANKED AND THE CARDIOHELP DEVICE WAS REPLACED. NEW INFORMATION WAS RECEIVED ON (B)(6) 2025 THAT THE PATIENT REMAINED FOR 3.5 HOURS ON THE REPLACED CARDIOHELP DEVICE BEFORE THE PATIENT PASSED AWAY. THE GETINGE SERVICE TECHNICIAN CONFIRMED THAT THERE WAS NO FAULT ON THE SECOND HARDWARE DEVICE NOR ON THE DISPOSABLE USED. THE PATIENT PASSED HAD MASSIVE BLEEDING, WHICH WAS BEFORE AND AFTER COMING ON ECMO. FURTHER IT WAS DIFFICULT TO CANNULATE THE PATIENT. THE PATIENT PASSED AWAY DUE TO BLEEDING AND LOW BLOOD PRESSURE. AS THE PATIENT PASSED AWAY AND THE CARDIOHELP DEVICE WAS REPLACED DURING TREATMENT, A REPORT IS REQUIRED. NEW INFORMATION RECEIVED ON (B)(6) 2025 THAT THE PATIENT DID NOT DIE DUE TO RELATED WRONG USER HANDLING ON THE DEVICE, BUT BECAUSE THE PATIENT WAS VERY SICK WHEN HE ARRIVED IN THE HOSPITAL. FURTHER THE LUCAS DEVICE WAS ON THE PATIENT FOR TWO HOURS AND LEAD TO A LOT IF INTERNAL BLEEDING. COMPLAINT ID#: (B)(4).Event Description: COMPLAINT ID: (B)(4). Manufacturer Narrative: IT WAS REPORTED THAT THE PATIENT ARRIVED IN THE HOSPITAL WITH A LUCAS PUMP PUMPING FOR ONE HOUR AND THE PATIENT HAS A CLOSED LEFT ANTERIOR DESCENDING (LAD). IT WAS HARD TO CANNULATE THE PATIENT AS THE PATIENT WAS BLEEDING ACCORDING TO THE CUSTOMER FROM ¿EVERYWHERE¿. THIS WAS CAUSED DUE TO THE LONG-TERM USE OF THE LUCAS PUMP. THE PATIENT WAS PUT ON ECMO. A CORONARY ARTERIOGRAPHY (KAG) WAS DONE AND THE LAD WAS OPENED. THE CUSTOMER WANTED TO DO A SMALL BYPASS WITH THE CANNULATION TO SECURE THE BLOOD FLOW TO THE LEG AND THEN THE STAFF CLAMPED THE HLS SET. A DEVICE DEFECTIVE ERROR MESSAGE OCCURRED, AS THE CUSTOMER HAS REMOVED THE SET WHILE THE PUMP WAS RUNNING. THE PATIENT WAS HAND CRANKED AND THE CARDIOHELP DEVICE WAS REPLACED. NEW INFORMATION WAS RECEIVED ON 2025-06-02 THAT THE PATIENT REMAINED FOR 3,5 HOURS ON THE REPLACED CARDIOHELP DEVICE BEFORE THE PATIENT PASSED AWAY. THE GETINGE SERVICE TECHNICIAN CONFIRMED THAT THERE WAS NO FAULT ON THE SECOND HARDWARE DEVICE NOR ON THE DISPOSABLE USED. THE PATIENT PASSED HAD MASSIVE BLEEDING, WHICH WAS BEFORE AND AFTER COMING ON ECMO. FURTHER IT WAS DIFFICULT TO CANNULATE THE PATIENT. THE PATIENT PASSED AWAY DUE TO BLEEDING AND LOW BLOOD PRESSURE. PATIENT DEATH WAS REPORTED. AS THE CARDIOHELP DEVICE WAS REPLACED DURING TREATMENT, A REPORT IS REQUIRED. NEW INFORMATION RECEIVED ON 2025-06-04 THAT THE PATIENT DID NOT DIE DUE TO RELATED WRONG USER HANDLING ON THE DEVICE, BUT BECAUSE THE PATIENT WAS VERY SICK WHEN HE ARRIVED IN THE HOSPITAL. FURTHER THE LUCAS DEVICE WAS ON THE PATIENT FOR TWO HOURS AND LEAD TO A LOT IF INTERNAL BLEEDING. A GETINGE FIELD SERVICE TECHNICIAN (FST) WAS SENT FOR INVESTIGATION ON 2025-05-27. NO PART WAS REPLACED. THE FAILURE OCCURRED DUE TO AN USER ERROR. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE LOG FILES OF THE REPORTED CARDIOHELP DEVICE WERE REVIEWED AND THE DEVICE DEFECTIVE COULD BE CONFIRMED ON THE DATE OF EVENT. FURTHER THE GETINGE FST CONFIRMED THAT THE USER ERROR COULD BE ADDITIONALLY CONFIRMED IN THE LOGFILES. THE ROOT CAUSE OF THE DEVICE DEFECTIVE ERROR IS AN USER RELATED FAULT CONFIRMED BY THE GETINGE SERVICE AND SALES UNIT (REFER TO COMMUNICATION GRID). THE DISPOSABLE WAS DETACHED WITHOUT REDUCING THE RPM. THE GETINGE FST CONFIRMED THAT THE HOSPITAL STAFF WILL RECEIVE AN TRAINING ON HOW TO HANDLE THE CARDIOHELP IN THIS KIND OF SITUATIONS. (REFER TO SERVICE ORDER). A MEDICAL REVIEW WAS PERFORMED ON 2025-07-02 WITH FOLLOWING CONCLUSION: "AFTER REVIEWING ALL AVAILABLE STATEMENTS AND CLINICAL INFORMATION, THE FOLLOWING CRITICAL SITUATION WAS DESCRIBED BY THE CUSTOMER: AFTER IMPLANTATION OF A VA-ECMO, DUE TO A PROLONGED CARDIAC ARREST WITH A 49-YEAR-OLD PATIENT, THE CLINICAL TEAM DECIDED TO CONNECT AN ANTEGRADE CANNULA TO PROVIDE SUFFICIENT PERFUSION OF THE LEG. THE CUSTOMER STATES, THAT THE CARDIOHELP-I AND THE USED HLS-SET WERE PRIMED CORRECTLY AND WORKED AS INTENDED, AFTER SUBSTITUTION OF VOLUME, DUE TO HYPOVOLEMIA. FOR THE INTEGRATION OF THIS CANNULA TO THE ESTABLISHED CIRCUIT, THE SYSTEM WAS CLAMPED, MAINTAINING THE CARDIOHELP-I TO APPROX. 3000RPM, AND THE NEW CANNULA ATTACHED TO THE ARTERIAL LINE. AFTER OPENING THE CLAMPS, THE CIRCULATION DID NOT RESTART. SO, THE CUSTOMER IMMEDIATELY CHANGED TO THE EMERGENCY DRIVE AND AFTERWARDS USED A SECOND CARDIOHELP-I TO REESTABLISH CONTINUOUS PERFUSION OF THE PATIENT. AFTER REVIEWING THE SERVICE POOL DATA AND THE SERVICE-REPORT IT IS CONFIRMED THAT A BACKFLOW OCCURRED, AS A ¿BACKFLOW PREVENTION¿-ALARM WAS DETECTED ON (B)(6) 2025 AT 18:07. THE LOG-DATA FOLLOWING THIS ALARM SUGGEST THAT THE HLS-SET WAS REMOVED FROM THE CARDIOHELP-I AND NO RESET OF THE INITIAL ALARM WAS PERFORMED. COMPARING THE INFORMATION OF THE SERVICE POOL WITH THE AVAILABLE DATA FROM THE USER POOL, IT CAN BE SEEN, THAT THE CUSTOMER UNLOCKED THE DEVICE MANUALLY AT 18:07, THE GLOBAL OVERDRIVE WAS ACTIVATED AT 18:10. THE BACK FLOW PREVENTION AND ZERO FLOW WERE SWITCHED OFF AT THE SAME TIME. THE REASON FOR THIS BACKFLOW ALARM REMAINS UNCLEAR. WHEN LOOKING AT POSSIBLE CAUSES FOR THE DETECTED BACKFLOW, TECHNICAL ISSUES CAN BE RULED OUT, AS THE SERVICE REPORT DOES NOT SHOW ANY DEFICIENCIES AND THE LOG DATA ALSO GIVE NO REASON TO BELIEVE, THAT THE FLOW-BUBBLE-SENSOR WAS POSITIONED FALSELY PRIOR TO THE IMPLEMENTATION OF THE DISTAL REPERFUSION CANNULA. GIVEN THE INFORMATION OF THE CUSTOMER, THAT THE CARDIOHELP-I WAS RUNNING AT 3000RPM, WHEN CLAMPED THERE ARE SOME POSSIBLE SITUATIONS, CREATING A BACKFLOW. 1) SUDDEN CLAMPING OF LINES WHILE THE PUMP WAS ACTIVE. CLAMPING THE VENOUS OR ARTERIAL LINES WHILE THE PUMP IS STILL OPERATING CAN RESULT IN SUDDEN PRESSURE DIFFERENTIALS AND TURBULENT FLOW, POTENTIALLY LEADING TO TRANSIENT NEGATIVE PRESSURES AT THE INLET OR OUTLET. THIS MAY CAUSE CAVITATION OR CONDITIONS INTERPRETED BY THE SYSTEM AS RISK FOR BACKFLOW. 2) TRANSIENT LOSS OF PRELOAD DUE TO BLEEDING OR AIR IN THE CIRCUIT. SEVERE HYPOVOLEMIA OR ENTRAINMENT OF AIR DUE TO MASSIVE BLEEDING CAN LEAD TO LOSS OF VENOUS RETURN (PRELOAD) INTO THE PUMP HEAD. EVEN AT HIGH RPMS, THE LACK OF SUFFICIENT VOLUME WOULD RESULT IN FLOW COLLAPSE, POTENTIALLY TRIGGERING BACKFLOW PREVENTION. 3) INCOMPLETE OCCLUSION OF THE ARTERIAL LINE OR CANNULA MALPOSITION. IMPROPERLY POSITIONED OR PARTIALLY OCCLUDED ARTERIAL CANNULAS CAN REDUCE AFTERLOAD RESISTANCE IN THE CIRCUIT, POTENTIALLY ALLOWING FOR RETROGRADE FLOW OR SYSTEM INSTABILITY. AS THE CUSTOMER STATED THAT THEY CLAMPED BOTH LINES (VENOUS AND ARTERIAL, THIS SEEMS TO BE LESS POSSIBLE) 4) MASSIVE BLEEDING AND RAPID VOLUME SHIFTS LARGE, RAPID SHIFTS IN INTRAVASCULAR VOLUME CAN DESTABILIZE INLET PRESSURE, PARTICULARLY DURING ECMO SUPPORT. THIS MAY AFFECT PUMP HEAD DYNAMICS AND POTENTIALLY INTERFERE WITH THE FLOW SENSING SYSTEM, TRIGGERING PROTECTIVE RESPONSES. (THIS MAY HAVE TRIGGERED A LOW PRESSURE ALARM ON THE PVEN BUT IS NOT SHOWN IN THE LOG DATA) SEEMINGLY, THE MOST LIKELY ROOT CAUSE, ACCORDING TO THE DATA PROVIDED BY THE CUSTOMER, SERVICE TECHNICIAN, AND LOG DATA WAS THE SUDDEN CLAMPING OF LINES WHILE THE PUMP WAS ACTIVE AT HIGH RPM IN COMBINATION WITH THE PATIENT¿S UNDERLYING HYPOVOLEMIA FROM MASSIVE BLEEDING. IT IS ASSUMED THAT THIS COMBINATION LED TO A TRANSIENT INSTABILITY IN THE CIRCUIT PRESSURES, CONTRIBUTING AN ACTUAL BACKFLOW SITUATION IN WHICH THE CARDIOHELP-I PERFORMED AS INTENDED AND THE BACKFLOW PREVENTION WAS ACTIVATED. GIVEN THE DATA PROVIDED, NO DEFINITE ROOT CAUSE FOR THE ACTIVATION OF THE BACKFLOW PREVENTION CAN BE DRAWN. THE IFU DEFINES A BACKFLOW AS A NEGATIVE FLOW FOR MORE THAN 3 SECONDS, CAUSING A ¿BACKFLOW PREVENTION¿ ALARM AND ACTIVATION OF THE ZERO-FLOW MODE. BY ACTIVATING THE ZERO-FLOW MODE THE CARDIOHELP-I AIMS TO MAINTAIN A FLOW OF 0 L/MIN (SEE IFU GRABS BELOW). WHEN THE ZERO-FLOW MODE IS IN OPERATION, IT IS CRUCIAL TO FOLLOW THE APPROPRIATE PROCEDURE ALLOWING THE CARDIOHELP-I TO LEAVE THE MODE AND RESTART THE PERFUSION. THIS SPECIFIC PROCEDURE INCLUDES THE REDUCTION OF THE FLOWRATE TO ZERO, PRESSING THE SAFETY BUTTON, AS WELL AS THE ZERO-FLOW BUTTON. REVIEWING THE INFORMATION PROVIDED BY THE CUSTOMER, NO CLEAR CAUSE FOR THIS BACKFLOW COULD BE IDENTIFIED. ANOTHER POSSIBILITY LEADING TO A BACKFLOW PREVENTION ALARM FOLLOWING ZERO-FLOW MODE ACTIVATION MAY BE INCORRECT VALUES FROM THE FLOW/BUBBLE SENSOR OR THE LEVEL SENSOR. HOWEVER, THE INSPECTION BY GETINGE FIELD SERVICES SUGGESTS THAT DEVICE WAS IN PROPER WORKING ORDER. WHEN CONSIDERING THE WHOLE CLINICAL SITUATION, PROLONGED RESUSCITATION, ACUTE LAD-OCCLUSION AND STENT-IMPLANTATION, UNCONTROLLED SEVERE HEMORRHAGE, AND LOW LVEF, THE CUSTOMER WAS LIKELY CONFRONTED WITH A HIGHLY CRITICAL AND ILL PATIENT. IT APPEARS THE ONGOING EXSANGUINATION AND PRONOUNCES DIFFUSE COAGULOPATHY MAY HAVE BEEN COMPLICATING SUPPORT, LEADING TO THE EXPIRATION OF THE PATIENT. IN CONCLUSION, THE CARDIOHELP SYSTEM APPEARS TO HAVE FUNCTIONED WITHIN ITS INTENDED DESIGN THROUGHOUT THE EVENT. BASED ON THE AVAILABLE DATA AND POST-EVENT ANALYSIS, NO TECHNICAL MALFUNCTION WAS IDENTIFIED. THE ACTIVATION OF SAFETY FEATURES - SUCH AS THE BACKFLOW PREVENTION - WAS CONSISTENT WITH CARDIOHELP-I DESIGN AND SENSOR INPUT. IT IS POSSIBLE THAT THE CRITICAL CLINICAL CONDITION OF THE PATIENT, INCLUDING PROLONGED CARDIAC ARREST, ONGOING MASSIVE HEMORRHAGE, AND THE COMPLEXITY OF THE PROCEDURAL ENVIRONMENT, MAY HAVE CONTRIBUTED TO THE OBSERVED CIRCUIT INSTABILITY AND ULTIMATELY INFLUENCING THE CLINICAL OUTCOME." ACCORDING TO THE INSTRUCTION FOR USE (IFU) OF THE CARDIOHELP DEVICE (CHAPTER ¿HIGH PRIORITY¿) IT IS STATED THAT THIS ERROR MESSAGE INFORM THE USER OF AN ERROR THAT REQUIRE THE SERVICE TO DIAGNOSE THE FAULT. THE CARDIOHELP SHOULD BE EXCHANGED AS QUICKLY AS POSSIBLE, IF THE FAILURE OCCURS DURING PATIENT TREATMENT. THE PERFUSION ON THE CARDIOHELP SHOULD CONTINUALLY BE CHECKED, AS WELL AS THE MONITORING OF THE PATIENT. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED ON 2025-06-02 FOR THE PERIOD OF 2018-02-15 TO 2025-05-27. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED PRODUCT AND FAILURE. THERE IS NO INDICATION THAT MANUFACTURING ISSUES OCCURRED DURING THIS TIME, THUS PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE DEVICE WAS MANUFACTURED ON 2018-02-15. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. ACCORDING TO THE RISK REVIEW THE REPORTED FAILURE IS BELOW THE ACCEPTANCE THRESHOLD ACCORDING TO THE RISK MANAGEMENT PLAN OF THE CARDIOHELP DEVICE. BASED ON THE RESULTS THE REPORTED FAILURE "DEVICE DEFECTIVE ERROR" COULD BE CONFIRMED, BUT WAS NOT A DEVICE RELATED FAULT. THIS COMPLAINT WAS FOUND IN THE DATABASE OF CUSTOMER COMPLAINTS FOR THE CARDIOHELP DEVICE AS A SINGLE EVENT (TIMEFRAME FROM 2024-05-27 TILL 2025-05-27). THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY'S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. Manufacturer Narrative: A GETINGE FIELD SERVICE TECHNICIAN (FST) WAS SENT FOR INVESTIGATION ON 2025-05-27. NO PART WAS REPLACED. THE FAILURE OCCURRED DUE TO AN USER ERROR. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE LOG FILES OF THE REPORTED CARDIOHELP DEVICE WERE REVIEWED AND THE DEVICE DEFECTIVE COULD BE CONFIRMED ON THE DATE OF EVENT. FURTHER THE GETINGE FST CONFIRMED THAT THE USER ERROR WAS VISIBLE IN THE LOGFILES. THE GETINGE FST CONFIRMED THAT THE HOSPITAL STAFF WILL RECEIVE AN TRAINING ON HOW TO HANDLE THE CARDIOHELP IN THIS KIND OF SITUATIONS. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE DENMARK MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K133598. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR CARDIOHELP WITH CATALOG NUMBER 701048012.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22139867&pc=OZD | 1220648-2025-28962 | 2025-05-27 04:00 | Death | ABIOMED, INC. | 2025-06-04 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON IMPELLA CP SUPPORT WITH INTRA-AORTIC BALLOON PUMP (IABP). DURING SUPPORT, THE PATIENT HAD FREQUENT FATAL ARRHYTHMIAS. THE CONTINUATION OF TREATMENT WAS ABANDONED AND THE PATIENT EXPIRED. THE PATIENT'S OVERALL CONDITION WAS POOR TO BEGIN WITH, AND THE PATIENT HAD A STRONG TENDENCY TOWARD COLLAPSE AND INFECTION PRIOR TO IMPELLA BEING INSERTED. THE PATIENT HAD BEEN SUPPORTED WITH IABP FOR A LONG PERIOD OF TIME. IT WAS STATED THAT THE ARRHYTHMIA WAS NOT CAUSED BY THE IMPELLA, HOWEVER IT CANNOT BE CONFIRMED THAT THE IMPELLA DID NOT CONTRIBUTE TO THE PATIENT'S ARRHYTHMIAS. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22139869&pc=DSQ | 2916596-2025-03625 | 2025-05-16 04:00 | Death | THORATEC CORPORATION | 2025-06-04 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO REFRACTORY MULTI SYSTEM ORGAN FAILURE. THIS WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE CAUSE OF DEATH WAS REFRACTORY, PROGRESSIVE MULTI SYSTEM ORGAN FAILURE. THERE WERE NO DEVICE ISSUES THAT LEAD TO MULTI SYSTEM ORGAN FAILURE. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS RUNNING AS EXPECTED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22139982&pc=OZD | 1220648-2025-28963 | 2025-05-09 04:00 | Death | ABIOMED, INC. | 2025-06-04 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT REPORTEDLY CALLED EMERGENCY MEDICAL SERVICES DUE TO NAUSEA AND VOMITING. THE PATIENT AS TAKEN TO THE NEAREST EMERGENCY ROOM. CODE ST- SEGMENT ELEVATION MYOCARDIAL INFARCTION WAS CALLED ON THE PATIENT. THE PATIENT DETERIORATED AND THE DECISION WAS MADE TO PLACE AN IMPELLA CP. DURING IMPELLA CP SUPPORT, THE PATIENT EXPERIENCED ATRIAL FIBRILATION THAT WAS IN THE 120'S. ADDITIONAL INFORMATION NOTED CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR. IT APPEARS SOME TIME AFTER THE WIRE WAS INTRODUCED, THE IMPELLA WAS IMPLANTED, AND THE PATIENT DEVELOPED ATRIAL FIBRILLATION. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22140141&pc=OZP | 2032227-2025-185631 | 2023-01-17 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-04 04:00 | OZP | PUMP MMT-1780KPK 670G PATHWAY BLACK MG | Break | Hyperglycemia | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2023 AT HOME. MEDTRONIC LEGAL RECEIVED INFORMATION FROM THE ATTORNEY OF THE FAMILY ON (B)(6) 2025. MEDTRONIC RECEIVED NOTICE OF A DEVICE/PRODUCT LEGAL HOLD NOTIFICATION, AND THE REPORTER ALLEGES THAT THE USE OF ONE OR MORE MEDTRONIC MINIMED 600 SERIES INSULIN PUMPS WITH A DAMAGED, MISSING, OR BROKEN RETAINER RING CAUSED THE CLAIMANT TO SUFFER PERSONAL INJURIES, AND THE CUSTOMER WAS HOSPITALIZED. THE ALLEGATION DID NOT SPECIFY THE PUMP IDENTIFIER (SERIAL NUMBER) AND THEREFORE ALL 600 SERIES INSULIN PUMPS IN POSSESSION OF THE CLAIMANT, INCLUDING THIS PUMP, ARE CONSIDERED POTENTIALLY WITHIN THE SCOPE OF THE REPORT PENDING CONFIRMATION OF THE AFFECTED SERIAL NUMBER(S). ALL RELATED REPORTS WILL BE UPDATED ACCORDINGLY WHEN AND IF THE AFFECTED SERIAL NUMBER IS IDENTIFIED. THE CAUSE OF DEATH WAS DIABETIC KETOACIDOSIS. IT WAS UNKNOWN WHETHER THE INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-1780KL, MMT-1780KPK, UNOMEDICAL. TROUBLESHOOTING WAS NOT PERFORMED. IT WAS UNKNOWN WHETHER THE CUSTOMER USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT, AND IT WAS UNKNOWN WHETHER THE SMARTGUARD/AUTO MODE OF THE INSULIN PUMP WAS USED OR NOT. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KL. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KPK. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. THIS IS 2 OF 2 MEDWATCH REPORTS REGARDING THIS EVENT. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22140168&pc=OZP | 2032227-2025-185630 | 2023-01-17 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-04 04:00 | OZP | 670G INSULIN PUMP MMT-1780KL | Break | Hyperglycemia | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2023 AT HOME. MEDTRONIC LEGAL RECEIVED INFORMATION FROM THE ATTORNEY OF THE FAMILY ON (B)(6) 2025. MEDTRONIC RECEIVED NOTICE OF A DEVICE/PRODUCT LEGAL HOLD NOTIFICATION, AND THE REPORTER ALLEGES THAT THE USE OF ONE OR MORE MEDTRONIC MINIMED 600 SERIES INSULIN PUMPS WITH A DAMAGED, MISSING, OR BROKEN RETAINER RING CAUSED THE CLAIMANT TO SUFFER PERSONAL INJURIES, AND THE CUSTOMER WAS HOSPITALIZED. THE ALLEGATION DID NOT SPECIFY THE PUMP IDENTIFIER (SERIAL NUMBER) AND THEREFORE ALL 600 SERIES INSULIN PUMPS IN POSSESSION OF THE CLAIMANT, INCLUDING THIS PUMP, ARE CONSIDERED POTENTIALLY WITHIN THE SCOPE OF THE REPORT PENDING CONFIRMATION OF THE AFFECTED SERIAL NUMBER(S). ALL RELATED REPORTS WILL BE UPDATED ACCORDINGLY WHEN AND IF THE AFFECTED SERIAL NUMBER IS IDENTIFIED. THE CAUSE OF DEATH WAS DIABETIC KETOACIDOSIS. IT WAS UNKNOWN WHETHER THE INSULIN PUMP WAS WORN AT THE TIME OF PASSING. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-332A, MMT-1780KL, MMT-1780KPK, UNOMEDICAL. TROUBLESHOOTING WAS NOT PERFORMED. IT WAS UNKNOWN WHETHER THE CUSTOMER USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT, AND IT WAS UNKNOWN WHETHER THE SMARTGUARD/AUTO MODE OF THE INSULIN PUMP WAS USED OR NOT. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KL. NO PRODUCT RETURN IS REQUIRED FOR MMT-1780KPK. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. THIS IS 1 OF 2 MEDWATCH REPORTS REGARDING THIS EVENT. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22140413&pc=CBK | MW5171059 | 2025-05-17 04:00 | Death | VIASYS (CAREFUSION) VELA / ZOLL PALM SPRINGS | 2025-06-04 04:00 | CBK | VIASYS VELA VENTILATOR | Defective Alarm | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED BY LISA THAT "I HAVE AN EVENT WHERE THE VENT ALARMED WITH AN "INOPERABLE" STATUS AND ALARMS FOR APPROX 4 MINUTES PRIOR TO A PATIENT CODE."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22140467&pc=DTQ | 3013876692-2025-00043 | 2025-05-27 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-06-04 04:00 | DTQ | CARDIOHELP | Device Handling Problem | Hemorrhage/Bleeding; Asystole | K133598 |  | 1 | Event Description: THE EVENT OCCURRED IN DENMARK DURING TREATMENT. IT WAS REPORTED THAT THE PATIENT ARRIVED IN THE HOSPITAL WITH A LUCAS PUMP PUMPING FOR ONE HOUR AND THE PATIENT HAS A CLOSED LEFT ANTERIOR DESCENDING (LAD). IT WAS HARD TO CANNULATE THE PATIENT AS THE PATIENT WAS BLEEDING ACCORDING TO THE CUSTOMER FROM ¿EVERYWHERE¿. THIS WAS CAUSED DUE TO THE LONG-TERM USE OF THE LUCAS PUMP. THE PATIENT WAS PUT ON ECMO. A CORONARY ARTERIOGRAPHY (KAG) WAS DONE AND THE LAD WAS OPENED. THE CUSTOMER WANTED TO DO A SMALL BYPASS WITH THE CANNULATION TO SECURE THE BLOOD FLOW TO THE LEG AND THEN THE STAFF CLAMPED THE HLS SET. A DEVICE DEFECTIVE ERROR MESSAGE OCCURRED, AS THE CUSTOMER HAS REMOVED THE SET WHILE THE PUMP WAS RUNNING. THE PATIENT WAS HAND CRANKED AND THE CARDIOHELP DEVICE WAS REPLACED. NEW INFORMATION WAS RECEIVED ON (B)(6) 2025 THAT THE PATIENT REMAINED FOR 3,5 HOURS ON THE REPLACED CARDIOHELP DEVICE BEFORE THE PATIENT PASSED AWAY. THE GETINGE SERVICE TECHNICIAN CONFIRMED THAT THERE WAS NO FAULT ON THE SECOND HARDWARE DEVICE NOR ON THE DISPOSABLE USED. THE PATIENT PASSED HAD MASSIVE BLEEDING, WHICH WAS BEFORE AND AFTER COMING ON ECMO. FURTHER IT WAS DIFFICULT TO CANNULATE THE PATIENT. THE PATIENT PASSED AWAY DUE TO BLEEDING AND LOW BLOOD PRESSURE. AS THE PATIENT PASSED AWAY AND THE CARDIOHELP DEVICE WAS REPLACED DURING TREATMENT, A REPORT IS REQUIRED. NEW INFORMATION RECEIVED ON (B)(6) 2025 THAT THE PATIENT DID NOT DIE DUE TO RELATED WRONG USER HANDLING ON THE DEVICE, BUT BECAUSE THE PATIENT WAS VERY SICK WHEN HE ARRIVED IN THE HOSPITAL. FURTHER THE LUCAS DEVICE WAS ON THE PATIENT FOR TWO HOURS AND LEAD TO A LOT IF INTERNAL BLEEDING. COMPLAINT ID# (B)(4).Event Description: COMPLAINT ID (B)(4). Manufacturer Narrative: A GETINGE FIELD SERVICE TECHNICIAN (FST) WAS SENT FOR INVESTIGATION ON 2025-05-27. NO PART WAS REPLACED. THE FAILURE OCCURRED DUE TO AN USER ERROR. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE LOG FILES OF THE REPORTED CARDIOHELP DEVICE WERE REVIEWED AND THE DEVICE DEFECTIVE COULD BE CONFIRMED ON THE DATE OF EVENT. FURTHER THE GETINGE FST CONFIRMED THAT THE USER ERROR WAS VISIBLE IN THE LOGFILES. THE GETINGE FST CONFIRMED THAT THE HOSPITAL STAFF WILL RECEIVE AN TRAINING ON HOW TO HANDLE THE CARDIOHELP IN THIS KIND OF SITUATIONS. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE DENMARK MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K133598. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR CARDIOHELP WITH CATALOG NUMBER 701048012. Manufacturer Narrative: IT WAS REPORTED THAT THE PATIENT ARRIVED IN THE HOSPITAL WITH A LUCAS PUMP PUMPING FOR ONE HOUR AND THE PATIENT HAS A CLOSED LEFT ANTERIOR DESCENDING (LAD). IT WAS HARD TO CANNULATE THE PATIENT AS THE PATIENT WAS BLEEDING ACCORDING TO THE CUSTOMER FROM ¿EVERYWHERE¿. THIS WAS CAUSED DUE TO THE LONG-TERM USE OF THE LUCAS PUMP. THE PATIENT WAS PUT ON ECMO. A CORONARY ARTERIOGRAPHY (KAG) WAS DONE AND THE LAD WAS OPENED. THE CUSTOMER WANTED TO DO A SMALL BYPASS WITH THE CANNULATION TO SECURE THE BLOOD FLOW TO THE LEG AND THEN THE STAFF CLAMPED THE HLS SET. A DEVICE DEFECTIVE ERROR MESSAGE OCCURRED, AS THE CUSTOMER HAS REMOVED THE SET WHILE THE PUMP WAS RUNNING. THE PATIENT WAS HAND CRANKED AND THE CARDIOHELP DEVICE WAS REPLACED. NEW INFORMATION WAS RECEIVED ON 2025-06-02 THAT THE PATIENT REMAINED FOR 3,5 HOURS ON THE REPLACED CARDIOHELP DEVICE BEFORE THE PATIENT PASSED AWAY. THE GETINGE SERVICE TECHNICIAN CONFIRMED THAT THERE WAS NO FAULT ON THE SECOND HARDWARE DEVICE NOR ON THE DISPOSABLE USED. THE PATIENT PASSED HAD MASSIVE BLEEDING, WHICH WAS BEFORE AND AFTER COMING ON ECMO. FURTHER IT WAS DIFFICULT TO CANNULATE THE PATIENT. THE PATIENT PASSED AWAY DUE TO BLEEDING AND LOW BLOOD PRESSURE. PATIENT DEATH WAS REPORTED. AS THE CARDIOHELP DEVICE WAS REPLACED DURING TREATMENT, A REPORT IS REQUIRED. NEW INFORMATION RECEIVED ON 2025-06-04 THAT THE PATIENT DID NOT DIE DUE TO RELATED WRONG USER HANDLING ON THE DEVICE, BUT BECAUSE THE PATIENT WAS VERY SICK WHEN HE ARRIVED IN THE HOSPITAL. FURTHER THE LUCAS DEVICE WAS ON THE PATIENT FOR TWO HOURS AND LEAD TO A LOT IF INTERNAL BLEEDING. A GETINGE FIELD SERVICE TECHNICIAN (FST) WAS SENT FOR INVESTIGATION ON 2025-05-27. NO PART WAS REPLACED. THE FAILURE OCCURRED DUE TO AN USER ERROR. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE LOG FILES OF THE REPORTED CARDIOHELP DEVICE WERE REVIEWED AND THE DEVICE DEFECTIVE COULD BE CONFIRMED ON THE DATE OF EVENT. FURTHER THE GETINGE FST CONFIRMED THAT THE USER ERROR COULD BE ADDITIONALLY CONFIRMED IN THE LOGFILES. THE ROOT CAUSE OF THE DEVICE DEFECTIVE ERROR IS AN USER RELATED FAULT CONFIRMED BY THE GETINGE SERVICE AND SALES UNIT (REFER TO COMMUNICATION GRID). THE DISPOSABLE WAS DETACHED WITHOUT REDUCING THE RPM. THE GETINGE FST CONFIRMED THAT THE HOSPITAL STAFF WILL RECEIVE AN TRAINING ON HOW TO HANDLE THE CARDIOHELP IN THIS KIND OF SITUATIONS. (REFER TO SERVICE ORDER). A MEDICAL REVIEW WAS PERFORMED ON 2025-07-02 WITH FOLLOWING CONCLUSION: "AFTER REVIEWING ALL AVAILABLE STATEMENTS AND CLINICAL INFORMATION, THE FOLLOWING CRITICAL SITUATION WAS DESCRIBED BY THE CUSTOMER: AFTER IMPLANTATION OF A VA-ECMO, DUE TO A PROLONGED CARDIAC ARREST WITH A 49-YEAR-OLD PATIENT, THE CLINICAL TEAM DECIDED TO CONNECT AN ANTEGRADE CANNULA TO PROVIDE SUFFICIENT PERFUSION OF THE LEG. THE CUSTOMER STATES, THAT THE CARDIOHELP-I AND THE USED HLS-SET WERE PRIMED CORRECTLY AND WORKED AS INTENDED, AFTER SUBSTITUTION OF VOLUME, DUE TO HYPOVOLEMIA. FOR THE INTEGRATION OF THIS CANNULA TO THE ESTABLISHED CIRCUIT, THE SYSTEM WAS CLAMPED, MAINTAINING THE CARDIOHELP-I TO APPROX. 3000RPM, AND THE NEW CANNULA ATTACHED TO THE ARTERIAL LINE. AFTER OPENING THE CLAMPS, THE CIRCULATION DID NOT RESTART. SO, THE CUSTOMER IMMEDIATELY CHANGED TO THE EMERGENCY DRIVE AND AFTERWARDS USED A SECOND CARDIOHELP-I TO REESTABLISH CONTINUOUS PERFUSION OF THE PATIENT. AFTER REVIEWING THE SERVICE POOL DATA AND THE SERVICE-REPORT IT IS CONFIRMED THAT A BACKFLOW OCCURRED, AS A ¿BACKFLOW PREVENTION¿-ALARM WAS DETECTED ON THE (B)(6) 2025 AT 18:07. THE LOG-DATA FOLLOWING THIS ALARM SUGGEST THAT THE HLS-SET WAS REMOVED FROM THE CARDIOHELP-I AND NO RESET OF THE INITIAL ALARM WAS PERFORMED. COMPARING THE INFORMATION OF THE SERVICE POOL WITH THE AVAILABLE DATA FROM THE USER POOL, IT CAN BE SEEN, THAT THE CUSTOMER UNLOCKED THE DEVICE MANUALLY AT 18:07, THE GLOBAL OVERDRIVE WAS ACTIVATED AT 18:10. THE BACK FLOW PREVENTION AND ZERO FLOW WERE SWITCHED OFF AT THE SAME TIME. THE REASON FOR THIS BACKFLOW ALARM REMAINS UNCLEAR. WHEN LOOKING AT POSSIBLE CAUSES FOR THE DETECTED BACKFLOW, TECHNICAL ISSUES CAN BE RULED OUT, AS THE SERVICE REPORT DOES NOT SHOW ANY DEFICIENCIES AND THE LOG DATA ALSO GIVE NO REASON TO BELIEVE, THAT THE FLOW-BUBBLE-SENSOR WAS POSITIONED FALSELY PRIOR TO THE IMPLEMENTATION OF THE DISTAL REPERFUSION CANNULA. GIVEN THE INFORMATION OF THE CUSTOMER, THAT THE CARDIOHELP-I WAS RUNNING AT 3000RPM, WHEN CLAMPED THERE ARE SOME POSSIBLE SITUATIONS, CREATING A BACKFLOW. 1)SUDDEN CLAMPING OF LINES WHILE THE PUMP WAS ACTIVE. CLAMPING THE VENOUS OR ARTERIAL LINES WHILE THE PUMP IS STILL OPERATING CAN RESULT IN SUDDEN PRESSURE DIFFERENTIALS AND TURBULENT FLOW, POTENTIALLY LEADING TO TRANSIENT NEGATIVE PRESSURES AT THE INLET OR OUTLET. THIS MAY CAUSE CAVITATION OR CONDITIONS INTERPRETED BY THE SYSTEM AS RISK FOR BACKFLOW. 2)TRANSIENT LOSS OF PRELOAD DUE TO BLEEDING OR AIR IN THE CIRCUIT. SEVERE HYPOVOLEMIA OR ENTRAINMENT OF AIR DUE TO MASSIVE BLEEDING CAN LEAD TO LOSS OF VENOUS RETURN (PRELOAD) INTO THE PUMP HEAD. EVEN AT HIGH RPMS, THE LACK OF SUFFICIENT VOLUME WOULD RESULT IN FLOW COLLAPSE, POTENTIALLY TRIGGERING BACKFLOW PREVENTION. 3)INCOMPLETE OCCLUSION OF THE ARTERIAL LINE OR CANNULA MALPOSITION. IMPROPERLY POSITIONED OR PARTIALLY OCCLUDED ARTERIAL CANNULAS CAN REDUCE AFTERLOAD RESISTANCE IN THE CIRCUIT, POTENTIALLY ALLOWING FOR RETROGRADE FLOW OR SYSTEM INSTABILITY. AS THE CUSTOMER STATED THAT THEY CLAMPED BOTH LINES (VENOUS AND ARTERIAL, THIS SEEMS TO BE LESS POSSIBLE) 4) MASSIVE BLEEDING AND RAPID VOLUME SHIFTS LARGE, RAPID SHIFTS IN INTRAVASCULAR VOLUME CAN DESTABILIZE INLET PRESSURE, PARTICULARLY DURING ECMO SUPPORT. THIS MAY AFFECT PUMP HEAD DYNAMICS AND POTENTIALLY INTERFERE WITH THE FLOW SENSING SYSTEM, TRIGGERING PROTECTIVE RESPONSES. (THIS MAY HAVE TRIGGERED A LOWPRESSURE ALARM ON THE PVEN BUT IS NOT SHOWN IN THE LOG DATA) SEEMINGLY, THE MOST LIKELY ROOT CAUSE, ACCORDING TO THE DATA PROVIDED BY THE CUSTOMER, SERVICE TECHNICIAN, AND LOG DATA WAS THE SUDDEN CLAMPING OF LINES WHILE THE PUMP WAS ACTIVE AT HIGH RPM IN COMBINATION WITH THE PATIENT¿S UNDERLYING HYPOVOLEMIA FROM MASSIVE BLEEDING. IT IS ASSUMED THAT THIS COMBINATION LED TO A TRANSIENT INSTABILITY IN THE CIRCUIT PRESSURES, CONTRIBUTING AN ACTUAL BACKFLOW SITUATION IN WHICH THE CARDIOHELP-I PERFORMED AS INTENDED AND THE BACKFLOW PREVENTION WAS ACTIVATED. GIVEN THE DATA PROVIDED, NO DEFINITE ROOT CAUSE FOR THE ACTIVATION OF THE BACKFLOW PREVENTION CAN BE DRAWN. THE IFU DEFINES A BACKFLOW AS A NEGATIVE FLOW FOR MORE THAN 3 SECONDS, CAUSING A ¿BACKFLOW PREVENTION¿ ALARM AND ACTIVATION OF THE ZERO-FLOW MODE. BY ACTIVATING THE ZERO-FLOW MODE THE CARDIOHELP-I AIMS TO MAINTAIN A FLOW OF 0 L/MIN (SEE IFU GRABS BELOW). WHEN THE ZERO-FLOW MODE IS IN OPERATION, IT IS CRUCIAL TO FOLLOW THE APPROPRIATE PROCEDURE ALLOWING THE CARDIOHELP-I TO LEAVE THE MODE AND RESTART THE PERFUSION. THIS SPECIFIC PROCEDURE INCLUDES THE REDUCTION OF THE FLOWRATE TO ZERO, PRESSING THE SAFETY BUTTON, AS WELL AS THE ZERO-FLOW BUTTON. REVIEWING THE INFORMATION PROVIDED BY THE CUSTOMER, NO CLEAR CAUSE FOR THIS BACKFLOW COULD BE IDENTIFIED. ANOTHER POSSIBILITY LEADING TO A BACKFLOW PREVENTION ALARM FOLLOWING ZERO-FLOW MODE ACTIVATION MAY BE INCORRECT VALUES FROM THE FLOW/BUBBLE SENSOR OR THE LEVEL SENSOR. HOWEVER, THE INSPECTION BY GETINGE FIELD SERVICES SUGGESTS THAT DEVICE WAS IN PROPER WORKING ORDER. WHEN CONSIDERING THE WHOLE CLINICAL SITUATION, PROLONGED RESUSCITATION, ACUTE LAD-OCCLUSION AND STENT-IMPLANTATION, UNCONTROLLED SEVERE HEMORRHAGE, AND LOW LVEF, THE CUSTOMER WAS LIKELY CONFRONTED WITH A HIGHLY CRITICAL AND ILL PATIENT. IT APPEARS THE ONGOING EXSANGUINATION AND PRONOUNCES DIFFUSE COAGULOPATHY MAY HAVE BEEN COMPLICATING SUPPORT, LEADING TO THE EXPIRATION OF THE PATIENT. IN CONCLUSION, THE CARDIOHELP SYSTEM APPEARS TO HAVE FUNCTIONED WITHIN ITS INTENDED DESIGN THROUGHOUT THE EVENT. BASED ON THE AVAILABLE DATA AND POST-EVENT ANALYSIS, NO TECHNICAL MALFUNCTION WAS IDENTIFIED. THE ACTIVATION OF SAFETY FEATURES - SUCH AS THE BACKFLOW PREVENTION - WAS CONSISTENT WITH CARDIOHELP-I DESIGN AND SENSOR INPUT. IT IS POSSIBLE THAT THE CRITICAL CLINICAL CONDITION OF THE PATIENT, INCLUDING PROLONGED CARDIAC ARREST, ONGOING MASSIVE HEMORRHAGE, AND THE COMPLEXITY OF THE PROCEDURAL ENVIRONMENT, MAY HAVE CONTRIBUTED TO THE OBSERVED CIRCUIT INSTABILITY AND ULTIMATELY INFLUENCING THE CLINICAL OUTCOME." ACCORDING TO THE INSTRUCTION FOR USE (IFU) OF THE CARDIOHELP DEVICE (CHAPTER ¿HIGH PRIORITY¿) IT IS STATED THAT THIS ERROR MESSAGE INFORM THE USER OF AN ERROR THAT REQUIRE THE SERVICE TO DIAGNOSE THE FAULT. THE CARDIOHELP SHOULD BE EXCHANGED AS QUICKLY AS POSSIBLE, IF THE FAILURE OCCURS DURING PATIENT TREATMENT. THE PERFUSION ON THE CARDIOHELP SHOULD CONTINUALLY BE CHECKED, AS WELL AS THE MONITORING OF THE PATIENT. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED ON 2025-06-02 FOR THE PERIOD OF 2018-02-15 TO 2025-05-27. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED PRODUCT AND FAILURE. THERE IS NO INDICATION THAT MANUFACTURING ISSUES OCCURRED DURING THIS TIME, THUS PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE DEVICE WAS MANUFACTURED ON 2018-02-15. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. ACCORDING TO THE RISK REVIEW THE REPORTED FAILURE IS BELOW THE ACCEPTANCE THRESHOLD ACCORDING TO THE RISK MANAGEMENT PLAN OF THE CARDIOHELP DEVICE. BASED ON THE RESULTS THE REPORTED FAILURE "DEVICE DEFECTIVE ERROR" COULD BE CONFIRMED, BUT WAS NOT A DEVICE RELATED FAULT. THIS COMPLAINT WAS FOUND IN THE DATABASE OF CUSTOMER COMPLAINTS FOR THE CARDIOHELP DEVICE AS A SINGLE EVENT (TIMEFRAME FROM 2024-05-27 TILL 2025-05-27). THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY'S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22140603&pc=MHX | 9610816-2025-000454 | 2025-05-03 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-06-04 04:00 | MHX | INTELLIVUE MX800 PATIENT MONITOR | Insufficient Information | Insufficient Information | K150310 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT PASSED AWAY WHILE CONNECTED TO THE MONITOR. IT WAS INDICATED THE MONITOR WAS DEFECTIVE OR IMPAIRED; HOWEVER, THE SPECIFIC DEVICE ISSUE WAS NOT PROVIDED. THE DEVICE WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. Manufacturer Narrative: A PHILIPS SENIOR FIELD SPECIALIST WENT TO THE CUSTOMER'S SITE AND TESTED THE DEVICE. THE SPECIFIED ISSUE COULD NOT BE REPRODUCED. ALARM FUNCTIONS WERE CHECKED AND LOGS, ALARM AUDIT, CONFIGURATIONS WERE PULLED FOR FURTHER INVESTIGATION. NO OTHER PROBLEMS WERE DETECTED DURING THE TEST. THE DEVICE IS FULLY FUNCTIONAL. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. E1; REPORTER INSTITUTION PHONE NUMBER (B)(6). E1; REPORTER PHONE NUMBER (B)(6). THE MANUFACTURING DATE FOR THE REPORTED DEVICE IS PRIOR TO (B)(6) 2016 SO NO UDI REQUIRED. Manufacturer Narrative: THE AVAILABLE INFORMATION WAS REVIEWED BY THE PHILIPS PMS CLINICAL EXPERT. THE INFORMATION INDICATED AN ECG OR SPO2 RELATED ALARM WAS LIKELY ACKNOWLEDGED AND THERE WAS A DELAY IN TREATMENT. THE PATIENT WAS FOUND WITH ECG AND SPO2 UNPLUGGED. A LOG REVIEW REVEALED YELLOW ALARMS FOLLOWED BY TECHNICAL INOP ALARMS FOR 'LEADS OFF' WERE OCCURRING PRIOR TO THE PATIENT'S DEATH. THE ALARMS WERE LATER ACKNOWLEDGED AT THE CENTRAL STATION. THE PATIENT MONITOR AND CENTRAL STATION GENERATED ALL ALARMS ACCORDING TO THE CONFIGURATION. BASED ON THIS INFORMATION, THERE DOES NOT APPEAR TO BE A DEVICE MALFUNCTION THAT CAUSED OR CONTRIBUTED TO THE REPORTED DEATH; HOWEVER, USER ERRORS RELATED TO ALARM MANAGEMENT AND CLINICAL WORKFLOW MAY HAVE BEEN FACTORS. ADDITIONAL INFORMATION HAS BEEN REQUESTED, A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: D2B FDA PRODUCT CODE WAS CORRECTED FROM DSI TO MHX. D4 MODEL NUMBER WAS CORRECTED FROM 865240 TO INTELLIVUE PATIENT MONITOR MX800. H6 CODING WAS UPDATED. NO ADDITIONAL INFORMATION WAS RECEIVED, 9610816-2025-000454-001 SHOULD HAVE BEEN THE FINAL REPORT. THE INVESTIGATION RESULTS WERE REPORTED IN THE INITIAL REPORT AND FOLLOW-UP REPORT 9610816-2025-000454-001.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22140606&pc=BZD | 2518422-2025-046566 | 2024-02-20 05:00 | Death | RESPIRONICS, INC. | 2025-06-04 04:00 | BZD | DS2ADV AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DS2ADV AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22140824&pc=OZD | 1220648-2025-28971 | 2025-05-11 04:00 | Death | ABIOMED, INC. | 2025-06-04 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WITH HEART FAILURE/CARDIOGENIC SHOCK WAS PLACED ON IMPELLA CP FOR SUPPORT THROUGH A PLANNED VALVE REPLACEMENT. DURING SUPPORT, IT WAS NOTED THAT A PATIENT HAD AN ARREST. IN ADDITION, HEPARIN WAS STOPPED DUE TO A GASTROINTESTINAL BLEED. THE FAMILY DECIDED ON COMFORT CARE AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22141114&pc=OZD | 1220648-2025-28973 | 2025-05-09 04:00 | Death | ABIOMED, INC. | 2025-06-04 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT'S RIGHT LEG BECAME MOLTED AND COLD. THE PATIENT BECAME MOTTLED IN ALL EXTREMITIES AND ALSO CHEST AND BACK. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST SYSTEM INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22141124&pc=LNQ | 3027386225-2025-00057 | 2021-08-17 04:00 | Death | ENTERRA MEDICAL, INC. | 2025-06-04 04:00 | LNQ | ENTERRA II IPG | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | H990014 |  | 1 | Event Description: PATIENT DEATH - DEVICE EXPLANTED AS PART OF AUTOPSY. CAUSE OF DEATH WAS COMPLICATIONS OF POSTURAL ORTHOSTATIC TACHYCARDIA SYNDROME (NOT ENTERRA DEVICE RELATED) PRODUCT RETURNED TO MEDTRONIC FOR ANALYSIS. CHECKED IN AT MDT (B)(6) 2025 MDT RECORD ID: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22142085&pc=DSQ | 2916596-2025-03507 | 2025-05-20 04:00 | Death | THORATEC CORPORATION | 2025-06-04 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON HOSPICE. THE PATIENT WAS FAILING TO THRIVE, SO THEY WERE PLACED ON HOSPICE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. D, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22142154&pc=NIK | 2017865-2025-75407 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NIK | GALLANT VR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED UPON PRODUCT RECEIPT THAT THE PATIENT WITH THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR SYSTEM HAD PASSED AWAY DUE TO CARDIAC AMYLOIDOSIS, CARDIAC FAILURE, AND MULTI-ORGAN FAILURE. NO FURTHER INFORMATION WAS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22142155&pc=OZD | 1220648-2025-28981 | 2025-05-13 04:00 | Death | ABIOMED, INC. | 2025-06-04 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Sepsis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ENDURED SUSPECTED SEPSIS DURING IMPELLA 5.5 SUPPORT. SUPPORT CONTINUED WITH THE IMPLANTED PUMP DESPITE THE NOTED CONDITION. CARE WAS WITHDRAWN ROUGHLY SIX DAYS LATER AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & CP WITH SMARTASSIST, SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS: ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22142156&pc=NVY | 2017865-2025-75409 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NVY | DURATA STS OPTIM ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: IT WAS REPORTED UPON PRODUCT RECEIPT THAT THE PATIENT WITH THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR SYSTEM HAD PASSED AWAY DUE TO CARDIAC AMYLOIDOSIS, CARDIAC FAILURE, AND MULTI-ORGAN FAILURE. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. ONLY THE CONNECTOR PORTION OF A PARTIAL LEAD WAS RETURNED IN ONE PIECE FOR ANALYSIS. VISUAL INSPECTION AND ELECTRICAL TESTING WERE NORMAL WITH NO ANOMALIES FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22142293&pc=NPT | 2015691-2025-04403 | 2025-05-13 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-04 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER VALVE | Difficult to Insert; Material Puncture/Hole; Difficult to Remove; Physical Resistance/Sticking | Perforation of Vessels | K200258 |  | 1 | Event Description: AS REPORTED BY EDWARDS JAPANESE AFFILIATE, DURING A TRANSFEMORAL TAVR PROCEDURE USING A 23 MM SAPIEN 3 ULTRA RESILIA (S3UR) VALVE, A STENT STRUT PEELED BACK, PREVENTING DEVICE REMOVAL. DESPITE CALCIFICATION IN THE LEFT COMMON ILIAC ARTERY (CIA) WITH A MINIMUM LUMINAL DIAMETER OF 5.9 MM, POBA WAS PERFORMED WITH AN 8 MM BALLOON BEFORE INSERTING THE 14 FR ESHEATH+. THE ESHEATH+ WAS INSERTED WITH SOME RESISTANCE BUT PASSED THROUGH. INTRALIPOS WAS USED AS A LUBRICANT, AND THE DELIVERY SYSTEM WAS INSERTED. THE S3UR VALVE PASSED THROUGH THE CALCIFIED CIA AREA WITH USUAL RESISTANCE, BUT A STENT STRUT PEELED BACK AFTER EXITING THE SHEATH, LEADING TO THE PROCEDURE'S ABORTION. ATTEMPTS TO WITHDRAW THE SHEATH AND DELIVERY SYSTEM AS ONE UNIT FAILED AS THE VALVE GOT STUCK AT THE CIA CALCIFICATION. BLOOD PRESSURE DROPPED TO 50/30 MMHG, PROMPTING ECMO INITIATION AND CARDIAC MASSAGE. ECMO WAS STARTED 10 MINUTES LATER, BUT ADEQUATE FLOW WAS NOT ACHIEVED, AND BLOOD PRESSURE REMAINED AROUND 60/40 MMHG. ANGIOGRAPHY REVEALED CIA BLEEDING, BUT HEMOSTASIS ATTEMPTS FAILED. SURGERY WAS DEEMED NECESSARY BUT NOT FEASIBLE DUE TO THE PATIENT'S AGE AND HYPOTENSION. THE FAMILY CONSENTED TO NO FURTHER INVASIVE TREATMENT, AND THE PATIENT WAS RETURNED TO THE ICU WITH THE DEVICES IN PLACE. THE PATIENT PASSED AWAY THE SAME DAY, AND POST-MORTEM EXAMINATION REVEALED A HOLE IN THE SHEATH WHERE THE STENT STRUT HAD PEELED BACK. Manufacturer Narrative: THIS IS 1 OF 3 REPORTS. THE INVESTIGATION IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS, SECTIONS G6, H2, H6: TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS AND H11 HAS BEEN ADDED. THE EVENT(S) REPORTED IS/ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THE COMPLAINTS FOR DIFFICULTY INTRODUCING SHEATH, SHEATH PUNCTURED, AND INABILITY TO REMOVE SHEATH, WERE UNABLE TO BE CONFIRMED WHILE THE COMPLAINT FOR DIFFICULTY ADVANCING THROUGH SHEATH WAS CONFIRMED THROUGH IMAGING EVALUATION. A REVIEW OF THE DHR DID NOT PROVIDE ANY INDICATION THAT A MANUFACTURING NON-CONFORMANCE WOULD HAVE CONTRIBUTED TO THE COMPLAINT EVENT. A REVIEW OF IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. FURTHERMORE, NO ABNORMALITIES WERE REPORTED DURING DEVICE UNPACKING OR PREPARATION. AS REPORTED, 'A STENT STRUT OF THE VALVE WAS PEELED BACK AND THE DEVICES WAS NOT ABLE TO BE REMOVED. ['] THE ESHEATH+ WAS INSERTED WITH SOME RESISTANCE ['] THE S3UR VALVE SUCCESSFULLY PASSED THROUGH THE CALCIFIED AREA OF THE CIA WITHOUT SIGNIFICANT RESISTANCE (THERE WAS USUAL RESISTANCE BUT NOT SIGNIFICANT ['] IT WAS FOUND THAT A STENT STRUT OF INFLOW SIDE HAD PEELED BACK JUST AFTER THE VALVE EXITING THE SHEATH ['] THE VALVE BECAME STUCK AT THE CALCIFICATION OF THE CIA. THE SHEATH AND THE VALVE REMAINED IMMOBILE, WHILE ONLY THE DELIVERY SYSTEM COULD BE PULLED BACK AND THE VALVE WAS DISLODGED FROM THE CRIMPED POSITION.' THE PATIENT REPORTEDLY HAD 'CALCIFICATION PROTRUDING INTO THE LUMEN OF THE LEFT COMMON ILIAC ARTERY' AND CONSIDERED 'MODERATE' COMPARED TO THE RIGHT ACCESS VESSEL WHICH WAS CONSIDERED 'SEVERE'. PER THE TRAINING MANUAL, 'ACCESS CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF THE SHEATH SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION, SEVERE TORTUOSITY OR VESSEL DIAMETER LESS THAN THE MINIMUM RECOMMENDED SHOULD BE CAREFULLY ASSESSED PRIOR TO THE PROCEDURE.' IN THIS CASE, IT IS LIKELY THAT RESISTANCE WAS EXPERIENCED WHEN INSERTING THE SHEATH DUE TO 'PROTRUDING' CALCIFICATION. SHARP, CALCIFIED NODULES WITHIN THE PATIENT ACCESS VESSEL CAN ACT AS PHYSICAL OBSTACLES LEADING TO HIGHER PUSH FORCE AND/OR DEVICE MANIPULATION, WHICH MAY RESULT IN STRUT DAMAGE AT THE VALVE INFLOW SIDE. A DAMAGED STRUT MAY LEAD TO RESISTANCE IN ADVANCING THE DELIVERY SYSTEM THROUGH THE SHEATH. ADDITIONALLY, THE PUNCTURE POTENTIALLY OCCURRED DUE TO THE DAMAGED STRUT AS IT WAS REPORTED 'A HOLE (OR POSSIBLY A TEAR) WAS FOUND IN THE SHEATH AROUND THE AREA WHERE THE STENT STRUT HAD PEELED BACK.' WHILE THE DELIVERY SYSTEM WAS ABLE TO BE WITHDRAWN, FURTHER REMOVAL OF THE SHEATH WITH THE DISLODGED CRIMPED VALVE WAS UNABLE TO BE PERFORMED. IT IS POSSIBLE THE 'PROTRUDING' CALCIFICATION ACTED AS A PHYSICAL OBSTACLE, INDUCING FRICTION AGAINST THE SHEATH SHAFT IN WITHDRAWAL ATTEMPTS. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION) LIKELY CONTRIBUTED TO THE DIFFICULTY IN INTRODUCING THE SHEATH AND WITHDRAWING THE SHEATH. PATIENT FACTORS (CALCIFICATION) AND/OR PROCEDURAL FACTORS (ALTERED CRIMP PROFILE) MAY HAVE CONTRIBUTED TO THE RESISTANCE IN ADVANCING THE DELIVERY SYSTEM THROUGH THE SHEATH. PROCEDURAL FACTORS (VALVE STRUT INTERACTION WITH SHEATH) LIKELY CONTRIBUTED TO THE ALLEGED PUNCTURE ON THE SHEATH. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22142312&pc=NPT | 2015691-2025-04404 | 2025-05-13 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-04 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA VALVE | Material Deformation | Insufficient Information | P140031 |  | 1 | Event Description: AS REPORTED BY EDWARDS JAPANESE AFFILIATE, DURING A TRANSFEMORAL TAVR PROCEDURE USING A 23 MM SAPIEN 3 ULTRA RESILIA (S3UR) VALVE, A STENT STRUT PEELED BACK, PREVENTING DEVICE REMOVAL. DESPITE CALCIFICATION IN THE LEFT COMMON ILIAC ARTERY (CIA) WITH A MINIMUM LUMINAL DIAMETER OF 5.9 MM, POBA WAS PERFORMED WITH AN 8 MM BALLOON BEFORE INSERTING THE 14 FR ESHEATH+. THE ESHEATH+ WAS INSERTED WITH SOME RESISTANCE BUT PASSED THROUGH. INTRALIPOS WAS USED AS A LUBRICANT, AND THE DELIVERY SYSTEM WAS INSERTED. THE S3UR VALVE PASSED THROUGH THE CALCIFIED CIA AREA WITH USUAL RESISTANCE, BUT A STENT STRUT PEELED BACK AFTER EXITING THE SHEATH, LEADING TO THE PROCEDURE'S ABORTION. ATTEMPTS TO WITHDRAW THE SHEATH AND DELIVERY SYSTEM AS ONE UNIT FAILED AS THE VALVE GOT STUCK AT THE CIA CALCIFICATION. BLOOD PRESSURE DROPPED TO 50/30 MMHG, PROMPTING ECMO INITIATION AND CARDIAC MASSAGE. ECMO WAS STARTED 10 MINUTES LATER, BUT ADEQUATE FLOW WAS NOT ACHIEVED, AND BLOOD PRESSURE REMAINED AROUND 60/40 MMHG. ANGIOGRAPHY REVEALED CIA BLEEDING, BUT HEMOSTASIS ATTEMPTS FAILED. SURGERY WAS DEEMED NECESSARY BUT NOT FEASIBLE DUE TO THE PATIENT'S AGE AND HYPOTENSION. THE FAMILY CONSENTED TO NO FURTHER INVASIVE TREATMENT, AND THE PATIENT WAS RETURNED TO THE ICU WITH THE DEVICES IN PLACE. THE PATIENT PASSED AWAY THE SAME DAY, AND POST-MORTEM EXAMINATION REVEALED A HOLE IN THE SHEATH WHERE THE STENT STRUT HAD PEELED BACK. Manufacturer Narrative: THIS IS 2 OF 3 REPORTS. INVESTIGATION IS ONGOING. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS, SECTIONS G6, H2, H6: COMPONENT CODES, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS, H10, AND H11 HAS BEEN ADDED. THE EVENT(S) REPORTED IS/ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATION'S HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH AND POTENTIAL VALVE FRAME DAMAGE USING THE S3U/S3UR VALVE CONFIGURATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCE'S. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, UNDERSIZED VESSELS, AND/OR STEEP INSERTION ANGLE. IN ADDITION, VALVE FRAME AND/OR SHEATH DAMAGE CAN BE A RESULT OF INCREASED PUSH FORCE AND ANY EXCESSIVE DEVICE MANIPULATION OR SHEATH-VALVE INTERACTION. THE COMPLAINT FOR FRAME DAMAGE WAS CONFIRMED BASED ON THE IMAGERY PROVIDED. AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (CALCIFICATION) AND/OR PROCEDURAL FACTORS (HIGH PUSH FORCE) LIKELY CONTRIBUTED TO THE EVENT AS REPORTED AND OBSERVED IN IMAGING EVALUATION. AS REPORTED, 'THE ACCESS VESSEL MINIMUM LUMINAL DIAMETER MEASURED 5.9 MM WITH MODERATE CALCIFICATION' AND 'IT WAS OBSERVED THAT THE STENT STRUT HAD PEELED BACK WHILE PASSING THROUGH THE CALCIFIED AREA OF THE CIA.' PER IMAGING EVALUATION, THE PROVIDED FLUOROSCOPY VIDEO OF CRIMPED VALVE BEING ADVANCED THROUGH THE SHEATH SHOWED THE VALVE JUMP AS IT PASSED THROUGH THE REPORTEDLY CALCIFIED SECTION OF THE VESSEL LEADING TO THE STRUT BENDING BACKWARDS, SUGGESTING THAT HIGH PUSH FORCE WAS APPLIED TO ADVANCE THE VALVE THROUGH THE CALCIFIED PATIENT ANATOMY. CALCIFICATION CAN REDUCE THE VESSEL LUMEN DIAMETER AND MAY INCREASE RESTRICTION LEADING TO RESISTANCE. CALCIFICATION CAN ALSO RESULT IN THE CREATION OF SUB-OPTIMAL ANGLES DURING DELIVERY SYSTEM INSERTION THAT MAY LEAD TO RESISTANCE. HIGH PUSH FORCE CAN LEAD TO THE VALVE STRUTS INTERACTING WITH THE SHEATH SHAFT AND RESULT IN THE STRUT DAMAGE AT THE VALVE INFLOW SIDE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22142332&pc=DRA | 2124215-2025-36646 | 2025-05-09 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-04 04:00 | DRA | FARADRIVE STEERABLE SHEATH CLEAR | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Tachycardia; Cardiac Perforation; Pericardial Effusion; Asystole |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY DURING A PULSED FIELD ABLATION PROCEDURE. THE PATIENT WAS PREPPED FOR A PULMONARY VEIN ISOLATION PROCEDURE TO TREAT ATRIAL FIBRILLATION. AT 8:50 AM, ACCESS WAS ESTABLISHED IN BOTH FEMORAL VEINS, AND ICE IMAGING WAS INITIATED. THE GUIDEWIRE AND VERSACROSS TRANSSEPTAL SHEATH WERE ADVANCED, WITH A TRANSSEPTAL PUNCTURE PERFORMED BY 8:59 AM. SHORTLY AFTER, THE PATIENT EXPERIENCED A SUSTAINED HEART RATE OF 215 BPM, FOLLOWED BY THE OBSERVATION OF PERICARDIAL EFFUSION AT 9:05 AM. A PERICARDIOCENTESIS WAS INITIATED, BUT THE PATIENT WENT INTO ASYSTOLE, PROMPTING CARDIOPULMONARY RESUSCITATION (CPR) AND A CODE BLUE ALARM. DESPITE EFFORTS INCLUDING CHEST OPENING WITH A STERNAL SAW, BLOOD TRANSFUSION, CARDIAC MASSAGE, AND MEDICATION, THE PATIENT REMAINED UNRESPONSIVE. ADDITIONALLY, WHEN THE CHEST WAS OPENED UP, THEY NOTICED ONE PERFORATION IN THE RIGHT VENTRICLE (RV) THAT THEY BELIEVED MAY HAVE BEEN CAUSED BY THE TAP FOR THE PERICARDIOCENTESIS. THE LEFT APPENDAGE HAD A COUPLE OF PERFORATIONS AS WELL THAT THEY BELIEVE WERE FROM THE DILATOR. BYPASS WAS INITIATED AT 10:30 AM, AND A 90-VOLT SHOCK WAS ADMINISTERED AT 11:45 AM. UNFORTUNATELY, THE PATIENT WAS DECLARED DECEASED AT 12:35 PM. AN AUTOPSY HAD NOT BEEN PERFORMED, HOWEVER IT WAS SUSPECTED THAT THE PATIENT DEATH WAS CAUSED BY COMPLICATIONS DUE TO PERFORATION OF CARDIAC TISSUE. NO ABLATION HAD BEEN PERFORMED, ONLY MAPPED THE CASE. IT WAS SUSPECTED THAT THE FARADRIVE AND THE VERSACROSS WIRE WERE THE CAUSE OF THE INJURY. IT WAS ALSO SUSPECTED THAT THE FARADRIVE DILATOR PERFORATED THE LEFT APPENDAGE WHILE GOING TRANSSEPTAL. A FARADRIVE SHEATH AND A BOSTON CORONARY SINUS (CS) WERE IN THE BODY WHEN THE EVENT OCCURRED. THE FARADRIVE STEERABLE SHEATH IS NOT EXPECTED TO RETURN AS IT WAS DISPOSED. IT WAS FURTHER REPORTED THAT THE SHEATH WAS LEFT INSIDE OF THAT PATIENT AFTER DEATH AND WAS KEPT FOR THE HOSPITAL FOR ANALYSIS. THE LOCATION OF THE SHEATH IS NOT KNOWN AND MAY PRESUME THAT IT HAS BEEN DISPOSED. NO AUTOPSY WAS PERFORMED. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22142487&pc=OZD | 1220648-2025-28983 | 2025-05-10 04:00 | Death | ABIOMED, INC. | 2025-06-04 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED 70 YEAR OLD MALE PATIENT ON THE IMPELLA CP AND DURING SUPPORT, THE PATIENT DEVELOPED HEMOLYSIS WITH TINGED URINE. WHEN THE HEMOLYSIS STARTED TO IMPROVE AND THE URINE WAS CLEARING UP, THE FOLLOWING DAY 1 UNIT OF BLOOD WAS GIVEN FOR HEMOGLOBIN OF 7.6. IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE PATIENT'S OUTCOME WAS REVIEWED BY ABIOMED'S MEDICAL SAFETY OFFICER (MSO) AND IT WAS DETERMINED THAT THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR TO THE PATIENT OUTCOME. THE PATIENT DEVELOPED HEMOLYSIS; SEVERE HEMOLYSIS CAN CAUSE MULTI-ORGAN FAILURE. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22142563&pc=MKJ | 3030677-2025-001447 | 2025-04-14 04:00 | Death | PHILIPS NORTH AMERICA LLC | 2025-06-04 04:00 | MKJ | PHILIPS HEARTSTART FRX | Device Handling Problem | Cardiac Arrest | P180028 |  | 1 | Event Description: IT HAS BEEN REPORTED THAT WHEN THE DEVICE WAS ATTEMPTED TO BE USED FOR A PATIENT USE EVENT, THE PADS WERE EXPIRED AND THE USER WAS NOT ABLE TO USE THE DEVICE. THE PATIENT DID NOT SURVIVE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22142803&pc=OZD | 1220648-2025-28986 | 2025-05-09 04:00 | Death | ABIOMED, INC. | 2025-06-04 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia; Renal Failure | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP AND EXPERIENCED LIMB ISCHEMIA AND RENAL FAILURE. ADDITIONALLY, THERE WAS UNRELATED BLEEDING FROM THE NASOGASTRIC TUBE AND THE ENDOTRACHEAL TUBE AND HEPARIN WAS WITHHELD. ULTIMATELY CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22142844&pc=DSQ | 2916596-2025-03338 | 2025-05-11 04:00 | Death | THORATEC CORPORATION | 2025-06-04 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiomyopathy; Embolism/Embolus; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure; Unspecified Respiratory Problem | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT DUE TO RESPIRATORY DISTRESS. THE PATIENT WAS FOUND TO HAVE SINUS OF VALSALVA THROMBUS THAT OCCLUDED THE LEFT MAIN CORONARY ARTERY (LMCA), LEFT ANTERIOR DESCENDING ARTERY (LAD), AND THE CIRCUMFLEX ARTERY (CX) SYSTEM THAT LED TO RIGHT VENTRICULAR (RV) FAILURE AND SYSTEMIC EMBOLIZATION. MECHANICAL CIRCULATORY SUPPORT (MCS) WAS NOT ESCALATED DUE TO FUNCTIONAL BASELINE STATUS. THE FAMILY ELECTED FOR WITHDRAWAL OF CARE. THE PATIENT PASSED AWAY DUE TO STAGE D CARDIOMYOPATHY. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. THE PUMP WOULD NOT BE EXPLANTED OR RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. HE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, RESPIRATORY FAILURE, THROMBOEMBOLISM, AND DEATH. SECTION 6 OF THE IFU ENTITLED ¿PATIENT CARE AND MANAGEMENT¿ LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SECTION 6 ALSO PROVIDES INFORMATION REGARDING ANTICOAGULATION, INCLUDING RECOMMENDED INR VALUES. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22142983&pc=NPT | 2025587-2025-04058 | 2025-05-14 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-04 04:00 | NPT | EVOLUT PRO TRANSCATHETER AORTIC VALVE | Gradient Increase; Insufficient Information; Patient Device Interaction Problem; Thickening of Material; Device Stenosis | Aortic Valve Stenosis; Dyspnea; Cardiogenic Shock; Respiratory Failure; Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 YEARS AND 7 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED WHICH REVEALED THAT THE VALVE FAILED DUE TO AORTIC VALVE STENOSIS. ADDITIONALLY, THROMBUS/PANNUS FORMATION ON THE VALVE LEAFLETS WAS OBSERVED. IT WAS REPORTED THAT THE PATIENT EXPERIENCED HEART FAILURE SYMPTOMS/HEMODYNAMIC INSTABILITY AND WAS HOSPITALIZED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT SUBSEQUENTLY PASSED AWAY.Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 YEARS AND 7 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED WHICH REVEALED THAT THE VALVE FAILED DUE TO AORTIC VALVE STENOSIS. ADDITIONALLY, THROMBUS/PANNUS FORMATION ON THE VALVE LEAFLETS WAS OBSERVED. IT WAS REPORTED THAT THE PATIENT EXPERIENCED HEART FAILURE SYMPTOMS/HEMODYNAMIC INSTABILITY AND WAS HOSPITALIZED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE VALVE IMPLANT DEPTH IN THE NATIVE ANNULUS WAS 4.2 MM ON THE LEFT CORONARY CUSP (LCC) AND 4.3 MM ON THE RIGHT CORONARY CUSP (RCC). AFTER THE PROCEDURE, ON DISCHARGE OF THE PATIENT, THE MEAN AORTIC GRADIENT WAS 11 MM HG. THE MEAN GRADIENT WAS 34 MM HG WHEN THE THROMBUS WAS DETECTED. OF NOTE, THE PATIENT HAD A HISTORY OF PELVIC FRACTURE APPROXIMATELY TWO MONTHS PRIOR. THIS MAY HAVE CONTRIBUTED TO THROMBUS FORMATION. THE LAST THREE INTERNATIONAL NORMALIZED RATIO (INR) RESULTS BEFORE THE THROMBUS WAS DETECTED WERE 1.0, 1.2, AND 1.0. THE MOST RECENT INR WAS MEASURED APPROXIMATELY 1 YEAR PRIOR TO THE THROMBUS. THE PATIENT HAD BEEN TAKING 81 MG OF ACETYLSALICYLIC ACID AFTER THE VALVE WAS IMPLANTED. AFTER THE THROMBUS WAS OBSERVED, ANTICOAGULANT THERAPY WAS INITIATED. THE HEART FAILURE SYMPTOMS NOTED INCLUDED WORSENING DYSPNEA, ACUTE HYPOXIC RESPIRATORY FAILURE, HYPOXIA, AND TACHYPNEA. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT SUBSEQUENTLY PASSED AWAY. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE PATIENT DETERIORATED DUE TO ACUTE HYPOXIC RESPIRATORY FAILURE, WORSENING HEART FAILURE, AND CARDIOGENIC SHOCK. NO DOCUMENTATION NOTED ONE SPECIFIC CAUSE OF DEATH.Event Description: IT WAS REPORTED THAT APPROXIMATELY 5 YEARS AND 7 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED WHICH REVEALED THAT THE VALVE FAILED DUE TO AORTIC VALVE STENOSIS. ADDITIONALLY, THROMBUS/PANNUS FORMATION ON THE VALVE LEAFLETS WAS OBSERVED. IT WAS REPORTED THAT THE PATIENT EXPERIENCED HEART FAILURE SYMPTOMS/HEMODYNAMIC INSTABILITY AND WAS HOSPITALIZED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE VALVE IMPLANT DEPTH IN THE NATIVE ANNULUS WAS 4.2 MM ON THE LEFT CORONARY CUSP (LCC) AND 4.3 MM ON THE RIGHT CORONARY CUSP (RCC). AFTER THE PROCEDURE, ON DISCHARGE OF THE PATIENT, THE MEAN AORTIC GRADIENT WAS 11 MM HG. THE MEAN GRADIENT WAS 34 MM HG WHEN THE THROMBUS WAS DETECTED. OF NOTE, THE PATIENT HAD A HISTORY OF PELVIC FRACTURE APPROXIMATELY TWO MONTHS PRIOR. THIS MAY HAVE CONTRIBUTED TO THROMBUS FORMATION. THE LAST THREE INTERNATIONAL NORMALIZED RATIO (INR) RESULTS BEFORE THE THROMBUS WAS DETECTED WERE 1.0, 1.2, AND 1.0. THE MOST RECENT INR WAS MEASURED APPROXIMATELY 1 YEAR PRIOR TO THE THROMBUS. THE PATIENT HAD BEEN TAKING 81 MG OF ACETYLSALICYLIC ACID AFTER THE VALVE WAS IMPLANTED. AFTER THE THROMBUS WAS OBSERVED, ANTICOAGULANT THERAPY WAS INITIATED. THE HEART FAILURE SYMPTOMS NOTED INCLUDED WORSENING DYSPNEA, ACUTE HYPOXIC RESPIRATORY FAILURE, HYPOXIA, AND TACHYPNEA. Manufacturer Narrative: UPDATED DATA: B5. ADDED SECOND PARAGRAPH. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. ADDED FOURTH PARAGRAPH. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B2. B5. D4. H1. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22143017&pc=DSQ | 2916596-2025-03604 | 2025-05-13 04:00 | Death | THORATEC CORPORATION | 2025-06-04 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE ANY ADDITIONAL INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22143079&pc=PYX | 1220648-2025-28989 | 2025-05-15 04:00 | Death | ABIOMED, INC. | 2025-06-04 04:00 | PYX | IMPELLA RP FLEX | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA RP FLEX SUPPORT AND DURING SUPPORT, ANOTHER SUTURE WAS APPLIED TO ASSIST WITH OOZING. THE PATIENT CONTINUED TO DECLINE SO THE FAMILY OPTED TO WITHDRAW SUPPORT AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RETURNED, AND THE INVESTIGATION HAS BEEN COMPLETED. THE ROOT CAUSE OF ACCESS SITE BLEEDING WAS MOST LIKELY USE ISSUE SINCE AN ADDITIONAL SUTURE WAS NEEDED TO ASSIST THE BLEED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA RP FLEX WITH SMARTASSIST INSTRUCTIONS FOR USE & CLINICAL REFERENCE MANUAL SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA RP FLEX SMARTASSIST SYSTEM CATHETER USE OF THE REPOSITIONING SHEATH AND THE 23 FR PEEL-AWAY INTRODUCER ¿MAKE SURE THERE IS NO BLEEDING AT THE TRANSITION FROM THE REPOSITIONING SHEATH TO THE INTERNAL JUGULAR VEIN. CLOSE AND DRESS THE WOUND.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22143195&pc=NVN | 2017865-2025-75430 | 2024-02-11 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NVN | TENDRIL MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD WITH ONLY THE CONNECTOR PORTION WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22143197&pc=LWP | 2017865-2025-75429 | 2024-02-11 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO PATIENT DEATH. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. THE RESULTS OF THE ELECTRICAL AND STRESS/ENVIRONMENTAL TESTS WERE PERFORMED TO INDICATE THAT THE PACER IS EXHIBITING NORMAL DEVICE CHARACTERISTICS. A LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22143205&pc=NVN | 2017865-2025-75431 | 2024-02-11 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NVN | TENDRIL MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD WITH ONLY THE CONNECTOR PORTION WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22143425&pc=DSQ | 2916596-2025-03698 | 2025-05-10 04:00 | Death | THORATEC CORPORATION | 2025-06-04 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT OUTCOME, THE PATIENT PASSED AWAY DUE TO A STROKE. THE CAUSE OF DEATH WAS NOT PROVIDED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22143509&pc=DSQ | 2916596-2025-03432 | 2025-04-25 04:00 | Death | THORATEC CORPORATION | 2025-06-04 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hematoma; Intracranial Hemorrhage; Convulsion/Seizure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS FOUND "CONSCIOUS CHANGE AND SEIZURE" IN THE WARD. A COMPUTED TOMOGRAPHY (CT) SCAN SHOWED SUBDURAL HEMATOMA. THE PATIENT'S FAMILY REFUSED SURGERY AND CHOSE TO FORGO RESUSCITATION. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT'S INTERNATIONAL NORMALIZED RATIO (INR) WAS WITHIN THE ACCEPTABLE RANGE, HOWEVER RELATION TO THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) COULD NOT BE RULED OUT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22143578&pc=BZD | 2518422-2025-106369 | 2025-05-28 04:00 | Death | RESPIRONICS, INC. | 2025-06-04 04:00 | BZD | DS2ADV AUTO CPAP W/HUMID CELL/BT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K200480 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A USER OF A DREAMSTATION 2 ADVANCED AUTO CPAP DEVICE PASSED AWAY LAST YEAR. THE PATIENT'S SPOUSE IS REQUESTING INFORMATION ON WHAT TO DO WITH THE PATIENT'S UNUSED DEVICES. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22143719&pc=OZD | 1220648-2025-28993 | 2025-05-11 04:00 | Death | ABIOMED, INC. | 2025-06-04 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Ischemia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA PC LOST PULSES IN THE LEG AND EXHIBITED HEMOLYSIS. THE PUMP WAS REPOSITIONED BUT THE HEMOLYSIS DID NOT RESOLVE. REPERFUSION OF THE CANNULA RESOLVED THE LIMB ISCHEMIA. SUPPORT WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: A4 AND A5 ARE UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22143894&pc=NPT | 2015691-2025-04424 | 2025-05-14 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-04 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY OUR BRAZILIAN AFFILIATES, JUST FOLLOWING IMPLANT OF A 26MM SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IMPLANT IN THE AORTIC POSITION BY TRANSFEMORAL APPROACH, THE PATIENT PRESENTED ANNULAR RUPTURE. A CARDIAC WINDOW WAS PERFORMED BY THE SURGERY TEAM, WHERE ACCORDING TO THE REPORT, THE PATIENT'S HEART WALL WAS FRIABLE AND WHEN ATTEMPTING TO SUTURE THE LESION, THE TISSUE WAS LACERATING. THE PATIENT PASSED AWAY IN THE OPERATING ROOM. Manufacturer Narrative: THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THAT PRE-EXISTING PATIENT FACTORS COULD HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22143994&pc=OZD | 1220648-2025-28996 | 2025-05-11 04:00 | Death | ABIOMED, INC. | 2025-06-04 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT WHILE IMPLANTED WITH AN IMPELLA CP THE PATIENT EXPERIENCED ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE. THREE FAILED CARDIOVERSION ATTEMPTS WERE MADE. THE PATIENT WAS TRANSFUSED A BOLUS OF AMIODARONE AND THE IMPELLA WAS REPOSITIONED. THE PATIENT ALSO SUSTAINED A GASTROINTESTINAL BLEED AND CRITICALLY LOW HEMOGLOBIN LEVELS. THE PATIENT'S FAMILY CHOSE TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22144043&pc=EOQ | 2955842-2025-23846 | 2025-04-30 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-06-04 04:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNEVENTFUL ION ENDOLUMINAL LUNG BIOPSY PROCEDURE. THE PHYSICIAN REPORTED THAT THE PATIENT WAS STABLE THE ENTIRE TIME AND WAS EXTUBATED. THE BIOPSIED LESION WAS 5 CM IN SIZE, LOCATED IN THE LEFT UPPER LOBE FAR FROM THE PLEURA, AND WAS DETERMINED TO BE NECROTIC TISSUE. APPROXIMATELY 2.5 HOURS POST-PROCEDURE THE PHYSICIAN AND ANESTHESIOLOGIST WERE NOTIFIED THAT THE PATIENT WAS STILL DROWSY AND WAS NOT FOLLOWING COMMANDS; THE PATIENT WAS OTHERWISE HEMODYNAMICALLY STABLE. THE PATIENT WAS TRANSFERRED TO THE EMERGENCY DEPARTMENT FOR A STROKE WORK UP. THE PATIENT WAS INTUBATED AT SOME POINT OVERNIGHT WHILE IN THE ICU FOR DECLINING MENTAL STATUS AND EXPIRED AT AN UNSPECIFIED TIME LATER. THE CAUSE OF DEATH WAS UNKNOWN. AN AUTOPSY IS PENDING. PER THE PHYSICIAN, THE ION SYSTEM DID NOT EXHIBIT ANY ISSUES OR UNEXPECTED BEHAVIORS THAT CONTRIBUTED TO THE DEATH. NO ADDITIONAL INFORMATION WAS PROVIDED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE SHOWED THERE WERE NO RELATED SYSTEM ERRORS TO HAVE OCCURRED DURING THE PROCEDURE THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT WITH A HISTORY OF ENDOMETRIAL CANCER UNDERWENT AN ION LUNG BIOPSY. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE. POST EXTUBATION IN RECOVERY THE PATIENT¿S SENSORIUM REMAINED ALTERED 2.5 HOURS POST PROCEDURE PROMPTING WORK UP FOR A STROKE AND ADMISSION TO THE ICU. THE PATIENT WAS INTUBATED OVERNIGHT FOR WORSENING MENTAL STATUS BUT REMAINED OTHERWISE HEMODYNAMICALLY STABLE. THE PATIENT EVENTUALLY DIED AND AN AUTOPSY IS PENDING AND EFFORTS AT OBTAINING THE RESULTS ARE ONGOING. AT THIS JUNCTURE THE CAUSE OF DEATH IN THE SETTING OF AN ALTERED SENSORIUM POST PROCEDURE IS UNCLEAR BUT THERE IS NO EVIDENCE OF MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS, OR ACCESSORIES. BASED ON THE AVAILABLE DATA THE EVENTS WERE LIKELY PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. SECTION B2: PATIENT DATE OF DEATH: INTUBATION WAS THE DAY AFTER THE PROCEDURE (05/01/2025); HOWEVER, THE DATE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: ADDITIONAL INFORMATION: SECTION B5: THE PHYSICIAN REPORTED THAT APPROXIMATELY 2.5 HOURS POST-PROCEDURE, THE PATIENT WAS NOTED TO BE HYPERTENSIVE AND REMAINED DROWSY, PROMPTING TRANSFER TO THE EMERGENCY ROOM FOR FURTHER EVALUATION. PER THE PHYSICIAN, THE PATIENT WAS NEVER HYPOXIC. THE ONLY CHANGE IN VITAL SIGNS WAS HYPERTENSION, WHILE THE PATIENT REMAINED ON ROOM AIR. A BRAIN, HEAD, AND NECK CT SCAN, AND ALL LABS WERE UNREMARKABLE EXCEPT FOR MILDLY ELEVATED LACTATE. AN MRI OF THE BRAIN SHOWED ABNORMALITY MAINLY ON THE RIGHT SIDE AND WAS SUGGESTIVE OF ANOXIC BRAIN INJURY. THE AUTOPSY WAS DESCRIBED AS "NOT VERY REVEALING," AND THE REPORTED CAUSE OF DEATH WAS HYPOXIA WITH NO SECONDARY CAUSES. THE PHYSICIAN REPORTED THE DEATH WAS NOT RELATED TO THE ION PROCEDURE. A RE-REVIEW OF THE EVENT WAS CONDUCTED BY AN ISI MEDICAL OFFICER, WHO CONCLUDED THAT THE PATIENT WITH A HISTORY OF ENDOMETRIAL CANCER UNDERWENT AN ION LUNG BIOPSY OF A LUNG LESION. THE BIOPSY WAS COMPLETED WITHOUT ISSUE IN ABOUT 35 MINUTES. 2.5 HOURS POST EXTUBATION THE PATIENT¿S SENSORIUM REMAINED ALTERED AND WAS NOT FOLLOWING COMMANDS BUT OTHERWISE HEMODYNAMICALLY STABLE. WORK UP FOR STROKE WAS NEGATIVE INCLUDING A CT OF THE HEAD AND NECK. LABS WERE NOTABLE FOR AN ELEVATED LACTATE IN THE SETTING OF CONTINUED HEMODYNAMIC STABILITY. MRI PERFORMED LATER THE SAME DAY WAS SUGGESTIVE OF ANOXIC BRAIN INJURY PREDOMINANTLY ON THE RIGHT SIDE. DUE TO WORSENING MENTAL STATUS THE PATIENT WAS INTUBATED OVERNIGHT IN THE ICU. THE PATIENT ULTIMATELY DIED. THERE WAS NO MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. THE CAUSE OF DEATH WAS ATTRIBUTED TO BRAIN HYPOXIA BUT THE ETIOLOGY REMAINED UNCLEAR DESPITE WORK UP WHICH INCLUDED AN AUTOPSY. BASED ON THE AVAILABLE DATA THE EVENT WAS PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE 227 (1.08%) WITH 5 (0.02%) TRANSIENT ISCHEMIC ATTACKS AND 4 (0.02%) TOTAL DEATHS BUT NO STROKES. ANOTHER MULTICENTER STUDY REPORTED 13 (2.2%) COMPLICATIONS WITH NO STROKES AND NO DEATHS ASSOCIATED WITH 581 CASES. A RECENT PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 BRONCHOSCOPIES REPORTED 1 ASSOCIATED DEATH (0.08%) AND NO STROKES. ANOTHER PROSPECTIVE SERIES OF 415 ION PROCEDURES REPORTED 1 CVA (0.2%). THE RISK OF STROKE ASSOCIATED WITH GENERAL ANESTHESIA HAS BEEN REPORTED TO VARY FROM 0.1-1.0% IN NON-CARDIAC, NON-NEUROLOGIC, AND NON-MAJOR SURGERY AND AS HIGH AS 10% IN THE SETTING OF BRAIN OR HIGH-RISK CARDIOVASCULAR SURGERY. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. BATEMAN BT, SCHUMACHER HC, WANG S, SHAEFI S, BERMAN MF. PERIOPERATIVE ACUTE ISCHEMIC STROKE IN NONCARDIAC AND NONVASCULAR SURGERY: INCIDENCE, RISK FACTORS, AND OUTCOMES. ANESTHESIOLOGY. 2009. SELIM M. PERIOPERATIVE STROKE. NEW ENGLAND JOURNAL OF MEDICINE. 2007.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22144271&pc=DSQ | 3007042319-2025-01229 | 2025-05-24 04:00 | Death | HEARTWARE, INC. | 2025-06-04 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE WAS FOUND DECEASED, WITH THE CAUSE OF DEATH UNKNOWN. THERE WAS NO KNOWN PUMP ISSUE LEADING TO DEATH. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS THE VAD REMAINS IMPLANTED. THE COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS WAS NOT PERFORMED SINCE LOG FILES WERE NOT AVAILABLE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH IS UNKNOWN. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22144279&pc=FKX | 0008030665-2025-01267 | 2025-05-22 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-06-04 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Dyspnea | K181108 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. THE PATIENT¿S DAUGHTER REPORTED THAT AS THEY WERE ATTEMPTED TO DISCONNECT THE PATIENT FROM THE CYCLER AFTER THE EVENT OCCURRED, FLUID DID COME OUT AS SHE WAS REMOVING THE CASSETTE FROM CASSETTE DOOR. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT WAS IN TREATMENT ON (B)(6) 2025 CONNECTED TO THE LIBERTY SELECT CYCLER. THE PATIENT HAD TO GET UP AND USE THE RESTROOM. THE PATIENT WAS IN THE RESTROOM WHEN SHE EXPRESSED TO HER SPOUSE THAT SHE COULD NOT BREATHE. THE PATIENT¿S SPOUSE CALLED 911. EMERGENCY MEDICAL SERVICES (EMS) RESPONDED. THE PATIENT HAD BECOME UNRESPONSIVE PRIOR TO EMS ARRIVAL. EMS ATTEMPTED RESUSCITATION, BUT IT WAS UNSUCCESSFUL. THE PATIENT WAS PRONOUNCED DECEASED IN THE HOME. THE CAUSE OF THE PATIENT¿S DEATH IS UNKNOWN. THERE WERE NO ISSUES OR REPORTED ALARMS DURING THE TREATMENT PRIOR TO THE PATIENT¿S DEATH. THE FLUID LEAK REPORTED DURING THE DISCONNECTION OF THE PATIENT AND THE SUPPLIES OCCURRED BECAUSE OF THE PATIENT¿S DAUGHTER NOT KNOWING WHAT SHE WAS DOING. THE PATIENT¿S DAUGHTER WAS ATTEMPTING TO DISCONNECT THE CASSETTE, AND SHE CAUSED THE FLUID TO LEAK. THERE WAS NO LEAK INDICATED DURING THE PATIENT¿S TREATMENT. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS AND UTILIZATION OF THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF NOT BEING ABLE TO BREATH, BECOMING UNRESPONSIVE, AND DEATH. HOWEVER, THERE IS NO DOCUMENTATION OF A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE REPORTED FLUID LEAK WAS A RESULT OF AN ERROR THAT THE PATIENT¿S DAUGHTER MADE WHEN ATTEMPTING TO DISCONNECT THE SUPPLIES FROM THE CYCLER. THERE WERE NO REPORTED ALARMS OR ISSUES WITH THE TREATMENT PRIOR TO THE PATIENT EVENT. ALTHOUGH THE CAUSE OF DEATH WAS NOT CONFIRMED, SUDDEN CARDIAC DEATH (SCD) IS RESPONSIBLE FOR THE MAJORITY OF CARDIOVASCULAR-RELATED DEATHS IN PATIENTS WITH ESRD WITH STUDIES REPORTING THAT UP TO 25 % OF ALL DEATHS IN THIS HIGH-RISK POPULATION IS ATTRIBUTABLE TO SCD. BASED ON THE AVAILABLE INFORMATION AND NO EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22144297&pc=GEX | 2124215-2025-35369 | 2025-01-16 05:00 | Death | LUMENIS LTD | 2025-06-04 04:00 | GEX | PULSE | Adverse Event Without Identified Device or Use Problem | Pneumonia; Sepsis; Multiple Organ Failure | K170121 |  | 1 | Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC VIA AN ARTICLE PUBLISHED IN THE WORLD JOURNAL OF UROLOGY THAT A COMPARATIVE RETROSPECTIVE ANALYSIS WAS CONDUCTED TO EVALUATE THE SAFETY AND EFFICACY OF HOLEP IN PATIENTS AGED OVER 85 YEARS WITH INDWELLING CATHETER (IDC). A OF A TOTAL OF 144 PATIENTS, WITH 71 IN THE GROUP OF PATIENTS AGED 85 OR GREATER AND 73 IN THE CONTROL GROUP WITH PATIENTS UNDER 70 YEARS OF AGE, WERE EVALUATED FOR RATE OF POSTOPERATIVE SPONTANEOUS MICTURITION, WITH SUCCESS DEFINED AS THE ABILITY TO RESUME SATISFACTORY SPONTANEOUS VOIDING WITHOUT SIGNIFICANT POST-VOID RESIDUAL VOLUME. MEDICAL FILES OF PATIENTS OPERATED BETWEEN JUNE 2012 AND APRIL 2020, AT TWO ACADEMIC CENTERS IN FRANCE, WERE REVIEWED TO RETRIEVE THOSE WITH HISTORY OF ACUTE OR CHRONIC URINARY RETENTION AND TRIAL WITHOUT CATHETER (TWOC) FAILURE REQUIRING MAINTENANCE OF IDC AT THE TIME OF SURGERY. FOLLOW-UP ASSESSMENTS WERE CONDUCTED AT BASELINE OF 3 MONTHS, 6 MONTHS, AND 12 MONTHS. FOLLOWING THE PROCEDURES, THE FOLLOWING COMPLICATIONS WERE REPORTED FOR THE OLDER PATIENTS GROUP: 8 PATIENTS REQUIRED BLOOD TRANSFUSIONS, 17 PATIENTS HAD TWOC FAILURE DURING HOSPITALIZATION, AT 3-MONTHS EVALUATION 5 PATIENTS COULD NOT VOID SPONTANEOUSLY, AND AT 12-MONTHS EVALUATION 2 PATIENTS REMAINED WITH AN IDC, 8 PATIENTS HAD URINARY INCONTINENCE, AND 10 PATIENTS HAD DIED. FOLLOWING THE PROCEDURES, THE FOLLOWING COMPLICATIONS WERE REPORTED FOR THE YOUNGER PATIENTS GROUP: 7 PATIENTS REQUIRED BLOOD TRANSFUSIONS, 7 PATIENTS HAD TWOC FAILURE DURING HOSPITALIZATION, AND AT 12-MONTHS EVALUATION 4 PATIENTS HAD URINARY INCONTINENCE AND 2 PATIENTS HAD DIED. IT WAS NOTED THAT ONLY ONE DEATH WAS DIRECTLY RELATED TO SURGERY (SEVERE SEPSIS DUE TO PNEUMONIA, COMPLICATED BY MULTIORGAN FAILURE, LEADING TO DEATH ON POSTOPERATIVE DAY 7). THIS REPORT IS FOR THE DEATH DIRECTLY RELATED TO SURGERY REPORTED IN THE ARTICLE. REFERENCE LITERATURE ARTICLE "COMPARATIVE STUDY OF HOLEP IN ELDERLY PATIENTS WITH INDWELLING CATHETERS: A RETROSPECTIVE DUAL-CENTER STUDY" INCLUDED WITH THIS REPORT FOR A FULL LISTING OF PHYSICIANS AND HEALTH CARE FACILITIES. Manufacturer Narrative: DATE OF DEATH (B2) WAS APPROXIMATED BASED ON AN ESTIMATED EARLIEST PROCEDURE DATE NOTED IN THE STUDY AND OTHER INFORMATION PROVIDED REGARDING THE DEATH EVENT. KLEIN, C., ANRACT, J., PINAR, U., LACROIX, X., MANSOUR, R., ROBERT, G., DELONGCHAMPS, N.B. (2025). "COMPARATIVE STUDY OF HOLEP IN ELDERLY PATIENTS WITH INDWELLING CATHETERS: A RETROSPECTIVE DUAL-CENTER STUDY". WORLD JOURNAL OF UROLOGY 43, 75. HTTPS://DOI.ORG/10.1007/S00345-024-05437-9 DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22144306&pc=DSQ | 2916596-2025-03560 | 2025-04-27 04:00 | Death | THORATEC CORPORATION | 2025-06-04 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; High Blood Pressure/ Hypertension; Unspecified Infection; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HAVING SURGERY AND HAD LOW FLOWS ALARMS DURING IT. THE CUSTOMER SENT IN LOG FILES TO BE REVIEWED TO MAKE SURE THERE WEREN'T ANY OTHER ISSUES GOING ON AND TO SEE HOW LONG THEY HAVE BEEN HAVING LOW FLOW EVENTS. THE PATIENT HAD A PUMP EXCHANGE ON (B)(6) 2024 (MFR#2916596-2024-01705). THE LOG FILES WERE REVIEWED AND CAPTURED MULTIPLE INTERMITTENT LOW FLOW ALARMS AND MOMENTARY LOW FLOW ESTIMATES WHEN THE CALCULATED PULSATILITY INDEX (PI) WAS ELEVATED STARTING ON (B)(6) 2025. THE ESTIMATED FLOW WAS FLUCTUATING BELOW 2.5 LPM. SOME OF THESE EVENTS WERE BRIEF AND DID NOT GENERATE AN ALARM. THE ESTIMATED FLOWS WERE AVERAGING FROM 2.3 TO 2.4 LPM AND PI IS AVERAGING FROM 4.2 TO 10.6 DURING THESE EVENTS. THESE SEEMED TO BE PHYSIOLOGICAL IN NATURE. THERE WERE NO OTHER NOTABLE ALARMS OR PARAMETER CHANGES. THE DEVICE WAS OPERATING AS EXPECTED. THE PATIENT HAD AN INCISION AND DRAINAGE (I&D) WITH A WOUND VAC PLACEMENT FOR THEIR PRE-EXISTING DRIVELINE EXIT SITE INFECTION (MFR#2916596-2024-02570). THE LOW FLOW ALARMS WERE DUE TO THE PATIENT WAS NOTHING BY MOUTH FOR THREE DAYS AND HAD OCCASIONAL HYPERTENSION. THE LOW FLOW ALARMS RESOLVED ABOUT A DAY AFTER SURGERY AND THE PATIENT WAS GIVEN INTRAVENOUS FLUIDS AND MEAN ARTERIAL PRESSURE WAS BETTER CONTROLLED. THE PATIENT HAD NOT REPORTED AN SYMPTOMS DURING THE ALARMS. THE PATIENT DECIDED TO GO ON HOSPICE THE EVENING OF (B)(6) 2025. THE PUMP WAS TURNED OFF ON (B)(6) 2025. THE PUMP WOULD NOT BE RETURNED AS THE PATIENT WAS A PRISONER AND THERE WOULD NOT BE AN AUTOPSY.Event Description: IT WAS REPORTED THAT THE PATIENT HAD CARDIAC ARREST AFTER DISCONTINUATION OF THE LEFT VENTRICULAR ASSIST DEVICE (LVAD). THIS WAS NOT CONSIDERED DEVICE RELATED AS THE PATIENT HAD MULTIPLE CO-MORBIDITIES AND DECIDED TO GO ON HOSPICE AND HAVE THE LVAD DISCONTINUED. Manufacturer Narrative: SECTION H6 - HEALTH EFFECT CODES: CORRECTED MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED INFECTION AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SUBMITTED LOG FILES CONTAINED NO ATYPICAL ALARMS AND APPEARED TO SHOW THE PUMP OPERATING AS INTENDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C, IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU LISTS INFECTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT,¿ LISTS INFECTION AS A POTENTIAL LATE POST-IMPLANT COMPLICATION THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. CARE INSTRUCTIONS IN REGARD TO PREVENTING INFECTION ARE PROVIDED IN VARIOUS SECTIONS OF THE INSTRUCTIONS FOR USE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22144335&pc=DSQ | 3007042319-2025-01230 | 2025-05-27 04:00 | Death | HEARTWARE, INC. | 2025-06-04 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Fall; Intracranial Hemorrhage; Brain Injury; Swelling/ Edema | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE EXPERIENCED A FALL, RESULTING IN TRAUMATIC BRAIN INJURY AND ALTERED MENTAL STATUS. ON ADMISSION, HEAD COMPUTED TOMOGRAPHY IDENTIFIED AN ACUTE LEFT FRONTAL LOBE SUBDURAL HEMORRHAGE, RIGHT PARIETAL INTRACRANIAL HEMORRHAGE, CEREBRAL EDEMA, AND A LEFT TO RIGHT MIDLINE SHIFT OF 0.7 CM. THE PATIENT HAD AN ELEVATED INTERNATIONAL NORMALIZED RATIO OF 2.7 ON ADMISSION AND HAD BEEN TAKING 2 MG OF ANTICOAGULANT DAILY PRIOR TO THE EVENT. REPEAT HEAD COMPUTED TOMOGRAPHY ON THE FOLLOWING DAY WAS UNCHANGED. THE PATIENT RECEIVED SUPPORTIVE CARE AND WAS NOT INTUBATED. DEATH OCCURRED AS A RESULT OF THE INJURIES. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINS IMPLANTED. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, HEMORRHAGIC CEREBROVASCULAR ACCIDENT AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22144496&pc=QZK | 2017233-2025-06272 | 2025-05-09 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-04 04:00 | QZK | GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Obstruction/Occlusion | P230023 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT A TAMBE ENDOVASCULAR PROCEDURE UTILIZING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS, AND GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX). REPORTEDLY, PHYSICIAN WAS UNABLE TO CANNULATE THE RIGHT AND LEFT RENAL ARTERY. THEY WERE ABLE TO GET A WIRE INTO THE LEFT RENAL ARTERY AND HAD NO ISSUES WITH THE THORACOABDOMINAL BRANCHED (TAMBE) GRAFT'S PORTALS. PRIMARILY, THE ISSUE WAS THE ANGLE OF PATIENT'S RENAL ARTERIES THAT WAS PROVING DIFFICULT TO CANNULATE AND GET ENOUGH PURCHASE IN THE LEFT RENAL ARTERY. PHYSICIAN TRIED TO INSERT THE BXB057902A (VBX) STENT AND PULLED THE WIRE BACK AS THERE WAS NOT ENOUGH SUPPORT, SO REMOVED THE VBX STENT GRAFT. PHYSICIAN THEN TRIED TO RE-POSITION THE GUIDEWIRE MULTIPLE TIMES AND WENT IN DIFFERENT DIRECTIONS AT THE OSTIUM OF THE LEFT RENAL ARTERY, WHICH CREATED A DISSECTION AS A RESULT, THUS, PHYSICIAN STOPPED AT THAT ATTEMPT. PHYSICIAN WAS NEVER ABLE TO GET A WIRE INTO THE RIGHT RENAL ARTERY (BY ENDOVASCULAR APPROACH). THE PHYSICIAN THEN EXPOSED THE ABDOMINAL AORTA AND BOTH RENAL ARTERIES, CANNULATED THE RIGHT RENAL ARTERY EXTERNALLY AND ATTEMPTED TO SNARE A GUIDEWIRE INSERTED INTO THE RIGHT RENAL ARTERY IN A RETROGRADE FASHION FROM A SNARE THAT WAS DELIVERED THROUGH THE RIGHT RENAL PORTAL OF THE TAMBE DEVICE. HE WAS NOT ABLE TO SNARE THE WIRE FROM THE RENAL ARTERY SO HE REVERSED THE WIRE AND SNARE AND TRIED TO SNARE FROM THE RIGHT RENAL ARTERY, UNSUCCESSFULLY. AT THIS POINT, HE PROCEEDED WITH PLUGGING BOTH RENAL PORTALS OF THE TAMBE DEVICE AND CONTINUED ON WITH THE REMAINDER OF THE TAMBE DEPLOYMENT, INCLUDING THE DISTAL MAIN BODY, BILATERAL ILIAC GRAFTS AND A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (CTAGAC) IN THE THORACIC AORTA. A COMPLETION ANGIO WAS PERFORMED WHICH SHOWED PATENT CELIAC AND SMA AND PLUGGED RENAL PORTALS. THE PHYSICIAN THEN PROCEEDED WITH THE BILATERAL RENAL BYPASS PROCEDURE. THE PATIENT WAS CLOSED UP WITH ALL DEVICES SUCCESSFULLY IMPLANTED AND BOTH RENALS WERE BYPASSED WITH THE PROCEDURE ENDING SUCCESSFULLY, HOWEVER, THIS WAS A PROLONGED PROCEDURE LASTING MORE THAN 9 HOURS. WHILE THE PATIENT WAS BEING SHIFTED OVER TO BED, THE PATIENT CODED. CPR WAS GIVEN BUT THE PATIENT PASSED AWAY. CAUSE OF DEATH IS PERSISTENT ACIDOTIC STATE. Manufacturer Narrative: H6: CODE C19 - A REVIEW OF THE MANUFACTURING RECORDS INDICATED THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE B20 -THE DEVICE REMAINS IMPLANTED AND WAS THEREFORE NOT AVAILABLE FOR ENGINEERING EVALUATION BY GORE. ACCORDING TO THE GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), POSSIBLE ADVERSE EVENTS AND COMPLICATIONS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE BUT ARE NOT LIMITED TO OCCLUSION OF DEVICE OR NATIVE VESSEL, SURGICAL INTERVENTION AND DEATH. IMAGING EVALUATION SUMMARY: THE IMAGING EVALUATION PERFORMED BY A CLINICAL IMAGING SPECIALIST SHOWED THE FOLLOWING: TWO TIME-POINTS AVAILABLE FOR EVALUATION: PRE-IMPLANTATION CTA¿S DATED (B)(6) 2024. PRE-IMPLANTATION CTA DATED (B)(6) 2024 SHOWS: AXIAL IMAGES SHOW A HIGHLY TORTUOUS RIGHT RENAL ARTERY (RRA) AND RIGHT ACCESSORY RENAL ARTERY (RARA) WITH DIAMETERS <4MM. THE RARA APPEARS TO ORIGINATE OFF THE RRA ~1.5MM DISTAL THE OSTIUM. THE PROXIMAL LEFT RENAL ARTERY (LRA) IS ALSO TORTUOUS AND IMMEDIATELY CURVES DISTALLY AFTER THE VESSEL OSTIUM. THE LRA APPEARS TO BE DISSECTED. DIAMETERS IN THE LRA APPEAR TO BE <4MM. MAXIMUM PROXIMAL AND DISTAL AAA DIAMETERS APPEAR TO BE 65.8MM AND 53.2MM. PRE-IMPLANTATION CTA DATED (B)(6) 2024 SHOW THE VERY SAME FINDINGS. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22144571&pc=DSQ | 2916596-2025-03420 | 2025-05-10 04:00 | Death | THORATEC CORPORATION | 2025-06-04 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SPONTANEOUS INTRACRANIAL BLEEDING (ICB). THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. THERE WERE NO ANTICOAGULATION ISSUES, NO DEVICE RELATED ISSUES WHICH LEAD TO THE OUTCOME WERE MENTIONED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22144740&pc=KDJ | 1416980-2025-03196 |  | Death | VANTIVE US HEALTHCARE LLC | 2025-06-04 04:00 | KDJ | NI | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Peritonitis | NI |  | 1 | Event Description: A PERITONEAL DIALYSIS (PD) PATIENT EXPERIENCED RECURRING PERITONITIS MANIFESTED BY RECURRING CLOUDY PD EFFLUENT. THE CAUSE OF PERITONITIS WAS UNKNOWN. IT WAS REPORTED THE PATIENT WAS HOSPITALIZED AND DISCHARGED FOR THE EVENTS. THE PATIENT WAS TREATED WITH CEFTAZIDIME INJECTION (1GM, ONCE DAILY, INTRAPERITONEAL, DISCONTINUED) FOR RECURRING PERITONITIS. IT WAS FURTHER REPORTED THAT THE PATIENT EXPERIENCED CARDIAC ARREST AND SUBSEQUENTLY PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED AS CARDIAC ARREST. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. IT WAS REPORTED THAT THE PATIENT WAS NOT RECOVERED FROM THE RECURRENT PERITONITIS. PD THERAPY WAS DISCONTINUED AND THE PATIENT WAS SWITCHED TO HEMODIALYSIS (HD) BEFORE THE EVENT OF CARDIAC ARREST. SAME HD WAS ONGOING PRIOR TO DEATH OR AT THE TIME OF DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED AND THE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22145110&pc=DSQ | 2916596-2025-03654 | 2025-05-13 04:00 | Death | THORATEC CORPORATION | 2025-06-04 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO HOSPITAL PRIVACY POLICY, NO ADDITIONAL INFORMATION WOULD BE PROVIDED. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22145123&pc=NVN | 2017865-2025-75357 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS CARDIOPULMONARY ARREST AND CONGESTIVE HEART FAILURE. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED FOR ANALYSIS. ELECTRICAL TESTING THAT COULD BE MEASURED DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22145129&pc=NVN | 2017865-2025-75356 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS CARDIOPULMONARY ARREST AND CONGESTIVE HEART FAILURE. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED FOR ANALYSIS. ELECTRICAL TESTING THAT COULD BE MEASURED DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22145136&pc=LWP | 2017865-2025-75355 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS CARDIOPULMONARY ARREST AND CONGESTIVE HEART FAILURE. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO THE PATIENT'S DEATH. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. THE RESULTS OF THE ELECTRICAL AND STRESS/ENVIRONMENTAL TESTS WERE PERFORMED TO INDICATE THAT THE PACER IS EXHIBITING NORMAL DEVICE CHARACTERISTICS. A LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22145174&pc=OZD | 1220648-2025-29007 | 2025-05-30 04:00 | Death | ABIOMED, INC. | 2025-06-04 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Vascular Dissection | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A CP PUMP PLACED TO SUPPORT A BUT THERE WAS AN AORTIC DISSECTION OBSERVED AND PATIENT EXPIRED. THE PATIENT RECEIVED E-CPR ON ADMISSION AND HIS PROGNOSIS WAS POOR SO THE DECISION MADE TO STOP THE PUMP SUPPORT. PATIENT EXPIRED FROM THE MULTI-ORGAN FAILURE AND AORTIC DISSECTION. NO PUMP MALFUNCTION. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI SECTION: WARNINGS & CAUTIONS: CAUTIONS. ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22145203&pc=LGW | 3006630150-2025-03942 | 2025-04-28 04:00 | Death | BOSTON SCIENTIFIC NEUROMODULATION CORPORATION | 2025-06-04 04:00 | LGW | WAVEWRITER ALPHA 16 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030017 |  | 1 | Event Description: IT WAS REPORTED THAT THE SPINAL CORD STIMULATION (SCS) PATIENT HAS DIED AND THE CAUSE OF DEATH IS UNKNOWN. THE PATIENT HAD RECENTLY BEEN HOSPITALIZED DUE TO AN INFECTION AT THE LEAD SITE AS REPORTED IN MFR. REPORT 3006630150-2025-03086 HOWEVER, IT IS UNKNOWN IF THE PATIENTS DEATH IS RELATED TO THE IMPLANTED SCS SYSTEM OR SCS PROCEDURE. ADDITIONALLY, THE PATIENT HAD DIABETES WHICH WAS A PRE-EXISTING CONDITION HOWEVER, IT IS ALSO UNKNOWN IF THIS CONDITION CONTRIBUTED TO THE DEATH. Manufacturer Narrative: ADDITIONAL SUSPECT MEDICAL DEVICE COMPONENT INVOLVED IN THE EVENT: PRODUCT FAMILY: SCS-LINEAR LEADS, UPN: M365SC2218700, MODEL: SC-2218-70, SERIAL: (B)(6), BATCH: 7104827, UDI: (B)(4). PRODUCT FAMILY: SCS-LINEAR LEADS, UPN: M365SC2218700, MODEL: SC-2218-70, SERIAL: (B)(6), BATCH: 7104783, UDI: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22145297&pc=MVK | 3008642652-2025-05482 | 2025-05-01 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-04 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 20:43:29 ON (B)(6) 2025. AT 03:16:53 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. AT 03:17:37, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. AT 03:22:22, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. AT 03:23:30, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT AND ELECTRODE LEAD FALLOFF. THE ELECTRODE BELT WAS DISCONNECTED AT 03:24:20 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22145304&pc=MVK | 3008642652-2025-05483 | 2025-04-08 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-04 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Incorrect Interpretation of Signal | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED FOUR INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 19:23:03 ON (B)(6) 2025. AT 22:40:47, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SVT @ 160 BPM WITH PVCS/NSVT. AT 22:48:33, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS SVT @ 150 BPM WITH PVCS. AT 22:49:00, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. SVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SVT @ 150 BPM. POST SHOCK RHYTHM WAS SVT @ 150 BPM. AT 22:49:27, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. SVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS SVT @ 150 BPM WITH PVCS. POST SHOCK RHYTHM WAS SVT @ 150 BPM WITH PVCS. AT 22:50:56, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS AF @ 130 BPM WITH NSVT/PVCS. AT 22:52:00, THE PATIENT RECEIVED THE FOURTH INAPPROPRIATE TREATMENT. NSVT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS NSVT. POST SHOCK RHYTHM WAS SINUS TACHYCARDIA @ 140 BPM. THE ELECTRODE BELT WAS DISCONNECTED AT 23:32:17 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22145327&pc=NIK | 2017865-2025-75446 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NIK | GALLANT VR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: D9 DATE SHOULD INDICATE MAY 16, 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22145328&pc=NVY | 2017865-2025-75447 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NVY | OPTISURE ACTIVE FIXATION, DF-4 CONNECTOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. VISUAL INSPECTION OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE. ELECTRICAL TESTING WAS NORMAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22145937&pc=LWP | 2017865-2025-75463 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED DUE TO ACUTE CARDIOPULMONARY ARREST. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO PATIENT¿S DEATH. THE DEVICE WAS RECEIVED WITH NORMAL TELEMETRY AND OUTPUT. ANALYSIS PERFORMED DID NOT IDENTIFY ANY FUNCTIONAL ISSUES. VISUAL INSPECTION OF THE EPOXY HEADER WAS PERFORMED, INDICATING AN UNUSUAL APPEARANCE OF THE EPOXY. MANUFACTURING INVESTIGATION WAS ALSO PERFORMED TO ASSESS THE EPOXY CONDITION, WHICH DETERMINED THAT THIS CONDITION IS ACCEPTABLE AND NOT PROBLEMATIC. LONGEVITY ASSESSMENT FOUND THE DEVICE WAS OPERATING ABOVE THE ELECTIVE REPLACEMENT INDICATOR (ERI) VOLTAGE WITH APPROPRIATE REMAINING LONGEVITY. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED, AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22145939&pc=NVN | 2017865-2025-75464 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED DUE TO ACUTE CARDIOPULMONARY ARREST. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING WAS NORMAL. VISUAL INSPECTION OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22146176&pc=NVN | 2017865-2025-75469 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: THE LEAD WAS RETURNED FOR PATIENT DEATH. AS RECEIVED ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING WAS NORMAL. VISUAL ANALYSIS DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22146183&pc=NVN | 2017865-2025-75470 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: THE LEAD WAS RETURNED FOR PATIENT DEATH. AS RECEIVED ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING WAS NORMAL. VISUAL ANALYSIS DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22146268&pc=NVN | 2017865-2025-75472 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NVN | ULTIPACE¿ LEAD, 52CM MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED DUE TO ACUTE RESPIRATORY FAILURE AND COMPLETE HEART BLOCK. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES. ELECTRICAL TESTING WAS NORMAL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22146269&pc=LWP | 2017865-2025-75471 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED DUE TO ACUTE RESPIRATORY FAILURE AND COMPLETE HEART BLOCK. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: DEVICE RETURNED FOR ANALYSIS DUE TO PATIENT DEATH. ANALYSIS PERFORMED, INCLUDING ELECTRICAL, MECHANICAL AND ENVIRONMENTAL TESTS INDICATED THAT THE DEVICE EXHIBITED NORMAL DEVICE CHARACTERISTICS WITH NO ANOMALIES. LONGEVITY ASSESSMENT WAS PERFORMED AND DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22146297&pc=NVN | 2017865-2025-75473 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NVN | ULTIPACE¿ LEAD, 58CM MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED DUE TO ACUTE RESPIRATORY FAILURE AND COMPLETE HEART BLOCK. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, A PARTIAL CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING WAS NORMAL. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR THE CUT END WHICH IS CONSISTENT WITH PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22146302&pc=LWP | 2017865-2025-75468 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO PATIENT EXPIRED. THE DEVICE WAS RECEIVED WITH A BATTERY LEVEL WITHIN THE NORMAL RANGE OF OPERATION. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. AFTER ALL ELECTRICAL TESTS PERFORMED, INCLUDING THERMAL AND MECHANICAL STRESS TESTING THE DEVICE EXHIBITS NORMAL CHARACTERISTICS. LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY. NO ANOMALIES WERE FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22146329&pc=LWP | 2017865-2025-75442 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF PENDING CONFIRMATION. Manufacturer Narrative: THE DEVICE WAS RETURNED DUE TO THE PATIENT'S DEATH. THE DEVICE IMAGE WAS SAVED SUCCESSFULLY. THE RESULTS OF THE ELECTRICAL AND STRESS/ENVIRONMENTAL TESTS WERE PERFORMED TO INDICATE THAT THE PACER IS EXHIBITING NORMAL DEVICE CHARACTERISTICS. A LONGEVITY ASSESSMENT WAS PERFORMED, AND THE DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22146330&pc=NVN | 2017865-2025-75444 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF PENDING CONFIRMATION. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22146331&pc=NVN | 2017865-2025-75443 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-04 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF PENDING CONFIRMATION. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD FOUND NO ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22129461&pc=BZD | 2518422-2025-046295 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-06-03 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Cancer; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem; Skin Inflammation/ Irritation | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, CANCER, LIVER CANCER AND RECTA L CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22129840&pc=PNJ | 9612164-2025-02695 | 2025-05-12 04:00 | Death | MEDTRONIC IRELAND | 2025-06-03 04:00 | PNJ | MICRA¿ AV2 | Deformation Due to Compressive Stress; Adverse Event Without Identified Device or Use Problem; Positioning Problem | Bacterial Infection; Respiratory Failure; Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. IT WAS REPORTED THAT THERE WAS OUT OF PLANE DEFLECTION ON THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) DELIVERY SYSTEM DURING AN ATTEMPTED IMPLANT. IT WAS NOTED IT WAS ALSO KINKED. IT WAS NOTED THE PATIENT PASSED AWAY THREE DAYS POST THE ATTEMPTED IMPLANT. IT WAS REPORTED THAT THERE WAS A SUCCESSFUL IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) ON THE SAME DAY.Event Description: IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS AND HYPER-CARBIC RESPIRATORY FAILURE AND METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS BACTEREMIA. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22130200&pc=OZP | 2032227-2025-184239 | 2025-04-28 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-03 04:00 | OZP | PUMP MMT-1884L MM780G BLE 3.0 MG/ML | Appropriate Term/Code Not Available | Hypoglycemia; Insufficient Information | P160017 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025 AND PUMP WAS NOT WORKING. THE CAUSE OF DEATH WAS UNKNOWN. THE CUSTOMER¿S BLOOD GLUCOSE WAS UNKNOWN. THE EVENT INVOLVED PRODUCT(S) MMT-332A, MMT-1884L, MMT-399A. TROUBLESHOOTING WAS NOT PERFORMED. IT WAS UNKNOWN WHETHER THE CUSTOMER HAD USED THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT AND WHETHER THE SMARTGUARD AUTO MODE OF THE INSULIN PUMP WAS USED. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A, MMT-1884L, MMT-399A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22130277&pc=OZP | 2032227-2025-184300 | 2025-03-29 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-03 04:00 | OZP | MINIMED 770G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Diabetic Ketoacidosis; Gastrointestinal Regurgitation | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY AT HOME ON (B)(6) 2025. THE CUSTOMER WAS NOT ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS DIABETES, DIABETIC KETOACIDOSIS AND GASTROPARESIS. CUSTOMER BLOOD GLUCOSE WAS 600 MG/DL AT THE TIME OF INCIDENT. THE LAST KNOWN PRODUCT(S) FOR THIS CUSTOMER ARE AS FOLLOWS: MMT-1880, MMT-213A, MMT-332A. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING, CUSTOMER WORN INSULIN PUMP AT THE TIME OF PASSING. CUSTOMER WAS USING THE INSULIN PUMP SYSTEM WITHIN 48HRS OF REPORTED EVENT AND CUSTOMER WAS NOT USING SENSORS. THERE IS NO MENTION OF THE AUTO MODE FEATURE AT THE TIME OF THE EVENT. NO PRODUCT RETURN IS REQUIRED FOR MMT-1880. NO PRODUCT RETURN IS REQUIRED FOR MMT-213A. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22130308&pc=BZD | 2518422-2025-046400 | 2023-10-02 04:00 | Death | RESPIRONICS, INC. | 2025-06-03 04:00 | BZD | BIPAP PRO | Degraded | Headache; Pulmonary Dysfunction; Dizziness; Unspecified Respiratory Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22130436&pc=BZD | 2518422-2025-046463 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-06-03 04:00 | BZD | REMSTAR PLUS SERIES ASSY | Degraded | Cancer | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22130450&pc=OZP | 2032227-2025-184349 | 2024-11-24 05:00 | Death | MEDTRONIC MINIMED | 2025-06-03 04:00 | OZP | UNKNOWN NGP | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Renal Failure |  |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER PASSED AWAY AT THE NURSING HOME ON (B)(6) 2024. THE CUSTOMER WAS ADMITTED TO A HOSPITAL PRIOR TO THE REPORTED INCIDENT. THE CAUSE OF DEATH WAS DUE TO KIDNEY DISEASE AND DIABETES. THE EVENT INVOLVED PRODUCT(S), AN UNKNOWN INSULIN PUMP. THE CUSTOMER¿S BLOOD GLUCOSE WAS UNKNOWN. THE CUSTOMER WAS WEARING THE INSULIN PUMP AT THE TIME OF DEATH. THE CUSTOMER WAS NOT USING SENSORS. IT IS UNKNOWN IF THE INSULIN PUMP WILL BE RETURNED FOR PRODUCT ANALYSIS. Manufacturer Narrative: CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22130692&pc=NPT | 2025587-2025-04014 | 2025-05-27 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-03 04:00 | NPT | EVOLUT FX PLUS VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Ventricular Fibrillation; Heart Block | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO EMERGENCY VALVE-IN-VALVE IMPLANT OF A TRANSCATHETER AORTIC VALVE, INTO A FAILED SURGICAL VALVE THAT CAUSED ACUTE PULMONARY EDEMA AND SEVERE AORTIC REGURGITATION, THE PATIENT HAD TO BE INTUBATED TO PERFORM THE PROCEDURE. THE PROCEDURE WAS CONSIDERED HIGH-RISK DUE TO THE PATIENT BEING UNSTABLE AND NEEDING CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) TO SUPPORT BREATHING. AFTER CROSSING THE SURGICAL VALVE, THE PATIENT BECAME VERY UNSTABLE. COMPLETE HEART BLOCK (CHB) DEVELOPED AND SIGNIFICANT HEMODYNAMIC COMPROMISE OCCURRED. THE VALVE WAS DEPLOYED SUCCESSFULLY, BUT THE PATIENT WENT INTO ASYSTOLE AND LOST CARDIAC OUTPUT. ONE SEQUENCE OF VENTRICULAR FIBRILLATION (VF) OCCURRED, AND THE PATIENT WAS SHOCKED OUT OF IT. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED FOR 30 MINUTES WITHOUT SUCCESS, AND THE PATIENT WAS DECLARED DECEASED. THE CAUSE OF DEATH WAS REPORTED TO BE ASYSTOLE DUE TO HEMODYNAMIC COMPROMISE AND SEVERE PULMONARY EDEMA FROM AORTIC REGURGITATION OF THE FAILED SURGICAL VALVE.Event Description: IT WAS REPORTED THAT PRIOR TO EMERGENCY VALVE-IN-VALVE IMPLANT OF A TRANSCATHETER AORTIC VALVE, INTO A FAILED SURGICAL VALVE THAT CAUSED ACUTE PULMONARY EDEMA AND SEVERE AORTIC REGURGITATION, THE PATIENT HAD TO BE INTUBATED TO PERFORM THE PROCEDURE. THE PROCEDURE WAS CONSIDERED HIGH-RISK DUE TO THE PATIENT BEING UNSTABLE AND NEEDING CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) TO SUPPORT BREATHING. AFTER CROSSING THE SURGICAL VALVE, THE PATIENT BECAME VERY UNSTABLE. COMPLETE HEART BLOCK (CHB) DEVELOPED AND SIGNIFICANT HEMODYNAMIC COMPROMISE OCCURRED. THE VALVE WAS DEPLOYED SUCCESSFULLY, BUT THE PATIENT WENT INTO ASYSTOLE AND LOST CARDIAC OUTPUT. ONE SEQUENCE OF VENTRICULAR FIBRILLATION (VF) OCCURRED, AND THE PATIENT WAS SHOCKED OUT OF IT. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED FOR 30 MINUTES WITHOUT SUCCESS, AND THE PATIENT WAS DECLARED DECEASED. THE CAUSE OF DEATH WAS REPORTED TO BE ASYSTOLE DUE TO HEMODYNAMIC COMPROMISE AND SEVERE PULMONARY EDEMA FROM AORTIC REGURGITATION OF THE FAILED SURGICAL VALVE. ADDITIONAL INFORMATION WAS RECEIVED WHICH CONFIRMED THAT THE SURGICAL VALVE WAS A NON-MEDTRONIC DEVICE. PER THE PHYSICIAN, NEITHER THE VALVE NOR THE DELIVERY CATHETER SYSTEM CAUSED OR CONTRIBUTED TO THE DEATH.Event Description: IT WAS REPORTED THAT PRIOR TO EMERGENCY VALVE-IN-VALVE IMPLANT OF A TRANSCATHETER AORTIC VALVE, INTO A FAILED SURGICAL VALVE THAT CAUSED ACUTE PULMONARY EDEMA AND SEVERE AORTIC REGURGITATION, THE PATIENT HAD TO BE INTUBATED TO PERFORM THE PROCEDURE. THE PROCEDURE WAS CONSIDERED HIGH-RISK DUE TO THE PATIENT BEING UNSTABLE AND NEEDING CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) TO SUPPORT BREATHING. AFTER CROSSING THE SURGICAL VALVE, THE PATIENT BECAME VERY UNSTABLE. COMPLETE HEART BLOCK (CHB) DEVELOPED AND SIGNIFICANT HEMODYNAMIC COMPROMISE OCCURRED. THE VALVE WAS DEPLOYED SUCCESSFULLY, BUT THE PATIENT WENT INTO ASYSTOLE AND LOST CARDIAC OUTPUT. ONE SEQUENCE OF VENTRICULAR FIBRILLATION (VF) OCCURRED, AND THE PATIENT WAS SHOCKED OUT OF IT. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED FOR 30 MINUTES WITHOUT SUCCESS, AND THE PATIENT WAS DECLARED DECEASED. THE CAUSE OF DEATH WAS REPORTED TO BE ASYSTOLE DUE TO HEMODYNAMIC COMPROMISE AND SEVERE PULMONARY EDEMA FROM AORTIC REGURGITATION OF THE FAILED SURGICAL VALVE. ADDITIONAL INFORMATION WAS RECEIVED WHICH CONFIRMED THAT THE SURGICAL VALVE WAS A NON-MEDTRONIC DEVICE. PER THE PHYSICIAN, NEITHER THE VALVE NOR THE DELIVERY CATHETER SYSTEM CAUSED OR CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED FROM THE PHYSICIAN WHICH CONFIRMED THAT THE NON-MEDTRONIC SURGICAL VALVE WAS A 23 MM PERIMOUNT (CARPENTIER-EDWARDS) AND WAS IMPLANTED TWELVE YEARS AGO (EXACT DATE NOT PROVIDED). THE PHYSICIAN ALSO CONFIRMED THAT THE FAILURE MODE OF THE PERIMOUNT VALVE WAS REGURGITATION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME EVOLUT FX DCS; PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012724252); PRODUCT TYPE: 0195-HEART VALVES; OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012728261); PRODUCT TYPE: 0195-HEART VALVES; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 (THIRD PARAGRAPH). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. SECOND PARAGRAPH ADDED H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22131046&pc=DSQ | 2916596-2025-03296 | 2025-05-09 04:00 | Death | THORATEC CORPORATION | 2025-06-03 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS AND MULTI-SYSTEM ORGAN FAILURE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE HEARTMATE 3 DEVICE WAS WORKING AS EXPECTED FOR THE WHOLE TIME OF SUPPORT. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION.Event Description: DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE FURTHER PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. B, AND PATIENT HANDBOOK, REV. B, ARE CURRENTLY AVAILABLE. THE IFU LISTS ADVERSE EVENTS, INCLUDING SEPSIS AND MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. ALTHOUGH ONLY SEPSIS WAS REPORTED, THE IFU LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION. THE IFU ALSO PROVIDES THE SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22131123&pc=DSQ | 2916596-2025-03335 | 2025-04-19 04:00 | Death | THORATEC CORPORATION | 2025-06-03 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Dyspnea; Tachycardia; Hypervolemia; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD FREQUENT VENTRICULAR TACHYCARDIA (VT) AND DID NOT THRIVE MUCH. THE PATIENT PRESENTED WITH VOLUME OVERLOAD AND SHORTNESS OF BREATH. THERE WAS PERSISTENT VASOPLEGIA WITH NO CLEAR ETIOLOGY FOUND. THE PATIENT'S RIGHT VENTRICLE WAS FAILING AND THEY WERE ON MAX VASOPRESSORS, SO THEIR FAMILY DECIDED IT WAS THE APPROPRIATE TIME TO WITHDRAW LIFE SUPPORT. COMFORT MEASURES WERE STARTED AND VASOPRESSORS STOPPED. EVENTUALLY, THE VENTRICULAR ASSIST DEVICE (VAD) STOPPED, AND THE PATIENT PASSED AWAY. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND OPERATED AS EXPECTED. THE PUMP WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS RIGHT HEART FAILURE, CARDIAC ARRHYTHMIA, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿RIGHT HEART FAILURE¿) OUTLINES INDICATIONS OF RIGHT HEART FAILURE AS WELL AS POSSIBLE TREATMENTS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22131553&pc=DSQ | 2916596-2025-03384 | 2025-05-02 04:00 | Death | THORATEC CORPORATION | 2025-06-03 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Respiratory Acidosis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CHRONIC RENAL FAILURE WITH SEVERE METABOLIC ACIDOSIS. THE DEATH WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DEATH OF THE PATIENT WAS NOT DEVICE RELATED. THE DEVICE WAS NOT EXPLANTED, AND AN AUTOPSY WAS NOT PERFORMED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE RENAL DYSFUNCTION WAS NOT DETERMINED TO BE RELATED OR CAUSED BY THE DEVICE OR DEVICE THERAPY. THEREFORE, THE HOSPITAL WILL NOT SHARE FURTHER INFORMATION REGARDING ONSET, SYMPTOMS NOR TREATMENT OF THAT DISEASE. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RENAL DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22131695&pc=DSQ | 2916596-2025-03474 | 2025-05-13 04:00 | Death | THORATEC CORPORATION | 2025-06-03 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Fall; Intracranial Hemorrhage; Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT A COMPUTED TOMOGRAPHY (CT) SCAN ON (B)(6) 2025 SHOWED SUBDURAL HEMATOMA WITH BRAIN STERN HERNIATION, THOUGHT TO BE FROM A MECHANICAL FALL. THE DECISION WAS MADE TO TURN OFF THE PUMP AND TERMINALLY EXTUBATE. THE PATIENT PASSED AWAY THE SAME DAY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE LVAS REPORTEDLY WORKED AS INTENDED, AND IT WAS REPORTED THAT THE PATIENT¿S DEATH WAS NOT DEVICE RELATED. THE DEVICE WILL REPORTEDLY NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS BLEEDING, STROKE, AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22131732&pc=DSQ | 2916596-2025-03655 | 2023-06-09 04:00 | Death | THORATEC CORPORATION | 2025-06-03 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED BY THE CUSTOMER.Event Description: IT WAS LATER REPORTED THAT THE PATIENT WAS FOUND AT HOME NOT BREATHING AND WAS PRONOUNCED DEAD AT HOME. IT WAS SAID THAT DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED AND THAT THE DEVICE WAS OPERATING AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS INSTRUCTIONS FOR USE (IFU), REV. A, IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22131918&pc=LYJ | 1644487-2025-10201 | 2025-05-08 04:00 | Death | HOUSTON | 2025-06-03 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22132059&pc=MIH | 3013164176-2025-02539 | 2025-05-02 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-03 04:00 | MIH | GORE® EXCLUDER® AAA ENDOPROSTHESIS | Insufficient Information | Stroke/CVA | P020004 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, A PATIENT WAS TREATED USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) SYSTEM. A GORE® EXCLUDER® AAA ENDOPROSTHESIS CONTRALATERAL LEG WAS ALSO USED IN THE AORTA. PROXIMAL ACCESS WAS GAINED THROUGH A GORE® DRYSEAL FLEX INTRODUCER SHEATH IN THE RIGHT SUBCLAVIAN ARTERY. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. EARLY ON (B)(6) 2025, WITHIN 24 HOURS OF THE INITIAL PROCEDURE, THE PATIENT HAD SOME WATERSHED INFARCTS. A CT OF THE CHEST SHOWED NOTHING UNUSUAL. ON (B)(6) 2025, THE PATIENT STARTED HAVING SEIZURES AND WAS INTUBATED. ACCORDING TO THE FSA, THERE WAS NO NOTABLE THROMBUS. POST-OPERATIVE IMAGES ARE NOT AVAILABLE.Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, A PATIENT WAS TREATED USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) SYSTEM. A GORE® EXCLUDER® AAA ENDOPROSTHESIS CONTRALATERAL LEG WAS ALSO USED IN THE AORTA. PROXIMAL ACCESS WAS GAINED THROUGH A GORE® DRYSEAL FLEX INTRODUCER SHEATH IN THE RIGHT SUBCLAVIAN ARTERY. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. EARLY ON (B)(6) 2025, WITHIN 24 HOURS OF THE INITIAL PROCEDURE, THE PATIENT HAD SOME WATERSHED INFARCTS. A CT OF THE CHEST SHOWED NOTHING UNUSUAL. ON (B)(6) 2025, THE PATIENT STARTED HAVING SEIZURES AND WAS INTUBATED. ACCORDING TO THE FSA, THERE WAS NO NOTABLE THROMBUS. ON (B)(6) 2025, THE PHYSICIAN REPORTED THAT THE PATIENT HAD DIED, SAYING THAT THE PATIENT NEVER RECOVERED FROM THE STROKE. Manufacturer Narrative: B2. OUTCOMES UPDATED TO DEATH. B4. EVENT DESCRIPTION UPDATED. G3/G4. DATE MANUFACTURER RECEIVED BY MANUFACTURER IS THE DATE WE LEARNED OF THE DEATH. H1. TYPE OF REPORTABLE EVENT UPDATED TO DEATH. H.6. HEALTH EFFECT - IMPACT CODE - ADDED F02. Manufacturer Narrative: H.6. INVESTIGATION FINDINGS CODE C20: THE DEVICE REMAINS IMPLANTED AND IS NOT AVAILABLE FOR ANALYSIS. H.6. INVESTIGATION FINDINGS CODE C19: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H.6. INVESTIGATION CONCLUSIONS CODE D12: IT SHOULD BE NOTED THAT, PER THE GORE® EXCLUDER® AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE, POSSIBLE ADVERSE EVENTS AND COMPLICATIONS ASSOCIATED WITH THE USE OF THE GORE® EXCLUDER® AAA ENDOPROSTHESIS THAT MAY OCCUR INCLUDE, BUT ARE NOT LIMITED TO: STROKE. . W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY HAVE NOT BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22132544&pc=DTK | 0001625425-2025-01023 | 2025-06-02 04:00 | Death | ARGON MEDICAL DEVICES | 2025-06-03 04:00 | DTK | OPTION ELITE RETRIEVABLE VENA CAVA FILTER | Difficult to Fold, Unfold or Collapse; Migration | Foreign Body In Patient | K133243 |  | 1 | Event Description: A CUSTOMER REPORTED "ON THE AFTERNOON OF JUNE 2ND, THE DOCTOR USED AN ARC-SHAPED PUSHER TO PUSH THE FILTER OUT OF THE FILTER STORAGE COMPARTMENT, BUT THE FILTER DID NOT OPEN. IT THEN MIGRATED TO THE ATRIUM AND FINALLY OPENED THERE. THE DOCTOR ATTEMPTED TO REMOVE THE FILTER, BUT IT ENDED IN FAILURE. THE PATIENT WAS ADMITTED TO THE ICU AND DIED AROUND THE EARLY MORNING OF THE 3RD." Manufacturer Narrative: THE SAMPLE IS NOT AVAILABLE FOR RETURN ALTHOUGH AN IMAGE WAS PROVIDED. THE INVESTIGATION IS CURRENTLY IN PROGRESS. A FOLLOW-UP REPORT WILL BE SUBMITTED UPON INVESTIGATION COMPLETION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22132599&pc=OUT | 2029214-2025-01293 | 2025-05-23 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-06-03 04:00 | OUT | PIPELINE VANTAGE WITH SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Vomiting; Loss of consciousness; Hemorrhagic Stroke | P100018 |  | 1 | Event Description: MEDTRONIC RECEIVED A REPORT THAT THE PATIENT WAS UNDERGOING TREATMENT OF AN AMORPHOUS, UNRUPTURED LEFT VERTEBRAL ARTERY ANEURYSM WITH A MAX DIAMETER OF 4.5MM AND A 3.9MM NECK DIAMETER. IT WAS NOTED THE PATIENT'S VESSEL TORTUOSITY WAS SEVERE. THE LANDING ZONE WAS 1.7MM DISTALLY AND 2MM PROXIMALLY. DUAL ANTIPLATELET TREATMENT (DAPT) WAS ADMINISTERED. THE PRU LEVEL WAS UNKNOWN. IT WAS REPORTED THAT THE ACCESS PAST THE ANEURYSM WAS NOT ABLE TO BE ACHIEVED FROM THE LEFT SIDE. ATTEMPTED COIL PLACEMENT BUT COILS WOULD NOT STAY IN THE ANEURYSM. THE PHYSICIAN DECIDED TO CROSS FROM THE RIGHT VERTEBRAL ARTERY AND DEPLOY A PIPELINE VANTAGE. THE DEVICE WAS DEPLOYED WITHOUT INCIDENT. AFTER THE PROCEDURE WAS CONCLUDED, THE PATIENT VOMITED AND BECAME UNRESPONSIVE IN RECOVERY. A CT SCAN WAS CONDUCTED, AND A HEMORRHAGE WAS SEEN. EXTERNAL VENTRICULAR DRAIN (EVD) WAS PLACED, AND THE PATIENT BECAME RESPONSIVE. APPROX 30 MINS LATER, THE PATIENT BECAME UNRESPONSIVE AGAIN AND THE PUPILS WERE FIXED. THE ANGIOGRAPHIC RESULT POST PROCEDURE SHOWED GOOD APPOSITION, AND NO COMPLICATIONS NOTED. THE PIPELINE USED FOR AN INDICATION THAT IS NOT APPROVED (OFF-LABEL). THE REPORTED DEVICE AND ANY ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THE PATIENT DECEASED.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED CATHETER WAS AN SL10 J SHAPE. COILS ATTEMPTED WERE FC-3.5-6, STRYKER TETRA 4X8. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22132622&pc=NVN | 2649622-2025-14943 | 2025-05-16 04:00 | Death | MPRI | 2025-06-03 04:00 | NVN | SELECTSECURE MRI SURESCAN | Over-Sensing; Under-Sensing | Syncope/Fainting | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. THE PATIENT EXPERIENCED SYNCOPE. THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING AND OVERSENSING DURING VENTRICULAR ARRHYTHMIA. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22132672&pc=NIK | 9614453-2025-02144 | 2025-04-28 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-06-03 04:00 | NIK | COBALT¿ XT HF CRT-D MRI SURESCAN¿ | Insufficient Information | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED WHILE INTUBATED POST-TOE AMPUTATION SURGERY. THE PATIENT'S FAMILY MEMBER REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) MAY HAVE BEEN "DAMAGED" OR NOT WORKING PROPERLY DURING THE PATIENT'S PASSING AND BELIEVED THE DEVICE MAY HAVE CAUSED THE PATIENT'S DEATH. ADDITIONAL INFORMATION WAS REQUESTED BUT COULD NOT BE OBTAINED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22133895&pc=NVZ | 2017865-2025-75315 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-03 04:00 | NVZ | ASSURITY RF DR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P880086 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. THE PACEMAKER AND LEADS WERE EXPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22133896&pc=NVN | 2017865-2025-75316 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-03 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. THE PACEMAKER AND LEADS WERE EXPLANTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22133906&pc=NVN | 2017865-2025-75317 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-03 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. THE PACEMAKER AND LEADS WERE EXPLANTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22134137&pc=DSQ | 2916596-2025-03543 | 2025-01-03 05:00 | Death | THORATEC CORPORATION | 2025-06-03 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE PATIENT OUTCOME, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO HOSPITAL PRIVACY POLICY, NO ADDITIONAL INFORMATION WOULD BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22134373&pc=NVN | 2017865-2025-75326 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-03 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. THE PACEMAKER AND LEADS WERE EXPLANTED. Manufacturer Narrative: AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED FOR ANALYSIS. ELECTRICAL TESTING WAS NORMAL. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE. NO OTHER ANOMALIES WERE FOUND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22134374&pc=NVN | 2017865-2025-75327 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-03 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. THE PACEMAKER AND LEADS WERE EXPLANTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DECEASED. AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES EXPECT FOR PROCEDURE DAMAGE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22134375&pc=NVZ | 2017865-2025-75325 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-03 04:00 | NVZ | ASSURITY RF DR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P880086 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. THE PACEMAKER AND LEADS WERE EXPLANTED. Manufacturer Narrative: DEVICE RETURNED FOR ANALYSIS DUE TO PATIENT DEATH. ANALYSIS PERFORMED, INCLUDING ELECTRICAL, MECHANICAL AND ENVIRONMENTAL TESTS INDICATED THAT THE DEVICE EXHIBITED NORMAL DEVICE CHARACTERISTICS WITH NO ANOMALIES. LONGEVITY ASSESSMENT WAS PERFORMED AND DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22134579&pc=NVZ | 2017865-2025-75329 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-03 04:00 | NVZ | ENDURITY DR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P880086 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: DEVICE RETURNED FOR ANALYSIS DUE TO PATIENT DEATH. ANALYSIS PERFORMED, INCLUDING ELECTRICAL, MECHANICAL AND ENVIRONMENTAL TESTS INDICATED THAT THE DEVICE EXHIBITED NORMAL DEVICE CHARACTERISTICS WITH NO ANOMALIES. LONGEVITY ASSESSMENT WAS PERFORMED AND DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22134600&pc=MVK | 3008642652-2025-05418 | 2025-05-09 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-03 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 21:08:12 ON (B)(6) 2025. AT 01:22:16 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. AT 01:23:36, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL, CPR/MOTION ARTIFACT, AND ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH INTERMITTENT CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. THE ELECTRODE BELT WAS DISCONNECTED AT 01:54:32 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22134834&pc=DSQ | 2916596-2025-03302 | 2025-05-07 04:00 | Death | THORATEC CORPORATION | 2025-06-03 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Respiratory Insufficiency | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RESPIRATORY ISSUES POST RESPIRATORY SYNCYTIAL VIRUS (RSV) AND FAILURE TO THRIVE. THE FAMILY MADE THE DECISION TO WITHDRAW CARE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22135121&pc=OZD | 1220648-2025-28953 | 2025-05-09 04:00 | Death | ABIOMED, INC. | 2025-06-03 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT WHILE IMPLANTED WITH AN IMPELLA CP, THE PATIENT EXHIBITED SIGNS OF HEMOLYSIS. THE PATIENT WAS TRANSFUSED TWO UNITS OF PACKED RED BLOOD CELLS. ULTIMATELY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22135360&pc=OZD | 1220648-2025-28955 | 2025-05-06 04:00 | Death | ABIOMED, INC. | 2025-06-03 04:00 | OZD | IMPELLA CP | Restricted Flow rate | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT FOLLOWING IMPELLA CP IMPLANT, LOW FLOWS WERE ENCOUNTERED. THE PUMP WAS REPLACED IN RESPONSE TO THE FAILURE. THERE WAS NO INDICATION OF A CLOT ON REMOVAL. THE PATIENT'S OUTCOME WAS UPDATED TO EXPIRED FOLLOWING THE EVENTS. ADDITIONAL INFORMATION WAS REQUESTED BUT NOT PROVIDED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22135552&pc=DSQ | 2916596-2025-03303 | 2025-05-09 04:00 | Death | THORATEC CORPORATION | 2025-06-03 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Pneumonia; Sepsis; Septic Shock | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. THE PATIENT REMAINED HOSPITALIZED FOR PNEUMONIA TREATMENT POST THE IMPLANT AND WAS BEING TREATED WITH ANTIBIOTICS. THE PATIENT PASSED AWAY DUE TO SEPTIC SHOCK ON (B)(6) 2025. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED AND THE DEVICE OPERATED AS EXPECTED BEFORE THE PATIENT HAD THE SEPTIC SHOCK. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22136114&pc=DSQ | 2916596-2025-03389 | 2025-05-10 04:00 | Death | THORATEC CORPORATION | 2025-06-03 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. IT WAS COMMUNICATED THAT DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WILL NOT COMMUNICATE ANY FURTHER PATIENT OUTCOME INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22136193&pc=DQX | 2126666-2025-00027 | 2025-05-21 04:00 | Death | LAKE REGION MEDICAL | 2025-06-03 04:00 | DQX | SAFARI2¿ | Use of Device Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Perforation; Pericardial Effusion | K151244 |  | 1 | Event Description: EVENT DESCRIPTION: DURING THE TAVI PROCEDURE, WE WERE IN THE VALVE RELEASE STAGE IN PART 2. WHEN AN ANESTHESIOLOGIST REPORTED THAT THE PATIENT'S BLOOD PRESSURE WAS DROPPING, AT THAT MOMENT THE CONCERN WAS TO MAINTAIN TENSION AND RELEASE THE VALVE COMPLETELY, WE RELEASED THE VALVE AND THEN THE ECHO DETECTED PERICARDIAL EFFUSION. THE CHANGE WAS MADE; THEY CALLED THE HEART MUSCLE AND OPENED THE PATIENT. BLOOD WAS DETECTED DEEP IN THE LEFT VENTRICLE, THE CATHETER MANAGED TO CONTAIN IT BUT THEN ANOTHER PERFORATION APPEARED IN THE RIGHT VENTRICLE, TEARING THE DISTAL PART OF THE HEART. PATIENT DIED ACCIDENTALLY. LOOKING AT THE FACTS, IMAGES AND WHAT THE APPEAL SAID, THE SAFARI GUIDE WIRE PIERCED THE LEFT VENTRICLE CAUSING ALL OF THIS. PATIENT PRESENT AT TIME OF EVENT? YES. PATIENT COMPLICATIONS: PERICARDIAL EFFUSION. IN THE PHYSICIAN'S OPINION, DID THE DEVICE OR PROCEDURE CAUSE OR CONTRIBUTE TO THE PATIENT COMPLICATION? YES. PLEASE DESCRIBE THE WAY IN WHICH THE DEVICE OR PROCEDURE CAUSED OR CONTRIBUTED TO THE PATIENT COMPLICATION? THE SAFARI GUIDE WIRE ENDED UP PIERCING THE LEFT VENTRICLE AND CAUSING PERICARDIAL EFFUSION AFTER VALVE DEPLOYMENT. MEDICAL OR SURGICAL INTERVENTIONS: A CARDIAC SURGEON WAS CALLED TO INTERVENE THE LEFT VENTRICLE. PATIENT OUTCOME: DEATH. PROCEDURE NUMBER: PN-2666356. PATIENT STATUS: THE PATIENT IS DECEASED WITH AN OFFICIAL CAUSE OF DEATH OF PERICARDIAL EFFUSION. 29MAY2025: QUESTION: IS THIS A NEW OR EXPERIENCED USER? ANSWER: EXPERIENCED. QUESTION: WERE THERE ANY EARLY SIGNS OF COMPLICATIONS BEFORE THE DROP IN BLOOD PRESSURE AND PERICARDIAL EFFUSION WERE IDENTIFIED? ANSWER: NO. QUESTION: WERE THERE ANY OTHER PROCEDURE COMPLICATIONS PRIOR TO THE DROP IN BLOOD PRESSURE AND PERICARDIAL EFFUSION? ANSWER: NO. QUESTION: WAS IT DETERMINED VIA IMAGING OR IS THERE IMAGING AVAILABLE TO REVIEW AND DETERMINE WHEN DURING THE PROCEDURE THE PERFORATION OF THE LEFT VENTRICLE OCCURRED? ANSWER: PROCEDURAL ANGIOGRAPHIC MEDIA WILL BE SUBMITTED TO BSC. QUESTION: AFTER THE INITIAL RELEASE OF THE VALVE THE PATIENT BECAME HYPOTENSIVE. WHAT DO THEY ATTRIBUTE THE HYPOTENSION TO, AT THIS STAGE? ANSWER: UNKNOWN. QUESTION: THE REPORT STATES THAT LOOKING AT THE FACTS, IMAGES, AND APPEAL - THE SAFARI WIRE PIERCED THE LEFT VENTRICLE CAUSING ALL OF THIS. CAN YOU PROVIDE THE IMAGING OF THE WIRE POSITION AND THE PROCEDURAL STEP WHEN THEY BELIEVE THE PIERCING OCCURRED? ANSWER: PROCEDURAL ANGIOGRAPHIC MEDIA WILL BE SUBMITTED TO BSC. QUESTION: THE REPORTS STATES THAT THE SAFARI WIRE CAUSED THE PERFORATION OF THE LEFT VENTRICLE. WHAT DOES THE PHYSICIAN/END USER THINK CAUSED THE PERFORATION IN THE RIGHT VENTRICLE? ANSWER: UNKNOWN. QUESTION: THE REPORT SUGGESTS THE TEAR IN THE RIGHT VENTRICLE TORE THE DISTAL PART OF THE HEART, PATIENT DIED ACCIDENTALLY' WHAT DOES ACCIDENTALLY MEAN IN THE CONTEXT OF THIS REPORTED DEATH ALLEGEDLY DUE TO THE SAFARI WIRE? ANSWER: THE PATIENT'S DEATH WAS NOT INTENTIONAL. QUESTION: WHEN THE PATIENT BECAME HYPOTENSIVE AFTER THE INITIAL RELEASE THE MAIN CONCERN WAS TO MAINTAIN TENSION AS THE PATIENT BLOOD PRESSURE WAS DROPPING. MAINTAIN TENSION OF WHAT? WAS IT A DEVICE THAT WAS UNDER TENSION'? ANSWER: YES. QUESTION: IF YES, WHAT DEVICE AND DO THEY HAVE IMAGING OF THE DEVICE UNDER TENSION TO CONFIRM CORRECT POSITIONING THROUGHOUT THE COMPLETE RELEASE OF THE VALVE (RELEASE STAGE PART 2)? ANSWER: YES, THE ACCURATE NEO2 SYSTEM. QUESTION: DO THEY HAVE CINE OR FLUOR IMAGING OF THE WIRE POSITION AS THEY RELEASED THE VALVE COMPLETELY (RELEASE STAGE PART 2)? ANSWER: PROCEDURAL ANGIOGRAPHIC MEDIA WILL BE SUBMITTED TO BSC. QUESTION: LISTING AND MODEL OF OTHER DEVICES USED DURING THE PROCEDURE, INCLUDE THE MODEL, BRAND NAME, SIZES, ETC. ANSWER: ISLEEVE, ACURATE NEO2 VALVE, ACURATE NEO2 DELIVERY SYSTEM. QUESTION: WAS THERE ANY RESISTANCE NOTED DURING ADVANCING THE SAFARI2 GUIDEWIRE AND/OR PRIOR TO THE VALVE RELEASE? ANSWER: NO RESISTANCE OR DIFFICULTIES ADVANCING SAFARI2. QUESTION: WAS BAV PERFORMED OVER THE WIRE AND IF SO, WAS IT COMPETED SUCCESSFULLY WITHOUT ANY COMPLICATIONS? ANSWER: YES. QUESTION: WAS THERE ANY DAMAGE NOTED TO THE GUIDEWIRE AND/OR OTHER DEVICES PRIOR TO OR AFTER THE PROCEDURE? ANSWER: NO.Event Description: ON 06 JUN 2025: QUESTION: HOW MUCH BLOOD WAS DRAINED FROM LEFT VENTRICLE? ANSWER: I DON'T HAVE THAT INFORMATION, ESPECIALLY BECAUSE WHILE THEY WERE DRAINING THE HEART SURGEON QUICKLY ENTERED THE FIELD. QUESTION: WAS THERE BLOOD IN THE RIGHT VENTRICLE? WAS IT DRAINED? HOW MUCH WAS DRAINED? ANSWER: AS FAR AS I KNOW, THE EXIT WAS IN THE LEFT VENTRICLE. AND THEY DIDN'T DRAIN IT. QUESTION: RAPID PACING -WHAT RIGHT TEMP PACING WIRE WAS USED? ANSWER: THEY USED THE SAFARI AS A GUIDE FOR PACING. QUESTION: LISTING OF ALL INTERVENTION(S) PERFORMED? ANSWER: AFTER THE ECHOCARDIOGRAPHER MENTIONED THE PERICARDIAL EFFUSION, THE CARDIAC SURGEON ENTERED THE ROOM AND BEGAN THE SURGERY, TRYING TO FIX WHERE THE PUNCTURE HAD OCCURRED, WHICH WAS IN THE LEFT VENTRICLE, WHICH HE HIMSELF MENTIONED WAS PROBABLY CAUSED BY THE SAFARI GUIDE PUTTING TOO MUCH STRAIN ON HIM. QUESTION: WAS THE CURVE OF THE SAFARI2 GUIDEWIRE CONSTRAINED WITHIN A CATHETER DURING INSERTION INTO AND WITHDRAWAL FROM THE BODY OR TREATMENT SITE? ANSWER: NO, AND THERE WASN'T ANY DAMAGED PART EITHER. QUESTION: WAS THE WIRE POSITION MAINTAINED WHILE TRACKING OR ADVANCING THE CATHETER OR OTHER DEVICES OVER THE WIRE? ANSWER: YES, BUT FROM WHAT THE DR MENTIONED, HE BELIEVES THAT THE FIRST OPERATOR PLACED TOO MUCH TENSION ON THE GUIDE WIRE WHEN RELEASING THE VALVE. QUESTION: DID THE END USER MONITOR THE WIRE POSITION VIA FLUOROSCOPY THROUGHOUT THE PROCEDURE FOR PROPER PLACEMENT OF THE CURVE AND DISTAL TIP? ANSWER: YES, THEY DID, BUT DURING THE RELEASE OF THE VALVE IN THE FINAL STAGE, THEY WERE VERY CONCERNED ABOUT THE POSITIONING OF THE VALVE AND ENDED UP PUTTING A LOT OF TENSION ON THE GUIDE WIRE DUE TO THEIR ERROR, I BELIEVE.Event Description: ADDITIONAL INFORMATION AND MEDIA WERE RECEIVED FROM THE DISTRIBUTOR: QUESTION: WAS ECHO TAKEN PRE TO OR AFTER FULL VALVE DEPLOYMENT? ANSWER: IT WAS CARRIED OUT THROUGHOUT THE PROCEDURE. QUESTION: WAS MORE THAN ONE ECHO PERFORMED? ANSWER: JUST ONE ECHO PERFORMED. SUPPLEMENTAL INFORMATION RECEIVED REPORTED, 'THE ECHOCARDIOGRAPHER WAS IN THE ROOM DURING THE ENTIRE PROCEDURE, BUT HE WAS MORE FOCUSED ON THE FINAL PART OF THE PROCEDURE, CHECKING THE RESULTS AND POSITIONING. WHEN I SAID ONLY ONE ECHO WAS PERFORMED, I WAS REFERRING TO ONLY ONE PROFESSIONAL.' QUESTION: WAS CPR OR CARDIAC MASSAGE PERFORMED? ANSWER: YES THEY DID. QUESTION: IT WAS INITIALLY REPORTED: 'WHEN AN ANESTHESIOLOGIST REPORTED THAT THE PATIENT'S BLOOD PRESSURE WAS DROPPING, AT THAT MOMENT THE CONCERN WAS TO MAINTAIN TENSION AND RELEASE THE VALVE COMPLETELY, WE RELEASED THE VALVE AND THEN THE ECHO DETECTED PERICARDIAL EFFUSION. THE CHANGE WAS MADE, THEY CALLED THE HEART MUSCLE AND OPENED THE PATIENT.' CAN YOU PLEASE CLARIFY OR RESTATE THE UNDERLINED SENTENCE? ANSWER: I REMEMBER THAT MOMENT, THEY REMOVED THE DELIVERY SYSTEM AND QUICKLY CALLED THE CARDIAC SURGEON. QUESTION: IT WAS INITIALLY REPORTED: 'LOOKING AT THE FACTS, IMAGES AND WHAT THE APPEAL SAID, THE SAFARI GUIDE WIRE PIERCED THE LEFT VENTRICLE CAUSING ALL OF THIS.' WHAT DOES 'THE APPEAL' REFER TO? ANSWER: THEY BELIEVE THE SAFARI PUNCTURED THE VENTRICLE DUE TO EXCESSIVE FORCE. QUESTION: CAN YOU CONFIRM IF THE WIRE WAS CONSTRAINED WITHIN A CATHETER DURING INSERTION? ANSWER: NO, THE SAFARI HAD NO DEFECTS. QUESTION: WHAT ARE THE LEVELS OF EXPERIENCE FOR: ANSWERS: A. DR- STARTING THERAPY HE IS A NEW IMPLANTER OF TAVI. B. DR.-HAVE SOME EXPERTISE IN TAVI (ACURATE AND ANOTHER'S VALVES). C. HEART SURGEON - EXPERT. QUESTION: CAN YOU CONFIRM NO ADDITIONAL PROCEDURAL MEDIA IS AVAILABLE? ANSWER: THEY DO NOT PROVIDE IT DUE TO HOSPITAL COMPLIANCE ISSUES. Manufacturer Narrative: PRODUCT WAS REPORTED AS LOST AND WAS NOT RECEIVED FOR EVALUATION; THEREFORE, NO VISUAL OR PHYSICAL ANALYSIS OF THE PRODUCT COULD BE PERFORMED. A REVIEW OF THE DEVICE HISTORY RECORDS OF THE SPECIMEN DEVICE DOES NOT PRESENT ANY INDICATION OF MANUFACTURING DEFECT OR ANOMALY THAT COULD HAVE IMPACTED ON THE EVENT AS REPORTED. DURING MANUFACTURING AND PACKAGING OF THIS PRODUCT THE OPERATORS ARE INSTRUCTED TO 100% VISUALLY INSPECT FOR ANY OBVIOUS DEFECT PRIOR TO SHIPMENT. THE HISTORY RECORDS INDICATE THIS PRODUCT WAS FINAL INSPECTION TESTED AT LAKE REGION MEDICAL AND WAS DETERMINED TO BE ACCEPTABLE. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE (DFU) WARNINGS SECTION: MONITOR WIRE POSITION THROUGHOUT THE PROCEDURE FOR PROPER PLACEMENT OF THE CURVE AND DISTAL TIP. WHEN ADVANCING OR REMOVING THE GUIDEWIRE, ALWAYS USE FLUOROSCOPIC GUIDANCE WITH RADIOGRAPHIC EQUIPMENT THAT PROVIDES HIGH-RESOLUTION IMAGES. NEVER POSITION THE GUIDEWIRE BLINDLY, AS THIS MAY RESULT IN MISPLACEMENT, DISSECTION OR PERFORATION. NEVER ADVANCE THE GUIDEWIRE AGAINST RESISTANCE WITHOUT FIRST DETERMINING THE REASON FOR RESISTANCE UNDER FLUOROSCOPY. EXCESSIVE FORCE AGAINST RESISTANCE MAY RESULT IN DAMAGE TO THE CATHETER OR VESSEL/ORGAN. CARE SHOULD BE TAKEN WHEN ADVANCING THE GUIDEWIRE AFTER DEVICE DEPLOYMENT. THE WIRE SHOULD ONLY BE INTRODUCED INTO OR WITHDRAWN FROM THE VENTRICLE THROUGH A CATHETER ALREADY POSITIONED IN THE VENTRICLE. THE CURVE OF THE SAFARI2 GUIDEWIRE SHOULD BE CONSTRAINED WITH IN A CATHETER DURING INSERTION INTO OR WITHDRAWAL FROM THE BODY OR TREATMENT SITE. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE: 1. ENSURE A CATHETER IS APPROPRIATELY PLACED IN THE VENTRICLE OR OTHER INTENDED TREATMENT AREA. 2. CAREFULLY REMOVE THE SAFARI2 GUIDEWIRE, FROM THE HOOP BY GRASPING THE PROXIMAL END OF THE J-STRAIGHTENER SEPARATING THE STRAIGHTENER FROM THE HOOP. 3. AFTER INSPECTION OF THE SAFARI2 GUIDEWIRE, ADVANCE THE J-STRAIGHTENER OVER THE DISTAL SECTION OF THE SAFARI2 WIRE TO STRAIGHTEN THE CURVE. 4. INSERT THE SAFARI2 GUIDEWIRE INTO THE HUB OF THE CATHETER WITH THE AID OF THE J- STRAIGHTENER. 5. CAREFULLY ADVANCE THE DISTAL TIP OF THE SAFARI2 GUIDEWIRE INTO THE LUMEN OF THE CATHETER. 6. REMOVE THE SAFARI2 GUIDEWIRE J- STRAIGHTENER BY WITHDRAWING IT OVER THE LENGTH OF THE SAFARI2 GUIDEWIRE. 7. ADVANCED THE SAFARI2 GUIDEWIRE THROUGH THE CATHETER UNDER FLUOROSCOPIC GUIDANCE. 8. CONFIRM SAFARI2 GUIDEWIRE CURVE POSITION TO ASSURE SAFARI2 GUIDEWIRE PLACEMENT AND STABILITY. 9. MAINTAIN WIRE POSITION WHILE TRACKING OR ADVANCING A CATHETER OR DEVICE OVER THE WIRE. 10. TO REMOVE THE GUIDEWIRE FROM THE TREATMENT AREA: A. A CATHETER MUST BE ADVANCED INTO THE TREATMENT AREA OVER THE SAFARI2 GUIDEWIRE PRIOR TO WITHDRAWING THE WIRE. B. THE SAFARI2 GUIDEWIRE IS PULLED BACK COMPLETELY INTO THE CATHETER. C. THE SAFARI2 GUIDEWIRE MAY THEN BE WITHDRAWN FROM THE CATHETER. AT THIS TIME, IT IS NOT POSSIBLE TO ASSIGN A DEFINITIVE ROOT CAUSE FOR THE EVENT AS REPORTED. BASED ON THE INFORMATION PROVIDED, IT APPEARS THAT CLINICAL AND/OR PROCEDURAL FACTORS HAVE CONTRIBUTED TO THE EVENT. CARDIAC PERFORATION, PERICARDIAL EFFUSION, AND DEATH ARE KNOWN POTENTIAL ADVERSE EVENTS AND ARE LISTED IN THE SAFARI2 INSTRUCTIONS FOR USE (IFU). SHOULD ADDITIONAL INFORMATION BE RECEIVED, A FOLLOW UP MEDWATCH REPORT WILL BE SUBMITTED. ALTHOUGH ANNEX G INSTRUCTIONS STATE THAT WITH STAND-ALONE DEVICES ANNEX G TERM SHOULD NOT BE USED, CODE 4755 WAS SELECTED DUE TO COMPONENT CODE (G) SHOWING UP ON THE MISSING DATA REPORT. Manufacturer Narrative: THIS MEDWATCH REPORT IS AN UPDATE TO THE PREVIOUS REPORT TO INCLUDE THE ADDITIONAL INFORMATION RECEIVED ON 06JUN2025 IN SECTION B(5). ADDITIONALLY, SECTION H(11) HAS BEEN UPDATED TO INCLUDE THE INTENDED USE STATEMENT FROM THE DEVICE INSTRUCTIONS FOR USE. PRODUCT WAS REPORTED AS LOST AND WAS NOT RECEIVED FOR EVALUATION; THEREFORE, NO VISUAL OR PHYSICAL ANALYSIS OF THE PRODUCT COULD BE PERFORMED. A REVIEW OF THE DEVICE HISTORY RECORDS OF THE SPECIMEN DEVICE DOES NOT PRESENT ANY INDICATION OF MANUFACTURING DEFECT OR ANOMALY THAT COULD HAVE IMPACTED ON THE EVENT AS REPORTED. DURING MANUFACTURING AND PACKAGING OF THIS PRODUCT THE OPERATORS ARE INSTRUCTED TO 100% VISUALLY INSPECT FOR ANY OBVIOUS DEFECT PRIOR TO SHIPMENT. THE HISTORY RECORDS INDICATE THIS PRODUCT WAS FINAL INSPECTION TESTED AT LAKE REGION MEDICAL AND WAS DETERMINED TO BE ACCEPTABLE. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE (DFU), INTENDED USE: · THE SAFARI2¿ GUIDEWIRE IS INTENDED TO FACILITATE THE INTRODUCTION AND PLACEMENT OF INTERVENTIONAL DEVICES WITHIN THE CHAMBERS OF THE HEART, INCLUDING THOSE USED IN TRANSCATHETER AORTIC VALVE PROCEDURES. .AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE (DFU) WARNINGS SECTION: ·MONITOR WIRE POSITION THROUGHOUT THE PROCEDURE FOR PROPER PLACEMENT OF THE CURVE AND DISTAL TIP. ·WHEN ADVANCING OR REMOVING THE GUIDEWIRE, ALWAYS USE FLUOROSCOPIC GUIDANCE WITH RADIOGRAPHIC EQUIPMENT THAT PROVIDES HIGH-RESOLUTION IMAGES. NEVER POSITION THE GUIDEWIRE BLINDLY, AS THIS MAY RESULT IN MISPLACEMENT, DISSECTION OR PERFORATION. ·NEVER ADVANCE THE GUIDEWIRE AGAINST RESISTANCE WITHOUT FIRST DETERMINING THE REASON FOR RESISTANCE UNDER FLUOROSCOPY. EXCESSIVE FORCE AGAINST RESISTANCE MAY RESULT IN DAMAGE TO THE CATHETER OR VESSEL/ORGAN. CARE SHOULD BE TAKEN WHEN ADVANCING THE GUIDEWIRE AFTER DEVICE DEPLOYMENT. ·THE WIRE SHOULD ONLY BE INTRODUCED INTO OR WITHDRAWN FROM THE VENTRICLE THROUGH A CATHETER ALREADY POSITIONED IN THE VENTRICLE. ·THE CURVE OF THE SAFARI2 GUIDEWIRE SHOULD BE CONSTRAINED WITH IN A CATHETER DURING INSERTION INTO OR WITHDRAWAL FROM THE BODY OR TREATMENT SITE. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE: 1.ENSURE A CATHETER IS APPROPRIATELY PLACED IN THE VENTRICLE OR OTHER INTENDED TREATMENT AREA. 2.CAREFULLY REMOVE THE SAFARI2 GUIDEWIRE, FROM THE HOOP BY GRASPING THE PROXIMAL END OF THE J-STRAIGHTENER SEPARATING THE STRAIGHTENER FROM THE HOOP. 3.AFTER INSPECTION OF THE SAFARI2 GUIDEWIRE, ADVANCE THE J-STRAIGHTENER OVER THE DISTAL SECTION OF THE SAFARI2 WIRE TO STRAIGHTEN THE CURVE. 4.INSERT THE SAFARI2 GUIDEWIRE INTO THE HUB OF THE CATHETER WITH THE AID OF THE J- STRAIGHTENER. 5.CAREFULLY ADVANCE THE DISTAL TIP OF THE SAFARI2 GUIDEWIRE INTO THE LUMEN OF THE CATHETER. 6.REMOVE THE SAFARI2 GUIDEWIRE J- STRAIGHTENER BY WITHDRAWING IT OVER THE LENGTH OF THE SAFARI2 GUIDEWIRE. 7.ADVANCED THE SAFARI2 GUIDEWIRE THROUGH THE CATHETER UNDER FLUOROSCOPIC GUIDANCE. 8.CONFIRM SAFARI2 GUIDEWIRE CURVE POSITION TO ASSURE SAFARI2 GUIDEWIRE PLACEMENT AND STABILITY. 9.MAINTAIN WIRE POSITION WHILE TRACKING OR ADVANCING A CATHETER OR DEVICE OVER THE WIRE. 10.TO REMOVE THE GUIDEWIRE FROM THE TREATMENT AREA: A. A CATHETER MUST BE ADVANCED INTO THE TREATMENT AREA OVER THE SAFARI2 GUIDEWIRE PRIOR TO WITHDRAWING THE WIRE. B. THE SAFARI2 GUIDEWIRE IS PULLED BACK COMPLETELY INTO THE CATHETER. C. THE SAFARI2 GUIDEWIRE MAY THEN BE WITHDRAWN FROM THE CATHETER. AT THIS TIME, IT IS NOT POSSIBLE TO ASSIGN A DEFINITIVE ROOT CAUSE FOR THE EVENT AS REPORTED. BASED ON THE INFORMATION PROVIDED, IT APPEARS THAT CLINICAL AND/OR PROCEDURAL FACTORS HAVE CONTRIBUTED TO THE EVENT. SHOULD ADDITIONAL INFORMATION BE RECEIVED, A FOLLOW UP MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS MEDWATCH REPORT IS AN UPDATE TO THE PREVIOUS REPORT TO INCLUDE THE RESULTS AND CONCLUSIONS (H11), TO UPDATE ADVERSE EVENT PROBLEM CODES H6 (A), (B), (C), AND (D), AND THE RESPONSE TO FDA REQUEST (SEE FILE ATTACHMENTS). AN EVALUATION OF THE AVAILABLE INFORMATION AND IMAGING PROVIDED BY THE END USER WAS UNDERTAKEN BY AN INTERVENTIONAL CARDIOLOGIST (MB BCH BAO MD PHD MRCPI FESC) ON BEHALF OF THE MANUFACTURER (INTEGER/LAKE REGION MEDICAL). THE DR CONCLUDED THAT THE SAFARI2 GUIDEWIRE IS A SAFE AND EFFECTIVE DEVICE THAT IS PROVEN TO REDUCE THE RISK OF VENTRICULAR INJURY COMPARED TO NON-PRE-SHAPED GUIDEWIRES. DUE CARE AND DILIGENCE ARE REQUIRED DURING THE USE OF THE DEVICE TO ENSURE THAT THE NATURAL CURVATURE OF THE WIRE IS MAINTAINED. EXCESSIVE FORCE WILL RESULT IN DEFORMATION OF THE GUIDEWIRE AND INTERACTION WITH THE MYOCARDIUM AND RISKS INJURY. THE SAFARI2 GUIDEWIRE INSTRUCTIONS FOR USE WAS REVIEWED. APPROPRIATE INSTRUCTION, PRECAUTION AND WARNINGS ARE PROVIDED TO THE USER. FROM REVIEW OF THE VIDEO PROVIDED FROM THE USER FACILITY DUE CARE WAS NOT EXERCISED BY THE USER IN THIS CASE. THE CONCLUSION, BOTH THE END USER FACILITY AND IN-HOUSE PROCTOR IN THIS CASE, AND THE INTERVENTIONAL CARDIOLOGIST (MB BCH BAO MD PHD MRCPI FESC), WHO VIEWED THE INFORMATION AND IMAGING ON BEHALF OF INTEGER/LAKE REGION MEDICAL, CONCLUDED THAT BASED ON THE AVAILABLE DATA, DUE CARE WAS NOT EXERCISED BY THE USER IN THIS CASE BY NOT MAINTAINING FULL VISUALISATION OF THE SAFARI2 GUIDEWIRE'S DISTAL END. Manufacturer Narrative: THIS MEDWATCH REPORT IS AN UPDATE TO THE PREVIOUS REPORT TO INCLUDE THE ADDITIONAL INFORMATION RECEIVED FROM THE DISTRIBUTOR IN SECTION B(5). MANUFACTURER'S INVESTIGATION AND THE RESPONSE TO FDA REQUEST IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSION WILL BE PROVIDED IN A SUPPLEMENTAL FOLLOW UP REPORT AND THE RESPONSE TO FDA REQUEST WILL BE PROVIDED VIA EMAIL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22136579&pc=LWS | 2649622-2025-15004 | 2025-05-01 04:00 | Death | MPRI | 2025-06-03 04:00 | LWS | TRANSVENE CS/SVC LEAD | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Tachycardia | P980050 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AFTER SUFFERING A VENTRICULAR TACHYCARDIA (VT) ARREST ON THE DAY OF THE RIGHT VENTRICULAR (RV) LEAD IMPLANT PROCEDURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22136582&pc=MGB | 2024168-2025-05676 | 2025-05-16 04:00 | Death | ABBOTT VASCULAR | 2025-06-03 04:00 | MGB | UNK PERCLOSE | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P960043 |  | 1 | Event Description: A META-ANALYSIS OF COMPARATIVE STUDIES WAS PERFORMED BETWEEN THE DUAL PROGLIDE VERSUS PROGLIDE AND ANGIO-SEAL FOR FEMORAL ACCESS HEMOSTASIAS IN TRANSCATHETER AORTIC VALVE REPLACEMENT. IN CONCLUSION, OUR ANALYSIS SUGGESTS THAT AS-PP, COMPARED TO DUAL PP, COULD HAVE SIGNIFICANT CLINICAL BENEFITS IN TF-TAVR WITHOUT RAISING SAFETY CONCERNS. THESE FINDINGS CAN AID CLINICIANS IN DECISION-MAKING AND SERVE AS A SOLID FOUNDATION FOR FUTURE LONGITUDINAL STUDIES TO FURTHER VALIDATE THE POTENTIAL UTILITY OF THE HYBRID AS-PP APPROACH IN PATIENTS UNDERGOING TF-TAVR. THE COMPLICATIONS LINKED TO THE PERCLOSE PROGLIDE DEVICES INCLUDED DEATH, OCCLUSION, BLEEDING, UNPLANNED ENDOVASCULAR INTERVENTION, SURGICAL INTERVENTION, ACCESS SITE VASCULAR COMPLICATIONS, AND ADDITIONAL DEVICE. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED. A REVIEW OF THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) REVIEW WAS NOT PERFORMED BECAUSE A SPECIFIC FAILURE MODE FOR THIS COMPLAINT WAS NOT PROVIDED. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY WAS NOT PERFORMED AS A SIMILARITY COULD NOT BE DETERMINED. BASED ON THE CASE INFORMATION, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS OF DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. A DEFINITIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE PATIENT EFFECT OF DEATH IS LISTED IN THE PROSTYLE INSTRUCTIONS FOR USE (IFU), AS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH USE OF THE SUTURE MEDIATED CLOSURE DEVICES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTION REPORTED IN THE ARTICLE ARE CAPTURED UNDER SEPARATE MEDWATCH REPORT. B2: DATE OF DEATH IS ESTIMATED B3: DATE OF EVENT IS ESTIMATED D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. ATTACHMENT ARTICLE TITLE: "DUAL PROGLIDE VERSUS PROGLIDE AND ANGIO-SEAL FOR FEMORAL ACCESS HEMOSTASIAS IN TRANSCATHETER AORTIC VALVE REPLACEMENT: A META-ANALYSIS OF UP-TO-DATE EVIDENCE".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22136605&pc=NPW | 2015691-2025-04387 |  | Death | EDWARDS LIFESCIENCES | 2025-06-03 04:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Patient Device Interaction Problem | Thrombosis/Thrombus; Insufficient Information | P230013 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS LIFESCIENCES AFFILIATE IN UNITE KINGDOM, EDWARDS RECEIVED NOTIFICATION OF A 56MM EVOQUE PROCEDURE WHERE APPROXIMATELY TWO MONTHS POST-PROCEDURE, VALVE PRESENTED WITH THROMBUS/LEAFLET THICKENED. THE ANTITHROMBOTIC THERAPY WAS CHANGED TO VITAMIN K ANTAGONIST AFTER THE PROCEDURE. THERE WAS CLINICAL EVIDENCE OF EVOQUE LEAFLET THROMBOSIS/ HALT AND WAS MEDICALLY MANAGED WITH IV HEPARIN. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: ADDITIONAL INFORMATION: THE MEDICATION WAS CHANGED TO WARFARIN AND HELD BY GP/PRIMARY CARE WITHOUT NOTIFICATION. PATIENT WAS NOTED TO BE SYMPTOMATIC, AND WAS ADMITTED WITH HYPOXIA OF UNKNOWN CAUSE TO LOCAL HOSPITAL, TREATED FOR CHEST INFECTION THEN REFERRED TO OTHER HOSPITAL 2 WEEKS LATER (SUBTHERAPEUTIC INR AT LOCAL HOSPITAL) AND DIAGNOSED WITH HYPOXIA SECONDARY TO RIGHT TO LEFT SHUNT THROUGH PFO DUE TO TRICUSPID STENOSIS SECONDARY TO VALVE THROMBOSIS. IT WAS PRESUMED PULMONARY EMBOLISM AS CAUSE OF DEATH DESPITE TREATMENT WITH HEPARIN INFUSION AT THE HOSPITAL WHERE THE PATIENT WAS TRANSFERRED. AS PER MEDICAL OPINION, THE PERCEIVED ROOT CAUSE OF THE EVENT WAS THAT WARFARIN WAS STOPPED IN PRIMARY CARE. THE COMPLAINT FOR "THROMBUS FORMATION (DEVICE) - POST PROCEDURE" WAS CONFIRMED WITH EMPIRICAL EVIDENCE VIA INFORMATION RECEIVED BY EDWARDS. BASED ON THE DEVICE HISTORY RECORD REVIEW, THERE WERE NO NON-CONFORMANCE'S RELATED TO THE ISSUE OBSERVED AND THE DEVICE PASSED ALL MANUFACTURING SPECIFICATIONS. POSSIBLE CONTRIBUTING FACTORS TO THE THROMBUS FORMATION CAN INCLUDE PATIENT FACTORS (ANTICOAGULANT REGIMENT, UNDERLYING COAGULOPATHY), PROCEDURAL FACTORS, OR A COMBINATION OF THESE FACTORS. NEVERTHELESS, A DEFINITIVE ROOT CAUSE CANNOT BE ESTABLISHED. SINCE THERE ARE NO CONFIRMED PRODUCT OR LABELING/IFU/TRAINING MATERIAL NON-CONFORMANCE'S AFFECTING DISTRIBUTED PRODUCT AND NO OTHER TRIGGERS HAVE BEEN MET, NO PREVENTATIVE OR CORRECTIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22137010&pc=FRN | 9615050-2025-00274 | 2025-05-03 04:00 | Death | ICU MEDICAL COSTA RICA LTD. | 2025-06-03 04:00 | FRN | PLUM 360¿ INFUSER | Adverse Event Without Identified Device or Use Problem; Unexpected Shutdown | Cardiac Arrest | K161469 |  | 1 | Event Description: THE COMPLAINT/EVENT INVOLVED A PLUM 360¿ INFUSER THAT SHUT DOWN DURING PATIENT CARE AND THE PATIENT SUBSEQUENTLY EXPIRED AFTER A SECOND CODE WAS CALLED. THE INCIDENT OCCURRED AT APPROXIMATELY 1230 AND THE UPDATED TIMELINE WAS AROUND 1333 WITH LAST ALARM. THE DRUG LIBRARY IN USE AT THE TIME WAS WITH VERSION 6.32. THE CUSTOMER PULLED THE EVENT LOGS AND IT APPEARED AS THOUGH THE PUMP HAD ALARMED LOW BATTERY. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. Manufacturer Narrative: ALTHOUGH THE DEVICE WAS REQUESTED MULTIPLE TIMES TO BE RETURNED FOR EVALUATION, IT HAS NOT BEEN RECEIVED AT THIS TIME. Manufacturer Narrative: ENGINEERING CONFIRMS THAT THE DEVICE POWERED OFF DURING INFUSION BY 10:10 APPROXIMATELY. 12:30 BECAUSE OF A DEPLETED BATTERY. THE PROBABLE CAUSE OF THE DEPLETED BATTERY IS THE CUSTOMER NOT LETTING THE BATTERY GET FULLY CHARGED BY PLUGGING INTO THE AC MAIN AND FREQUENTLY SWITCHING THE POWER TO THE BATTERY WHEN AT LEVEL 0, LEADING TO MULTIPLE ABRUPT SHUTDOWNS. ALSO, AFTER THE BATTERY WAS REPLACED, THE PUMP WAS NOT ALLOWED TO GO THROUGH A FULL 12-HOUR CHARGE CYCLE BEFORE ANY CLINICAL USE. BUT SINCE THE BATTERY DRAINED FROM LEVEL 4 TO LEVEL 0 IN LESS THAN 4 HOURS, IT INDICATES A BAD BATTERY CONDITION. ENGINEERING RECOMMENDS THAT THE SERVICE CENTER PERFORM PREVENTIVE MAINTENANCE TESTS. CONCLUSION: IN CONCLUSION, THE DEVICE WAS NOT RETURNED TO THE SERVICE HUB FOR ANALYSIS AND EVALUATION, AND NO VISUAL INSPECTION AND FUNCTIONAL TESTING COULD BE PERFORMED TO DETERMINE IF THE DEVICE HAD A ROLE IN THE ADVERSE EVENT. D9 - DEVICE AVAILABLE FOR EVALUATION: NO. CORRECTION: B1: ADVERSE EVENT, MEDICAL DEVICE PROBLEM CODE: A24.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22137130&pc=NPT | 2025587-2025-04044 | 2025-04-23 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-03 04:00 | NPT | EVOLUT FX VALVE | Malposition of Device; Device Dislodged or Dislocated; Patient Device Interaction Problem | Arrhythmia; Bradycardia; Non specific EKG/ECG Changes; Heart Block; Aspiration Pneumonitis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF A TRANSCATHETER VALVE, THE VALVE WAS IMPLANTED AT A SLIGHTLY DEEP IMPLANT DEPTH. ONE DAY FOLLOWING THE IMPLANT PROCEDURE, LEFT BUNDLE BRANCH BLOCK (LBBB) WAS OBSERVED. ADDITIONALLY, 6 DAYS FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT DEVELOPED BRADYCARDIA THAT PROGRESSED INTO COMPLETE HEART BLOCK (CHB) AND TORSADE DE POINTES. SUBSEQUENTLY, A PERMANENT PACEMAKER (MICRA) WAS IMPLANTED 9 DAYS FOLLOWING THE IMPLANT OF THE VALVE; HOWEVER, THE PATIENT DIED 14 DAYS FOLLOWING THE IMPLANT OF THE VALVE DUE TO ASPIRATION PNEUMONIA.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF A TRANSCATHETER VALVE, THE VALVE WAS IMPLANTED AT A SLIGHTLY DEEP IMPLANT DEPTH. ONE DAY FOLLOWING THE IMPLANT PROCEDURE, LEFT BUNDLE BRANCH BLOCK (LBBB) WAS OBSERVED. ADDITIONALLY, 6 DAYS FOLLOWING THE IMPLANT OF THE VALVE, THE PATIENT DEVELOPED BRADYCARDIA THAT PROGRESSED INTO COMPLETE HEART BLOCK (CHB) AND TORSADE DE POINTES. SUBSEQUENTLY, A PERMANENT PACEMAKER WAS IMPLANTED 9 DAYS FOLLOWING THE IMPLANT OF THE VALVE; HOWEVER, THE PATIENT DIED 14 DAYS FOLLOWING THE IMPLANT OF THE VALVE DUE TO ASPIRATION PNEUMONIA. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE VALVE DEMONSTRATED A TENDENCY TO DISLODGE WHICH CONTRIBUTED TO THE SLIGHTLY DEEP IMPLANT DEPTH. PER THE PHYSICIAN, THE DEEP IMPLANT DEPTH OF THE 34 MM VALVE WAS UNINTENTIONAL.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED USING A 20 MM NON-MEDTRONIC BALLOON. THE DEPLOYMENT STARTING POINT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER, AND THE IMPLANT DEPTH PRIOR TO VALVE DISLODGEMENT WAS 4 MM ON THE NON-CORONARY CUSP (NCC) AND 5 MM ON THE LEFT CORONARY CUSP (LCC). THE VALVE SUBSEQUENTLY DISLODGED VENTRICULAR. AFTER VALVE DISLODGEMENT, THE IMPLANT DEPTH ON THE NCC WAS 6 MM AND THE IMPLANT DEPTH ON THE LCC WAS 8 MM. ADDITIONALLY, A NON-MEDTRONIC GUIDEWIRE WAS USED DURING THE PROCEDURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME EVOLUT FX DCS; PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012362236); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE UPDATED DATA: B5. ADDED SECOND PARAGRAPH. H6. H11. SYSTEM REPORTED DCS. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-34 (LOT: 0012362236); PRODUCT TYPE: 0195-HEART VALVES; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22137354&pc=OJX | 2649622-2025-15006 | 2025-04-11 04:00 | Death | MPRI | 2025-06-03 04:00 | OJX | ATTAIN STABILITY QUAD MRI SURESCAN | Low impedance | Cardiac Arrest; Insufficient Information | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING ON THE CURRENT ELECTROGRAMS (EGM) RESULTING IN INAPPROPRIATE PACING AND THE DEVICE FALSELY CLASSIFYING AT/AF EPISODES AS TERMINATED. IT WAS ALSO NOTED THAT THE ATRIAL LEAD AND LEFT VENTRICULAR (LV) LEAD TRIGGERED WARNINGS FOR LOW PACING IMPEDANCE MEASUREMENTS. THE PATIENT PASSED AWAY THE DAY AFTER THE UNDERSENSING WAS OBSERVED.Event Description: IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS CARDIORESPIRATORY ARREST. Manufacturer Narrative: CONTINUATION OF D10: 5076-52 LEAD IMPLANTED: (B)(6) 2021. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22137398&pc=DTB | 2649622-2025-15007 | 2025-04-11 04:00 | Death | MPRI | 2025-06-03 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Pacing Problem; Under-Sensing; Low impedance | Cardiac Arrest; Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING ON THE CURRENT ELECTROGRAMS (EGM) RESULTING IN INAPPROPRIATE PACING AND THE DEVICE FALSELY CLASSIFYING AT/AF EPISODES AS TERMINATED. IT WAS ALSO NOTED THAT THE ATRIAL LEAD AND LEFT VENTRICULAR(LV) LEAD TRIGGERED WARNINGS FOR LOW PACING IMPEDANCE MEASUREMENTS. THE PATIENT PASSED AWAY THE DAY AFTER THE UNDERSENSING WAS OBSERVED.Event Description: IT WAS FURTHER REPORTED THAT THE CAUSE OF DEATH WAS CARDIORESPIRATORY ARREST. Manufacturer Narrative: CONTINUATION OF D10: 5076-58 LEAD IMPLANTED: (B)(6) 2021; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22137492&pc=NGV | 2135147-2025-02937 | 2025-05-09 04:00 | Death | ABBOTT MEDICAL | 2025-06-03 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 22MM AMPLATZER AMULET OCCLUDER WAS SUCCESSFULLY IMPLANTED IN A PATIENT WITH PRE-EXISTING RENAL FAILURE. THERE WAS NO DIFFICULTY EXPERIENCED IMPLANTING THE 22MM AMPLATZER AMULET OCCLUDER THE PATIENT WAS DISCHARGED FOLLOWING THE IMPLANT. 3 HOURS POST-DISCHARGE, IT WAS NOTED THAT THE PATIENT WAS FOUND DECEASED. THERE WERE NO SYMPTOMS NOTED BY FAMILY MEMBERS. THE CAUSE OF DEATH IS UNKNOWN. IT'S NOTED THAT THE PATIENT WAS ON DIALYSIS PRIOR TO PROCEDURE. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2 - DATE OF DEATH: DEATH DATE IS ESTIMATED AS IT IS NOT CONFIRMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22137795&pc=LWP | 2124215-2025-36404 | 2025-05-08 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-03 04:00 | LWP | ACUITY? X4 SPIRAL S | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P010012 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT CODED AND DIED DURING THE IMPLANT PROCEDURE. THE RIGHT VENTRICULAR (RV) LEAD WAS IMPLANTED WITHOUT INCIDENT. THEN THIS LEFT VENTRICULAR (LV) LEAD WAS IMPLANTED BUT DURING THRESHOLD TESTING, THE PATIENT WENT INTO PULSELESS ELECTRICAL ACTIVITY (PEA). A CODE WAS CALLED, BUT THE PATIENT DIED. THE LEADS WERE NOT EXPLANTED AND WILL NOT BE RETURNED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22138001&pc=DSQ | 2916596-2025-03340 | 2025-05-09 04:00 | Death | THORATEC CORPORATION | 2025-06-03 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Hypervolemia; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RIGHT HEART FAILURE (RHF). THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. IT WAS NOTED THAT RHF EXISTED PRE-LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT AND A DEVICE RELATED ISSUE DID NOT CONTRIBUTE TO RHF. THE PATIENT HAD EXHIBITED VOLUME OVERLOAD. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE (HM) 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED RIGHT HEART FAILURE (RHF) AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22138355&pc=DYB | 2017865-2025-75363 | 2025-05-14 04:00 | Death | ABBOTT | 2025-06-03 04:00 | DYB | AVEIR¿ INTRODUCER KIT ¿ 50 CM | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K211989 |  | 1 | Event Description: A LEADLESS PACEMAKER (LP) SYSTEM IMPLANT PROCEDURE WAS PERFORMED WITH NO COMPLICATIONS OR ALLEGED PRODUCT MALFUNCTION. THE PATIENT WAS DISCARDED AND WHILE AT HOME, WAS REPORTED TO HAVE BEEN COUGHING AND BEGAN BLEEDING FROM THE FEMORAL ACCESS. THE PATIENT WAS BROUGHT TO EMERGENCY SERVICES BY AMBULANCE. ATTEMPTS TO STABILIZE PATIENT WERE UNSUCCESSFUL AND THE PATIENT PASSED AWAY. THE INITIAL OPINION OF THE PHYSICIAN WAS THAT THE BLEEDING MAY HAVE BEEN CAUSED BY ISSUE WITH ACCESS SITE OR COMPRESSION FOLLOWING PROCEDURE, HOWEVER THE CORONER'S REPORT HAS NOT BEEN SHARED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22121445&pc=BZD | 2518422-2025-045966 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-06-02 04:00 | BZD | REMSTAR AUTO SERIES AFLEX ASSY | Degraded | Pulmonary Dysfunction; Unspecified Heart Problem; Unspecified Respiratory Problem; Insufficient Information | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, HEART DISEASE, OTHER (UNSPECIFIED EVENT) AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22121493&pc=BZD | 2518422-2025-045808 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-06-02 04:00 | BZD | REMSTAR PRO C-FLEX+ | Degraded | Cancer; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, LIVER CANCER (FATAL) AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22121594&pc=BZD | 2518422-2025-045993 | 2023-08-15 04:00 | Death | RESPIRONICS, INC. | 2025-06-02 04:00 | BZD | BIPAP AUTO SERIES ASSY | Degraded | Cancer | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND TONSIL CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22121704&pc=BZD | 2518422-2025-046069 | 2024-01-22 05:00 | Death | RESPIRONICS, INC. | 2025-06-02 04:00 | BZD | REMSTAR AUTO SERIES AFLEX ASSY | Degraded | Cancer | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY CANCER RESULTING IN DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22121734&pc=BZD | 2518422-2025-046033 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-06-02 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Pulmonary Dysfunction; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22121862&pc=BZD | 2518422-2025-046124 | 2023-12-01 05:00 | Death | RESPIRONICS, INC. | 2025-06-02 04:00 | BZD | REMSTAR PRO C-FLEX+ | Degraded | Pulmonary Embolism; Cancer; Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, OTHER (UNSPECIFIED EVENT), PROSTATE CANCER AND PULMONARY EMBOLISM. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22121891&pc=BZD | 2518422-2025-046087 | 2023-10-12 04:00 | Death | RESPIRONICS, INC. | 2025-06-02 04:00 | BZD | DREAMSTATION CPAP | Degraded | Pulmonary Dysfunction; Unspecified Respiratory Problem; Unspecified Kidney or Urinary Problem | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE / TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, CUSTOMER COMPLAINT WAS NOT CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22121906&pc=DSP | 2249723-2025-0002415 | 2025-04-29 04:00 | Death | DATASCOPE CORP. | 2025-06-02 04:00 | DSP | CS300 INTRA-AORTIC BALLOON PUMP, PORTUGUESE, 220V | No Device Output | Insufficient Information | K063525 |  | 1 | Event Description: IT WAS REPORTED THAT CS300 INTRA-AORTIC BALLOON PUMP (IABP) EXPERIENCED FAILURE IN FIBER OPTICAL SENSOR MODULE. BALLOON BURSTED INSIDE THE PATIENT. THERE ARE NO REPORTS OF COMPLICATIONS DIRECTLY RELATED TO THE BALLOON, BUT THE PATIENT DIED. Manufacturer Narrative: DUE TO CHARACTER LIMITATION IN E1, FULL INITIAL REPORTER NAME (B)(6). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22122146&pc=OZD | 1220648-2025-28880 | 2025-05-10 04:00 | Death | ABIOMED, INC. | 2025-06-02 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WAS ON DAY TWO OF IMPELLA CP SUPPORT AND DEVELOPED HEMATOMA IN THE ACCESS SITE. ONE UNIT OF PACKED RED BLOOD CELLS WAS ADMINISTERED. IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED WHILE ON SUPPORT. THERE WAS NO OTHER INFORMATION PROVIDED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22122198&pc=OZD | 1220648-2025-28882 | 2025-05-04 04:00 | Death | ABIOMED, INC. | 2025-06-02 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AN EMERGENCY CHEST EXPLORATION WAS REQUIRED DURING IMPELLA 5.5 SUPPORT TO CONTROL A BLEED. THE PUMP WAS REPOSITIONED FOLLOWING THE INTERVENTION AND SUPPORT CONTINUED WITH THE IMPLANTED DEVICE. DURING SUPPORT THE PATIENT WELL SUPPORTED ON IMPELLA 5.5 ON P9 WITH 5.0 LPM NOTED. PATIENT HAD 6 BEATS OF VENTRICULAR TACHYCARDIA OVERNIGHT. THE FAMILY WITHDRAW CARE AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN. SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22122576&pc=OZD | 1220648-2025-28891 | 2025-05-08 04:00 | Death | ABIOMED, INC. | 2025-06-02 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Shock | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT UPON REMOVAL OF THE PEEL-AWAY SHEATH, THE PATIENT BEGAN TO BLEED FROM THE ACCESS SITE OF THE IMPELLA CP. MANUAL COMPRESSION AND FEMOSTOP WERE UTILIZED TO MANAGE THE CONDITION. FOUR UNITS OF BLOOD WERE ALSO GIVEN. HOWEVER, THE BLEED PERSISTED AND THE PATIENT WENT INTO HEMORRHAGIC SHOCK. THE PATIENT PASSED AWAY THE FOLLOWING DAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22122648&pc=QFG | 3004464228-2025-23658 | 2025-05-06 04:00 | Death | INSULET CORPORATION | 2025-06-02 04:00 | QFG | OMNIPOD 5 POD | Adverse Event Without Identified Device or Use Problem | Hypoglycemia; Myocardial Infarction | K203768 |  | 1 | Event Description: IT WAS REPORTED ON (B)(6) 2025 BY THE PATIENT'S SISTER THAT THE PATIENT HAD PASSED AWAY. THE DATE OF THE PATIENT'S PASSING WAS NOT PROVIDED. THE PATIENT'S SISTER REPORTED; WHILE WEARING A POD, THE PATIENT EXPERIENCED A HYPOGLYCEMIC EPISODE. BLOOD GLUCOSE LEVELS DECLINED TO 43 MG/DL. LENGTH OF TIME THE POD WAS WORN WAS NOT PROVIDED. THE PATIENT'S SISTER REPORTED WHILE SHE WAS ASSISTING THE PATIENT WITH TREATING HER HYPOGLYCEMIA (SPECIFIC TREATMENT WAS NOT PROVIDED), THE PATIENT EXPERIENCED A HEART ATTACK. THE PATIENT WAS ALSO REPORTED TO HAVE AN IMMUNE DISORDER, CVID (COMMON VARIABLE IMMUNE DISORDER). THE PATIENT WAS TRANSPORTED BY AMBULANCE TO THE HOSPITAL, THE BLOOD GLUCOSE LEVEL HAD INCREASED TO 90 MG/DL, HOWEVER THE PATIENT REPORTEDLY DID NOT HAVE A PULSE WHILE IN TRANSPORT. THE PATIENT WAS PRONOUNCED DEAD UPON ARRIVAL TO THE HOSPITAL. NO ALLEGATIONS OF A DEVICE MALFUNCTION WERE REPORTED. Manufacturer Narrative: THE PHYSICAL DEVICE IS NOT EXPECTED TO BE RETURNED FOR INVESTIGATION. USER-INITIATED LOG FILES FROM THE CONTROLLER SERIAL NUMBER IN THE COMPLAINT WERE NOT AVAILABLE IN THE CLOUD SYSTEM. CLOUD DATA LINKED TO THE PATIENT ID WAS REVIEWED FOR THE PERIOD OF 01APR2025 TO 06MAY2025. THE LAST DATA ENTRY OF THE PATIENT HISTORY BUFFER (PHB) WAS FOUND TO BE AT 07:00:48 ON (B)(6) 2025. THE PHB DATA SHOWED THAT THE AUTOMATED ALGORITHM WAS ABLE TO APPROPRIATELY ADAPT THE AUTOMATED INSULIN AMOUNTS BASED ON THE ESTIMATED GLUCOSE VALUES AND USER INPUTS. THE PHB DATA CONFIRMED THAT THE USER'S BLOOD GLUCOSE VALUES REACHED URGENT LOW LEVELS AT 06:50:48 ON (B)(6) 2025 AND STAYED THERE UNTIL 07:00:58. THE DATA SHOWED THAT THE AUTOMATED ALGORITHM STOPPED THE DELIVERY OF AUTOMATED INSULIN AT 04:15:48 WHEN THE POD RECEIVED A BLOOD GLUCOSE LEVEL OF 119 MG/DL. THIS WAS DUE TO A STEADY DECREASE IN BLOOD GLUCOSE LEVELS. THE SYSTEM WAS CORRECTLY REDUCING AND STOPPED AUTOMATED INSULIN DELIVERY APPROPRIATELY AS THE ESTIMATED GLUCOSE VALUES DECREASED FROM A HIGH OF 338 MG/DL AT 18:20:49 ON (B)(6) 2025 DOWN TO THE USER'S TARGET OF 110 MG/DL. THE AUTOMATED ALGORITHM WAS DETERMINED TO FUNCTION AS INTENDED. CLOUD - LOCKED DOWN/SMARTPHONE LOCKDOWN CLOUD - OMNIPOD 5 SOFTWARE APP VERSION 3.1.1 CLOUD - SMARTPHONE OPERATING SYSTEM N5004L-AM-Q-MV01602-06-01.06 CLOUD - SMARTPHONE HARDWARE N5004L CLOUD - CGM SENSOR TYPE G6 *PLEASE NOTE, THAT SECTION D IS CAPTURING THE DEVICE IDENTIFIERS AS REPORTED BY THE COMPLAINANT. THIS MAY NOT ALIGN TO THE DEVICE CONFIGURATION REPORTED IN H11, AS THIS DATA IS PULLED FROM OUR CLOUD BASED ON THE REPORTED DATE OF EVENT. Manufacturer Narrative: PLEASE DISREGARD PREVIOUSLY SUBMITTED (B)(4)/ REFERENCE REPORT NUMBER 3004464228-2025-23658-00. THIS WAS NOT A REPORTABLE MEDICAL EVENT RELATED TO AN INSULET DEVICE, AS THERE WAS NO ALLEGATION OR INDICATION THAT OUR DEVICE CAUSED OR CONTRIBUTED TO THE MEDICAL EVENT. THE EMDR REPORT WAS SUBMITTED IN ERROR. EMS (EMERGENCY MEDICAL SERVICES) WERE CONTACTED INITIALLY DUE TO THE PATIENT HAVING HYPOGLYCEMIA WITH A BLOOD GLUCOSE OF 43 MG/DL; HOWEVER, WHEN EMS ARRIVED THE PATIENT'S BLOOD GLUCOSE WAS 90 MG/DL, AND NO TREATMENT WAS PROVIDED. THE PATIENT LATER EXPERIENCED A HEART ATTACK AND WAS TRANSPORTED TO THE HOSPITAL. WHILE IN ROUTE TO THE HOSPITAL, THE PATIENT PASSED AWAY FROM THE HEART ATTACK. REVIEW OF CLOUD DATA FOUND OUR SYSTEM WAS CORRECTLY REDUCING AND STOPPED AUTOMATED INSULIN DELIVERY APPROPRIATELY AND THE AUTOMATED ALGORITHM WAS DETERMINED TO FUNCTION AS INTENDED. OUR PRODUCT DID NOT CAUSE OR CONTRIBUTE TO THE MEDICAL EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22122653&pc=OZD | 1220648-2025-28893 | 2025-05-08 04:00 | Death | ABIOMED, INC. | 2025-06-02 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Tachycardia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT PRIOR TO THE IMPELLA 5.5 BEING INSERTED, THE PATIENT WENT INTO CONTINUOUS VENTRICULAR TACHYCARDIA (VT). AFTER THE IMPELLA 5.5 WAS PLACED THIS EPISODE CONTINUED. MULTIPLE SHOCKS, AMIODARONE BOLUSES/DRIPS, LIDOCAINE DRIP AND ELECTROLYTES REPLACEMENTS WERE GIVEN. THIS LASTED FOR APPROXIMATELY 30-45 MINUTES. PADDLES WERE USED WHICH WERE INEFFECTIVE. NEW PADS WERE PLACED IN WHICH A NICE SHOCK WAS DELIVERED CONVERTING THE PATIENT TO SINUS RHYTHM. THE PATIENT CONTINUED WITH VT DURING IMPELLA 5.5 SUPPORT. IT WAS FURTHER REPORTED THAT THE PATIENT HAD EXTENSIVE HEMORRHAGIC STROKE WITH MIDLINE SHIFT AND EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA USED AS AN ASSOCIATED FACTOR TO THE EVENT OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)." SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: "TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22122657&pc=NIQ | 2124215-2025-35772 | 2025-05-06 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-02 04:00 | NIQ | PROMUS PREMIER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. A 2.75 X 38 MM PROMUS PREMIER WAS IMPLANTED DURING PERCUTANEOUS TRANSLUMINAL CORONARY ANGIOPLASTY TO THE LEFT ANTERIOR DESCENDING ARTERY. A FEW HOURS AFTER THE PROCEDURE, THE PATIENT WENT INTO CARDIAC ARREST AND DIED. Manufacturer Narrative: GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22122739&pc=QZK | 2017233-2025-06259 | 2025-05-02 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-02 04:00 | QZK | GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS | Insufficient Information | Stroke/CVA | P230023 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, A PATIENT WAS TREATED USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) SYSTEM. A GORE® EXCLUDER® AAA ENDOPROSTHESIS CONTRALATERAL LEG WAS ALSO USED IN THE AORTA. PROXIMAL ACCESS WAS GAINED THROUGH A GORE® DRYSEAL FLEX INTRODUCER SHEATH IN THE RIGHT SUBCLAVIAN ARTERY. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. EARLY ON (B)(6) 2025, WITHIN 24 HOURS OF THE INITIAL PROCEDURE, THE PATIENT HAD SOME WATERSHED INFARCTS. A CT OF THE CHEST SHOWED NOTHING UNUSUAL. ON (B)(6) 2025, THE PATIENT STARTED HAVING SEIZURES AND WAS INTUBATED. ACCORDING TO THE FSA, THERE WAS NO NOTABLE THROMBUS. POST-OPERATIVE IMAGES ARE NOT AVAILABLE.Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, A PATIENT WAS TREATED USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) SYSTEM. A GORE® EXCLUDER® AAA ENDOPROSTHESIS CONTRALATERAL LEG WAS ALSO USED IN THE AORTA.. PROXIMAL ACCESS WAS GAINED THROUGH A GORE® DRYSEAL FLEX INTRODUCER SHEATH IN THE RIGHT SUBCLAVIAN ARTERY. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. EARLY ON (B)(6) 2025, WITHIN 24 HOURS OF THE INITIAL PROCEDURE, THE PATIENT HAD SOME WATERSHED INFARCTS. A CT OF THE CHEST SHOWED NOTHING UNUSUAL. ON (B)(6) 2025, THE PATIENT STARTED HAVING SEIZURES AND WAS INTUBATED. ACCORDING TO THE FSA, THERE WAS NO NOTABLE THROMBUS. ON (B)(6) 2025, THE PHYSICIAN REPORTED THAT THE PATIENT HAD DIED, SAYING THAT THE PATIENT NEVER RECOVERED FROM THE STROKE. Manufacturer Narrative: B2. OUTCOMES UPDATED TO DEATH. B4. EVENT DESCRIPTION UPDATED. G3/G4. DATE MANUFACTURER RECEIVED BY MANUFACTURER IS THE DATE WE LEARNED OF THE DEATH. H1. TYPE OF REPORTABLE EVENT UPDATED TO DEATH. H.6. HEALTH EFFECT - IMPACT CODE - ADDED F02. Manufacturer Narrative: H.6. INVESTIGATION FINDINGS CODE C20: THE DEVICE REMAINS IMPLANTED AND IS NOT AVAILABLE FOR ANALYSIS. H.6. INVESTIGATION FINDINGS CODE C19: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H.6. INVESTIGATION CONCLUSIONS CODE D12: IT SHOULD BE NOTED THAT, PER THE GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) INSTRUCTIONS FOR USE, POSSIBLE ADVERSE EVENTS AND COMPLICATIONS THAT MAY OCCUR WITH THE USE OF THE GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS INCLUDE: STROKE. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22122800&pc=DYB | 3007284313-2025-04005 | 2025-05-02 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-02 04:00 | DYB | GORE® DRYSEAL FLEX INTRODUCER SHEATH | Insufficient Information | Stroke/CVA | K160254 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, A PATIENT WAS TREATED USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) SYSTEM. A GORE® EXCLUDER® AAA ENDOPROSTHESIS CONTRALATERAL LEG WAS ALSO USED IN THE AORTA. PROXIMAL ACCESS WAS GAINED THROUGH A GORE® DRYSEAL FLEX INTRODUCER SHEATH IN THE RIGHT SUBCLAVIAN ARTERY. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. EARLY ON (B)(6) 2025, WITHIN 24 HOURS OF THE INITIAL PROCEDURE, THE PATIENT HAD SOME WATERSHED INFARCTS. A CT OF THE CHEST SHOWED NOTHING UNUSUAL. ON (B)(6) 2025, THE PATIENT STARTED HAVING SEIZURES AND WAS INTUBATED. ACCORDING TO THE FSA, THERE WAS NO NOTABLE THROMBUS. POST-OPERATIVE IMAGES ARE NOT AVAILABLE.Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, A PATIENT WAS TREATED USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) SYSTEM. A GORE® EXCLUDER® AAA ENDOPROSTHESIS CONTRALATERAL LEG WAS ALSO USED IN THE AORTA.. PROXIMAL ACCESS WAS GAINED THROUGH A GORE® DRYSEAL FLEX INTRODUCER SHEATH IN THE RIGHT SUBCLAVIAN ARTERY. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. EARLY ON (B)(6) 2025, WITHIN 24 HOURS OF THE INITIAL PROCEDURE, THE PATIENT HAD SOME WATERSHED INFARCTS. A CT OF THE CHEST SHOWED NOTHING UNUSUAL. ON (B)(6) 2025, THE PATIENT STARTED HAVING SEIZURES AND WAS INTUBATED. ACCORDING TO THE FSA, THERE WAS NO NOTABLE THROMBUS. ON (B)(6) 2025, THE PHYSICIAN REPORTED THAT THE PATIENT HAD DIED, SAYING THAT THE PATIENT NEVER RECOVERED FROM THE STROKE. Manufacturer Narrative: H.6. INVESTIGATION FINDINGS CODE C20: THE DEVICE WAS DISCARDED AT THE TREATING FACILITY AND WAS THEREFORE NOT AVAILABLE FOR ANALYSIS. H.6. TYPE OF INVESTIGATION CODE B22: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE CONDUCTED BECAUSE THE SERIAL/LOT NUMBER REMAINS UNKNOWN. H.6. INVESTIGATION CONCLUSIONS CODE D12: IT SHOULD BE NOTED THAT, PER THE GORE® DRYSEAL FLEX INTRODUCER SHEATH INSTRUCTIONS FOR USE, COMPLICATIONS ASSOCIATED WITH THE USE OF THE GORE® DRYSEAL FLEX INTRODUCER SHEATH THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO: STROKE. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT. Manufacturer Narrative: B2. OUTCOMES UPDATED TO DEATH. B4. EVENT DESCRIPTION UPDATED. G3/G4. DATE MANUFACTURER RECEIVED BY MANUFACTURER IS THE DATE WE LEARNED OF THE DEATH. H1. TYPE OF REPORTABLE EVENT UPDATED TO DEATH. H.6. HEALTH EFFECT - IMPACT CODE - ADDED F02.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22122901&pc=NIQ | 2124215-2025-35782 | 2025-05-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-02 04:00 | NIQ | PROMUS ELITE | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus; Heart Failure/Congestive Heart Failure |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE TARGET LESION WAS LOCATED IN THE LEFT ANTERIOR DESCENDING ARTERY. THE LESION WAS PRE-DILATED WITH A 3.50 X 12 MM BALLOON. A 3.50 X 38 MM PROMUS ELITE WAS THEN IMPLANTED AND STENT THROMBUS FORMED AT THE SITE OF THE LESION. THE PATIENT WENT INTO HEART FAILURE. CARDIOPULMONARY RESUSCITATION WAS PERFORMED UNSUCCESSFULLY AND THE PATIENT DIED ON THE TABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22122992&pc=BZD | 2518422-2025-046108 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-06-02 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Cancer | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING PANCREATIC CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22123323&pc=GEI | 1717344-2025-00700 | 2024-07-16 04:00 | Death | COVIDIEN MFG DC BOULDER | 2025-06-02 04:00 | GEI | UNKNOWN LIGASURE INSTRUMENT | Energy Output Problem | Hemorrhage/Bleeding; Hypovolemic Shock |  |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN OPEN TOTAL ABDOMINAL HYSTERECTOMY AND BILATERAL SALPINGO-OOPHORECTOMY, THE PATIENT GOT INTO HYPOVOLEMIC SHOCK THE FOLLOWING MORNING. THE BLEEDING VESSELS WERE IDENTIFIED WHEN THE PATIENT WAS TAKEN BACK TO THEATRE AS AN EMERGENCY. A MASSIVE TRANSFUSION PROTOCOL (MTP) WAS INITIATED, RESUSCITATED AND GAVE FLUIDS AND THEN TAKEN BY THE AMBULANCE CREW TO TRANSFER TO A DIFFERENT HOSPITAL WHERE THEY OPENED THE PATIENT UP AND PERFORMED HEMOSTASIS BUT THE PATIENT SUBSEQUENTLY DIED OF HYPOVOLEMIC SHOCK. AN INQUEST WAS MADE BY THE CORONER, AND IT WAS IDENTIFIED THAT THE PATIENT EXPERIENCED A DELAYED INTRA-ABDOMINAL BLEED FROM 2 OF THE 9 BLOODS VESSELS THAT WERE SEALED WITH THE DEVICE HAD FAILED. Manufacturer Narrative: D10 CONCOMITANT PRODUCT: VLFT10GEN, VLFT10GEN FT SERIES ENERGY PLATFORMX1 (SERIAL #: (B)(6)). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22123325&pc=BZD | 2518422-2025-046207 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-06-02 04:00 | BZD | REMSTAR AUTO A-FLEX | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22123765&pc=DXF | 2124215-2025-35479 | 2025-05-09 04:00 | Death | BAYLIS MEDICALE CIE INC | 2025-06-02 04:00 | DXF | VERSACROSS ACCESS SOLUTION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Tachycardia; Cardiac Perforation; Pericardial Effusion; Asystole | K150709 |  | 1 | Event Description: IT WAS REPORTED A PATIENT DEATH OCCURRED. DURING A PULMONARY VEIN ISOLATION ABLATION PROCEDURE, A VERSACROSS ACCESS SOLUTION KIT WAS SELECTED FOR USE. THE PHYSICIAN OBTAINED ACCESS IN BOTH RIGHT AND LEFT FEMORAL VEIN. INTRACARDIAC ECHOCARDIOGRAPHY (ICE) IMAGINING WAS INSERTED IN LEFT FEMORAL VEIN. THE GUIDEWIRE WAS ADVANCED TO SUPERIOR VENA CAVA (SVC) AND THEN THE VERSACROSS TRANSSEPTAL SHEATH AS WELL AS THE RF WIRE. THE TRANSSEPTAL PUNCTURE (TSP) AND THE RF WIRE WERE ADVANCED IN THE LEFT ATRIUM. THERE WAS NO DIFFICULTY NOTED WITH THE TSP WORKFLOW. THE PATIENT HAD AN INCREASED HEART RATE OF 215 BPM THAT SUSTAINED FOR A MINUTE. THEN PERICARDIAL EFFUSION (PE) WAS NOTED ON ICE IMAGING AND A PERICARDIOCENTESIS BEGAN. CARDIOPULMONARY RESUSCITATION (CPR) WAS STARTED AND CODE BLUE ALARM WAS CALLED. THE PATIENT WAS ASYSTOLE. THE PHYSICIAN OPENED PATIENT CHEST, AND BLOOD TRANSFUSION BEGAN. CARDIAC MESSAGE WAS PERFORMED. SODIUM BICARB WAS GIVEN THROUGH THE SHEATH. VASOPRESSORS GIVEN BY IC PHYSICIAN. AORTIC CANNULATION WAS ADMINISTERED AND VENOUS PERFUSION BEGAN. THE BYPASS SURGERY WAS INITIATED. A 90 VOLT SHOCK ADMINISTERED TO HEART, HOWEVER LATER THE PATIENT WAS DECLARED DEAD. THE AUTOPSY HAS NOT BEEN PERFORMED YET HOWEVER IT IS EXPECTED THAT DEATH WAS CAUSED BY COMPLICATIONS DUE TO PERFORATION OF CARDIAC TISSUE. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). THE OFFICIAL CAUSE OF DEATH WAS PERICARDIAL EFFUSION. THE PATIENT HAD A SMALL LEFT ATRIUM. THERE ARE NO PRE-EXISTING PATIENT CONDITIONS THAT MAY HAVE CAUSED A COMPLICATION DURING THE PROCEDURE. MEDWATCH REFERENCE MW5170596. IT WAS ADDITIONALLY REPORTED THAT IT WAS PULLED OUT OVER A LITER OF FLUID. THE PATIENT CONTINUED TO BLEED. WHEN THE PATIENT'S CHEST WAS OPENED UP AND IT WAS NOTICED ONE PERFORATION IN THE RIGHT VENTRICLE (RV) THAT THEY BELIEVED MAY HAVE BEEN CAUSED BY THE TAP FOR THE PERICARDIOCENTESIS. THE LEFT APPENDAGE HAD A COUPLE OF PERFORATIONS AS WELL THAT THEY BELIEVE WAS FROM THE DILATOR. THE CALLER STATED THEY ASSUMED THE FARADRIVE CATHETER AND THE VERSACROSS WIRE WERE THE CAUSE OF THE INJURY. FARADRIVE SHEATH AND A BOSTON CORONARY SINUS (CS) WERE IN THE BODY WHEN THE EVENT OCCURRED. THE CALLER STATED THEY BELIEVE THE FARADRIVE DILATOR PERFORATED THE LEFT APPENDAGE WHILE GOING TRANSSEPTAL. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22123868&pc=OZD | 1220648-2025-28886 | 2025-05-08 04:00 | Death | ABIOMED, INC. | 2025-06-02 04:00 | OZD | IMPELLA CP | Infusion or Flow Problem | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP WAS MOVED BY NURSING STAFF AND SUBSEQUENTLY, THE PUMP'S LV WAVEFORM FLIPPED. THE PATIENT'S FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22124031&pc=NPT | 2015691-2025-04351 | 2025-01-22 05:00 | Death | EDWARDS LIFESCIENCES | 2025-06-02 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Difficult to Advance | Rupture | P140031 |  | 1 | Event Description: AS REPORTED THROUGH A REVIEW OF MEDICAL ARTICLE " A HOLY SITUATION: AN UNEXPECTED COMPLICATION DURING A TAVI PROCEDURE", AFTER AN UNKNOWN AMOUNT OF TIME, A PATIENT UNDERWENT INTERVENTION FOR A 26MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION DUE TO SEVERE AORTIC STENOSIS. PRE-DILATION WAS PERFORMED WITH A 22MM BALLOON. AFTER THE VALVE WAS IMPLANTED, MODERATE REGURGITATION (UNKNOWN WHETHER IT WAS PARAVALVULAR OR CENTRAL) WAS NOTED. POST-DILATION WITH THE SAME BALLOON WAS ATTEMPTED, ENCOUNTERING DIFFICULTIES WITH BALLOON PROGRESSION. AT THAT TIME, THE PATIENT PRESENTED WITH HEMODYNAMIC COLLAPSE AND CARDIAC ARREST. AN ECHO SHOWED NO PERICARDIAL EFFUSION, EXCLUDING ANNULAR RUPTURE. WHILE PERFORMING ADVANCED LIFE SUPPORT MANEUVERS, AORTOGRAPHY WAS PERFORMED SHOWING AN AORTIC RUPTURE AT THE TRANSITION BETWEEN THE AORTIC ARCH AND THE DESCENDING AORTA. TAMPONADE WITH A 28MM NON-EDWARDS BALLOON WAS PERFORMED AND A THORACIC ENDOVASCULAR AORTIC REPAIR (TEVAR) 28MM WAS IMPLANTED WITH SUBSEQUENT DILATION USING A BALLOON. WHILE RESOLUTION OF THE ACTIVE RUPTURE WAS ACHIEVED ANGIOGRAPHICALLY AND DESPITE 55 MINUTES OF RESUSCITATION MANEUVERS, CARDIAC RHYTHM COULD NOT BE RESTORED AND THE PATIENT EXPIRED. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR THE COMPLETION OF THE INVESTIGATION. THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATION'S INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. PROVIDED IMAGERY WAS EVALUATED, AND THE FOLLOWING OBSERVATIONS WERE NOTED: A SEVERE CALCIFICATION OF THE AORTIC ARCH. AORTIC RUPTURE AT THE TRANSITION BETWEEN THE ARCH AND THE DESCENDING AORTA. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, DIFFICULTY OR INABILITY IN TRACKING THE DELIVERY SYSTEM TO THE LANDING ZONE, INCLUDING TRAVERSING THE AORTIC ARCH AND CROSSING THE NATIVE ANNULUS/BIOPROSTHESIS HAS NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCE'S. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PATIENT FACTORS, POOR GUIDEWIRE MANAGEMENT, DAMAGED GUIDEWIRE, AND/OR EXCESSIVE MANIPULATION OF THE FLEX WHEEL. THE COMPLAINT FOR DIFFICULT TO TRACK SYSTEM THROUGH ANATOMY' WAS EMPIRICALLY CONFIRMED BASED ON EVALUATION OF PROVIDED MEDICAL RECORDS. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION) LIKELY CONTRIBUTED TO THE EVENT AS PER MEDICAL RECORDS THERE WERE A SEVERE DEGREE OF CALCIFICATION AT THE AORTIC ARCH. PATIENT FACTORS (I.E. CALCIFICATION, TORTUOSITY, SMALL VESSEL SIZE, PRE-EXISTING VASCULAR STENT) MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERES WITH PASSAGE OF THE DELIVERY SYSTEM AND CAUSES DIFFICULTY DURING TRACKING/CROSSING. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22124693&pc=DSQ | 2916596-2025-03096 | 2022-10-01 04:00 | Death | THORATEC CORPORATION | 2025-06-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE "VENTRICULAR ARRHYTHMIAS IN PATIENTS WITH CARDIAC SARCOIDOSIS FOLLOWING LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION" THAT THE HEARTMATE 3 MAY BE ASSOCIATED WITH VENTRICULAR ARRHYTHMIAS (VA), THE NEED TO UNDERGO A HEART TRANSPLANT AND DEATH. THIS RETROSPECTIVE SINGLE CENTER STUDY EVALUATED 543 PATIENTS WHO UNDERWENT DURABLE LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION FROM 01OCT2006 TO 01OCT2022. 39 OF THESE PATIENTS HAS CARDIAC SARCOIDOSIS (CS) WITH A MEAN AGE OF 51 AND THE REMAINING 504 PATIENTS HAD OTHER NON-ISCHEMIC CARDIOMYOPATHIES (NICMS) WITH A MEAN AGE OF 55. OF THE PATIENTS WITH CS, 26 WERE MALE, 15 HAD A HISTORY OF VENTRICULAR TACHYCARDIA (VT), 5 HAD A HISTORY OF VT ABLATION, 28 HAD AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) PRESENT AT IMPLANT, 6 HAD AN ICD SHOCK IN THE 7 DAYS PRIOR TO IMPLANT, 6 HAD A PREOPERATIVE ELECTRICAL STORM, AND 19 HAD A HEARTMATE 3 IMPLANTED. OF THE PATIENTS WITH NICMS, 346 WERE MALE, 170 HAD A HISTORY OF VENTRICULAR TACHYCARDIA (VT), 18 HAD A HISTORY OF VT ABLATION, 337 HAD AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) PRESENT AT IMPLANT, 42 HAD AN ICD SHOCK IN THE 7 DAYS PRIOR TO IMPLANT, 17 HAD A PREOPERATIVE ELECTRICAL STORM, AND 216 HAD A HEARTMATE 3 IMPLANTED. A MAJORITY OF PATIENTS WITH CS HAD AN INTERAGENCY REGISTRY FOR MECHANICALLY ASSISTED CIRCULATORY SUPPORT (INTERMACS) SCORE OF 1-2 AT THE TIME OF LVAD IMPLANTATION (N=25,64.1%), AND 28.2% OF PATIENTS WITH CS RECEIVED LVAD IMPLANTATION AS DESTINATION THERAPY. PRIOR TO LVAD IMPLANTATION, 20 PATIENTS (51.3) WERE RECEIVING IMMUNOSUPPRESSION. 322 OF PATIENTS WITH NCIMS HAD AN INTERMACS SCORE OF 1-2 AT THE TIME OF LVAD IMPLANTATION, AND 40.2% OF PATIENTS WITH NCIMS RECEIVED LVAD IMPLANTATION AS DESTINATION THERAPY. THERE WAS NO SIGNIFICANT DIFFERENCE IN THE PREOPERATIVE INCIDENCE OF SUSTAINED VAS BETWEEN PATIENTS WITH CS AND PATIENTS WITH OTHER NICMS (VAS:38.5% VS 33.7%; P=0.212). AFTER LVAD IMPLANTATION, EARLY VENTRICULAR ARRHYTHMIAS (LESS THAN OR EQUAL TO 30 DAYS AFTER LVAD IMPLANTATION) AND LATE VENTRICULAR ARRHYTHMIAS (GREATER THAN 30 DAYS AFTER LVAD IMPLANTATION OCCURRED MORE COMMONLY IN PATIENTS WITH CS COMPARED TO PATIENTS WITH OTHER NICMS (EARLY VENTRICULAR ARRHYTHMIAS 23.1% VS 10.9%; P+0.023; LATE VENTRICULAR ARRHYTHMIAS: 39.5% VS 24.6%; P=0.044). THIS RESULT OCCURRED DESPITE THERE BEING NO SIGNIFICANT DIFFERENCES BETWEEN THE PERCENTAGE OF PATIENTS WITH CS AND WITH OTHER NICMS WHO RECEIVED ANTIARRHYTHMIC MEDICATIONS FOLLOWING LVAD IMPLANTATION (CS: 46.2% VS NICM: 32.9%; P=0.113). AFTER LVAD IMPLANTATION, 23 PATIENTS WITH CS (59.0%) RECEIVED IMMUNOSUPPRESSION. THERE WERE NO SIGNIFICANT DIFFERENCES IN 30-DAY MORTALITY RATES BETWEEN THE 2 GROUPS AFTER LVAD IMPLANTATION (CS: 0% VS NICM: 6.2%; P=0.111). AMONG PATIENTS SURVIVING GREATER THAN 30 DAYS, ALL-CAUSE MORTALITY WAS COMPARABLE IN BOTH GROUPS (CS: 43.6% VS NICM: 47.1%; P=0.679). THERE WAS NO SIGNIFICANT DIFFERENCE IN TRANSPLANT FREE SURVIVAL RATES WITHOUT VENTRICULAR ARRHYTHMIAS IN THE 2 GROUPS (HR 0.66295% CI 0.389-1.127; P=0.5128). Manufacturer Narrative: SECTION B: REPORT TYPE AND OUTCOME CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF THE EVENT HAS BEEN ENTERED AS 01OCT2022 AS THE DATA WAS COLLECTED BETWEEN 01OCT2006 AND 01OCT2022. GEORGETOWN UNIVERSITY SCHOOL OF MEDICINE, WASHINGTON, D.C.; MEDSTAR HEART AND VASCULAR INSTITUTE, MEDSTAR WASHINGTON HOSPITAL CENTER, WASHINGTON D.C.; HELMSLEY ELECTROPHYSIOLOGY CENTER, DEPARTMENT OF CARDIOLOGY, ICAHN SCHOOL OF MEDICINE AT MOUNT SINAI, NEW YORK, NY BERMUDEZ, F., RIZK, A. A., SHAH, M. H., SHEIKH, F. H., & OATES, C. P. (2025). VENTRICULAR ARRHYTHMIAS IN PATIENTS WITH CARDIAC SARCOIDOSIS FOLLOWING LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION. JOURNAL OF CARDIAC FAILURE, 31(2), 485-488. DOI:HTTP://DX.DOI.ORG/10.1016/J.CARDFAIL.2024.10.442. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22124987&pc=OZD | 1220648-2025-28899 | 2025-05-10 04:00 | Death | ABIOMED, INC. | 2025-06-02 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP AND DURING SUPPORT, BLEEDING NOTED AROUND THE PLACEMENT SHEATH, STAFF DECIDED TO NOT REMOVE THE PEEL AWAY AND PATIENT CONTINUED TO BLEED AT BEDSIDE. BLOOD PRODUCTS WERE GIVEN TO THE PATIENT. ADDITIONALLY, COMPUTED TOMOGRAPHY SCAN SHOWED LARGE AMOUNTS OF SWELLING IN THE PATIENT'S BRAIN AND BRAIN BLEEDING. THE PATIENT WAS TAKEN OFF OF HEPARIN AT THAT TIME. SUPPORT CONTINUED UNTIL CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22125141&pc=OZD | 1220648-2025-28900 | 2025-05-07 04:00 | Death | ABIOMED, INC. | 2025-06-02 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT SUFFERED VENTRICULAR FIBRILLATION WHILE IMPLANTED WITH AN IMPELLA 5.5. CARDIOPULMONARY RESUSCITATION WAS ATTEMPTED BUT THE PATIENT ULTIMATELY PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22125148&pc=NPT | 2025587-2025-03996 | 2025-01-30 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-02 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 YEAR AND 10 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT PASSED AWAY AT HOME. THE CAUSE OF DEATH IS UNKNOWN. THE RELATIONSHIP BETWEEN THE DEATH AND THE DEVICE OR PROCEDURE IS UNKNOWN. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22125291&pc=OZD | 1220648-2025-28901 | 2025-05-10 04:00 | Death | ABIOMED, INC. | 2025-06-02 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia; Thrombocytopenia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT DURING IMPELLA CP PREP, COMPLICATIONS WERE ENCOUNTERED WITH A UTILIZED PROSTYLE RESULTING IN MANUAL PRESSURE BEING HELD AND TWO UNITS OF BLOOD BEING GIVEN. FOLLOWING IMPLANT, A FLUORO REVEALED NO DISTAL FLOW PAST THE IMPELLA SHEATH. A FEM-FEM BYPASS WAS PERFORMED TO ALLOW REPERFUSION OF THE EXTREMITY. ONCE THE POSITIONING SHEATH WAS PLACED, OOZING CONTINUED AT THE SITE. MANUAL PRESSURE, FEMOSTOP, AND A HEMOSTATIC PATCH WERE UTILIZED TO ACHIEVE HEMOSTASIS. TWO DAYS INTO SUPPORT, HEPARIN WAS STOPPED DUE TO CONCERNS FOR LOW PLATELETS AND THROMBOCYTOPENIA. SUPPORT CONTINUED FOR ONE MORE DAY UNTIL THE DECISION WAS MADE BY THE PATIENT'S FAMILY TO WITHDRAW CARE AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22125417&pc=GZF | 3010676138-2025-00115 | 2025-05-07 04:00 | Death | CURONIX LLC | 2025-06-02 04:00 | GZF | FREEDOM PERIPHERAL NERVE STIMULATOR | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K171366 |  | 1 | Event Description: THE COMMERCIAL TEAM MEMBER WAS NOTIFIED OF THE PATIENT'S DEATH WHICH OCCURRED ON (B)(6) 2025. Manufacturer Narrative: THE OTHER ADVERSE EVENTS ISSUES QUESTIONNAIRE WAS REVIEWED. THE CAUSE OF THE PATIENT'S DEATH IS DUE TO A BRAIN HEMORRHAGE WHICH WAS UNRELATED TO THE CURONIX DEVICE. THERE ARE NO ALLEGED DEFICIENCIES ASSOCIATED WITH THE CURONIX DEVICE. A CURONIX REPRESENTATIVE CONDUCTED A REVIEW OF STERILIZATION AND PACKAGING RECORDS FOR THE RESPECTIVE PRODUCT LOT; CURONIX HAS CONFIRMED THAT THE PRODUCT WAS DELIVERED STERILE, VALIDATED STERILIZATION PARAMETERS WERE USED, AND STERILE BARRIERS WERE VERIFIED TO BE INTACT FOLLOWING PACKAGING. THE STIMULATOR IS USED TO TREAT PAIN. THE CAUSE OF THE PATIENT'S DEATH IS UNRELATED TO THE CURONIX DEVICE. THE PROVIDED INFORMATION DOES NOT EVIDENCE THAT THE CURONIX DEVICE CONTRIBUTED TO THE ISSUE (NO FAULT FOUND). RATES REVIEWED AT THE MOST RECENT COMPLAINT TRENDING MEETING DO NOT INDICATE A SIGNIFICANT INCREASE IN OTHER ADVERSE EVENTS ISSUES. OTHER ADVERSE EVENTS ISSUE RATES REMAIN ACCEPTABLY LOW; THUS, A CAPA IS NOT REQUIRED. OTHER ADVERSE EVENTS ISSUE RATES WILL CONTINUE TO BE TRACKED AND TRENDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22125514&pc=BTR | 9681384-2025-00626 | 2025-05-16 04:00 | Death | COVIDIEN AG JUAREZ (MMJ) | 2025-06-02 04:00 | BTR | UNSPECIFIED ENDOTRACH TUBE | Disconnection | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT, WHILE IN USE ON A PATIENT WITH A ENDOTRACHEAL TUBE (ETT), THIS 980 VENTILATOR DID NOT ALARM WHEN THE ETT TUBE BECAME DISLODGED. THE ETT BALLOON INFLATION WAS VERIFIED 20 MINUTES PRIOR TO THE INCIDENT BUT WHEN THE NURSE MONITOR ALARMED AND PATIENT WAS CHECKED, THE ETT TUBE WAS DEFLATED. THE PATIENT EXPIRED IMMEDIATELY AFTER THE INCIDENT. Manufacturer Narrative: D10 CONCOMITANT PRODUCT: 980X3ENDIUU, 980 VENT; CMPRSSR CONFIG; US; LATAM, LOT# 35B2001523. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22125818&pc=BZD | 2518422-2025-046312 | 2023-10-09 04:00 | Death | RESPIRONICS, INC. | 2025-06-02 04:00 | BZD | BIPAP AUTO SERIES ASSY | Degraded | Asthma; Headache; Nausea; Pulmonary Dysfunction; Vomiting; Dizziness; Unspecified Respiratory Problem; Unspecified Hepatic or Biliary Problem; Unspecified Kidney or Urinary Problem; Skin Inflammation/ Irritation | K091319 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, ASTHMA, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE AND REDUCED CARDIOPULMONARY RESERVE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22126490&pc=DSQ | 2916596-2025-03522 | 2025-04-07 04:00 | Death | THORATEC CORPORATION | 2025-06-02 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO UNKNOWN ETIOLOGY. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO HOSPITAL PRIVACY POLICY, NO ADDITIONAL INFORMATION WOULD BE PROVIDED. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22126574&pc=OZD | 1220648-2025-28912 | 2025-05-13 04:00 | Death | ABIOMED, INC. | 2025-06-02 04:00 | OZD | IMPELLA CP | Pumping Stopped; Infusion or Flow Problem | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AN IMPLANTED IMPELLA CP PUMP STOPPED UNEXPECTEDLY DURING PATIENT SUPPORT. THE PUMP WAS RESTARTED ON P-4, BUT THE MEAN, MAX, AND MINIMUM FLOWS WERE ALL AT 3.4 L/MIN. THE PUMP WAS SUBSEQUENTLY EXPLANTED AS A RESULT OF THE NOTED FAILURES. LATER, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: A5 IS UNKNOWN. THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22126685&pc=BZD | 2518422-2025-106289 | 2025-05-12 04:00 | Death | RESPIRONICS, INC. | 2025-06-02 04:00 | BZD | REP DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS INFORMED THAT A PATIENT USING A REP DREAMSTATION AUTO BIPAP DEVICE PASSED AWAY ON (B)(6) 2024. ACCORDING TO THE PATIENT¿S WIFE, HE WAS NOT USING THE DEVICE AT THE TIME OF HIS DEATH, AS HE WAS UNDER HOSPICE CARE. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE THIRD-PARTY SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND THE DEVICE WAS WORKING PROPERLY. NO DEVICE ERRORS OR MALFUNCTIONS WERE OBSERVED. THE DEVICE WAS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22126951&pc=DSQ | 2916596-2025-03252 | 2025-04-17 04:00 | Death | THORATEC CORPORATION | 2025-06-02 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Muscle Weakness; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY WHILE ON PALLIATIVE CARE/END OF LIFE. THE PATIENT WAS VERY WEAK AND PALLIATIVE CARE WAS REQUESTED. THE PUMP PERFORMED AS INTENDED. THIS WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. THE PATIENT WAS IN A WHEELCHAIR WITH NO STRENGTH LEADING TO PALLIATIVE CARE.Event Description: THE PATIENT WAS IN A WHEELCHAIR DUE TO HAVING NO MUSCLE STRENGTH. THIS STARTED 1 YEAR AGO. THE DEVICE WAS DECOMMISSIONED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THERE WERE NO DEVICE RELATED ISSUES WITH (B)(6) REPORTED OR IDENTIFIED THROUGH THIS EVALUATION. IT WAS DETERMINED THAT THE REPORTED INFORMATION DID NOT MEET THE REQUIREMENTS OF A COMPLAINT. THE PATIENT WAS IN A WHEELCHAIR WITH NO STRENGTH LEADING TO PALLIATIVE CARE. THE PATIENT WAS IN A WHEELCHAIR DUE TO HAVING NO MUSCLE STRENGTH. THE DEVICE OPERATED AS INTENDED UNTIL BEING DECOMMISSIONED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22127259&pc=DSQ | 2916596-2025-03322 | 2025-04-26 04:00 | Death | THORATEC CORPORATION | 2025-06-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO WITHDRAWAL OF CARE AND RIGHT HEART FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED RIGHT HEART FAILURE AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS RIGHT HEART FAILURE AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAD. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22127503&pc=DSQ | 3007042319-2025-01216 | 2025-05-29 04:00 | Death | HEARTWARE, INC. | 2025-06-02 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Stopped; Power Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WHO HAD RECEIVED A VENTRICULAR ASSIST DEVICE (VAD) PATIENT DIED DUE TO UNKNOWN CAUSES. THE CAUSE OF DEATH REMAINED UNCLEAR DESPITE AN AUTOPSY BEING PERFORMED WHICH DID NOT IDENTIFY ANY PATHOLOGICAL FINDINGS. PER LOG FILE REVIEW THERE WAS A CONTROLLER LOSS OF POWER AND THE VAD WAS NOT RUNNING. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (B)(6) AND THE CONTROLLER (B)(6) WERE NOT RETURNED FOR EVALUATION AS THE VAD REMAINS IMPLANTED AND THE CONTROLLER WAS REMOVED FROM SERVICE. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS REVEALED A CONTROLLER POWER-UP AND ASSOCIATED PUMP START EVENT LOGGED ON 25/MAY/2025 AT 17:03:44, INDICATING A LOSS OF POWER TO THE CONTROLLER. THE DATA POINT RECORDED PRIOR TO THE LOSS OF POWER INDICATED THAT (B)(6) WAS CONNECTED TO POWER PORT ONE (1) WITH 24% RELATIVE STATE OF CHARGE (RSOC) AND (B)(6) WAS CONNECTED TO POWER PORT TWO (2) WITH 88% RSOC. THE DATA POINT RECORDED AFTER THE LOSS OF POWER INDICATED THAT (B)(6) WAS CONNECTED TO POWER PORT ONE (1) AND (B)(6) WAS CONNECTED TO POWER PORT TWO (2). THERE WERE NO ANOMALIES LEADING UP TO THE LOSS OF POWER EVENT. THE CONTROLLER WAS WITHOUT POWER FOR THIRTY-SIX (36) SECONDS. ADDITIONALLY, LOG FILE ANALYSIS REVEALED NO ALARMS WITHIN THE ANALYZED PERIOD. AS A RESULT, THE REPORTED CONTROLLER LOSS OF POWER EVENT WAS CONFIRMED. THE REPORTED VAD STOP EVENT COULD NOT BE CONFIRMED; HOWEVER, IT IS LIKELY THAT THE LOSS OF POWER WAS PERCEIVED AS THE REPORTED VAD STOP EVENT. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH IS UNKNOWN. THE MOST LIKELY ROOT CAUSE OF THE LOSS OF POWER TO THE CONTROLLER CAN BE ATTRIBUTED TO A DISCONNECTION OF BOTH POWER SOURCES AND/OR TO AN INTERMITTENT DISCONNECTION ON ONE OR BOTH POWER SOURCES. CAPA PR00551638 INVESTIGATED CONTROLLER LOSSES OF POWER. EVEN THOUGH THIS CAPA IS NOW CLOSED, (B)(6) FALLS IN SCOPE OF THIS CAPA. PER THE INSTRUCTIONS FOR USE, DEATH IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. ADDITIONAL PRODUCTS: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM CONTROLLER 2.0 D4: MODEL #: 1420 / CATALOG #: 1420 / EXPIRATION DATE: UNKNOWN / SERIAL OR LOT#: (B)(6); D9: NO H3: NO H4: MFG DATE: UNKNOWN H5: NO H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22127624&pc=LWP | 2017865-2025-75240 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-02 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: DEVICE RETURNED FOR ANALYSIS DUE TO PATIENT DECEASED. ANALYSIS PERFORMED, INCLUDING ELECTRICAL, MECHANICAL AND ENVIRONMENTAL TESTS INDICATED THAT THE DEVICE EXHIBITED NORMAL DEVICE CHARACTERISTICS WITH NO ANOMALIES. LONGEVITY ASSESSMENT WAS PERFORMED, AND DEVICE WAS IN THE NORMAL RANGE OF OPERATION WITH APPROPRIATE REMAINING LONGEVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22127638&pc=NVN | 2017865-2025-75241 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-02 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: FINAL ANALYSIS NOTES AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF ANY CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES EXCEPT FOR PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22127645&pc=NVN | 2017865-2025-75242 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-02 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPIRED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THE DEATH WAS RELATED TO THE DEVICE. IT WAS REPORTED THAT THE CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: FINAL ANALYSIS NOTES AS RECEIVED, ONLY THE CONNECTOR PORTION OF THE LEAD WAS RETURNED IN ONE PIECE. VISUAL INSPECTION OF THE PARTIAL LEAD DID NOT FIND ANY ANOMALIES. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. Manufacturer Narrative: CORRECTION: SECTION B1: IN THE INITIAL REPORT, "ADVERSE EVENT" SHOULD HAVE BEEN SELECTED ONLY INSTEAD OF "PRODUCT PROBLEM". "PRODUCT PROBLEM" WAS SELECTED IN ERROR AND SHOULD HAVE BEEN BLANK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22128070&pc=LYJ | 1644487-2025-10198 | 2025-05-07 04:00 | Death | LIVANOVA USA, INC. | 2025-06-02 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY, THE CAUSE OF DEATH AND RELATIONSHIP TO VNS/SEIZURES IS UNKNOWN. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22128628&pc=LWS | 2124215-2025-36087 | 2021-05-14 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-02 04:00 | LWS | EMBLEM? MRI S-ICD | Signal Artifact/Noise; High impedance; Over-Sensing; Pacing Problem; Failure to Read Input Signal; Inaccurate Synchronization; Under-Sensing; Use of Device Problem | Ventricular Fibrillation | P110042 |  | 1 | Event Description: IT WAS REPORTED THAT THIS SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (S-ICD) DEVICE AND ELECTRODE STORED AN EPISODE SHOWING SMART PASS WAS DISABLED. A REVIEW OF THE EPISODE SHOWED A NON-PHYSIOLOGIC NOISE ARTIFACT THAT WAS OVERSENSED. ADDITIONALLY, SMALL AMPLITUDE R-WAVES WERE NOTED. TECHNICAL SERVICES RECOMMENDED SEEING THE PATIENT FOR TROUBLESHOOTING AND TO PERFORM X-RAYS. THE DEVICE WAS CURRENTLY PROGRAMMED TO THE SECONDARY SENSE VECTOR. X-RAYS WERE NOT CONCLUSIVE FOR AN ELECTRODE FRACTURE. TROUBLESHOOTING WITH PATIENT POSTURES AND MOVEMENTS WAS PERFORMED AND NOISE WAS PRODUCED WHEN THE PATIENT TURNED THEIR UPPER TORSO TO THE LEFT AND TO THE RIGHT. TECHNICAL SERVICES REVIEWED THE DATA AND CONFIRMED THE SHOCK IMPEDANCE MEASUREMENTS WERE WITHIN NORMAL RANGE. THE PRIMARY SENSE VECTOR APPEARED THE BEST OPTION. THE HEALTH CARE PROFESSIONAL (HCP) CONFIRMED THE DEVICE WAS PROGRAMMED TO PRIMARY AND THE BEEPING TONES WERE DEMONSTRATED FOR THE PATIENT, AND THE S-ICD SYSTEM WOULD CONTINUE TO BE MONITORED IN THE REMOTE MONITORING SYSTEM. FOUR YEARS LATER, IT WAS REPORTED THIS PATIENT HAD PASSED AWAY. THE PATIENT HAD BEEN NON-COMPLIANT WITH REMOTE MONITORING TRANSMISSIONS AND IN-CLINIC APPOINTMENTS. A DEVICE INTERROGATION POST-MORTEM SHOWED THE PATIENT HAD RECEIVED APPROPRIATE SHOCKS FROM THE DEVICE ON THE DATE OF DEATH, AND SOME SHOCKS HAD HIGH, OUT-OF-RANGE IMPEDANCE MEASUREMENTS. ADDITIONALLY, DURING THE SHOCKS, SOME UNDERSENSING AND OVERSENSING WAS OBSERVED, WITH SOME INHIBITION OF POST-SHOCK PACING. ONE EPISODE SHOWED THAT A DELIVERED SHOCK INDUCED A VENTRICULAR FIBRILLATION (VF). THE SYSTEM WAS EXPLANTED AND RETURNED FOR ANALYSIS. AN ELECTRODE FRACTURE WAS SUSPECTED.Event Description: IT WAS REPORTED THAT THIS SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (S-ICD) DEVICE AND ELECTRODE STORED AN EPISODE SHOWING SMART PASS WAS DISABLED. A REVIEW OF THE EPISODE SHOWED A NON-PHYSIOLOGIC NOISE ARTIFACT THAT WAS OVERSENSED. ADDITIONALLY, SMALL AMPLITUDE R-WAVES WERE NOTED. TECHNICAL SERVICES RECOMMENDED SEEING THE PATIENT FOR TROUBLESHOOTING AND TO PERFORM X-RAYS. THE DEVICE WAS CURRENTLY PROGRAMMED TO THE SECONDARY SENSE VECTOR. X-RAYS WERE NOT CONCLUSIVE FOR AN ELECTRODE FRACTURE. TROUBLESHOOTING WITH PATIENT POSTURES AND MOVEMENTS WAS PERFORMED AND NOISE WAS PRODUCED WHEN THE PATIENT TURNED THEIR UPPER TORSO TO THE LEFT AND TO THE RIGHT. TECHNICAL SERVICES REVIEWED THE DATA AND CONFIRMED THE SHOCK IMPEDANCE MEASUREMENTS WERE WITHIN NORMAL RANGE. THE PRIMARY SENSE VECTOR APPEARED THE BEST OPTION. THE HEALTH CARE PROFESSIONAL (HCP) CONFIRMED THE DEVICE WAS PROGRAMMED TO PRIMARY AND THE BEEPING TONES WERE DEMONSTRATED FOR THE PATIENT, AND THE S-ICD SYSTEM WOULD CONTINUE TO BE MONITORED IN THE REMOTE MONITORING SYSTEM. FOUR YEARS LATER, IT WAS REPORTED THIS PATIENT HAD PASSED AWAY. THE PATIENT HAD BEEN NON-COMPLIANT WITH REMOTE MONITORING TRANSMISSIONS AND IN-CLINIC APPOINTMENTS. A DEVICE INTERROGATION POST-MORTEM SHOWED THE PATIENT HAD RECEIVED APPROPRIATE SHOCKS FROM THE DEVICE ON THE DATE OF DEATH, AND SOME SHOCKS HAD HIGH, OUT-OF-RANGE IMPEDANCE MEASUREMENTS. ADDITIONALLY, DURING THE SHOCKS, SOME UNDERSENSING AND OVERSENSING WAS OBSERVED, WITH SOME INHIBITION OF POST-SHOCK PACING. ONE EPISODE SHOWED THAT A DELIVERED SHOCK INDUCED A VENTRICULAR FIBRILLATION (VF). THE SYSTEM WAS EXPLANTED AND RETURNED FOR ANALYSIS. AN ELECTRODE FRACTURE WAS SUSPECTED. Manufacturer Narrative: THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS. Manufacturer Narrative: THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS. UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS DEVICE WAS INSPECTED AND ANALYZED. VISUAL EXAMINATION IDENTIFIED NO ANOMALIES. THE DEVICE WAS ABLE TO BE INTERROGATED AND A MEMORY DOWNLOAD WAS PERFORMED SUCCESSFULLY. THE DEVICE WAS THEN EXPOSED TO SIMULATED HEART LOAD CONDITIONS, AND THE DEFIBRILLATION AND SENSING FUNCTIONS WERE TESTED. THE DEVICE OPERATED APPROPRIATELY, ACCORDING TO ITS PERFORMANCE SPECIFICATIONS WITH NO OUT-OF-RANGE MEASUREMENTS OR INTERRUPTIONS IN THERAPY OUTPUT. ANALYSIS DID NOT IDENTIFY ANY DEVICE CHARACTERISTICS THAT WOULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED CLINICAL OBSERVATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22128630&pc=LWS | 2124215-2025-36089 | 2021-05-14 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-02 04:00 | LWS | EMBLEM? S-ICD | Signal Artifact/Noise; Fracture; High impedance; Over-Sensing; Pacing Problem; Failure to Read Input Signal; Under-Sensing | Ventricular Fibrillation | P110042 |  | 1 | Event Description: IT WAS REPORTED THAT THIS SUBCUTANEOUS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (S-ICD) DEVICE AND ELECTRODE STORED AN EPISODE SHOWING SMART PASS WAS DISABLED. A REVIEW OF THE EPISODE SHOWED A NON-PHYSIOLOGIC NOISE ARTIFACT THAT WAS OVERSENSED. ADDITIONALLY, SMALL AMPLITUDE R-WAVES WERE NOTED. TECHNICAL SERVICES RECOMMENDED SEEING THE PATIENT FOR TROUBLESHOOTING AND TO PERFORM X-RAYS. THE DEVICE WAS CURRENTLY PROGRAMMED TO THE SECONDARY SENSE VECTOR. X-RAYS WERE NOT CONCLUSIVE FOR AN ELECTRODE FRACTURE. TROUBLESHOOTING WITH PATIENT POSTURES AND MOVEMENTS WAS PERFORMED AND NOISE WAS PRODUCED WHEN THE PATIENT TURNED THEIR UPPER TORSO TO THE LEFT AND TO THE RIGHT. TECHNICAL SERVICES REVIEWED THE DATA AND CONFIRMED THE SHOCK IMPEDANCE MEASUREMENTS WERE WITHIN NORMAL RANGE. THE PRIMARY SENSE VECTOR APPEARED THE BEST OPTION. THE HEALTH CARE PROFESSIONAL (HCP) CONFIRMED THE DEVICE WAS PROGRAMMED TO PRIMARY AND THE BEEPING TONES WERE DEMONSTRATED FOR THE PATIENT, AND THE S-ICD SYSTEM WOULD CONTINUE TO BE MONITORED IN THE REMOTE MONITORING SYSTEM. FOUR YEARS LATER, IT WAS REPORTED THIS PATIENT HAD PASSED AWAY. THE PATIENT HAD BEEN NON-COMPLIANT WITH REMOTE MONITORING TRANSMISSIONS AND IN-CLINIC APPOINTMENTS. A DEVICE INTERROGATION POST-MORTEM SHOWED THE PATIENT HAD RECEIVED APPROPRIATE SHOCKS FROM THE DEVICE ON THE DATE OF DEATH, AND SOME SHOCKS HAD HIGH, OUT-OF-RANGE IMPEDANCE MEASUREMENTS. ADDITIONALLY, DURING THE SHOCKS, SOME UNDERSENSING AND OVERSENSING WAS OBSERVED, WITH SOME INHIBITION OF POST-SHOCK PACING. ONE EPISODE SHOWED THAT A DELIVERED SHOCK INDUCED A VENTRICULAR FIBRILLATION (VF). THE SYSTEM WAS EXPLANTED AND RETURNED FOR ANALYSIS. AN ELECTRODE FRACTURE WAS SUSPECTED. Manufacturer Narrative: THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22128632&pc=NIQ | 2124215-2025-31001 | 2025-05-05 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-02 04:00 | NIQ | SYNERGY | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Chest Pain; Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH CORONARY ARTERY DISEASE AND UNDERWENT PERCUTANEOUS CORONARY INTERVENTION. THE 95% STENOSED TARGET LESION WAS LOCATED IN THE SEVERELY TORTUOUS AND MODERATELY CALCIFIED LEFT ANTERIOR DESCENDING ARTERY (LAD). AFTER A WELL-VESSEL PREPARATION, A 3.00 X 38 SYNERGY DRUG-ELUTING STENT WAS DEPLOYED IN LAD. SECOND AND THIRD NON-BOSTON SCIENTIFIC STENTS WERE ALSO DEPLOYED RESPECTIVELY IN THE LEFT CIRCUMFLEX ARTERY AND THE RIGHT CORONARY ARTERY. IT WAS NOTED THAT THE PATIENT CONDITION WAS BAD ON THE CARDIAC CATHETERIZATION TABLE. FOLLOWING THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO THE CARDIAC CARE UNIT. AFTER SOME TIME, THE PATIENT DEVELOPED HYPOTENSION, EXPERIENCED CHEST PAIN, AND WENT INTO CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS IMMEDIATELY PERFORMED BUT THE PATIENT DIED. THE OFFICIAL CAUSE OF DEATH WAS CARDIAC ARREST, AND AN AUTOPSY WAS NOT PERFORMED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22129002&pc=DSQ | 2916596-2025-03210 | 2023-07-01 04:00 | Death | THORATEC CORPORATION | 2025-06-02 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE, "RELATIONSHIP OF PAD AND USE OF TMCS ON OUTCOMES POST-LVAD: A CASE SERIES OF 197 PATIENTS" THAT THE HEARTMATE 3 (HM3) LEFT VENTRICULAR ASSIST DEVICE (LVAD) MAY BE ASSOCIATED WITH STROKE, BLEEDING, RENAL FAILURE, VASCULAR COMPLICATION, AND DEATH. THE PURPOSE OF THE STUDY WAS TO QUANTIFY THE IMPACT OF PERIPHERAL ARTERY DISEASE (PAD) ON VASCULAR COMPLICATIONS, PERIOPERATIVE MORBIDITY, AND SURVIVAL IN LVAD RECIPIENTS. THE RETROSPECTIVE COHORT INCLUDED 197 PATIENTS WHO UNDERWENT HM3 LVAD IMPLANT BETWEEN SEP2015 AND JUL2023. THE COHORT WAS DIVIDED INTO THOSE WITH SEVERE PAD AND THOSE WITHOUT, AS WELL AS THOSE THAT REQUIRED THE USE OF TEMPORARY MECHANICAL CIRCULATORY SUPPORT (TMCS). 179 PATIENTS (90%) HAD PREOPERATIVE LOWER EXTREMITY ARTERIAL DUPLEX IMAGING TO EVALUATE FOR PERIPHERAL VASCULAR DISEASE. SEVERE PAD WAS DEFINED AS HAVING ANY OF THE FOLLOWING: TOE PRESSURE <50 MMHG, INFLOW DISEASE, OR OUTFLOW DISEASE. OF THE 197 PATIENTS WHO UNDERWENT HM3 IMPLANTATION, 22 (11%) WERE IDENTIFIED AS HAVING SEVERE PAD SONOGRAPHICALLY. PATIENTS WITH SEVERE PAD WERE MORE LIKELY TO HAVE ISCHEMIC CARDIOMYOPATHY (35% VS. 68%, P=0.014). THERE WERE NO DIFFERENCES IN AGE, BODY MASS INDEX, INTERAGENCY REGISTRY FOR MECHANICALLY ASSISTED CIRCULATORY SUPPORT (INTERMACS) CLASSIFICATION, CARDIOPULMONARY BYPASS TIMES, OR USE OF PERIOPERATIVE TMCS. OF NOTE, FIVE PATIENTS OF THE PAD COHORT HAD A PREOPERATIVE HISTORY OF CLAUDICATION. PRIMARY OUTCOMES WERE CHARACTERIZED AS THE INCIDENCE OF VASCULAR COMPLICATIONS, DEFINED AS THE DEVELOPMENT OF DRY GANGRENE, NEED FOR LOWER EXTREMITY AMPUTATION, OR REVASCULARIZATION. SECONDARY OUTCOMES INCLUDED STROKES, RENAL FAILURE, DEATH DURING IMPLANT HOSPITALIZATION, AND OVERALL SURVIVAL (OS). REGARDING POSTOPERATIVE COMPLICATIONS, 12 PATIENTS DEVELOPED VASCULAR COMPLICATIONS; AND THESE WERE MORE COMMON IN THE GROUP OF PATIENTS WITH SEVERE PAD (23% VS. 4%, P = 0.005). THERE WAS NO DIFFERENCE IN OTHER COMPLICATIONS SUCH AS RENAL FAILURE, STROKE, BLEEDING, LENGTH OF STAY, OR DEATH DURING IMPLANT HOSPITALIZATION. 7 PATIENTS WHO DID NOT HAVE PREOPERATIVE VASCULAR DISEASE DEVELOPED A VASCULAR COMPLICATION, OFTEN IN THE SETTING OF SEVERE SHOCK OR USE OF TMCS. AN UNADJUSTED ANALYSIS SHOWED THAT PAD WAS AN INDEPENDENT PREDICTOR OF VASCULAR COMPLICATIONS, WITH AN IRR OF 5.68 (95% CONFIDENCE INTERVAL [CI]: 1.80¿17.90, P = 0.003). OTHER SIGNIFICANT PREDICTORS INCLUDED POSTOPERATIVE MECHANICAL CIRCULATORY SUPPORT (MCS), WITH AN IRR OF 9.35 (95% CI: 2.81¿31.05, P < 0.001) AND FEMORAL CANNULATION, WITH AN IRR OF 4.04 (95% CI: 1.09¿14.93, P = 0.036). ONLY PREOPERATIVE MCS WAS IDENTIFIED AS A SIGNIFICANT FACTOR, WITH AN IRR OF 8.15 (95% CI: 1.51¿44.10, P = 0.015). THE MEDIAN SURVIVAL FOR THE COHORT WAS 64.6 MONTHS. THERE WAS NO DIFFERENCE IN OS (HAZARD RATIO [HR]: 1.25, 95% CI: 0.62¿2.54, P = 0.536). A SIGNIFICANTLY HIGHER MORTALITY WAS NOTED IN PATIENTS WHO HAD TMCS (HR: 2.00, 95% CI: 1.17¿3.44, P = 0.011). OF THE SECONDARY COMPLICATIONS: THERE WERE 4 PATIENTS THAT EXPERIENCED STROKE (3 WITHOUT PAD AND 1 WITH PAD, P=0.378). 16 PATIENTS EXPERIENCED BLEEDING (14 PATIENTS WITHOUT PAD AND 2 WITH PAD, P=0.694). 41 PATIENTS EXPERIENCED RENAL FAILURE (37 PATIENTS WITHOUT PAD AND 4 WITH PAD, P=0.79). 22 PATIENTS EXPERIENCED DEATH DURING IMPLANT HOSPITALIZATION (20 WITHOUT PAD AND 2 WITH PAD, P=0.514). 12 PATIENTS EXPERIENCED VASCULAR COMPLICATIONS (7 WITHOUT PAD AND 5 WITH PAD, P=0.005). THE AVERAGE NUMBER OF DAYS SPENT IN THE INTENSIVE CARE UNIT (ICU) WAS 12 DAYS, P=0.727. THE AVERAGE NUMBER OF HOURS SPENT ON VENTILATION WAS 29 HOURS, P=0.925. OF THE 12 PATIENTS THAT EXPERIENCED VASCULAR COMPLICATIONS: - PATIENT 1 HAD PAD AND CLAUDICATION PREOPERATIVELY. THEY DEVELOPED VASCULAR COMPLICATIONS 250 DAYS AFTER LVAD IMPLANT. THEY HAD ACUTE LIMB ISCHEMIA AND A THROMBOSED LOWER EXTREMITY GRAFT. THEY DID NOT REQUIRE TMCS. - PATIENT 2 HAD PAD AND DID NOT HAVE CLAUDICATION PREOPERATIVELY. THEY DEVELOPED VASCULAR COMPLICATIONS 427 DAYS AFTER LVAD IMPLANT. THEY HAD A NONHEALING ULCER REQUIRING ANGIOPLASTY. THEY DID NOT REQUIRE TMCS. - PATIENT 3 DID NOT HAVE PAD AND DID NOT EXPERIENCE CLAUDICATION PREOPERATIVELY. THEY DEVELOPED VASCULAR COMPLICATIONS 273 DAYS AFTER LVAD IMPLANT. THEY HAD ACUTE LIMB ISCHEMIA, LEFT VENTRICULAR THROMBUS, AND SHOCK. THEY DID NOT REQUIRE TMCS. - PATIENT 4 DID NOT HAVE PAD AND DID NOT HAVE CLAUDICATION PREOPERATIVELY. THEY DEVELOPED VASCULAR COMPLICATIONS 19 DAYS AFTER LVAD IMPLANT AND HAD CRITICAL LIMB ISCHEMIA IN THE SETTING OF PUMP THROMBOSIS. THEY DID NOT REQUIRE TMCS. - PATIENT 5 DID NOT HAVE PAD AND DID NOT HAVE CLAUDICATION PREOPERATIVELY. THEY DEVELOPED VASCULAR COMPLICATIONS 10 DAYS AFTER LVAD IMPLANT WITH DRY GANGRENE IN THE SETTING OF SHOCK. THEY DID NOT REQUIRE TMCS. - PATIENT 6 DID NOT HAVE PAD AND DID NOT HAVE CLAUDICATION PREOPERATIVELY. THEY DEVELOPED VASCULAR COMPLICATIONS 720 DAYS AFTER LVAD IMPLANT WITH RAY AMPUTATION FOR A NONHEALING WOUND. THEY REQUIRED PRE AND INTRAOPERATIVE TMCS (CENTRAL BIVAD, AND CENTRAL RVAD) - IT WAS UNKNOWN IF PATIENT 7 HAD PAD AND THEY DID NOT HAVE CLAUDICATION PREOPERATIVELY. THEY DEVELOPED VASCULAR COMPLICATION BOTH 91 AND 175 DAYS AFTER LVAD IMPLANT. THEY HAD BILATERAL TRANS METATARSAL AMPUTATION FOR DRY GANGRENE AND REQUIRED A POSTOPERATIVE CENTRAL RVAD. - PATIENT 8 HAD PAD AND HAD CLAUDICATION PREOPERATIVELY. THEY DEVELOPED VASCULAR COMPLICATIONS 57 DAYS AFTER LVAD IMPLANT WITH BILATERAL DRY GANGRENE IN THE SETTING OF SHOCK. THEY DID NOT REQUIRE TMCS. - PATIENT 9 HAD PAD AND DID NOT HAVE CLAUDICATION PREOPERATIVELY. THEY DEVELOPED VASCULAR COMPLICATIONS 0 DAYS AFTER LVAD IMPLANT WITH CRITICAL LIMB ISCHEMIA AT THE FEMORAL CANNULATION SITE, WHICH PROGRESSED TO BILATERAL ISCHEMIA IN THE SETTING OF SEVERE SHOCK. THEY REQUIRED THE PROTEK DUO RVAD POSTOPERATIVELY. - PATIENT 10 DID NOT HAVE PAD AND DID NOT HAVE CLAUDICATION PREOPERATIVELY. THEY DEVELOPED VASCULAR COMPLICATIONS 10 DAYS AFTER LVAD IMPLANT WITH ISCHEMIA IN THE SETTING OF SHOCK. THEY REQUIRED PREOPERATIVE IMPELLA 5.5 AND POSTOPERATIVE PROTEK DUO. - PATIENT 11 HAD PAD AND CLAUDICATION PREOPERATIVELY. THEY DEVELOPED VASCULAR COMPLICATIONS 247 DAYS AFTER LVAD IMPLANT WITH A NONHEALING WOUND. THEY DID NOT REQUIRE TMCS. - IT WAS UNKNOWN IF PATIENT 12 HAD PAD AND THEY DID NOT HAVE CLAUDICATION PREOPERATIVELY. THEY DEVELOPED VASCULAR COMPLICATIONS 10 DAYS AFTER LVAD IMPLANT. THEY HAD EMERGENCY CANNULATION AT BEDSIDE LEADING TO LIMB ISCHEMIA REQUIRING THROMBECTOMY. THEY HAD A POSTOPERATIVE FEMORAL EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) CANNULATION (AT BEDSIDE). THE RISK OF DEVELOPING A VASCULAR SURGERY COMPLICATION POST- LVAD IMPLANTATION IN PATIENTS WITH SEVERE PAD IS 22% VS. 4% IN PATIENTS WITHOUT SEVERE PAD, BUT THIS MAY BE AN ACCEPTABLE RISK TO SOME PATIENTS WHO HAVE NO OTHER SURGICAL ALTERNATIVE TO HEART REPLACEMENT THERAPIES. THESE FINDINGS UNDERSCORE THE IMPORTANCE OF PREOPERATIVE RISK STRATIFICATION AND TAILORED PERIOPERATIVE MANAGEMENT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OR OTHER IDENTIFYING INFORMATION OF THE PRODUCTS WERE NOT REPORTED AND WERE NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. SECTION B3: DATE OF EVENT WAS ESTIMATED AND HAS BEEN ENTERED AS 01JUL2023 AS PATIENTS WERE IMPLANTED BETWEEN (B)(6) 2015 AND (B)(6) 2023. FROM (B)(6). TASOUDIS, P., WILLIAMS, M., SUKHAVASI, A., BUCKERIDGE, S., MOSTERTZ, W., BYKU, M., & MERLO, A. (2025). RELATIONSHIP OF PAD AND USE OF TMCS ON OUTCOMES POST-LVAD: A CASE SERIES OF 197 PATIENTS. ASAIO JOURNAL. HTTPS://DOI.ORG/10.1097/MAT.0000000000002428. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22129309&pc=GDW | 2955842-2025-23311 | 2025-04-24 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-06-02 04:00 | GDW | SUREFORM | Failure to Form Staple | Unspecified Tissue Injury; Insufficient Information | K173721 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A DA VINCI-ASSISTED ROUX-EN-Y GASTRIC BYPASS AND HIATAL HERNIA REPAIR. THE PROCEDURE WAS DESCRIBED AS CHALLENGING DUE TO THE PATIENT¿S WEIGHT; HOWEVER, THERE WERE NO SPECIFIC COMPLICATIONS REPORTED AND THE PROCEDURE WAS COMPLETED WITHOUT ANY DA VINCI SYSTEM OR INSTRUMENT ISSUES. THE INTUITIVE CLINICAL SALES REPRESENTATIVE WAS LATER INFORMED BY THE SURGEON THAT ONE-WEEK POSTOPERATIVE, THE PATIENT WAS HOSPITALIZED FOR UNSPECIFIED SYMPTOMS. AN OPEN LAPAROTOMY WAS PERFORMED WHERE A LEAK WAS IDENTIFIED AT THE JEJUNOJEJUNOSTOMY (JJ) ANASTOMOSIS. AT AN UNSPECIFIED TIME, THE PATIENT REPORTEDLY CODED DUE TO POSTOPERATIVE COMPLICATIONS AND EXPIRED APPROXIMATELY 2 WEEKS AFTER THE INITIAL ROBOTIC SURGERY. NO FURTHER DETAILS WERE PROVIDED. ADDITIONAL INFORMATION WAS REQUESTED; A RESPONSE HAS NOT BEEN RECEIVED. Manufacturer Narrative: REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS OCCURRED DURING THE PROCEDURE THAT WOULD BE RELATED TO THE REPORTED EVENT. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, FENESTRATED BIPOLAR FORCEPS, MEGA SUTURECUT NEEDLE DRIVER, MONOPOLAR CURVED SCISSORS, TIP-UP FENESTRATED GRASPER, VESSEL SEALER EXTEND, SUCTION IRRIGATOR, SUREFORM 60 STAPLER WITH WHITE RELOADS. A SITE HISTORY REVIEW SHOWS NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THESE PRODUCTS. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE SHOWED NO RELATED NON-CONFORMANCES WERE IDENTIFIED. A REVIEW OF THE SUREFORM 60 STAPLER LOG FOUND THAT THE INSTRUMENT WAS INSTALLED FIVE TIMES AND FIRED 5 WHITE RELOADS. THERE WERE NO INCOMPLETE CLAMPS OR UNCLAMPS BY THIS INSTRUMENT AND ALL FIRINGS WERE COMPLETED. THERE WERE NO STAPLER RELATED ERRORS IN THE LOGS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD MORBID OBESITY WITH A BMI OF 74 AND UNDERWENT A GASTRIC BYPASS WITH HIATAL HERNIA REPAIR. DETAILS ON HOW THE PROCEDURE WAS PERFORMED, AND WHICH INSTRUMENTS WERE USED AT THE ANASTOMOSIS WERE NOT PROVIDED. NO INTRAOPERATIVE ISSUES WERE REPORTED. THE PATIENT RETURNED ONE WEEK LATER AND WAS FOUND TO HAVE A LEAK. THE LOCATION OF THE LEAK WAS IDENTIFIED AT THE JEJUNOJEJUNOSTOMY (JJ) SITE. THE PATIENT SUBSEQUENTLY CODED AND DIED. WHEN THEY DIED AND THE EVENTS THAT LED TO THE DEATH ARE NOT PROVIDED. THE CAUSE OF DEATH IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22129353&pc=NAY | 2955842-2025-23422 | 2025-05-07 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-06-02 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED MALIGNANT HYSTERECTOMY SURGICAL PROCEDURE FOR ENDOMETRIAL CANCER, THE PATIENT CODED AND EXPIRED. THE SURGEON REPORTED THAT THE HYSTERECTOMY PORTION OF THE CASE WAS COMPLETED. DURING THE COLPOTOMY VAGINAL CUFF CLOSURE THE PATIENT CODED; THE OR TEAM EMERGENTLY REMOVED THE INSTRUMENTS, UNDOCKED THE SYSTEM AND PERFORMED CARDIOPULMONARY RESUSCITATION WITHOUT SUCCESS, AND DURING WHICH THERE WAS SOME UNSPECIFIED CHEST BLEEDING. THE SURGEON STATED HE BELIEVES THAT THE DA VINCI SYSTEM AND INSTRUMENTS HAD NOTHING TO DO WITH THE EVENT AS THERE WERE NO ISSUES REPORTED DURING THE HYSTERECTOMY PORTION; HE BELIEVES THAT THE PATIENT HAD AN UNDIAGNOSED CLOT, SUFFERED A PULMONARY EMBOLISM AND EXPIRED. NO ADDITIONAL DETAILS WERE PROVIDED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 0 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, CADIERE FORCEPS, FENESTRATED FORCE BIPOLAR. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT SUDDENLY CODED DURING A MALIGNANT HYSTERECTOMY PROCEDURE. NO ISSUES OR COMPLAINTS WERE REPORTED WITH ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS UP TO THIS POINT. THEY EMERGENTLY UNDOCKED AND ATTEMPTED CPR BUT THE PATIENT DIED DESPITE THESE EFFORTS. THE CAUSE OF DEATH WAS ATTRIBUTED TO A PULMONARY EMBOLISM. NO ALLEGATION WAS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22313352&pc=MHX | 9610816-2025-000514 | 2025-05-28 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-06-24 04:00 | MHX | INTELLIVUE MX550 PATIENT MONITOR | Defective Alarm | Insufficient Information | K131872 |  | 1 | Event Description: IT WAS REPORTED THE MONITOR DID NOT GENERATE ALARMS FOR APPROXIMATELY 25 MINUTES AND THE PATIENT PASSED AWAY. ENGINEERING REVIEWED THE CLINICAL AUDIT LOGS, REVEALING BOTH VISUAL AND AUDIBLE RED AND YELLOW ALARMS HAD OCCURRED AT BOTH THE MONITOR AND CENTRAL STATION. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22313410&pc=DTB | 2649622-2025-16951 |  | Death | MPRI | 2025-06-24 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Atrial Fibrillation; Cardiac Arrest; Unspecified Infection; Pain; Sepsis | P930039 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE PATIENT WAS HOSPITALIZED DUE TO SEPSIS FOR FORTY SEVEN DAYS AND WAS GRADUALLY ABLE TO GET UP AND WALK AND DID NOT HAVE THE PAIN ANYMORE. THE PATIENT THEN DEVELOPED AN INFECTION IN THE FOOT AND THERE WAS CONCERN THAT IT MAY DEVELOP INTO A BONE INFECTION. THE PATIENT TOOK A FALL ATTEMPTING TO CLIMB INTO BED AND WHACKED THE NIGHTSTAND ON THEIR WAY DOWN AND THE PATIENT REPRESENTATIVE EXPRESSED CONCERN THAT THE IMPLANTABLE PULSE GENERATOR (IPG) MAY HAVE BEEN DAMAGED BY THE VERY HARD FALL. RESCUE EFFORTS WERE MADE AND IT WAS SUSPECTED THAT THE PATIENT EXPERIENCED CARDIAC ARREST AS THE IPG WAS TREATING ATRIAL FIBRILLATION (AF). THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: W1DR01 IPG DOI: (B)(6) 2025 97714 PAIN STIM IPG DOI: (B)(6) 2018, 977A260 PAIN STIM LEAD DOI: (B)(6) 2018, 977A260 PAIN STIM LEAD DOI: (B)(6) 2018. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22313412&pc=DTB | 2649622-2025-16952 |  | Death | MPRI | 2025-06-24 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Appropriate Term/Code Not Available | Atrial Fibrillation; Cardiac Arrest; Unspecified Infection; Pain; Sepsis | P930039 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE PATIENT WAS HOSPITALIZED DUE TO SEPSIS FOR FORTY SEVEN DAYS AND WAS GRADUALLY ABLE TO GET UP AND WALK AND DID NOT HAVE THE PAIN ANYMORE. THE PATIENT THEN DEVELOPED AN INFECTION IN THE FOOT AND THERE WAS CONCERN THAT IT MAY DEVELOP INTO A BONE INFECTION. THE PATIENT TOOK A FALL ATTEMPTING TO CLIMB INTO BED AND WHACKED THE NIGHTSTAND ON THEIR WAY DOWN AND THE PATIENT REPRESENTATIVE EXPRESSED CONCERN THAT THE IMPLANTABLE PULSE GENERATOR (IPG) MAY HAVE BEEN DAMAGED BY THE VERY HARD FALL. RESCUE EFFORTS WERE MADE AND IT WAS SUSPECTED THAT THE PATIENT EXPERIENCED CARDIAC ARREST AS THE IPG WAS TREATING ATRIAL FIBRILLATION (AF). THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: W1DR01 IPG DOI: (B)(6) 2025; 97714 PAIN STIM IPG DOI: (B)(6) 2018, 977A260 PAIN STIM LEAD DOI: (B)(6) 2018, 977A260 PAIN STIM LEAD DOI: (B)(6) 2018. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22313415&pc=NVZ | 3008973940-2025-04463 |  | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-06-24 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Appropriate Term/Code Not Available | Atrial Fibrillation; Cardiac Arrest; Unspecified Infection; Pain; Sepsis | P980035 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT REPRESENTATIVE THAT THE PATIENT WAS HOSPITALIZED DUE TO SEPSIS FOR FORTY-SEVEN DAYS AND WAS GRADUALLY ABLE TO GET UP AND WALK AND DID NOT HAVE THE PAIN ANYMORE. THE PATIENT THEN DEVELOPED AN INFECTION IN THE FOOT AND THERE WAS CONCERN THAT IT MAY DEVELOP INTO A BONE INFECTION. THE PATIENT TOOK A FALL ATTEMPTING TO CLIMB INTO BED AND WHACKED THE NIGHTSTAND ON THEIR WAY DOWN AND THE PATIENT REPRESENTATIVE EXPRESSED CONCERN THAT THE IMPLANTABLE PULSE GENERATOR (IPG) MAY HAVE BEEN DAMAGED BY THE VERY HARD FALL. RESCUE EFFORTS WERE MADE AND IT WAS SUSPECTED THAT THAT THE PATIENT EXPERIENCED CARDIAC ARREST AS THE IPG WAS TREATING ATRIAL FIBRILLATION (AF). THE PATIENT PASSED AWAY. Manufacturer Narrative: CONTINUATION OF D10: 97714 PAIN STIM IPG DOI: (B)(6) 2018, 977A260 PAIN STIM LEAD DOI: (B)(6) 2018, 977A260 PAIN STIM LEAD DOI: (B)(6) 2018. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED: B1 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22313641&pc=LJT | 3006260740-2025-04211 | 2025-05-30 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-06-24 04:00 | LJT | UNKNOWN POWER PORT | Fracture; Material Separation; Device Dislodged or Dislocated; Migration | Insufficient Information; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | UNKNOWN |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME POST A PORT PLACEMENT, THE PORT FRACTURED, DISLODGED AND MIGRATED. IT WAS FURTHER REPORTED THAT PATIENT ALLEGEDLY DEVELOPED WITH LACERATION ON THE BLOOD VESSELS. REPORTEDLY, THE DEFECTIVE PORT WAS REMOVED. HOWEVER, THE PATIENT EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22317654&pc=OZD | 1220648-2025-29717 | 2025-06-16 04:00 | Death | ABIOMED, INC. | 2025-06-24 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT AN IMPELLA CP WAS INSERTED AFTER PERCUTANEOUS CORONARY INTERVENTION (PCI), BUT PERFORATION BY THE 0.014 INCH GUIDEWIRE HAD OCCURRED DURINGTHE PCI. THE PERFORATION WAS NEAR THE TIP OF LEFT ANTERIOR DESCENDING ARTERY. THE TREATMENT FOR THE PERFORATION WAS USING A MICRO CATHETER AND HEMOSTATIS PERFORMED WITH MICROTHROMBUS AND SPONGE. THE GUIDEWIRE WAS NOT DAMAGED. THE PERFORATION WAS THE MAIN CAUSE OF DETERIORATION IN HEMODYNAMICS, LEADING TO DEATH DUE TO MULTIPLE ORGAN FAILURE. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED STATING THAT DURING THE PCI PROCEDURE, A PERFORATION TO LEFT ANTERIOR DESCENDING ARTERY (LAD) WAS NOTED. THIS PERFORATION WAS CAUSED BY A PCI GUIDEWIRE(NON-ABIOMED), LEADING TO DEATH DUE TO MULTIPLE ORGAN FAILURE. THERE WAS NO ALLEGATION OF A INJURY OR MALFUNCTION ATTRIBUTED TO THE USE OF THE IMPELLA. THE DEATH IS UNRELATED. D10-CONCOMITANT DEVICE ADDED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED STATING THAT DURING THE PCI PROCEDURE, A PERFORATION TO LEFT ANTERIOR DESCENDING ARTERY (LAD) WAS NOTED. THIS PERFORATION WAS CAUSED BY A PCI GUIDEWIRE (NON-ABIOMED), LEADING TO DEATH DUE TO MULTIPLE ORGAN FAILURE. THERE WAS NO ALLEGATION OF A INJURY OR MALFUNCTION ATTRIBUTED TO THE USE OF THE IMPELLA. THE DEATH IS UNRELATED. THERE IS CURRENTLY NO ALLEGATION OF A MALFUNCTION OR SERIOUS INJURY ASSOCIATED WITH THE IMPLANTED PUMP. A REGULATORY REPORT IS NO LONGER NECESSARY. PREVIOUSLY SUBMITTED MANUFACTURER DEVICE REPORT 1220648-2025-29717 IS NO LONGER APPLICABLE. H1 CHANGED TO NA. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22317717&pc=OZD | 1220648-2025-29718 | 2025-06-19 04:00 | Death | ABIOMED, INC. | 2025-06-24 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Application Program Problem | Cardiovascular Insufficiency; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THE AUTOMATED IMPELLA CONTROLLER (AIC) HAD A "CONTROLLER ERROR ALARM (YELLOW)" ALARM SOUNDED TEMPORARILY AFTER USING THE IMPELLA AND WHILE RETURNING TO THE INTENSIVE CARE UNIT. THE POSITION WAVEFORM DISPLAY DISAPPEARED. EVEN AFTER TAKING MEASURES SUCH AS RE-INSERTING THE CONNECTION CABLE, THE ALARM CONTINUED TO APPEAR FREQUENTLY. THE AIC WAS REPLACED. THE PATIENT'S VITAL SIGNS REMAINED STABLE DURING REPLACMENT.Event Description: THE PATIENT EXPIRED DUE TO PREEXISTING CARDIAC CONDITION AND MULTIPLE ORGAN FAILURE. HOWEVER, IT IS UNCLEAR IF THE AUTOMATED IMPELLA CONTROLLER (AIC) ISSUE CONTRIBUTED TO THE CAUSE OF DEATH. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: DEVICE RETURNED ON JUL-14-2025 AND A DEVICE EVALUATION IS UNDERWAY. SHOULD ANY NEW/RELEVANT INFORMATION BECOME AVAILABLE, WE WILL UPDATE THE RECORD B1 ADVERSE EVENT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29718. B2 DEATH AND DATE OF DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29718. B5 PATIENT OUTCOME WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29718. H1 TYPE OF REPORT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29718. H6 HEALTH EFFECT - CLINICAL CODE 4582 WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29718. H6 HEALTH EFFECT - IMPACT CODE 1802 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29718.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22317869&pc=BZD | 2518422-2025-047669 | 2022-11-28 05:00 | Death | RESPIRONICS, INC. | 2025-06-24 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22317892&pc=OZD | 1220648-2025-29720 | 2025-05-31 04:00 | Death | ABIOMED, INC. | 2025-06-24 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT HAD ACCESS SITE BLEEDING THAT WAS MANAGED WITH TABOTAMP. HOWEVER, THE PATIENT EXPIRED WHILE ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22317990&pc=LDD | 3003832357-2025-000468 | 2025-06-18 04:00 | Death | REMOTE DIAGNOSTIC TECHNOLOGIES LTD. | 2025-06-24 04:00 | LDD | TEMPUS LS | Erratic or Intermittent Display | Cardiac Arrest; No Clinical Signs, Symptoms or Conditions | K200849 |  | 1 | Event Description: THIS REPORT IS BASED ON INFORMATION PROVIDED BY REMOTE SERVICE ENGINEER RSE AND IS CURRENTLY UNDER INVESTIGATION BY THE PHILIPS COMPLAINT HANDLING TEAM. PHILIPS RECEIVED A COMPLAINT ON THE TEMPUS LS MANUAL INDICATING THE SCREEN FAILED TO PROPERLY SHOW ECG (ELECTROCARDIOGRAM) RHYTHM WHILE ON A CALL. THE SCREEN WAS CLEANED AFTER RETURNING TO STATION, AND AS SOON AS PRESSURE WAS APPLIED, SCREEN BEGAN TO FUNCTION PROPERLY. THE PATIENT WAS IN CARDIAC ARREST AND BEING MONITORED THROUGH THE DEFIBRILLATOR PADS. THE ECG (ELECTROCARDIOGRAM) RHYTHM WAS SUCCESSFULLY DISPLAYED ON THE TEMPUS PRO, BUT NOT THE TEMPUS LS MANUAL, AS SUCH, DID NOT IMPACT THE CARE OR OUTCOME OF THIS PATIENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22318126&pc=OZD | 1220648-2025-29724 | 2025-06-05 04:00 | Death | ABIOMED, INC. | 2025-06-24 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure | Fever; Unspecified Infection | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT WAS ON IMPELLA 5.5 SUPPORT. DURING SUPPORT THE PATIENT BECAME FEBRILE AND THE SOURCE OF THE INFECTION REMAINS UNCLEAR. AT THE SAME TIME, HIGH PURGE PRESSURE WAS EXPERIENCED FOR WHICH TISSUE PLASMINOGEN ACTIVATOR WAS USED. LATER IT WAS REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF INFECTION/SEPSIS AND HIGH PURGE PRESSURE HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. INFECTION/SEPSIS: THE ROOT CAUSE OF THE INFECTION/SEPSIS WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. HIGH PURGE PRESSURE: DATA LOGS WERE REVIEWED AND A PURGE PRESSURE RISE OF APPROXIMATELY 7 HOURS WAS OBSERVED BEFORE "HIGH PURGE PRESSURE" ALARMS WERE TRIGGERED. NO MOTOR CURRENT SPIKES OBSERVED PRIOR TO PURGE PRESSURE RISE. PURGE PRESSURE REMAINED HIGH FOR APPROXIMATELY 7 HOURS BEFORE SLOWLY RESOLVING. CLINICAL DETAILS NOTED THAT PURGE PRESSURE HIGH ALARMS WERE TRIGGERED AND LYSIS THERAPY WAS ADDED TO PURGE AND PURGE FLOWS IMPROVED. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE WAS MOST LIKELY DUE TO BIOMATERIAL BUILD-UP BASED ON RETURNED DATA LOG ANALYSIS AND CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST. SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER. ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS. ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22318175&pc=NKE | 9614453-2025-02408 | 2024-07-14 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-06-24 04:00 | NKE | PERCEPTA QUAD CRT-P MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Sepsis; Drug Resistant Bacterial Infection | P010015 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS AND HEART FAILURE WITH REDUCED EJECTION FRACTION.Event Description: IT WAS FURTHER REPORTED THAT THE SOURCE OF SEPSIS WAS DUE TO A METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS INFECTION AT AN UNKNOWN LOCATION. IT WAS FURTHER REPORTED THAT THE PATIENT'S DEATH WAS NOT BELIEVED TO BE LINKED TO THE CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. THE DEVICE WAS RETURNED AND SCREENING ANALYSIS WAS PERFORMED, BUT NO ISSUE WAS IDENTIFIED REQUIRING FULL ANALYSIS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22318185&pc=DTB | 2649622-2025-17000 | 2024-07-14 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-06-24 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Sepsis; Drug Resistant Bacterial Infection | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS AND HEART FAILURE WITH REDUCED EJECTION FRACTION.Event Description: IT WAS FURTHER REPORTED THAT THE SOURCE OF SEPSIS WAS DUE TO A METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS INFECTION AT AN UNKNOWN L OCATION. IT WAS FURTHER REPORTED THAT THE THE PATIENT'S DEATH WAS NOT BELIEVED TO BE LINKED TO THE CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22318210&pc=DXZ | 22318210 | 2025-01-14 05:00 | Death | LIFENET HEALTH | 2025-06-24 04:00 | DXZ | CARDIOGRAFT 22MM AORTIC VALVE ALLOGRAFT CRYOPRESERVE ADULT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Fungal Infection; Asystole |  |  | 1 | Event Description: A PATIENT HAD CARDIAC SURGERY WITH THE USE OF A HARVESTED HOMOGRAFT FROM A TISSUE DONOR. THE PATIENT STAYED IN THE HOSPITAL FOR ABOUT 2 WEEKS POSTOPERATIVELY AND WAS THEN DISCHARGED TO AN ACUTE REHAB FACILITY. APPROXIMATELY 3 WEEKS LATER, THE PATIENT RETURNED TO THE HOSPITAL AND WAS DIAGNOSED WITH ENDOCARDITIS. DESPITE BEING TREATED WITH ANTIBIOTICS, THE PATIENT'S CONDITION CONTINUED TO DETERIORATE. THE PATIENT ULTIMATELY EXPERIENCED CARDIAC ARREST, AND RESUSCITATION ATTEMPTS WERE UNSUCCESSFUL. AN AUTOPSY REVEALED THE PRESENCE OF ASPERGILLUS FUMIGATUS. THE SOURCE OF THE INFECTION WAS THOROUGHLY INVESTIGATED AND INCLUDED THE CONSIDERATION THAT THE DONOR TISSUE MIGHT HAVE BEEN INFECTED BEFORE IMPLANTATION. HOWEVER, THE INVESTIGATION COULD NOT CONFIRM THE EXACT SOURCE OF THE INFECTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22318248&pc=OJX | 2649622-2025-17006 | 2024-07-14 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-06-24 04:00 | OJX | ATTAIN PERFORMA STRAIGHT MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Sepsis; Drug Resistant Bacterial Infection | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS AND HEART FAILURE WITH REDUCED EJECTION FRACTION.Event Description: IT WAS FURTHER REPORTED THAT THE SOURCE OF SEPSIS WAS DUE TO A METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS INFECTION AT AN UNKNOWN LOCATION. IT WAS FURTHER REPORTED THAT THE PATIENT'S DEATH WAS NOT BELIEVED TO BE LINKED TO THE CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22318258&pc=NVN | 2649622-2025-17007 | 2024-07-14 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-06-24 04:00 | NVN | SELECTSECURE MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Sepsis; Drug Resistant Bacterial Infection | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS AND HEART FAILURE WITH REDUCED EJECTION FRACTION.Event Description: IT WAS FURTHER REPORTED THAT THE SOURCE OF SEPSIS WAS DUE TO A METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS INFECTION AT AN UNKNOWN LOCATION. IT WAS FURTHER REPORTED THAT THE PATIENT'S DEATH WAS NOT BELIEVED TO BE LINKED TO THE CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P) SYSTEM. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22318526&pc=DYE | 2025587-2025-04437 | 2025-06-07 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-24 04:00 | DYE | MOSAIC MITRAL BIOPROSTHETIC HEART VALVE | Insufficient Information; Patient Device Interaction Problem | Septic Shock; Cardiogenic Shock; Insufficient Information | P990064 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT 9 DAYS POST IMPLANT OF THIS MOSAIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE MOSAIC VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE CAUSE OF DEATH WAS SEPTIC SHOCK AND CARDIOGENIC SHOCK. IT WAS STATED THAT THE PATIENT'S SEVERE RIGHT VENTRICULAR FAILURE CONTRIBUTED TO THE DEATH AND THE PHYSICIAN DOES NOT BELIEVE THAT MEDTRONIC PRODUCT CAUSED OR CONTRIBUTED TO THIS DEATH. Manufacturer Narrative: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT CHANGED THE EVENT DESCRIPTION AND DEVICE RELATEDNESS OF THIS PREVIOUSLY REPORTED EVENT. THERE IS NO EVIDENCE TO SUGGEST THE DEVICE CAUSED OR CONTRIBUTED TO A DEATH. A4, B5 AND H6 UPDATED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22318618&pc=OZD | 1220648-2025-29730 | 2025-06-07 04:00 | Death | ABIOMED, INC. | 2025-06-24 04:00 | OZD | IMPELLA CP | Positioning Failure | Hematoma; Pulmonary Edema; Tachycardia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT PRESENTED WITH CHEST PAINS AND ON ELECTROCARDIOGRAM SHOWED ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION. THE PATIENT WAS TAKEN TO THE CARDIAC CATHETERIZATION LABORATORY FOR EMERGENT PERCUTANEOUS CORONARY INTERVENTION (PCI). IT WAS REPORTED THAT SHORTLY AFTER ARRIVING TO CATHETERIZATION LABORATORY, THE PATIENT WENT INTO VENTRICULAR FIBRILLATION ARREST AND REQUIRED CARDIOPULMONARY RESUSCITATION (CPR) AND WAS DEFIBRILLATED FIVE TIMES. THE PATIENT WAS INTUBATED AND PCI WAS PERFORMED TO LEFT ANTERIOR DESCENDING ARTERY. RIGHT HEART CATHETERIZATION PERFORMED AND THE IMPELLA CP INSERTED TO RIGHT FEMORAL ARTERY. DURING SUPPORT, PULSES WERE CONFIRMED AND UROLOGY CONSULT ORDERED FOR FOLEY PLACEMENT. HEPARIN WAS HELD ON ARRIVAL DUE TO RIGHT GROIN HEMATOMA. THE PUMP WAS PULLED BACK UNDER ECHOCARDIOGRAM GUIDANCE. IT WAS NOTED THAT THEY WERE UNABLE TO GET THE FOLEY CATHETER IN PLACE DESPITE MULTIPLE ATTEMPTS. THE PATIENT HAD VISIBLE BLOOD AT THE URETHRAL MEATUS. HOWEVER, THE NEXT DAY, UROLOGY PLACED A FOLEY WITH SCOPE AT BEDSIDE. ALSO NOTED WAS THAT THE CRITICAL CARE TEAM REPOSITIONED THE DEVICE WITH ECHOCARDIOGRAM. ULTIMATELY, THE PATIENT DECOMPENSATED SEVERELY INTO FLASH PULMONARY EDEMA UPON INDUCTION FOR INTUBATION, EXPERIENCE TWO ROUNDS OF VENTRICULAR TACHYCARDIA REQUIRING DEFIBRILLATION AND CPR FOR NINE MINUTES. RETURN OF SPONTANEOUS CIRCULATION WAS OBTAINED. HOWEVER, THE PATIENT WAS NOT RESPONDING TO CHEMICAL OR MECHANICAL THERAPY. THE DECISION WAS MADE TO DISCONTINUE CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING, VF/VT/AF/AT, AND POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ACCESS SITE BLEEDING: CLINICAL STATES THE HEMATOMA TO IMPELLA GROIN SITE ON ARRIVAL TO HILLCREST MEDICAL CENTER AND PATIENT WAS COUGHING A LOT. THE CAUSE OF THE ACCESS SITE BLEEDING WAS MOST LIKELY PATIENT MOVEMENT RELATED AS THE HEMATOMA OCCURRED AFTER PATIENT TRANSFERRED HOSPITALS AND WAS COUGHING A LOT. POSITIONING: CLINICAL STATES THE PUMP WAS PULLED BACK UNDER ECHO GUIDANCE TO 3.9 CM ON DAY 1 OF SUPPORT AND AGAIN CRITICAL CARE REPOSITIONED THE DEVICE BEDSIDE WITH ECHO 2 DAYS LATER. THE CAUSE OF THE POSITIONING ISSUE WAS NOT DETERMINED DUE TO LACK OF CLINICAL DETAILS. VT: CLINICAL STATES THAT THE PT DECOMPENSATED SEVERELY INTO FLASH PULMONARY EDEMA UPON INDUCTION FOR INTUBATION, EXPERIENCE TWO ROUNDS OF VT REQUIRING DEFIBRILLATION, CPR FOR NINE MINUTES. IN ORDER TO MAKE A ROOT CAUSE DETERMINATION, ALL OF THE CLINICAL DETAILS OUTLINED IN (B)(4) WOULD HAVE TO BE PROVIDED. THE ROOT CAUSE OF THE VT/VF/AT/AF WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22319184&pc=DTE | MW5171785 | 2025-06-11 04:00 | Death | MEDTRONIC INC. | 2025-06-24 04:00 | DTE | PACEMAKER GENERATOR BOX | Adverse Event Without Identified Device or Use Problem | Aspiration/Inhalation; Ventricular Fibrillation |  |  | 1 | Event Description: A 55-YEAR-OLD MALE WHO WAS ALERT, EXTUBATED, AND WALKING ON HIS FIRST DAY FOLLOWING OPEN HEART SURGERY, AORTIC VALVE REPLACEMENT, ANEURYSM REPAIR EXPERIENCED R ON T WAVE EVENT, VFIB ARREST FOLLOWING PACEMAKER DISCHARGE. PATIENT WAS RESUSCITATED AND ASPIRATED DURING INTUBATION. HE DEVELOPED FURTHER COMPLICATIONS AND DIED ON HIS 6TH POST-OP DAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22319270&pc=LWP | 2017865-2025-89046 | 2025-05-24 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-24 04:00 | LWP | ASSURITY MRI | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P140033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED DUE TO HYPERTENSIVE AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE. IT WAS UNKNOWN IF THE PATIENT'S IMPLANTED SYSTEM CONTRIBUTED TO THEIR DEATH. THE PATIENT'S PACEMAKER, RIGHT ATRIAL (RA) LEAD, AND RIGHT VENTRICULAR (RV) LEAD WERE ALL SUCCESSFULLY EXPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22319274&pc=NVN | 2017865-2025-89047 | 2025-05-24 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-24 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED DUE TO HYPERTENSIVE AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE. IT WAS UNKNOWN IF THE PATIENT'S IMPLANTED SYSTEM CONTRIBUTED TO THEIR DEATH. THE PATIENT'S PACEMAKER, RIGHT ATRIAL (RA) LEAD, AND RIGHT VENTRICULAR (RV) LEAD WERE ALL SUCCESSFULLY EXPLANTED. Manufacturer Narrative: THE REPORTED EVENT WAS THAT THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD WAS RETURNED IN ONE PIECE FOR ANALYSIS. DURING ANALYSIS, A FAILURE EVENT WAS OBSERVED WHICH WAS UNRELATED TO THE REPORTED EVENT. FINAL ANALYSIS FOUND AN EXTERNAL INSULATION ABRASION CONSISTENT WITH FRICTION TO DEVICE CAN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22319292&pc=NVN | 2017865-2025-89048 | 2025-05-24 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-24 04:00 | NVN | TENDRIL STS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED DUE TO HYPERTENSIVE AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE. IT WAS UNKNOWN IF THE PATIENT'S IMPLANTED SYSTEM CONTRIBUTED TO THEIR DEATH. THE PATIENT'S PACEMAKER, RIGHT ATRIAL (RA) LEAD, AND RIGHT VENTRICULAR (RV) LEAD WERE ALL SUCCESSFULLY EXPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22319398&pc=NPT | 2015691-2025-05115 | 2025-02-09 05:00 | Death | EDWARDS LIFESCIENCES | 2025-06-24 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 |  | 1 | Event Description: AS REPORTED THROUGH THE JAPANESE TAVI REGISTRY, A 23MM SAPIEN 3 VALVE WAS TRANSFEMORALLY DEPLOYED IN THE AORTIC ANNULUS. ON POSTOPERATIVE DAY (POD) 3, THE PATIENT DIED DUE TO CEREBRAL INFARCTION. THE OCCURRENCE DATE OF CEREBRAL INFARCTION WAS UNKNOWN. THE DEVICE REMAINS IMPLANTED. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE EXACT CAUSE OF THE REPORTED EVENT COULD NOT BE DETERMINED, BUT INVESTIGATION RESULTS SUGGEST THAT FACTORS/MECHANISMS MENTIONED ABOVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22319989&pc=FRN | 2016493-2025-89616 | 2025-05-27 04:00 | Death | CAREFUSION SD | 2025-06-24 04:00 | FRN | ALARIS SYSTEM | Excess Flow or Over-Infusion | Insufficient Information | K133532 |  | 1 | Event Description: IT WAS INITIALLY REPORTED AN OVER INFUSION EVENT THAT OCCURRED ON (B)(6) 2025. PER REPORT, THE PATIENT "DIED." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE DEVICE LOGS TO IDENTIFY WHICH MEDICATIONS WERE INFUSED FROM 3:00 PM ON (B)(6) 2025 TO 1:30 AM ON (B)(6) 2025. NO FURTHER INFORMATION WAS PROVIDED. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "THE REASON TO DO THIS (REVIEW OF DEVICE LOGS) ONLY TO CHECK THE MEDICATIONS AND THE RATE OF IT GIVEN BY THE MACHINE. I¿M NOT BELIEVE THAT THE DEVICE CAUSED THE DEATH BUT MAYBE HUMAN MISTAKE." MULTIPOLE ATTEMPTS HAVE BEEN MADE REQUESTING FOR THE DETAILS OF THE EVENT. HOWEVER, NO FURTHER INFORMATION HAS BEEN PROVIDED TO DATE. Manufacturer Narrative: THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22320203&pc=FRN | 2016493-2025-89617 | 2025-05-27 04:00 | Death | CAREFUSION SD | 2025-06-24 04:00 | FRN | ALARIS SYSTEM | Excess Flow or Over-Infusion | Insufficient Information | K133532 |  | 1 | Event Description: IT WAS INITIALLY REPORTED AN OVER INFUSION EVENT THAT OCCURRED ON 27 MAY 2025. PER REPORT, THE PATIENT "DIED." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE DEVICE LOGS TO IDENTIFY WHICH MEDICATIONS WERE INFUSED FROM 3:00 PM ON (B)(6) 2025 TO 1:30 AM ON (B)(6) 2025. NO FURTHER INFORMATION WAS PROVIDED. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "THE REASON TO DO THIS (REVIEW OF DEVICE LOGS) ONLY TO CHECK THE MEDICATIONS AND THE RATE OF IT GIVEN BY THE MACHINE. I¿M NOT BELIEVE THAT THE DEVICE CAUSED THE DEATH BUT MAYBE HUMAN MISTAKE." MULTIPOLE ATTEMPTS HAVE BEEN MADE REQUESTING FOR THE DETAILS OF THE EVENT. HOWEVER, NO FURTHER INFORMATION HAS BEEN PROVIDED TO DATE. Manufacturer Narrative: THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22320206&pc=NPT | 2015691-2025-05126 | 2025-06-05 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-24 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Malposition of Device; Adverse Event Without Identified Device or Use Problem; Central Regurgitation | Rupture | P140031 |  | 1 | Event Description: AS REPORTED, IT WAS A CASE OF AN IMPLANT OF A 23MM SAPIEN 3 ULTRA VALVE, IN AORTIC POSITION VIA TRANSFEMORAL APPROACH. THE PATIENT PRESENTED WITH NATIVE AORTIC INSUFFICIENCY. DUE TO MULTIPLE COMORBIDITIES, INCLUDING PARKINSON'S DISEASE, THE PROCEDURE WAS CONDUCTED UNDER GENERAL ANESTHESIA. THE VALVE WAS DEPLOYED WITH AN ADDITIONAL +2CC VOLUME TO ENHANCE ANCHORING; HOWEVER, THE FINAL POSITION WAS NOTED TO BE SLIGHTLY LOW. IMMEDIATELY FOLLOWING DEPLOYMENT, THE PATIENT EXPERIENCED SUDDEN HEMODYNAMIC COLLAPSE. AORTOGRAPHY REVEALED MASSIVE CENTRAL AORTIC REGURGITATION. WHILE A SECOND VALVE WAS BEING PREPARED FOR A POSSIBLE VALVE-IN-VALVE IMPLANTATION, AN ECHOCARDIOGRAPHY WAS PERFORMED, SHOWING A LARGE PERICARDIAL EFFUSION. EMERGENT PERICARDIAL DRAINAGE YIELDED OVER 500 ML OF BLOOD, BUT BLEEDING PERSISTED. THE ECHOCARDIOGRAPHIC FINDINGS AND CLINICAL PRESENTATION WERE HIGHLY SUGGESTIVE OF ACUTE AORTIC DISSECTION SECONDARY TO ANNULAR RUPTURE. GIVEN THE PATIENT'S HEMODYNAMIC INSTABILITY, ONGOING HEMORRHAGE, AND POOR SURGICAL CANDIDACY, FURTHER INTERVENTION WAS DEEMED UNFEASIBLE. UNFORTUNATELY, THE PATIENT PASSED AWAY. POST-EVENT IMAGING CONFIRMED ANNULAR RUPTURE LEADING TO AORTIC DISSECTION. ACCORDING TO MEDICAL OPINION, THE FRAGILITY OF THE AORTIC TISSUE AND ANATOMICAL FACTORS LIKELY CONTRIBUTED TO THE ADVERSE OUTCOME. Manufacturer Narrative: THE INVESTIGATION IS ONGOING. PER THE INSTRUCTIONS FOR USE (IFU), VALVE MALPOSITION REQUIRING INTERVENTION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO MALPOSITION, INCLUDING IMPROPER POSITIONING BEFORE DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, PRESERVED EJECTION FRACTION, SIGNIFICANT LANDING ZONE CALCIFICATION, LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, THE RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE IFU CAUTIONS THAT INCORRECT SIZING OF THE VALVE MAY LEAD TO PARAVALVULAR LEAK, MIGRATION, OR EMBOLIZATION. THE DEVICE TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR IN THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR MALPOSITION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BALLOON VALVULOPLASTY MAY INDICATE POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. AORTIC DISSECTION OR RUPTURE MAY OCCUR WHEN MULTIPLE ATTEMPTS ARE MADE TO CROSS THE STENOTIC NATIVE VALVE, AND/OR WHEN EXCESSIVE FORCE IS USED. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). THE THV TRAINING MANUALS GUIDE TO FACILITATE SAFE CROSSING OF THE NATIVE OR FAILING VALVE, INCLUDING CAMERA PROJECTIONS, HANDLING DURING ADVANCEMENT, AND TROUBLESHOOTING TECHNIQUES IF DIFFICULTY IS ENCOUNTERED. AS STATED, EXCESSIVE FORCE SHOULD NOT BE USED WHEN THE DEVICE HAS DIFFICULTY CROSSING THE STENOTIC VALVE. ADDING TENSION TO THE WIRE, PULLING BACK THE SYSTEM TO RE-ORIENT THE VALVE, AS NEEDED, AND TORQUING THE FLEX CATHETER MAY HELP SOLVE THE PROBLEM. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THE FRAGILITY OF THE AORTIC TISSUE AND ANATOMICAL FACTORS LIKELY CONTRIBUTED TO THE ADVERSE OUTCOME. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR CORRECTION AND ADDITIONAL INFORMATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: CORRECTED H.6 INVESTIGATION CONCLUSIONS AND INVESTIGATION FINDINGS. ADDED NEW INFORMATION TO H.6 TYPE OF INVESTIGATION. THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE IFU, CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, CENTRAL REGURGITATION EVENTS POST-DEPLOYMENT WITH OR WITHOUT REPORT RELATED TO LEAFLET MOBILITY FOR THE S3 AND S3U VALVES HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO PROCEDURAL FACTORS (MALPOSITION, SLOW RECOVERY OF BLOOD FLOW, LEAFLET IMPINGEMENT DUE TO CALCIFICATION OR GUIDEWIRE, OVER-INFLATION, POST-DILATION, AND/OR UNDER-EXPANSION). THE COMPLAINT WAS CONFIRMED BASED ON EVALUATION OF THE PROVIDED IMAGERY. AVAILABLE INFORMATION SUGGESTS THAT PROCEDURAL FACTORS (MALPOSITION, THV, OVER-INFLATION) LIKELY CONTRIBUTED TO THE EVENT AS PROVIDED INFORMATION INDICATED THAT THE VALVE WAS IMPLANTED WITH AN ADDITIONAL +2CC OF VOLUME, AND A LOW FINAL VALVE POSITION. VALVE MALPOSITIONING DURING DEPLOYMENT MAY RESULT IN THE POSSIBILITY OF THE NATIVE LEAFLETS HANGING OVER AND COVERING THE OUTFLOW OF THE VALVE, RESULTING IN REDUCED COAPTATION OF THE LEAFLETS. THE LEAFLETS ARE IN A NORMALLY OPEN POSITION AND REQUIRE THE BACK FLOW/PRESSURE GENERATED TO INITIATE LEAFLET CLOSURE. THERE ARE MULTIPLE PATIENT AND/OR PROCEDURAL FACTORS THAT CONTRIBUTE TO MALPOSITION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, NARROW OR CALCIFIED STJ, MINIMALLY OR BULKY/SEVERELY CALCIFIED LEAFLETS, PRESERVED EJECTION FRACTION, SIGNIFICANT MITRAL ANNULAR CALCIFICATION (MAC), LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. ADDITIONALLY, DEPLOYING THE VALVE USING MORE THAN THE PRESCRIBED VOLUME OR ATTEMPTING A SECOND INFLATION MAY RESULT IN INABILITY FOR THE VALVE LEAFLETS TO PROPERLY FUNCTION. THE INCREASE IN VALVE DIAMETER DUE TO OVERINFLATION OR POST DILATION MAY PREVENT PROPER MOTION OF THE THV LEAFLETS RESULTING IN THV REGURGITATION. ADDITIONALLY, PER THE TRAINING MANUAL, CENTRAL REGURGITATION IS AN ASSOCIATED RISK OF POST-DILATION. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: THE DEVICE WAS USED IN OFF-LABEL IMPLANTATION IN A PURE AORTIC REGURGITATION NATIVE VALVE. AS THERE ARE NO SPECIFIC IFU OR TRAINING MATERIALS RELATED TO PURE AORTIC REGURGITATION PROCEDURES, THE AVAILABLE TRAINING MATERIALS WERE REVIEWED ONLY FOR INFORMATION POTENTIALLY RELEVANT TO THE DEVICE USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22320354&pc=OZD | 1220648-2025-29737 | 2025-06-07 04:00 | Death | ABIOMED, INC. | 2025-06-24 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD A IMPELLA CP PUMP PLACED FOR THE PATIENT ADMITTED IN WITH ACUTE MYOCARDIAL INFARCTION AND CARDIOGENIC SHOCK. THE PUMP WAS PLACED BUT THE PATIENT WAS EXHIBITING SIGNS OF HEMOLYSIS WITHIN FIRST HOURS OF SUPPORT. THE TEAM CALLED FOR AIR TRANSPORT TO SEND PATIENT TO TERTIARY CENTER BUT, THE PATIENT EXPIRED WHILE CODING PRIOR TO AIR TRANSPORT. THE FAMILY STOPPED THE CODE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22320457&pc=QBJ | 3004753838-2025-161638 | 2025-06-18 04:00 | Death | DEXCOM, INC. | 2025-06-24 04:00 | QBJ | DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | Appropriate Term/Code Not Available | Insufficient Information | K213919 |  | 1 | Event Description: IT WAS REPORTED VIA WEB SELF-SERVICE THAT A PATIENT PASSED AWAY. THE REPORTER STATED, "MY BOYFRIEND DIED BECAUSE HIS DEXCOM WAS HOOKED TO A PUMP.". NO OTHER EVENT DETAILS WERE PROVIDED AND FOLLOW UP ATTEMPTS TO THE REPORTER FOR ADDITIONAL EVENT DETAILS WERE UNSUCCESSFUL. NO PRODUCT OR DATA WAS PROVIDED FOR INVESTIGATION. THE ALLEGATION AND PROBABLE CAUSE COULD NOT BE DETERMINED. NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4) H6 MEDICAL DEVICE PROBLEM CODE - 3191 - PATIENT WAS UNABLE TO IDENTIFY DEVICE ISSUE THEREFORE A MORE SPECIFIC CODE COULD NOT BE SELECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22320500&pc=OZD | 1220648-2025-29739 | 2025-06-05 04:00 | Death | ABIOMED, INC. | 2025-06-24 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT (MCS) AND EXPERIENCED BLEEDING AT THE ACCESS SITE, ATRIAL FIBRILLATION AND SUBSEQUENTLY EXPIRED. THE PATIENT WITH LEFT VENTRICULAR EJECTION FRACTION OF APPROXIMATELY 25% AND AORTIC VALVE AREA OF 0.4 CENTIMETERS SQUARED WAS ADMITTED TO THE HOSPITAL AND EXPERIENCED CARDIOGENIC SHOCK ALONGSIDE HYPOTENSION AND RESPIRATORY FAILURE. TWO DAYS LATER, THE PATIENT UNDERWENT AORTIC VALVULOPLASTY AND IMPELLA CP INSERTION. THE FOLLOWING DAY, BLEEDING AT THE IMPELLA INSERTION SITE WAS NOTED. THE PATIENT RECEIVED ONE UNIT OF RED BLOOD CELLS DUE TO A DROP IN HEMOGLOBIN. THE DRESSING WAS NOTED TO HAVE BEEN CHANGED SEVERAL TIMES AND HEMOSTATIC DRESSING WAS APPLIED. THE NEXT DAY, IT WAS NOTED THAT A DECISION WAS MADE BY THE PATIENT'S SON TO NOT ESCALATE CARE AND WITHDRAW CARE IF NO IMPROVEMENT WERE MADE. THE PUMP WAS REPOSITIONED WITH HOPES OF IMPROVEMENT. SUPPORT WAS CONTINUED AND THE PATIENT REMAINED CRITICALLY ILL. THE PATIENT'S BLOOD PRESSURE WAS EXTREMELY LABILE, REMAINED ON HIGH DOSES OF PRESSORS AND PERFORMANCE LEVEL P-9. THE PLAN WAS TO STAY COURSE UNTIL THE SON MADE A FINAL DECISION. THE PATIENT WAS TAKEN OFF LEVOPHED OVER THE WEEKEND WITH IMPROVEMENT BUT WENT INTO ATRIAL FIBRILLATION WITH RAPID VENTRICULAR RESPONSE THE FOLLOWING MORNING AND NEEDED LEVOPHED SUPPORT. THE PHYSICIAN NOTED IF SUPPORT WAS CONTINUED THE PATIENT WOULD NEED CARDIOVERTED AND BE PLACED ON CONTINUOUS RENAL REPLACEMENT THERAPY. HOWEVER, THIS DAY THE PATIENT WOULD EXPIRE. CARE WAS WITHDRAWN. Manufacturer Narrative: ALTHOUGH THE PATIENT'S FAMILY MEMBER DECLINED TO CONTINUE WITH ESCALATION IN CARE AND SUBSEQUENTLY WITHDREW CARE, THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA DEVICE USED AS AN ASSOCIATED FACTOR TO THE OVERALL OUTCOME (GIVEN THE IMPELLA RELATED COMPLICATIONS). THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22320517&pc=NPT | 2025587-2025-04446 | 2025-05-29 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-24 04:00 | NPT | EVOLUT FX PLUS VALVE | Patient Device Interaction Problem | Obstruction/Occlusion; Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE (TAVR), AFTER THE VALVE WAS IMPLANTED, CORONARY ACCESS THROUGH THE VALVE AND CORONARY ARTERY BYPASS GRAFTING (CABG) WAS ATTEMPTED; HOWEVER, THE PATIENT DIED DUE TO A CORONARY OCCLUSION FROM A POTENTIAL AORTIC DISSECTION. SUBSEQUENTLY, THE VALVE WAS EXPLANTED.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE (TAVR), AFTER THE VALVE WAS IMPLANTED, CORONARY ACCESS THROUGH THE VALVE AND CORONARY ARTERY BYPASS GRAFTING (CABG) WAS ATTEMPTED; HOWEVER, THE PATIENT DIED DUE TO A CORONARY OCCLUSION FROM A POTENTIAL AORTIC DISSECTION. SUBSEQUENTLY, THE VALVE WAS EXPLANTED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT, PER THE PHYSICIAN, THE DELIVERY CATHETER SYSTEM WAS NOT A CONTRIBUTING FACTOR TO THE DEATH. UPON OPENING THE CHEST, THE CARDIOVASCULAR SURGEON CONFIRMED AN AORTIC DISSECTION.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE (TAVR), AFTER THE VALVE WAS IMPLANTED, CORONARY ACCESS THROUGH THE VALVE AND CORONARY ARTERY BYPASS GRAFTING (CABG) WAS ATTEMPTED; HOWEVER, THE PATIENT DIED DUE TO A CORONARY OCCLUSION FROM A POTENTIAL AORTIC DISSECTION. SUBSEQUENTLY, THE VALVE WAS EXPLANTED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT, PER THE PHYSICIAN, THE DELIVERY CATHETER SYSTEM WAS NOT A CONTRIBUTING FACTOR TO THE DEATH. UPON OPENING THE CHEST, THE CARDIOVASCULAR SURGEON CONFIRMED AN AORTIC DISSECTION. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE CAUSE OF THE INJURY IS UNKNOWN. PER THE PHYSICIAN, IT CANNOT BE CONFIRMED WHETHER THE NON-MEDTRONIC GUIDEWIRE CONTRIBUTED TO THE DISSECTION. THE DELIVERY CATHETER SYSTEM CANNOT BE EXCLUDED AS A CONTRIBUTING FACTOR. Manufacturer Narrative: CONTINUATION OF D10: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME EVOLUT FX DCS; PRODUCT ID D-EVOLUTFX-2329 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE UPDATED DATA: B5. ADDED THIRD PARAGRAPH. H6. H11. SYSTEM REPORT DCS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: {{D-EVOULTFX-2329}}; PRODUCT LOT/SERIAL NUMBER {{UNKNOWN}}; PRODUCT TYPE: {{0195 HEART VALVES}}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. ADDED SECOND PARAGRAPH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22321061&pc=KRD | 2124215-2025-41604 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-24 04:00 | KRD | DC BEAD | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED VIA LITERATURE THAT DEATH OCCURRED. PURPOSE: ABLATIVE THERAPIES FOR PRIMARY AND SECONDARY LIVER MALIGNANCIES ARE INCREASINGLY ADOPTED IN CURRENT GUIDELINES. NEVERTHELESS, SURGICAL RESECTION REMAINS THE GOLD STANDARD IN MOST CURATIVE THERAPY SETTINGS. HOWEVER, THERE ARE INDICATIONS FOR ABLATIVE TREATMENTS SUCH AS THERMAL ABLATION, HIGH-CONFORMAL RADIATION, AND LOCO-REGIONAL THERAPIES. WE INVESTIGATED COMPLICATIONS AND MORTALITY AFTER ABLATIVE TREATMENT IN A LARGE, UNSELECTED STUDY COHORT. THIS STUDY EVALUATES TREATMENT DATA FROM THE GERMAN DRG-BASED HOSPITAL REIMBURSEMENT SYSTEM OF ABLATIVE PROCEDURES OF PRIMARY AND SECONDARY LIVER MALIGNANCIES PERFORMED BETWEEN 2005 AND 2020. OVER 15 YEARS, THERE WERE 4374 ABLATIVE PROCEDURES OF THE LIVER PERFORMED. PATIENTS UNDERWENT 3085 HIGH-DOSE-RATE BRACHYTHERAPY (HDR- BT) (70.5%), 296 THERMAL ABLATIONS (6.8%), 111 TRANSARTERIAL CHEMOEMBOLIZATION (TACE)(2.5%), AND 882 SELECTIVE INTERNAL RADIATION THERAPY (SIRT) (20.2%). EVALUATION OF THE TREATMENT DATA FOCUSED ON PRE-EXISTING MEDICAL CONDITIONS, LENGTH OF HOSPITALIZATION, MAJOR COMPLICATIONS DURING THE HOSPITALIZATION, AND IN-HOSPITAL DEATH. PROCEDURE SUMMARY - TRANSARTERIAL CHEMOEMBOLIZATION (TACE): WE INSERTED A 5F DIAGNOSTIC ANGIOGRAPHIC CATHETER USING TRANSFEMORAL ACCESS AND A COAXIAL NON-BOSTON SCIENTIFIC 2.0F MICROCATHETER TO DISPLAY TUMOR-SUPPLYING ARTERIES SELECTIVELY. WE INJECTED DC BEADS M1 70 -150 MICROMETRE PARTICLES, WHICH WERE LOADED WITH 50 -75 MG OF DOXORUBICIN (DEBDOX). NONLOADED MICROPARTICLES OF DIFFERENT SIZES FROM BEAD BLOCK (100-300 MICRONS AND 300-500 MICRONS) WERE USED TO REACH FLOW STAGNATION IF THE INJECTION OF DEBDOX WAS INSUFFICIENT. EVENT SUMMARY: THERE WERE 5 IN-HOSPITAL DEATHS IN TREATMENT CASES WITH TACE. Manufacturer Narrative: MARCH, CHRISTINE, ET AL. "MORTALITY AND POSTINTERVENTIONAL COMPLICATIONS AFTER ABLATIVE TREATMENT OF LIVER MALIGNANCIES: A COHORT STUDY OF 4374 PATIENTS." BRACHYTHERAPY 23.6 (2024): 743-750. B3: DATE OF EVENT: USED FIRST DAY IN THE FIRST MONTH OF THE YEAR OF THE ARTICLE. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC AS THE EVENT WAS REPORTED THROUGH LITERATURE. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22321064&pc=KRD | 2124215-2025-41603 | 2024-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-24 04:00 | KRD | BEAD BLOCK | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K150876 |  | 1 | Event Description: IT WAS REPORTED VIA LITERATURE THAT DEATH OCCURRED. PURPOSE: ABLATIVE THERAPIES FOR PRIMARY AND SECONDARY LIVER MALIGNANCIES ARE INCREASINGLY ADOPTED IN CURRENT GUIDELINES. NEVERTHELESS, SURGICAL RESECTION REMAINS THE GOLD STANDARD IN MOST CURATIVE THERAPY SETTINGS. HOWEVER, THERE ARE INDICATIONS FOR ABLATIVE TREATMENTS SUCH AS THERMAL ABLATION, HIGH-CONFORMAL RADIATION, AND LOCO-REGIONAL THERAPIES. WE INVESTIGATED COMPLICATIONS AND MORTALITY AFTER ABLATIVE TREATMENT IN A LARGE, UNSELECTED STUDY COHORT. THIS STUDY EVALUATES TREATMENT DATA FROM THE GERMAN DRG-BASED HOSPITAL REIMBURSEMENT SYSTEM OF ABLATIVE PROCEDURES OF PRIMARY AND SECONDARY LIVER MALIGNANCIES PERFORMED BETWEEN 2005 AND 2020. OVER 15 YEARS, THERE WERE 4374 ABLATIVE PROCEDURES OF THE LIVER PERFORMED. PATIENTS UNDERWENT 3085 HIGH-DOSE-RATE BRACHYTHERAPY (HDR- BT) (70.5%), 296 THERMAL ABLATIONS (6.8%), 111 TRANSARTERIAL CHEMOEMBOLIZATION (TACE)(2.5%), AND 882 SELECTIVE INTERNAL RADIATION THERAPY (SIRT) (20.2%). EVALUATION OF THE TREATMENT DATA FOCUSED ON PRE-EXISTING MEDICAL CONDITIONS, LENGTH OF HOSPITALIZATION, MAJOR COMPLICATIONS DURING THE HOSPITALIZATION, AND IN-HOSPITAL DEATH. PROCEDURE SUMMARY - TRANSARTERIAL CHEMOEMBOLIZATION (TACE): WE INSERTED A 5F DIAGNOSTIC ANGIOGRAPHIC CATHETER USING TRANSFEMORAL ACCESS AND A COAXIAL NON-BOSTON SCIENTIFIC 2.0F MICROCATHETER TO DISPLAY TUMOR-SUPPLYING ARTERIES SELECTIVELY. WE INJECTED DC BEADS M1 70 -150 MICROMETRE PARTICLES, WHICH WERE LOADED WITH 50 -75 MG OF DOXORUBICIN (DEBDOX). NONLOADED MICROPARTICLES OF DIFFERENT SIZES FROM BEAD BLOCK (100-300 MICRONS AND 300-500 MICRONS) WERE USED TO REACH FLOW STAGNATION IF THE INJECTION OF DEBDOX WAS INSUFFICIENT. EVENT SUMMARY THERE WERE 5 IN-HOSPITAL DEATHS IN TREATMENT CASES WITH TACE. Manufacturer Narrative: MARCH, CHRISTINE, ET AL. "MORTALITY AND POSTINTERVENTIONAL COMPLICATIONS AFTER ABLATIVE TREATMENT OF LIVER MALIGNANCIES: A COHORT STUDY OF 4374 PATIENTS." BRACHYTHERAPY 23.6 (2024): 743-750. B3: DATE OF EVENT: USED FIRST DAY IN THE FIRST MONTH OF THE YEAR OF THE ARTICLE. D2B: PRO CODE (PRODUCT CODE): HCG. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC AS THE EVENT WAS REPORTED THROUGH LITERATURE. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22321262&pc=OZD | 1220648-2025-29748 | 2025-06-08 04:00 | Death | ABIOMED, INC. | 2025-06-24 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT HAD VENTRICULAR FIBRILLATION, AND CODED. CHEST COMPRESSIONS WERE PERFORMED AND RESUSCITATION WAS ACHIEVED AND SUPPORT CONTINUED. THE PATIENT EVENTUALLY EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM AND IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22321275&pc=DSQ | 2916596-2025-03757 | 2025-05-27 04:00 | Death | THORATEC CORPORATION | 2025-06-24 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WILL NOT BE EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: G2 - REPORT SOURCE: CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. DUE TO PATIENT PRIVACY LAWS, THE MANAGING HOSPITAL WOULD NOT COMMUNICATE ANY ADDITIONAL INFORMATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22321317&pc=OZD | 1220648-2025-29749 | 2025-06-08 04:00 | Death | ABIOMED, INC. | 2025-06-24 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND RP FLEX DEVICES FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE TEAM ESCALATED TO A VA ECMO FROM THE RP. THE CP SUPPORTED DESPITE BLEEDING FROM ALL ACCESS SITES AND CONCERNS OF LIMB ISCHEMIA FOR 12 HOURS BEFORE THE FAMILY MADE THE CHOICE TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ACCESS SITE BLEEDING : BLEEDING FROM ALL ACCESS SITES. THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED SINCE THERE WAS BLEEDING FROM ALL ACCESS SITES. LIMB ISCHEMIA : THE ROOT CAUSE OF THE LIMB ISCHEMIA WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. H6 HEALTH EFFECT - IMPACT CODE 1802 AND COMPONENT CODE 3038 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29749. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMART ASSIST SYSTEM INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER. ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22321909&pc=NVZ | 2017865-2025-87899 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-24 04:00 | NVZ | ZEPHYR XL DR | Use of Device Problem; Adverse Event Without Identified Device or Use Problem | Insufficient Information | P880086 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FOR AN UNKNOWN CAUSE. UPON DEVICE INTERROGATION, THE PACEMAKER WAS FOUND TO BE ON BACKUP MODE. THERE WERE CONCERNS IF THE DEVICE COULD HAD STOPPED WORKING. SINCE ELECTROCARDIOGRAMS WERE NOT ENABLED, A RELATIONSHIP BETWEEN THE DEVICE AND THE PATIENT'S DEATH COULD NOT BE ESTABLISHED. THE CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: CORRECTION FOR H6 CODING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22322008&pc=LWR | 2135147-2025-03392 | 2001-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-06-24 04:00 | LWR | EPIC¿ VALVE (AORTIC) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P040021 |  | 1 | Event Description: THE ARTICLE, "ANTICOAGULATION THERAPY AND RELATED OUTCOMES AMONG ASIAN PATIENTS AFTER BIOPROSTHETIC VALVE REPLACEMENT", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, MULTICENTER STUDY TO IDENTIFY THE OPTIMAL INTERNATIONAL NORMALIZED RATIO (INR) RANGE FOR ASIAN PATIENTS DURING A 1-YEAR FOLLOW-UP AFTER TISSUE VALVE REPLACEMENT. DEVICES INCLUDED IN THE STUDY WERE HANCOCK II, ST. JUDE EPIC, TRIFECTA, MOSAIC, MAGNA EASE/PERIMOUNT, AND UNKNOWN. THE ARTICLE CONCLUDED THAT FOR PATIENTS WHO UNDERWENT TISSUE VALVE REPLACEMENT, THE BLEEDING RISK WAS ELEVATED WHEN THE INR WAS GREATER THAN 2.0, BUT THE RISK OF THROMBOEMBOLIC EVENT INCREASED ONLY WHEN THE INR WAS LOWER THAN 1.84 IN THE ATRIAL FIBRILLATION (AF) GROUP AND 1.7 IN THE NON-AF GROUP, REGARDLESS OF WHETHER THE PATIENT RECEIVED AORTIC VALVE REPLACEMENT (AVR), MITRAL VALVE REPLACEMENT (MVR), OR DOUBLE VALVE REPLACEMENT (DVR). [THE PRIMARY AUTHOR WAS TANG-YU LIU, AMERICAN COLLEGE OF CARDIOLOGY, WASHINGTON, D.C., U.S.A. THE CORRESPONDING AUTHOR WAS SHAO-WEI CHEN, DIVISION OF THORACIC AND CARDIOVASCULAR SURGERY, DEPARTMENT OF SURGERY, CHANG GUNG MEMORIAL HOSPITAL, LINKOU MEDICAL CENTER, CHANG GUNG UNIVERSITY, NO. 5 FUXING STREET, GUISHAN DISTRICT, TAOYUAN CITY 33305, TAIWAN, WITH CORRESPONDING EMAIL: JOSEPHCHEN0314@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS FROM 01 JANUARY 2001 TO 31 DECEMBER 2018. A TOTAL OF 1059 PATIENTS WERE INCLUDED IN THE STUDY. OF THE 1192 TOTAL DEVICES INVOLVED, 640 (53.7%) WERE ABBOTT. THE AVERAGE AGE WAS 65.5 YEARS, THE AVERAGE WEIGHT WAS 58.9KG, AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED CHRONIC OBSTRUCTIVE PULMONARY DISEASE, CHRONIC LIVER DISEASE, CHRONIC KIDNEY DISEASE, HYPERTENSION, HYPERLIPIDEMIA, DIABETES, INFECTIVE ENDOCARDITIS, RHEUMATIC HEART DISEASE, LUNG EDEMA, GOUT, PERIPHERAL ARTERY DISEASE, MALIGNANCY, PRIOR ISCHEMIC STROKE, PRIOR MYOCARDIAL INFARCTION, HISTORY OF BLEEDING (GASTROINTESTINAL, INTRACRANIAL, MAJOR), AND HEART FAILURE HOSPITALIZATION.Event Description: N/A. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: ANTICOAGULATION THERAPY AND RELATED OUTCOMES AMONG ASIAN PATIENTS AFTER BIOPROSTHETIC VALVE REPLACEMENT. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF ANTICOAGULATION THERAPY AND RELATED OUTCOMES AMONG ASIAN PATIENTS AFTER BIOPROSTHETIC VALVE REPLACEMENT WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING CHRONIC OBSTRUCTIVE PULMONARY DISEASE, CHRONIC LIVER DISEASE, CHRONIC KIDNEY DISEASE, HYPERTENSION, HYPERLIPIDEMIA, DIABETES, INFECTIVE ENDOCARDITIS, RHEUMATIC HEART DISEASE, LUNG EDEMA, GOUT, PERIPHERAL ARTERY DISEASE, MALIGNANCY, PRIOR ISCHEMIC STROKE, PRIOR MYOCARDIAL INFARCTION, HISTORY OF BLEEDING (GASTROINTESTINAL, INTRACRANIAL, MAJOR), AND HEART FAILURE HOSPITALIZATION. SOME OF THE COMPLICATIONS REPORTED WERE STROKE, THROMBUS, HEMORRHAGE AND DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: ANTICOAGULATION THERAPY AND RELATED OUTCOMES AMONG ASIAN PATIENTS AFTER BIOPROSTHETIC VALVE REPLACEMENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22322016&pc=DSQ | 2916596-2025-03996 | 2025-06-04 04:00 | Death | THORATEC CORPORATION | 2025-06-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Nausea; Pneumonia; Vomiting; Viral Infection; Weight Changes; Heart Failure/Congestive Heart Failure; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED ON (B)(6) 2025 FOR NAUSEA AND VOMITING, AND RECURRENT PNEUMONIA (PNA) WITH PARAINFLUENZA VIRUS. THE PATIENT WAS ALSO NOTED TO HAVE EXPERIENCED WEIGHT GAIN AND DECOMPENSATION. THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO PNA AND RIGHT HEART FAILURE (RHF). THE PATIENT HAD MODERATE TO SEVERE RIGHT HEART FAILURE PRIOR TO LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT. THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO RIGHT HEART FAILURE. PNA WAS NOT CAUSED BY RHF AND RHF DID NOT LEAD TO PNA. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: CORRECTED DATA: SECTION H6: HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: IT WAS ORIGINALLY REPORTED THAT THE PATIENT EXPIRED DUE TO PNEUMONIA AND RIGHT HEART FAILURE. HOWEVER, FURTHER INFORMATION COMMUNICATED BY THE ACCOUNT REVEALED THAT THE RIGHT HEART FAILURE EXISTED PRIOR TO LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANT, AND THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE RIGHT HEART FAILURE. THE PATIENT'S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED, AND THE DEVICE OPERATED AS EXPECTED. IT WAS COMMUNICATED THAT THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22322017&pc=PNJ | 2017865-2025-87845 | 2025-06-12 04:00 | Death | ABBOTT | 2025-06-24 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Material Too Rigid or Stiff | Arrhythmia; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P150035 |  | 1 | Event Description: DURING AN IMPLANT PROCEDURE, CONTRAST WAS OBSERVED TO BE IN THE PERICARDIAL SPACE AND A CARDIAC PERFORATION OF THE ANTERIOR GROOVE RESULTING IN TAMPONADE, BILATERAL FIXATED MYDRIASIS, AND CARDIAC ARREST WERE OBSERVED DUE TO THE DELIVERY CATHETER. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. FOLLOWING CPR, THE HEART WAS SUTURED AND THE CATHETER WAS REMOVED FROM THE PATIENT. THE EVENT OCCURRED PRIOR TO FIXATION AND THE CAUSE OF THE EVENT WAS ALLEGED TO BE DUE TO THE CATHETER BEING TOO STIFF. THE LEADLESS PACEMAKER WAS NOT IMPLANTED. THE PATIENT WAS HOSPITALIZED AND WILL CONTINUE TO BE MONITORED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT EXPERIENCED CEREBRAL HYPOXIA FOLLOWED BY PATHOLOGICAL AWAKENING UPON CESSATION OF SEDATION, CHARACTERIZED BY SEIZURES. PERSISTENT HYPERLACTATEMIA WAS OBSERVED DESPITE AMINERGIC SUPPORT. FURTHERMORE, SUPRATENTORIAL CEREBRAL EDEMA WAS OBSERVED ON A BRAIN SCAN. FOLLOWING THE PROCEDURE, TREATMENT WAS DISCONTINUED AND THE PATIENT PASSED AWAY. Manufacturer Narrative: THE REPORTED EVENT WAS PERFORATION AND THE CATHETER WAS TOO RIGID. AS RECEIVED, A CATHETER WAS RETURNED IN ONE PIECE FOR ANALYSIS WITH DEVICE ATTACHED. VISUAL EXAMINATION OF THE CATHETER DID NOT FIND ANY ANOMALIES RELATED TO RIGIDITY. X-RAY EXAMINATION FOUND THE HYPO TUBE IS FRACTURED ADJACENT TO THE HANDLE. THE ASSOCIATED DEVICE WAS ABLE TO UN-DOCK AND RELEASE WITHOUT DIFFICULTY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22322126&pc=NPT | 2025587-2025-04461 | 2025-06-12 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-24 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Embolism/Embolus | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF THE EVOLUTFX-26 VALVE, A PRE-IMPLANT BALLOON DILATION WITH AN 18X40 NON-MEDTRONIC BALLOON WAS PERFORMED. FOLLOWING THE VALVE IMPLANT AND DURING CLOSURE OF THE VASCULAR ACCESS, A POSSIBLE OCCLUSION OR DISSECTION IN THE RIGHT COMMON FEMORAL ARTERY WAS IDENTIFIED. ANGIOPLASTY WITH A PERIPHERAL BALLOON WAS PERFORMED, AND A COATED STENT WAS IMPLANTED IN THE RIGHT COMMON FEMORAL ARTERY. AN ECHOCARDIOGRAM WAS PERFORMED AND SHOWED A MEAN GRADIENT OF 3 MMHG. DURING OROTRACHEAL EXTUBATION, THE PATIENT EXPERIENCED ELECTRICAL ACTIVITY WITHOUT A PULSE. CARDIOPULMONARY RESUSCITATION WAS PERFORMED BUT WAS UNSUCCESSFUL AND THE PATIENT SUBSEQUENTLY DIED. THE HEART RHYTHM ACTIVITY WAS RECORDED ON THE MONITOR BEFORE, DURING, AND AFTER THE EVENT WITHOUT SHOWING A COMPLETE HEART BLOCK. LABORATORY TESTS SHOWED NO CHANGES, AND THE PATIENT HAD NORMAL PARAMETERS FOR EXTUBATION, INCLUDING NORMAL BLOOD PRESSURE AND HEART RATE, AND WAS AWAKE AND COMMUNICATING. THE PROBABLE DIAGNOSIS WAS PULMONARY EMBOLISM, ALTHOUGH THIS WAS UNCONFIRMED.Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF THE EVOLUTFX-26 VALVE, A PRE-IMPLANT BALLOON DILATION WITH AN 18X40 NON-MEDTRONIC (ATLAS) BALLOON WAS PERFORMED. FOLLOWING THE VALVE IMPLANT AND DURING CLOSURE OF THE VASCULAR ACCESS, A POSSIBLE OCCLUSION OR DISSECTION IN THE RIGHT COMMON FEMORAL ARTERY WAS IDENTIFIED. ANGIOPLASTY WITH A PERIPHERAL BALLOON WAS PERFORMED, AND A COATED STENT WAS IMPLANTED IN THE RIGHT COMMON FEMORAL ARTERY. AN ECHOCARDIOGRAM WAS PERFORMED AND SHOWED A MEAN GRADIENT OF 3 MMHG. DURING OROTRACHEAL EXTUBATION, THE PATIENT EXPERIENCED ELECTRICAL ACTIVITY WITHOUT A PULSE. CARDIOPULMONARY RESUSCITATION WAS PERFORMED BUT WAS UNSUCCESSFUL AND THE PATIENT SUBSEQUENTLY DIED. THE HEART RHYTHM ACTIVITY WAS RECORDED ON THE MONITOR BEFORE, DURING, AND AFTER THE EVENT WITHOUT SHOWING A COMPLETE HEART BLOCK. LABORATORY TESTS SHOWED NO CHANGES, AND THE PATIENT HAD NORMAL PARAMETERS FOR EXTUBATION, INCLUDING NORMAL BLOOD PRESSURE AND HEART RATE, AND WAS AWAKE AND COMMUNICATING. THE PROBABLE DIAGNOSIS WAS PULMONARY EMBOLISM, ALTHOUGH THIS WAS UNCONFIRMED. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT VASCULAR ACCESS FOR DELIVERY CATHETER SYSTEM (DCS) INSERTION WAS ACHIEVED AT THE RIGHT COMMON FEMORAL ARTERY WHICH HAD A MINIMUM DIAMETER OF 6MM. PER THE PHYSICIAN, THE NON-MEDTRONIC (PERCLOSE) VASCULAR CLOSURE SYSTEM WAS THE CAUSE OF THE OCCLUSION AND/OR DISSECTION AT THE ACCESS SITE; THE DCS DID NOT CONTRIBUTE TO THE ACCESS SITE INJURY. ADDITIONALLY, THE PHYSICIAN INDICATED THERE WAS NO ANOMALOUS PATIENT ANATOMY THAT CONTRIBUTED TO THE INJURY AT THE ACCESS SITE. THE PHYSICIAN ALSO REPORTED NEITHER THE MEDTRONIC VALVE NOR THE DCS CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: L-EVOLUTFX-2329, (LOT: 0012405245); PRODUCT TYPE: 0195-HEART VALVES; NON-IMPLANTABLE DEVICE PRODUCT ID: D-EVOLUTFX-2329, (0012483145); PRODUCT TYPE: 0195-HEART VALVES; NON-IMPLANTABLE DEVICE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22322131&pc=NPT | 2025587-2025-04462 | 2025-06-12 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-24 04:00 | NPT | EVOLUT FX DCS | Insufficient Information; Patient Device Interaction Problem | Cardiac Arrest; Vascular Dissection; Embolism/Embolus | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF THE EVOLUTFX-26 VALVE, A PRE-IMPLANT BALLOON DILATION WITH AN 18X40 NON-MEDTRONIC BALLOON WAS PERFORMED. FOLLOWING THE VALVE IMPLANT AND DURING CLOSURE OF THE VASCULAR ACCESS, A POSSIBLE OCCLUSION OR DISSECTION IN THE RIGHT COMMON FEMORAL ARTERY WAS IDENTIFIED. ANGIOPLASTY WITH A PERIPHERAL BALLOON WAS PERFORMED, AND A COATED STENT WAS IMPLANTED IN THE RIGHT COMMON FEMORAL ARTERY. AN ECHOCARDIOGRAM WAS PERFORMED AND SHOWED A MEAN GRADIENT OF 3 MMHG. DURING OROTRACHEAL EXTUBATION, THE PATIENT EXPERIENCED ELECTRICAL ACTIVITY WITHOUT A PULSE. CARDIOPULMONARY RESUSCITATION WAS PERFORMED BUT WAS UNSUCCESSFUL AND THE PATIENT SUBSEQUENTLY DIED. THE HEART RHYTHM ACTIVITY WAS RECORDED ON THE MONITOR BEFORE, DURING, AND AFTER THE EVENT WITHOUT SHOWING A COMPLETE HEART BLOCK. LABORATORY TESTS SHOWED NO CHANGES, AND THE PATIENT HAD NORMAL PARAMETERS FOR EXTUBATION, INCLUDING NORMAL BLOOD PRESSURE AND HEART RATE, AND WAS AWAKE AND COMMUNICATING. THE PROBABLE DIAGNOSIS WAS PULMONARY EMBOLISM, ALTHOUGH THIS WAS UNCONFIRMED.Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF THE EVOLUTFX-26 VALVE, A PRE-IMPLANT BALLOON DILATION WITH AN 18X40 NON-MEDTRONIC (ATLAS) BALLOON WAS PERFORMED. FOLLOWING THE VALVE IMPLANT AND DURING CLOSURE OF THE VASCULAR ACCESS, A POSSIBLE OCCLUSION OR DISSECTION IN THE RIGHT COMMON FEMORAL ARTERY WAS IDENTIFIED. ANGIOPLASTY WITH A PERIPHERAL BALLOON WAS PERFORMED, AND A COATED STENT WAS IMPLANTED IN THE RIGHT COMMON FEMORAL ARTERY. AN ECHOCARDIOGRAM WAS PERFORMED AND SHOWED A MEAN GRADIENT OF 3 MMHG. DURING OROTRACHEAL EXTUBATION, THE PATIENT EXPERIENCED ELECTRICAL ACTIVITY WITHOUT A PULSE. CARDIOPULMONARY RESUSCITATION WAS PERFORMED BUT WAS UNSUCCESSFUL AND THE PATIENT SUBSEQUENTLY DIED. THE HEART RHYTHM ACTIVITY WAS RECORDED ON THE MONITOR BEFORE, DURING, AND AFTER THE EVENT WITHOUT SHOWING A COMPLETE HEART BLOCK. LABORATORY TESTS SHOWED NO CHANGES, AND THE PATIENT HAD NORMAL PARAMETERS FOR EXTUBATION, INCLUDING NORMAL BLOOD PRESSURE AND HEART RATE, AND WAS AWAKE AND COMMUNICATING. THE PROBABLE DIAGNOSIS WAS PULMONARY EMBOLISM, ALTHOUGH THIS WAS UNCONFIRMED. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT VASCULAR ACCESS FOR DELIVERY CATHETER SYSTEM (DCS) INSERTION WAS ACHIEVED AT THE RIGHT COMMON FEMORAL ARTERY WHICH HAD A MINIMUM DIAMETER OF 6MM. PER THE PHYSICIAN, THE NON-MEDTRONIC (PERCLOSE) VASCULAR CLOSURE SYSTEM WAS THE CAUSE OF THE OCCLUSION AND/OR DISSECTION AT THE ACCESS SITE; THE DCS DID NOT CONTRIBUTE TO THE ACCESS SITE INJURY. ADDITIONALLY, THE PHYSICIAN INDICATED THERE WAS NO ANOMALOUS PATIENT ANATOMY THAT CONTRIBUTED TO THE INJURY AT THE ACCESS SITE. THE PHYSICIAN ALSO REPORTED NEITHER THE MEDTRONIC VALVE NOR THE DCS CONTRIBUTED TO THE PATIENT'S DEATH. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: L-EVOLUTFX-2329 (LOT: 0012405245); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE. PRODUCT ID: EVOLUTFX-26 ((B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22322227&pc=LWR | 2135147-2025-03394 | 2001-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-06-24 04:00 | LWR | TRIFECTA¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100029 |  | 1 | Event Description: THE ARTICLE, "ANTICOAGULATION THERAPY AND RELATED OUTCOMES AMONG ASIAN PATIENTS AFTER BIOPROSTHETIC VALVE REPLACEMENT", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, MULTICENTER STUDY TO IDENTIFY THE OPTIMAL INTERNATIONAL NORMALIZED RATIO (INR) RANGE FOR ASIAN PATIENTS DURING A 1-YEAR FOLLOW-UP AFTER TISSUE VALVE REPLACEMENT. DEVICES INCLUDED IN THE STUDY WERE HANCOCK II, ST. JUDE EPIC, TRIFECTA, MOSAIC, MAGNA EASE/PERIMOUNT, AND UNKNOWN. THE ARTICLE CONCLUDED THAT FOR PATIENTS WHO UNDERWENT TISSUE VALVE REPLACEMENT, THE BLEEDING RISK WAS ELEVATED WHEN THE INR WAS GREATER THAN 2.0, BUT THE RISK OF THROMBOEMBOLIC EVENT INCREASED ONLY WHEN THE INR WAS LOWER THAN 1.84 IN THE ATRIAL FIBRILLATION (AF) GROUP AND 1.7 IN THE NON-AF GROUP, REGARDLESS OF WHETHER THE PATIENT RECEIVED AORTIC VALVE REPLACEMENT (AVR), MITRAL VALVE REPLACEMENT (MVR), OR DOUBLE VALVE REPLACEMENT (DVR). [THE PRIMARY AUTHOR WAS TANG-YU LIU, AMERICAN COLLEGE OF CARDIOLOGY, WASHINGTON, D.C., U.S.A. THE CORRESPONDING AUTHOR WAS SHAO-WEI CHEN, DIVISION OF THORACIC AND CARDIOVASCULAR SURGERY, DEPARTMENT OF SURGERY, CHANG GUNG MEMORIAL HOSPITAL, LINKOU MEDICAL CENTER, CHANG GUNG UNIVERSITY, NO. 5 FUXING STREET, GUISHAN DISTRICT, TAOYUAN CITY 33305, TAIWAN, WITH CORRESPONDING EMAIL: JOSEPHCHEN0314@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS FROM 01 JANUARY 2001 TO 31 DECEMBER 2018. A TOTAL OF 1059 PATIENTS WERE INCLUDED IN THE STUDY. OF THE 1192 TOTAL DEVICES INVOLVED, 640 (53.7%) WERE ABBOTT. THE AVERAGE AGE WAS 65.5 YEARS, THE AVERAGE WEIGHT WAS 58.9KG, AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED CHRONIC OBSTRUCTIVE PULMONARY DISEASE, CHRONIC LIVER DISEASE, CHRONIC KIDNEY DISEASE, HYPERTENSION, HYPERLIPIDEMIA, DIABETES, INFECTIVE ENDOCARDITIS, RHEUMATIC HEART DISEASE, LUNG EDEMA, GOUT, PERIPHERAL ARTERY DISEASE, MALIGNANCY, PRIOR ISCHEMIC STROKE, PRIOR MYOCARDIAL INFARCTION, HISTORY OF BLEEDING (GASTROINTESTINAL, INTRACRANIAL, MAJOR), AND HEART FAILURE HOSPITALIZATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF ANTICOAGULATION THERAPY AND RELATED OUTCOMES AMONG ASIAN PATIENTS AFTER BIOPROSTHETIC VALVE REPLACEMENT WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING CHRONIC OBSTRUCTIVE PULMONARY DISEASE, CHRONIC LIVER DISEASE, CHRONIC KIDNEY DISEASE, HYPERTENSION, HYPERLIPIDEMIA, DIABETES, INFECTIVE ENDOCARDITIS, RHEUMATIC HEART DISEASE, LUNG EDEMA, GOUT, PERIPHERAL ARTERY DISEASE, MALIGNANCY, PRIOR ISCHEMIC STROKE, PRIOR MYOCARDIAL INFARCTION, HISTORY OF BLEEDING (GASTROINTESTINAL, INTRACRANIAL, MAJOR), AND HEART FAILURE HOSPITALIZATION. SOME OF THE COMPLICATIONS REPORTED WERE EMBOLISM, THROMBUS, HEMORRHAGE, STROKE AND DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. CODES REMOVED: B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: ANTICOAGULATION THERAPY AND RELATED OUTCOMES AMONG ASIAN PATIENTS AFTER BIOPROSTHETIC VALVE REPLACEMENT. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: ANTICOAGULATION THERAPY AND RELATED OUTCOMES AMONG ASIAN PATIENTS AFTER BIOPROSTHETIC VALVE REPLACEMENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22322260&pc=FKX | 0008030665-2025-01422 | 2025-06-02 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-06-24 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K222318 |  | 1 | Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED TO FRESENIUS CUSTOMER SUPPORT THAT A PATIENT PASSED AWAY ON THE CYCLER. THE CAUSE OF DEATH IS UNKNOWN. A RED BAG PICKUP FROM THE CLINIC PROCESS WAS REQUESTED. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DISCHARGE SUMMARY, HOSPITAL RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, TIMELINE OF EVENTS, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL.Event Description: A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED TO FRESENIUS CUSTOMER SUPPORT THAT A PATIENT PASSED AWAY ON THE CYCLER. THE CAUSE OF DEATH IS UNKNOWN. A RED BAG PICKUP FROM THE CLINIC PROCESS WAS REQUESTED. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DISCHARGE SUMMARY, HOSPITAL RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, TIMELINE OF EVENTS, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, G4, H3. PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER WITH NO PHYSICAL DAMAGE NOTED. ALTHOUGH THERE WAS NO EVIDENCE OF DRIED FLUID PRESENT WITHIN THE CASSETTE COMPARTMENT, THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED THE CASSETTE MEMBRANE. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THE CYCLER UNDERWENT AND PASSED A SYSTEM AIR LEAK TEST, VALVE ACTUATION TEST, AND TEACH PUMPS TEST. THERE WERE NO VISUAL DISCREPANCIES ENCOUNTERED DURING THE INTERNAL INSPECTION. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED. Manufacturer Narrative: CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENT OCCURRED. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. IT SHOULD BE NOTED THAT LIMITED CLINICAL FOLLOW-UP PRECLUDED A MORE COMPREHENSIVE ASSESSMENT/INVESTIGATION. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CONTRIBUTORY ROLE IN THE PATIENT¿S EXPIRATION. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION OCCURRED. HOWEVER, GIVEN THE UNKNOWN CAUSE OF DEATH, PAST MEDICAL HISTORY, SUMMARY/TIMELINE OF EVENTS, DISCHARGE SUMMARY, AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT DISASSOCIATE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22322399&pc=OZD | 1220648-2025-29755 | 2025-06-09 04:00 | Death | ABIOMED, INC. | 2025-06-24 04:00 | OZD | IMPELLA CP | Detachment of Device or Device Component | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: PRODUCT DAMAGE WAS REPORTED BY A CUSTOMER IMPLANTING AN IMPELLA CP. THE STERILE SLEEVE TORE OFF THE DISTAL END OF THE DEVICE AND THE PUMP BECAME UNSTERILE. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22322601&pc=DSQ | 2916596-2025-03971 | 2025-06-04 04:00 | Death | THORATEC CORPORATION | 2025-06-24 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT BE EXPLANTED. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22322611&pc=DSQ | 2916596-2025-03975 | 2025-06-04 04:00 | Death | THORATEC CORPORATION | 2025-06-24 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT BE EXPLANTED. Manufacturer Narrative: A1-A6: PATIENT INFORMATION PENDING FOLLOW-UP WITH HOSPITAL. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22322619&pc=OZD | 1220648-2025-29759 | 2025-06-09 04:00 | Death | ABIOMED, INC. | 2025-06-24 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Ventricular Fibrillation; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP ENCOUNTERED A THROMBUS, LOW PUMP FLOW AND VENTRICULAR FIBRILLATION (VF). AFTER 10 MINUTES OF SUPPORT, THE IMPELLA BEGAN TO SUCTION WITH NO FLOW. THE PATIENT BECAME HYPOTENSIVE AND ARRESTED. A THROMBUS WAS FOUND OCCLUDING THE AORTIC ROOT, FOR WHICH CHEST COMPRESSION WAS PERFORMED. THE THROMBUS MIGRATED AND PUMP FLOWS RETURNED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF THROMBUS, LOW PUMP FLOW, AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE DATA LOG SHOWS SEVERAL SUCTION ALARMS AND LOW PUMP FLOW DURING THE CASE START, WITH AN OBSERVED TRENDY PUMP FLOW AND PLACEMENT SIGNAL TREND. PUMP FLOW WAS RECOVERED AND REMAINED WITHIN SPECIFICATIONS FOR THE REST OF THE CASE RUN. THERE WAS NO INDICATION OF BIOMATERIAL INTERACTION WITH THE IMPELLER DURING THE CASE. ACCORDING TO THE CLINICAL NOTES, FLOW RETURNED TO NORMAL AFTER A LARGE THROMBUS MIGRATED FROM THE AORTIC ROOT. THE CAUSE OF THE LOW PUMP FLOW ISSUE WAS NOT DETERMINED WITHOUT RETURNED PRODUCT INVESTIGATION. THE ROOT CAUSE OF THE THROMBUS WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. THE ROOT CAUSE OF THE VT/VF/AT/AF WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. B1 PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29759. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST SYSTEM LISTS VENTRICULAR ARRYTHMIA AS AN ADVERSE EVENT THAT OCCURRED WHILE IN USE WITH THE IMPELLA DEVICE. IT STATES THAT "ONLY A FRACTION OF THESE RATES WERE ATTRIBUTED TO THE IMPELLA DEVICE AND THE EVENTS RESOLVED WITHOUT RESIDUAL EFFECT IN MOST OF THE CASES, UNLESS THE EVENT OF DEATH OCCURRED. OVERALL, THE RESULTS DID NOT SHOW ANY EVIDENCE OF INCREASED MORBIDITY ASSOCIATED WITH THE IMPELLA SUPPORT IN CARDIOMYOPATHY, MYOCARDITIS, AND PERIPARTUM CARDIOMYOPATHY (PPCM) PATIENTS." IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22322832&pc=DSQ | 2916596-2025-03899 | 2025-05-29 04:00 | Death | THORATEC CORPORATION | 2025-06-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY SECONDARY TO RIGHT VENTRICULAR FAILURE (RVF). NO ADDITIONAL INFORMATION WAS PROVIDED AT THE TIME. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION B3: DATE OF DEATH CORRECTED MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT AND THE RETURN STATUS OF PUMP; HOWEVER, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RVAD PLACEMENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22322869&pc=FKX | 0008030665-2025-01438 | 2025-06-12 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-06-24 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K181108 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING TREATMENT ON (B)(6) 2025. THE PATIENT REPORTEDLY AWOKE GASPING FOR AIR AND PASSED SHORTLY THEREAFTER. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DISCHARGE SUMMARY, HOSPITAL RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, TIMELINE OF EVENTS, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL. Manufacturer Narrative: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENT OCCURRED. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. IT SHOULD BE NOTED THAT LIMITED CLINICAL FOLLOW-UP PRECLUDED A MORE COMPREHENSIVE ASSESSMENT/INVESTIGATION. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CONTRIBUTORY ROLE IN THE PATIENT¿S EXPIRATION. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION OCCURRED. HOWEVER, GIVEN THE UNKNOWN CAUSE OF DEATH, PAST MEDICAL HISTORY, SUMMARY/TIMELINE OF EVENTS, DISCHARGE SUMMARY, AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT EXCLUDE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22322977&pc=DSQ | 2916596-2025-03915 | 2025-02-17 05:00 | Death | THORATEC CORPORATION | 2025-06-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Inflammation; Shock; Respiratory Failure; Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO ACUTE-ON-CHRONIC HYPOXEMIC HYPERCAPNIC RESPIRATORY FAILURE, RIGHT VENTRICULAR FAILURE (RVF), SYSTEMIC INFLAMMATORY RESPONSE SYNDROME (SIRS) INFLAMMATION, AND DISTRIBUTIVE SHOCK. THE DEVICE WOULD NOT BE SENT BACK.Event Description: IT WAS REPORTED THAT THE PATIENT HAD A TRACHEOSTOMY ON (B)(6) 2025, UNRELATED TO LVAD PROCEDURE. THE PATIENT HAD PRE-EXISTING DILATED RIGHT VENTRICLE AND MILDLY DEPRESSED RIGHT VENTRICULAR SYSTOLIC FUNCTION. THE PATIENT EXHIBITED SYMPTOMS OF HYPOTENSION AND MULTI SYSTEM ORGAN FAILURE. THE PATIENT CONTINUED TO REQUIRE INTERMITTENT POSITIVE PRESSURE VENTILATION, WORSENING RV FAILURE REQUIRING ESCALATION IN PHARMACOLOGIC RV SUPPORT DUE TO CHRONIC RV FAILURE, SIRS INFLAMMATION, AND DISTRIBUTIVE SHOCK.Event Description: IT WAS REPORTED THAT A DEVICE RELATED ISSUE DID NOT CAUSE OR CONTRIBUTE TO THE RIGHT HEART FAILURE, RESPIRATORY FAILURE, SIRS, OR SHOCK. THE DEVICE OPERATED AS EXPECTED AND WAS NOT CONSIDERED RELATED TO THE DEATH. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS RESPIRATORY FAILURE, RIGHT HEART FAILURE, AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: CORRECTED DATA: SECTION B5: EVENT DESCRIPTION CORRECTED. SECTION H6: HEALTH EFFECT - IMPACT CODE CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22323005&pc=DSQ | 2916596-2025-03916 | 2024-10-01 04:00 | Death | THORATEC CORPORATION | 2025-06-24 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED FOR SEPSIS ON (B)(6) 2024 FROM SERRATIA BLOOD CULTURE BACTEREMIA AND WAS DEBILITATED. THE PATIENT AND FAMILY DECIDED TO NOT SEEK CURATIVE TREATMENT AND GO HOME ON HOSPICE CARE. WHILE ON HOSPICE CARE, THE PASSED AWAY AT HOME ON (B)(6) 2025. THE PUMP WAS ON AT THE TIME OF DEATH AND WORKING AS INTENDED. THE PUMP WOULD NOT BE EXPLANTED OR RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22323168&pc=MIH | 2247858-2025-00153 | 2024-11-25 05:00 | Death | BOLTON MEDICAL, INC. | 2025-06-24 04:00 | MIH | RELAY PRO NBS THORACIC STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem | Respiratory Tract Infection | P200045 |  | 1 | Event Description: "POSTOPERATIVE RESPIRATORY FAILURE DUE TO EXTRINSIC BRONCHIUS COMPRESSION + INFECTION AE STARTED ON (B)(6) 2024. SAE WAS CLASSED AS SERIOUS, SERIOUS CLASSIFICATION AS DEATH, MEDICAL OR SURGICAL INTERVENTION WAS PERFORMED TO PREVENT LIFE-THREATENING ILLNESS OR INJURY OR PERMANENT IMPAIRMENT TO BODY STRUCTURE OR A BODY FUNCTION RELATED TO PROCEDURE UNDETERMINED AS RELATED TO DEVICE NOT DETERMINED AS RELATED TO PRE-EXISTING CONDITION UNANTICIPATED EVENT IN RESPONSE TO AE THORACOTOMY AND DISTAL DESCENDING THORACIC AORTA BYPASS TO LSA + TEVAR ZONEZ 0 TO 4 CARRIED OUT PLUS OTHER INTERVENTION BRONCHOSCOPY AND BRONCHIAL STENTING ATTEMPT". PATIENT OUTCOME: "OUTCOME OF EVENT WAS DEATH ON (B)(6) 2025 END EVENT WAS DEATH". Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH TWO DEVICES. DEVICE 1 IS BEING REPORTED UNDER , AND DEVICE 2 IS BEING REPORTED UNDER MDR-2247858-2025-00154. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22323169&pc=MIH | 2247858-2025-00154 | 2024-11-25 05:00 | Death | BOLTON MEDICAL, INC. | 2025-06-24 04:00 | MIH | RELAY PRO NBS THORACIC STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem | Respiratory Tract Infection | P200045 |  | 1 | Event Description: "POSTOPERATIVE RESPIRATORY FAILURE DUE TO EXTRINSIC BRONCHIUS COMPRESSION + INFECTION AE STARTED ON (B)(6) 2024. SAE WAS CLASSED AS SERIOUS, SERIOUS CLASSIFICATION AS DEATH, MEDICAL OR SURGICAL INTERVENTION WAS PERFORMED TO PREVENT LIFE-THREATENING ILLNESS OR INJURY OR PERMANENT IMPAIRMENT TO BODY STRUCTURE OR A BODY FUNCTION, RELATED TO PROCEDURE, UNDETERMINED AS RELATED TO DEVICE, NOT DETERMINED AS RELATED TO PRE-EXISTING CONDITION, UNANTICIPATED EVENT. IN RESPONSE TO AE THORACOTOMY AND DISTAL DESCENDING THORACIC AORTA BYPASS TO LSA + TEVAR ZONEZ 0 TO 4 CARRIED OUT PLUS OTHER INTERVENTION BRONCHOSCOPY AND BRONCHIAL STENTING ATTEMPT". PATIENT OUTCOME: "OUTCOME OF EVENT WAS DEATH ON (B)(6) 2025 END EVENT WAS DEATH". Manufacturer Narrative: THIS COMPLAINT WAS INVOLVED WITH TWO DEVICES. DEVICE 1 IS BEING REPORTED UNDER MDR-2247858-2025-00153, AND DEVICE 2 IS BEING REPORTED UNDER. BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22323366&pc=NPT | 2135147-2025-03403 | 2025-05-30 04:00 | Death | ABBOTT MEDICAL | 2025-06-24 04:00 | NPT | FLEXNAV DELIVERY SYSTEM | Use of Device Problem | Low Blood Pressure/ Hypotension | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 29MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM (DS). DURING THE PROCEDURE, THE PATIENT HAD A PRE-BAV PERFORMED, AFTER WHICH THE DS WAS BEING BROUGHT UP AROUND THE ARCH. AT SOME POINT DURING THIS, THE PATIENTS BLOOD PRESSURE DROPPED AND THERE WAS A CHANGE IN THE ECG. THE BLOOD PRESSURE WAS TREATED MEDICALLY WHILE THE VALVE WAS POSITIONED. AT 80% DEPLOYMENT, THEY REALIZED THAT THE PATIENT CONTINUED TO CRASH AND SO THEY REMOVED THE VALVE AND DS. IT WAS NOTED THAT THE PATIENT HAD EXPERIENCED AND AORTIC DISSECTION AND WAS PLACED ON ECMO. FURTHER INTERVENTION WAS GIVEN AS AN OPTION, BUT ULTIMATELY DENIED. THE PATIENT PASSED AWAY AS A RESULT. IT'S THOUGHT THAT DUE TO THE EXTENSIVE CALCIUM THROUGHOUT THE AORTA, THAT THE SYSTEM DISRUPTED A PIECE OF THE PLAQUE AND CAUSED A FLAP WHICH WENT DOWN AND COVERED THE RCA.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 29MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM (DS). DURING THE PROCEDURE, THE PATIENT HAD A PRE-BAV PERFORMED, AFTER WHICH THE DS WAS BEING BROUGHT UP AROUND THE ARCH. AT SOME POINT DURING THIS, THE PATIENTS BLOOD PRESSURE DROPPED AND THERE WAS A CHANGE IN THE ECG. THE BLOOD PRESSURE WAS TREATED MEDICALLY WHILE THE VALVE WAS POSITIONED. AT 80% DEPLOYMENT, THEY REALIZED THAT THE PATIENT CONTINUED TO CRASH AND SO THEY REMOVED THE VALVE AND DS. THE PATIENT WAS PLACED ON ECMO, AFTER QQHICH ECHO WAS PERFORMED AND SHOWED THAT THE PATIENT HAD EXPERIENCED AN AORTIC DISSECTION. FURTHER INTERVENTION WAS GIVEN AS AN OPTION, BUT ULTIMATELY DENIED. THE PATIENT PASSED AWAY AS A RESULT. IT'S THOUGHT THAT DUE TO THE EXTENSIVE CALCIUM THROUGHOUT THE AORTA, THAT THE SYSTEM DISRUPTED A PIECE OF THE PLAQUE AND CAUSED A FLAP WHICH WENT DOWN AND COVERED THE RCA. Manufacturer Narrative: THE DEVICE IS NOT RETURNING FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: IT WAS REPORTED THAT DURING NAVITOR PROCEDURE PATIENT'S BLOOD PRESSURE DROPPED AND WAS MEDICALLY TREATED WHILE THE VALVE WAS POSITIONED. THE VALVE WAS DEPLOYED AT 80%, HOWEVER, DUE THE PATIENTS HEMODYNAMIC STATE DETERIORATED AND IT WAS DECIDED TO RECAPTURE THE VALVE AND REMOVE THE DS FROM THE BODY. THE PATIENT EXPERIENCED AN AORTIC DISSECTION. REPORTEDLY, THE PATIENT EXPIRED THE SAME DAY OF PROCEDURE. THE DEVICE WAS NOT RETURNED TO ABBOTT FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. INFORMATION FROM FIELD INDICATED THAT DUE TO THE EXTENSIVE CALCIUM THROUGHOUT THE AORTA, THAT THE SYSTEM DISRUPTED A PIECE OF THE PLAQUE AND CAUSED A FLAP WHICH WENT DOWN AND COVERED THE RCA. BASED ON THE INFORMATION AVAILABLE, THE REPORTED AORTIC DISSECTION IS POSSIBLY RELATED TO THE ANATOMICAL AND PROCEDURAL CONDITION (EXTENSIVE CALCIUM THROUGHOUT THE AORTA). HOWEVER, IT IS DIFFICULT TO DETERMINE WITH CERTAINTY THE EXACT CAUSE OF THE REPORTED EVENT. THE CAUSE OF PATIENT DEATH COULD NOT BE DETERMINED. THERE WERE NO COMPLAINTS ASSOCIATED WITH ANY OTHER DEVICES FROM THE LOT. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22323863&pc=MHX | 3009156722-2025-00009 | 2025-05-27 04:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-06-24 04:00 | MHX | PANORAMA CENTRAL STATION WITH TELEMETRY | Device Handling Problem | Insufficient Information | K150082 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, BETWEEN 6:30 AM AND 7:30 AM, CLINICAL STAFF DID NOT HEAR AN ALARM AT THE WORKSTATION (PANORAMA CENTRAL STATION). IN ADDITION, IT WAS REPORTED THE PATIENT WAS ONLY CONNECTED TO PATIENT MONITORING (MINDRAY V12). THE PATIENT EXPIRED. Manufacturer Narrative: A SIMULATOR WAS CONNECTED TO AN EMERGENCY DEPARTMENT MONITOR, TRIGGERED AN ALARM, AND CONFIRMED THAT THERE WAS NO AUDIO COMING OUT OF THE WORKSTATION. SYSTEM LOGS WERE NOT ABLE TO BE COLLECTED AS THE SYSTEM INDICATED A BAD HDD. CUSTOMER BIOMED, WITH THE ASSISTANCE OF MINDRAY FIELD SERVICE, INVESTIGATED THE AUDIO PROBLEMS OF THE CENTRAL STATION AND IT WAS FOUND THAT THE STAFF NOT HEARING THE AUDIBLE ALARM WAS CAUSED BY THE INCORRECT CONNECTION OF THE SPEAKER CABLES TO THE KVM IN THE SERVER ROOM. THIS LED TO THE AUDIO OF THE ALARM NOT SOUNDING AT THE NURSE'S STATION. IT IS LIKELY THE EQUIPMENT WAS ALTERED AFTER THE MANUFACTURER INSTALLATION. THE MINDRAY REPRESENTATIVE ADJUSTED THE CABLES THAT WERE INCORRECTLY CONNECTED. HE MOVED THE TWO CABLES MISTAKENLY INSERTED INTO THE MICROPHONE PORTS TO THE CORRESPONDING SPEAKER PORTS. AFTER THE ADJUSTMENT WAS COMPLETED, THE AUDIO ALARMS FUNCTIONED PER SPECIFICATIONS. MINDRAY PANORAMA CENTRAL STATION PERFORMED PER SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22323866&pc=MHX | 2221819-2025-00009 | 2025-05-27 04:00 | Death | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 2025-06-24 04:00 | MHX | PANORAMA CENTRAL STATION WITH TELEMETRY | Device Handling Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, BETWEEN 6:30 AM AND 7:30 AM, CLINICAL STAFF DID NOT HEAR AN ALARM AT THE WORKSTATION (PANORAMA CENTRAL STATION). IN ADDITION, IT WAS REPORTED THE PATIENT WAS ONLY CONNECTED TO PATIENT MONITORING (MINDRAY V12). THE PATIENT EXPIRED. Manufacturer Narrative: A SIMULATOR WAS CONNECTED TO AN EMERGENCY DEPARTMENT MONITOR, TRIGGERED AN ALARM, AND CONFIRMED THAT THERE WAS NO AUDIO COMING OUT OF THE WORKSTATION. SYSTEM LOGS WERE NOT ABLE TO BE COLLECTED AS THE SYSTEM INDICATED A BAD HDD. CUSTOMER BIOMED, WITH THE ASSISTANCE OF MINDRAY FIELD SERVICE, INVESTIGATED THE AUDIO PROBLEMS OF THE CENTRAL STATION AND IT WAS FOUND THAT THE STAFF NOT HEARING THE AUDIBLE ALARM WAS CAUSED BY THE INCORRECT CONNECTION OF THE SPEAKER CABLES TO THE KVM IN THE SERVER ROOM. THIS LED TO THE AUDIO OF THE ALARM NOT SOUNDING AT THE NURSE'S STATION. IT IS LIKELY THE EQUIPMENT WAS ALTERED AFTER THE MANUFACTURER INSTALLATION. THE MINDRAY REPRESENTATIVE ADJUSTED THE CABLES THAT WERE INCORRECTLY CONNECTED. HE MOVED THE TWO CABLES MISTAKENLY INSERTED INTO THE MICROPHONE PORTS TO THE CORRESPONDING SPEAKER PORTS. AFTER THE ADJUSTMENT WAS COMPLETED, THE AUDIO ALARMS FUNCTIONED PER SPECIFICATIONS. MINDRAY PANORAMA CENTRAL STATION PERFORMED PER SPECIFICATIONS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22327767&pc=MOZ | MW5171876 |  | Death | FERRING PHARMACEUTICALS INC. | 2025-06-24 04:00 | MOZ | EUFLEXXA 20 MG/2ML (6=3 SYR | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: PT'S NURSE (B)(6) REPORTS THE PT HAS PASSED AWAY. NO OTHER INFO WAS PROVIDED. FREQ: INJECT CONTENT OF ONE SYRINGE INTRA-ARTICULARLY INTO RIGHT KNEE ONCE WEEKLY FOR THREE WEEKS. (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22305121&pc=FRN | 2016493-2025-89239 | 2025-05-27 04:00 | Death | CAREFUSION SD | 2025-06-23 04:00 | FRN | ALARIS SYSTEM | Excess Flow or Over-Infusion | Insufficient Information | K133532 |  | 1 | Event Description: IT WAS INITIALLY REPORTED AN OVER INFUSION EVENT THAT OCCURRED ON (B)(6) 2025. PER REPORT, THE PATIENT "DIED." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE DEVICE LOGS TO IDENTIFY WHICH MEDICATIONS WERE INFUSED FROM 3:00 PM ON (B)(6) 2025 TO 1:30 AM ON (B)(6) 2025. NO FURTHER INFORMATION WAS PROVIDED. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "THE REASON TO DO THIS (REVIEW OF DEVICE LOGS) ONLY TO CHECK THE MEDICATIONS AND THE RATE OF IT GIVEN BY THE MACHINE. I¿M NOT BELIEVE THAT THE DEVICE CAUSED THE DEATH BUT MAYBE HUMAN MISTAKE." MULTIPOLE ATTEMPTS HAVE BEEN MADE REQUESTING FOR THE DETAILS OF THE EVENT. HOWEVER, NO FURTHER INFORMATION HAS BEEN PROVIDED TO DATE. Manufacturer Narrative: THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22305233&pc=BZD | 2518422-2025-047598 | 2024-04-26 04:00 | Death | RESPIRONICS, INC. | 2025-06-23 04:00 | BZD | DREAMSTATION CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED OF DEATH. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22305670&pc=FRN | 2016493-2025-89242 | 2025-05-27 04:00 | Death | CAREFUSION SD | 2025-06-23 04:00 | FRN | ALARIS SYSTEM | Excess Flow or Over-Infusion | Insufficient Information | K133532 |  | 1 | Event Description: IT WAS INITIALLY REPORTED AN OVER INFUSION EVENT THAT OCCURRED ON (B)(6) 2025. PER REPORT, THE PATIENT "DIED." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE DEVICE LOGS TO IDENTIFY WHICH MEDICATIONS WERE INFUSED FROM 3:00 PM ON (B)(6) 2025 TO 1:30 AM ON (B)(6) 2025. NO FURTHER INFORMATION WAS PROVIDED. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "THE REASON TO DO THIS (REVIEW OF DEVICE LOGS) ONLY TO CHECK THE MEDICATIONS AND THE RATE OF IT GIVEN BY THE MACHINE. I¿M NOT BELIEVE THAT THE DEVICE CAUSED THE DEATH BUT MAYBE HUMAN MISTAKE." MULTIPOLE ATTEMPTS HAVE BEEN MADE REQUESTING FOR THE DETAILS OF THE EVENT. HOWEVER, NO FURTHER INFORMATION HAS BEEN PROVIDED TO DATE. Manufacturer Narrative: THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22306131&pc=FRN | 2016493-2025-89247 | 2025-05-27 04:00 | Death | CAREFUSION SD | 2025-06-23 04:00 | FRN | ALARIS SYSTEM | Excess Flow or Over-Infusion | Insufficient Information | K133532 |  | 1 | Event Description: IT WAS INITIALLY REPORTED AN OVER INFUSION EVENT THAT OCCURRED ON (B)(6) 2025. PER REPORT, THE PATIENT "DIED." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE DEVICE LOGS TO IDENTIFY WHICH MEDICATIONS WERE INFUSED FROM 3:00 PM ON (B)(6) 2025 TO 1:30 AM ON (B)(6) 2025. NO FURTHER INFORMATION WAS PROVIDED. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "THE REASON TO DO THIS (REVIEW OF DEVICE LOGS) ONLY TO CHECK THE MEDICATIONS AND THE RATE OF IT GIVEN BY THE MACHINE. I¿M NOT BELIEVE THAT THE DEVICE CAUSED THE DEATH BUT MAYBE HUMAN MISTAKE." MULTIPOLE ATTEMPTS HAVE BEEN MADE REQUESTING FOR THE DETAILS OF THE EVENT. HOWEVER, NO FURTHER INFORMATION HAS BEEN PROVIDED TO DATE. Manufacturer Narrative: THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22306400&pc=DSQ | 2916596-2025-04161 | 2025-06-08 04:00 | Death | THORATEC CORPORATION | 2025-06-23 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE PATIENT OUTCOME, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED BY CUSTOMER ACCORDING TO DATA PRIVACY LAW. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. THE DEVICE WILL NOT BE EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: SECTION A: PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22306608&pc=GDW | 1219930-2025-02688 | 2025-03-09 05:00 | Death | COVIDIEN LP LLC NORTH HAVEN | 2025-06-23 04:00 | GDW | UNKNOWN ENDO GIA SULU | Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Unintended Radiation Exposure |  |  | 1 | Event Description: ACCORDING TO THE LITERATURE, A RETROSPECTIVE STUDY INCLUDING 81 CASES OF LOCALLY ADVANCED ESOPHAGEAL SQUAMOUS CELL CARCINOMA UNDERGOING EITHER ROBOTIC OR CONVENTIONAL MINIMALLY INVASIVE ESOPHAGECTOMY WITH CURATIVE INTENT WAS COMPLETED BETWEEN 2018 AND 2022. ENDO GIA RADIAL RELOAD WITH TRISTAPLE TECHNOLOGY WAS USED TO CREATE THE GASTRIC CONDUIT. LINEAR ANASTOMOSES WERE ACHIEVED USING 45MM AND 60MM LINEAR ENDO GIA LOADS WITH TRISTAPLE TECHNOLOGY. COMPLICATIONS INCLUDED ANASTOMOTIC LEAKAGE. INTERVENTIONS LISTED INCLUDED COMPUTED TOMOGRAPHY SCANS OR OTHER RADIOLOGICAL IMAGING. OTHER COMPLICATIONS NOT DEVICE RELATED INCLUDE PNEUMONIA, ARRHYTHMIA, AND RECURRENT NERVE PALSY. ADDITIONALLY, IN-HOSPITAL DEATHS OCCURRED IN TWO CASES WITH REASONS UNSPECIFIED. SHORT-AND MIDDLE-TERM OUTCOMES OF ROBOT-ASSISTED MINIMALLY INVASIVE ESOPHAGECTOMY FOR HIGHLY LOCALLY ADVANCED ESOPHAGEAL CANCER WITH STAGE CT3 BORDERLINE AND CT4B AT INITIAL DIAGNOSIS SURGICAL ENDOSCOPY (2025) 39:2994¿3005 HTTPS://DOI.ORG/10.1007/S00464-025-11666-9 TAKEO FUJITA TAKFUJIT@EAST.NCC.GO.JP. Manufacturer Narrative: D10 CONCOMITANT PRODUCTS: UNKNOWN EGIA SU, UNKNOWN ENDO GIA SULU (LOT#:UNKNOWN) TAKEO FUJITA. "SHORT-AND MIDDLE-TERM OUTCOMES OF ROBOT-ASSISTED MINIMALLY INVASIVE ESOPHAGECTOMY FOR HIGHLY LOCALLY ADVANCED ESOPHAGEAL CANCER WITH STAGE CT3 BORDERLINE AND CT4B AT INITIAL DIAGNOSIS", 2025, HTTPS://DOI.ORG/10.1007/S00464-025-11666-9. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22307389&pc=NPT | 2015691-2025-05068 | 2025-05-31 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-23 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Perivalvular Leak; Adverse Event Without Identified Device or Use Problem | Rupture; Insufficient Information | P140031 |  | 1 | Event Description: AS REPORTED BY OUR MALAYSIAN AFFILIATES, DURING IMPLANT OF A 23MM SAPIEN 3 VALVE IN AORTIC POSITION BY TRANSFEMORAL APPROACH, AFTER VALVE PRE-DILATION WITH A 18MM NON-EDWARDS BALLOON AND DURING INSERTION OF THE COMMANDER DELIVERY SYSTEM (DS) WITH THE VALVE IN THE ESHEATH+ AND AFTER THE VALVE PASSED THE HEMOSTASIS VALVES, THE WIRE JUMPED OUT OF THE LEFT VENTRICLE DUE TO THE PATIENT SUDDENLY MOVING WHILE THE VALVE WAS PARTIALLY IN THE ESHEATH+. DURING THE WITHDRAWAL OF THE SYSTEM, WHILE ATTEMPTING TO RETRIEVE FROM THE ESHEATH+, THE VALVE MOVED TOWARDS THE TIP OF THE COMMANDER. THE SYSTEM WAS DISCARDED, AND A NEW ONE WAS PREPARED FOR REPLACEMENT. WITH THE SECOND SYSTEM, THE VALVE WAS POSITIONED TO AIM FOR AN 80/20 POSITION, HOWEVER THE VALVE WAS DEPLOYED WITH PACING LOSS OF CAPTURE NOTED. THE PATIENT'S BLOOD PRESSURE SIGNIFICANTLY DROPPED AND TORRENTIAL PARAVALVULAR LEAK (PVL) WAS NOTED. SUSPECTING PVL, POST-DILATION WAS ADDED WITH +1 CC (INITIAL INFLATION VOLUME WAS 23 + 1 CC), BUT AN ANNULAR/LEFT VENTRICULAR OUTFLOW TRACT (LVOT) RUPTURE WAS NOTED, WITH SIGNIFICANT PERICARDIAL EFFUSION. THE MEDICAL TEAM SUSPECTED THAT THE BLOOD PRESSURE DROP WAS CAUSED BY THE ANNULAR RUPTURE. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION AND COULD NOT BE RESCUED DESPITE CPR AND SHOCKS AND WAS DECLARED DEAD LATER THAT DAY. Manufacturer Narrative: THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE A LOSS OF PACING CAPTURE COULD HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. PER THE INSTRUCTIONS FOR USE (IFU), PARAVALVULAR LEAK (PVL) IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE AND THE LANDING ZONE, AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE. SOME PVL IS NOT UNCOMMON POST-DEPLOYMENT. MANY CASES ARE MILD TO MODERATE, AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS. OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION. THE PATIENT SCREENING MANUAL AND THE PROCEDURE DIDACTIC IDENTIFY SEVERAL PROCEDURAL AND ANATOMICAL FACTORS WHICH COULD CONTRIBUTE TO PVL, INCLUDING DEVICE MALPOSITION, INACCURATE MEASUREMENT OF THE LANDING ZONE, UNEVEN DISTRIBUTION OF CALCIUM ON THE LANDING ZONE, BULKY OR SEVERE CALCIFICATION, AN ELLIPTICAL ANNULUS SHAPE, AND VALVE UNDER-SIZING. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). THE DEVICE TRAINING MANUALS INSTRUCTS THE OPERATOR ON PROPER IMAGING SCREENING REQUIREMENTS, INCLUDING THE USE OF GOOD QUALITY ECHOCARDIOGRAPHY AND/OR COMPUTED TOMOGRAPHY (CT) TO APPROPRIATELY MEASURE THE LANDING ZONE, ASSESS THE CONTENT AND DISTRIBUTION OF CALCIUM, AND OTHER PATIENT ANATOMICAL FACTORS. MULTIPLE IMAGING MODALITIES SHOULD BE CONSIDERED DURING VALVE SIZE SELECTION. CONTRAINDICATIONS, IMPORTANT CONSIDERATIONS WHEN ASSESSING THE VALVE, AND CHOOSING THE PROPER TRANSCATHETER HEART VALVE ARE ALSO DISCUSSED. THE DEVICE TRAINING MANUALS ALSO INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING PATIENTS SCREENING AND PROCEDURAL CONSIDERATIONS. DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES ARE ALSO INCLUDED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE THE ANNULAR INJURY CAUSED BY THE PACING MISCAPTURE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR A CORRECTION TO THE INVESTIGATION FINDINGS AND NARRATIVE: H6 AND H11. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES ON CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. PER THE INSTRUCTIONS FOR USE (IFU), PARAVALVULAR LEAK (PVL) IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BIOPROSTHETIC HEART VALVES AND THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. PARAVALVULAR LEAK REFERS TO BLOOD FLOWING THROUGH A CHANNEL BETWEEN THE STRUCTURE OF THE IMPLANTED VALVE AND THE LANDING ZONE, AS A RESULT OF A LACK OF APPROPRIATE SEALING OF THE VALVE TO THE TARGET SITE. SOME PVL IS NOT UNCOMMON POST-DEPLOYMENT. MANY CASES ARE MILD TO MODERATE, AND EITHER RESOLVE OVER TIME OR DO NOT CAUSE SYMPTOMS. OTHERS MAY BE MORE CLINICALLY SIGNIFICANT AND REQUIRE INTERVENTION. THE PATIENT SCREENING MANUAL AND THE PROCEDURE DIDACTIC IDENTIFY SEVERAL PROCEDURAL AND ANATOMICAL FACTORS WHICH COULD CONTRIBUTE TO PVL, INCLUDING DEVICE MALPOSITION, INACCURATE MEASUREMENT OF THE LANDING ZONE, UNEVEN DISTRIBUTION OF CALCIUM ON THE LANDING ZONE, BULKY OR SEVERE CALCIFICATION, AN ELLIPTICAL ANNULUS SHAPE, AND VALVE UNDER-SIZING. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVES (ALL MODELS). THE DEVICE TRAINING MANUALS INSTRUCTS THE OPERATOR ON PROPER IMAGING SCREENING REQUIREMENTS, INCLUDING THE USE OF GOOD QUALITY ECHOCARDIOGRAPHY AND/OR COMPUTED TOMOGRAPHY (CT) TO APPROPRIATELY MEASURE THE LANDING ZONE, ASSESS THE CONTENT AND DISTRIBUTION OF CALCIUM, AND OTHER PATIENT ANATOMICAL FACTORS. MULTIPLE IMAGING MODALITIES SHOULD BE CONSIDERED DURING VALVE SIZE SELECTION. CONTRAINDICATIONS, IMPORTANT CONSIDERATIONS WHEN ASSESSING THE VALVE, AND CHOOSING THE PROPER TRANSCATHETER HEART VALVE ARE ALSO DISCUSSED. THE DEVICE TRAINING MANUALS ALSO INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING PATIENTS SCREENING AND PROCEDURAL CONSIDERATIONS. DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES ARE ALSO INCLUDED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS THAT PROCEDURAL FACTORS (PACING MISCAPTURE CONTRIBUTING TO THE ANNULAR RUPTURE AND THIS RUPTURE TO THE PVL) CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22307727&pc=MAL | 1640201-2025-0000013 | 2025-05-30 04:00 | Death | INTERVASCULAR SAS | 2025-06-23 04:00 | MAL | HEMASHIELD PLATINUM | Off-Label Use | Hemorrhage/Bleeding | K021213 |  | 1 | Event Description: A SURVEY WAS PERFORMED BY EXTRACORPOREAL LIFE SUPPORT ORGANIZATION (ELSO) REGARDING THE USE OF VASCULAR GRAFTS WITH GRAFT-FACILITATED ACCESS FOR VEINO ATERIAL-EXTRACOPOREAL MEMBRAN OXYGENATION (VA-ECMO) AND MECHANICAL CIRCULATORY SUPPORT (MCS). AS PART OF THESE TYPE OF PROCEDURES, ADVERSE EVENTS WERE REPORTED. THIS REPORT RELATES THE CASES OF DEATH MENTIONED IN THE SURVEY. BELOW IS A SUMMARY OF THE SURVEY. OBJECTIVE OF THE STUDY: THE FOCUS OF THIS SURVEY WAS ON THE USE OF SIDE GRAFTS ANASTOMOSED TO THE AXILLARY OR SUBCLAVIAN ARTERY AS PORT ACCESS FOR ARTERIAL PERFUSION IN V-A ECLS OR FOR IMPELLA OR FOR A COMBINATION OF SUCH DEVICES (REFERRED TO AS ECPELLA OR ECMELLA). EVENT DESCRIPTION: THE SURVEY DESCRIBED THAT MOST RESPONDENTS FOR BOTH GLOBALLY AND WITHIN THE US HAD EXPERIENCED AN ADVERSE EVENT AT THE GRAFT SITE. THIS REPRESENTED 53.2% GLOBALLY AND 64.3% IN THE US OF THE 62 PEOPLE WHO RESPOND TO THE QUESTIONS. GLOBALLY, 8.3% AND 15.8% IN US OF THE 36 RESPONDENTS, EXPERIENCED BLEEDING EVENTS WHICH LED TO THE DEATH OF THE PATIENT. GLOBALLY, IT REPRESENTS APPROXIMATELY 3 DEATHS. TO BE NOTED THAT DATA FROM THIS SURVEY ARE AGGREGATED DATA FROM A VOLUNTARY SURVEY, WITH NO PATIENT-INDIVIDUAL CLINICAL DATA, NO VERIFICATION OF FITTING COMPLIANCE, AND NO COMPARATIVE ANALYSIS BY PRODUCT TYPE. THERE IS NO CORRELATION BETWEEN A SPECIFIC GRAFT AND REPORTED ADVERSE EVENTS. COMPLAINT#: (B)(4). THIS REPORT IS THE FIRST IN A SERIES OF 4. Manufacturer Narrative: (4111/3233) ATTEMPTS TO CONTACT THE SURVEY AUTHOR TO GET MORE INFORMATION OF THE ADVERSE EVENTS REPORTED AS PART OF THE SURVEY WILL BE PERFORMED. (11) THE INVESTIGATION IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE SENT UPON COMPLETION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22307960&pc=MHX | 9610816-2025-000444 | 2025-05-25 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-06-23 04:00 | MHX | INTELLIVUE MULTI-MEASUREMENT MODULE X3 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K171801 |  | 1 | Event Description: IT WAS REPORTED FOLLOWING THE CODE OF A DNR PATIENT, IT WAS INDICATED THE MONITOR WAS NOT FUNCTIONING CORRECTLY. THE CAUSE OF DEATH OF THIS DNR PATIENT REMAINS UNKNOWN. THE ENGINEER WAS ONSITE AT THE TIME AND REVIEWED DATA AT THE CENTRAL STATION, REVEALING EXTREME BRADYCARDIA AND ASYSTOLE ALARMS. THERE WAS ARTIFACT ON SOME WAVEFORMS, LIKELY RELATED TO PATIENT MOVEMENT DUE TO AGITATION. THE NURSE HAD INDICATED THE MONITOR WAS NOT WORKING RIGHT BUT LATER CLARIFIED THERE WERE TOO MANY ALARMS, AND THE REPORTED MONITOR ISSUE DID NOT AFFECT THE OUTCOME IN ANY WAY. THE STAFF BEGAN CPR ON THE PATIENT BUT HALTED CPR AFTER AN OLD DNR ORDER WAS LOCATED IN THE CHART. Manufacturer Narrative: A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE FOLLOWING FUNCTIONAL/DIAGNOSTIC TESTING AND COMMUNICATIONS WERE PERFORMED; THE PHILIPS CLINICAL CONSULTANT STATED THE CAUSE OF DEATH OF THIS DNR PATIENT REMAINS UNKNOWN. THE ENGINEER WAS ONSITE AT THE TIME AND REVIEWED DATA AT THE CENTRAL STATION, REVEALING EXTREME BRADYCARDIA AND ASYSTOLE ALARMS. THERE WAS ARTIFACT ON SOME WAVEFORMS, LIKELY RELATED TO PATIENT MOVEMENT DUE TO AGITATION. THE NURSE HAD INDICATED THE MONITOR WAS NOT WORKING RIGHT BUT LATER CLARIFIED THERE WERE TOO MANY ALARMS, AND THE REPORTED MONITOR ISSUE DID NOT AFFECT THE OUTCOME IN ANY WAY. THE STAFF BEGAN CPR ON THE PATIENT BUT HALTED CPR AFTER AN OLD DNR ORDER WAS LOCATED IN THE CHART. REVIEW OF THE DATA REVEALED NO MALFUNCTION. IT WAS ALSO CONFIRMED FROM THE PHILIPS TECHNICAL CONSULTANT REVIEWED THE ALARM WAVEFORM THAT WAS AVAILABLE AND STATED THAT THE ALARMS WERE NOT ARTIFACT ALARMS. BASED ON THIS INFORMATION, THERE DOES NOT APPEAR TO BE ANY DEVICE MALFUNCTION; HOWEVER, A COMBINATION OF ALARM MANAGEMENT AND CLINICAL WORKFLOW DURING USE OF THE DEVICE MAY HAVE BEEN FACTORS IN THE PATIENT OUTCOME. BASED ON THE INFORMATION AVAILABLE THE CAUSE OF THE REPORTED PROBLEM WAS UNABLE TO BE DETERMINED. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE REMAINS IN USE AT THE CUSTOMER'S SITE. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22308286&pc=PNJ | 2017865-2025-87878 | 2025-05-29 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-23 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Discomfort; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR A SCHEDULED PROCEDURE. DURING THE PROCEDURE, FIXATION WAS ATTEMPTED MULTIPLE TIMES DUE TO LOSS OF CAPTURE. DURING FIXATION, THERE WAS A RAPID INCREASE IN IMPEDANCE AND ST SEGMENT BECAME NEGATIVE INDICATING POTENTIAL RV PERFORATION. SHORTLY AFTER THE PROCEDURE, THE PATIENT WAS EXPERIENCING CHEST PAIN AND THEN STARTED TO HAVE A HEMODYNAMIC COLLAPSE. A PERICARDIOCENTESIS WAS PERFORMED, THE PATIENT WAS UNABLE TO BE RESUSCITATED, AND THE PATIENT PASSED SHORTLY AFTER. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22308472&pc=DSQ | 3007042319-2025-01307 | 2025-06-16 04:00 | Death | HEARTWARE, INC. | 2025-06-23 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Increase in Suction; Power Problem; Pumping Problem | Arrhythmia; Dehydration; Tachycardia; Loss of consciousness; Multiple Organ Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS IN THE CLINIC FOR A ROUTINE APPOINTMENT, AND A LOW FLOW ALARM ON THE VENTRICULAR ASSIST DEVICE (VAD) WAS NOTED DURING LOG FILE REVIEW. IT WAS ALSO NOTED THAT THERE WERE AREAS OF POTENTIAL SUCTION. IT WAS SUSPECTED THAT THE PATIENT HAD ARRHYTHMIAS AND SOME DEHYDRATION THAT CAUSED THE LOW FLOWS AND SUCTION. THE VAD REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT SUPPORT WAS DISCONTINUED DUE TO MULTI ORGAN SYSTEM FAILURE AND THE PATIENT SUBSEQUENTLY EXPIRED. WHEN THE PATIENT WAS FOUND DOWN THEIR BODY TEMPERATURE ON ARRIVAL TO THE EMERGENCY ROOM WAS 108 DEGREES FAHRENHEIT AND THERE WERE SECOND DEGREE BURNS ON THEIR BODY FROM BEING OUTSIDE FOR AN UNKNOWN PERIOD OF TIME. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL INFORMATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ANALYSIS AND INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) AND ONE (1) CONTROLLER (B)(6) WERE NOT RETURNED FOR EVALUATION AS THEY REMAIN IN USE. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO MULTI ORGAN SYSTEM FAILURE. REVIEW OF CONTROLLER LOG FILES REVEALED TRANSIENT DECREASES IN POWER CONSUMPTION AND ESTIMATED FLOWS ALONG WITH SUCTION EVENTS STARTING ON 16/JUN/2025 ACCOMPANIED BY EIGHT (8) LOW FLOW ALARMS LOGGED SINCE ON (B)(6) 2025. REVIEW OF THE CONTROLLER EVENT FILE REVEALED ONE (1) CONTROLLER POWER-UP EVENT WITH AN ASSOCIATED MOTOR START EVENT LOGGED ON 5/JUL/2025 AT 17:57:41, INDICATING A LOSS OF POWER TO THE CONTROLLER. THE DATA POINT LOGGED PRIOR TO THE LOSS OF POWER REVEALED THAT NO POWER WAS CONNECTED TO POWER PORT ONE (1) AND THE ADAPTER WAS CONNECTED TO POWER PORT TWO (2). THE DATA POINT RECORDED AFTER THE LOSS OF POWER REVEALED THE ADAPTER WAS CONNECTED TO POWER PORT ONE (1) AND THAT (B)(6) WAS CONNECTED TO POWER PORT TWO (2). THE CONTROLLER WAS WITHOUT POWER FOR EIGHT (8) SECONDS. NO ANOMALIES WERE RECORDED LEADING UP TO THE LOSS OF POWER. AS A RESULT, THE REPORTED EVENTS WERE CONFIRMED. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE RISK DOCUMENTATION, POSSIBLE CAUSES OF THE REPORTED LOW FLOW AND SUCTION EVENTS MAY BE ATTRIBUTED TO MULTIPLE FACTORS INCLUDING BUT NOT LIMITED TO THROMBUS AT THE INFLOW CANNULA/OUTFLOW GRAFT, CONSTRICTION AT THE OUTFLOW GRAFT, POOR VAD FILLING, AND/OR INAPPROPRIATE PUMP ROTATIONAL SPEED. THE MOST LIKELY ROOT CAUSE OF THE LOSS OF POWER TO THE CONTROLLER CAN BE ATTRIBUTED TO A DISCONNECTION OF BOTH POWER SOURCES AND/OR TO AN INTERMITTENT DISCONNECTION ON ONE OR BOTH POWER SOURCES. CAPA: PR00551638 INVESTIGATED CONTROLLER LOSSES OF POWER. EVEN THOUGH THIS CAPA IS NOW CLOSED, (B)(6) FALLS IN SCOPE OF THIS CAPA. PER THE INSTRUCTIONS FOR USE, CARDIAC ARRHYTHMIA, MULTI-ORGAN FAILURE, AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. ADDITIONAL PRODUCT: D1: HEARTWARE VENTRICULAR ASSIST SYSTEM ¿ CONTROLLER 2.0 D4: SERIAL#: (B)(6), H3: YES, H6: THE CODES PRESENT IN SECTION H6 CORRESPOND TO COMPONENTS/PRODUCTS THAT COMPRISE THE REPORTED EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22308516&pc=DSQ | 2916596-2025-03898 | 2025-05-30 04:00 | Death | THORATEC CORPORATION | 2025-06-23 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Multiple Organ Failure; Drug Resistant Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE PATIENT DEVELOPED SEPSIS THEN MULTIPLE ORGAN FAILURE. THE HEALTHCARE PROVIDER BELIEVED THIS WAS NOT RELATED TO THE DEVICE. THE ONSET DATE OF INFECTION WAS NOT PROVIDED BY THE HEALTHCARE PROVIDER. THE PATIENT WAS FOUND TO HAVE INFECTION IN BOTH SPUTUM AND BLOOD. THE CULTURE IDENTIFIED IN BOTH WAS MULTIDRUG-RESISTANT KLEBSIELLA PNEUMONIAE (MDCKP). IT WAS NOT PROVIDED HOW INFECTION WAS TREATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER: (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. B ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, SEPSIS, INFECTION, AND MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, AND HEPATIC DYSFUNCTION), THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION, AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22308696&pc=OZD | 1220648-2025-29703 | 2025-06-06 04:00 | Death | ABIOMED, INC. | 2025-06-23 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT HAD VENTRICULAR TACHYCARDIA (VT)/VENTRICULAR FIBRILLATION (VF). THE PATIENT HAD SIX (6) EPISODES, ALL REQUIRING SHOCKS AND CARDIOPULMONARY RESUSCITATION. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22308760&pc=NVZ | 2017865-2025-88982 | 2025-06-17 04:00 | Death | ABBOTT | 2025-06-23 04:00 | NVZ | ENDURITY¿ CORE DR, JPD | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P880086 |  | 1 | Event Description: IT WAS REPORTED, THE PATIENT HAD PASSED AWAY. IT IS UNKNOWN IF THE DEATH IS RELATED TO THE ABBOTT DEVICE. THE DEVICE WAS EXPLANTED AT THE REQUEST OF LOCAL AUTHORITIES. ADDITIONAL INFORMATION WAS REQUESTED BUT WAS UNABLE TO BE PROVIDED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22308861&pc=OZD | 1220648-2025-29704 | 2025-06-06 04:00 | Death | ABIOMED, INC. | 2025-06-23 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT CODED FOR ABOUT ONE MINUTE. CPR STARTED AND ONE OF EPI GIVEN BEFORE RETURN OF SPONTANEOUS CIRCULATION (ROSC) WAS ACHIEVED. LEFT MAIN RECEIVED 1 DRUG-ELUTING STENT (DES) AND CX 1 DES. THE PATIENT'S URINE WAS RED AND ALL LABS HEMOLYZING. THE PUMP WAS REPOSITIONED, AND P LEVEL WAS REDUCED. ADDITIONALLY, THE PATIENT WENT BACK INTO AFIB WITH RAPID VENTRICULAR RESPONSE (RVR), REQUIRING INCREASED IMPELLA SUPPORT. THE PATIENT REQUIRED INCREASED PRESSOR, HAVING OXYGENATION ISSUES. THE CARE WAS ULTIMATELY WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA CP WITH SMARTASSIST SYSTEM INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22308879&pc=MCX | 3004742232-2025-00176 | 2025-05-31 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-06-23 04:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Obstruction/Occlusion; Embolism/Embolus | P130005 |  | 1 | Event Description: IT WAS REPORTED AN ORBITAL ATHERECTOMY PROCEDURE WAS BEING PREPARED BY FIRST TAKING TPI AND INTRA-AORTIC BALLOON PUMP (IABP) SUPPORT. THE LESION WAS CROSSED WITH A NON-ABBOTT GUIDE WIRE. THE LESION WAS PRE-DILATED WITH A NC 1.2 X 12 MM BALLOON AND UPGRADED TO A 2 X 12 MM BALLOON. THE NON-ABBOTT GUIDE WIRE WAS EXCHANGED FOR A VIPERWIRE ADVANCE GUIDE WIRE WITH FLEX TIP IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD). THE DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS ADVANCED AND FIVE TREATMENTS WERE PERFORMED. POST TREATMENT, SLOW FLOW WAS OBSERVED. SUDDENLY, THE PATIENT'S BLOOD PRESSURE DROPPED. THE PATIENT WAS INTUBATED AND SENT FOR CORONARY ARTERY BYPASS GRAFT SURGERY (CABG). THE PARTICULATE SIZE POST ORBITAL ATHERECTOMY LED TO THE SLOW FLOW, WHICH SUBSEQUENTLY LED TO THE DROP IN BLOOD PRESSURE. CABG WAS NOT AGREED TO BY THE PATIENT, AND THEY WERE SENT FOR HRPTCA. THE SURGERY WAS SUCCESSFUL, HOWEVER, THE PATIENT EXPIRED.Event Description: IT WAS REPORTED AN ORBITAL ATHERECTOMY PROCEDURE WAS BEING PREPARED BY FIRST TAKING TPI AND INTRA-AORTIC BALLOON PUMP (IABP) SUPPORT. THE LESION WAS CROSSED WITH A NON-ABBOTT GUIDE WIRE. THE LESION WAS PRE-DILATED WITH A NC 1.2 X 12 MM BALLOON, AND UPGRADED TO A 2 X 12 MM BALLOON. THE NON-ABBOTT GUIDE WIRE WAS EXCHANGED FOR A VIPERWIRE ADVANCE GUIDE WIRE WITH FLEX TIP IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD). THE DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS ADVANCED AND FIVE TREATMENTS WERE PERFORMED. POST TREATMENT, SLOW FLOW WAS OBSERVED. SUDDENLY, THE PATIENT'S BLOOD PRESSURE DROPPED. THE PATIENT WAS INTUBATED AND SENT FOR CORONARY ARTERY BYPASS GRAFT SURGERY (CABG). NO ADDITIONAL INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING PARTICULATE SIZE POST ORBITAL ATHERECTOMY LED TO THE SLOW FLOW, WHICH SUBSEQUENTLY LED TO THE DROP IN BLOOD PRESSURE. CABG WAS NOT AGREED TO BY THE PATIENT, AND THEY WERE SENT FOR HRPTCA. THE SURGERY WAS SUCCESSFUL, HOWEVER, THE PATIENT EXPIRED. IN THE PHYSICIAN'S OPINION, THE PATIENT DEATH WAS PRIMARILY DUE TO SLOW FLOW IN THE LAD. THE SECONDARY CAUSE WAS NO SUPPORT IN THE RIGHT CORONARY ARTERY (RCA) AND LEFT CIRCUMFLEX ARTERY (LCX). IT WAS INDICATED ORBITAL ATHERECTOMY MAY HAVE CONTRIBUTED TO THE PATIENT DEATH. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED PATIENT HARMS INCLUDING EMBOLISM, OBSTRUCTION, LOW BLOOD PRESSURE, SURGICAL INTERVENTION AND DEATH APPEAR TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE IT IS POSSIBLE THAT THE REPORTED PATIENT EMBOLISM, OBSTRUCTION, LOW BLOOD PRESSURE, SURGICAL INTERVENTION AND DEATH ARE A RESULT OF THE PATIENT¿S DISEASE SEVERITY COMBINED WITH USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. UPDATED FIELD: B5, H6 (CODES 4118, 3233, AND 11 NO LONGER APPLY). Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22309272&pc=NPT | 2015691-2025-05081 | 2025-05-29 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-23 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED BY OUR AFFILIATES IN THE UNITED KINGDOM, THIS WAS A CASE WITH A 23MM SAPIEN 3 ULTRA RESILIA VALVE, IN AORTIC POSITION BY TRANSFEMORAL APPROACH. DURING THE PROCEDURE, THE VALVE WAS IMPLANTED AS INTENDED, WITH AN ADDITIONAL 2 ML OF VOLUME, AND SLIGHTLY HIGHER AS IT WAS A FUNCTIONAL BICUSPID VALVE. HOWEVER, THE POST-DEPLOYMENT AORTOGRAM REVEALED CENTRAL AORTIC REGURGITATION, AND THE PATIENT DEVELOPED ST-SEGMENT ELEVATION AND CHEST PAIN. A DECISION WAS MADE TO IMPLANT A 26 MM SAPIEN 3 ULTRA RESILIA VALVE, WHICH WAS DEPLOYED AS INTENDED WITH 2 ML LESS THAN NOMINAL VOLUME. THIS RESULTED IN TRIVIAL AORTIC REGURGITATION. HOWEVER, ECHOCARDIOGRAPHY REVEALED A PERICARDIAL EFFUSION, PROMPTING PERICARDIAL DRAINAGE. THE PATIENT CONDITION WAS CLOSELY MONITORED BY THE ANESTHESIOLOGIST. AT THIS STAGE, THE LEFT MAIN CORONARY ARTERY WAS ASSESSED TO RULE OUT SEQUESTRATION OR OCCLUSION. THE LEFT MAIN STEM (LMS) WAS FOUND TO BE PATENT. DESPITE THIS, THE PATIENT CONTINUED TO BLEED, AND IMAGING REVEALED ANNULAR/LEFT VENTRICULAR OUTFLOW TRACT (LVOT) RUPTURE. SURGICAL BACKUP WAS CONSULTED BUT DETERMINED THAT THE PATIENT WAS NOT A CANDIDATE FOR SURGICAL BAILOUT. A SECOND 26 MM SAPIEN 3 ULTRA RESILIA VALVE WAS IMPLANTED AT A DEEPER POSITION TO SEAL THE RUPTURE WITHOUT SUCCESS. FOLLOWING THIS, THE PATIENT'S SYSTOLIC BLOOD PRESSURE DROPPED TO 45 MMHG. RESUSCITATION WAS UNSUCCESSFUL, AND TIME OF DEATH WAS DECLARED 25 MINUTES LATER. THE CAUSE OF DEATH WAS THE RUPTURE. THE PERCEIVED ROOT CAUSE OF THE CENTRAL REGURGITATION WAS THE HEAVILY CALCIFIED VALVE. THE PERCEIVED ROOT CAUSE OF THE RUPTURE WAS THE USE OF PREDNISONE, HEAVILY CALCIFIED VALVE AND FRAIL PATIENT. Manufacturer Narrative: THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. PLEASE REFERENCE RELATED MANUFACTURER REPORT NO: 2015691-2025-05029. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES ON CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE PERCEIVED ROOT CAUSE OF THE RUPTURE WAS THE USE OF PREDNISONE, HEAVILY CALCIFIED VALVE AND FRAIL PATIENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22309297&pc=LWS | 2124215-2025-41180 | 2025-05-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-23 04:00 | LWS | VIGILANT? EL ICD VR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960040 |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS IMPLANTED ON (B)(6) 2025 AND THE PATIENT WAS DISCHARGED TO HOME AFTER THE PROCEDURE. A LATITUDE TRANSMISSION WAS PERFORMED THAT DAY WITH NO DEVICE ISSUES OBSERVED. ON THE NEXT DAY, THE PATIENT WAS FOUND DECEASED. THE PATIENT'S PHYSICIAN WAS NOT NOTIFIED ABOUT THE DEATH UNTIL A WEEK LATER AND HAD NO INFORMATION ABOUT THE CAUSE OF DEATH; IT WAS SUSPECTED THAT THE DEVICE WAS BURIED WITH THE PATIENT BEFORE THE DEVICE PHYSICIAN WAS MADE AWARE, SO NO POST-MORTEM DEVICE INTERROGATION WAS PERFORMED. NO PROCEDURE-RELATED COMPLICATIONS WERE REPORTED, BUT WITH NO FURTHER INFORMATION, BOSTON SCIENTIFIC IS UNABLE TO RULE OUT THE IMPLANT PROCEDURE AS A CONTRIBUTORY CAUSE OF THE PATIENT'S DEATH. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22309493&pc=NPT | 2015691-2025-05083 | 2025-05-29 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-23 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Difficult to Advance |  | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING A TRANSFEMORAL TAVR PROCEDURE USING A 23 MM COMMANDER DELIVERY SYSTEM, THE CTS INITIATED THE CASE AND SUCCESSFULLY OBTAINED ACCESS AT ALL REQUIRED POINTS. THE 14FR ESHEATH WAS INTRODUCED WITHOUT ISSUE. THE 23 MM SAPIEN VALVE WAS CROSSED, AND A STIFF WIRE WAS POSITIONED IN THE VENTRICLE. THE COMMANDER DELIVERY SYSTEM AND VALVE WERE ADVANCED THROUGH THE ESHEATH WITH PROPER ALIGNMENT. WHILE THE CTS WAS CONTROLLING THE COMMANDER, FLEXION WAS APPLIED, BUT RESISTANCE WAS ENCOUNTERED AT THE AORTIC ARCH JUST PAST THE RIGHT SUBCLAVIAN ARTERY. THE CTS CONTINUED TO ADVANCE THE SYSTEM, WHICH THEN ABRUPTLY JUMPED FORWARD. THE VALVE WAS POSITIONED ACROSS THE ANNULUS AND ALIGNED IN THE COPLANAR VIEW. AT THIS POINT, THE PATIENT'S BLOOD PRESSURE DROPPED TO APPROXIMATELY 60/30 MMHG. THE VALVE WAS DEPLOYED UNDER RAPID PACING WITHOUT COMPLICATION, ACHIEVING A 90/10 POSITION WITH NO PARAVALVULAR LEAK AND A MEAN GRADIENT OF 3 MMHG. THE PATIENT RECEIVED FLUIDS AND MEDICATIONS TO STABILIZE BLOOD PRESSURE. A HEMOGLOBIN CHECK REVEALED A DROP FROM APPROXIMATELY 12 G/DL TO 8 G/DL. THE INTERVENTIONAL CARDIOLOGIST (IC) AND CTS REVIEWED ANGIOGRAPHIC IMAGES AND SUSPECTED THE VALVE MAY HAVE ENCOUNTERED PLAQUE WHILE NAVIGATING THE ARCH. DUE TO THE HEMOGLOBIN DROP AND ONGOING CONCERN, A CT SCAN WAS ORDERED. AFTER THE PATIENT WAS TRANSFERRED FOR IMAGING, FURTHER REVIEW OF THE ANGIOGRAMS REVEALED A DISSECTION IN THE AORTIC ARCH, LIKELY OCCURRING AS THE SYSTEM PASSED THROUGH AND BEFORE VALVE DEPLOYMENT. THE CT CONFIRMED A POSITIVE DISSECTION IN THE ARCH. THE PATIENT REMAINED HEMODYNAMICALLY STABLE WITH NO EVIDENCE OF ROOT LEAKAGE. HOWEVER, LOCALIZED BLOOD LOSS WAS NOTED NEAR THE RIGHT GROIN ACCESS SITE. PRESSURE WAS MAINTAINED AT THE SITE, AND GIVEN THE PATIENT'S STABLE CONDITION, NO FURTHER INTERVENTION IS PLANNED AT THIS TIME. PER FEEDBACK FROM THE FCS, THE PATIENT EXPIRED ON THE SAME DAY AS THE PROCEDURE. NO DAMAGE WAS OBSERVED TO THE VALVE OR DELIVERY SYSTEM AS IT PASSED OVER THE ARCH. THE PERCEIVED ROOT CAUSE OF THE BLOOD LOSS IN THE GROIN AREA APPEARS TO BE MULTIPLE NEEDLE STICKS DURING ATTEMPTS TO GAIN ACCESS. THERE IS NO INDICATION THAT AN EDWARDS DEVICE CONTRIBUTED TO THE BLOOD LOSS. THE FCS IS UNAWARE OF ANY BLOOD PRODUCTS BEING ADMINISTERED. THE CAUSE OF DEATH REMAINS UNKNOWN AT THIS TIME. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED FOR APPROXIMATELY 30 MINUTES BEFORE THE PATIENT WAS PRONOUNCED DECEASED. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR THE COMPLETED ENGINEERING EVALUATION. AN ADDITION WAS MADE TO SECTION B.1 AND H.11. SECTIONS B.4, G.3, G.6, AND H.2 HAVE BEEN UPDATED. CORRECTIONS HAVE BEEN MADE TO H.6: CLINICAL CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSION. THE EVENT REPORTED IS ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THE INSTRUCTIONS FOR USE (IFU), AND TRAINING MANUALS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, DIFFICULTY OR INABILITY IN TRACKING THE DELIVERY SYSTEM TO THE LANDING ZONE, INCLUDING TRAVERSING THE AORTIC ARCH AND CROSSING THE NATIVE ANNULUS/BIOPROSTHESIS HAS NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAS HISTORICALLY BEEN DUE TO PATIENT FACTORS, POOR GUIDEWIRE MANAGEMENT, DAMAGED GUIDEWIRE, AND/OR EXCESSIVE MANIPULATION OF THE FLEX WHEEL. THE COMPLAINT FOR DIFFICULT TO TRACK SYSTEM THROUGH ANATOMY WAS CONFIRMED EMPIRICALLY. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (PLAQUE) LIKELY CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT ''THE VALVE MAY HAVE ENCOUNTERED PLAQUE WHILE NAVIGATING THE ARCH.'' PATIENT FACTORS (I.E. CALCIFICATION) MAY CAUSE CONSTRAINED SECTIONS OF THE ANATOMY THAT INTERFERE WITH PASSAGE OF THE DELIVERY SYSTEM AND CAUSE DIFFICULTY DURING TRACKING/CROSSING. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22309525&pc=DTB | 2017865-2025-88992 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-23 04:00 | DTB | PERMANENT PACEMAKER ELECTRODE | Adverse Event Without Identified Device or Use Problem | Unspecified Infection |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS LEAD RELATED. THE CAUSE OF DEATH WAS PYELONEPHRITIS WITH SEPSIS. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: D6B SHOULD BE (B)(6) 1991. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22309605&pc=LWP | 2017865-2025-88993 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-23 04:00 | LWP | SOLUS | Adverse Event Without Identified Device or Use Problem | Unspecified Infection | P880086 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAS DECEASED. THERE IS NO KNOWN ALLEGATION FROM A HEALTH CARE PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS PYELONEPHRITIS WITH SEPSIS. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22309893&pc=NPT | 2025587-2025-04407 | 2025-06-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-23 04:00 | NPT | EVOLUT PRO PLUS VALVE | Failure to Advance; Patient Device Interaction Problem | Perforation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC VALVE, THE VALVE WAS RELEASED AT THE WRONG TIME AND CAUSED A BAD IMPLANT POSITION. WHEN REMOVING THE PIGTAIL CATHETER, THE VALVE WAS PUSHED AND CAUSED THE VALVE TO DISLODGE UP. A SECOND VALVE WAS DEPLOYED, THE VALVE WAS RAISED WITH THE DELIVERY CATHETER SYSTEM (DCS) AND BECAME HOOKED ON THE FIRST VALVE, CAUSING IT TO BE PUSHED AND THE AORTA BECAME PERFORATED. THE PATIENT SUBSEQUENTLY DIED DUE TO THE AORTIC PERFORATION. ADDITIONAL INFORMATION WAS RECEIVED THAT THE SECOND VALVE WAS NOT IMPLANTED AS IT NEVER PASSED THE AORTIC ROOT. PER THE PHYSICIAN, THE SECOND VALVE PUSHED THEFIRST VALVE IN THE AORTA. THERE WAS NO TREATMENT FOR THE PERFORATION PRIOR TO THE DEATH.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC VALVE, THE VALVE WAS RELEASED AT THE WRONG TIME AND CAUSED A BAD IMPLANT POSITION. WHEN REMOVING THE PIGTAIL CATHETER, THE VALVE WAS PUSHED AND CAUSED THE VALVE TO DISLODGE UP. A SECOND VALVE WAS DEPLOYED, THE VALVE WAS RAISED WITH THE DELIVERY CATHETER SYSTEM (DCS) AND BECAME HOOKED ON THE FIRST VALVE, CAUSING IT TO BE PUSHED AND THE AORTA BECAME PERFORATED. THE PATIENT SUBSEQUENTLY DIED DUE TO THE AORTIC PERFORATION. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVPROP23-29 (LOT: 0012663964); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE ; EXPLANT DATE PRODUCT ID L-EVPROP23-29 (LOT: 0012623710); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE ; EXPLANT DATE PRODUCT ID L-EVPROP23-29 (LOT: 0012483153); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE ; EXPLANT DATE PRODUCT ID EVPROPLUS-29 ((B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT PRO PLUS DCS; PRODUCT ID D-EVPROP23-29 (LOT: 0012617743); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE; EXPLANT DATE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: PRO+ DELIVERY CATHETER SYSTEM (DCS), PRODUCT ID: D-EVPROP23-29, LOT: 0012617743, USE BY DATE: 2027-01-09, UDI: (B)(4). UPDATED: B5, D4, H6, H11. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED D.4 AND H.4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22309895&pc=NPT | 2025587-2025-04406 | 2025-06-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-23 04:00 | NPT | EVOLUT PRO PLUS VALVE | Device Dislodged or Dislocated; Patient Device Interaction Problem | Perforation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC VALVE, THE VALVE WAS RELEASED AT THE WRONG TIME AND CAUSED A BAD IMPLANT POSITION. WHEN REMOVING THE PIGTAIL CATHETER, THE VALVE WAS PUSHED AND CAUSED THE VALVE TO DISLODGE UP. A SECOND VALVE WAS DEPLOYED, THE VALVE WAS RAISED WITH THE DELIVERY CATHETER SYSTEM (DCS) AND BECAME HOOKED ON THE FIRST VALVE, CAUSING IT TO BE PUSHED AND THE AORTA BECAME PERFORATED. THE PATIENT SUBSEQUENTLY DIED DUE TO THE AORTIC PERFORATION.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC VALVE, THE VALVE WAS RELEASED AT THE WRONG TIME AND CAUSED A BAD IMPLANT POSITION. WHEN REMOVING THE PIGTAIL CATHETER, THE VALVE WAS PUSHED AND CAUSED THE VALVE TO DISLODGE UP. A SECOND VALVE WAS DEPLOYED, THE VALVE WAS RAISED WITH THE DELIVERY CATHETER SYSTEM (DCS) AND BECAME HOOKED ON THE FIRST VALVE, CAUSING IT TO BE PUSHED AND THE AORTA BECAME PERFORATED. THE PATIENT SUBSEQUENTLY DIED DUE TO THE AORTIC PERFORATION. ADDITIONAL INFORMATION WAS RECEIVED THAT THE SECOND VALVE WAS NOT IMPLANTED AS IT NEVER PASSED THE AORTIC ROOT. PER THE PHYSICIAN, THE SECOND VALVE PUSHED THE FIRST VALVE IN THE AORTA. THERE WAS NO TREATMENT FOR THE PERFORATION PRIOR TO THE DEATH. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVPROP23-29 (LOT: 0012663964); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE ; EXPLANT DATE PRODUCT ID L-EVPROP23-29 (LOT: 0012623710); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE ; EXPLANT DATE PRODUCT ID L-EVPROP23-29 (LOT: 0012483153); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE ; EXPLANT DATE PRODUCT ID EVPROPLUS-29 (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE PRODUCT ID D-EVPROP23-29 (LOT: 0012617743); PRODUCT TYPE: 0195-HEART VALVES; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B5, D4, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22310034&pc=MGB | 2024168-2025-06410 | 2025-05-28 04:00 | Death | ABBOTT VASCULAR INC. | 2025-06-23 04:00 | MGB | PERCLOSE PROGLIDE¿ | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P960043 |  | 1 | Event Description: THE STUDY RETROSPECTIVELY COMPARED THE SAFETY AND EFFICACY OF THE 14F MANTA (PLUG BASED VESSEL CLOSURE DEVICE OR PB-VCD) TO THE PERCLOSE PROGLIDE (SUTURE BASED VESSEL CLOSURE DEVICE OR SB-VCD) FOLLOWING THE USE OF 14F LOW-PROFILE TRANSCATHETER HEART VALVE DELIVERY SHEATH DURING TAVR. THE STUDY ALSO COMPARED THE FINDINGS TO PREVIOUS DESCRIBED RATES OF VASCULAR COMPLICATIONS AFTER LARGER 18F VESSEL CLOSURE DEVICE USE. THE STUDY CONCLUDED THAT 14F PB-VCD MANTAS ARE ASSOCIATED WITH HIGHER RATES OF MINOR VASCULAR COMPLICATIONS COMPARED TO 14F SB-VCD PROGLIDES. REPORTEDLY, FAILURE TO ACHIEVE HEMOSTASIS AND OTHER UNSPECIFIED FAILURES OCCURRED WITH THE PROGLIDE, RESULTING IN DEATH, UNSPECIFIED MAJOR AND MINOR VASCULAR COMPLICATIONS, BLEEDING, OCCLUSION, DISSECTION AND HOSPITALIZATION. INTERVENTIONS AND TREATMENTS FOR THE ADVERSE PATIENT EFFECTS INCLUDED UNSPECIFIED ENDOVASCULAR MEASURES, MANUAL COMPRESSION, AND SURGERY. ADDITIONAL INFORMATION CAN BE FOUND IN THE ARTICLE TITLED, "PROPENSITY MATCHED ANALYSIS OF 14 FRENCH SUTURE- VERSUS PLUG-BASED VASCULAR CLOSURE DURING TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR)." Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2- THE DATE OF DEATH IS ESTIMATED AS (B)(6) 2025 B3- THE DATE OF EVENT IS ESTIMATED AS 05/28/2025 D4- THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. ATTACHMENT ARTICLE TITLED: "PROPENSITY MATCHED ANALYSIS OF 14 FRENCH SUTURE- VERSUS PLUG-BASED VASCULAR CLOSURE DURING TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR)". THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS REPORTED IN THE ARTICLE ARE CAPTURED UNDER SEPARATE A MEDWATCH REPORT. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. SUMMARY OF PATIENT DEATH: BASED ON THE ARTICLE INFORMATION, A CONCLUSIVE CAUSE FOR THE REPORTED DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. BASED ON THE INFORMATION REPORTED THROUGH THE RESEARCH ARTICLE, A DEFINITIVE CAUSE FOR THE REPORTED DEATH AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22310874&pc=LPB | 2029046-2025-02057 | 2025-05-29 04:00 | Death | BIOSENSE WEBSTER INC | 2025-06-23 04:00 | LPB | THERMOCOOL® SMART TOUCH® SF UNI-DIRECTIONAL NAVIGATION CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P030031 |  | 1 | Event Description: THE REPORT INDICATED THAT DURING AFIB ABLATION PROCEDURE WITH A THERMOCOOL® SMART TOUCH® SF UNI-DIRECTIONAL NAVIGATION CATHETER, THE PATIENT EXPERIENCED BLOOD PRESSURE DROP 50/40, AND PERICARDIAL EFFUSION DISCOVERED WITH INTRACARDIAC ECHO. HEPARIN WAS STOPPED AND PROTAMINE WAS GIVEN. A PERICARDIOCENTESIS WAS PERFORMED BY THE PHYSICIAN, REMOVING ABOUT 1300CC'S. OPEN REPAIR WAS THEN PERFORMED IN THE LAB TO CORRECT THE PERFORATION. THE PATIENT EXPIRED DURING THE SURGERY. Manufacturer Narrative: ON 18-JUL-2025, THE PRODUCT INVESTIGATION WAS COMPLETED AS THE COMPLAINT DEVICE WAS NOT RETURNED. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER 31568972L AND NO INTERNAL ACTION RELATED TO THE COMPLAINT WAS FOUND DURING THE REVIEW. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. BASED ON THE INFORMATION AVAILABLE, THERE IS NO INDICATION THAT A DESIGN OR MANUFACTURING ISSUE HAS CAUSED THE REPORTED COMPLAINT CONDITION. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REFERENCE NUMBER: (B)(4). Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC. OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22311667&pc=LWS | 2649622-2025-16936 | 2024-07-29 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-06-23 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED T-WAVE OVERSENSING (TWOS). THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT IS DECEASED AND THE T-WAVE OVERSENSING (TWOS) EPISODES OCCURRED APPROXIMATELY FIVE MONTHS PRIOR TO THE PATIENT'S DEATH. THE DAY BEFORE THE PATIENT PASSED AWAY, THE PATIENT WAS NOTED TO BE IN VENTRICULAR FIBRILLATION (VF) AND MULTIPLE HIGH VOLTAGE THERAPY SHOCKS WERE DELIVERED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22312461&pc=NPT | 2025587-2025-04430 | 2025-06-16 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-23 04:00 | NPT | EVOLUT FX PLUS VALVE | Device Dislodged or Dislocated; Patient Device Interaction Problem; Central Regurgitation | Ventricular Fibrillation; Cardiogenic Shock; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE PROCEDURE, A PATIENT WITH A MEDICAL HISTORY OF SEVERE PERIPHERAL OCCLUSIVE DISEASE, SEVERE AORTIC VALVE STENOSIS (OPENING AREA 0.4), AND SEVERE PULMONARY EDEMA UNDERWENT MULTIPLE VASCULAR PUNCTURES, INCLUDING THE RIGHT COMMON VEIN, RIGHT COMMON ARTERY, LEFT AXILLARY ARTERY, AND LEFT BRACHIAL ARTERY. SHEATHS WERE INSERTED UNDER ULTRASOUND GUIDANCE, AND A PIGTAIL CATHETER WAS PLACED IN THE ASCENDING AORTA. A PACEMAKER WAS POSITIONED IN THE SEPTAL RIGHT VENTRICLE. TWO NON-MEDTRONIC CLOSURE SYSTEMS WERE PLACED IN THE LEFT AXILLARY ARTERY, AND A NON-MEDTRONIC GUIDEWIRE WAS INSERTED INTO THE DESCENDING AORTA AS A SAFETY WIRE. RETROGRADE PROBING OF THE LEFT VENTRICLE VIA THE AORTIC VALVE WAS PERFORMED WITH AN AL 1 CATHETER AND STRAIGHT NON-MEDTRONIC WIRE. IMMEDIATELY AFTER PASSAGE OF THE WIRE, REFRACTORY VENTRICULAR FIBRILLATION OCCURRED. RESUSCITATION WAS INITIATED, AND BALLOON VALVULOPLASTY WAS PERFORMED. THE EVOLUT FX+ 26 MILLIMETER VALVE WAS INSERTED AND DEPLOYED WHILE RESUSCITATION CONTINUED, AND MEDICATION WAS ADMINISTERED ACCORDING TO PROTOCOL. PERIPHERAL EXTRACORPOREAL MEMBRANE OXYGENATION (ECCMO) WAS INITIATED, AND AN ARTERIAL PUNCTURE OF THE LEFT GROIN WAS PERFORMED. AORTOGRAPHY REVEALED HIGH-GRADE AORTIC REGURGITATION AFTER THE FIRST VALVE DEPLOYMENT. A SECOND 23 MM NON-MEDTRONIC VALVE WAS IMPLANTED, AND SUBSEQUENT AORTOGRAPHY SHOWED CONTINUED MODERATE AORTIC REGURGITATION. DESPITE THESE INTERVENTIONS, THERE WAS NO ADEQUATE CIRCULATION UNDER ECCMO, AND THE PATIENT EXPERIENCED CARDIOGENIC SHOCK, RESULTING IN DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT FOLLOWING THE VALVE IMPLANT, SEVERE CENTRAL AORTIC REGURGITATION WAS OBSERVED AS THE VALVE WAS DISLODGING DURING THE REANIMATION AND THE VALVE WAS DEPLOYED UNDER REANIMATION. PER THE PHYSICIAN, THE VALVE COULD BE EXCLUDED AS A CONTRIBUTING CAUSE OF DEATH. PER THE PHYSICIAN, THE PATIENT'S UNDER LYING MEDICAL CONDITIONS CONTRIBUTED THE PATIENT DEATH. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012768940); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012786156); PRODUCT TYPE: 0195-HEART VALVES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA D4 H2 H3 H6 ADDITIONAL CODES CONCLUSION: FLUOROSCOPIC CINE LOOPS OF THE PRE-BALLOON DILATION AND IMPLANT PROCEDURE WERE PROVIDED FOR REVIEW OF THE EVENT. PATIENT SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE EVOLUT FX+ SYSTEM INCLUDE, BUT ARE NOT LIMITED TO, THE FOLLOWING: ¿ DEATH, CARDIAC ARREST, CARDIOGENIC SHOCK, PROSTHETIC VALVE DYSFUNCTION (REGURGITATION OR STENOSIS), MALPOSITION (EITHER TOO HIGH OR TOO LOW)/MALPLACEMENT DURING FLUOROSCOPIC LOAD INSPECTION, NO INFLOW CROWN-OVERLAP AND LATER, NO VALVE-INFOLDING COULD BE OBSERVED DURING DEPLOYMENT. THIS MAKES THE OBSERVED REGURGITATION ORIGINATING FROM THOSE CAUSES HIGHLY UNLIKELY. WHILE AN IMAGE SHOWED A LAMINAR AORTIC FLOW WITH A SMOOTH ASCENDING AORTA WALL, ANOTHER IMAGE SHOWED A SLOW MOVING FLOW WITH UNUSUAL FLOW PATTERNS. THE LOW FLOW MAY WELL HAVE BEEN CAUSED BY THE REPORTED VENTRICULAR FIBRILLATION (VF), WHICH HAD SET IN AFTER CROSSING THE VALVE WITH THE SUPPORTIVE GUIDE WIRE. HOWEVER, ANOTHER IMAGE SHOWED A STRUCTURE MOVING IN THE BLOOD STREAM JUST NEXT TO THE PIGTAIL. WHEN THE VALVE IS DEPLOYED TO ALMOST 80%, AN UNUSUAL INDENT IN THE ASCENDING AORTA WALL IS VISIBLE WHICH WAS NOT THERE IN THE BEGINNING IMAGE. THERE WAS THE POSSIBILITY THAT EITHER THE GUIDEWIRE CROSSING THE VALVE, THE PRE-BALLOON DILATION OR THE BIOPROSTHESIS INSERTION, DAMAGED THE AORTIC WALL. THE REPORTED VF MAY HAVE BEEN ONE CONSEQUENCE OF THIS INJURY. IN A WORST-CASE SCENARIO, AN AORTIC DISSECTION MIGHT HAVE FORMED WHICH EXTENDED OVER THE AORTIC VALVE INTO THE VENTRICLE, CREATING A SHUNT THROUGH WHICH BLOOD FLOWED INTO THE VENTRICLE IN DIASTOLE IN A PARAVALVULAR LEAK (PVL) FASHION. ANOTHER IMAGE FEATURED A DARK ¿SACK¿ ON THE SIDE OF THE VALVE OPPOSITE TO THE PIGTAIL AND THE NEXT IMAGE SHOWED A FLOW THAT APPEARED TO FLOW THROUGH THE LEFT AORTIC VALVE NADIR. THIS POSSIBILITY IS SUPPORTED BY THE FACT THAT NEITHER THE PVL FOR THE EVOLUT NOR EDWARDS COULD EFFECTIVELY BE TREATED. ANOTHER, POSSIBLE CAUSE FOR THE HIGH-DEGREE PVL FORMATION DURING THE EVOLUT DEPLOYMENT IS THE FACT THAT THE IMPLANTERS DID NOT ADHERE TO MEDTRONIC BEST PRACTICE REGARDING DEPLOYMENT STARTING POSITION AND IMPLANT DEPTH. THE VALVE ENDED UP MUCH TOO DEEP (NON-CORONARY CUSP (NCC) DEPTH CA. 6-8 MM). THIS REDUCES THE RADIAL FORCE APPLIED TO THE ANNULUS AND CAN CAUSE PVL AS WELL. WITHOUT MORE DETAILED INFORMATION, IT¿S NOT POSSIBLE TO DETERMINE THE ROOT CAUSE OF THE PVL WHICH LED TO SNARING THE VALVE AND IMPLANTING A SECOND NON-MEDTRONIC VALVE. THE ROOT CAUSE OF THE REFRACTORY VF APPEARS TO HAVE BEEN AN INTERACTION OF THE GUIDE WIRE WITH THE CONDUCTION SYSTEM. DOWN THE LINE, THIS WAS PROBABLY THE CAUSE FOR THE LOW FLOW AND LATER CARDIOGENIC SHOCK WITH SUBSEQUENTIALLY LEAD TO THE DEATH OF THE PATIENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22312489&pc=LOD | 0002249697-2025-00633 | 2025-05-29 04:00 | Death | STRYKER ORTHOPAEDICS-MAHWAH | 2025-06-23 04:00 | LOD | SIMPLEX P-JAPANESE TWIN PACK | Insufficient Information | Pulmonary Embolism |  |  | 1 | Event Description: OCCURRED DURING BHA SURGERY. DURING SIMPLEX CEMENT INJECTION, BLOOD PRESSURE DROPPED RAPIDLY DUE TO SUSPECTED PULMONARY EMBOLISM, AND THERE WAS A POSSIBILITY OF CARDIAC ARREST. RESUSCITATION (CARDIAC MASSAGE, ETC.) WAS ATTEMPTED, BUT THE PATIENT DID NOT REGAIN CONSCIOUSNESS, AND THE FAMILY DECIDED NOT TO CONTINUE LIFE-SUSTAINING MEASURES. Manufacturer Narrative: IT WAS NOTED THAT THE DEVICE IS NOT AVAILABLE FOR EVALUATION. IF ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE PROVIDED IN A SUPPLEMENTAL REPORT UPON COMPLETION OF THE INVESTIGATION. REVIEW OF THE DEVICE HISTORY RECORDS INDICATE DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO REPORTED DISCREPANCIES. THERE HAVE BEEN NO OTHER EVENTS FOR THE LOT REFERENCED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22312573&pc=NAY | 2955842-2025-26592 |  | Death | INTUITIVE SURGICAL, INC | 2025-06-23 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Sepsis; Insufficient Information | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN UNSPECIFIED DA VINCI-ASSISTED SURGICAL PROCEDURE, THE PATIENT DEVELOPED CHOLANGITIS CAUSED BY STENT PLACEMENT AND LATER EXPIRED. THE SURGEON INFORMED THE INTUITIVE CLINICAL SALES REPRESENTATIVE OF THE PATIENT DEATH. THE SURGEON STATED THAT THE PATIENT DEATH WAS UNRELATED TO THE PROCEDURE THAT WAS PERFORMED, HOWEVER, NO ADDITIONAL EVENT DETAILS WERE PROVIDED. ADDITIONAL INFORMATION WAS REQUESTED; HOWEVER, NO RESPONSE HAS BEEN RECEIVED.Event Description: REFER TO H11 FOR FOLLOW-UP INFORMATION. Manufacturer Narrative: INSUFFICIENT INFORMATION IS AVAILABLE FOR PROCEDURE VERIFICATION; THEREFORE, NO DA VINCI SYSTEM OR INSTRUMENT LOGS ARE AVAILABLE FOR REVIEW. CONFIRMATION OF THE SPECIFIC PROCEDURE HAS BEEN REQUESTED. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT REPORTEDLY DIED FROM CHOLANGITIS AFTER A STENT PLACEMENT FOLLOWING AN UNNAMED PROCEDURE. NO OTHER INFORMATION IS PROVIDED ABOUT THE PROCEDURE, THE REASON FOR THE STENT, THE ISSUES OR COMPLICATIONS THAT MAY HAVE OCCURRED LEADING TO THE STENT PLACEMENT, NOR THE ISSUES OR COMPLICATIONS THAT LED TO THE DEATH. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. Manufacturer Narrative: ADDITIONAL INFORMATION B5: IT WAS REPORTED THAT FOLLOWING A DA VINCI-ASSISTED SUBTOTAL CHOLECYSTECTOMY WITH COMMON BILE DUCT (CBD) EXPLORATION, THE PATIENT DEVELOPED CHOLANGITIS AND SUBSEQUENTLY EXPIRED. THE INDICATION FOR SURGERY WAS JAUNDICE SECONDARY TO MIRIZZI SYNDROME, CAUSED BY A LARGE GALLSTONE COMPRESSING THE CBD. A CBD STENT HAD BEEN PLACED THREE MONTHS PRIOR TO THE PROCEDURE TO RELIEVE THE OBSTRUCTION. INTRAOPERATIVELY, THERE WERE NO COMPLICATIONS; HOWEVER, THE PROCEDURE WAS NOTED TO BE CHALLENGING. AS A RESULT, THE GALLBLADDER WAS PARTIALLY REMOVED, ALONG WITH THE OBSTRUCTING STONE. THE CYSTIC DUCT LUMEN WAS VISUALIZED, THE CBD STENT WAS CONFIRMED IN PLACE, AND THE CYSTIC DUCT WAS SUTURED CLOSED. THE PROCEDURE WAS SUCCESSFULLY COMPLETED ROBOTICALLY. THE PATIENT EXPERIENCED NO IMMEDIATE POST-OPERATIVE COMPLICATIONS AND WAS DISCHARGED HOME IN STABLE CONDITION. THE SURGEON WAS NOT MADE AWARE OF ANY CONCERNS UNTIL RECEIVING A PHONE CALL, AND WAS INFORMED THE PATIENT WAS FEELING UNWELL WITH UPPER RESPIRATORY SYMPTOMS. AS A RESULT, THE PATIENT REQUESTED TO POSTPONE THE SCHEDULED ERCP ON POST-OPERATIVE DAY (POD) 8, WHICH HAD BEEN PLANNED FOR THE CBD STENT REMOVAL. THE PATIENT LATER EXPIRED ON POD 18. AN AUTOPSY WAS PERFORMED, AND THE CAUSE OF DEATH WAS CONFIRMED TO BE SEPSIS SECONDARY TO CHOLANGITIS. THE SURGEON DOES NOT BELIEVE THAT A DA VINCI SYSTEM, INSTRUMENT, AND/OR ACCESSORY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE SURGEON BELIEVES THE PATIENT¿S DEATH MAY HAVE BEEN PREVENTABLE HAD THEY PRESENTED FOR THE SCHEDULED ERCP AND STENT REMOVAL. THE SURGEON ALSO NOTED THAT THE RISK OF CHOLANGITIS INCREASES SIGNIFICANTLY WHEN A BILIARY STENT REMAINS IN PLACE FOR LONGER THAN THREE MONTHS; IN THIS CASE, THE LAST STENT EXCHANGE HAD OCCURRED OVER THREE MONTHS PRIOR. ADDITIONAL INFORMATION H10: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 4 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, LARGE NEEDLE DRIVER, PROGRASP FORCEPS, FENESTRATED BIPOLAR FORCEPS, SUCTION IRRIGATOR. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE ABOUT THE INSTRUMENTS USED DURING THIS PROCEDURE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A BILIARY STENT PRIOR TO THE PROCEDURE INDICATING SOME DEGREE OF BILE DUCT OBSTRUCTION. THEY UNDERWENT A CHALLENGING CHOLECYSTECTOMY WHERE ONLY A PORTION OF THE GALLBLADDER COULD BE SAFELY REMOVED DUE TO THE DIFFICULT ANATOMY. NO ALLEGATION WAS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT DURING THAT PROCEDURE. POSTOPERATIVELY, THERE WAS THE INTENTION TO REMOVE THE PREOPERATIVELY PLACED STENT VIA ERCP ON POD 8 BUT THE PATIENT DEFERRED THE PROCEDURE BECAUSE THEY WERE NOT FEELING WELL. THEY SUBSEQUENTLY DIED ON POD 18 FROM CHOLANGITIS WHICH WAS CONFIRMED ON AUTOPSY. CHOLANGITIS IS A WELL-KNOWN COMPLICATION FROM BILIARY STENTS AND CAN BE QUITE SEVERE(1). BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. 1) EVERETT BT, NAUD S, ZUBARIK RS. RISK FACTORS FOR THE DEVELOPMENT OF STENT-ASSOCIATED CHOLANGITIS FOLLOWING ENDOSCOPIC BILIARY STENT PLACEMENT. DIG DIS SCI. 2019 AUG;64(8):2300-2307. DOI: 10.1007/S10620-019-05533-6. EPUB 2019 FEB 20. PMID: 30788687.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22304201&pc=OZP | 2032227-2025-195058 | 2025-05-15 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-22 04:00 | OZP | MINIMED 780G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Coma | P160017 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2025, AND THE CAUSE OF DEATH WAS DIABETIC COMA. THE CUSTOMER REPORTED THE HYPERGLYCEMIA, HYPOGLYCEMIA, AND COMA. THE EVENT INVOLVED PRODUCTS MMT-342G, MMT-431A, AND MMT-1884. TROUBLESHOOTING WAS PERFORMED AND FOUND THAT THE DOCUMENT OF ANY HEALTH ISSUES OR ILLNESS THAT MAY HAVE CONTRIBUTED OR LED UP TO PASSING WAS THE DIABETIC COMA. THE CUSTOMER WAS STATED THAT THE PUMP WAS WORN AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-342G AND MMT-431A. MMT-1884 WAS REQUESTED, AND THE CUSTOMER'S RESPONSE WAS THAT THE DEVICE WILL BE RETURNED. THE CUSTOMER WILL DISCONTINUE THE USE OF THE DEVICE. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. THE DEVICE WILL BE RETURNED FOR ANALYSIS AND FURTHER INFORMATION WILL FOLLOW ONCE THE ANALYSIS HAS BEEN COMPLETED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22304483&pc=OZD | 1220648-2025-29686 | 2025-06-05 04:00 | Death | ABIOMED, INC. | 2025-06-22 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT (UNKNOWN ETHNICITY) WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT (MCS) EXPERIENCED VENTRICULAR TACHYCARDIA AND WOULD SUBSEQUENTLY EXPIRE. THE PATIENT UNDERWENT A LEFT AND RIGHT HEART CATHETERIZATION, AND THE DECISION WAS MADE TO CONSULT SURGERY. AS THE PATIENT WAITED FOR SURGERY CONSULT, THE PATIENT CODED. CARDIOPULMONARY RESUSCITATION WAS INITIATED, AND ONE SHOCK WAS DELIVERED. THE PATIENT WAS TRANSPORTED TO THE CATHETERIZATION LAB FOR THE IMPELLA CP DEVICE IMPLANT VIA THE LEFT FEMORAL ARTERY. PERFORMANCE LEVEL WAS RAMPED UP TO P-5. THE PATIENT HAD TWO EPISODES OF VENTRICULAR TACHYCARDIA NECESSITATING SHOCKS. A PULMONARY CATHETER WAS PLACE AND THE PATIENT WAS RETURN TO THE INTENSIVE CARE UNIT IN STABLE CONDITION. HOWEVER, SOME TIME THEREAFTER THE PATIENT EXPIRED. FURTHER DETAILS WERE UNNOTED. TOTAL IMPELLA MCS TIME WAS 18 HOURS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22304520&pc=OZD | 1220648-2025-29689 | 2025-06-06 04:00 | Death | ABIOMED, INC. | 2025-06-22 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP HAD HEMOLYZED LABS AND RED-TINGED URINE DESPITE REPOSITIONING THE PUMP. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS AND THROMBUS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. HEMOLYSIS: NO PRODUCT WAS RETURNED FOR ANALYSIS. CLINICAL MENTION HEMOLYZED LABS WHICH CONTINUED DESPITE PUMP REPOSITIONING. NO PATIENT ABNORMALITIES WERE NOTED. THE DATA LOGS SHOW A MOTOR CURRENT SPIKE ON 6/6 A LITTLE AFTER MIDNIGHT BEFORE HEMOLYZED LABS WERE REPORTED LATER IN THE MORNING. THERE WAS A MOTOR CURRENT SPIKE WITH SPEED DEVIATION CHARACTERISTIC OF AN INGESTION EVENT. THE ROOT CAUSE OF THE HEMOLYSIS WAS MOST LIKELY BIOMATERIAL AS EVIDENCED BY THE INGESTION PATTERN OBSERVED. THE FAILURE MODE THROMBUS WAS ADDED AS A FINDING DURING THE INVESTIGATION. THROMBUS: THE ROOT CAUSE OF THROMBUS WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22304551&pc=OZD | 1220648-2025-29691 | 2025-06-06 04:00 | Death | ABIOMED, INC. | 2025-06-22 04:00 | OZD | IMPELLA 5.5 | Failure to Advance; Patient-Device Incompatibility | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT DELIVERY ISSUES WERE ENCOUNTERED WHEN ATTEMPTING TO IMPLANT AN IMPELLA 5.5. RESISTANCE WAS MET AT ANASTAMOSIS SO THE PUMP WAS REMOVED, FLUSHED, AND IMPLANTATION WAS SUCCESSFUL. THE PATIENT ALSO EXHIBITED BLEEDING AT THE SURGICAL SITE. THE SITE WAS SURGICALLY REPAIRED. THE PATIENT'S FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22304569&pc=OZD | 1220648-2025-29693 | 2025-06-06 04:00 | Death | ABIOMED, INC. | 2025-06-22 04:00 | OZD | IMPELLA CP | Malposition of Device; Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP SUPPORT EXPERIENCED BLEEDING FROM THE RADIOFREQUENCY ABLATION SITE, AND THE PATIENT'S HEMOGLOBIN DROPPED FROM 12 TO 5. ADDITIONALLY, THE PUMP REPOSITIONING SHEATH HAD BACKED OUT DESPITE BEING SUTURED IN PLACE. FOUR UNITS OF RED BLOOD CELLS WERE TRANSFUSED AND THE SHEATH WAS REPOSITIONED TO RESOLVE THE ISSUES. THE NEXT DAY THE PATIENT EXPERIENCED A RETROPERITONEAL BLEED AND ADDITIONAL BLOOD PRODUCTS WERE TRANSFUSED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ACCESS SITE BLEEDING: THE ROOT CAUSE OF THE ACCESS SITE BLEEDING IS DETERMINED TO BE USE ISSUE BECAUSE BLEEDING WAS RESOLVED BY ADVANCING THE REPO SHEATH BACK INTO PLACE AND ANGLE MATCHING. VASCULAR DAMAGE - PERIPHERAL VESSEL: THE ROOT CAUSE OF THE VASCULAR DAMAGE WAS UNABLE TO BE DETERMINED AS INSUFFICIENT CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED FOR ANALYSIS. B1 PRODUCT PROBLEM WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29693. H6 MEDICAL DEVICE PROBLEM CODE 2616 WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29693. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR EVALUATION. SHOULD THE DEVICE OR ANY ADDITIONAL INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22304577&pc=OZD | 1220648-2025-29696 | 2025-06-06 04:00 | Death | ABIOMED, INC. | 2025-06-22 04:00 | OZD | IMPELLA CP | Malposition of Device | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED BLEEDING AS THE IMPELLA PEEL AWAY SHEATH WAS ACCIDENTLY PULLED BACK. A LARGE HEMATOMA DEVELOPED AS A RESULT. THE PATIENT'S FAMILY WITHDREW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION POSITIONING ISSUE: THE CAUSE OF THE POSITIONING ISSUE MOST LIKELY USE RELATED AS THE PEEL-AWAY WAS ACCIDENTALLY PULLED BACK DURING CARDIOPULMONARY RESUSCITATION (CPR). ACCESS SITE BLEEDING: THE CAUSE OF THE ACCESS SITE BLEEDING ISSUE MOST LIKELY WAS USE RELATED AS THE PEEL-AWAY WAS ACCIDENTALLY PULLED BACK DURING CPR LEADING TO HEMATOMA. B2 INTERVENTION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29696. H6 HEALTH EFFECT - IMPACT CODE 4641 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29696.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22303546&pc=OZD | 1220648-2025-29639 | 2025-06-04 04:00 | Death | ABIOMED, INC. | 2025-06-21 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND EXHIBITED A MOTTLED AND COOL LOWER EXTREMITY. THE LIMB LOST DOPPLER PULSE AND SHOWED WEAK NON-PULSATILE FLOW. THE PUMP WAS WEANED TO P4 TO RESTORE DOPPLER PULSES, COLOR, AND TEMPERATURE. DUE TO THE PATIENT'S POOR PROGNOSIS, THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22303556&pc=OZD | 1220648-2025-29642 | 2025-06-02 04:00 | Death | ABIOMED, INC. | 2025-06-21 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Gastrointestinal Hemorrhage | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT DEVELOPED A GASTROINTESTINAL BLEED. ANTICOAGULATION WAS WITHHELD AND SUPPORT CONTINUED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22303654&pc=PYX | 1220648-2025-29653 | 2025-06-03 04:00 | Death | ABIOMED, INC. | 2025-06-21 04:00 | PYX | IMPELLA RP FLEX | Patient-Device Incompatibility; Deformation Due to Compressive Stress | Cardiovascular Insufficiency | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED THE PATIENT HAD AN ATTEMPTED IMPELLA RP FLEX IMPLANT AND THERE WAS KINKING AND BUCKLING AT THE INLET AND MOTOR HOUSING. NOT ABLE TO ADVANCE CATHETER ANY FURTHER WITHOUT THIS ISSUE. DECISION MADE TO START OVER WITH A NEW CATHETER WITH FEAR OF DAMAGE FROM THE KINKING AND BUCKLING. FURTHER DETAILS STATE THAT ONCE THE BI-PELLA WAS PLACED, THE PATIENT EXPIRED ON THE TABLE. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: RP SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ARRHYTHMIA, ATRIAL FIBRILLATION, BLEEDING, CARDIAC TAMPONADE, CARDIOGENIC SHOCK, DEATH, DEVICE MALFUNCTION, HEMOLYSIS, HEPATIC FAILURE, INSERTION SITE INFECTION, PERFORATION, PHLEGMASIA CERULEA DOLENS (A SEVERE FORM OF DEEP VENOUS THROMBOSIS), PULMONARY VALVE INSUFFICIENCY, RESPIRATORY DYSFUNCTION, SEPSIS, THROMBOCYTOPENIA, THROMBOTIC VASCULAR (NON-CENTRAL NERVOUS SYSTEM COMPLICATION, TRICUSPID VALVE INJURY, CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION), VENOUS THROMBOSIS, VENTRICULAR FIBRILLATION AND/OR TACHYCARDIA.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿DO NOT ADVANCE OR WITHDRAW THE IMPELLA RP FLEX WITH SMARTASSIST SYSTEM CATHETER AGAINST RESISTANCE WITHOUT USING FLUOROSCOPY TO DETERMINE THE CAUSE OF THE RESISTANCE. DOING SO COULD RESULT IN SEPARATION OF THE CATHETER OR GUIDEWIRE TIP, DAMAGE TO THE CATHETER OR VESSEL, OR PERFORATION.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES AND PRODUCT DAMAGE(CANNULA KINK) HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. DELIVERY ISSUE: THE ROOT CAUSE OF DELIVERY ISSUE IS DETERMINED TO BE USE ISSUE BECAUSE THE NEW RP FLEX WAS INSERTED SUCCESSFULLY. PRODUCT DAMAGE(CANNULA KINK): THE ROOT CAUSE OF PRODUCT DAMAGE, CANNULA KINK IS DETERMINED TO BE USE ISSUE BECAUSE THE NEW RP FLEX WAS IMPLANTED SUCCESSFULLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22303766&pc=BZD | 2518422-2025-106775 | 2025-06-13 04:00 | Death | RESPIRONICS, INC. | 2025-06-21 04:00 | BZD | PHILIPS CPAP DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22303767&pc=CAF | 2518422-2025-106774 | 2025-06-10 04:00 | Death | RESPIRONICS, INC. | 2025-06-21 04:00 | CAF | PHILIPS NEBULIZER DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A NEBULIZER DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22294721&pc=MGB | 3013394970-2025-00326 | 2024-11-18 05:00 | Death | TERUMO MEDICAL CORPORATION | 2025-06-20 04:00 | MGB | ANGIO-SEAL | Failure to Seal | Retroperitoneal Hemorrhage | P930038 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THEY HAD A CASE WHERE AN ANGIOGRAM WAS PERFORMED USING THE FEMORAL APPROACH. THE PUNCTURE WAS SEALED USING AN ANGIO-SEAL DEVICE FOLLOWING A FLURO WHICH CONFIRMED THE PUNCTURE WAS ABOVE THE BIFURCATION. THE FOLLOWING DAY, THE PATIENT EXPERIENCED GROIN PAIN AND COLLAPSE WITH HYPOTENSION AFTER STRAINING ON THE TOILET, AND A NATIONAL EARLY WARNING SCORE (NEWS) EMERGENCY CALL WAS ACTIVATED. SHE WAS STABILIZED AND TRANSFERRED FROM THE CARDIOLOGY WARD (G9) TO THE CORONARY CARE UNIT (CCU) FOR CLOSER MONITORING, AND A COMPUTED TOMOGRAPHY (CT) SCAN WAS REQUESTED. SHE WAS TAKEN FOR THE COMPUTED TOMOGRAPHY (CT) SCAN AND ON ARRIVAL BACK IN THE CORONARY CARE UNIT (CCU), THE PATIENT REPORTED FURTHER GROIN PAIN AND BECAME HYPOTENSIVE. THE RADIOLOGIST PHONED THROUGH THE COMPUTED TOMOGRAPHY (CT) SCAN REPORT SHOWING AN ACTIVE RETROPERITONEAL BLEED FROM THE FEMORAL PUNCTURE SITE. A NEWS CALL WAS ACTIVATED, AND VASCULAR AND INTERVENTIONAL RADIOLOGY ADVICE SOUGHT. MANUAL PRESSURE WAS APPLIED TO THE SITE OF THE FEMORAL PUNCTURE. AFTER DISCUSSION BETWEEN THE CARDIOLOGY, VASCULAR SURGERY AND INTERVENTIONAL RADIOLOGY CONSULTANTS, THE PLAN WAS FOR INTERVENTIONAL RADIOLOGY GUIDED INTERVENTION TO TRY TO CONTROL THE BLEEDING. A MASSIVE HEMORRHAGE CALL WAS ACTIVATED, AND BLOOD TRANSFUSIONS WERE ADMINISTERED. THE PATIENT WAS TRANSFERRED TO RADIOLOGY FOR INTERVENTIONAL RADIOLOGY REPAIR WITH THE VASCULAR AND ANAESTHETIC TEAM. ON ARRIVAL TO RADIOLOGY THE PATIENT BECAME UNRESPONSIVE AND EXPERIENCED A CARDIAC ARREST AND DIED. HER DEATH WAS REFERRED TO THE CORONER. Manufacturer Narrative: E3: OCCUPATION: SPECIALIST NURSE THE PRODUCTION LOT # WAS NOT PROVIDED BY THE USER FACILITY, WHICH PREVENTED A MEANINGFUL REVIEW OF THE DEVICE HISTORY RECORD. THE ACTUAL DEVICE IS NOT AVAILABLE FOR RETURN. THE INVESTIGATION IS CURRENTLY ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. Manufacturer Narrative: THIS REPORT IS BEING SENT AS FOLLOW-UP NO. 1 TO PROVIDE THE COMPLETED INVESTIGATION RESULTS. THE ACTUAL DEVICE WAS NOT RETURNED; THEREFORE, AN EVALUATION OF THE ACTUAL DEVICE WAS UNABLE TO BE CONDUCTED. THE SAMPLE WAS NOT AVAILABLE FOR EVALUATION. THE COMPLAINT WAS CONFIRMED SINCE A PATIENT DEATH OCCURRED. THE EXACT ROOT CAUSE CANNOT BE DETERMINED. THE RELATIONSHIP OF THE DEVICE TO THE REPORTED EVENT CANNOT BE SUBSTANTIATED BASED ON AVAILABLE INFORMATION. THE DEVICE HISTORY RECORD (DHR) WAS UNABLE TO BE REVIEWED SINCE A VALID LOT NUMBER WAS NOT IDENTIFIED. THERE IS NO INDICATION THAT ANY MANUFACTURING ISSUES MAY HAVE LED TO THIS EVENT. CURRENTLY NO ACTION IS RECOMMENDED SINCE THIS RISK EVALUATION IS WITHIN THE PREDETERMINED LIMITS IN THE FAILURE MODE AND EFFECTS ANALYSIS (FMEA)/HAZARD BASED RISK TABLE (HBRT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22294790&pc=NKM | 2135147-2025-03299 | 2025-05-19 04:00 | Death | ABBOTT MEDICAL | 2025-06-20 04:00 | NKM | MITRACLIP | Incomplete Coaptation | Cardiogenic Shock; Heart Failure/Congestive Heart Failure; Mitral Valve Insufficiency/ Regurgitation | P100009 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT DEGENERATIVE MITRAL REGURGITATION (MR) WITH A GRADE OF 4, PROLAPSED POSTERIOR LEAFLET, LONG LEAFLETS, AND PROMINENT CHORDAE. A MITRACLIP XTW (50121A1112) WAS INSERTED WITHOUT ISSUES. HOWEVER, WHILE IN THE VALVE, IT WAS OBSERVED THAT ONE GRIPPER WAS NOT FUNCTIONING AS INTENDED. TROUBLESHOOTING WAS PERFORMED, BUT THE ISSUE CONTINUED TO OCCUR. THEREFORE, THE CLIP WAS REMOVED AND REPLACED. THREE CLIPS WERE THEN INSERTED AND DEPLOYED ON THE MITRAL VALVE, REDUCING MR TO A GRADE OF 1-2. ON (B)(6) 2025, THE PATIENT PRESENTED TO THE HOSPITAL WITH TACHYCARDIA, WITH RATES UP TO 200 BEATS PER MINUTE (BPM). THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) AND THE PATIENT WAS TREATED WITH DRUG THERAPY. AFTER SUCCESSFUL DRUG THERAPY, A TRANSTHORACIC ECHOCARDIOGRAM (TTE) WAS PERFORMED AND REVEALED THAT THE SECOND CLIP (50218A1064) PREVIOUSLY IMPLANTED, HAD DETACHED FROM THE ANTERIOR LEAFLET AND REMAINED ATTACHED TO THE POSTERIOR LEAFLET (SINGLE LEAFLET DEVICE ATTACHMENT/SLDA), CAUSING MR TO INCREASE. IT WAS NOTED THAT AT THIS TIME, THE PATIENT WAS HEMODYNAMICALLY STABLE. ON (B)(6) 2025, THE PATIENT'S CONDITIONS WORSENED. THE PATIENT EXPERIENCED CARDIOGENIC SHOCK DUE TO RIGHT VENTRICULAR FAILURE. THE PATIENT WAS TRANSFERRED TO SURGERY, WHERE A MITRAL VALVE REPLACEMENT AND TRICUSPID VALVE ANNULOPLASTY WAS PERFORMED. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED DURING THE SURGICAL INTERVENTION AND REVEALED THAT THE THIRD IMPLANTED CLIP (50218A1063) HAD DETACHED FROM THE POSTERIOR LEAFLET AND REMAINED ATTACHED TO THE ANTERIOR LEAFLET (SINGLE LEAFLET DEVICE ATTACHMENT/SLDA), CAUSING MR TO INCREASE. IT WAS NOTED THAT THE FIRST IMPLANTED CLIP REMAINED STABLE AND ATTACHED TO BOTH LEAFLES. AFTER THE SURGICAL INTERVENTION, THE PATIENT DIED DUE TO RIGHT VENTRICULAR FAILURE. IT WAS NOTED THAT PRE-EXISTING TO THE PATIENT'S DEATH, THE PATIENT HAD A SIGNIFICANTLY ELEVATED PULMONARY ARTERY PRESSURE (PAP) OF 50MMHG.Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT DEGENERATIVE MITRAL REGURGITATION (MR) WITH A GRADE OF 4, PROLAPSED POSTERIOR LEAFLET, LONG LEAFLETS, AND PROMINENT CHORDAE. A MITRACLIP XTW (50121A1112) WAS INSERTED WITHOUT ISSUES. HOWEVER, WHILE IN THE VALVE, IT WAS OBSERVED THAT ONE GRIPPER WAS NOT FUNCTIONING AS INTENDED. TROUBLESHOOTING WAS PERFORMED, BUT THE ISSUE CONTINUED TO OCCUR. THEREFORE, THE CLIP WAS REMOVED AND REPLACED. THREE CLIPS WERE THEN INSERTED AND DEPLOYED ON THE MITRAL VALVE, REDUCING MR TO A GRADE OF 1-2. ON (B)(6) 2025, THE PATIENT PRESENTED TO THE HOSPITAL WITH TACHYCARDIA, WITH RATES UP TO 200 BEATS PER MINUTE (BPM). THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) AND THE PATIENT WAS TREATED WITH DRUG THERAPY. AFTER SUCCESSFUL DRUG THERAPY, A TRANSTHORACIC ECHOCARDIOGRAM (TTE) WAS PERFORMED AND REVEALED THAT THE SECOND CLIP (50218A1064) PREVIOUSLY IMPLANTED, HAD DETACHED FROM THE ANTERIOR LEAFLET AND REMAINED ATTACHED TO THE POSTERIOR LEAFLET (SINGLE LEAFLET DEVICE ATTACHMENT/SLDA), CAUSING MR TO INCREASE. IT WAS NOTED THAT AT THIS TIME, THE PATIENT WAS HEMODYNAMICALLY STABLE. ON (B)(6) 2025, THE PATIENT'S CONDITIONS WORSENED. THE PATIENT EXPERIENCED CARDIOGENIC SHOCK DUE TO RIGHT VENTRICULAR FAILURE. THE PATIENT WAS TRANSFERRED TO SURGERY, WHERE A MITRAL VALVE REPLACEMENT AND TRICUSPID VALVE ANNULOPLASTY WAS PERFORMED. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED DURING THE SURGICAL INTERVENTION AND REVEALED THAT THE THIRD IMPLANTED CLIP (50218A1063) HAD DETACHED FROM THE POSTERIOR LEAFLET AND REMAINED ATTACHED TO THE ANTERIOR LEAFLET (SINGLE LEAFLET DEVICE ATTACHMENT/SLDA), CAUSING MR TO INCREASE. IT WAS NOTED THAT THE FIRST IMPLANTED CLIP REMAINED STABLE AND ATTACHED TO BOTH LEAFLES. AFTER THE SURGICAL INTERVENTION, THE PATIENT DIED DUE TO RIGHT VENTRICULAR FAILURE. IT WAS NOTED THAT PRE-EXISTING TO THE PATIENT'S DEATH, THE PATIENT HAD A SIGNIFICANTLY ELEVATED PULMONARY ARTERY PRESSURE (PAP) OF 50MMHG. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY DID NOT INDICATE A LOT-SPECIFIC QUALITY ISSUE. BASED ON AVAILABLE INFORMATION, THE REPORTED INCOMPLETE COAPTATION (SINGLE LEAFLET DEVICE ATTACHMENT/SLDA) RESULTING IN MITRAL VALVE INSUFFICIENCY/REGURGITATION APPEARS TO BE RELATED TO A PROCEDURAL CIRCUMSTANCE. ADDITIONALLY, THE REPORTED PATIENT EFFECT OF HEART FAILURE, DEATHS, SHOCK, CARDIOGENIC, AND MITRAL VALVE INSUFFICIENCY/REGURGITATION ARE KNOWN POSSIBLE COMPLICATION ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED SHOCK, CARDIOGENIC WAS DUE TO THE RIGHT VENTRICULAR FAILURE. THE REPORTED DEATH WAS RELATED TO PRE-EXISTING COMORBIDITIES-CARDIOGENIC SHOCK AND RIGHT VENTRICULAR FAILURE. HEART FAILURE COULD NOT BE DETERMINED. THE REPORTED HOSPITALIZATION AND SURGICAL INTERVENTION WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22296062&pc=MNS | 2518422-2025-047604 | 2023-03-24 04:00 | Death | RESPIRONICS, INC. | 2025-06-20 04:00 | MNS | DREAMSTATION BIPAP AUTOSV | Degraded | Fibrosis | K090539 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH FROM PULMONARY FIBROSIS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22296577&pc=NPT | 2015691-2025-05039 | 2025-03-08 05:00 | Death | EDWARDS LIFESCIENCES | 2025-06-20 04:00 | NPT | EDWARDS LIFESCIENCES SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Endocarditis | P140031 |  | 1 | Event Description: IT WAS REPORTED BY THE JAPANESE TAVI REGISTRY THAT APPROXIMATELY 30 DAYS POST TRANSFEMORAL TAVR WITH A 29 MM SAPIEN 3 ULTRA RESILIA VALVE THE PATIENT EXPIRED DUE TO INFECTIVE ENDOCARDITIS. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), INFECTIONS SUCH AS SEPTICEMIA AND ENDOCARDITIS ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. ENDOCARDITIS IS AN INFECTION OF A NATIVE OR PROSTHETIC VALVE, IS TREATED WITH ANTIBIOTICS, AND MAY REQUIRE VALVE REPLACEMENT IF ANTIBIOTIC THERAPY IS NOT EFFECTIVE. CAUSES OF PROSTHETIC VALVE ENDOCARDITIS ARE WELL DOCUMENTED IN THE LITERATURE AND ARE TYPICALLY CLASSIFIED AS EARLY (<60 DAYS) OR LATE (>60 DAYS). THE TWO CATEGORIES SHOW DEFINITE DIFFERENCES IN CLINICAL FEATURES, MICROBIAL PATTERNS, AND MORTALITY. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A TECHNICAL SUMMARY WRITTEN BY EW, EARLY PROSTHETIC VALVE ENDOCARDITIS IS USUALLY CAUSED BY PERIOPERATIVE BACTERIAL CONTAMINATION OF THE VALVE. IN EARLY CASES OF PROSTHETIC ENDOCARDITIS, SUBSEQUENT INFECTIONS MAY BE RELATED TO CONTAMINATION AT THE TIME OF SURGERY (POOR STERILE TECHNIQUE). ADDITIONAL CAUSES OF EARLY PVE CAN OCCUR FROM OTHER SOURCES OF INFECTION AND NOT FROM THE VALVE (E.G., DENTAL TREATMENT OR SURGICAL PROCEDURES INVOLVING THE UPPER RESPIRATORY TRACT, UTIS, PNEUMONIA). ADDITIONALLY, A NON-CONFORMANCE IN THE STERILITY OR PACKAGING PROCESSES OF THE VALVE MAY ALSO MANIFEST IN THE EARLY POST-OPERATIVE PERIOD CAUSING EARLY PVE. HOWEVER, THERE ARE SEVERAL TYPES OF BACTERIA THAT WOULD NOT SURVIVE THE EDWARDS STERILIZATION PROCESS. EDWARDS LIFESCIENCES PRODUCES AND PROVIDES STERILE TISSUE BIOPROSTHESIS TO ITS CUSTOMERS BY FOLLOWING CAREFULLY DESIGNED ROBUST STERILIZATION PROCESSES. THESE MANUFACTURING PROCESSES HAVE BEEN VALIDATED AND DEMONSTRATED TO CONSISTENTLY PROVIDE A SIGNIFICANT SAFETY FACTOR FROM WHICH MICROORGANISMS COULD NOT SURVIVE. MICROBIOLOGY AND PROCESS MONITORING IS ROUTINELY REVIEWED WITHIN QUALITY SYSTEMS TO MAINTAIN STERILITY CONTROL. VALIDATED TESTING HAS DEMONSTRATED THAT MICROORGANISMS COULD NOT SURVIVE EDWARDS MULTI-STAGE PROCESSING WITH ENHANCED STERILANT OR HEATED GLUTARALDEHYDE TERMINAL STERILANT SOLUTION. THESE MULTIPLE, REDUNDANT MANUFACTURING CONTROLS ENSURE THE STERILITY OF EDWARDS VALVES AS PROVIDED TO CUSTOMERS. THEREFORE, THE PROBABILITY OF ENDOCARDITIS RELATED TO EDWARDS BIOPROSTHESIS IS REMOTE. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ALTHOUGH THE EXACT CAUSE AND SOURCE OF THE ENDOCARDITIS CANNOT BE DETERMINED, DUE TO THE LIMITED INFORMATION PROVIDED. HOWEVER, IT IS POSSIBLE THAT UNKNOWN PROCEDURAL FACTORS OR PATIENT FACTORS MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22296587&pc=OZD | 1220648-2025-29603 | 2025-05-29 04:00 | Death | ABIOMED, INC. | 2025-06-20 04:00 | OZD | IMPELLA CP | Deformation Due to Compressive Stress | Hemolysis; Perforation of Vessels | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT DURING IMPELLA CP INSERTION, THERE WAS INITIALLY DIFFICULTY CROSSING THE AORTIC VALVE (AOV) WITH THE WIRE INTO THE LEFT VENTRICLE (LV) DUE TO SEVERE AORTIC STENOSIS. AS THE PHYSICIAN ATTEMPTED TO PUSH ACROSS, THE IMPELLA CANNULA KINKED AND COULD NOT MOVE FURTHER ACROSS THE AOV. ON FLUOROSCOPY, THE OUTFLOW WAS JUST ACROSS THE VALVE. THE PHYSICIAN THEN PERFORMED SINGLE ACCESS THROUGH THE14 FRENCH PEEL-AWAY TO GET A WIRE ACROSS AND ATTEMPTED VALVULOPLASTY FOR BETTER DELIVERY OF THE IMPELLA ACROSS THE VALVE. AS THE PHYSICIAN MANAGED TO GET BALLOON ACROSS THE AOV, IT PUSHED THE IMPELLA CP IN FURTHER TO A BETTER POSITION AND UNKINKED THE CANNULA. WIRES AND BALLOON WERE REMOVED AND THE IMPELLA WAS STARTED WITH GOOD FLOWS. ADDITIONALLY, UPON ARRIVAL AT THE INTENSTIVE CARE UNIT, HEMOLYSIS WAS NOTED IN THE FOLEY CATHETER. THE IMPELLA P-LEVEL WAS TURNED DOWN. THE NEXT DAY, THE IMPELLA WAS REPOSITIONED AND URINE STARTED TO CLEAR. IT WAS FURTHER REPORTED THAT THE PATIENT STARTED DISPLAYING SIGNS AND SYMPTOMS OF A GASTROINTESTINAL BLEED. THE FAMILY MADE THE DECISION TO WITHDRAW CARE, AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF PRODUCT DAMAGE (CANNULA KINK), HEMOLYSIS, AND VASCULAR DAMAGE - PERIPHERAL VESSEL HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. THE DATA LOG DID NOT RECORD ANY ABNORMALITIES RELATED TO PUMP FLOW, MOTOR CURRENT, OR POSITIONING ISSUES. SOME DISCREPANCIES IN THE PLACEMENT SIGNAL WERE NOTED WITHOUT AFFECTING THE PUMP'S PERFORMANCE DURING THE HEMOLYSIS EVENT. THE DATA LOG WAS NOT CONCLUSIVE WITHOUT THE RETURNED PRODUCT INVESTIGATION AND SUFFICIENT CLINICAL INFORMATION TO DERIVE THE ROOT CAUSE OF THE HEMOLYSIS. THE CAUSE OF THE HEMOLYSIS ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED. THE CAUSE OF THE CANNULA KINK ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED TO SEVERE AORTIC STENOSIS. THE CAUSE OF THE VASCULAR DAMAGE ISSUE WAS NOT DETERMINED SINCE SUFFICIENT PRODUCT WAS NOT RETURNED AND SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED. THE RELATION TO IMPELLA AND GASTROINTESTINAL BLEED IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22296645&pc=OZD | 1220648-2025-29604 | 2025-05-29 04:00 | Death | ABIOMED, INC. | 2025-06-20 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT CAME IN WITH A FLUTTER AND AN EJECTION FRACTION OF 5%. THE IMPELLA 5.5 WAS SUPPORTING WHEN VENTRICULAR TACHYCARDIA RUNS OCCURRED WHEN THE PATIENT LEANED FORWARD OR SOMETIMES TURNED AND TORSADES OCCURRED AS WELL. SHOCKING WAS REQUIRED. THE PUMP SUPPORTED FOR OVER TEN DAYS WHEN CARE WAS WITHDRAWN AND THE PUMP WAS EXPLANTED. THE PATIENT WENT ON COMFORT CARE AND EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22296674&pc=LWS | 9614453-2025-02374 | 2025-04-29 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-06-20 04:00 | LWS | COBALT¿ XT VR MRI SURESCAN¿ | Appropriate Term/Code Not Available | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IS DECEASED. A REPRESENTATIVE FOR THE PATIENT STATED THAT THEY BELIEVED THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) DID NOT WORK LIKE IT WAS SUPPOSED TO WHEN THE PATIENT PASSED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION WERE UNSUCCESSFUL. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22296678&pc=OZO | 2032227-2025-194547 | 2024-12-10 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-20 04:00 | OZO | 630G INSULIN PUMP MMT-1715K 630G BLACK MG | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Coma | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON 10 DEC 2024. THE EVENT INVOLVED PRODUCTS UNOMEDICAL, MMT-1715K AND MMT-332A. TROUBLESHOOTING WAS NOT PERFORMED. THE CAUSE OF DEATH WAS DUE TO DIABETIC COMA. THE BLOOD GLUCOSE VALUE WAS 1000 MG/DL. IT WAS UNKNOWN WHETHER THE CUSTOMER WAS USING THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT AND THE AUTO MODE/SMART GUARD FEATURE AT THE TIME OF THE EVENT. IT WAS UNKNOWN WHETHER THE PUMP WAS WORN AT TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL, MMT-1715K AND MMT-332A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22296696&pc=LJT | 3006260740-2025-04093 | 2025-05-27 04:00 | Death | C.R. BARD, INC. (BASD) -3006260740 | 2025-06-20 04:00 | LJT | M.R.I IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | Fracture | Insufficient Information | K873213 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME AFTER PORT PLACEMENT PROCEDURE; THE PORT CATHETER HAD FRACTURED. IT WAS FURTHER REPORTED THAT PATIENT WAS EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. Manufacturer Narrative: H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS 01-JAN-1900 FOR THE MDR SUBMISSION REQUIREMENT. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22296966&pc=FRN | 2016493-2025-88916 | 2025-05-24 04:00 | Death | CAREFUSION SD | 2025-06-20 04:00 | FRN | ALARIS SYSTEM | Contamination; Excess Flow or Over-Infusion; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K211218 |  | 1 | Event Description: IT WAS REPORTED THAT THE PUMP WAS INTENDED TO INFUSE AT A RATE OF 4 ML/HOUR. HOWEVER, MEDICATION BAG WAS FOUND TO BE EMPTY. FURTHER INFORMATION WAS RECEIVED THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT MORPHINE 100 ML WAS PROGRAMMED ON (B)(6) 2025 TO INFUSE AT 4 ML/HOUR AND DOUBLE VERIFIED. THE CLINICIAN NOTED AN AIR-IN-LINE ALARM AND EMPTY MEDICATION BAG APPROXIMATELY ONE (1) HOUR AND TWENTY (20) MINUTES LATER. THERE WAS NO REPORTED CHANGE IN THE PATIENT'S CONDITION, AND NO INTERVENTIONS WERE PROVIDED. HOWEVER, IT WAS REPORTED THE PATIENT PASSED AWAY ON (B)(6) 2025. CUSTOMER RISK MANAGEMENT STATED, "THE EVENT REPORTED WAS EQUIPMENT MALFUNCTION; HOWEVER, AT THIS POINT NOT FELT TO HAVE CONTRIBUTED TO THE DEATH."Event Description: IT WAS REPORTED THAT THE PUMP WAS INTENDED TO INFUSE AT A RATE OF 4ML/HOUR. HOWEVER, MEDICATION BAG WAS FOUND TO BE EMPTY. FURTHER INFORMATION WAS RECEIVED THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT MORPHINE 100ML WAS PROGRAMMED ON 24MAY2025 TO INFUSE AT 4ML/HOUR AND DOUBLE VERIFIED. THE CLINICIAN NOTED AN AIR-IN-LINE ALARM AND EMPTY MEDICATION BAG APPROXIMATELY ONE (1) HOUR AND TWENTY (20) MINUTES LATER. THERE WAS NO REPORTED CHANGE IN THE PATIENT'S CONDITION, AND NO INTERVENTIONS WERE PROVIDED. HOWEVER, IT WAS REPORTED THE PATIENT PASSED AWAY ON (B)(6) 2025. CUSTOMER RISK MANAGEMENT STATED "THE EVENT REPORTED WAS EQUIPMENT MALFUNCTION HOWEVER AT THIS POINT NOT FELT TO HAVE CONTRIBUTED TO THE DEATH." Manufacturer Narrative: BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE FAILURE INVESTIGATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. Manufacturer Narrative: OMIT : B21 - TYPE OF INVESTIGATION NOT YET DETERMINED, C21 - RESULTS PENDING COMPLETION OF INVESTIGATION, D16 - CONCLUSION NOT YET AVAILABLE. CORRECTION: MEDICAL DEVICE SERIAL #, UNIQUE IDENTIFIER (UDI) #. ADDITIONAL INFORMATION: DEVICE EVAL BY MANUFACTURER?, IMDRF ANNEX A, B, C, D, G CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. IT WAS DETERMINED THROUGH INVESTIGATION OF THE RETURNED DEVICES THAT THE INITIALLY REPORTED SUSPECT DEVICE WITH SERIAL NUMBER# (B)(6) IS A CONCOMITANT AND THIS FILE HAS CAPTURED THE CORRECT SUSPECT DEVICE WITH SERIAL NUMBER# (B)(6) AS PER INVESTIGATION REPORT. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER. INVESTIGATION SUMMARY: THE REPORTED INCIDENT OF AN OVER INFUSION OF MORPHINE COULD NOT BE CONFIRMED THROUGH REVIEW OF THE LOGS AND WAS NOT REPLICATED DURING DEVICE TESTING. THE INSPECTION AND TESTING OF THE PUMP MODULE DID NOT FIND ANY EXISTING MALFUNCTIONS THAT COULD HAVE CONTRIBUTED TO THE REPORTED INCIDENT. THE SUSPECT PUMP MODULE WAS FOUND TO BE INFUSING WITHIN SPECIFICATIONS AND DID NOT SHOW ANY SIGNS OF UNREGULATED FLOW OCCURRING. ALL INSPECTED PARTS WERE MANUFACTURED BY BD. REVIEW OF THE PCU EVENT LOG DID NOT SHOW THE FACILITY-REPORTED DEVICE SN (B)(6) TO HAVE BEEN ACTIVELY INFUSING ON THE DAY OF THE INCIDENT. INSTEAD, SN (B)(6) WAS OBSERVED TO HAVE ACTIVITY MATCHING THE REPORTED INFUSION OF MORPHINE, AND PUMP MODULE SN (B)(6) WAS MARKED AS THE SUSPECT DEVICE. THE SYSTEM WAS POWERED ON (B)(6) 2025 AT 2:56 PM, ¿NO¿ WAS SELECTED FOR NEW PATIENT AND ¿YES¿ WAS SELECTED TO CONFIRM THE CURRENT PROFILE ¿MED-SURG/TELE¿. THE INITIAL INFUSION FOR THE DRUG ¿MORPHINE¿ WAS PROGRAMMED AT A CONCENTRATION OF 100 MG / 100 ML AND STARTED ON (B)(6) 2025 AT 3:01 PM. THE RATE WAS PROGRAMMED AT 4 ML/HR, CALCULATING THE DOSE AT 4 MG/HR, AND THE VOLUME TO BE INFUSED (VTBI) WAS 100 ML. THE TAMPER SWITCH WAS PRESSED, LOCKING THE PANEL. AT 4:21 PM, AN AIR-IN-LINE ALARM WAS OBSERVED. THERE WERE ATTEMPTS TO RESTART THE INFUSION AND TURNING OFF THE UNIT BUT THE PANEL WAS LOCKED. AT 4:23 THE TAMPER SWITCH WAS PRESSED, UNLOCKING THE PANEL. BETWEEN 4:25 PM AND 4:28 PM, TWO (2) SAFETY CLAMP OPEN ALARMS WERE OBSERVED, WITH ONE LASTING UP TO THREE (3) SECONDS. THE INFUSION WAS THEN RESTARTED. THE INFUSION WAS PAUSED AT 4:30 PM AND THE SYSTEM WAS POWERED OFF AT 4:31 PM. THE TOTAL VOLUME INFUSED WAS CALCULATED AT 5.523 ML. IT IS NOT POSSIBLE TO DETERMINE WHAT THE ACTUAL VOLUME IS WITHIN AN IV CONTAINER AT THE START AND ENDING OF AN INFUSION FROM A REVIEW OF A PUMP MODULE OR A PCU EVENT LOG. THIS IS NOT A FUNCTION OF THE EVENT LOG; IT ONLY CAN REFLECT THE INPUTTED INFORMATION IT RECEIVES EITHER MANUALLY OR REMOTELY WITH RESPECT TO VOLUME TO BE DELIVERED. ROOT CAUSE: THE REPORTED INCIDENT OF AN OVER INFUSION OF MORPHINE COULD NOT BE CONFIRMED THROUGH REVIEW OF THE LOGS AND WAS NOT REPLICATED DURING DEVICE TESTING. THEREFORE, THE ROOT CAUSE FOR THE REPORTED OVER INFUSION OF MORPHINE WAS NOT DETERMINED. NO ANOMALIES WERE OBSERVED WITH THE PUMP MODULE THAT WOULD CONTRIBUTE TO THE REPORTED EVENT. THE RETURNED PUMP MODULE WAS FOUND TO BE INFUSING WITHIN SPECIFICATION. NOTE THAT THIS REPORT LISTS IMDRF ANNEX A1801, G04037, C0601, D15 CODES NOT ASSOCIATED WITH THE REPORTED EVENT BUT IDENTIFIED AS REPORTABLE MALFUNCTIONS OBSERVED ON THE DEVICE DURING INVESTIGATION ARE UNRELATED TO THE REPORTED ISSUE. THESE OTHER FAILURES PRESUMPTIVELY DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED ISSUE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22297257&pc=CBK | 2518422-2025-106746 | 2025-05-02 04:00 | Death | RESPIRONICS, INC. | 2025-06-20 04:00 | CBK | TRILOGY 100 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K083526 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT "A PATIENT PASSED AWAY IN THE UNIT" WHILE ON A TRILOGY 100 DEVICE. THE CUSTOMER REQUESTED THE ACTION LOG OF ALARMS DURING THE PATIENT'S USE OF THE DEVICE. DURING THE HOSPITALIZATION PERIOD, THE DEVICE ISSUED DISCONNECT AND DESATURATION ALARMS, AND THE PATIENT WAS ON A MONITOR FOR VITAL SIGNS. DESPITE RESUSCITATION EFFORTS, THE PATIENT DID NOT RECOVER. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THIS ISSUE THE CUSTOMER REPORTING THE PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2025, FOR A STROKE WHICH OCCURRED DAYS PRIOR TO THE PATIENT'S PASSING ON (B)(6) 2025. THE TIME OF THE PATIENT'S PASSING REMAINS UNCERTAIN. THE CAUSE OF DEATH IS UNKNOWN, HAS BEEN REVIEWED BY THE CLINICAL EXPERT AND DETERMINED THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. THE VENTILATOR SETTINGS REPORTED WERE MODE: CIRCUIT PASSIVE AVAP, PRESSURE SETTINGS: 45, MINIMUM OF 35, APAP SETTINGS: 25, BREATH RATE: 13 BREATHS PER MINUTE, INSPIRATORY TIME: 1.8 SECONDS, TRIGGER TYPE: AUTO TRACK, RISE TIME: 4. THE CUSTOMER STATES THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THE REASON FOR USING THE DEVICE IS OAS AND SEVERE OBESITY. THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22297405&pc=NKM | 2135147-2025-03316 | 2018-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-06-20 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "IMPLICATION OF THE DOSE OF MINERALOCORTICOID RECEPTOR ANTAGONIST FOLLOWING TRANSCATHETER EDGE-TO-EDGE MITRAL VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE PROGNOSTIC IMPACTS OF MRA DOSING IN OLDER PATIENTS RECEIVING TEER FOR SECONDARY MITRAL REGURGITATION. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT IN OLDER PATIENTS WHO UNDERWENT TEER FOR SECONDARY MITRAL REGURGITATION CAUSED BY SYSTOLIC HEART FAILURE, EVEN A LOW-DOSE MRA WAS ASSOCIATED WITH IMPROVED CLINICAL OUTCOMES COMPARED WITH NO MRA ADMINISTRATION. HOWEVER, FURTHER UP-TITRATION OF THE MRA DOSE DID NOT RESULT IN ADDITIONAL IMPROVEMENTS IN CLINICAL OUTCOMES. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS TERUHIKO IMAMURA AT TOYAMA UNIVERSITY HOSPITAL INSTITUTION WITH CORRESPONDING EMAIL TEIMAMU@MED.U-TOYAMA.AC.JP]. THE TIME FRAME OF THE STUDY WAS APRIL 2018 TO JUNE 2023. A TOTAL OF 2,026 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 2,026 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED DYSLIPIDEMIA, DIABETES, ATRIAL FIBRILLATION, HISTORY OF VENTRICULAR TACHYCARDIA AND STROKE. (B)(6), UNK MITRACLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. ATTACHMENT: ARTICLE TITLED "IMPLICATION OF THE DOSE OF MINERALOCORTICOID RECEPTOR ANTAGONIST FOLLOWING TRANSCATHETER EDGE-TO-EDGE MITRAL VALVE REPAIR". Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DYSLIPIDEMIA, DIABETES, ATRIAL FIBRILLATION, HISTORY OF VENTRICULAR TACHYCARDIA AND STROKE. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22297421&pc=DSQ | 2916596-2025-03786 | 2024-10-05 04:00 | Death | THORATEC CORPORATION | 2025-06-20 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO STROKE. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. NO FURTHER INFORMATION WAS KNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22297430&pc=LWQ | 1649833-2025-00021 | 2025-05-24 04:00 | Death | ON-X LIFE TECHNOLOGIES, INC. | 2025-06-20 04:00 | LWQ | ON-X PROSTHETIC MITRAL VALVE WITH STANDARD SEWING RING (ONXM) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P000037 |  | 1 | Event Description: ACCORDING TO THE INITIAL NOTIFICATION ON 05/25/2025, "MITRAL VALVE ONXM-25 WAS IMPLANTED AND VALVE GOT STUCK AND EXPLANTED THE VALVE. BUT UNFORTUNATELY, PATIENT DIED ON TABLE." Manufacturer Narrative: THIS INVESTIGATION IS CURRENTLY ONGOING. ANY ADDITIONAL INFORMATION WILL BE PROVIDED IN THE FOLLOW-UP REPORT. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22297440&pc=NKM | 2135147-2025-03319 | 2018-04-01 04:00 | Death | ABBOTT MEDICAL | 2025-06-20 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "VALIDATION OF MITRASCORE IN DIVERSE MITRAL REGURGITATION SUBTYPES: INSIGHTS FROM THE OCEAN-MITRAL REGISTRY" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE PERFORMANCE OF MITRASCORE ACROSS DIVERSE MR SUBTYPES. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT THIS WAS THE LARGEST EXTERNAL VALIDATION STUDY FOR MITRASCORE, DEMONSTRATING MODERATE DISCRIMINATION, GOOD CALIBRATION, AND EFFECTIVE RISK STRATIFICATION ACROSS DIVERSE MR TYPES. FRAILTY ASSESSMENT ENHANCES PREDICTIVE VALUE, ALTHOUGH ADDITIONAL VALIDATION IS REQUIRED. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS MASAHIKO ASAMI AT MITSUI MEMORIAL HOSPITAL INSTITUTION]. THE TIME FRAME OF THE STUDY WAS APRIL 2018 TO JUNE 2023. A TOTAL OF 3,764 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 3,764 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, PREVIOUS STROKE, DIABETES, HYPERTENSION, DYSLIPIDEMIA, ISCHEMIC HEART DISEASE (B)(6), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING ATRIAL FIBRILLATION, PREVIOUS STROKE, DIABETES, HYPERTENSION, DYSLIPIDEMIA, ISCHEMIC HEART DISEASE. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22297518&pc=MIH | 9612164-2025-03068 | 2025-05-02 04:00 | Death | MEDTRONIC IRELAND | 2025-06-20 04:00 | MIH | ENDURANT STENT GRAFT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ¿ A WIDER GRIP: ENDOANCHORS INCREASE PROXIMAL SEALING SURFACE DURING ENDOVASCULAR ANEURYSM REPAIR¿ THE TIME FRAME OF THIS STUDY WAS OVER A SEVEN YEAR PERIOD. 221 PATIENTS UNDERWENT EVAR WITH ENDURANT STENT GRAFTS. IN ANOTHER 31 PATIENTS, ENDOANCHORS WERE IMPLANTED ALONG WITH THE ENDURANT STENT GRAFTS. PATIENT MORTALITY WAS REPORTED IN THE PATIENT POPULATION (TOTAL 28.4%). IT WAS REPORTED ONE DEATH OCCURRED WITHIN 30 DAYS OF THE EVAR PROCEDURE. THE CAUSE OF DEATH WAS MULTIPLE ORGAN FAILURE IN A PATIENT WHO HAD A RUPTURED ANEURYSM. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS Manufacturer Narrative: MEDTRONIC RECEIVED THE FOLLOWING INFORMATION FROM A JOURNAL ARTICLE ENTITLED; ¿A WIDER GRIP: ENDOANCHORS INCREASE PROXIMAL SEALING SURFACE DURING ENDOVASCULAR ANEURYSM REPAIR¿ ACCARINO G, ESPOSITO D, CAMPOLMI M, TURCHINO D, SERRA R, MARCELLO BRACALE U ANN VASC SURG 2025; 119: 37¿45 HTTPS://DOI.ORG/10.1016/J.AVSG.2025.04.119. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22297529&pc=NPT | 2025587-2025-04356 | 2025-06-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-20 04:00 | NPT | EVOLUT PRO PLUS VALVE | Fluid/Blood Leak; Device Dislodged or Dislocated; Patient Device Interaction Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension; Myocardial Infarction; Aortic Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING A TRANSCATHETER AORTIC VALVE IMPLANTATION, DURING THE RELEASE OF THE VALVE, THE VALVE DISLODGED BY SEVERAL MILLIMETERS. AFTER THE PIGTAIL CATHETER WAS WITHDRAWN, IT WAS NOTED THAT THE VALVE DISLODGED EVEN MORE. THE PATIENT BECAME HYPOTENSIVE, AND AN OBSTRUCTION OF TWO CORONARY ARTERIES WAS OBSERVED. A BALLOON WAS INFLATED WITHIN THE VALVE TO MOVE THE VALVE TO THE ASCENDING AORTA AS A SNARE WAS NOT YET AVAILABLE. AFTER THE VALVE WAS RAISED, BLOOD FLOW WAS NOTED IN BOTH CORONARY ARTERIES ALONG WITH A RETURN OF CIRCULATION. THE VALVE WAS THEN SNARED FOR THE DELIVERY OF A SECOND VALVE, WHICH HAD DIFFICULTY, BUT WAS SUCCESSFULLY IMPLANTED AT A LOWER DEPTH. A BALLOON DILATION WAS PERFORMED USING A 20MM BALLOON, AND SUBSEQUENTLY, A CORONARY ARTERY BYPASS WAS PERFORMED. THE PATIENT DEVELOPED CARDIAC ARREST WITHOUT RETURN OF SPONTANEOUS CIRCULATION AND ULTIMATELY DIED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT A PRE-IMPLANT BALLOON DILATION WAS PERFORMED USING A 20X40 MM NON-MEDTRONIC SEMI-COMPLIANT BALLOON. IT WAS REPORTED THAT THERE WAS DIFFICULTY ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) THROUGH THE FIRST DISLODGED VALVE. THE PHYSICIAN BELIEVED THAT THE LEAFLET OF THE FIRST VALVE WAS MAKING IT DIFFICULT FOR THE DCS TO PASS, WHICH WAS ASSOCIATED WITH SMALL ANATOMY AND LACK OF A NON-MEDTRONIC GUIDEWIRE. THE SNARE THAT WAS USED TO HOLD THE DISLODGED VALVE IN PLACE DURING THE SECOND VALVE PLACEMENT MEASURED 15 MM. IT WAS CONFIRMED THAT THE PATIENT EXPERIENCED CARDIAC ARREST FOLLOWING THE DISLODGEMENT OF THE FIRST VALVE. THE POST-IMPLANT BALLOON DILATION WAS PERFORMED AFTER THE SECOND VALVE WAS IMPLANTED BECAUSE "SIGNIFICANT" PARAVALVULAR LEAK (PVL) WAS OBSERVED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS DUE TO THE PATIENT'S PREVIOUS CLINICAL CONDITION ASSOCIATED WITH INFARCTION AND SEVERE AORTIC REGURGITATION WHILE AWAITING THE IMPLANTATION OF THE SECOND VALVE. PER THE PHYSICIAN, THE FIRST VALVE COULD NOT BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH, BUT THE SECOND VALVE COULD BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH. ADDITIONALLY, THE DCS COULD BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF THE MYOCARDIAL INFARCTION WAS DUE TO THE DISLODGE OF THE FIRST VALVE THAT OBSTRUCTED THE CORONARY ARTERIES. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVPROP23-29 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID EVPROPLUS-26 (J342926); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. CORRECTED DATA: D4. FULL UDI ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW: ONE MEDIA FILE WAS PROVIDED FOR REVIEW. THE PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. FLUOROSCOPIC VALVE LOAD INSPECTION WAS PERFORMED AND EVEN THOUGH THE IMAGE IS DARK, IT APPEARS TO BE A GOOD LOAD. A PRE BALLOON AORTIC VALVULOPLASTY WAS PERFORMED PRIOR TO THE VALVE IMPLANT ATTEMPT. THE VALVE DEPTH PRIOR TO RELEASE WAS APPROXIMATELY 3MM ON THE NON-CORONARY CUSP AND 6-7MM ON THE LEFT CORONARY CUSP. DURING THE FINAL RELEASE, THE VALVE VISIBLY MIGRATES TO A SHALLOWER POSITION TO A DEPTH OF 0MM. AFTER REMOVAL OF THE DELIVERY CATHETER SYSTEM, THE VALVE APPEARS TO MIGRATE FURTHER AORTIC. A SUBSEQUENT ANGIOGRAM REVEALED COMPLETE ABSENCE OF CORONARY FLOW. A BALLOON WAS USED TO MOVE THE VALVE TO THE ASCENDING AORTA TO RESTORE CORONARY FLOW. AS STATED IN THE INSTRUCTIONS FOR USE (IFU): PHYSICIANS SHOULD USE JUDGMENT WHEN CONSIDERING REPOSITIONING A FULLY DEPLOYED BIOPROSTHESIS (FOR EXAMPLE, USING A SNARE, BALLOON, AND/OR FORCEPS). REPOSITIONING THE BIOPROSTHESIS IS NOT RECOMMENDED, EXCEPT IN CASES WHERE IMMINENT SERIOUS HARM OR DEATH IS POSSIBLE (FOR EXAMPLE, CORONARY OCCLUSION). REPOSITIONING OF A DEPLOYED VALVE MAY CAUSE AORTIC ROOT DAMA GE, CORONARY ARTERY DAMAGE, MYOCARDIAL DAMAGE, VASCULAR COMPLICATIONS, PROSTHETIC VALVE DYSFUNCTION (INCLUDING DEVICE MALPOSITION), EMBOLIZATION, STROKE, AND/OR EMERGENT SURGERY. AFTER THE VALVE WAS MOVED, MINIMAL FLOW TO THE LEFT CORONARY ARTERY AND SIGNIFICANT AORTIC REGURGITATION WERE NOTED AND CARDIOPULMONARY RESUSCITATION WAS INITIATED. A SECOND VALVE WAS IMPLANTED, AND POST TAVI CORONARY ACCESS WAS PERFORMED CONFIRMING FLOW TO BOTH THE RIGHT AND LEFT CORONARY ARTERIES. A POST IMPLANT DILATATION WAS PERFORMED BUT NO SUBSEQUENT ANGIOGRAM WAS PERFORMED TO ENSURE CORONARY FLOW REMAINED AFTER THE POST DILATATION. IT WAS REPORTED THAT A CORONARY ARTERY BYPASS WAS PERFORMED BUT THE PATIENT EVENTUALLY DEVELOPED CARDIAC ARREST AND DIED. AS STATED IN THE INSTRUCTIONS FOR USE (IFU), POTENTIAL RISKS ASSOCIATED WITH THE IMPLANTATION OF THE VALVE MAY INCLUDE, BUT ARE NOT LIMITED TO, THE FOLLOWING: DEATH; MYOCARDIAL INFARCTION, CARDIAC ARREST, CARDIOGENIC SHOCK, CARDIAC TAMPONADE; CORONARY OCCLUSION, OBSTRUCTION, OR VESSEL SPASM (INCLUDING ACUTE CORONARY CLOSURE); EMERGENT SURGICAL OR TRANSCATHETER INTERVENTION (FOR EXAMPLE, CORONARY ARTERY BYPASS, HEART VALVE REPLACEMENT, VALVE EXPLANT, PERCUTANEOUS CORONARY INTERVENTION [PCI], BALLOON VALVULOPLASTY); PROSTHETIC VALVE MIGRATION/EMBOLIZATION. UPDATED DATA: B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22297530&pc=OTD | 9612164-2025-03069 | 2025-05-02 04:00 | Death | MEDTRONIC IRELAND | 2025-06-20 04:00 | OTD | ENDO ANCHOR SYSTEM - HELI-FX AAA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K140036 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING ¿A WIDER GRIP: ENDOANCHORS INCREASE PROXIMAL SEALING SURFACE DURING ENDOVASCULAR ANEURYSM REPAIR¿ THE TIME FRAME OF THIS STUDY WAS OVER A SEVEN YEAR PERIOD. 221 PATIENTS UNDERWENT EVAR WITH ENDURANT STENT GRAFTS. IN ANOTHER 31 PATIENTS, ENDOANCHORS WERE IMPLANTED ALONG WITH THE ENDURANT STENT GRAFTS. PATIENT MORTALITY WAS REPORTED IN THE PATIENT POPULATION (TOTAL 28.4%). IT WAS REPORTED ONE DEATH OCCURRED WITHIN 30 DAYS OF THE EVAR PROCEDURE. THE CAUSE OF DEATH WAS MULTIPLE ORGAN FAILURE IN A PATIENT WHO HAD A RUPTURED ANEURYSM. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS Manufacturer Narrative: MEDTRONIC RECEIVED THE FOLLOWING INFORMATION FROM A JOURNAL ARTICLE ENTITLED; ¿A WIDER GRIP: ENDOANCHORS INCREASE PROXIMAL SEALING SURFACE DURING ENDOVASCULAR ANEURYSM REPAIR¿ ACCARINO G, ESPOSITO D, CAMPOLMI M, TURCHINO D, SERRA R, MARCELLO BRACALE U ANN VASC SURG 2025; 119: 37¿45, HTTPS://DOI.ORG/10.1016/J.AVSG.2025.04.119. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22297534&pc=NPT | 2025587-2025-04357 | 2025-06-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-20 04:00 | NPT | EVOLUT PRO PLUS VALVE | Perivalvular Leak; Patient Device Interaction Problem | Cardiac Arrest; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING A TRANSCATHETER AORTIC VALVE IMPLANTATION, DURING THE RELEASE OF THE VALVE, THE VALVE DISLODGED BY SEVERAL MILLIMETERS. AFTER THE PIGTAIL CATHETER WAS WITHDRAWN, IT WAS NOTED THAT THE VALVE DISLODGED EVEN MORE. THE PATIENT BECAME HYPOTENSIVE, AND AN OBSTRUCTION OF TWO CORONARY ARTERIES WAS OBSERVED. A BALLOON WAS INFLATED WITHIN THE VALVE TO MOVE THE VALVE TO THE ASCENDING AORTA AS A SNARE WAS NOT YET AVAILABLE. AFTER THE VALVE WAS RAISED, BLOOD FLOW WAS NOTED IN BOTH CORONARY ARTERIES ALONG WITH A RETURN OF CIRCULATION. THE VALVE WAS THEN SNARED FOR THE DELIVERY OF A SECOND VALVE, WHICH HAD DIFFICULTY, BUT WAS SUCCESSFULLY IMPLANTED AT A LOWER DEPTH. A BALLOON DILATION WAS PERFORMED USING A 20MM BALLOON, AND SUBSEQUENTLY, A CORONARY ARTERY BYPASS WAS PERFORMED. THE PATIENT DEVELOPED CARDIAC ARREST WITHOUT RETURN OF SPONTANEOUS CIRCULATION AND ULTIMATELY DIED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT A PRE-IMPLANT BALLOON DILATION WAS PERFORMED USING A 20X40 MM NON-MEDTRONIC SEMI-COMPLIANT BALLOON. IT WAS REPORTED THAT THERE WAS DIFFICULTY ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) THROUGH THE FIRST DISLODGED VALVE. THE PHYSICIAN BELIEVED THAT THE LEAFLET OF THE FIRST VALVE WAS MAKING IT DIFFICULT FOR THE DCS TO PASS, WHICH WAS ASSOCIATED WITH SMALL ANATOMY AND LACK OF A NON-MEDTRONIC GUIDEWIRE. THE SNARE THAT WAS USED TO HOLD THE DISLODGED VALVE IN PLACE DURING THE SECOND VALVE PLACEMENT MEASURED 15 MM. IT WAS CONFIRMED THAT THE PATIENT EXPERIENCED CARDIAC ARREST FOLLOWING THE DISLODGEMENT OF THE FIRST VALVE. THE POST-IMPLANT BALLOON DILATION WAS PERFORMED AFTER THE SECOND VALVE WAS IMPLANTED BECAUSE "SIGNIFICANT" PARAVALVULAR LEAK (PVL) WAS OBSERVED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS DUE TO THE PATIENT'S PREVIOUS CLINICAL CONDITION ASSOCIATED WITH INFARCTION AND SEVERE AORTIC REGURGITATION WHILE AWAITING THE IMPLANTATION OF THE SECOND VALVE. PER THE PHYSICIAN, THE FIRST VALVE COULD NOT BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH, BUT THE SECOND VALVE COULD BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH. ADDITIONALLY, THE DCS COULD BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF THE MYOCARDIAL INFARCTION WAS DUE TO THE DISLODGE OF THE FIRST VALVE THAT OBSTRUCTED THE CORONARY ARTERIES. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVPROP23-29 (LOT: UNKNOWN); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID EVPROPLUS-26 (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. CORRECTED DATA: D4. FULL UDI ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW: ONE MEDIA FILE WAS PROVIDED FOR REVIEW. THE PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. FLUOROSCOPIC VALVE LOAD INSPECTION WAS PERFORMED AND EVEN THOUGH THE IMAGE IS DARK, IT APPEARS TO BE A GOOD LOAD. A PRE BALLOON AORTIC VALVULOPLASTY WAS PERFORMED PRIOR TO THE VALVE IMPLANT ATTEMPT. THE VALVE DEPTH PRIOR TO RELEASE WAS APPROXIMATELY 3MM ON THE NON-CORONARY CUSP AND 6-7MM ON THE LEFT CORONARY CUSP. DURING THE FINAL RELEASE, THE VALVE VISIBLY MIGRATES TO A SHALLOWER POSITION TO A DEPTH OF 0MM. AFTER REMOVAL OF THE DELIVERY CATHETER SYSTEM, THE VALVE APPEARS TO MIGRATE FURTHER AORTIC. A SUBSEQUENT ANGIOGRAM REVEALED COMPLETE ABSENCE OF CORONARY FLOW. A BALLOON WAS USED TO MOVE THE VALVE TO THE ASCENDING AORTA TO RESTORE CORONARY FLOW. AS STATED IN THE INSTRUCTIONS FOR USE (IFU): PHYSICIANS SHOULD USE JUDGMENT WHEN CONSIDERING REPOSITIONING A FULLY DEPLOYED BIOPROSTHESIS (FOR EXAMPLE, USING A SNARE, BALLOON, AND/OR FORCEPS). REPOSITIONING THE BIOPROSTHESIS IS NOT RECOMMENDED, EXCEPT IN CASES WHERE IMMINENT SERIOUS HARM OR DEATH IS POSSIBLE (FOR EXAMPLE, CORONARY OCCLUSION). REPOSITIONING OF A DEPLOYED VALVE MAY CAUSE AORTIC ROOT DAMA GE, CORONARY ARTERY DAMAGE, MYOCARDIAL DAMAGE, VASCULAR COMPLICATIONS, PROSTHETIC VALVE DYSFUNCTION (INCLUDING DEVICE MALPOSITION), EMBOLIZATION, STROKE, AND/OR EMERGENT SURGERY. AFTER THE VALVE WAS MOVED, MINIMAL FLOW TO THE LEFT CORONARY ARTERY AND SIGNIFICANT AORTIC REGURGITATION WERE NOTED AND CARDIOPULMONARY RESUSCITATION WAS INITIATED. A SECOND VALVE WAS IMPLANTED, AND POST TAVI CORONARY ACCESS WAS PERFORMED CONFIRMING FLOW TO BOTH THE RIGHT AND LEFT CORONARY ARTERIES. A POST IMPLANT DILATATION WAS PERFORMED BUT NO SUBSEQUENT ANGIOGRAM WAS PERFORMED TO ENSURE CORONARY FLOW REMAINED AFTER THE POST DILATATION. IT WAS REPORTED THAT A CORONARY ARTERY BYPASS WAS PERFORMED BUT THE PATIENT EVENTUALLY DEVELOPED CARDIAC ARREST AND DIED. AS STATED IN THE INSTRUCTIONS FOR USE (IFU), POTENTIAL RISKS ASSOCIATED WITH THE IMPLANTATION OF THE VALVE MAY INCLUDE, BUT ARE NOT LIMITED TO, THE FOLLOWING: DEATH; MYOCARDIAL INFARCTION, CARDIAC ARREST, CARDIOGENIC SHOCK, CARDIAC TAMPONADE; CORONARY OCCLUSION, OBSTRUCTION, OR VESSEL SPASM (INCLUDING ACUTE CORONARY CLOSURE); EMERGENT SURGICAL OR TRANSCATHETER INTERVENTION (FOR EXAMPLE, CORONARY ARTERY BYPASS, HEART VALVE REPLACEMENT, VALVE EXPLANT, PERCUTANEOUS CORONARY INTERVENTION [PCI], BALLOON VALVULOPLASTY); PROSTHETIC VALVE MIGRATION/EMBOLIZATION. UPDATED DATA: B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22297754&pc=OZD | 1220648-2025-29617 | 2025-06-04 04:00 | Death | ABIOMED, INC. | 2025-06-20 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED DISTAL ISCHEMIA IN THE RIGHT LEG. A DISTAL REPERFUSION CATHETER WAS PLACED IN THE RIGHT LEG TO REESTABLISH FLOW. THE PATIENT EXPIRED ON SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22297755&pc=NKM | 2135147-2025-03324 | 2016-02-01 05:00 | Death | ABBOTT MEDICAL | 2025-06-20 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "TRANSCATHETER EDGE-TO-EDGE REPAIR IN PATIENTS WITH PRIMARY TRICUSPID REGURGITATION" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE SAFETY AND FEASIBILITY OF TEER IN PATIENTS WITH PRIMARY TR. DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP, AND A NON-ABBOTT DEVICE. THE ARTICLE CONCLUDED THAT TEER IS A SAFE AND EFFECTIVE OPTION FOR PRIMARY TR, PROMOTING RIGHT HEART REVERSE REMODELING AND SYMPTOMATIC RELIEF, OFFERING A VITAL ALTERNATIVE TO SURGERY IN SELECTED PATIENTS. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS ATSUSHI SUGIURA AT UNIVERSITY HOSPITAL BONN INSTITUTION WITH CORRESPONDING EMAIL ATSUSHI.SUGIURA@UKBONN.DE]. THE TIME FRAME OF THE STUDY WAS FEBRUARY 2016 TO APRIL 2023. A TOTAL OF 114 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 75 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED DIABETES, HYPERTENSION, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, CORONARY ARTERY DISEASE, PRIOR VALVULAR SURGERY, ATRIAL FIBRILLATION. (B)(6), UNK MITRACLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, SINGLE LEAFLET DEVICE ATTACHMENT (SLDA), RECURRENT TR, HEART FAILURE, PROLONGED HOSPITALIZATION, MEDICAL INTERVENTION. (B)(6), UNK TRICLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, SINGLE LEAFLET DEVICE ATTACHMENT (SLDA), RECURRENT TR, HEART FAILURE, PROLONGED HOSPITALIZATION, MEDICAL INTERVENTION. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICES WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, HYPERTENSION, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, CORONARY ARTERY DISEASE, PRIOR VALVULAR SURGERY, ATRIAL FIBRILLATION. COMPLICATIONS REPORTED INCLUDED DEATH, SINGLE LEAFLET DEVICE ATTACHMENT (SLDA), RECURRENT TR, HEART FAILURE, PROLONGED HOSPITALIZATION, MEDICAL INTERVENTION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. THE REPORTED OFF-LABEL USE OF THE DEVICE WAS ASSOCIATED WITH THE MITRACLIP DEVICE BEING USED ON THE TRICUSPID VALVE TO TREAT TRICUSPID REGURGITATION (TR). A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22297756&pc=NPS | 2135147-2025-03327 | 2016-02-01 05:00 | Death | ABBOTT MEDICAL | 2025-06-20 04:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "TRANSCATHETER EDGE-TO-EDGE REPAIR IN PATIENTS WITH PRIMARY TRICUSPID REGURGITATION" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE SAFETY AND FEASIBILITY OF TEER IN PATIENTS WITH PRIMARY TR. DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP, AND A NON-ABBOTT DEVICE. THE ARTICLE CONCLUDED THAT TEER IS A SAFE AND EFFECTIVE OPTION FOR PRIMARY TR, PROMOTING RIGHT HEART REVERSE REMODELING AND SYMPTOMATIC RELIEF, OFFERING A VITAL ALTERNATIVE TO SURGERY IN SELECTED PATIENTS. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS ATSUSHI SUGIURA AT UNIVERSITY HOSPITAL BONN INSTITUTION WITH CORRESPONDING EMAIL ATSUSHI.SUGIURA@UKBONN.DE]. THE TIME FRAME OF THE STUDY WAS FEBRUARY 2016 TO APRIL 2023. A TOTAL OF 114 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 75 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED DIABETES, HYPERTENSION, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, CORONARY ARTERY DISEASE, PRIOR VALVULAR SURGERY, ATRIAL FIBRILLATION. (B)(6), UNK MITRACLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, SINGLE LEAFLET DEVICE ATTACHMENT (SLDA), RECURRENT TR, HEART FAILURE, PROLONGED HOSPITALIZATION, MEDICAL INTERVENTION. (B)(6), UNK TRICLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, SINGLE LEAFLET DEVICE ATTACHMENT (SLDA), RECURRENT TR, HEART FAILURE, PROLONGED HOSPITALIZATION, MEDICAL INTERVENTION. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP DEVICES WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, HYPERTENSION, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, CORONARY ARTERY DISEASE, PRIOR VALVULAR SURGERY, ATRIAL FIBRILLATION. COMPLICATIONS REPORTED INCLUDED DEATH, SINGLE LEAFLET DEVICE ATTACHMENT (SLDA), RECURRENT TR, HEART FAILURE, PROLONGED HOSPITALIZATION, MEDICAL INTERVENTION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22297938&pc=NPS | 2135147-2025-03332 | 2015-08-01 04:00 | Death | ABBOTT MEDICAL | 2025-06-20 04:00 | NPS | TRICLIP DELIVERY SYSTEM (TCDS) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P230007 |  | 1 | Event Description: THE ARTICLE "ASSOCIATION BETWEEN PSOAS MUSCLE AREA AND OUTCOMES AFTER TRANSCATHETER TRICUSPID VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE ASSOCIATION OF PMA WITH PROCEDURAL AND CLINICAL OUTCOMES FOLLOWING TTVR. DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP, AND NON-ABBOTT DEVICES. THE ARTICLE CONCLUDED THAT APPROXIMATELY 30% OF PATIENTS UNDERGOING TTVR WERE FOUND TO HAVE LOW PMA. PROCEDURAL OUTCOMES AND POSTPROCEDURAL SYMPTOMATIC IMPROVEMENTS WERE COMPARABLE BETWEEN PATIENTS WITH LOW AND HIGH PMA, WHILE PATIENTS WITH LOW PMA HAD POOR PROGNOSIS AFTER TTVR. ALTHOUGH TTVR MAY BE A SAFE THERAPEUTIC OPTION TO REDUCE TR AND IMPROVE HEART FAILURE SYMPTOMS IN PATIENTS WITH SARCOPENIA, THE PROGNOSIS AFTER TTVR REMAINS POOR IN THIS POPULATION. PRE-PROCEDURAL CT-BASED ASSESSMENT OF PMA MAY ENHANCE RISK STRATIFCATION AND SUPPORT BETTER CLINICAL DECISION-MAKING FOR TTVR. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS TETSU TANAKA AT UNIVERSITY HOSPITAL BONN INSTITUTION WITH CORRESPONDING EMAIL TA.CHI.TSU.TETSU@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS AUGUST 2015 TO MAY 2022. A TOTAL OF 163 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED DIABETES, HYPERTENSION, CORONARY ARTERY DISEASE, PREVIOUS MYOCARDIAL INFARCTION, CARDIAC SURGERY, AND STROKE, ATRIAL FIBRILLATION. (B)(6), UNK MITRACLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION. (B)(6), UNK TRICLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF TRICLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, HYPERTENSION, CORONARY ARTERY DISEASE, PREVIOUS MYOCARDIAL INFARCTION, CARDIAC SURGERY, AND STROKE, ATRIAL FIBRILLATION. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22297948&pc=NKM | 2135147-2025-03331 | 2015-08-01 04:00 | Death | ABBOTT MEDICAL | 2025-06-20 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "ASSOCIATION BETWEEN PSOAS MUSCLE AREA AND OUTCOMES AFTER TRANSCATHETER TRICUSPID VALVE REPAIR" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE ASSOCIATION OF PMA WITH PROCEDURAL AND CLINICAL OUTCOMES FOLLOWING TTVR. DEVICES MENTIONED INCLUDE MITRACLIP, TRICLIP, AND NON-ABBOTT DEVICES. THE ARTICLE CONCLUDED THAT APPROXIMATELY 30% OF PATIENTS UNDERGOING TTVR WERE FOUND TO HAVE LOW PMA. PROCEDURAL OUTCOMES AND POSTPROCEDURAL SYMPTOMATIC IMPROVEMENTS WERE COMPARABLE BETWEEN PATIENTS WITH LOW AND HIGH PMA, WHILE PATIENTS WITH LOW PMA HAD POOR PROGNOSIS AFTER TTVR. ALTHOUGH TTVR MAY BE A SAFE THERAPEUTIC OPTION TO REDUCE TR AND IMPROVE HEART FAILURE SYMPTOMS IN PATIENTS WITH SARCOPENIA, THE PROGNOSIS AFTER TTVR REMAINS POOR IN THIS POPULATION. PRE-PROCEDURAL CT-BASED ASSESSMENT OF PMA MAY ENHANCE RISK STRATIFCATION AND SUPPORT BETTER CLINICAL DECISION-MAKING FOR TTVR. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS TETSU TANAKA AT UNIVERSITY HOSPITAL BONN INSTITUTION WITH CORRESPONDING EMAIL TA.CHI.TSU.TETSU@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS AUGUST 2015 TO MAY 2022. A TOTAL OF 163 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH AN UNKNOWN AMOUNT RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED DIABETES, HYPERTENSION, CORONARY ARTERY DISEASE, PREVIOUS MYOCARDIAL INFARCTION, CARDIAC SURGERY, AND STROKE, ATRIAL FIBRILLATION. (B)(6), UNK MITRACLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION. (B)(6), UNK TRICLIP. PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, HYPERTENSION, CORONARY ARTERY DISEASE, PREVIOUS MYOCARDIAL INFARCTION, CARDIAC SURGERY, AND STROKE, ATRIAL FIBRILLATION. COMPLICATIONS REPORTED INCLUDED DEATH, HEART FAILURE, PROLONGED HOSPITALIZATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. THE REPORTED OFF-LABEL USE OF THE DEVICE WAS ASSOCIATED WITH THE MITRACLIP DEVICE BEING USED ON THE TRICUSPID VALVE TO TREAT TRICUSPID REGURGITATION (TR). A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22298431&pc=NPT | 2025587-2025-04365 | 2025-05-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-20 04:00 | NPT | EVOLUT R TRANSCATHETER AORTIC VALVE | Patient Device Interaction Problem; Thickening of Material; Central Regurgitation | Dyspnea; Hemorrhage/Bleeding; Pleural Effusion; Pulmonary Edema; Thrombosis/Thrombus; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 6 YEARS AND 6 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, AN ECHOCARDIOGRAM WAS PERFORMED THAT REVEALED THE VALVE WAS FUNCTIONING NORMALLY, AND NO REGURGITATION OR STENOSIS WAS OBSERVED. APPROXIMATELY 7 YEARS AND 9 MONTHS FOLLOWING THE IMPLANT THE VALVE, THE PATIENT PRESENTED TO THE HOSPITAL WITH DYSPNEA AND A LARGE PLEURAL EFFUSION FOR WHICH THORACENTESIS WAS PERFORMED. AN ECHOCARDIOGRAM WAS PERFORMED THE NEXT DAY THAT REVEALED MODERATE TO SEVERE AORTIC TRANSVALVULAR REGURGITATION WITH A PRESSURE HALF-TIME OF 149 MILLISECONDS (MS) AND DIASTOLIC FLOW REVERSAL, BUT NO EVIDENCE OF AORTIC VALVE STENOSIS. A TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS PERFORMED THAT REVEALED SEVERE AORTIC TRANSVALVULAR REGURGITATION. ADDITIONALLY, A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED THAT REVEALED POSSIBLE LEFT ATRIAL APPENDAGE (LAA) THROMBUS AND SLIGHT LEAFLET THICKENING. THE PATIENT WAS HOSPITALIZED AND WAS BEING EVALUATED FOR A VALVE-IN-VALVE TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT A DISSECTION WAS OBSERVED IN THE ASCENDING AORTA THAT COULD NOT BE SEPARATED FROM THE STERNUM WITHOUT ENTERING THE AORTA. GIVEN THE PATIENT'S MEDICAL HISTORY OF STERNOTOMY IN THE PAST, MANTLE RADIATION TO THE CHEST WITH SIGNIFICANT MEDIASTINAL FIBROSIS, IT WAS OPTED FOR THE PATIENT TO GO ON CARDIO PULMONARY BYPASS AT THAT POINT. THE ASCENDING AND HEMIARCH REPLACEMENT WAS PERFORMED DUE TO NARROWING AND NEAR CIRCUMFERENTIAL CALCIFICATION OF THE AORTIC ROOT, ASCENDING AND AORTIC ARCH. THE PATIENT WAS PLACED ON CARDIOPULMONARY BYPASS PRIOR TO THE EXPLANT OF THE VALVE DUE TO THE AORTIC VALVE DISEASE AND THE PATIENT'S MEDICAL HISTORY. IT WAS FURTHER REPORTED THAT THE SURGICAL PROCEDURE WAS COMPLICATED BY ANEMIA, ACTIVE HEMORRHAGE, SEVERE HYPOTHERMIA, SEVERE LACTIC ACIDOSIS, DIFFICULTY OXYGENATING DUE TO VOLUME OVERLOAD AND ACUTE FULMINANT PULMONARY EDEMA. FOLLOWING SURGERY, THE PATIENT CONTINUED TO DEMONSTRATE AGGRESSIVE HEMORRHAGE THROUGH CHEST TUBES. MULTIPLE BLOOD PRODUCTS WERE BEING INFUSED INCLUDING PACKED RED BLOOD CELLS (PRBCS), PLATELETS, FRESH FROZEN PLASMA (FFP), AND CRYOPRECIPITATE. THE PATIENT THEN DEMONSTRATED METABOLIC ACIDOSIS AND MULTIPLE PUSHES OF BICARBONATE WERE ADMINISTERED THROUGHOUT THE NIGHT. THE PATIENT CONTINUED TO DEMONSTRATE HIGH LACTIC ACIDOSIS AND SUBSEQUENTLY DIED IN THE INTENSIVE CARE UNIT (ICU) 1 DAY FOLLOWING THE EXPLANT OF THE VALVE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22298700&pc=NPT | 2025587-2025-04368 | 2025-06-04 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-20 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT HAD STENOSIS THROUGHOUT THE FEMORAL AND ILIAC STENTS WHICH HAD BEEN TREATED WITH COATED STENTS. WHILE ADVANCING THE DELIVERY CATHETER SYSTEM (DCS), THIS CAUSED BOTH THE STENTS TO MIGRATE OVER THE DCS AND GUIDEWIRE. THE AORTIC STENT DETACHED AND BECAME STUCK IN THE ILIAC ARTERY STENT. THE DCS BECAME DAMAGED AND THE SYSTEM NEEDED TO BE REPLACED. THROMBOSIS AT THE LEVEL OF THE ABDOMINAL AORTA WAS REPORTED AS THE STENT BECAME COMPACTED. THE PATIENT UNDERWENT SURGERY. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PROTRUSION WAS NOTED AFTER INSERTION INTO THE PATIENT. THE MINIMUM DIAMETER OF THE ACCESS VESSEL WAS 4.2/3.5 MM BUT WAS PREVIOUSLY TREATED WITH A STENT. A 14 FRENCH MEDTRONIC (SENTRANT) EXTERNAL INTRODUCER SHEATH WAS USED FOR THE PROCEDURE. THERE WAS DIFFICULTY INSERTING THE DCS AND IT WOULD NOT PASS THE ABDOMINAL AORTA. A PROCEDURAL DELAY OCCURRED AS A RESULT OF THE PROTRUSION. SURGERY WAS PERFORMED TO REMOVE THE THROMBOSED AORTIC STENT. A NEW VALVE WAS IMPLANTED. THE PATIENT DIED AFTER THE THIRD SURGERY.Event Description: MEDTRONIC RECEIVED INFORMATION THAT DURING IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT HAD STENOSIS THROUGHOUT THE FEMORAL AND ILIAC STENTS WHICH HAD BEEN TREATED WITH COATED STENTS. WHILE ADVANCING THE DELIVERY CATHETER SYSTEM (DCS), THIS CAUSED BOTH THE STENTS TO MIGRATE OVER THE DCS AND GUIDEWIRE. THE AORTIC STENT DETACHED AND BECAME STUCK IN THE ILIAC ARTERY STENT. THE DCS BECAME DAMAGED AND THE SYSTEM NEEDED TO BE REPLACED. THROMBOSIS AT THE LEVEL OF THE ABDOMINAL AORTA WAS REPORTED AS THE STENT BECAME COMPACTED. THE PATIENT UNDERWENT SURGERY. ADDITIONAL INFORMATION WAS RECEIVED THAT THE PROTRUSION WAS NOTED AFTER INSERTION INTO THE PATIENT. THE MINIMUM DIAMETER OF THE ACCESS VESSEL WAS 4.2/3.5 MM BUT WAS PREVIOUSLY TREATED WITH A STENT. A 14 FRENCH MEDTRONIC (SENTRANT) EXTERNAL INTRODUCER SHEATH WAS USED FOR THE PROCEDURE. THERE WAS DIFFICULTY INSERTING THE DCS AND IT WOULD NOT PASS THE ABDOMINAL AORTA. A PROCEDURAL DELAY OCCURRED AS A RESULT OF THE PROTRUSION. SURGERY WAS PERFORMED TO REMOVE THE THROMBOSED AORTIC STENT. A NEW VALVE WAS IMPLANTED. THE PATIENT DIED AFTER THE THIRD SURGERY. ADDITIONAL INFORMATION WAS RECEIVED THAT THE DEATH OCCURRED ONE DAY FOLLOWING THE VALVE IMPLANT PROCEDURE. THE CAUSE OF DEATH WAS NOT KNOWN, BUT THE PHYSICIAN ATTRIBUTED THE DEATH TO THE FIRST DCS BUT NOT THE IMPLANTED VALVE OR SECOND DCS.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT DURING THE VALVE IMPLANT, THREE SURGERIES WERE PERFORMED. THE FIRST TO BRIDGE THE FEMORAL ARTERIES, THE SECOND FOR THE ABDOMINAL THROMBOSIS AND THE THIRD FAILED WHEN THEY TRIED TO REMOVE THE BLOOD CLOTS. NON-MEDTRONIC STENT (BEGRAFT) WAS USED DURING THE PROCEDURE. DURING THE VALVE ADVANCEMENT, THE PROTRUSION OCCURRED WITH THE DELIVERY CATHETER SYSTEM (DCS) CAPSULE AS IT INTERACTED WITH THE ILIAC STENTS AND SUBSEQUENT AORTIC STENT. Manufacturer Narrative: UPDATED H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW: THREE STATIC IMAGES WERE PROVIDED FOR REVIEW. THE PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW THEREFORE IT IS NOT POSSIBLE TO CONFIRM MINIMUM VESSEL DIAMETERS. IT WAS REPORTED THAT THE PATIENT¿S FEMORAL AND ILIAC ARTERIES HAD STENOSIS AND STENTS WERE IMPLANTED. IMAGES SHOW THAT THE DELIVERY CATHETER SYSTEM (DCS) INTERACTED WITH THE PERIPHERAL STENTS CAUSING DISLODGEMENT OF THE STENTS AND DAMAGE OF THE DCS. IT WAS REPORTED THAT THE PROCEDURE WAS ABORTED, AND THE PATIENT UNDERWENT SURGERY. THE MEDWATCH FOR THE DCS REFERENCED IN THE IMAGE REVIEW PROVIDED IN H10. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B2, B5, D4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22298804&pc=NIK | 9614453-2025-02379 | 2025-05-25 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-06-20 04:00 | NIK | CLARIA MRI QUAD CRT-D SURESCAN | Device Sensing Problem | Cardiac Arrest | P010031 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT¿S FAMILY MEMBER THAT "SOME FUNKY THINGS WENT ON WITH THE DEVICE". THE FAMILY FURTHER NOTED THAT AFTER DEVICE SETTINGS WERE CHANGED, THE DEVICE HAD INTERMITTENT DETECTION ISSUES. ADDITIONALLY, THE FAMILY NOTED THAT THERE WERE ¿ISSUES WITH SETTINGS¿ RESULTING IN THE PATIENT NEEDING A CARDIOVERSION. THE PATIENT, WHO HAD A CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D), IS NOTED TO HAVE LATER DIED AND THE CAUSE OF DEATH WAS CARDIAC ARREST. ADDITIONAL INFORMATION WAS REQUESTED; HOWEVER, NONE WAS RECEIVED. Manufacturer Narrative: CONTINUATION OF D10: 439888 LEAD IMPLANTED ON (B)(6) 2022. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. CORRECTION: D4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22298842&pc=MIH | 1820334-2025-00743 |  | Death | COOK INC | 2025-06-20 04:00 | MIH | ZENITH AAA ENDOVASCULAR GRAFT ILIAC LEG | Obstruction of Flow | Cardiac Arrest; Thrombosis/Thrombus | P020018 |  | 1 | Event Description: COOK WAS NOTIFIED THAT A PATIENT WHO HAD A ZENITH AAA ENDOVASCULAR GRAFT ILIAC LEG IMPLANTED, EXPERIENCED TOTAL THROMBOSIS OF THE STENT GRAFT, CARDIAC ARREST, AND DEATH. THE INFORMATION WAS PROVIDED IN THE LITERATURE ARTICLE ¿SYSTEMATIC LONG-TERM FOLLOW UP AFTER ENDOVASCULAR ABDOMINAL AORTIC ANEURYSM REPAIR WITH THE ZENITH STENT GRAFT¿ THAT WAS PUBLISHED ON 27JUN2019 IN THE EUROPEAN JOURNAL OF VASCULAR AND ENDOVASCULAR SURGERY. ¿ENDOVASCULAR ANEURYSM REPAIR (EVAR) HAS BECOME THE METHOD OF CHOICE FOR THE TREATMENT OF ABDOMINAL AORTIC ANEURYSMS (AAA).1 SINCE THE INTRODUCTION OF EVAR OVER 20 YEARS AGO, MANY SUCCESSFUL SHORT-TERM RESULTS HAVE BEEN REPORTED, BUT LONG-TERM DATA ARE STILL LIMITED. FURTHERMORE, THE INCONSISTENCY IN REPORTING THE RESULTS OF INDIVIDUAL STUDIES MAKES THE INTERPRETATION OF THE FINDINGS TROUBLESOME. THIS IS TRUE ESPECIALLY WHEN A COHORT CONSISTS OF PATIENTS TREATED WITH MULTIPLE DEVICES WITH DIFFERENT TECHNICAL QUALITIES. CONSEQUENTLY, THIS MAY CARRY THE RISK OF UNDERESTIMATING THE LONG-TERM DURABILITY OF A PARTICULAR DEVICE. DESPITE ALL THE ADVANCES OF ENDOVASCULAR TECHNIQUES AND GRAFT MATERIALS, THE LATE FAILURE OF ENDOVASCULAR REPAIR DUE TO ENDOLEAKS PERSISTS. EVEN THOUGH AWARENESS OF THE NATURE OF DIFFERENT TYPES OF ENDOLEAK AND THEIR TREATMENT OPTIONS IS INCREASING, THE RE-INTERVENTION RATE REMAINS HIGHER THAN FOLLOWING OPEN REPAIR. ACCORDING TO THE AVAILABLE DATA, THE NUMBER OF NECESSARY RE-INTERVENTIONS AFTER THE PRIMARY PROCEDURE DOES NOT SEEM TO HAVE A NEGATIVE EFFECT ON THE AAA RELATED MORTALITY AFTER EVAR; THE TREATMENT MODALITY THEREFORE PROVIDES EFFECTIVE PROTECTION AGAINST AAA RUPTURE. AT THE SAME TIME, A SIMILAR POSITIVE IMPACT ON OVERALL MORTALITY HAS NOT BEEN DEMONSTRATED. THERE ARE SEVERAL REASONS FOR THIS. FIRSTLY, AS THE POPULATION IS AGEING AND THE PREVALENCE OF AAA HAS BEEN SHOWN TO BE HIGHEST AMONG THOSE OVER 80 YEARS OF AGE, ELDERLY PATIENTS WITH AAA AND MULTIPLE COMORBIDITIES ARE INCREASINGLY BEING TREATED. SECONDLY, IT HAS BEEN IMPLIED IN THE LITERATURE THAT EVAR MAY BE ASSOCIATED WITH INCREASED CANCER MORTALITY VS. OPEN REPAIR. THE JUSTIFICATION FOR TREATMENT MUST THEREFORE ALWAYS BE CAREFULLY CONSIDERED, DESPITE THE MINIMALLY INVASIVE NATURE OF THE MODALITY. THE PURPOSE OF THIS STUDY WAS TO REPORT THE 16-YEAR OVERALL SURVIVAL, AS WELL AS AAA RUPTURE, COMPLICATION, AND RE-INTERVENTION FREE SURVIVAL AFTER ELECTIVE EVAR WITH A ZENITH STENT GRAFT AT AN ACADEMIC INSTITUTION. MATERIAL AND METHODS: BETWEEN 2000 AND 2010, 282 PATIENTS, CONSTITUTING THE STUDY GROUP, UNDERWENT ELECTIVE EVAR WITH THE ZENITH STENT GRAFT (COOK, BLOOMINGTON, IN, USA) AT A SINGLE ACADEMIC INSTITUTION. PATIENTS WITH AN ISOLATED ILIAC ARTERY ANEURYSM OR PREVIOUS OPEN AORTIC SURGERY WERE EXCLUDED FROM THE STUDY. THE INDICATION FOR INITIAL TREATMENT WAS AN ANEURYSM WITH A DIAMETER OF GREATER THAN OR EQUAL TO 55 MM IN MEN AND GREATER THAN OR EQUAL TO 50 MM IN WOMEN. PATIENTS WITH AN INCREASE IN AAA DIAMETER OF GREATER THAN OR EQUAL TO 5 MM OVER A SIX MONTH PERIOD, AS WELL AS THOSE IN WHOM THE AAA BECAME SYMPTOMATIC, WERE ALSO TREATED. ALL PROCEDURES WERE PERFORMED BY A VASCULAR SURGEON TOGETHER WITH AN INTERVENTIONAL RADIOLOGIST IN A HYBRID SUITE. THE PATIENTS WERE MAINLY TREATED UNDER SPINAL ANESTHESIA, AND THE STENT GRAFTS WERE IMPLANTED THROUGH THE COMMON FEMORAL ARTERY. BIFURCATED GRAFTS WERE USED IN MOST OF THE EVAR PROCEDURES, BUT AN AORTO-UNI-ILIAC GRAFT WAS EMPLOYED IF ONE OF THE ILIAC ARTERIES WAS OCCLUDED OR SEVERELY STENOSED. AN ATTEMPT TO EMBOLISE A PATENT INFERIOR MESENTERIC ARTERY (IMA) WAS ALWAYS MADE PRIOR TO STENT GRAFT PLACEMENT. THE PATIENTS WERE SYSTEMATICALLY FOLLOWED AFTER THE PROCEDURE ACCORDING TO THE PROTOCOL IN TABLE 1. BASED ON THE AVAILABLE DATA AND THE AUTHORS¿ EXPERIENCE AT THE TIME, THE INITIAL SURVEILLANCE PROTOCOL WAS MODIFIED BY REPLACING THE ANNUAL COMPUTED TOMOGRAPHY ANGIOGRAPHY (CTA) SCAN WITH A COLOUR DUPLEX ULTRASONOGRAPHY (CDUS) IN 2005. THEREAFTER, ALL CDUS EXAMINATIONS WERE PERFORMED BY EXPERIENCED VASCULAR SURGEONS. THE CTA SCANS WERE CONTINUED FOR ALL PATIENTS, IF NOT CONTRAINDICATED, 24 MONTHS AFTER THE INITIAL PROCEDURE TO CONFIRM THE RELIABILITY OF THE CDUS EXAMINATIONS. FOR OBESE INDIVIDUALS AND FOR PATIENTS WITH A SUSPECTED COMPLICATION AT CDUS, AN ADDITIONAL CTA SCAN WAS PERFORMED. A PLAIN ABDOMINAL XRAY WAS TAKEN ANNUALLY UP UNTIL 2013. THE PATIENTS WERE FOLLOWED ANNUALLY UNTIL THE END OF MARCH 2017. AT THE END OF FOLLOW UP, THE STATUS OF EACH PATIENT WAS CONFIRMED (I.E., DEAD OR ALIVE), IN ADDITION TO OBTAINING INFORMATION ON POSSIBLE EVAR RELATED REINTERVENTIONS UP UNTIL THE END OF THE FOLLOW UP. FOR SOME INDIVIDUALS, LONG-TERM SURVEILLANCE WAS CARRIED OUT IN THE REFERRING HOSPITALS THEY HAD INITIALLY VISITED PRIOR TO THE ENDOVASCULAR AAA TREATMENT. THE REFERRING VASCULAR COLLEAGUES WERE CONTACTED ACCORDINGLY TO OBTAIN FOLLOW UP DATA. PATIENTS WERE CONTACTED IF A DELAY IN A FOLLOW UP EXAMINATION WAS NOTED. ALL CTAS WERE RE-ANALYSED FOR THE STUDY. THE PRIMARY OUTCOMES WERE OVERALL AND AAA RUPTURE FREE SURVIVAL, WHILE THE SECONDARY OUTCOMES INCLUDED GRAFT RELATED COMPLICATION FREE AND RE-INTERVENTION FREE SURVIVAL. DEATHS WERE ASCERTAINED BY RECORD LINKAGE BETWEEN THE STUDY POPULATION AND THE NATIONAL CAUSES OF DEATH REGISTER ON THE BASIS OF THE PERSONAL IDENTIFICATION CODE UNIQUE TO EVERY RESIDENT. SPSS 22.0 FOR WINDOWS WAS USED FOR STATISTICAL ANALYSIS (IBM, ARMONK, NY, USA). KAPLAN-MEIER SURVIVAL ANALYSIS WAS USED TO EXAMINE SURVIVAL. UNI- AND MULTIVARIABLE ANALYSES WERE USED TO ASSESS THE EFFECT OF AGE, SEX, AAA DIAMETER, NECK DIAMETER, NECK LENGTH, NECK ANGULATION, PATENT IMA, PRIMARY TYPE II ENDOLEAK, FOLLOWING THE INSTRUCTIONS FOR USE (IFU), AND OF BEING LOST TO FOLLOW UP ON ANY STENT GRAFT RELATED RE-INTERVENTION (DEFINED LATER IN TABLE 5) AND AAA RUPTURE. A PRIMARY ENDOLEAK WAS DEFINED AS AN ENDOLEAK OBSERVED IN THE PRIMARY PROCEDURE OR IN THE FIRST POST-OPERATIVE CTA ON THE SECOND OR THIRD DAY. STATISTICAL SIGNIFICANCE WAS INDICATED BY A P VALUE LESS THAN OR EQUAL TO .05. RESULTS: THE MEDIAN AGE OF THE COHORT WAS 76 YEARS (INTERQUARTILE RANGE [IQR] 11; RANGE 49-92 YEARS) AND MEAN ANEURYSM DIAMETER WAS 61 MM (RANGE 40-110 MM) AT THE TIME OF THE TREATMENT. THE IFU ISSUED FOR THE IMPLANTATION OF THE STENT GRAFT WERE FOLLOWED IN 90% OF THE PATIENTS. PATIENT CHARACTERISTICS ARE PRESENTED IN TABLE 2. THE MEDIAN FOLLOW UP WAS 76 MONTHS (IQR 69; RANGE 0-201 MONTHS). THE STATUS OF EACH PATIENT WAS CONFIRMED AT THE END OF THE FOLLOW UP (I.E., DEAD OR ALIVE). FOR EACH PATIENT, INFORMATION WAS AVAILABLE REGARDING THE LATEST ANNUAL FOLLOW UP EXAMINATION AND POSSIBLE EVAR RELATED RE-INTERVENTIONS AFTER THAT UP TO THE END OF FOLLOW UP. IN SOME CASES, THE FOLLOW UP INTERVAL WAS PROLONGED, OR THE SURVEILLANCE DISCONTINUED AFTER UNCOMPLICATED SURVEILLANCE, IF THE PATIENT WAS PHYSICALLY FRAIL OR IF LIFE EXPECTANCY SHORT. THESE PATIENTS WERE REPORTED AS LOST TO FOLLOW UP. SPINAL ANAESTHESIA WAS USED IN MOST CASES (N = 271;96%). A BIFURCATED STENT GRAFT WAS EMPLOYED IN 268 PATIENTS (95%) AND A UNI-ILIAC STENT IN 14 CASES (5%). THE IMA WAS PATENT IN 186 PATIENTS, AND WAS SUCCESSFULLY EMBOLIZED DURING THE INITIAL PROCEDURE PRIOR TO THE STENT GRAFT PLACEMENT IN 147 CASES (79%). THERE WERE NO INTRA-OPERATIVE DEATHS OR PRIMARY CONVERSIONS. THE THIRTY DAY MORTALITY WAS 1.8% (N = 5), AND THE CAUSES OF DEATH WERE BRAINSTEM INFARCTION (N =1), CARDIAC FAILURE (N = 3), AND CARDIAC ARREST DUE TO TOTAL STENT GRAFT THROMBOSIS (N = 1). THERE WAS NO SIGNIFICANT DIFFERENCE IN 30 DAY MORTALITY BETWEEN PATIENTS UNDER AND OVER 80 YEARS OF AGE (1.5% VS. 2.3%; P = .500). THE OVERALL SURVIVAL RATES WERE 93% (SE 0.02), 83% (SE 0.03), 61% (SE 0.08), 25% (SE 0.16), AND 9% (SE 0.19) AT 1, 2, 5, 10, AND 16 YEARS, RESPECTIVELY (FIG. 1). THE CORRESPONDING FIGURES FOR THOSE AGED GREATER THAN OR EQUAL TO 80 YEARS WERE 93% (SE 0.03), 80% (SE 0.04), 46% (SE 0.19), AND 4% (SE 0.42) AT 1, 2, 5, AND 10 YEARS, RESPECTIVELY. THERE WAS NO SIGNIFICANT DIFFERENCE IN SURVIVAL BETWEEN PATIENTS UNDER AND OVER 80 YEARS OF AGE AT TWO YEARS OF SURVEILLANCE (P = .529). THE MOST COMMON CAUSES OF DEATH WERE A CARDIAC EVENT (N =70; 32%) AND CANCER (N =48; 22%). THE AUTOPSY RATE WAS 26%. THERE WERE 10 (3.5%) AAA RUPTURES DURING FOLLOW UP (TABLE 3). ALL RUPTURES WERE CONFIRMED BY CTA (N = 8) OR AUTOPSY (N = 2). MEAN TIME TO RUPTURE WAS 81 MONTHS (RANGE 22 - 173 MONTHS). CONSEQUENTLY, THE CUMULATIVE AAA RUPTURE FREE SURVIVAL WAS 100%, 98% (SE 0.01), 96% (SE 0.02), AND 79% (SE 0.12) AT 1, 5, 10, AND 16 YEARS, RESPECTIVELY. EIGHT OF THE 10 PATIENTS WITH A RUPTURED AAA (RAAA) REACHED THE HOSPITAL, AND TWO OF THEM UNDERWENT AN EMERGENCY OPEN REPAIR. ONE OF THESE PATIENTS DIED IN THE INTENSIVE CARE UNIT TWO DAYS AFTER THE PROCEDURE, AND THE OTHER WAS STILL ALIVE AT THE END OF THE FOLLOW UP. IN THREE RAAA CASES, AN ADDITIONAL PROXIMAL CUFF WAS PLACED DURING AN EMERGENCY PROCEDURE, BUT ADEQUATE PROXIMAL SEALING WAS NOT ACHIEVED, AND THE PATIENTS DIED POST-OPERATIVELY. THEY WERE NOT CONSIDERED CANDIDATES FOR OPEN CONVERSION. IN THE REMAINING THREE CASES, ENDOVASCULAR REPAIR WAS NOT CONSIDERED TECHNICALLY POSSIBLE AND OPEN REPAIR WAS REFUSED OWING TO THE PATIENTS¿ HIGH AGE AND MULTIPLE COMORBIDITIES. DURING SURVEILLANCE, ONE PATIENT WITH A RAAA WAS FOUND TO HAVE A TYPE IB ENDOLEAK, BUT, AT THE AGE OF 90 YEARS, HE WAS SEVERELY DEMENTED AND NO FURTHER PROCEDURES WERE PERFORMED. IN ONE CASE, THE RUPTURE WAS UNEXPECTEDLY FOUND AT AUTOPSY 14 YEARS AFTER THE PRIMARY TREATMENT. THIS FEMALE PATIENT HAD REFUSED FURTHER FOLLOW UP AFTER 9 YEARS OF UNCOMPLICATED SURVEILLANCE, WITH A SHRUNKEN AAA DIAMETER OF 30 MM. THE NINE OTHER PATIENTS WERE UNDER SURVEILLANCE AT THE TIME OF THE RAAA. SEVEN OF THE 10 AAA RUPTURES WERE DUE TO TYPE I ENDOLEAK (IA: N = 5; IB: N= 2). SIX OF THESE PATIENTS HAD A PERSISTENT PRIMARY TYPE II ENDOLEAK IN EARLY SURVEILLANCE PRIOR TO THE DEVELOPMENT OF THE TYPE I ENDOLEAK. IN THREE OF THESE CASES, AS THE ANEURYSM SAC SHOWED ENLARGEMENT IN SURVEILLANCE, REPEATED UNSUCCESSFUL ATTEMPTS TO EMBOLISE THE TYPE II ENDOLEAK WERE MADE. IN THREE CASES, HOWEVER, THE TYPE II ENDOLEAK SEEMED STABLE, WITH NO DETECTABLE CHANGE IN ANEURYSM SAC SIZE DURING EARLY SURVEILLANCE, BUT AAA ENLARGEMENT WAS NOTED APPROXIMATELY THREE YEARS AFTER THE PRIMARY TREATMENT. BY THAT TIME, THESE PATIENTS HAD ALSO DEVELOPED A TYPE I ENDOLEAK BEFORE AAA RUPTURE. ONE RUPTURE PATIENT REFUSED FURTHER PROCEDURES FOR THE TYPE IA ENDOLEAK AND DIED THREE YEARS AFTER THE INITIAL PROCEDURE. IN THE CASE OF THREE OF THE 10 RUPTURES, NO SIGNIFICANT ENDOLEAK HAD BEEN NOTED DURING SURVEILLANCE. ONE PATIENT WITH A PRIMARY TYPE II ENDOLEAK SHOWED CLEAR SHRINKAGE OF THE ANEURYSM SAC DURING THE EARLY SURVEILLANCE. HOWEVER, AT 22 MONTHS, THE ANEURYSM RUPTURED. ANOTHER PATIENT WAS FOLLOWED FOR ENDOTENSION, WITH NO SIGNS OF AN ENDOLEAK, AND HE DIED OF A RAAA FIVE YEARS AFTER THE PRIMARY PROCEDURE. ONE RUPTURE, AS MENTIONED EARLIER, WAS FOUND AT AUTOPSY WITH NO PRIOR HISTORY OF AN ENDOLEAK. THE UNIVARIABLE ANALYSIS SHOWED A SIGNIFICANT ASSOCIATION BETWEEN PRIMARY TYPE II ENDOLEAK AND AAA RUPTURE (ODDS RATIO [OR] 4.09, 95% CI 1.12 - 14.9; P = .033). AN EVEN MORE SIGNIFICANT ASSOCIATION WAS FOUND BETWEEN PRE-OPERATIVE ANEURYSM SIZE AND LATE AAA RUPTURE (OR 1.06, 95% CI 1.02 - 1.11; P = .007). THESE RESULTS ALSO REMAINED CONSTANT IN THE MULTIVARIABLE ANALYSIS, WITH AN OR OF 4.65 (95% CI 1.14 - 18.9; P = .032) FOR TYPE II ENDOLEAK AND 1.07 (95% CI 1.02 - 1.13; P= .009) FOR PREOPERATIVE ANEURYSM SIZE. CONSEQUENTLY, AT THE END OF THE 16 YEAR FOLLOW UP, THE MEAN ANNUAL ANEURYSM RUPTURE RATE WAS 0.49% (TABLE 4). A TOTAL OF 168 GRAFT RELATED COMPLICATIONS WERE SEEN IN A TOTAL OF 119 PATIENTS (42%). ALL PRIMARY TYPE II ENDOLEAKS WERE TAKEN INTO ACCOUNT, EVEN IF THEY WERE DETECTED ONLY IN THE PERI-OPERATIVE ANGIOGRAM OR ON THE FIRST CTA. ALL COMPLICATIONS ARE LISTED IN TABLE 3. THE FREEDOM FROM ANY STENT GRAFT RELATED COMPLICATION WAS 68% (SE 0.03), 63% (SE 0.05), 58% (SE 0.09), 54% (SE 0.17), AND 52% (SE 0.21) AT 1, 2, 5, 10, AND 16 YEARS, RESPECTIVELY (FIG. 2). COMPLICATIONS SEEMED TO ACCUMULATE DURING THE EARLY YEARS OF SURVEILLANCE BUT WERE ENCOUNTERED UP TO 14 YEARS AFTER THE PRIMARY PROCEDURE. A TYPE II ENDOLEAK WAS THE MOST COMMON COMPLICATION (N = 79; 28%), WITH ALMOST HALF (48%) OF THEM SEALING SPONTANEOUSLY AND THE VAST MAJORITY (89%) OCCURRING DURING THE FIRST TWO YEARS OF FOLLOW UP. OF THOSE WITH A PATENT IMA (39 PATIENTS) AFTER THE INITIAL PROCEDURE, 18 (46%) DEVELOPED EITHER A TYPE II ENDOLEAK (N =14) OR ENDOTENSION (N = 4). MOST OF THESE PATIENTS (88%) WITH EITHER A TYPE II ENDOLEAK OR ENDOTENSION ALSO HAD A PATENT IMA ON SURVEILLANCE, AND ONLY TWO OF THEM REQUIRED LATER EMBOLISATION. IN ALL, ONLY 26 (33%) OF ALL TYPE II ENDOLEAKS REQUIRED ENDOVASCULAR TREATMENT DUE TO ANEURYSM SAC ENLARGEMENT OVER THE STUDY PERIOD. THE PRIMARY EMBOLIZATION STRATEGY WAS COIL EMBOLISATION; IF THAT PROVED UNSUCCESSFUL, A SECONDARY ATTEMPT WAS MADE WITH GLUE. A TOTAL OF 13 PATIENTS WERE DIAGNOSED WITH ENDOTENSION (GREATER THAN OR EQUAL TO 5 MM) WITHOUT A DETECTABLE ENDOLEAK DURING THE SURVEILLANCE. FOR ONE PATIENT, THERE WAS AN ATTEMPT TO TREAT THE ENDOTENSION BY RELINING THE OLD GRAFT, WHILE THE REST WERE FOLLOWED CAREFULLY. TWO OF THESE PATIENTS LATER DEVELOPED A TYPE I ENDOLEAK, AND THEY WERE TREATED AT THAT POINT WITH AN ADDITIONAL CUFF AND LIMB GRAFT. NONE OF THE PATIENTS WITH ENDOTENSION SUFFERED A RAAA. RE-INTERVENTION FOR A GRAFT RELATED COMPLICATION WAS REQUIRED FOR 61 PATIENTS (22%). ALL ADDITIONAL PROCEDURES ARE LISTED IN TABLE 5. THE RATE OF FREEDOM FROM GRAFT RELATED REINTERVENTIONS WAS 95% (SE 0.01), 91% (SE 0.03), 80% (SE 0.08), 73% (SE 0.11), AND 70% (SE 0.16) AT 1, 2, 5, 10, AND 16 YEARS, RESPECTIVELY (FIG. 3). ONLY FIVE (1.8%) LATE CONVERSIONS TO OPEN REPAIR WERE REQUIRED, TWO AS AN EMERGENCY REPAIR OF A RUPTURE AS MENTIONED EARLIER. TWO PATIENTS HAD A PERSISTENT TYPE I ENDOLEAK DESPITE THE PLACEMENT OF AN EXTRA PROXIMAL CUFF, AND CONVERSIONS WERE PERFORMED AT 12 AND 27 MONTHS, RESPECTIVELY. ONE PATIENT HAD A STENT FRACTURE AND TYPE III ENDOLEAK 149 MONTHS AFTER THE PRIMARY TREATMENT. THE OTHER CASE OF TYPE III ENDOLEAK DEVELOPED AS A RESULT OF LIMB GRAFT MIGRATION AND WAS SUCCESSFULLY TREATED WITH AN ADDITIONAL LIMB GRAFT AT 47 MONTHS. ALL ELECTIVE CONVERSIONS WERE UNCOMPLICATED. A PRIMARY TYPE II ENDOLEAK WAS FOUND TO BE THE ONLY INDEPENDENT SIGNIFICANT FACTOR TO INCREASE THE RISK OF ANY LATE RE-INTERVENTION (OR 6.79, 95% CI 3.61 - 12.79; P < .001). DURING FOLLOW UP, ANEURYSM SAC SHRINKAGE OF _5 MM WAS SEEN IN 61% OF THE CASES, WITH AT MEAN DECREASE OF 18 MM (RANGE 5 - 41 MM) IN MAXIMUM DIAMETER.¿ REFERENCE: VÄÄRÄMÄKI, S., SALENIUS, J.-P., PIMENOFF, G., UURTO, I., & SUOMINEN, V. (2019). SYSTEMATIC LONG-TERM FOLLOW UP AFTER ENDOVASCULAR ABDOMINAL AORTIC ANEURYSM REPAIR WITH THE ZENITH STENT GRAFT. EUROPEAN JOURNAL OF VASCULAR AND ENDOVASCULAR SURGERY, 58(2), 182¿188. HTTPS://DOI.ORG/10.1016/J.EJVS.2019.03.006. THE FOCUS OF THIS REPORT IS THE PATIENT WHO HAD ZENITH AAA ENDOVASCULAR GRAFT ILIAC LEG¿S IMPLANTED AND EXPERIENCED TOTAL THROMBOSIS OF STENT GRAFTS, CARDIAC ARREST, AND DEATH. Manufacturer Narrative: B2 - DATE OF DEATH: DATE OF DEATH IS UNKNOWN. THE DATE PROVIDED IS THE PUBLICATION DATE OF THE JOURNAL ARTICLE. B3 - DATE OF EVENT: 01/01/2000 TO 03/31/2017. D6A - IMPLANT DATE: (B)(6) 2000 TO (B)(6) 2010. E1 - PHONE: (B)(6). H3 - DEVICE EVALUATED BY MFG?: DEVICE NOT RETURNED TO MANUFACTURER. THIS REPORT INCLUDES INFORMATION KNOWN AT THIS TIME. A FOLLOW-UP REPORT WILL BE SUBMITTED SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803 AND IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT CONTAINED HEREIN IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED OR THAT A DEATH OR SERIOUS INJURY OCCURRED; NOR IS IT ADMISSION THAT ANY COOK DEVICE CAUSED, CONTRIBUTED TO, OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22298933&pc=LWS | 9614453-2025-02381 | 2023-05-29 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-06-20 04:00 | LWS | EVERA MRI XT DR SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY FOUR MONTHS AFTER IMPLANT OF AN IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) THE PATIENT PASSED AWAY. THE PATIENT'S CAUSE OF DEATH IS NOT KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22298956&pc=LWS | 2017865-2025-87850 | 2025-06-05 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-20 04:00 | LWS | NEUTRINO¿ NXT DR | Failure to Capture; Over-Sensing; Device Sensing Problem | No Clinical Signs, Symptoms or Conditions | P910023 |  | 1 | Event Description: IT WAS REPORTED THROUGH REMOTE TRANSMISSION THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR EXHIBITED OVERSENSING, LOW SENSING AMPLITUDES, AND FAILURE TO CAPTURE. IT WAS BELIEVED THAT THE LOSS OF CAPTURE ON TRANSMISSION WAS DUE TO THE PATIENT HAVING PASSED. CAUSE OF DEATH WAS UNKNOWN AND THERE WAS NO ALLEGATION THAT THE DEVICE WAS RELATED TO PATIENT DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22298988&pc=OAE | 3008452825-2025-00288 | 2025-06-12 04:00 | Death | ABBOTT MEDICAL | 2025-06-20 04:00 | OAE | TACTIFLEX¿ ABLATION CATHETER, SENSOR ENABLED¿ | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P220013 |  | 1 | Event Description: DURING AN ATRIAL FIBRILLATION PROCEDURE, A PERICARDIAL EFFUSION OCCURRED AND THE PATIENT EXPIRED. AT THE END OF ABLATION THE PATIENT BECAME HYPOTENSIVE. AN ULTRASOUND CONFIRMED THE EFFUSION AND A PERICARDIOCENTESIS WAS PERFORMED STABILIZING THE PATIENT. THE PATIENT BECAME HYPOTENSIVE ONCE AGAIN A FEW MINUTES LATER AND CARDIAC MASSAGE WAS ATTEMPTED, HOWEVER, THE PATIENT EXPIRED. IT IS ALLEGED THAT THE PERICARDIAL EFFUSION MOST LIKELY OCCURRED DURING ABLATION WITH THE ABLATION CATHETER. THERE WERE NO ALLEGED PERFORMANCE ISSUES WITH ANY ABBOTT DEVICES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22299019&pc=NTE | 2124215-2025-40232 | 2025-05-22 04:00 | Death | SILK ROAD MEDICAL | 2025-06-20 04:00 | NTE | ENROUTE TRANSCAROTID NEUROPROTECTION SYSTEM | Adverse Event Without Identified Device or Use Problem | Paralysis; Ischemia Stroke |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED AN EMBOLIC STROKE 2 HOURS POST TCAR, RESULTING IN ADDITIONAL INTERVENTION, HOSPITALIZATION, AND ULTIMATELY DEATH. AN ENROUTE TRANSCAROTID NEUROPROTECTION SYSTEM (NPS) WAS SELECTED FOR USE IN THE LEFT SIDED TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) PROCEDURE. THE PATIENT HAD A TOTAL OCCLUSION OF THE RIGHT SUBCLAVIAN. THE ARTERIAL LINE WAS NOT WORKING WELL, AND WAS MEASURING BLOOD PRESSURE AS FLAT, RANGING FROM 77-87. THE SURGEON ELECTED TO MONITOR BLOOD PRESSURE INSTEAD THROUGH THE BLOOD PRESSURE CUFF. THROUGHOUT THE CASE CUFF READINGS RANGED FROM 125-130. THE ARTERIAL AND VENOUS SHEATHS WERE PLACED, AND FLOW REVERSAL WAS CONFIRMED. A 8MM X 40MM ENROUTE STENT WAS ADVANCED, AND THEN A TCAR TIMEOUT WAS PERFORMED TO TAKE SOME IMAGES. THE SURGEON SUSPECTED THERE WAS A SMALL DISSECTION, AND ELECTED TO PLACE AN ADDITIONAL ENROUTE STENT, 7MM X 30MM, DISTALLY TO THE 8MM X 40MM, OVERLAPPING. FURTHER IMAGING WAS REVIEWED, AND THE SURGEON WAS SATISFIED WITH THE STENT PLACEMENT. THE TCAR PROCEDURE WAS COMPLETED, AND THE ARTERIOTOMY WAS CLOSED. THE PATIENT AWOKE AND WAS UNABLE TO MOVE HIS RIGHT ARM. THE SURGEON ELECTED TO RE-OPEN THE PATIENT AND PERFORM ANOTHER TCAR. HE BELIEVED THE STENTED PORTION NEEDED TO EXTEND INTO THE COMMON CAROTID ARTERY (CCA). THE SURGEON ACCESSED THE CCA, AND REVIEWED THE EXTERNAL AND INTERNAL CAROTID ARTERIES ALONG WITH THE BRAIN FROM IMAGING. HE STATED EVERYTHING LOOKED GOOD, AND THERE WERE NO BLOCKAGES IN THE STENTS OR THE CRANIAL VESSELS. HOWEVER, THE SURGEON DID SUSPECT THERE MIGHT BE A DISSECTION AT THE BIFURCATION OF THE CAROTID. A THIRD STENT WAS PLACED TO COVER THE DISSECTION. SUBSEQUENT IMAGING REVEALED A BLOCKAGE IN THE CRANIUM. A NON-BOSTON SCIENTIFIC THROMBECTOMY VACUUM WAS USED TO SUCTION OUT AND REMOVE THE CLOT. THE PATIENT WOKE UP POST-PROCEDURE, BUT WAS STILL UNABLE TO MOVE HIS RIGHT ARM. THE PATIENT WAS SENT TO THE INTENSIVE CARE UNIT (ICU). THE PATIENT ULTIMATELY EXPIRED OVER THE WEEKEND. NO FURTHER INFORMATION WAS AVAILABLE DESPITE REQUEST BY BOSTON SCIENTIFIC.Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED AN EMBOLIC STROKE 2 HOURS POST TCAR, RESULTING IN ADDITIONAL INTERVENTION, HOSPITALIZATION, AND ULTIMATELY DEATH. AN ENROUTE TRANSCAROTID NEUROPROTECTION SYSTEM (NPS) WAS SELECTED FOR USE IN THE LEFT SIDED TRANSCAROTID ARTERY REVASCULARIZATION (TCAR) PROCEDURE. THE PATIENT HAD A TOTAL OCCLUSION OF THE RIGHT SUBCLAVIAN. THE ARTERIAL LINE WAS NOT WORKING WELL, AND WAS MEASURING BLOOD PRESSURE AS FLAT, RANGING FROM 77-87. THE SURGEON ELECTED TO MONITOR BLOOD PRESSURE INSTEAD THROUGH THE BLOOD PRESSURE CUFF. THROUGHOUT THE CASE CUFF READINGS RANGED FROM 125-130. THE ARTERIAL AND VENOUS SHEATHS WERE PLACED, AND FLOW REVERSAL WAS CONFIRMED. A 8MM X 40MM ENROUTE STENT WAS ADVANCED, AND THEN A TCAR TIMEOUT WAS PERFORMED TO TAKE SOME IMAGES. THE SURGEON SUSPECTED THERE WAS A SMALL DISSECTION, AND ELECTED TO PLACE AN ADDITIONAL ENROUTE STENT, 7MM X 30MM, DISTALLY TO THE 8MM X 40MM, OVERLAPPING. FURTHER IMAGING WAS REVIEWED, AND THE SURGEON WAS SATISFIED WITH THE STENT PLACEMENT. THE TCAR PROCEDURE WAS COMPLETED, AND THE ARTERIOTOMY WAS CLOSED. THE PATIENT AWOKE AND WAS UNABLE TO MOVE HIS RIGHT ARM. THE SURGEON ELECTED TO RE-OPEN THE PATIENT AND PERFORM ANOTHER TCAR. HE BELIEVED THE STENTED PORTION NEEDED TO EXTEND INTO THE COMMON CAROTID ARTERY (CCA). THE SURGEON ACCESSED THE CCA, AND REVIEWED THE EXTERNAL AND INTERNAL CAROTID ARTERIES ALONG WITH THE BRAIN FROM IMAGING. HE STATED EVERYTHING LOOKED GOOD, AND THERE WERE NO BLOCKAGES IN THE STENTS OR THE CRANIAL VESSELS. HOWEVER, THE SURGEON DID SUSPECT THERE MIGHT BE A DISSECTION AT THE BIFURCATION OF THE CAROTID. A THIRD STENT WAS PLACED TO COVER THE DISSECTION. SUBSEQUENT IMAGING REVEALED A BLOCKAGE IN THE CRANIUM. A NON-BOSTON SCIENTIFIC THROMBECTOMY VACUUM WAS USED TO SUCTION OUT AND REMOVE THE CLOT. THE PATIENT WOKE UP POST-PROCEDURE, BUT WAS STILL UNABLE TO MOVE HIS RIGHT ARM. THE PATIENT WAS SENT TO THE ICU. THE PATIENT ULTIMATELY EXPIRED OVER THE WEEKEND. NO FURTHER INFORMATION WAS AVAILABLE DESPITE REQUEST BY BOSTON SCIENTIFIC. Manufacturer Narrative: B2 - DATE OF DEATH: DATE OF DEATH WAS APPROXIMATED TO (B)(6) 2025, AS THE PATIENT WAS REPORTED TO HAVE EXPIRED OVER THE WEEKEND FOLLOWING THE TCAR PROCEDURE. UPDATED FIELDS: H1 - TYPE OF REPORTABLE EVENT: UPDATED TO DEATH. H6 - EVALUATION METHOD CODES, EVALUATION CONCLUSION CODES. Manufacturer Narrative: B2 - DATE OF DEATH: DATE OF DEATH WAS APPROXIMATED TO (B)(6) 2025, AS THE PATIENT WAS REPORTED TO HAVE EXPIRED OVER THE WEEKEND FOLLOWING THE TCAR PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22299167&pc=OZD | 1220648-2025-29620 | 2025-05-07 04:00 | Death | ABIOMED, INC. | 2025-06-20 04:00 | OZD | IMPELLA 5.5 | Positioning Failure | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PUMP BECAME MISPOSITIONED SO THE PUMP WAS REMOVED AND READJUSTED, BUT THERE WERE MANY BLOOD CLOTS IN THE PUMP. THE PUMP WAS REPLACED. THE PATIENT EXPIRED ON SUPPORT OF THE REPLACEMENT PUMP. THE PATIENT DIED WITHOUT ANY IMPROVEMENT IN CIRCULATORY DYNAMICS. NO PROBLEMS WITH THE OPERATION OF THE REPLACEMENT DEVICE. THE PATIENT DIED DUE TO MULTIPLE ORGAN FAILURE. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA 5.5 DEVICE USED AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." . IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22299195&pc=MIH | 1820334-2025-00746 |  | Death | COOK INC | 2025-06-20 04:00 | MIH | MIH; SYSTEM, ENDOVASCULAR GRAFT, AORTIC ANEURYSM TREATMENT |  | Ruptured Aneurysm | P020018 |  | 1 | Event Description: COOK WAS NOTIFIED THAT A PATIENT WHO HAD AN UNKNOWN COOK ZENITH AAA ENDOVASCULAR GRAFT IMPLANTED, EXPERIENCED A TYPE 1A ENDOLEAK THAT RESULTED IN ABDOMINAL AORTIC ANEURYSM (AAA) RUPTURE. THE INFORMATION WAS PROVIDED IN THE LITERATURE ARTICLE ¿SYSTEMATIC LONG-TERM FOLLOW UP AFTER ENDOVASCULAR ABDOMINAL AORTIC ANEURYSM REPAIR WITH THE ZENITH STENT GRAFT¿ THAT WAS PUBLISHED ON 27JUN2019 IN THE EUROPEAN JOURNAL OF VASCULAR AND ENDOVASCULAR SURGERY. ¿ENDOVASCULAR ANEURYSM REPAIR (EVAR) HAS BECOME THE METHOD OF CHOICE FOR THE TREATMENT OF ABDOMINAL AORTIC ANEURYSMS (AAA).1 SINCE THE INTRODUCTION OF EVAR OVER 20 YEARS AGO, MANY SUCCESSFUL SHORT-TERM RESULTS HAVE BEEN REPORTED, BUT LONG-TERM DATA ARE STILL LIMITED. FURTHERMORE, THE INCONSISTENCY IN REPORTING THE RESULTS OF INDIVIDUAL STUDIES MAKES THE INTERPRETATION OF THE FINDINGS TROUBLESOME. THIS IS TRUE ESPECIALLY WHEN A COHORT CONSISTS OF PATIENTS TREATED WITH MULTIPLE DEVICES WITH DIFFERENT TECHNICAL QUALITIES. CONSEQUENTLY, THIS MAY CARRY THE RISK OF UNDERESTIMATING THE LONG-TERM DURABILITY OF A PARTICULAR DEVICE. DESPITE ALL THE ADVANCES OF ENDOVASCULAR TECHNIQUES AND GRAFT MATERIALS, THE LATE FAILURE OF ENDOVASCULAR REPAIR DUE TO ENDOLEAKS PERSISTS. EVEN THOUGH AWARENESS OF THE NATURE OF DIFFERENT TYPES OF ENDOLEAK AND THEIR TREATMENT OPTIONS IS INCREASING, THE RE-INTERVENTION RATE REMAINS HIGHER THAN FOLLOWING OPEN REPAIR. ACCORDING TO THE AVAILABLE DATA, THE NUMBER OF NECESSARY RE-INTERVENTIONS AFTER THE PRIMARY PROCEDURE DOES NOT SEEM TO HAVE A NEGATIVE EFFECT ON THE AAA RELATED MORTALITY AFTER EVAR; THE TREATMENT MODALITY THEREFORE PROVIDES EFFECTIVE PROTECTION AGAINST AAA RUPTURE. AT THE SAME TIME, A SIMILAR POSITIVE IMPACT ON OVERALL MORTALITY HAS NOT BEEN DEMONSTRATED. THERE ARE SEVERAL REASONS FOR THIS. FIRSTLY, AS THE POPULATION IS AGEING AND THE PREVALENCE OF AAA HAS BEEN SHOWN TO BE HIGHEST AMONG THOSE OVER 80 YEARS OF AGE, ELDERLY PATIENTS WITH AAA AND MULTIPLE COMORBIDITIES ARE INCREASINGLY BEING TREATED. SECONDLY, IT HAS BEEN IMPLIED IN THE LITERATURE THAT EVAR MAY BE ASSOCIATED WITH INCREASED CANCER MORTALITY VS. OPEN REPAIR. THE JUSTIFICATION FOR TREATMENT MUST THEREFORE ALWAYS BE CAREFULLY CONSIDERED, DESPITE THE MINIMALLY INVASIVE NATURE OF THE MODALITY. THE PURPOSE OF THIS STUDY WAS TO REPORT THE 16-YEAR OVERALL SURVIVAL, AS WELL AS AAA RUPTURE, COMPLICATION, AND RE-INTERVENTION FREE SURVIVAL AFTER ELECTIVE EVAR WITH A ZENITH STENT GRAFT AT AN ACADEMIC INSTITUTION. MATERIAL AND METHODS: BETWEEN 2000 AND 2010, 282 PATIENTS, CONSTITUTING THE STUDY GROUP, UNDERWENT ELECTIVE EVAR WITH THE ZENITH STENT GRAFT (COOK, BLOOMINGTON, IN, USA) AT A SINGLE ACADEMIC INSTITUTION. PATIENTS WITH AN ISOLATED ILIAC ARTERY ANEURYSM OR PREVIOUS OPEN AORTIC SURGERY WERE EXCLUDED FROM THE STUDY. THE INDICATION FOR INITIAL TREATMENT WAS AN ANEURYSM WITH A DIAMETER OF GREATER THAN OR EQUAL TO 55 MM IN MEN AND GREATER THAN OR EQUAL TO 50 MM IN WOMEN. PATIENTS WITH AN INCREASE IN AAA DIAMETER OF GREATER THAN OR EQUAL TO 5 MM OVER A SIX MONTH PERIOD, AS WELL AS THOSE IN WHOM THE AAA BECAME SYMPTOMATIC, WERE ALSO TREATED. ALL PROCEDURES WERE PERFORMED BY A VASCULAR SURGEON TOGETHER WITH AN INTERVENTIONAL RADIOLOGIST IN A HYBRID SUITE. THE PATIENTS WERE MAINLY TREATED UNDER SPINAL ANESTHESIA, AND THE STENT GRAFTS WERE IMPLANTED THROUGH THE COMMON FEMORAL ARTERY. BIFURCATED GRAFTS WERE USED IN MOST OF THE EVAR PROCEDURES, BUT AN AORTO-UNI-ILIAC GRAFT WAS EMPLOYED IF ONE OF THE ILIAC ARTERIES WAS OCCLUDED OR SEVERELY STENOSED. AN ATTEMPT TO EMBOLISE A PATENT INFERIOR MESENTERIC ARTERY (IMA) WAS ALWAYS MADE PRIOR TO STENT GRAFT PLACEMENT. THE PATIENTS WERE SYSTEMATICALLY FOLLOWED AFTER THE PROCEDURE ACCORDING TO THE PROTOCOL IN TABLE 1. BASED ON THE AVAILABLE DATA AND THE AUTHORS¿ EXPERIENCE AT THE TIME, THE INITIAL SURVEILLANCE PROTOCOL WAS MODIFIED BY REPLACING THE ANNUAL COMPUTED TOMOGRAPHY ANGIOGRAPHY (CTA) SCAN WITH A COLOUR DUPLEX ULTRASONOGRAPHY (CDUS) IN 2005. THEREAFTER, ALL CDUS EXAMINATIONS WERE PERFORMED BY EXPERIENCED VASCULAR SURGEONS. THE CTA SCANS WERE CONTINUED FOR ALL PATIENTS, IF NOT CONTRAINDICATED, 24 MONTHS AFTER THE INITIAL PROCEDURE TO CONFIRM THE RELIABILITY OF THE CDUS EXAMINATIONS. FOR OBESE INDIVIDUALS AND FOR PATIENTS WITH A SUSPECTED COMPLICATION AT CDUS, AN ADDITIONAL CTA SCAN WAS PERFORMED. A PLAIN ABDOMINAL XRAY WAS TAKEN ANNUALLY UP UNTIL 2013. THE PATIENTS WERE FOLLOWED ANNUALLY UNTIL THE END OF MARCH 2017. AT THE END OF FOLLOW UP, THE STATUS OF EACH PATIENT WAS CONFIRMED (I.E., DEAD OR ALIVE), IN ADDITION TO OBTAINING INFORMATION ON POSSIBLE EVAR RELATED REINTERVENTIONS UP UNTIL THE END OF THE FOLLOW UP. FOR SOME INDIVIDUALS, LONG-TERM SURVEILLANCE WAS CARRIED OUT IN THE REFERRING HOSPITALS THEY HAD INITIALLY VISITED PRIOR TO THE ENDOVASCULAR AAA TREATMENT. THE REFERRING VASCULAR COLLEAGUES WERE CONTACTED ACCORDINGLY TO OBTAIN FOLLOW UP DATA. PATIENTS WERE CONTACTED IF A DELAY IN A FOLLOW UP EXAMINATION WAS NOTED. ALL CTAS WERE RE-ANALYSED FOR THE STUDY. THE PRIMARY OUTCOMES WERE OVERALL AND AAA RUPTURE FREE SURVIVAL, WHILE THE SECONDARY OUTCOMES INCLUDED GRAFT RELATED COMPLICATION FREE AND RE-INTERVENTION FREE SURVIVAL. DEATHS WERE ASCERTAINED BY RECORD LINKAGE BETWEEN THE STUDY POPULATION AND THE NATIONAL CAUSES OF DEATH REGISTER ON THE BASIS OF THE PERSONAL IDENTIFICATION CODE UNIQUE TO EVERY RESIDENT. SPSS 22.0 FOR WINDOWS WAS USED FOR STATISTICAL ANALYSIS (IBM, ARMONK, NY, USA). KAPLAN-MEIER SURVIVAL ANALYSIS WAS USED TO EXAMINE SURVIVAL. UNI- AND MULTIVARIABLE ANALYSES WERE USED TO ASSESS THE EFFECT OF AGE, SEX, AAA DIAMETER, NECK DIAMETER, NECK LENGTH, NECK ANGULATION, PATENT IMA, PRIMARY TYPE II ENDOLEAK, FOLLOWING THE INSTRUCTIONS FOR USE (IFU), AND OF BEING LOST TO FOLLOW UP ON ANY STENT GRAFT RELATED RE-INTERVENTION (DEFINED LATER IN TABLE 5) AND AAA RUPTURE. A PRIMARY ENDOLEAK WAS DEFINED AS AN ENDOLEAK OBSERVED IN THE PRIMARY PROCEDURE OR IN THE FIRST POST-OPERATIVE CTA ON THE SECOND OR THIRD DAY. STATISTICAL SIGNIFICANCE WAS INDICATED BY A P VALUE LESS THAN OR EQUAL TO .05. RESULTS: THE MEDIAN AGE OF THE COHORT WAS 76 YEARS (INTERQUARTILE RANGE [IQR] 11; RANGE 49-92 YEARS) AND MEAN ANEURYSM DIAMETER WAS 61 MM (RANGE 40-110 MM) AT THE TIME OF THE TREATMENT. THE IFU ISSUED FOR THE IMPLANTATION OF THE STENT GRAFT WERE FOLLOWED IN 90% OF THE PATIENTS. PATIENT CHARACTERISTICS ARE PRESENTED IN TABLE 2. THE MEDIAN FOLLOW UP WAS 76 MONTHS (IQR 69; RANGE 0-201 MONTHS). THE STATUS OF EACH PATIENT WAS CONFIRMED AT THE END OF THE FOLLOW UP (I.E., DEAD OR ALIVE). FOR EACH PATIENT, INFORMATION WAS AVAILABLE REGARDING THE LATEST ANNUAL FOLLOW UP EXAMINATION AND POSSIBLE EVAR RELATED RE-INTERVENTIONS AFTER THAT UP TO THE END OF FOLLOW UP. IN SOME CASES, THE FOLLOW UP INTERVAL WAS PROLONGED, OR THE SURVEILLANCE DISCONTINUED AFTER UNCOMPLICATED SURVEILLANCE, IF THE PATIENT WAS PHYSICALLY FRAIL OR IF LIFE EXPECTANCY SHORT. THESE PATIENTS WERE REPORTED AS LOST TO FOLLOW UP. SPINAL ANAESTHESIA WAS USED IN MOST CASES (N = 271;96%). A BIFURCATED STENT GRAFT WAS EMPLOYED IN 268 PATIENTS (95%) AND A UNI-ILIAC STENT IN 14 CASES (5%). THE IMA WAS PATENT IN 186 PATIENTS, AND WAS SUCCESSFULLY EMBOLIZED DURING THE INITIAL PROCEDURE PRIOR TO THE STENT GRAFT PLACEMENT IN 147 CASES (79%). THERE WERE NO INTRA-OPERATIVE DEATHS OR PRIMARY CONVERSIONS. THE THIRTY DAY MORTALITY WAS 1.8% (N = 5), AND THE CAUSES OF DEATH WERE BRAINSTEM INFARCTION (N =1), CARDIAC FAILURE (N = 3), AND CARDIAC ARREST DUE TO TOTAL STENT GRAFT THROMBOSIS (N = 1). THERE WAS NO SIGNIFICANT DIFFERENCE IN 30 DAY MORTALITY BETWEEN PATIENTS UNDER AND OVER 80 YEARS OF AGE (1.5% VS. 2.3%; P = .500). THE OVERALL SURVIVAL RATES WERE 93% (SE 0.02), 83% (SE 0.03), 61% (SE 0.08), 25% (SE 0.16), AND 9% (SE 0.19) AT 1, 2, 5, 10, AND 16 YEARS, RESPECTIVELY (FIG. 1). THE CORRESPONDING FIGURES FOR THOSE AGED GREATER THAN OR EQUAL TO 80 YEARS WERE 93% (SE 0.03), 80% (SE 0.04), 46% (SE 0.19), AND 4% (SE 0.42) AT 1, 2, 5, AND 10 YEARS, RESPECTIVELY. THERE WAS NO SIGNIFICANT DIFFERENCE IN SURVIVAL BETWEEN PATIENTS UNDER AND OVER 80 YEARS OF AGE AT TWO YEARS OF SURVEILLANCE (P = .529). THE MOST COMMON CAUSES OF DEATH WERE A CARDIAC EVENT (N =70; 32%) AND CANCER (N =48; 22%). THE AUTOPSY RATE WAS 26%. THERE WERE 10 (3.5%) AAA RUPTURES DURING FOLLOW UP (TABLE 3). ALL RUPTURES WERE CONFIRMED BY CTA (N = 8) OR AUTOPSY (N = 2). MEAN TIME TO RUPTURE WAS 81 MONTHS (RANGE 22 - 173 MONTHS). CONSEQUENTLY, THE CUMULATIVE AAA RUPTURE FREE SURVIVAL WAS 100%, 98% (SE 0.01), 96% (SE 0.02), AND 79% (SE 0.12) AT 1, 5, 10, AND 16 YEARS, RESPECTIVELY. EIGHT OF THE 10 PATIENTS WITH A RUPTURED AAA (RAAA) REACHED THE HOSPITAL, AND TWO OF THEM UNDERWENT AN EMERGENCY OPEN REPAIR. ONE OF THESE PATIENTS DIED IN THE INTENSIVE CARE UNIT TWO DAYS AFTER THE PROCEDURE, AND THE OTHER WAS STILL ALIVE AT THE END OF THE FOLLOW UP. IN THREE RAAA CASES, AN ADDITIONAL PROXIMAL CUFF WAS PLACED DURING AN EMERGENCY PROCEDURE, BUT ADEQUATE PROXIMAL SEALING WAS NOT ACHIEVED, AND THE PATIENTS DIED POST-OPERATIVELY. THEY WERE NOT CONSIDERED CANDIDATES FOR OPEN CONVERSION. IN THE REMAINING THREE CASES, ENDOVASCULAR REPAIR WAS NOT CONSIDERED TECHNICALLY POSSIBLE AND OPEN REPAIR WAS REFUSED OWING TO THE PATIENTS¿ HIGH AGE AND MULTIPLE COMORBIDITIES. DURING SURVEILLANCE, ONE PATIENT WITH A RAAA WAS FOUND TO HAVE A TYPE IB ENDOLEAK, BUT, AT THE AGE OF 90 YEARS, HE WAS SEVERELY DEMENTED AND NO FURTHER PROCEDURES WERE PERFORMED. IN ONE CASE, THE RUPTURE WAS UNEXPECTEDLY FOUND AT AUTOPSY 14 YEARS AFTER THE PRIMARY TREATMENT. THIS FEMALE PATIENT HAD REFUSED FURTHER FOLLOW UP AFTER 9 YEARS OF UNCOMPLICATED SURVEILLANCE, WITH A SHRUNKEN AAA DIAMETER OF 30 MM. THE NINE OTHER PATIENTS WERE UNDER SURVEILLANCE AT THE TIME OF THE RAAA. SEVEN OF THE 10 AAA RUPTURES WERE DUE TO TYPE I ENDOLEAK (IA: N = 5; IB: N= 2). SIX OF THESE PATIENTS HAD A PERSISTENT PRIMARY TYPE II ENDOLEAK IN EARLY SURVEILLANCE PRIOR TO THE DEVELOPMENT OF THE TYPE I ENDOLEAK. IN THREE OF THESE CASES, AS THE ANEURYSM SAC SHOWED ENLARGEMENT IN SURVEILLANCE, REPEATED UNSUCCESSFUL ATTEMPTS TO EMBOLISE THE TYPE II ENDOLEAK WERE MADE. IN THREE CASES, HOWEVER, THE TYPE II ENDOLEAK SEEMED STABLE, WITH NO DETECTABLE CHANGE IN ANEURYSM SAC SIZE DURING EARLY SURVEILLANCE, BUT AAA ENLARGEMENT WAS NOTED APPROXIMATELY THREE YEARS AFTER THE PRIMARY TREATMENT. BY THAT TIME, THESE PATIENTS HAD ALSO DEVELOPED A TYPE I ENDOLEAK BEFORE AAA RUPTURE. ONE RUPTURE PATIENT REFUSED FURTHER PROCEDURES FOR THE TYPE IA ENDOLEAK AND DIED THREE YEARS AFTER THE INITIAL PROCEDURE. IN THE CASE OF THREE OF THE 10 RUPTURES, NO SIGNIFICANT ENDOLEAK HAD BEEN NOTED DURING SURVEILLANCE. ONE PATIENT WITH A PRIMARY TYPE II ENDOLEAK SHOWED CLEAR SHRINKAGE OF THE ANEURYSM SAC DURING THE EARLY SURVEILLANCE. HOWEVER, AT 22 MONTHS, THE ANEURYSM RUPTURED. ANOTHER PATIENT WAS FOLLOWED FOR ENDOTENSION, WITH NO SIGNS OF AN ENDOLEAK, AND HE DIED OF A RAAA FIVE YEARS AFTER THE PRIMARY PROCEDURE. ONE RUPTURE, AS MENTIONED EARLIER, WAS FOUND AT AUTOPSY WITH NO PRIOR HISTORY OF AN ENDOLEAK. THE UNIVARIABLE ANALYSIS SHOWED A SIGNIFICANT ASSOCIATION BETWEEN PRIMARY TYPE II ENDOLEAK AND AAA RUPTURE (ODDS RATIO [OR] 4.09, 95% CI 1.12 - 14.9; P = .033). AN EVEN MORE SIGNIFICANT ASSOCIATION WAS FOUND BETWEEN PRE-OPERATIVE ANEURYSM SIZE AND LATE AAA RUPTURE (OR 1.06, 95% CI 1.02 - 1.11; P = .007). THESE RESULTS ALSO REMAINED CONSTANT IN THE MULTIVARIABLE ANALYSIS, WITH AN OR OF 4.65 (95% CI 1.14 - 18.9; P = .032) FOR TYPE II ENDOLEAK AND 1.07 (95% CI 1.02 - 1.13; P= .009) FOR PREOPERATIVE ANEURYSM SIZE. CONSEQUENTLY, AT THE END OF THE 16 YEAR FOLLOW UP, THE MEAN ANNUAL ANEURYSM RUPTURE RATE WAS 0.49% (TABLE 4). A TOTAL OF 168 GRAFT RELATED COMPLICATIONS WERE SEEN IN A TOTAL OF 119 PATIENTS (42%). ALL PRIMARY TYPE II ENDOLEAKS WERE TAKEN INTO ACCOUNT, EVEN IF THEY WERE DETECTED ONLY IN THE PERI-OPERATIVE ANGIOGRAM OR ON THE FIRST CTA. ALL COMPLICATIONS ARE LISTED IN TABLE 3. THE FREEDOM FROM ANY STENT GRAFT RELATED COMPLICATION WAS 68% (SE 0.03), 63% (SE 0.05), 58% (SE 0.09), 54% (SE 0.17), AND 52% (SE 0.21) AT 1, 2, 5, 10, AND 16 YEARS, RESPECTIVELY (FIG. 2). COMPLICATIONS SEEMED TO ACCUMULATE DURING THE EARLY YEARS OF SURVEILLANCE BUT WERE ENCOUNTERED UP TO 14 YEARS AFTER THE PRIMARY PROCEDURE. A TYPE II ENDOLEAK WAS THE MOST COMMON COMPLICATION (N = 79; 28%), WITH ALMOST HALF (48%) OF THEM SEALING SPONTANEOUSLY AND THE VAST MAJORITY (89%) OCCURRING DURING THE FIRST TWO YEARS OF FOLLOW UP. OF THOSE WITH A PATENT IMA (39 PATIENTS) AFTER THE INITIAL PROCEDURE, 18 (46%) DEVELOPED EITHER A TYPE II ENDOLEAK (N =14) OR ENDOTENSION (N = 4). MOST OF THESE PATIENTS (88%) WITH EITHER A TYPE II ENDOLEAK OR ENDOTENSION ALSO HAD A PATENT IMA ON SURVEILLANCE, AND ONLY TWO OF THEM REQUIRED LATER EMBOLISATION. IN ALL, ONLY 26 (33%) OF ALL TYPE II ENDOLEAKS REQUIRED ENDOVASCULAR TREATMENT DUE TO ANEURYSM SAC ENLARGEMENT OVER THE STUDY PERIOD. THE PRIMARY EMBOLIZATION STRATEGY WAS COIL EMBOLISATION; IF THAT PROVED UNSUCCESSFUL, A SECONDARY ATTEMPT WAS MADE WITH GLUE. A TOTAL OF 13 PATIENTS WERE DIAGNOSED WITH ENDOTENSION (GREATER THAN OR EQUAL TO 5 MM) WITHOUT A DETECTABLE ENDOLEAK DURING THE SURVEILLANCE. FOR ONE PATIENT, THERE WAS AN ATTEMPT TO TREAT THE ENDOTENSION BY RELINING THE OLD GRAFT, WHILE THE REST WERE FOLLOWED CAREFULLY. TWO OF THESE PATIENTS LATER DEVELOPED A TYPE I ENDOLEAK, AND THEY WERE TREATED AT THAT POINT WITH AN ADDITIONAL CUFF AND LIMB GRAFT. NONE OF THE PATIENTS WITH ENDOTENSION SUFFERED A RAAA. RE-INTERVENTION FOR A GRAFT RELATED COMPLICATION WAS REQUIRED FOR 61 PATIENTS (22%). ALL ADDITIONAL PROCEDURES ARE LISTED IN TABLE 5. THE RATE OF FREEDOM FROM GRAFT RELATED REINTERVENTIONS WAS 95% (SE 0.01), 91% (SE 0.03), 80% (SE 0.08), 73% (SE 0.11), AND 70% (SE 0.16) AT 1, 2, 5, 10, AND 16 YEARS, RESPECTIVELY (FIG. 3). ONLY FIVE (1.8%) LATE CONVERSIONS TO OPEN REPAIR WERE REQUIRED, TWO AS AN EMERGENCY REPAIR OF A RUPTURE AS MENTIONED EARLIER. TWO PATIENTS HAD A PERSISTENT TYPE I ENDOLEAK DESPITE THE PLACEMENT OF AN EXTRA PROXIMAL CUFF, AND CONVERSIONS WERE PERFORMED AT 12 AND 27 MONTHS, RESPECTIVELY. ONE PATIENT HAD A STENT FRACTURE AND TYPE III ENDOLEAK 149 MONTHS AFTER THE PRIMARY TREATMENT. THE OTHER CASE OF TYPE III ENDOLEAK DEVELOPED AS A RESULT OF LIMB GRAFT MIGRATION AND WAS SUCCESSFULLY TREATED WITH AN ADDITIONAL LIMB GRAFT AT 47 MONTHS. ALL ELECTIVE CONVERSIONS WERE UNCOMPLICATED. A PRIMARY TYPE II ENDOLEAK WAS FOUND TO BE THE ONLY INDEPENDENT SIGNIFICANT FACTOR TO INCREASE THE RISK OF ANY LATE RE-INTERVENTION (OR 6.79, 95% CI 3.61 - 12.79; P < .001). DURING FOLLOW UP, ANEURYSM SAC SHRINKAGE OF _5 MM WAS SEEN IN 61% OF THE CASES, WITH AT MEAN DECREASE OF 18 MM (RANGE 5 - 41 MM) IN MAXIMUM DIAMETER.¿ REFERENCE: VÄÄRÄMÄKI, S., SALENIUS, J.-P., PIMENOFF, G., UURTO, I., & SUOMINEN, V. (2019). SYSTEMATIC LONG-TERM FOLLOW UP AFTER ENDOVASCULAR ABDOMINAL AORTIC ANEURYSM REPAIR WITH THE ZENITH STENT GRAFT. EUROPEAN JOURNAL OF VASCULAR AND ENDOVASCULAR SURGERY, 58(2), 182¿188. HTTPS://DOI.ORG/10.1016/J.EJVS.2019.03.006. THE FOCUS OF THIS REPORT IS THE PATIENT WHO HAD AN UNKNOWN COOK ZENITH AAA ENDOVASCULAR GRAFT IMPLANTED AND EXPERIENCED A TYPE 1 ENDOLEAK RESULTING IN AAA RUPTURE, OPEN CONVERSION, AND DEATH. Manufacturer Narrative: B2: DATE OF DEATH: DATE OF DEATH IS UNKNOWN. THE DATE PROVIDED IS THE PUBLICATION DATE OF THE JOURNAL ARTICLE. B3: DATE OF EVENT: 01/01/2000 TO 03/31/2017. D6A: IMPLANT DATE: 01/01/2000 TO 12/31/2010. E1: PHONE: (B)(6). H3: DEVICE EVALUATED BY MFG? DEVICE NOT RETURNED TO MANUFACTURER. THIS REPORT INCLUDES INFORMATION KNOWN AT THIS TIME. A FOLLOW-UP REPORT WILL BE SUBMITTED SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803 AND IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT CONTAINED HEREIN IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED OR THAT A DEATH OR SERIOUS INJURY OCCURRED; NOR IS IT ADMISSION THAT ANY COOK DEVICE CAUSED, CONTRIBUTED TO, OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22299517&pc=FKX | 0008030665-2025-01413 | 2025-05-12 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-06-20 04:00 | FKX | LIBERTY CYCLER SET, DUAL PATIENT CONNECT | Adverse Event Without Identified Device or Use Problem | Abdominal Pain; Dyspnea; Sepsis; Peritonitis; Swelling/ Edema | K173718 |  | 1 | Event Description: A REGISTERED NURSE (RN) REPORTED TO FRESENIUS THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY ON (B)(6) 2025. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT WAS SEEN IN THE PD CLINIC ON (B)(6) 2025 DUE TO ABDOMINAL PAIN AND CLOUDY PD EFFLUENT. PD CULTURES WERE OBTAINED. THE PATIENT WAS DIAGNOSED WITH PERITONITIS. THE PATIENT WAS INITIATED ON EMPIRIC ANTIBIOTIC THERAPY WITH INTRAPERITONEAL (IP) VANCOMYCIN AND IP CEFEPIME (DOSE, FREQUENCY, AND DURATION UNKNOWN). THE CULTURES WERE POSITIVE FOR PSEUDOMONAS (NO SPECIES PROVIDED). THE VANCOMYCIN WAS DISCONTINUED AND THE CEFEPIME WAS TO BE CONTINUED FOR 3 WEEKS. THE PATIENT¿S SYMPTOMS WERE NOT IMPROVING. ON (B)(6) 2025 THE PATIENT WAS ADMITTED TO THE HOSPITAL DUE TO SHORTNESS OF BREATH AND SWELLING (LOCATION NOT PROVIDED). THE PATIENT¿S PD CATHETER WAS REMOVED. THE PATIENT TRANSITIONED TO HEMODIALYSIS (HD). THE PATIENT CONTINUED TO WORSEN AND NOT MAKE ANY IMPROVEMENTS. THE PATIENT PASSED AWAY ON (B)(6) 2025 IN THE HOSPITAL. THE HOSPITAL DID NOT SEND OVER THE DEATH REPORT, BUT THE DISCHARGE RECORDS STATED THE CAUSE OF DEATH WAS SEPTICEMIA AND RESPIRATORY DISTRESS SYNDROME. NO ADDITIONAL INFORMATION WAS AVAILABLE AS THE PATIENT¿S RECORDS HAD BEEN CLOSED OUT OF THE CLINIC¿S SYSTEM. THE CAUSE OF THE PERITONITIS WAS NOT ESTABLISHED. Manufacturer Narrative: CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PD THERAPY AND UTILIZATION OF THE CYCLER SET AND THE PATIENT EVENT OF PSEUDOMONAS PERITONITIS WITH HOSPITALIZATION AND SUBSEQUENT DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY, OR CYCLER SET DEFECT REPORTED FOR THE EVENT. THE PATIENT WAS INITIALLY DIAGNOSED WITH PSEUDOMONAS PERITONITIS. PERITONITIS CAUSED BY PSEUDOMONAS SPECIES IS A PARTICULARLY COMMON AND SERIOUS COMPLICATION AND IS OFTEN ASSOCIATED WITH A POOR RESPONSE TO ANTIBIOTICS AS WELL AS A HIGH RATE OF CATHETER REMOVAL. ALTHOUGH THE CAUSE OF THE PERITONITIS WAS NOT CONFIRMED, PSEUDOMONAS IS A SPECIES THAT NORMALLY INHABITS SOIL, WATER, AND VEGETATION AND CAN BE ISOLATED FROM THE SKIN, THROAT, AND STOOL OF HEALTHY PERSONS. THE PATIENT¿S CONDITION WORSENED OVER TIME DESPITE BEING ON ANTIBIOTIC THERAPY. THE PD CATHETER WAS EVENTUALLY REMOVED APPROXIMATELY TWO WEEKS AFTER THE PERITONITIS DIAGNOSIS. ADDITIONALLY, THE PATIENT¿S ANTIBIOTICS WERE INITIALLY CHANGED TO ONE ANTIBIOTIC WITH CEFEPIME AFTER THE CULTURE RESULTS OF PSEUDOMONAS. THE INTERNATIONAL SOCIETY OF PERITONEAL DIALYSIS (ISPD) GUIDELINES FOR MANAGEMENT OF PD-RELATED PERITONITIS FOR PSEUDOMONAS RECOMMENDS THE USE OF DUAL ANTIBIOTIC THERAPY WITH DIFFERENT MECHANISMS OF ACTION AND TREATMENT FOR 3 WEEKS. WHEN THERE IS NO CLINICAL RESPONSE AFTER 4 DAYS OF ANTIBIOTIC TREATMENT, EARLY CATHETER REMOVAL SHOULD BE PREFERRED. THE PATIENT¿S WORSENING SYMPTOMS AND EVENTUAL PD CATHETER REMOVAL SUGGESTS THAT THE BACTERIA POSSIBLY DEVELOPED A BIOFILM WITHIN THE PD CATHETER AND THAT THE BACTERIA INVADED TISSUES SPREADING TO CAUSE THE EVENTUAL SEPTICEMIA WITH RESPIRATORY DISTRESS SYNDROME AND SUBSEQUENT DEATH. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22299757&pc=DSQ | 2916596-2025-04081 | 2025-06-12 04:00 | Death | THORATEC CORPORATION | 2025-06-20 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT |  | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO THE PATIENT NEVER RECOVERING AFTER IMPLANTATION. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED; THE PATIENT EXPIRED DUE TO ANOTHER DISEASE. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WOULD NOT RETURN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR(B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22300361&pc=DSQ | 2916596-2025-03902 | 2025-05-31 04:00 | Death | THORATEC CORPORATION | 2025-06-20 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THE PATIENT PASSED AWAY. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE DEVICE WILL NOT BE RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO HOSPITAL PRIVACY POLICY, THE ACCOUNT WILL NOT PROVIDE ADDITIONAL INFORMATION REGARDING THE EVENT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22300422&pc=OZD | 1220648-2025-29627 | 2025-06-05 04:00 | Death | ABIOMED, INC. | 2025-06-20 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Thrombocytopenia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED A DECREASE IN PLATELET COUNT. THE PATIENT HAD JUST BEGUN TO RECEIVE HEPARIN. THE HEPARIN WAS REPLACED WITH ARGATROBAN TO RESOLVE THE HEPARIN INDUCED THROMBOCYTOPENIA. THE PATIENT DECOMPENSATED SHORTLY AFTER AND EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22300986&pc=DSQ | 2916596-2025-03789 | 2023-09-07 04:00 | Death | THORATEC CORPORATION | 2025-06-20 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER: (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22301482&pc=GAM | 2210968-2025-07194 | 2024-02-23 05:00 | Death | ETHICON INC. | 2025-06-20 04:00 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Adverse Event Without Identified Device or Use Problem | Wound Dehiscence; Cardiac Arrest; Unspecified Infection; Respiratory Failure | K022269 |  | 1 | Event Description: TIMING OF COMPLICATIONS IN OPEN AIRWAY RECONSTRUCTION. THE AIM OF THIS STUDY IS TO INVESTIGATE THE INCIDENCE AND TIMING OF COMPLICATIONS THAT OCCUR AFTER TRACHEAL RESECTION (TR)/CRICOTRACHEAL RESECTION (CTR), TO BETTER CHARACTERIZE WHEN CLINICAL SUSPICION SHOULD REMAIN HIGHEST AND GUIDE POST-OPERATIVE MANAGEMENT. BETWEEN 2015 TO 2022, SIXTY-NINE CASES OF TR/CTR WERE REVIEWED. ALL PATIENTS WHO UNDERWENT TR/CTR DURING THIS TIME PERIOD WERE ANALYZED FOR TIMING OF COMPLICATIONS IN DAYS. DEMOGRAPHIC DATA AND CLINICAL FACTORS WERE COLLECTED FOR A SECONDARY AIM OF ASSESSING RISK FACTORS. A SUBSET OF NINE PATIENTS WERE EXCLUDED FROM RISK FACTOR ANALYSIS GIVEN INCOMPLETE MEDICAL RECORDS. CASES WERE PERFORMED BY FOUR SURGEONS, ALL USING THE SAME ANASTOMOSIS TECHNIQUE WITH FULL-THICKNESS INTERRUPTED VICRYL SUTURES. REPORTED COMPLICATIONS ARE THE FOLLOWING: VICRYL SUTURES (ETHICON) POST-SURGICAL NECK HEMATOMA OCCURRED BETWEEN 1 AND 5 DAYS AFTER SURGERY (N=4) TREATMENT: TWO WERE MANAGED WITH BEDSIDE INCISION AND DRAINAGE AND TWO UNDERWENT EVACUATION IN THE OPERATING ROOM. ANASTOMOTIC DEHISCENCE (WITH DEFINITION RANGING FROM SUBCUTANEOUS EMPHYSEMA TO BRONCHOSCOPICALLY-VISUALIZED AIR LEAK) OCCURRED ON AVERAGE OF 7.5 DAYS AFTER SURGERY (N=4) TREATMENT: TWO PATIENTS WERE MANAGED CONSERVATIVELY WITH ANTIBIOTICS AND LOCAL WOUND PACKING, AND TWO PATIENTS HAD SIGNIFICANT DEHISCENCE REQUIRING REVISION TRACHEOSTOMY. UNKNOWN EVENT ON FIRST OCCURRING 4.6 DAYS AFTER SURGERY ON AVERAGE (N=8) TREATMENT: NOT SPECIFIED BUT THERE WERE PATIENT HAD REQUIRED RE-INTUBATION AT LEAST ONCE WITH UNKNOWN CAUSE EARLY RESTENOSIS (N=1), ANASTOMOTIC DEHISCENCE (N=2), PROLONGED REINTUBATION IN THE SETTING OF ARDS (N=1) ON AVERAGE OCCURRING 10.3 DAYS POST-OP TREATMENT: REQUIRED REVISION TRACHEOSTOMY RESTENOSIS AT THE ANASTOMOTIC LINE DEVELOPED THIS DURING THEIR POST-OPERATIVE HOSPITALIZATION AT DAY 8 (N=1) TREATMENT: MANAGED WITH REINTUBATION FOLLOWED BY TRACHEOSTOMY IN THE ACUTE SETTING AND HAS GONE ON TO DEVELOP NEAR-COMPLETE SUPRASTOMAL/ANASTOMOTIC STENOSIS. RESTENOSIS AT THE ANASTOMOTIC LINE DEVELOPED BETWEEN 25 AND 92 DAYS POST-OP (N=6) TREATMENT: ALL SEEN DURING ROUTINE FOLLOW-UP APPOINTMENTS ON WHICH ONE PATIENT HAD ONLY MILD STENOSIS AND WAS OBSERVED WITHOUT INTERVENTION. THE OTHER FIVE ALL REQUIRED AT LEAST TWO FURTHER OPERATIVE ENDOSCOPIC DILATIONS AND INTRALESIONAL STEROID INJECTIONS WITH CONTROL OF THEIR SYMPTOMS. VICRYL SUTURES (ETHICON): DEATH (N=3) DUE TO SEVERE ARDS (N=2) AND ADVANCED AIRWAY INVOLVEMENT OF THYROID CARCINOMA DEVELOPED ANASTOMOTIC DEHISCENCE LIKELY DUE TO LOCAL INFECTION, AND ONE SUFFERED AIRWAY LOSS AND CARDIAC ARREST IN THE OPERATING ROOM AND WITHDRAWAL OF CARE 7 DAYS POST OPERATIVELY (N=1). TREATMENT: NONE. IN CONCLUSION, IN THIS SINGLE-CENTER STUDY, THE MAJORITY OF POST OPERATIVE COMPLICATIONS AFTER TR OR CTR, INCLUDING HEMATOMA AND ANASTOMOTIC DEHISCENCE, OCCURRED WITHIN THE FIRST 8 DAYS. THUS, PATIENTS WITH SIGNIFICANT MEDICAL COMORBIDITIES OR LONGER-SEGMENT TRACHEAL STENOSIS MAY BENEFIT FROM A PERIOD OF LONGER OBSERVATION IN THE HOSPITAL. MOST CASES OF RESTENOSIS WERE NOTED APPROXIMATELY 1¿ 3 MONTHS AFTER SURGERY AT ROUTINE FOLLOW-UP VISITS. THIS MAY INFORM CLINICAL DECISION-MAKING REGARDING PATIENT MONITORING IN THE ACUTE SETTING AS WELL AS POST-OPERATIVE SURVEILLANCE AFTER OPEN AIRWAY SURGERY. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DOES THE SURGEON BELIEVE THAT ANY OF THE ETHICON PRODUCTS INVOLVED CAUSED AND/OR CONTRIBUTED TO THE POST-OPERATIVE COMPLICATIONS/DEATH DESCRIBED IN THE ARTICLE? WHICH SPECIFIC ETHICON PRODUCTS HAVE BEEN USED DURING THE PROCEDURES (PRODUCT CODE, LOT NUMBER)? DOES THE SURGEON BELIEVE THERE WAS ANY DEFICIENCY WITH ANY OF THE ETHICON PRODUCTS USED IN THIS PROCEDURE? IF SO, PLEASE PROVIDE DETAILS. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE A COMPLAINT REFERENCE NUMBER. PATIENT DEMOGRAPHICS? HAS AN AUTOPSY BEEN PERFORMED FOR DEATH CASE? IF SO, CAN THE RESULTS BE SHARED? WOULD THE JOURNAL AUTHOR / PHYSICIAN LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. H6 COMPONENT CODE: G07002 ¿ DEVICE NOT RETURNED. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. CITATION: LARYNGOSCOPE. 2024 AUG;134(8):3527-3531. DOI: 10.1002/LARY.31362. EPUB 2024 FEB 23. PMID: 38390693. HTTPS://DOI.ORG/10.1002/LARY.31362. Manufacturer Narrative: PRODUCT COMPLAINT#: (B)(4). CORRECTED INFORMATION: B1, B2, H1: EVENT INFORMATION REVIEWED. THIS EVENT NO LONGER MEETS REPORTABILITY CRITERIA. THEREFORE, THIS MEDWATCH REPORT IS BEING VOIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22301484&pc=DSQ | 2916596-2025-03887 | 2025-06-04 04:00 | Death | THORATEC CORPORATION | 2025-06-20 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Respiratory Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY SECONDARY TO RESPIRATORY ARREST. IT WAS A SUDDEN, ACUTE SITUATION OF LOW FLOW ALARMS AND RESPIRATORY ARREST. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. THE DEVICE WAS NOT RETURNED FOR EVALUATION.Event Description: IT WAS ADDITIONALLY REPORTED THAT THE PATIENT WAS BEING CARED FOR AT HOME AT THE TIME OF THE EVENT. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS INVESTIGATION. ADDITIONALLY, THE REPORTED LOW FLOW ALARMS COULD NOT BE CONFIRMED AS LOG FILES WERE NOT AVAILABLE FOR REVIEW DUE TO THE EVENTS OCCURRING WHEN THE PATIENT WAS BEING CARED FOR AT HOME. THE ACCOUNT COMMUNICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THIS SECTION ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿HEARTMATE TOUCH COMMUNICATION SYSTEM¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LPM, A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. THIS SECTION ALSO OUTLINES SITUATIONS THAT CAN RESULT IN A LOW FLOW ALARM AND FURTHER EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING," AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿ OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22301503&pc=OZD | 1220648-2025-29631 | 2025-06-04 04:00 | Death | ABIOMED, INC. | 2025-06-20 04:00 | OZD | IMPELLA CP | Device Sensing Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THERE WERE FALSE IMPELLA IN AORTA ALARMS. THE POSITION WAS CONFIRMED ON ECHO. PLACEMENT SIGNAL MONITORING ALARM DISABLED. ADDITIONALLY, THE PATIENT HAD A LARGE INTERCRANIAL HEMATOMA CAUSING A MIDLINE LINE SHIFT AND ACTIVE HEMORRHAGING. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST SYSTEM INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22301876&pc=OZD | 1220648-2025-29636 | 2025-06-05 04:00 | Death | ABIOMED, INC. | 2025-06-20 04:00 | OZD | IMPELLA CP | Positioning Failure | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP SUPPORT AND DURING SUPPORT, THE PATIENT HAD A LEFT VENTRICLE (LV) THROMBUS. ADDITIONALLY, THE PUMP MEASURED AT 2.7CM AND WAS ROTATED JUST UNDER THE MITRAL VALVE. THE DOCTOR DID NOT WANT TO REPOSITION DUE TO THE LV THROMBUS. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION, SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH. IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22302136&pc=PNJ | 2017865-2025-87877 | 2025-05-29 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-20 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Failure to Capture; High impedance | Cardiac Tamponade; Discomfort; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR A SCHEDULED PROCEDURE. DURING THE PROCEDURE, FIXATION WAS ATTEMPTED MULTIPLE TIMES DUE TO LOSS OF CAPTURE. DURING FIXATION, THERE WAS A RAPID INCREASE IN IMPEDANCE AND ST SEGMENT BECAME NEGATIVE INDICATING POTENTIAL RV PERFORATION. SHORTLY AFTER THE PROCEDURE, THE PATIENT WAS EXPERIENCING CHEST PAIN AND THEN STARTED TO HAVE A HEMODYNAMIC COLLAPSE. A PERICARDIOCENTESIS WAS PERFORMED, THE PATIENT WAS UNABLE TO BE RESUSCITATED, AND THE PATIENT PASSED SHORTLY AFTER. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22302137&pc=BZD | 2518422-2025-106766 | 2025-06-06 04:00 | Death | RESPIRONICS, INC. | 2025-06-20 04:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: A RELATIVE OF THE PATIENT REPORTED THE PATIENT EXPIRED. THE REPORTER STATED THE PATIENT WAS NOT USING THE DEVICE AT THE TIME THE PATIENT PASSED AWAY. THE REPORTER STATED THE PATIENT WAS USING ANOTHER UNSPECIFIED DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22302139&pc=MVK | 3008642652-2025-06042 | 2025-05-28 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-20 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 04:17:58 ON (B)(6) 2025. AT 05:56:11, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. AT 05:57:36, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL, CPR/MOTION ARTIFACT, AND ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF/INTERMITTENT CARDIAC ACTIVITY. AT 05:58:56, AN ARRHYTHMIA WAS DETECTED. ECG WAS OBSCURED DUE TO CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. AT 06:00:10, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL, CPR/MOTION ARTIFACT, AND ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. AT 06:04:09, AN ARRHYTHMIA WAS DETECTED. ECG WAS OBSCURED DUE TO CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. AT 06:04:38, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL, CPR/MOTION ARTIFACT, AND ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. THE ELECTRODE BELT WAS DISCONNECTED AT 06:10:23 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22302141&pc=MVK | 3008642652-2025-06043 | 2025-05-28 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-20 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 07:08:42ON (B)(6) 2025. AT 23:47:22, AN ARRHYTHMIA WAS DETECTED. ECG WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 23:47:57, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS SINUS TACHYCARDIA @ 100 BPM WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 23:48:24, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT/INTERMITTENT CARDIAC ACTIVITY. POST SHOCK RHYTHM WAS SEVERE BRADYCARDIA @ 10 BPM WITH CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 23:48:25 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22302574&pc=BZD | 2518422-2025-106767 | 2025-06-09 04:00 | Death | RESPIRONICS, INC. | 2025-06-20 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Degraded | Insufficient Information | K113068 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THE PATIENT'S RELATIVE IS RETURNING THEIR DEVICE. THE CAUSE OF DEATH WAS NOT PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22302615&pc=MIH | 1820334-2025-00747 |  | Death | COOK INC | 2025-06-20 04:00 | MIH | ZENITH AAA ENDOVASCULAR GRAFT ILIAC LEG |  | Ruptured Aneurysm | P020018 |  | 1 | Event Description: COOK WAS NOTIFIED THAT A PATIENT WHO HAD ZENITH AAA ENDOVASCULAR GRAFT ILIAC LEG IMPLANTED AND EXPERIENCED A TYPE 1B ENDOLEAK RESULTING IN ABDOMINAL AORTIC ANEURYSM (AAA) RUPTURE. THE INFORMATION WAS PROVIDED IN THE LITERATURE ARTICLE ¿SYSTEMATIC LONG-TERM FOLLOW UP AFTER ENDOVASCULAR ABDOMINAL AORTIC ANEURYSM REPAIR WITH THE ZENITH STENT GRAFT¿ THAT WAS PUBLISHED ON 27JUN2019 IN THE EUROPEAN JOURNAL OF VASCULAR AND ENDOVASCULAR SURGERY. ¿ENDOVASCULAR ANEURYSM REPAIR (EVAR) HAS BECOME THE METHOD OF CHOICE FOR THE TREATMENT OF ABDOMINAL AORTIC ANEURYSMS (AAA).1 SINCE THE INTRODUCTION OF EVAR OVER 20 YEARS AGO, MANY SUCCESSFUL SHORT-TERM RESULTS HAVE BEEN REPORTED, BUT LONG-TERM DATA ARE STILL LIMITED. FURTHERMORE, THE INCONSISTENCY IN REPORTING THE RESULTS OF INDIVIDUAL STUDIES MAKES THE INTERPRETATION OF THE FINDINGS TROUBLESOME. THIS IS TRUE ESPECIALLY WHEN A COHORT CONSISTS OF PATIENTS TREATED WITH MULTIPLE DEVICES WITH DIFFERENT TECHNICAL QUALITIES. CONSEQUENTLY, THIS MAY CARRY THE RISK OF UNDERESTIMATING THE LONG-TERM DURABILITY OF A PARTICULAR DEVICE. DESPITE ALL THE ADVANCES OF ENDOVASCULAR TECHNIQUES AND GRAFT MATERIALS, THE LATE FAILURE OF ENDOVASCULAR REPAIR DUE TO ENDOLEAKS PERSISTS. EVEN THOUGH AWARENESS OF THE NATURE OF DIFFERENT TYPES OF ENDOLEAK AND THEIR TREATMENT OPTIONS IS INCREASING, THE RE-INTERVENTION RATE REMAINS HIGHER THAN FOLLOWING OPEN REPAIR. ACCORDING TO THE AVAILABLE DATA, THE NUMBER OF NECESSARY RE-INTERVENTIONS AFTER THE PRIMARY PROCEDURE DOES NOT SEEM TO HAVE A NEGATIVE EFFECT ON THE AAA RELATED MORTALITY AFTER EVAR; THE TREATMENT MODALITY THEREFORE PROVIDES EFFECTIVE PROTECTION AGAINST AAA RUPTURE. AT THE SAME TIME, A SIMILAR POSITIVE IMPACT ON OVERALL MORTALITY HAS NOT BEEN DEMONSTRATED. THERE ARE SEVERAL REASONS FOR THIS. FIRSTLY, AS THE POPULATION IS AGEING AND THE PREVALENCE OF AAA HAS BEEN SHOWN TO BE HIGHEST AMONG THOSE OVER 80 YEARS OF AGE, ELDERLY PATIENTS WITH AAA AND MULTIPLE COMORBIDITIES ARE INCREASINGLY BEING TREATED. SECONDLY, IT HAS BEEN IMPLIED IN THE LITERATURE THAT EVAR MAY BE ASSOCIATED WITH INCREASED CANCER MORTALITY VS. OPEN REPAIR. THE JUSTIFICATION FOR TREATMENT MUST THEREFORE ALWAYS BE CAREFULLY CONSIDERED, DESPITE THE MINIMALLY INVASIVE NATURE OF THE MODALITY. THE PURPOSE OF THIS STUDY WAS TO REPORT THE 16-YEAR OVERALL SURVIVAL, AS WELL AS AAA RUPTURE, COMPLICATION, AND RE-INTERVENTION FREE SURVIVAL AFTER ELECTIVE EVAR WITH A ZENITH STENT GRAFT AT AN ACADEMIC INSTITUTION. MATERIAL AND METHODS BETWEEN 2000 AND 2010, 282 PATIENTS, CONSTITUTING THE STUDY GROUP, UNDERWENT ELECTIVE EVAR WITH THE ZENITH STENT GRAFT (COOK, BLOOMINGTON, IN, USA) AT A SINGLE ACADEMIC INSTITUTION. PATIENTS WITH AN ISOLATED ILIAC ARTERY ANEURYSM OR PREVIOUS OPEN AORTIC SURGERY WERE EXCLUDED FROM THE STUDY. THE INDICATION FOR INITIAL TREATMENT WAS AN ANEURYSM WITH A DIAMETER OF GREATER THAN OR EQUAL TO 55 MM IN MEN AND GREATER THAN OR EQUAL TO 50 MM IN WOMEN. PATIENTS WITH AN INCREASE IN AAA DIAMETER OF GREATER THAN OR EQUAL TO 5 MM OVER A SIX MONTH PERIOD, AS WELL AS THOSE IN WHOM THE AAA BECAME SYMPTOMATIC, WERE ALSO TREATED. ALL PROCEDURES WERE PERFORMED BY A VASCULAR SURGEON TOGETHER WITH AN INTERVENTIONAL RADIOLOGIST IN A HYBRID SUITE. THE PATIENTS WERE MAINLY TREATED UNDER SPINAL ANESTHESIA, AND THE STENT GRAFTS WERE IMPLANTED THROUGH THE COMMON FEMORAL ARTERY. BIFURCATED GRAFTS WERE USED IN MOST OF THE EVAR PROCEDURES, BUT AN AORTO-UNI-ILIAC GRAFT WAS EMPLOYED IF ONE OF THE ILIAC ARTERIES WAS OCCLUDED OR SEVERELY STENOSED. AN ATTEMPT TO EMBOLISE A PATENT INFERIOR MESENTERIC ARTERY (IMA) WAS ALWAYS MADE PRIOR TO STENT GRAFT PLACEMENT. THE PATIENTS WERE SYSTEMATICALLY FOLLOWED AFTER THE PROCEDURE ACCORDING TO THE PROTOCOL IN TABLE 1. BASED ON THE AVAILABLE DATA AND THE AUTHORS¿ EXPERIENCE AT THE TIME, THE INITIAL SURVEILLANCE PROTOCOL WAS MODIFIED BY REPLACING THE ANNUAL COMPUTED TOMOGRAPHY ANGIOGRAPHY (CTA) SCAN WITH A COLOUR DUPLEX ULTRASONOGRAPHY (CDUS) IN 2005. THEREAFTER, ALL CDUS EXAMINATIONS WERE PERFORMED BY EXPERIENCED VASCULAR SURGEONS. THE CTA SCANS WERE CONTINUED FOR ALL PATIENTS, IF NOT CONTRAINDICATED, 24 MONTHS AFTER THE INITIAL PROCEDURE TO CONFIRM THE RELIABILITY OF THE CDUS EXAMINATIONS. FOR OBESE INDIVIDUALS AND FOR PATIENTS WITH A SUSPECTED COMPLICATION AT CDUS, AN ADDITIONAL CTA SCAN WAS PERFORMED. A PLAIN ABDOMINAL XRAY WAS TAKEN ANNUALLY UP UNTIL 2013. THE PATIENTS WERE FOLLOWED ANNUALLY UNTIL THE END OF (B)(6) 2017. AT THE END OF FOLLOW UP, THE STATUS OF EACH PATIENT WAS CONFIRMED (I.E., DEAD OR ALIVE), IN ADDITION TO OBTAINING INFORMATION ON POSSIBLE EVAR RELATED REINTERVENTIONS UP UNTIL THE END OF THE FOLLOW UP. FOR SOME INDIVIDUALS, LONG-TERM SURVEILLANCE WAS CARRIED OUT IN THE REFERRING HOSPITALS THEY HAD INITIALLY VISITED PRIOR TO THE ENDOVASCULAR AAA TREATMENT. THE REFERRING VASCULAR COLLEAGUES WERE CONTACTED ACCORDINGLY TO OBTAIN FOLLOW UP DATA. PATIENTS WERE CONTACTED IF A DELAY IN A FOLLOW UP EXAMINATION WAS NOTED. ALL CTAS WERE RE-ANALYSED FOR THE STUDY. THE PRIMARY OUTCOMES WERE OVERALL AND AAA RUPTURE FREE SURVIVAL, WHILE THE SECONDARY OUTCOMES INCLUDED GRAFT RELATED COMPLICATION FREE AND RE-INTERVENTION FREE SURVIVAL. DEATHS WERE ASCERTAINED BY RECORD LINKAGE BETWEEN THE STUDY POPULATION AND THE NATIONAL CAUSES OF DEATH REGISTER ON THE BASIS OF THE PERSONAL IDENTIFICATION CODE UNIQUE TO EVERY RESIDENT. SPSS 22.0 FOR WINDOWS WAS USED FOR STATISTICAL ANALYSIS (IBM, (B)(6) USA). KAPLAN-MEIER SURVIVAL ANALYSIS WAS USED TO EXAMINE SURVIVAL. UNI- AND MULTIVARIABLE ANALYSES WERE USED TO ASSESS THE EFFECT OF AGE, SEX, AAA DIAMETER, NECK DIAMETER, NECK LENGTH, NECK ANGULATION, PATENT IMA, PRIMARY TYPE II ENDOLEAK, FOLLOWING THE INSTRUCTIONS FOR USE (IFU), AND OF BEING LOST TO FOLLOW UP ON ANY STENT GRAFT RELATED RE-INTERVENTION (DEFINED LATER IN TABLE 5) AND AAA RUPTURE. A PRIMARY ENDOLEAK WAS DEFINED AS AN ENDOLEAK OBSERVED IN THE PRIMARY PROCEDURE OR IN THE FIRST POST-OPERATIVE CTA ON THE SECOND OR THIRD DAY. STATISTICAL SIGNIFICANCE WAS INDICATED BY A P VALUE LESS THAN OR EQUAL TO .05. RESULTS THE MEDIAN AGE OF THE COHORT WAS 76 YEARS (INTERQUARTILE RANGE [IQR] 11; RANGE 49-92 YEARS) AND MEAN ANEURYSM DIAMETER WAS 61 MM (RANGE 40-110 MM) AT THE TIME OF THE TREATMENT. THE IFU ISSUED FOR THE IMPLANTATION OF THE STENT GRAFT WERE FOLLOWED IN 90% OF THE PATIENTS. PATIENT CHARACTERISTICS ARE PRESENTED IN TABLE 2. THE MEDIAN FOLLOW UP WAS 76 MONTHS (IQR 69; RANGE 0-201 MONTHS). THE STATUS OF EACH PATIENT WAS CONFIRMED AT THE END OF THE FOLLOW UP (I.E., DEAD OR ALIVE). FOR EACH PATIENT, INFORMATION WAS AVAILABLE REGARDING THE LATEST ANNUAL FOLLOW UP EXAMINATION AND POSSIBLE EVAR RELATED RE-INTERVENTIONS AFTER THAT UP TO THE END OF FOLLOW UP. IN SOME CASES, THE FOLLOW UP INTERVAL WAS PROLONGED, OR THE SURVEILLANCE DISCONTINUED AFTER UNCOMPLICATED SURVEILLANCE, IF THE PATIENT WAS PHYSICALLY FRAIL OR IF LIFE EXPECTANCY SHORT. THESE PATIENTS WERE REPORTED AS LOST TO FOLLOW UP. SPINAL ANAESTHESIA WAS USED IN MOST CASES (N = 271;96%). A BIFURCATED STENT GRAFT WAS EMPLOYED IN 268 PATIENTS (95%) AND A UNI-ILIAC STENT IN 14 CASES (5%). THE IMA WAS PATENT IN 186 PATIENTS, AND WAS SUCCESSFULLY EMBOLIZED DURING THE INITIAL PROCEDURE PRIOR TO THE STENT GRAFT PLACEMENT IN 147 CASES (79%). THERE WERE NO INTRA-OPERATIVE DEATHS OR PRIMARY CONVERSIONS. THE THIRTY DAY MORTALITY WAS 1.8% (N = 5), AND THE CAUSES OF DEATH WERE BRAINSTEM INFARCTION (N =1), CARDIAC FAILURE (N = 3), AND CARDIAC ARREST DUE TO TOTAL STENT GRAFT THROMBOSIS (N = 1). THERE WAS NO SIGNIFICANT DIFFERENCE IN 30 DAY MORTALITY BETWEEN PATIENTS UNDER AND OVER 80 YEARS OF AGE (1.5% VS. 2.3%; P = .500). THE OVERALL SURVIVAL RATES WERE 93% (SE 0.02), 83% (SE 0.03), 61% (SE 0.08), 25% (SE 0.16), AND 9% (SE 0.19) AT 1, 2, 5, 10, AND 16 YEARS, RESPECTIVELY (FIG. 1). THE CORRESPONDING FIGURES FOR THOSE AGED GREATER THAN OR EQUAL TO 80 YEARS WERE 93% (SE 0.03), 80% (SE 0.04), 46% (SE 0.19), AND 4% (SE 0.42) AT 1, 2, 5, AND 10 YEARS, RESPECTIVELY. THERE WAS NO SIGNIFICANT DIFFERENCE IN SURVIVAL BETWEEN PATIENTS UNDER AND OVER 80 YEARS OF AGE AT TWO YEARS OF SURVEILLANCE (P = .529). THE MOST COMMON CAUSES OF DEATH WERE A CARDIAC EVENT (N =70; 32%) AND CANCER (N =48; 22%). THE AUTOPSY RATE WAS 26%. THERE WERE 10 (3.5%) AAA RUPTURES DURING FOLLOW UP (TABLE 3). ALL RUPTURES WERE CONFIRMED BY CTA (N = 8) OR AUTOPSY (N = 2). MEAN TIME TO RUPTURE WAS 81 MONTHS (RANGE 22 - 173 MONTHS). CONSEQUENTLY, THE CUMULATIVE AAA RUPTURE FREE SURVIVAL WAS 100%, 98% (SE 0.01), 96% (SE 0.02), AND 79% (SE 0.12) AT 1, 5, 10, AND 16 YEARS, RESPECTIVELY. EIGHT OF THE 10 PATIENTS WITH A RUPTURED AAA (RAAA) REACHED THE HOSPITAL, AND TWO OF THEM UNDERWENT AN EMERGENCY OPEN REPAIR. ONE OF THESE PATIENTS DIED IN THE INTENSIVE CARE UNIT TWO DAYS AFTER THE PROCEDURE, AND THE OTHER WAS STILL ALIVE AT THE END OF THE FOLLOW UP. IN THREE RAAA CASES, AN ADDITIONAL PROXIMAL CUFF WAS PLACED DURING AN EMERGENCY PROCEDURE, BUT ADEQUATE PROXIMAL SEALING WAS NOT ACHIEVED, AND THE PATIENTS DIED POST-OPERATIVELY. THEY WERE NOT CONSIDERED CANDIDATES FOR OPEN CONVERSION. IN THE REMAINING THREE CASES, ENDOVASCULAR REPAIR WAS NOT CONSIDERED TECHNICALLY POSSIBLE AND OPEN REPAIR WAS REFUSED OWING TO THE PATIENTS¿ HIGH AGE AND MULTIPLE COMORBIDITIES. DURING SURVEILLANCE, ONE PATIENT WITH A RAAA WAS FOUND TO HAVE A TYPE IB ENDOLEAK, BUT, AT THE AGE OF 90 YEARS, HE WAS SEVERELY DEMENTED AND NO FURTHER PROCEDURES WERE PERFORMED. IN ONE CASE, THE RUPTURE WAS UNEXPECTEDLY FOUND AT AUTOPSY 14 YEARS AFTER THE PRIMARY TREATMENT. THIS FEMALE PATIENT HAD REFUSED FURTHER FOLLOW UP AFTER 9 YEARS OF UNCOMPLICATED SURVEILLANCE, WITH A SHRUNKEN AAA DIAMETER OF 30 MM. THE NINE OTHER PATIENTS WERE UNDER SURVEILLANCE AT THE TIME OF THE RAAA. SEVEN OF THE 10 AAA RUPTURES WERE DUE TO TYPE I ENDOLEAK (IA: N = 5; IB: N= 2). SIX OF THESE PATIENTS HAD A PERSISTENT PRIMARY TYPE II ENDOLEAK IN EARLY SURVEILLANCE PRIOR TO THE DEVELOPMENT OF THE TYPE I ENDOLEAK. IN THREE OF THESE CASES, AS THE ANEURYSM SAC SHOWED ENLARGEMENT IN SURVEILLANCE, REPEATED UNSUCCESSFUL ATTEMPTS TO EMBOLISE THE TYPE II ENDOLEAK WERE MADE. IN THREE CASES, HOWEVER, THE TYPE II ENDOLEAK SEEMED STABLE, WITH NO DETECTABLE CHANGE IN ANEURYSM SAC SIZE DURING EARLY SURVEILLANCE, BUT AAA ENLARGEMENT WAS NOTED APPROXIMATELY THREE YEARS AFTER THE PRIMARY TREATMENT. BY THAT TIME, THESE PATIENTS HAD ALSO DEVELOPED A TYPE I ENDOLEAK BEFORE AAA RUPTURE. ONE RUPTURE PATIENT REFUSED FURTHER PROCEDURES FOR THE TYPE IA ENDOLEAK AND DIED THREE YEARS AFTER THE INITIAL PROCEDURE. IN THE CASE OF THREE OF THE 10 RUPTURES, NO SIGNIFICANT ENDOLEAK HAD BEEN NOTED DURING SURVEILLANCE. ONE PATIENT WITH A PRIMARY TYPE II ENDOLEAK SHOWED CLEAR SHRINKAGE OF THE ANEURYSM SAC DURING THE EARLY SURVEILLANCE. HOWEVER, AT 22 MONTHS, THE ANEURYSM RUPTURED. ANOTHER PATIENT WAS FOLLOWED FOR ENDOTENSION, WITH NO SIGNS OF AN ENDOLEAK, AND HE DIED OF A RAAA FIVE YEARS AFTER THE PRIMARY PROCEDURE. ONE RUPTURE, AS MENTIONED EARLIER, WAS FOUND AT AUTOPSY WITH NO PRIOR HISTORY OF AN ENDOLEAK. THE UNIVARIABLE ANALYSIS SHOWED A SIGNIFICANT ASSOCIATION BETWEEN PRIMARY TYPE II ENDOLEAK AND AAA RUPTURE (ODDS RATIO [OR] 4.09, 95% CI 1.12 - 14.9; P = .033). AN EVEN MORE SIGNIFICANT ASSOCIATION WAS FOUND BETWEEN PRE-OPERATIVE ANEURYSM SIZE AND LATE AAA RUPTURE (OR 1.06, 95% CI 1.02 - 1.11; P = .007). THESE RESULTS ALSO REMAINED CONSTANT IN THE MULTIVARIABLE ANALYSIS, WITH AN OR OF 4.65 (95% CI 1.14 - 18.9; P = .032) FOR TYPE II ENDOLEAK AND 1.07 (95% CI 1.02 - 1.13; P= .009) FOR PREOPERATIVE ANEURYSM SIZE. CONSEQUENTLY, AT THE END OF THE 16 YEARS FOLLOW UP; THE MEAN ANNUAL ANEURYSM RUPTURE RATE WAS 0.49% (TABLE 4). A TOTAL OF 168 GRAFT RELATED COMPLICATIONS WERE SEEN IN A TOTAL OF 119 PATIENTS (42%). ALL PRIMARY TYPE II ENDOLEAKS WERE TAKEN INTO ACCOUNT, EVEN IF THEY WERE DETECTED ONLY IN THE PERI-OPERATIVE ANGIOGRAM OR ON THE FIRST CTA. ALL COMPLICATIONS ARE LISTED IN TABLE 3. THE FREEDOM FROM ANY STENT GRAFT RELATED COMPLICATION WAS 68% (SE 0.03), 63% (SE 0.05), 58% (SE 0.09), 54% (SE 0.17), AND 52% (SE 0.21) AT 1, 2, 5, 10, AND 16 YEARS, RESPECTIVELY (FIG. 2). COMPLICATIONS SEEMED TO ACCUMULATE DURING THE EARLY YEARS OF SURVEILLANCE BUT WERE ENCOUNTERED UP TO 14 YEARS AFTER THE PRIMARY PROCEDURE. A TYPE II ENDOLEAK WAS THE MOST COMMON COMPLICATION (N = 79; 28%), WITH ALMOST HALF (48%) OF THEM SEALING SPONTANEOUSLY AND THE VAST MAJORITY (89%) OCCURRING DURING THE FIRST TWO YEARS OF FOLLOW UP. OF THOSE WITH A PATENT IMA (39 PATIENTS) AFTER THE INITIAL PROCEDURE, 18 (46%) DEVELOPED EITHER A TYPE II ENDOLEAK (N =14) OR ENDOTENSION (N = 4). MOST OF THESE PATIENTS (88%) WITH EITHER A TYPE II ENDOLEAK OR ENDOTENSION ALSO HAD A PATENT IMA ON SURVEILLANCE, AND ONLY TWO OF THEM REQUIRED LATER EMBOLISATION. IN ALL, ONLY 26 (33%) OF ALL TYPE II ENDOLEAKS REQUIRED ENDOVASCULAR TREATMENT DUE TO ANEURYSM SAC ENLARGEMENT OVER THE STUDY PERIOD. THE PRIMARY EMBOLIZATION STRATEGY WAS COIL EMBOLISATION; IF THAT PROVED UNSUCCESSFUL, A SECONDARY ATTEMPT WAS MADE WITH GLUE. A TOTAL OF 13 PATIENTS WERE DIAGNOSED WITH ENDOTENSION (GREATER THAN OR EQUAL TO 5 MM) WITHOUT A DETECTABLE ENDOLEAK DURING THE SURVEILLANCE. FOR ONE PATIENT, THERE WAS AN ATTEMPT TO TREAT THE ENDOTENSION BY RELINING THE OLD GRAFT, WHILE THE REST WERE FOLLOWED CAREFULLY. TWO OF THESE PATIENTS LATER DEVELOPED A TYPE I ENDOLEAK, AND THEY WERE TREATED AT THAT POINT WITH AN ADDITIONAL CUFF AND LIMB GRAFT. NONE OF THE PATIENTS WITH ENDOTENSION SUFFERED A RAAA. RE-INTERVENTION FOR A GRAFT RELATED COMPLICATION WAS REQUIRED FOR 61 PATIENTS (22%). ALL ADDITIONAL PROCEDURES ARE LISTED IN TABLE 5. THE RATE OF FREEDOM FROM GRAFT RELATED REINTERVENTIONS WAS 95% (SE 0.01), 91% (SE 0.03), 80% (SE 0.08), 73% (SE 0.11), AND 70% (SE 0.16) AT 1, 2, 5, 10, AND 16 YEARS, RESPECTIVELY (FIG. 3). ONLY FIVE (1.8%) LATE CONVERSIONS TO OPEN REPAIR WERE REQUIRED, TWO AS AN EMERGENCY REPAIR OF A RUPTURE AS MENTIONED EARLIER. TWO PATIENTS HAD A PERSISTENT TYPE I ENDOLEAK DESPITE THE PLACEMENT OF AN EXTRA PROXIMAL CUFF, AND CONVERSIONS WERE PERFORMED AT 12 AND 27 MONTHS, RESPECTIVELY. ONE PATIENT HAD A STENT FRACTURE AND TYPE III ENDOLEAK 149 MONTHS AFTER THE PRIMARY TREATMENT. THE OTHER CASE OF TYPE III ENDOLEAK DEVELOPED AS A RESULT OF LIMB GRAFT MIGRATION AND WAS SUCCESSFULLY TREATED WITH AN ADDITIONAL LIMB GRAFT AT 47 MONTHS. ALL ELECTIVE CONVERSIONS WERE UNCOMPLICATED. A PRIMARY TYPE II ENDOLEAK WAS FOUND TO BE THE ONLY INDEPENDENT SIGNIFICANT FACTOR TO INCREASE THE RISK OF ANY LATE RE-INTERVENTION (OR 6.79, 95% CI 3.61 - 12.79; P < .001). DURING FOLLOW UP, ANEURYSM SAC SHRINKAGE OF _5 MM WAS SEEN IN 61% OF THE CASES, WITH AT MEAN DECREASE OF 18 MM (RANGE 5 - 41 MM) IN MAXIMUM DIAMETER.¿ REFERENCE: VÄÄRÄMÄKI, S., SALENIUS, J.-P., PIMENOFF, G., UURTO, I., & SUOMINEN, V. (2019). SYSTEMATIC LONG-TERM FOLLOW UP AFTER ENDOVASCULAR ABDOMINAL AORTIC ANEURYSM REPAIR WITH THE ZENITH STENT GRAFT. EUROPEAN JOURNAL OF VASCULAR AND ENDOVASCULAR SURGERY, 58(2), 182¿188. HTTPS://DOI.ORG/10.1016/J.EJVS.2019.03.006. THE FOCUS OF THIS REPORT IS THE PATIENT WHO HAD A ZENITH AAA ENDOVASCULAR GRAFT ILIAC LEG IMPLANTED AND EXPERIENCED A TYPE 1B ENDOLEAK RESULTING IN ABDOMINAL AORTIC ANEURYSM (AAA) RUPTURE AND DEATH. Manufacturer Narrative: B2 - DATE OF DEATH: DATE OF DEATH IS UNKNOWN. THE DATE PROVIDED IS THE PUBLICATION DATE OF THE JOURNAL ARTICLE. B3 - DATE OF EVENT: 01/01/2000 TO 03/31/2017. D6A - IMPLANT DATE: (B)(6) 2000 TO (B)(6) 2010. E1 - PHONE: (B)(6). H3 - DEVICE EVALUATED BY MFG? DEVICE NOT RETURNED TO MANUFACTURER. THIS REPORT INCLUDES INFORMATION KNOWN AT THIS TIME. A FOLLOW-UP REPORT WILL BE SUBMITTED SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803 AND IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT CONTAINED HEREIN IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED OR THAT A DEATH OR SERIOUS INJURY OCCURRED; NOR IS IT ADMISSION THAT ANY COOK DEVICE CAUSED, CONTRIBUTED TO, OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22302993&pc=NIK | 3004209178-2025-10617 | 2024-11-03 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-20 04:00 | NIK | CLARIA MRI CRT-D SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P010031 |  | 1 | Event Description: IT WAS REPORTED, BY A REPRESENTATIVE OF THE PATIENT, THAT THE PATIENT HAD PASSED AWAY AND NOTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) "DID NOT FUNCTION PROPERLY". Manufacturer Narrative: CONTINUATION OF D10: 693565 LEAD, IMPLANTED: (B)(6) 2012. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22289080&pc=OZD | 1220648-2025-29510 | 2025-05-23 04:00 | Death | ABIOMED, INC. | 2025-06-19 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Tachycardia; Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT EXPERIENCED FIVE MINUTES OF VENTRICULAR TACHYCARDIA (VT)/VENTRICULAR FIBRILLATION (VF). THE PATIENT WAS DEFIBRILLATED SIX TIMES AND WAS GIVEN AMIODARONE AND LIDOCAINE BOLUSES. SUPPORT WAS CONTINUED FOR APPROXIMATELY A WEEK OR SO. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22289147&pc=DYE | 2025587-2025-04335 | 2024-10-06 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-19 04:00 | DYE | HANCOCK II AORTIC BIOPROSTHETIC HEART VALVE | Insufficient Information | Insufficient Information | P980043 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT 7 YEARS AND 8 MONTHS POST IMPLANT OF THIS 23MM AORTIC BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22289199&pc=OZD | 1220648-2025-29518 | 2025-05-26 04:00 | Death | ABIOMED, INC. | 2025-06-19 04:00 | OZD | IMPELLA CP | Malposition of Device | Hemolysis | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND THAT THE PUMP WAS POSITIONED DEEP WITH THE PIGTAIL FOLDING ON ITSELF. THE PUMP WAS RETRACTED TO THE AORTIC VALVE. THE PATIENT DEVELOPED HEMOLYSIS OVERNIGHT AND ONE LITER BOLUS WAS GIVEN WHICH ALLOWED THE PUMP FLOW TO BE INCREASED. THE HEMOLYSIS CONTINUED TO LINGER FROM THE POOR POSITIONING THAT WAS CORRECTED. THE PATIENT CONTINUED TO DETERIORATE WHICH LEAD TO A DISCUSSION WITH THE FAMILY DECIDING TO TRANSITION TO COMFORT CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿. SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿. IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿. SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS, INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES, MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿. SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22289275&pc=OZD | 1220648-2025-29527 | 2025-05-29 04:00 | Death | ABIOMED, INC. | 2025-06-19 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THE PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT DEVELOPED HEMOLYSIS. THE PATIENT HAD HEMOLYSIS IN THE EFFLUENT BAG AND THEIR PLASMA FREE HEMOGLOBIN WAS TRENDING UP. THE SOURCE OF THE HEMOLYSIS WAS NOT CONFIRMED. POSITION WAS VERIFIED ON ECHOCARDIOGRAM. SUPPORT CONTINUED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22289415&pc=OZD | 1220648-2025-29544 | 2025-06-01 04:00 | Death | ABIOMED, INC. | 2025-06-19 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT (UNKNOWN RACE AND ETHNICITY) WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. FOLLOWING THE IMPLANT, SAME DAY, THE PATIENT'S LEG BECAME MARBLED/MOTTLED DURING THE MCS. A FEM-FEM BYPASS WAS RECOMMENDED; HOWEVER, THE DECISION WAS MADE TO NOT INTERVENE. THE NEXT DAY THE PATIENT EXPIRED. ACCORDING TO HEAD PHYSICIAN, THERE WAS NO ISOLATED ISCHEMIA IN THE LEG WHERE THE IMPELLA WAS LOCATED IN THE FEMORAL ARTERY; THERE WAS A GENERALIZED CAPILLARY HYPOPERFUSION IN ALL TERMINAL VASCULAR AREAS DURING MAXIMALLY DOSED VASOPRESSOR THERAPY. Manufacturer Narrative: THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP DEVICE USED AS AN ASSOCIATED FACTOR GIVEN THE LIMB ISCHEMIA. HOWEVER, IT SHOULD BE NOTED ACCORDING TO THE PHYSICIAN, THE PATIENT'S CAUSE OF DEATH WAS PROLONGED CARDIOGENIC SHOCK (SCAI SHOCK STAGE E) FOLLOWING A MYOCARDIAL INFARCTION WITH PRE-EXISTING MULTIVESSEL CORONARY ARTERY DISEASE. NEITHER THE HYPOPERFUSION NOR THE FATAL OUTCOME CAN BE CAUSALLY ATTRIBUTED TO THE IMPELLA, BUT RATHER TO THE SEVERE CARDIOGENIC SHOCK. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22289466&pc=OZD | 1220648-2025-29551 | 2025-05-27 04:00 | Death | ABIOMED, INC. | 2025-06-19 04:00 | OZD | IMPELLA CP | Patient-Device Incompatibility; Deformation Due to Compressive Stress | Cardiovascular Insufficiency; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WHO WAS HEMODYNAMICALLY UNSTABLE, WITH HIGH-DOSE CATECHOLAMINES, WAS BEING IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE IMPLANT WAS ABORTED DUE TO THE PATIENT'S ANATOMY, AND THE PATIENT SUBSEQUENTLY EXPIRED.Event Description: IT WAS FURTHER REPORTED BY THE PHYSICIAN THAT THE IMPELLA IMPLANTATION WAS NOT SUCCESSFUL BECAUSE OF KINKING ABOVE THE SHEATH. IMPLANTATION WITH SHEATH 14F X25CM WAS NOT SUCCESSFUL. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI STATES THE FOLLOWING. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. Manufacturer Narrative: B5 UPDATED WITH ADDITIONAL INFORMATION THAT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29551. H6 HEALTH EFFECT - CLINICAL CODE 4582 WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29551. H6 MEDICAL DEVICE PROBLEM CODE 2889 AND COMPONENT CODE 4742 WERE INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29551.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22289477&pc=DSQ | 2916596-2025-03674 | 2025-05-25 04:00 | Death | THORATEC CORPORATION | 2025-06-19 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Disconnection; Pumping Stopped; Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT (ED) ON (B)(6) 2025 AFTER BEING FOUND DOWN AT HOME BY THEIR FAMILY. THE PATIENT ARRESTED AT HOME WHERE THE PATIENT RECEIVED RETURN OF SPONTANEOUS CIRCULATION (ROSC) BEFORE BEING TRANSPORTED TO THE HOSPITAL. IN THE ED, THE PATIENT ONCE AGAIN ARRESTED AND RECEIVED RESUSCITATION EFFORTS FOR AT LEAST 30 MINUTES BEFORE FAMILY AND MEDICAL STAFF MADE THE DECISION TO CEASE EFFORTS AND WITHDRAW CARE. THE PATIENT WAS PRONOUNCED DEAD AT 23:52 ON (B)(6) 2025 AND THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS TURNED OFF AT APPROXIMATELY 00:15 ON (B)(6) 2025 AFTER THE LVAD TEAM WAS CALLED FOR ASSISTANCE. ON INTERROGATION IT WAS NOTED THAT THE DRIVELINE WAS DISCONNECTED INITIALLY AT 22:19 ON (B)(6) 2025 AND REMAINED DISCONNECTED WITH PUMP OFF ALARMS FOR APPROXIMATELY 20 MINUTES, AT WHICH TIME IT APPEARED PARAMETERS RETURNED TO ACCEPTABLE RANGES BEFORE AGAIN TRIGGERING LOW FLOW HAZARDS AND LABILE PARAMETERS UNTIL THE TIME OF DEATH. THE PATIENT¿S BACKUP SYSTEM CONTROLLER WAS NOT AVAILABLE FOR FURTHER INFORMATION. IT WAS UNCLEAR WHY THE PUMP REMAINED OFF FOR 20 MINUTES BEFORE BEING RESTARTED. LOG FILES WERE SENT FOR REVIEW. THE PUMP STOP WAS CAUSED BY THE DRIVELINE BEING DISCONNECTED ON (B)(6) 2025 AT 22:19:34 AND RECONNECTED AT 22:39:24. ONCE THE DRIVELINE WAS RECONNECTED, THE PUMP RETURNED TO OPERATING AT THE SET SPEED. JUST PRIOR TO THE DRIVELINE DISCONNECT, THE PATIENT SWITCHED FROM BATTERY POWER TO THE MOBILE POWER UNIT (MPU) WITHOUT ANY ISSUES. ON (B)(6) 2025 AT 22:53:33, A LOW POWER HAZARD & MULTIPLE OTHER ASSOCIATED ALARM TYPES OCCURRED. THIS WAS CAUSED BY POWER TO THE MPU BEING BRIEFLY INTERRUPTED. ON (B)(6) 2025 BETWEEN 23:18:26 AND 0:10:47, INTERMITTENT LOW FLOW ALARMS OCCURRED. THERE DID NOT APPEAR TO BE ANY TYPE OF MECHANICAL CIRCULATORY SUPPORT (MCS) EQUIPMENT ISSUES THAT CAUSED THESE EVENTS. THE LOW FLOW EVENTS SEEMED TO BE A PATIENT HEMODYNAMICALLY RELATED ISSUE AND NOT A VENTRICULAR ASSIST DEVICE (VAD) RELATED ISSUE. THE FILE ENDS WITH THE DRIVELINE DISCONNECTED ON (B)(6) 2025 AT 00:13:47 AND THE PUMP STOPPED. THE MCS EQUIPMENT APPEARED TO HAVE OPERATED AS INTENDED. THE MPU WAS NOTED TO BE A PART OF THE RECALL. IT WAS ADVISED THAT THE MPU AND CONTROLLERS BE SENT BACK FOR ANALYSIS. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. ADDITIONALLY, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED A PUMP STOP EVENT ASSOCIATED WITH THE DRIVELINE BEING DISCONNECTED AS WELL AS MULTIPLE LOW FLOW ALARMS. A SPECIFIC CAUSE FOR THESE FINDINGS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CONTROLLER EVENT LOG FILE CAPTURED DRIVELINE DISCONNECT AND PUMP OFF ALARMS ON (B)(6) 2025 STARTING AT 22:19:34. FOLLOWING RECONNECTION OF THE DRIVELINE AT 22:39:24 ON THE SAME DAY, MULTIPLE LOW FLOW ALARMS WERE THEN CAPTURED BETWEEN 23:18:26 ON (B)(6) 2025 AND 00:10:47 ON (B)(6) 2025. A DRIVELINE DISCONNECT ALARM FOLLOWED BY A CONTROLLER SHUT DOWN WAS OBSERVED BETWEEN 00:13:47 AND 00:15:16 ON (B)(6) 2025. THE SYSTEM CONTROLLER WAS POWERED DOWN BY 00:15:18. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. SECTION 2 OF THE IFU, "SYSTEM OPERATIONS" (UNDER "SYSTEM CONTROLLER WARNINGS AND CAUTIONS"), STATES "CHECK THE SYSTEM CONTROLLER DRIVELINE CONNECTOR TO CONFIRM THAT THE DRIVELINE IS SECURELY INSERTED IN THE SOCKET. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, THE PUMP STOPS. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, PROMPTLY RECONNECT IT TO RESUME PUMP OPERATION." ADDITIONALLY, SECTION 2 OF THE IFU, UNDER "CONNECTING THE DRIVELINE TO THE SYSTEM CONTROLLER", PROVIDES INSTRUCTIONS ON CONNECTING/DISCONNECTING THE DRIVELINE TO/FROM THE SYSTEM CONTROLLER AND MAKING SURE THAT IT IS FULLY AND PROPERLY INSERTED INTO THE SYSTEM CONTROLLER SOCKET. SECTION 2 OF THE PATIENT HANDBOOK, ¿HOW YOUR HEART PUMP WORKS¿, ALSO ADVISES THE USER TO ¿CHECK THE SYSTEM CONTROLLER DRIVELINE CONNECTOR OFTEN TO CONFIRM THAT THE DRIVELINE IS SECURELY INSERTED IN THE SOCKET. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, THE PUMP WILL STOP. IF THE DRIVELINE DISCONNECTS FROM THE SYSTEM CONTROLLER, PROMPTLY RECONNECT IT TO RESUME PUMP OPERATION¿. THE PATIENT HANDBOOK ALSO EXPLAINS THAT THE PUMP CANNOT RUN WITHOUT POWER. SECTION 4 OF THE IFU, ¿SYSTEM MONITOR,¿ EXPLAINS THAT THE LOW FLOW HAZARD ALARM OCCURS WHEN PUMP FLOW IS LESS THAN 2.5 LPM. A 10-SECOND DELAY IS IMPOSED BETWEEN THE DETECTION OF THE LOW FLOW STATUS AND THE ACTIVATION OF THE ASSOCIATED AUDIO AND VISUAL INDICATORS ON THE SYSTEM CONTROLLER. CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 7 OF THE IFU, "ALARMS AND TROUBLESHOOTING", AND SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING", ADDRESS ALL SYSTEM ALARM CONDITIONS, INCLUDING THE DRIVELINE DISCONNECTED, LOW FLOW, AND PUMP OFF ALARMS, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22289660&pc=OZD | 1220648-2025-29560 | 2025-06-02 04:00 | Death | ABIOMED, INC. | 2025-06-19 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AN IMPLANTED IMPELLA 5.5 PUMP COULD NOT GO HIGHER THAN P-2 WITHOUT ENCOUNTERING SUCTION ISSUES. THE DECISION WAS MADE TO PLACE AN IMPELLA RP FLEX PUMP FOR ADDITIONAL SUPPORT, BUT THE PATIENT SUBSEQUENTLY PASSED AWAY FOLLOWING BIPELLA IMPLANT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. LOW PUMP FLOW: DATA LOG REVIEWED: NO MOTOR CURRENT (MC) SPIKE OUTSIDE OF P-LEVEL CHANGES AND INITIATION OF IMPELLA SUPPORT OBSERVED. SUCTION ALARMS OBSERVED. POSITION UNKNOWN ALARMS SEEN. PLACEMENT SIGNAL (PS) TRENDING DOWN. THE CAUSE OF LOW PUMP FLOW MOST LIKELY WAS PATIENT CONDITION RELATED DUE TO RIGHT VENTRICLE FAILURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22289686&pc=OZD | 1220648-2025-29563 | 2025-05-25 04:00 | Death | ABIOMED, INC. | 2025-06-19 04:00 | OZD | IMPELLA CP | Pumping Stopped; Infusion or Flow Problem | Hemolysis; Sepsis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT ON IMPELLA CP SUPPORT WAS SEPTIC WITH METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS. ADDITIONALLY, LABS WERE COMING BACK HEMOLYZED. THE PHYSICIAN TEAM WAS NOT CONCERNED DUE TO THE FACT THAT THE KIDNEYS WERE ALREADY UNRECOVERABLE, AND NO MORE REPOSITIONING EFFORTS WOULD BE MADE. THE P-LEVEL WAS DECREASED DUE TO HEMOLYSIS. IT WAS FURTHER REPORTED THAT THE IMPELLA CP HAD A SPIKE IN MOTOR CURRENT AND PUMP STOP OCCURRED.THE DEVICE WAS SUCCESSFULLY RESTARTED IN P2 AND RAMPED UP TO P7. SATURATED FLOWS WERE NOTICED. THE PUMP WAS RUNNING AT P7. THERE WERE NO ANTICOAGULANTS INFUSING RELATED TO THE PATIENTS DISEASE PROCESS. THE FAMILY ELECTED TO WITHDRAW CARE, AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF TEMPORARY PUMP STOP, INFECTION/SEPSIS, HEMOLYSIS, AND SATURATED FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. TEMPORARY PUMP STOP: DATA LOGS WERE REVIEWED AND A MOTOR CURRENT SPIKE WITH SPEED DEVIATION OCCURRED JUST BEFORE AN ALARM 13 "IMPELLA STOPPED - MOTOR CURRENT HIGH" ALARM WAS TRIGGERED. PUMP WAS ABLE TO BE RESTARTED, PURGE PRESSURE INCREASED AND PURGE FLOWS DECREASED AFTER RESTART. PUMP FLOWS ALSO BECAME FULLY SATURATED AND WERE ABLE TO REGAIN SOME PULSATILITY AS PUMP CONTINUED TO RUN. CLINICAL DETAILS NOTED THAT NURSE OBSERVED A MOTOR CURRENT SPIKE PRIOR TO PUMP STOP OCCURRING. PUMP WAS ABLE TO BE RESTARTED AT P-2 AND TURNED UP TO P-7. CLINICAL DETAILS ALSO NOTED THAT THE PATIENT WAS NOT ON ANY ANTICOAGULATION INFUSION AT THE TIME OF THE PUMP STOP. THE ROOT CAUSE OF THE TEMPORARY PUMP STOP COULD NOT BE DEFINITIVELY DETERMINED AS NO PRODUCT WAS RETURNED FOR EVALUATION TO CONFIRM PATTERN OBSERVED IN THE LOGS. HEMOLYSIS: DATA LOGS WERE REVIEWED AND LT LOG AT TIME OF REPORTED HEMOLYSIS SHOWED NO POSITIONING ISSUES, SUCTION ALARMS, OR MOTOR CURRENT SPIKES. CLINICAL DETAILS NOTED THAT PATIENTS LABS WERE HEMOLYZING. PATIENT HAD CHRONIC KIDNEY DISEASE (CKD) AND CLINICAL TEAM WAS AWARE THAT KIDNEYS WERE UNRECOVERABLE. CLINICAL TEAM OPTED TO NOT REPOSITION PUMP AT THAT TIME. THE ROOT CAUSE OF THE HEMOLYSIS WAS MOST LIKELY PATIENT CONDITION RELATED PER CLINICAL DETAILS. SATURATED FLOWS: DATA LOGS WERE REVIEWED AND LT LOG SHOWED THAT SATURATED FLOWS OCCURRED AFTER A MOTOR CURRENT SPIKE AND PUMP STOP. PRIOR TO MOTOR CURRENT SPIKE, PUMP FLOWS ARE WITHIN NORMAL RANGE, AFTER MOTOR CURRENT SPIKE AND PUMP STOP, PUMP FLOWS ARE FULLY SATURATED BUT REGAIN SOME PULSATILITY AS PUMP CONTINUES TO RUN. CLINICAL DETAILS NOTED THAT AFTER PUMP WAS RESTARTED AFTER PUMP STOP, FLOWS BECAME SATURATED. THE ROOT CAUSE OF THE SATURATED FLOWS COULD NOT BE DEFINITIVELY DETERMINED AS NO PRODUCT WAS RETURNED TO CONFIRM PATTERN OBSERVED IN THE DATA LOGS. INFECTION/SEPSIS: WITHOUT ADDITIONAL CLINICAL DETAILS THE ROOT CAUSE OF THE INFECTION CANNOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST. SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS. ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER. ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION. ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION. SECTION: ENTERING CARDIAC OUTPUT. ¿MAINTAINING ACT AT OR ABOVE 250 SECONDS WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22289707&pc=DSQ | 2916596-2025-03823 | 2025-06-01 04:00 | Death | THORATEC CORPORATION | 2025-06-19 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Coagulation Disorder; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT COMMUNICATED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE HOSPITAL COMMUNICATED NO INFORMATION RELATED TO THE REASON BEHIND THE DEATH OF THE PATENT. IT WAS NOTED THAT THE PATIENT GOT IMPLANTED IN ALREADY IN A VERY SEVERE SITUATION AFTER HE CRASHED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: ADDITIONAL INFORMATION WAS PROVIDED THAT PRIOR TO IMPLANT, THE PATIENT HAD KNOWN THROMBI IN THE LEFT VENTRICLE AND A DECISION WAS MADE TO IMPLANT THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) INSTEAD OF STABILIZING WITH AN IMPELLA. THE SURGEON BELIEVED THEY REMOVED ALL PRE-IMPLANT THROMBI. ALTHOUGH THERE WERE SIGNS OF IMPROVEMENT AND DESPITE HEPARIN INDUCED THROMBOCYTOPENIA (HIT), THE PATIENT PASSED AWAY. THE PATIENT DEVELOPED HEPARIN INDUCED THROMBOCYTOPENIA (HIT) DURING THE IMPLANTATION OF THE SYSTEM. IT WAS NOT COMMUNICATED IF THE PATIENT WAS ON HEPARIN PRIOR TO IMPLANT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22289798&pc=DSQ | 2916596-2025-04054 | 2024-04-07 04:00 | Death | THORATEC CORPORATION | 2025-06-19 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. IT WAS SAID THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS LATER REPORTED THAT THE CAUSE OF DEATH WAS UNKNOWN AND THAT THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REVISION, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22289934&pc=DSQ | 2916596-2025-03783 | 2019-06-03 04:00 | Death | THORATEC CORPORATION | 2025-06-19 04:00 | DSQ | THORATEC® HEARTMATE II®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO NATURAL CAUSES. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT BE EXPLANTED. Manufacturer Narrative: SECTION D4: EXPIRATION DATE AND MANUFACTURE DATE (H4) COULD NOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22290296&pc=QZI | 2124215-2025-40444 | 2025-03-08 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-19 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22290785&pc=LWS | 2649622-2025-16564 | 2025-05-25 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-06-19 04:00 | LWS | SPRINT QUATTRO SECURE MRI SURESCAN | Over-Sensing; Under-Sensing | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS UNDERSENSING ON THE RIGHT VENTRICULAR (RV) LEAD DURING POSSIBLE VENTRICULAR ARRYTHMIA. THERE WAS ALSO A RV LEAD INTEGRITY ALERT (LIA) FOR HIGH RATE NON-SUSTAINED EPISODES AND HIGH SENSING INTEGRITY COUNTER (SIC). IT WAS ALSO NOTED THAT THERE WAS UNDERSENSING ON THE RIGHT ATRIAL (RA) LEAD THAT APPEARED TO RESULT IN INAPPROPRIATE ATRIAL PACING. THE RV LEAD AND RA LEAD REMAIN IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT APPROXIMATELY FIVE DAYS LATER THE PATIENT EXPIRED. ADDITIONAL INFORMATION WAS REQUESTED BUT YIELD NO RESULTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: B1. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: DTPA2D1 CRT-D IMPLANTED: (B)(6) 2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22291374&pc=DSP | 2249723-2025-0002656 | 2025-06-04 04:00 | Death | DATASCOPE CORP. | 2025-06-19 04:00 | DSP | CARDIOSAVE HYBRID, TYPE B PLUG | Pressure Problem | Insufficient Information | K112372 |  | 1 | Event Description: IT WAS REPORTED THAT CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) WAS NOT SHOWING BALLOON PRESSURE DURING USE. THE PATIENT EXPIRED FOR REASONS OTHER THAN IABP THERAPY. Manufacturer Narrative: DUE TO CHARACTER LIMITATION IN E1, THE POSTAL CODE IS (B)(6). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22291523&pc=DTB | 3008973940-2025-04345 | 2025-05-28 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-06-19 04:00 | DTB | CAPSUREFIX NOVUS MRI SURESCAN | Intermittent Capture; Failure to Capture; Pacing Problem | Bradycardia; Non specific EKG/ECG Changes; Asystole; Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE WHEN ATTEMPTING TO CAPTURE THE LEFT BUNDLE BRANCH, THE PHYSICIAN UNFIXED THE ELECTRODE TO GET BETTER CAPTURE. LARGER PAUSES WERE OBSERVED AND THE PHYSICIAN ATTEMPTED TO FIX THE LEAD AND STIMULATE WITH THE PROGRAMMER ANALYZER. THE LEAD WAS FIXED AGAIN TO STIMULATE WITH INTERMITTENT CAPTURE WAS OBSERVED. THE PHYSICIAN HURRIED TO PLACE ANOTHER LEAD TO STIMULATE THE VENTRICLE. WHEN BOTH LEADS WERE STIMULATING LOSS OF CAPTURE WAS OBSERVED AND THE PATIENT EXPERIENCED SINUS ARREST. RESCUE EFFORTS WERE MADE WITH CARDIOPULMONARY RESUSCITATION AND AFTER THIRTY MINUTES IMAGING SHOWED THE HEART HAD ELECTRIC ACTIVITY HOWEVER IT DID NOT MOVE. THE PATIENT PASSED AWAY.Event Description: IT WAS NOTED THAT THE PATIENT EXPERIENCED ASYSTOLE, BRADYCARDIA AND PAUSES. IT WAS NOTED THAT THE HIGH THRESHOLDS AND LOSS OF CAPTURE WERE PATIENT CONDITION.Event Description: IT WAS REPORTED THAT THE PROGRAMMER WAS TURNED OF SPONTANEOUSLY AND NEVER TURNED ON. THE PATIENT WAS DECEASED ON THIS EVENT BUT PROGRAMMER WAS NOT RELATED TO THE DECEASE OF THE PATIENT. TROUBLESHOOTING HAS BEEN DONE BY CHECKING THE POWER CORD WITH OTHER PROGRAMMER AND WORKED NORMALLY AND REPLACED THE PROGRAMMER IMMEDIATELY WITH THE MOBILE PROGRAMMER SO THAT THE IMPLANT PROCEDURE WAS NOT INTERRUPTED BY THE CHANGE OF PROGRAMMER. THE PROGRAMMER IS EXPECTED TO BE RETURNED FOR EVALUATION. Manufacturer Narrative: CORRECTION: G3 FIELD FROM PREVIOUS FOLLOW-UP REPORT 001 WAS OMITTED. THE OMITTED DATE SHOULD HAVE BEEN 09-JUL-2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: G3 FIELD FROM PREVIOUS FOLLOW-UP REPORT 001 WAS OMITTED. THE OMITTED DATE SHOULD HAVE BEEN 25-MAY-2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: G3 FIELD FROM PREVIOUS FOLLOW-UP REPORT 002 WAS OMITTED. THE OMITTED DATE SHOULD HAVE BEEN 28-MAY-2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22291700&pc=NVN | 2649622-2025-16583 | 2025-05-28 04:00 | Death | MPRI | 2025-06-19 04:00 | NVN | SELECTSECURE MRI SURESCAN | Intermittent Capture; Failure to Capture; Pacing Problem | Bradycardia; Non specific EKG/ECG Changes; Asystole; Insufficient Information | P030036 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT PROCEDURE WHEN ATTEMPTING TO CAPTURE THE LEFT BUNDLE BRANCH, THE PHYSICIAN UNFIXED THE ELECTRODE TO GET BETTER CAPTURE. LARGER PAUSES WERE OBSERVED AND THE PHYSICIAN ATTEMPTED TO FIX THE LEAD AND STIMULATE WITH THE PROGRAMMER ANALYZER. THE LEAD WAS FIXED AGAIN TO STIMULATE WITH INTERMITTENT CAPTURE WAS OBSERVED. THE PHYSICIAN HURRIED TO PLACE ANOTHER LEAD TO STIMULATE THE VENTRICLE. WHEN BOTH LEADS WERE STIMULATING LOSS OF CAPTURE WAS OBSERVED AND THE PATIENT EXPERIENCED SINUS ARREST. RESCUE EFFORTS WERE MADE WITH CARDIOPULMONARY RESUSCITATION AND AFTER THIRTY MINUTES IMAGING SHOWED THE HEART HAD ELECTRIC ACTIVITY HOWEVER IT DID NOT MOVE. THE PATIENT PASSED AWAY.Event Description: IT WAS NOTED THAT THE PATIENT EXPERIENCED ASYSTOLE, BRADYCARDIA AND PAUSES. IT WAS NOTED THAT THE HIGH THRESHOLDS AND LOSS OF CAPTURE WERE PATIENT CONDITION.Event Description: IT WAS REPORTED THAT THE PROGRAMMER WAS TURNED OF SPONTANEOUSLY AND NEVER TURNED ON. THE PATIENT WAS DECEASED ON THIS EVENT BUT PROGRAMMER WAS NOT RELATED TO THE DECEASE OF THE PATIENT. TROUBLESHOOTING HAS BEEN DONE BY CHECKING THE POWER CORD WITH OTHER PROGRAMMER AND WORKED NORMALLY AND REPLACED THE PROGRAMMER IMMEDIATELY WITH THE MOBILE PROGRAMMER SO THAT THE IMPLANT PROCEDURE WAS NOT INTERRUPTED BY THE CHANGE OF PROGRAMMER. THE PROGRAMMER IS EXPECTED TO BE RETURNED FOR EVALUATION. Manufacturer Narrative: CORRECTION: G3 FIELD FROM PREVIOUS FOLLOW-UP REPORT 001 WAS OMITTED. THE OMITTED DATE SHOULD HAVE BEEN 25-MAY-2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: G3 FIELD FROM PREVIOUS FOLLOW-UP REPORT 001 WAS OMITTED. THE OMITTED DATE SHOULD HAVE BEEN 09-JULY-2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: G3 FIELD FROM PREVIOUS FOLLOW-UP REPORT 002 WAS OMITTED. THE OMITTED DATE SHOULD HAVE BEEN 28-MAY-2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22291826&pc=OZD | 1220648-2025-29516 | 2025-06-02 04:00 | Death | ABIOMED, INC. | 2025-06-19 04:00 | OZD | IMPELLA 5.5 | Improper or Incorrect Procedure or Method | Cardiac Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT DURING IMPELLA 5.5 INSERTION, NO WIRE WAS USED AND THE HEART WAS DIGITALLY MANIPULATED CAUSING A LEFT VENTRICLE PERFORATION. THE LEFT VENTRICLE WAS PATCHED AND THE IMPELLA 5.5 WAS PLACED AND TURNED ON. THE PATIENT ENDED UP EXPIRING AND THE DEVICE STAYED IN THE PATIENT. Manufacturer Narrative: THE INVESTIGATION OF VENTRICLE PERFORATION HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED FOR EVALUATION. CLINICAL DETAILS NOTED THAT DURING DIRECT ACCESS PUMP INSERTION WITHOUT A WIRE, THE LV WAS PERFORATED AS HEART WAS BEING MANIPULATED BY SURGEON. THE ROOT CAUSE OF THE VENTRICLE PERFORATION WAS MOST LIKELY A USE RELATED ISSUE AS THE HEART SHIFTED DURING INSERTION PROCEDURE. B2 DEATH AND DATE OF DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29516. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AS IT WAS KEPT INSIDE THE DECEASED PATIENT. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ IMPELLA 5.5 WITH SMARTASSIST SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22291873&pc=DTB | 2649622-2025-16602 | 2025-06-03 04:00 | Death | MPRI | 2025-06-19 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Decreased Sensitivity; Capturing Problem | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED ONE WEEK POST IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM, AS PER THE HEALTHCARE PROFESSIONAL, THE PATIENT'S P-WAVES DECREASED AND THE RIGHT ATRIAL (RA) LEAD THRESHOLD HAS INCREASED SINCE IMPLANT. THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS REPORTED ONE WEEK POST IMPLANT OF THE IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM, AS PER THE HEALTHCARE PROFESSIONAL, THE PATIENT'S P-WAVES DECREASED AND THE RIGHT ATRIAL (RA) LEAD THRESHOLD HAS INCREASED SINCE IMPLANT. THE LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. 2025-07-03: IT WAS FURTHER REPORTED THE PATIENT PASSED AWAY APPROXIMATELY TWO WEEKS POST IMPLANT OF THE IPG SYSTEM. ADDITIONAL INFOR MATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22291894&pc=LJS | 2245270-2025-00055 | 2025-05-15 04:00 | Death | VYGON USA | 2025-06-19 04:00 | LJS | PREMICATH | Fluid/Blood Leak | Insufficient Information | K954302 |  | 1 | Event Description: PICC LINE WAS LEAKING AT THE HUB, NEEDLESS ACCESS DEVICE (NAD) CHANGED SEVERAL TIMES, T-CONNECTOR ADDED. LINE WAS LEAKING SIGNIFICANTLY. WE TRIED TROUBLE SHOOTING,USING NUMEROUS DEVICES LIKE (NADS, T-CONNECTORS, VALGUARDS) NOTHING HELPED.Event Description: PICC LINE WAS LEAKING AT THE HUB, NEEDLESS ACCESS DEVICE (NAD) CHANGED SEVERAL TIMES, T-CONNECTOR ADDED. LINE WAS LEAKING SIGNIFICANTLY. WE TRIED TROUBLE SHOOTING,USING NUMEROUS DEVICES LIKE (NADS, T-CONNECTORS, VALGUARDS) NOTHING HELPED. Manufacturer Narrative: ALTHOUGH THE MALFUNCTIONING DEVICE WILL NOT BE RETURNED TO VYGON, THE DETAILS OF THE MALFUNCTION WILL BE EVALUATED AS PART OF THE COMPLAINT INVESTIGATION. THE RESULTS OF THE INVESTIGATION ARE STILL PENDING AND WILL BE COMMUNICATED TO THE FDA. Manufacturer Narrative: ALTHOUGH THE FAULTY SAMPLE WAS NOT RETURNED, THE CUSTOMER PROVIDED ONE OPENED BLISTER AND TWO STERILE UNOPENED BLISTERS FOR ANALYSIS. THE OPENED BLISTER CONTAINED THE CATHETER BUT DID NOT INCLUDE THE NEEDLE-FREE DEVICE, AS SHOWN IN THE IMAGE PROVIDED BY THE CUSTOMER. ADDITIONALLY, IMAGES WERE PROVIDED SHOWING LEAKAGE AT THE CONNECTION BETWEEN THE CATHETER AND THE LUER LOCK HUB, WHERE A NEEDLE-FREE CONNECTOR (NOT A VYGON GERMANY GMBH PRODUCT) WAS ATTACHED TO THE CATHETER. THE OPENED BLISTER WAS SUCCESSFULLY FLUSHED WITH WATER, AND NO LEAKAGE WAS OBSERVED-EVEN AFTER REMOVING THE STYLET OR CLAMPING THE CATHETER TUBE TO INCREASE INTERNAL PRESSURE. ADDITIONALLY, ONE OF THE STERILE BLISTERS WAS OPENED TO TEST ANOTHER CATHETER, AND AGAIN, NO LEAKAGE WAS DETECTED. THEREFORE, WE WERE UNABLE TO REPRODUCE THE ISSUE DESCRIBED IN THE COMPLAINT AND ARE CONSEQUENTLY UNABLE TO DETERMINE ITS CAUSE. THE CUSTOMER SHOULD DIRECT THIS COMPLAINT TO THE MANUFACTURER OF THE NEEDLE-FREE DEVICE (NFD) SHOWN IN THE IMAGE. THE ISSUE MAY BE DUE TO THE NFD NOT BEING SECURELY ATTACHED TO THE CATHETER'S LL-HUB OR BECAUSE THE NFD IS AN INCOMPATIBLE ACCESSORY. A REVIEW OF THE COMPONENT BATCH HISTORY RECORDS WAS PERFORMED, AND NO DEVIATIONS WERE FOUND. THE BATCHES COMPLIED WITH THEIR SPECIFICATIONS AND WERE RELEASED. EACH CATHETER IS FLOW AND LEAK TESTED DURING PRODUCTION. THE TENSILE FORCE AND DIMENSIONS OF CATHETER AND SET COMPONENTS ARE RANDOMLY CHECKED. THE LL-HUBS ARE RANDOMLY TESTED IN ACCORDANCE WITH ISO 80369-7. A 2-YEAR REVIEW OF VYGON USA'S COMPLAINT DATA SHOWS THREE COMPLAINTS FOR LOT 24K011D RELATED TO LEAKING CATHETERS, ALL ORIGINATING FROM THIS FACILITY. INVESTIGATIONS INTO ALL THREE COMPLAINTS DETERMINED THAT THE LEAKS WERE NOT CAUSED BY A MANUFACTURING DEFECT. CORRECTIVE ACTION: BASED ON THE INVESTIGATION THE RETUNED SAMPLES WERE SUCCESSFULLY FLUSHED WITH WATER, AND NO LEAKAGE WAS DETECTED. THEREFORE, NO FURTHER CORRECTIVE ACTION HAS BEEN INITIATED BY VYGON GERMANY'S QUALITY MANAGEMENT AT THIS TIME. THE CUSTOMER SHOULD DIRECT THIS COMPLAINT TO THE MANUFACTURER OF THE NEEDLE-FREE DEVICE (NFD) SHOWN IN THE IMAGE PROVIDED BY THE CUSTOMER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22291925&pc=OZD | 1220648-2025-29517 | 2025-05-26 04:00 | Death | ABIOMED, INC. | 2025-06-19 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND EXPERIENCED VENTRICULAR FIBRILLATION WHICH REQUIRED DEFIBRILLATION. TWO DAYS LATER, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22291998&pc=MCW | 3004742232-2025-00171 | 2025-05-28 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-06-19 04:00 | MCW | DIAMONDBACK PERIPHERAL | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels; Cardiac Tamponade | P130005 |  | 1 | Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED FOR TREATMENT. DURING TREATMENT, A PERFORATION OCCURRED. A CARDIAC MASSAGE WAS PERFORMED, HOWEVER THIS WAS UNSUCCESSFUL AND SURGICAL INTERVENTION WAS REQUIRED.Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED FOR TREATMENT. DURING TREATMENT, A PERFORATION OCCURRED. A CARDIAC MASSAGE WAS PERFORMED, HOWEVER THIS WAS UNSUCCESSFUL AND SURGICAL INTERVENTION WAS REQUIRED. NO ADDITIONAL INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING BALLOON ANGIOPLASTY WAS UNABLE TO PASS THE LESION, WHICH LED TO ORBITAL ATHERECTOMY BEING PERFORMED. TWO TREATMENTS WERE PERFORMED PRIOR TO THE PERFORATION. THE PATIENT EXPERIENCED A PERFORATION-INDUCED TAMPONADE-RELATED HEMODYNAMIC DETERIORATION. CARDIOPULMONARY RESUSCITATION WAS PERFORMED. TWO PERICARDIAL DRAINS WERE ATTEMPTED. ECMO WAS PLACED. THE PATIENT EXPIRED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED PATIENT DEATH, CARDIAC TAMPONADE, PERFORATION OF VESSELS , SURGICAL INTERVENTION, AND UNEXPECTED MEDICAL INTERVENTION APPEAR TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE IT IS POSSIBLE THAT THE REPORTED PATIENT PERFORATION OF VESSELS , SURGICAL INTERVENTION, AND UNEXPECTED MEDICAL INTERVENTION ARE A RESULT OF THE PATIENT¿S DISEASE SEVERITY AND/OR USE TECHNIQUES EMPLOYED; HOWEVER, SINCE THE DEVICE WAS NOT RETURNED THIS COULD NOT BE CONFIRMED. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. UPDATED - B2, H1, H6 (CODES 2226, 1802 ADDED. CODES 4118, 3233, AND 11 NO LONGER APPLY).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22292294&pc=KRA | 2029214-2025-01431 | 2025-06-17 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-06-19 04:00 | KRA | REBAR | Adverse Event Without Identified Device or Use Problem | Nervous System Injury | K093750 |  | 1 | Event Description: MEDTRONIC RECEIVED REPORT THAT ISSUES OCCURRED WHEN USING A REACT-71 ASPIRATION CATHETER DURING A THROMBECTOMY PROCEDURE. IT WAS REPORTED THAT THE REACT-71 AND ALL ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). CATHETER F FLUSHED WAS ADMINISTERED AS INDICATED IN THE IFU. THE REACT-71 COULD NOT PUSH UP WHEN TURNING AND THEN ALSO COULD NOT BE PULLED OUT. THE SURGEON PUSHED THE GAMX TO GO UPPER AND WAS ABLE TO PULL OUT THE REACT-71. WHEN REMOVED, IT WAS FOUND THAT THE DISTAL END OF THE REACT-71 WAS KINKED SO IT WAS REPLACED WITH A FIRMER CATHETER TO CONTINUE THE PROCEDURE. IT WAS NOTED THAT A REBAR MICROCATHETER WAS ALSO USED INT HE PROCEDURE THOUGH NOT DEVICE ISSUE WAS REPORTED REGARDING THE REPORT. IT WAS NOTED THAT THERE WAS NO PATIENT HARM OR INJURY ASSOCIATED WITH THE DEVICE ISSUES. HOWEVER, THE PATIENT DIED THE SAME DAY, (B)(6) 2025.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED CONDITION BEING TREATED WAS ACUTE INTRACRANIAL LARGE ARTERY OCCLUSION. AUTOPSY NOT AVAILABLE. CAUSE OF DEATH WAS NOT DETERMINED. DEATH NOT THOUGHT TO BE RELATED TO REACT-71 CATHETER. UNKNOWN IF DEATH IS RELATED TO PROCEDURE OR PATIENT BASELINE CONDITION. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: REACT-71 (B609931). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22292363&pc=DQX | 2124215-2025-40148 | 2005-01-01 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-19 04:00 | DQX | AMPLATZ SUPER STIFF | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Hemorrhage/Bleeding; Myocardial Infarction; Pneumonia; Renal Failure; Sepsis; Peritonitis; No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC VIA AN ARTICLE PUBLISHED IN THE BRACHYTHERAPY JOURNAL. THE STUDY WAS AIMED TO INVESTIGATE COMPLICATIONS AND MORTALITY AFTER ABLATIVE TREATMENT IN A LARGE, UNSELECTED STUDY COHORT. 4,374 LIVER ABLATIVE PROCEDURES FOR BOTH PRIMARY AND SECONDARY LIVER MALIGNANCIES CONDUCTED BETWEEN 2005 AND 2020 WERE INCLUDED IN THE STUDY. AN AMPLATZ SUPER STIFF GUIDEWIRE WAS USED TO REMOVE A COAXIAL NEEDLE PREVIOUSLY USED TO PUNCTURE THE INTRAHEPATIC LESION THE PATIENT WAS BEING TREATED FOR. THE MOST FREQUENTLY OBSERVED COMPLICATIONS INCLUDED BLOOD TRANSFUSIONS, EXTENDED VENTILATION EXCEEDING 48 HOURS, SEPSIS, PERITONITIS, AND PNEUMONIA. WITHIN THE STUDY COHORT, PULMONARY EMBOLISM, ACUTE RENAL FAILURE, GASTROINTESTINAL BLEEDING, AND PNEUMONIA WERE THE PREDOMINANT COMPLICATIONS ENCOUNTERED DURING HOSPITALIZATION. THE LENGTH OF STAY (LOS) WAS NOTABLY EXTENDED IN CASES WHERE PERITONITIS, PNEUMONIA, SEPSIS, ACUTE RENAL FAILURE, AND PULMONARY EMBOLISM. OVERALL, THERE WERE ONLY A LIMITED NUMBER OF RECORDED INSTANCES OF IN-HOSPITAL MORTALITY, TOTALING 9 OUT OF 4,374 CASES. THESE FATALITIES WERE LINKED TO PRIOR PARTIAL LIVER RESECTION, COAGULOPATHY, AND ACUTE RENAL FAILURE. PRIOR PARTIAL LIVER RESECTION WAS RECOGNIZED AS AN INDEPENDENT RISK FACTOR FOR COMPLICATIONS AND WAS ASSOCIATED WITH IN-HOSPITAL MORTALITY. SEVERE COMPLICATIONS AND IN-HOSPITAL DEATHS WERE REPORTED SIGNIFICANTLY MORE FREQUENTLY FOLLOWING TACE TREATMENT COMPARED TO ALL OTHER TREATMENT MODALITIES. THE INCREASED OCCURRENCE OF COMPLICATIONS AND IN-HOSPITAL MORTALITY AFTER TACE TREATMENT MAY BE ASSOCIATED WITH PRE-EXISTING LIVER CIRRHOSIS.Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC VIA AN ARTICLE PUBLISHED IN THE BRACHYTHERAPY JOURNAL. THE STUDY WAS AIMED TO INVESTIGATE COMPLICATIONS AND MORTALITY AFTER ABLATIVE TREATMENT IN A LARGE, UNSELECTED STUDY COHORT. 4,374 LIVER ABLATIVE PROCEDURES FOR BOTH PRIMARY AND SECONDARY LIVER MALIGNANCIES CONDUCTED BETWEEN 2005 AND 2020 WERE INCLUDED IN THE STUDY. AN AMPLATZ SUPER STIFF GUIDEWIRE WAS INSERTED AFTER AN 18 G COAXIAL NEEDLE WAS USED TO PUNCTURE THE INTRAHEPATIC LESION, FOLLOWED BY THE REMOVAL OF THE NEEDLE AND THE INSERTION OF A NON-BOSTON SCIENTIFIC (BSC) 6F ANGIOGRAPHIC CATHETER SHEATH. FINALLY, A 6 F NON-BSC IRRADIATION CATHETER WAS PLACED WITHIN THE CATHETER SHEATH. IN LARGER OR IRREGULAR-SHAPED LESIONS, MULTIPLE CATHETERS WERE INSERTED. AFTER INSERTION, THE IMAGING DATA (AXIAL SLICES WITH 3 MM THICKNESS) WERE TRANSFERRED TO THE NON-BSC IRRADIATION PLANNING SYSTEM FOR HIGH-DOSE-RATE BRACHYTHERAPY (HDR-BT). THE MOST FREQUENTLY OBSERVED COMPLICATIONS INCLUDED BLOOD TRANSFUSIONS, EXTENDED VENTILATION EXCEEDING 48 HOURS, SEPSIS, PERITONITIS, AND PNEUMONIA. WITHIN THE STUDY COHORT, PULMONARY EMBOLISM, ACUTE RENAL FAILURE, GASTROINTESTINAL BLEEDING, AND PNEUMONIA WERE THE PREDOMINANT COMPLICATIONS ENCOUNTERED DURING HOSPITALIZATION. THE LENGTH OF STAY (LOS) WAS NOTABLY EXTENDED IN CASES WHERE PERITONITIS, PNEUMONIA, SEPSIS, ACUTE RENAL FAILURE, AND PULMONARY EMBOLISM. OVERALL, THERE WERE ONLY A LIMITED NUMBER OF RECORDED INSTANCES OF IN-HOSPITAL MORTALITY, TOTALING 9 OUT OF 4,374 CASES. THESE FATALITIES WERE LINKED TO PRIOR PARTIAL LIVER RESECTION, COAGULOPATHY, AND ACUTE RENAL FAILURE. PRIOR PARTIAL LIVER RESECTION WAS RECOGNIZED AS AN INDEPENDENT RISK FACTOR FOR COMPLICATIONS AND WAS ASSOCIATED WITH IN-HOSPITAL MORTALITY. SEVERE COMPLICATIONS AND IN-HOSPITAL DEATHS WERE REPORTED SIGNIFICANTLY MORE FREQUENTLY FOLLOWING TACE TREATMENT COMPARED TO ALL OTHER TREATMENT MODALITIES. THE INCREASED OCCURRENCE OF COMPLICATIONS AND IN-HOSPITAL MORTALITY AFTER TACE TREATMENT MAY BE ASSOCIATED WITH PRE-EXISTING LIVER CIRRHOSIS. UPON FURTHER REVIEW, IT WAS DETERMINED THAT THIS REPORT WAS SUBMITTED IN ERROR. THERE ARE NO ADVERSE EVENTS ATTRIBUTABLE TO THE AMPLATZ SUPER STIFF GUIDEWIRE. Manufacturer Narrative: BLOCK B3: EXACT DATE OF EVENT IS UNKNOWN; THEREFORE, FIRST DAY OF TREATMENT MONTH/YEAR HAS BEEN SELECTED. BLOCK D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK E1: INITIAL REPORTER FACILITY NAME: (B)(6). BLOCK G2: LITERATURE SOURCE: MARCH, C., THORMANN, M., HASS, P., GEORGIADES, M., SENSSE, M., HERRMANN, T., OMARI, J., PECH, M., DAMM, R. MORTALITY AND POSTINTERVENTIONAL COMPLICATIONS AFTER ABLATIVE TREATMENT OF LIVER MALIGNANCIES: A COHORT STUDY OF 4374 PATIENTS. BRACHYTHERAPY JOURNAL 23 (2024) 743-750, HTTPS://DOI.ORG/10.1016/J.BRACHY.2024.07.001. BLOCK H6: IMDRF PATIENT CODE E1024 CAPTURES THE REPORTABLE EVENT OF PERITONITIS. IMDRF PATIENT CODE E0306 CAPTURES THE REPORTABLE EVENT OF SEPSIS. IMDRF PATIENT CODE E050303 CAPTURES THE REPORTABLE EVENT OF PULMONARY EMBOLISM. IMDRF PATIENT CODE E061202 CAPTURES THE REPORTABLE EVENT OF MYOCARDIAL INFARCTION. IMDRF PATIENT CODE E130501CAPTURES THE REPORTABLE EVENT OF RENAL FAILURE. IMDRF PATIENT CODE E0506 CAPTURES THE REPORTABLE EVENT OF HEMORRHAGE. IMDRF PATIENT CODE E0733 CAPTURES THE REPORTABLE EVENT OF PNEUMONIA. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF DEATH. IMDRF IMPACT CODE F1901CAPTURES THE REPORTABLE EVENT OF ADDITIONAL SURGERY. IMDRF IMPACT CODE F08 CAPTURES THE REPORTABLE EVENT OF HOSPITALIZATION. IMDRF IMPACT CODE F2302 CAPTURES THE REPORTABLE EVENT OF BLOOD TRANSFUSION. IMDRF IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF UNEXPECTED MEDICAL INTERVENTION. Manufacturer Narrative: CORRECTIONS TO B5 DESCRIBE EVENT OR PROBLEM, D2A COMMON DEVICE NAME, D2B PRO CODE (PRODUCT CODE), G1 MANUFACTURER CONTACT, AND H6 PATIENT CODES, IMPACT CODES, AND EVALUATION CONCLUSION CODES. BLOCK B3: EXACT DATE OF EVENT IS UNKNOWN; THEREFORE, FIRST DAY OF TREATMENT MONTH/YEAR HAS BEEN SELECTED. BLOCK D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK E1: INITIAL REPORTER FACILITY NAM: (B)(6). BLOCK G2: LITERATURE SOURCE: MARCH, C., THORMANN, M., HASS, P., GEORGIADES, M., SENSSE, M., HERRMANN, T., OMARI, J., PECH, M., DAMM, R. MORTALITY AND POSTINTERVENTIONAL COMPLICATIONS AFTER ABLATIVE TREATMENT OF LIVER MALIGNANCIES: A COHORT STUDY OF 4374 PATIENTS. BRACHYTHERAPY JOURNAL 23 (2024) 743-750, HTTPS://DOI.ORG/10.1016/J.BRACHY.2024.07.001.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22292426&pc=PYX | 1220648-2025-29533 | 2025-05-31 04:00 | Death | ABIOMED, INC. | 2025-06-19 04:00 | PYX | IMPELLA RP FLEX | Adverse Event Without Identified Device or Use Problem | Thrombocytopenia | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH KNOWN CARDIAC HISTORY AND END-STAGE RENAL DISEASE (ESRD) WAS IMPLANTED WITH AN IMPELLA RP FLEX DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THERE WAS A DIAGNOSIS MADE OF HEPARIN INDUCED THROMBOCYTOPENIA (HIT) THAT PROMPTED HEPARIN REMOVAL FROM SYSTEMIC IV, THE PURGE WAS NOT RUNNING WITH HEPARIN BUT RATHER SODIUM BICARB. THE PUMP SUPPORTED FOR 4 DAYS AND THEN THE PATIENT EXPIRED FROM CARE. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22292467&pc=OAE | 3008452825-2025-00291 |  | Death | ABBOTT MEDICAL | 2025-06-19 04:00 | OAE | TACTIFLEX¿ ABLATION CATHETER, SENSOR ENABLED¿ | Adverse Event Without Identified Device or Use Problem | Perforation of Esophagus; Insufficient Information | P220013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED 2-3 WEEKS AFTER A PULMONARY VEIN ISOLATION PROCEDURE THAT WAS PERFORMED ON (B)(6) 2024. THE CAUSE OF DEATH IS UNKNOWN. THERE WERE NO PERFORMANCE ISSUES DURING THE PROCEDURE WITH ANY ABBOTT DEVICE. FURTHER INFORMATION HAS BEEN REQUESTED BUT IS NOT AVAILABLE AT THIS TIME.Event Description: THIS REPORT INCLUDES AN UPDATE TO THE HEALTH EFFECT CLINICAL CODE. Manufacturer Narrative: ADDITIONAL INFORMATION: B5, G3, H2. CORRECTED DATA: H6. Manufacturer Narrative: ADDITIONAL INFORMATION: D9, G3, H2, H3, H6 THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE TACTISYS LOG FILES RETURNED WERE IN A FILE FORMAT THAT WAS UNABLE TO BE ANALYZED BY THE LOG VIEWER PROGRAM. THEREFORE, NO LOG FILE ANALYSIS WAS PERFORMED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED. DUE TO UNKNOWN PROCEDURAL CONDITIONS, WE ARE UNABLE TO CONCLUSIVELY DETERMINE THE CAUSE OF THE REPORTED EVENT, AND THE CAUSE REMAINS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22292869&pc=DSQ | 2916596-2025-03710 | 2025-05-28 04:00 | Death | THORATEC CORPORATION | 2025-06-19 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. THE DEVICE OPERATED AS INTENDED. THE PATIENT HAD A DIFFICULT POSTOPERATIVE COURSE, AND THE DEVICE WAS DEACTIVATED ON (B)(6) 2025. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22292970&pc=OZD | 1220648-2025-29558 | 2025-05-23 04:00 | Death | ABIOMED, INC. | 2025-06-19 04:00 | OZD | IMPELLA CP | Failure to Prime | Cardiovascular Insufficiency; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AN IMPELLA CP PUMP FAILED TO PROPERLY PRIME DURING PREPARATION FOR IMPLANT. THE PUMP WAS REPLACED TO RESOLVE THE ISSUE. NO PATIENT HARM WAS REPORTED.Event Description: THE MEDICAL SAFETY OFFICER REVIEWED AND DETERMINED THE FOLLOWING: THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP PUMP 1 USED AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THERE WAS AN INTERRUPTION IN THE PROCEDURE AS A RESULT OF THE DEVICE MALFUNCTION (NO PURGE NOTED EXITING THE PUMP). HOWEVER, IT MUST BE NOTED CARE WAS WITHDRAWAL BECAUSE THE PATIENT NEVER QUITE RECOVERED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ Manufacturer Narrative: THE INVESTIGATION OF PRIMING ISSUE HAS BEEN COMPLETED. BASED ON PURGE PRESSURE AND FLOWS BEING WITHIN SPECIFICATION AT THE BEGINNING OF THE CASE IN FIELD DATA LOGS AND RETURNED PRODUCTS BEING ABLE TO PRIME WITHOUT ISSUE, THE ROOT CAUSE OF THE PRIMING ISSUE WAS DETERMINED TO MOST LIKELY BE NO PROBLEM FOUND. THE MEDICAL SAFETY OFFICER REVIEWED AND DETERMINED THE FOLLOWING: THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP PUMP 1 USED AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. THERE WAS AN INTERRUPTION IN THE PROCEDURE AS A RESULT OF THE DEVICE MALFUNCTION (NO PURGE NOTED EXITING THE PUMP). HOWEVER, IT MUST BE NOTED CARE WAS WITHDRAWAL BECAUSE THE PATIENT NEVER QUITE RECOVERED. D9 DEVICE RETURNED TO MANUFACTURER DATE ADDED B1 ADVERSE EVENT WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29558. B2 PATIENT OUTCOME WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29558. B5 NARRATIVE OF PATIENT OUTCOME AND MEDICAL SAFETY OFFICER REVIEW WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29558. D6B EXPLANT DATE WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29558. H1 TYPE OF REPORT WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29558. H6 HEALTH EFFECT - CLINICAL CODE 4582 WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29558. NEW CODE ENTERED. H6 HEALTH EFFECT - IMPACT CODE 1802 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29558.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22293023&pc=DQX | 2126666-2025-00029 | 2025-05-20 04:00 | Death | LAKE REGION MEDICAL | 2025-06-19 04:00 | DQX | CIRCULO | Use of Device Problem; Insufficient Information | Arrhythmia; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | K221575 |  | 1 | Event Description: PROCEDURE DESCRIPTION IT WAS REPORTED THAT ON (B)(6) 2025, A 27MM NAVITOR VISION VALVE WAS CHOSEN FOR A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE USING AN UNKNOWN DELIVERY SYSTEM. DURING PROCEDURE, THE PATIENT HAD SUDDENLY NO AORTIC PRESSURE (ASYSTOLE) AND DESPITE REANIMATION THE PATIENT DID NOT RECOVER. A TRANSTHORACIC ECHOCARDIOGRAM (TTE) SHOWED MASSIVE PERICARDIAL EFFUSION (TAMPONADE AND A PERICARDIOCENTESIS WAS PERFORMED. AFTER THAT, NO RECOVERY TO THE PATIENT WAS NOTED AND HAD PERSISTENT VENTRICULAR FIBRILLATION. THEY DECIDED TO DO A STERNOTOMY; A RUPTURE OF THE LEFT VENTRICLE APPEARED TO BE THE PROBLEM DUE TO PERFORATION WITH THE TAVI GUIDEWIRE. THE PATIENT PASSED AWAY DURING THE PROCEDURE. HAS THE HEALTHCARE PROFESSIONAL STATED THAT THEY BELIEVE THE PATIENT OR USER EXPERIENCED ADVERSE HEALTH CONSEQUENCES DUE TO THE PERFORMANCE OF THE PRODUCT?: NO. PATIENT STATUS: DEAD. PRODUCT REPLACEMENT: WAS THE PRODUCT REPLACED?: NO. QUESTION: WAS THERE AN ABBOTT EMPLOYEE PRESENT FOR DURING PROCEDURE? ANSWER: YES. QUESTION: WAS THE DELIVERY SYSTEM USED AN ABBOTT PRODUCT? IF YES, PLEASE PROVIDE MODEL AND BATCH/LOT NUMBER. ANSWER: FNAV-DS-LG, LOT 10572353. QUESTION: AT WHAT POINT OF THE PROCEDURE, "THE PATIENT HAD SUDDENLY NO AORTIC PRESSURE (ASYSTOLE) AND DESPITE REANIMATION THE PATIENT DID NOT RECOVER." WAS NOTED? ANSWER: DURING DEPLOYMENT OF VALVE ( AT 80%) JUST AFTER 1ST RESHEATH QUESTION: WAS A PERICARDIAL EFFUSION NOTED PRIOR TO PROCEDURE? ANSWER: NO. QUESTION: DOES THE USER BELIEVE AN ABBOTT DEVICE CAUSED THE PERICARDIAL EFFUSION? ANSWER: YES, THE CIRCULO GUIDEWIRE HAS CAUSED A PERFORATION IN THE LV. QUESTION: IF THE USER BELIEVES AN ABBOTT DEVICE CAUSED THE PERICARDIAL EFFUSION, PLEASE SPECIFY WHICH ONE. ANSWER: CIRCULO-XS, LOT UNKNOWN. QUESTION: IF THE PERICARDIAL EFFUSION WAS NOT RELATED TO THE ABBOTT DEVICE, WHAT DOES THE USER BELIEVE CAUSED THE PERICARDIAL EFFUSION? ANSWER: PERFORATION. QUESTION: DID THE PHYSICIAN(S) CONCLUDE THAT CARDIAC TAMPONADE WAS PRESENT? ANSWER: YES, BY TTE CONFIRMED. QUESTION: WAS THE VALVE EVER RE-SHEATHED MORE THAN TWO TIMES? IF YES, IS THE PHYSICIAN AWARE OF THE IFU WARNING AGAINST THIS? ANSWER: 1 TIME. QUESTION: DURING THE PROCEDURE, WHAT WERE THE PATIENT'S ACTIVATED CLOTTING TIME (ACT) LEVELS? ANSWER: AROUND 300-350. QUESTION: DURING THE PROCEDURE, WAS HEPARIN ADMINISTERED? IF SO, HOW MUCH? ANSWER: YES, 100 IE/KG (STANDARD PROTOCOL). QUESTION: DURING THE PROCEDURE, WAS THERE ANY DIFFICULTY IMPLANTING THE DEVICE, SUCH AS MULTIPLE MANIPULATIONS? ANSWER: NO, WITH 1ST ATTEMPT JUST TOO HIGH AT THE LCC ( IN LAO VIEW), SO NEED FOR RESHEATHING. MAYBE TOO MUCH PUSH ON THE GW BY THE 2ND IMPLANTER? QUESTION: PLEASE PROVIDE THE SERIAL NUMBER OF THE 27MM NAVITOR VISION VALVE? ANSWER: (B)(6). QUESTION: PLEASE CONFIRM WAS THE 27MM NAVITOR VISION VALVE REMAIN IMPLANTED? ANSWER: YES. QUESTION: DID THEY ADMINISTER ANY PROTAMINE OR REVERSAL AGENT DURING PROCEDURE OR AFTER THE PROCEDURE? ANSWER: NO. QUESTION: WERE THERE MULTIPLE WIRE OR DELIVERY SHEATH EXCHANGES? ANSWER: NO, NOT MORE THAN USUAL PROCEDURE. QUESTION: WHAT WAS THE CAUSE OF DEATH? ANSWER: RUPTURE OF THE FREE WALL OF THE LV. QUESTION: WHAT WAS THE PATIENT'S CONDITION PRIOR TO THE IMPLANT PROCEDURE? ANSWER: GOOD. QUESTION: WHAT WAS THE PATIENT'S CONDITION AFTER THE IMPLANT PROCEDURE? ANSWER: SHE DIED JUST AFTER THE IMPLANTATION OF THE VALVE (IN THE CATHLAB). QUESTION: DOES THE IMPLANTER CLASSIFY THIS DEATH AS BEING RELATED TO THE PERFORMANCE OF THE IMPLANTED DEVICE, OR LIKELY RELATED TO THE PROCEDURE AND THE PATIENT'S COMORBIDITIES? (I.E. INQUIRE ABOUT THE FUNCTION OF THE IMPLANTED PRODUCT IF APPLICABLE). ANSWER: THE PROCEDURE (WRONG MANIPULATION GW). QUESTION: COULD YOU PLEASE PROVIDE OPERATIVE NOTES FROM THE PROCEDURE, WITH IDENTIFIABLE PATIENT INFORMATION REDACTED IF APPLICABLE? ANSWER: PREDIL WITH A 22MM BALLOON. START DEPLOYMENT IN CUSP OVERLAP VIEW. 1 RESHEATH BECAUSE OF TOO HIGH POSITION LCC IN LAA VIEW. JUST AFTER RESHEATH AND IMPLANTATION AT 80%, (B)(6) > REANIMATION, INTUBATION. TTE SHOWED TAMPONADE AND PERICARDIAL EFFUSION> PERICARDIAL PUNCTURE> NO RECOVERY> STERNOTOMY> (B)(6) DIED. QUESTION: PLEASE PROVIDE IMPLANT PROCEDURAL IMAGING, INCLUDING ECHOCARDIOGRAPHY AND ANGIOGRAPHY? DICTATED REPORTS OF ANY THERAPEUTIC INTERVENTIONS? IF NOT AVAILABLE, PLEASE PROVIDE A JUSTIFICATION. (PLEASE REDACT ANY IDENTIFIABLE PATIENT INFORMATION.) ANSWER: MEDIA WAS PROVIDED WITH THE COMPLAINT. ADDITIONAL INFORMATION RECEIVED FROM THE DISTRIBUTOR ON (B)(6) 2025: QUESTION: WHAT IS THE CONTEXT OF THE IMAGE? WAS THIS BEFORE, DURING, OR AFTER CPR? ANSWER: AFTER. QUESTION: WAS THE GUIDEWIRE LOCATION VERIFIED PRIOR TO DEPLOYMENT OF THE VALVE? ANSWER: YES. QUESTION: WAS THE DISTAL CURVE OF THE GUIDEWIRE VISUALIZED DURING DEPLOYMENT OF THE VALVE? ANSWER: YES. QUESTION: HAS ABBOTT'S MEDICAL TEAM REVIEWED THE MEDIA, AND CAN YOU PROVIDE A SUMMARY OF FINDINGS? ANSWER: GUIDEWIRE MUST HAVE PERFORATED THE LV DURING FIRST DEPLOYMENT OR RESHEATHING THE 1ST TIME. QUESTION: PLEASE LIST ALL THE OTHER DEVICES USED DURING THE PROCEDURE, INCLUDE THE MODEL, BRAND NAME, SIZES, ETC.? ANSWER: NVTR-LS-LG+: (B)(6) / NVRO-27: (B)(6)/ FNAV-DS-LG: (B)(6)/ CIRCULO-XS. QUESTION: PLEASE CONFIRM WAS THE CURVE OF THE CIRCULO GUIDEWIRE CONSTRAINED WITHIN A CATHETER DURING INSERTION INTO AND WITHDRAWAL FROM THE BODY OR TREATMENT SITE? ANSWER: YES. QUESTION: PLEASE CONFIRM WAS THE WIRE POSITION MAINTAINED WHILE TRACKING OR ADVANCING A CATHETER OR DEVICE OVER THE WIRE? ANSWER: YES, HOWEVER THERE COULD HAVE BEEN RESIDUAL PUSH. QUESTION: COULD YOU PLEASE VERIFY IF CIRCULO IS STILL AVAILABLE FOR ANALYSIS AND IF YES, WHAT IS THE PROJECTED DATE OF THE DEVICE RETURN TO INTEGER/LAKE REGION MEDICAL? ANSWER: NO. QUESTION: IF CIRCULO IS NOT AVAILABLE TO BE RETURNED, WHAT IS THE REASON FOR NO RETURN? ANSWER: IN THE PERIPROCEDURAL RUSH THE GW WAS THROWN AWAY. QUESTION: IS THIS A NEW USER OR EXPERIENCE USER? ANSWER: EXPERIENCED. QUESTIONS REGARDING THE NAVITOR VALVE: QUESTION: PLEASE CONFIRM WAS THERE ANY ISSUES NOTED ON THE NAVITOR VALVE? ANSWER: NO. QUESTION: WAS THE NAVITOR VALVE PERFORMED AS INTENDED? ANSWER: YES. Manufacturer Narrative: THIS MEDWATCH REPORT IS AN UPDATE TO THE PREVIOUS REPORT. THE DEVICE INVOLVED IN THE EVENT WAS DISCARDED THEREFORE, NO PHYSICAL OR VISUAL ANALYSIS OF THE PRODUCT COULD BE PERFORMED. A REVIEW OF THE DEVICE HISTORY RECORDS OF THE SPECIMEN DEVICE DOES NOT PRESENT ANY INDICATION OF MANUFACTURING DEFECT OR ANOMALY THAT COULD HAVE IMPACTED ON THE EVENT AS REPORTED. DURING MANUFACTURING AND PACKAGING OF THIS PRODUCT THE OPERATORS ARE INSTRUCTED TO 100% VISUALLY INSPECT FOR ANY OBVIOUS DEFECT PRIOR TO SHIPMENT. THE HISTORY RECORDS INDICATE THIS PRODUCT WAS FINAL INSPECTION TESTED AT LAKE REGION MEDICAL AND WAS DETERMINED TO BE ACCEPTABLE. A FURTHER EVALUATION OF THE AVAILABLE AND PROVIDED INFORMATION INCLUDING IMAGING WAS UNDERTAKEN BY AN INTERVENTIONAL CARDIOLOGIST ON BEHALF OF THE MANUFACTURER. SUMMARY PROVIDED BY THE INTERVENTIONAL CARDIOLOGIST: 1. FROM THE CINE IMAGES PROVIDED, IT APPEARS THAT THE LEFT VENTRICLE SIZE IS SMALL. IN PATIENTS WITH A SMALL LEFT VENTRICLE, THERE IS ALWAYS A HIGHER RISK FOR THE WIRE TO BE POSITIONED TOO DEEP IN THE LEFT VENTRICLE DURING TAVR. 2. IN ONE BAV IMAGE, THE LEFT VENTRICLE APEX IS NOT VISIBLE. THE GUIDEWIRE APPEARS TO BE POSITIONED MUCH DEEPER IN COMPARISON TO THE OTHER IMAGES PROVIDED FOR REVIEW. 3. ANOTHER IMAGE PROVIDED SHOWS THE TAVR DEVICE LOADED. FROM THE TIME OF POSITIONING THE GUIDEWIRE AND UNTIL THE TAVR VALVE GETS LOADED AND DEPLOYED, THE WIRE CAN MOVE TOO DEEP INTO THELEFT VENTRICLE DURING TAVR VALVE ADVANCEMENT WITHOUT MAINTAINING CONTROL OF THE GUIDEWIRE. THAT PARTICULAR LOSS OF CONTROL OF GUIDEWIRE AND INADVERTENT ADVANCEMENT OF THE WIRE INTO THE LEFT VENTRICLE IS NOT DOCUMENTED. WHEN THE OPERATORS REALIZE THE WIRE GOT INADVERTENTLY ADVANCED INTO THE LEFT VENTRICLE TOO DEEP, THE IMMEDIATE REACTION IS TO PULL THE WIRE BACK INTO THE IDEAL POSITION. THAT SEQUENCE, THE MAJORITY OF TIME, HAPPENS DURING REGULAR FLUORO IMAGING RATHER THAN CINE IMAGING. IT IS EXTREMELY HARD TO DOCUMENT IT. THIS INADVERTENT ADVANCEMENT OF THE GUIDEWIRE IS OBVIOUS IN THE LEFT VENTRICLE ANGIOGRAM WHERE THE PIGTAIL CATHETER HAD BEEN ADVANCED DEEP INTO THE LEFT VENTRICLE ALMOST DOUBLE THE USUAL LENGTH. 4. CONCLUSION FROM PROVIDED INFORMATION: THE LEFT VENTRICLE ANGIOGRAM INDICATES CONTRAST APPEARS TO BE LEAVING THE LEFT VENTRICLE AT THE APEX WHICH REINFORCES MY IMPRESSION THAT THERE IS A HIGH CHANCE THAT THE GUIDEWIRE, WHICH WAS POSITIONED DEEP IN THE LEFT VENTRICLE TO START WITH. DURING THE SUBSEQUENT STEPS OF THE TAVR PROCEDURE, THE WIRE WAS INADVERTENTLY ADVANCED MUCH FURTHER, CAUSING LEFT VENTRICLE PERFORATION. THE FOLLOWING INFORMATION FROM THE USER INFORMATION, WARNINGS, AND PROCEDURE SECTIONS ARE INCLUDED IN THE DEVICE INSTRUCTIONS FOR USE AND INFORM THE USERS OF STEPS TO TAKE TO PREVENT VENTRICLE PERFORATION FROM OCCURRING. AS INDICATED IN THE USER INFORMATION: THE GUIDEWIRE SHOULD BE USED ONLY BY PHYSICIANS OR CLINICIANS TRAINED IN THE INTRODUCTION AND PLACEMENT OF INTERVENTIONAL DEVICES WITHIN THE CHAMBERS OF THE HEART, INCLUDING THOSE USED DURING TRANSCATHETER AORTIC VALVE PROCEDURES. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE WARNINGS: 1. THE GUIDEWIRE SHOULD BE USED ONLY BY PHYSICIANS TRAINED IN THE INTRODUCTION AND PLACEMENT OF INTERVENTIONAL DEVICES INCLUDING THOSE USED WITHIN TRANSCATHETER AORTIC VALVE PROCEDURES. 2. CAREFULLY READ ALL INSTRUCTIONS PRIOR TO USE. OBSERVE ALL WARNINGS AND PRECAUTIONS. FAILURE TO DO SO MAY RESULT IN COMPLICATIONS. 4. MONITOR GUIDEWIRE POSITION THROUGHOUT THE PROCEDURE FOR PROPER PLACEMENT OF CURVE AND DISTAL TIP. 9. WHEN ADVANCING OR REMOVING THE GUIDEWIRE, ALWAYS USE FLUOROSCOPIC GUIDANCE WITH RADIOGRAPHIC EQUIPMENT THAT PROVIDES HIGH- RESOLUTION IMAGES. NEVER POSITION THE GUIDEWIRE BLINDLY, AS THIS MAY RESULT IN MISPLACEMENT, DISSECTION OR PERFORATION. 11. NEVER ADVANCE THE GUIDEWIRE AGAINST THE RESISTANCE WITHOUT FIRST DETERMINING THE REASON FOR RESISTANCE UNDER FLUOROSCOPY. EXCESSIVE FORCE AGAINST RESISTANCE MAY RESULT IN DAMAGE TO THE CATHETER OR VESSEL/ORGAN. CARE SHOULD BE TAKEN WHEN ADVANCING A GUIDEWIRE AFTER DEVICE DEPLOYMENT. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE PROCEDURE: 10. ADVANCE THE INTERVENTIONAL DEVICE OVER THE GUIDEWIRE WHILE MAINTAINING GUIDEWIRE POSITION. A. VISIBLY MONITOR THE GUIDEWIRE DOUBLE CURVE WHEN ADVANCING OR WITHDRAWING A CATHETER OR INTERVENTIONAL DEVICE OVER THE WIRE. IF RESISTANCE IS MET WHILE ADVANCING THE DEVICE OVER THE WIRE STOP; EVALUATE THE CAUSE BEFORE PROCEEDING (USE FLUOROSCOPY IF NECESSARY). WE REVIEWED THE ASSOCIATED RISK DOCUMENTATION RELATIVE TO THIS PRODUCT AND CONFIRMED THAT THIS INCIDENT IS LISTED AND THAT IT IS TRENDING WITHIN THE ESTABLISHED OCCURRENCE THRESHOLD. LAKE REGION'S EFFECTIVENESS OF DESIGN VERIFICATION ACTIVITIES BRINGS THE OVERALL RISK DOWN TO AS FAR AS POSSIBLE. ARRHYTHMIA, PERICARDIAL EFFUSION, CARDIAC TAMPONADE, CARDIAC PERFORATION, AND DEATH ARE KNOWN POTENTIAL ADVERSE EVENTS AND ARE LISTED IN THE CIRCULO DEVICE INSTRUCTIONS FOR USE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FOLLOW UP REPORT. THE LOT NUMBER FOR THE DEVICE IS UNKNOWN AT THE TIME OF THIS REPORT; THEREFORE, SECTION D4 COULD NOT BE COMPLETED, INCLUDING THE UDI NUMBER. ADDITIONAL EVENT INFORMATION HAS BEEN REQUESTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22293120&pc=DSQ | 2916596-2025-03961 | 2025-06-04 04:00 | Death | THORATEC CORPORATION | 2025-06-19 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO BRAIN DEATH. AN AUTOPSY WAS NOT PERFORMED AND THE DEVICE WAS NOT EXPLANTED.Event Description: DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT DEATH. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS INVESTIGATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING NEUROLOGIC DYSFUNCTION AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS NEUROLOGIC DYSFUNCTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22293139&pc=OZD | 1220648-2025-29564 | 2025-06-02 04:00 | Death | ABIOMED, INC. | 2025-06-19 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A CP PUMP PLACED ALONG WITH ECMO TO SUPPORT THE PATIENT WITH KNOWN CARDIAC AND SYSTEMIC COMORBIDITIES. THE PUMP WAS SUPPORTING WHEN THE LIMB WITH BOTH ECMO AND CP CANNULA BECAME ISCHEMIC. THE MEDICAL THERAPY DOSING MAY HAVE CONTRIBUTED TO THE ISCHEMIA, AND THE PUMP WAS NOT EXPLANTED. AFTER JUST UNDER 5 DAYS THE PUMP SUPPORT WAS STOPPED/WITHDRAWN DUE TO UNDERLYING CONDITION AND ANOXIC BRAIN INJURY. THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22293214&pc=DSQ | 2916596-2025-03716 | 2025-03-26 04:00 | Death | THORATEC CORPORATION | 2025-06-19 04:00 | DSQ | THORATEC¿¿ HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Wound Dehiscence; Hemorrhage/Bleeding; Sepsis; Post Operative Wound Infection; Bowel Perforation; Heart Failure/Congestive Heart Failure; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN.Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT REQUIRED RIGHT VENTRICULAR ASSIST DEVICE (RVAD) SUPPORT AND HAD BEEN ON DIALYSIS PRIOR TO IMPLANT. THE PATIENT DEVELOPED A SEPTIC BOWEL FROM A SMALL BOWEL PERFORATION AND REQUIRED MULTIPLE ABDOMINAL SURGERIES INCLUDING END-ILEOSTOMY. THE CAUSE OF THE BOWEL PERFORATION WAS UNKNOWN. THE PATIENT CONTINUED TO HAVE INTRA-ABDOMINAL BLEEDING AND SEPSIS. THE PATIENT THEN DEVELOPED A STERNAL INCISION DEHISCENCE WITH MEDIASTINITIS. THE PATIENT ASKED TO BE PLACED ON HOSPICE CARE. THE PATIENT'S DEATH WAS NOT THOUGHT TO BE DIRECTLY RELATED TO THE LVAD. THE LVAD OPERATED AS EXPECTED. Manufacturer Narrative: SECTION H6: HEALTH EFFECT - CLINICAL CODE CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. FURTHER, A SPECIFIC CAUSE FOR THE PATIENT¿S SEPSIS COULD NOT BE CONCLUSIVELY DETERMINED. THE PUMP WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, BLEEDING, SEPSIS, AND WOUND DEHISCENCE, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RVAD PLACEMENT. SECTION 6 (UNDER ¿ANTICOAGULATION¿) ALSO PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INR RANGE, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: B5 NARRATIVE INFORMATION UPDATED. H6 - CODES UPDATED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22293353&pc=OZD | 1220648-2025-29569 | 2025-05-30 04:00 | Death | ABIOMED, INC. | 2025-06-19 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT DISTAL PULSES WERE NOT FOUND FOLLOWING IMPELLA CP IMPLANT. THE STAFF OPTED TO CONTINUE WITH SUPPORT WHILE MONITORING THE CONDITION. THE PATIENT PASSED AWAY THE FOLLOWING DAY DURING SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22293479&pc=DSQ | 2916596-2025-04057 | 2025-05-24 04:00 | Death | THORATEC CORPORATION | 2025-06-19 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Adult Respiratory Distress Syndrome; Hypoxia; Pneumonia; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS INTUBATED AND HAD POOR PROGNOSIS. THE PATIENT PASSED AWAY.Event Description: IT WAS LATER REPORTED THAT THE PATIENT HAD ACUTE RESPIRATORY DISTRESS SYNDROME (ARDS) LEADING UP TO THEIR DEATH. A COMPUTED TOMOGRAPHY OF THEIR CHEST REVEALED CONCERN FOR ORGANIZING PNEUMONIA AND ARDS OF UNKNOWN ETIOLOGY. RESPIRATORY CULTURES WERE TAKEN AND REMAINED NEGATIVE. DEATH WAS NOT CONSIDERED TO BE DEVICE RELATES AS IT WAS ALSO SAID THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: CORRECTED DATA: SECTION G6: REPORT TYPE CORRECTED TO INCLUDE 30-DAY. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22293517&pc=DYB | 2015691-2025-05014 | 2025-05-29 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-19 04:00 | DYB | EDWARDS ESHEATH+ INTRODUCER SET | Adverse Event Without Identified Device or Use Problem |  | K200258 |  | 1 | Event Description: ACCORDING TO THE TERRITORY MANAGER (TM), DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE UTILIZING A 26MM SAPIEN 3 ULTRA VALVE VIA TRANSFEMORAL APPROACH, THE ESHEATH WAS SLIGHTLY RETRACTED TO ALLOW FOR VALVE ALIGNMENT PRIOR TO THE FIRST BEND OF ABDOMINAL AORTIC TORTUOSITY. VALVE ALIGNMENT WAS ACHIEVED WITHOUT COMPLICATION. ADVANCEMENT OF THE DELIVERY SYSTEM REQUIRED INCREASED FORCE DUE TO THE TORTUOSITY, AND THE INITIAL ATTEMPT WAS UNSUCCESSFUL IN NAVIGATING THE AORTIC ARCH. A LUNDERQUIST WIRE WAS SUBSEQUENTLY INTRODUCED THROUGH THE PIGTAIL CATHETER ON THE CONTRALATERAL SIDE, WHICH FACILITATED SUCCESSFUL TRAVERSAL OF THE TORTUOUS ANATOMY AND PASSAGE OVER THE ARCH. SIGNIFICANT FLEXING OF THE DELIVERY SYSTEM WAS OBSERVED DURING ADVANCEMENT. AS THE VALVE NEARED THE ANNULUS, ANESTHESIA REPORTED A SUDDEN DROP IN THE PATIENT'S BLOOD PRESSURE. AN ECHOCARDIOGRAM REVEALED NO PERICARDIAL EFFUSION, AND PHYSICAL EXAMINATION OF THE GROIN AND ABDOMEN WAS UNREMARKABLE. AN ABDOMINAL AORTOGRAM IDENTIFIED CONTRAST EXTRAVASATION NEAR THE FIRST BEND OF THE AORTA, LIKELY AT THE PROXIMAL END OF THE SHEATH. BLOOD PRODUCTS WERE ADMINISTERED. A CODA BALLOON WAS DEPLOYED IN AN ATTEMPT TO TAMPONADE THE EXTRAVASATION, WITH PROLONGED INFLATION AND INTERMITTENT AORTOGRAPHY SHOWING REDUCED LEAKAGE. VASCULAR SURGERY WAS CONSULTED BUT DEEMED THE PATIENT INOPERABLE FOR BOTH SURGICAL AND ENDOVASCULAR REPAIR. THE PATIENT WAS STABILIZED AND TRANSFERRED TO THE ICU. ALTHOUGH INITIALLY AWAKE AND RESPONSIVE, THE PATIENT EXPERIENCED A SUDDEN DROP IN BLOOD PRESSURE AND, DESPITE SUPPORTIVE MEASURES, WAS PRONOUNCED DECEASED APPROXIMATELY ONE HOUR AFTER LEAVING THE CATHETERIZATION LAB. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. AORTIC DISSECTION OR RUPTURE MAY OCCUR WHEN MULTIPLE ATTEMPTS ARE MADE TO CROSS THE STENOTIC NATIVE VALVE, AND/OR WHEN EXCESSIVE FORCE IS USED. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). THE THV TRAINING MANUALS GUIDE TO FACILITATE SAFE CROSSING OF THE NATIVE OR FAILING VALVE, INCLUDING CAMERA PROJECTIONS, HANDLING DURING ADVANCEMENT, AND TROUBLESHOOTING TECHNIQUES IF DIFFICULTY IS ENCOUNTERED. AS STATED, EXCESSIVE FORCE SHOULD NOT BE USED WHEN THE DEVICE HAS DIFFICULTY CROSSING THE STENOTIC VALVE. ADDING TENSION TO THE WIRE, PULLING BACK THE SYSTEM TO RE-ORIENT THE VALVE, AS NEEDED, AND TORQUING THE FLEX CATHETER MAY HELP SOLVE THE PROBLEM. IN THIS CASE, SIGNIFICANT FLEXING OF THE DELIVERY SYSTEM WAS OBSERVED DURING ADVANCEMENT. AS THE VALVE APPROACHED THE ANNULUS, ANESTHESIA REPORTED A SUDDEN DROP IN THE PATIENT'S BLOOD PRESSURE. AN ECHOCARDIOGRAM REVEALED NO PERICARDIAL EFFUSION, AND PHYSICAL EXAMINATION OF THE GROIN AND ABDOMEN WAS UNREMARKABLE. HOWEVER, AN ABDOMINAL AORTOGRAM IDENTIFIED CONTRAST EXTRAVASATION NEAR THE FIRST BEND OF THE AORTA, LIKELY AT THE PROXIMAL END OF THE SHEATH. THESE FINDINGS SUGGEST THAT MOVEMENT OF THE SHEATH TIP PARTICULARLY DURING EFFORTS TO NAVIGATE THE TORTUOUS ABDOMINAL AORTIC ANATOMY WITH THE COMMANDER DELIVERY SYSTEM MAY HAVE BEEN A CONTRIBUTING FACTOR TO THE VASCULAR INJURY. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22293693&pc=DSQ | 2916596-2025-03711 | 2025-05-25 04:00 | Death | THORATEC CORPORATION | 2025-06-19 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Respiratory Failure; Epistaxis | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS HOSPITALIZED IN THE MORNING OF (B)(6) 2025 DUE TO RECURRENT AND PERSISTENT EPISTAXIS. THE PATIENT INITIALLY PRESENTED TO THE EMERGENCY ROOM (ER) AND WHEN THE RHINO ROCKET INTERVENTION DID NOT CEASE THE BLEEDING, THEY WERE ADMITTED FOR OBSERVATION AND EAR NOSE AND THROAT (ENT) CONSULTATION. ONCE ENT WAS CONSULTED, THE BLEEDING HAD STOPPED AND THE PATIENT WAS DISCHARGED WITH THE INSTRUCTIONS TO RETURN THREE DAYS LATER FOR REMOVAL OF THE RHINO ROCKET AND FURTHER ASSESSMENT/INTERVENTION AND RECOMMENDATIONS. HEMOGLOBIN AND HEMATOCRIT WERE STABLE AT 11.7/35.2 AND MET THE CRITERIA FOR DISCHARGE. THE PATIENT'S PARTNER CALLED EMERGENCY MEDICAL SERVICES (EMS) DUE TO RECURRENT EPISTAXIS THAT NIGHT. THE PATIENT WAS LOADED IN AN AMBULANCE AND A SHORT TIME LATER WENT UNRESPONSIVE WITHOUT CLEAR REASON. THE RHINO ROCKET REMAINED IN PLACE. THE PATIENT HAD COMPLETE RESPIRATORY FAILURE AND NO NEUROLOGIC SIGNS OF LIFE. AN I-GEL WAS PLACED FOR VENTILATORY EFFORTS. IT WAS UNCLEAR WHETHER THEY HAD AN UNDERLYING RHYTHM INITIALLY. THE PATIENT ARRIVED AT THE ER FOR EVALUATION AND WAS NOTED TO HAVE THEIR LEFT VENTRICULAR ASSIST DEVICE (LVAD) ALARMING WITH LOW FLOW ALARMS. THE PATIENT WAS DO NOT RESUSCITATE/DO NOT INTUBATE (DNR/DNI) STATUS AND THIS WAS CONFIRMED BY THE PATIENT'S PARTNER UPON ARRIVAL, SO ALL RESUSCITATIVE EFFORTS STOPPED. THE DEVICE WAS INTERROGATED SHOWING LOW FLOW ALARMS STARTING AT 18:32, BUT IT WAS NOTED THE TIME OF ALARMS WAS TECHNICALLY 19:32 DUE TO THE CLOCK BEING OFF AN HOUR. THIS CONTINUED UNTIL 19:22 (20:22) WHEN THE DEVICE WAS EVENTUALLY DEACTIVATED. TIME OF DEATH WAS OFFICIALLY CALLED AT 20:37 BY THE TEAM ONCE THE DEVICE WAS FULLY DEACTIVATED AND ABSENCE OF HEART TONES AND RESPIRATIONS WERE CONFIRMED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. FURTHER, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE COULD NOT BE CONCLUSIVELY DETERMINED. OF NOTE, SOME OF THE CAPTURED LOW FLOW EVENTS APPEARED TO BE CONSISTENT WITH THE DISCONTINUATION OF DEVICE SUPPORT. THE SUBMITTED LOG FILE CONTAINED EVENTS FROM 01MAY2025 THROUGH (B)(6) 2025. LOW FLOW EVENTS ASSOCIATED WITH LOW FLOW HAZARD ALARMS WERE CAPTURED ON THE REPORTED EVENT DATE OF (B)(6) 2025 BEGINNING AT 04:15:36. ON (B)(6) 2025 AT 19:28:27, THE DRIVELINE AND POWER CABLES WERE DISCONNECTED, CONSISTENT WITH THE REPORTED EVENTS AND THE DISCONTINUATION OF LEFT VENTRICULAR ASSIST DEVICE (LVAD) SUPPORT. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE PUMP APPEARED TO FUNCTION AS INTENDED AT THE FIXED SPEED. THE PUMP WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LVAS IFU AND THE HEARTMATE II PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE, BLEEDING, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. THIS SECTION ALSO ADDRESSES PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿HEARTMATE TOUCH COMMUNICATION SYSTEM¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. THIS SECTION ALSO OUTLINES SITUATIONS THAT CAN RESULT IN A LOW FLOW ALARM AND FURTHER EXPLAINS THAT CHANGES IN PATIENT CONDITIONS CAN RESULT IN LOW FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. SECTION 5 OF THE PATIENT HANDBOOK, "ALARMS AND TROUBLESHOOTING," AND SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, OUTLINE SYSTEM CONTROLLER ALARMS, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH THEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22293770&pc=FOS | 2245270-2025-00057 | 2025-06-08 04:00 | Death | VYGON | 2025-06-19 04:00 | FOS | UMBILICAL CATHETER | Malposition of Device; Positioning Problem | Extravasation; Foreign Body In Patient | K921352 |  | 1 | Event Description: AN UMBILICAL CATHETER WAS INSERTED AT DAY 0 ON A PREMATURE NEWBORN OF 28 WEEKS. 2 DAYS LATER AN EXTRAVASATION OF PARENTERAL WAS NOTICED. THE CATHETER WAS LOCATED IN THE LIVER. THE PREMATURE DIED ON (B)(6) 2025. Manufacturer Narrative: THE INVOLVED UMBILICAL CATHETER IS NOT AVAILABLE. FROM THE DESCRIPTION, THIS SERIOUS ADVERSE EVENT IS CLEARLY RELATED TO THE MALPOSITION OF THE UMBILICAL CATHETER IN THE LIVER. IT IS NOT TRACED TO A DEFECT OF OUR DEVICE. THE BATCH REVIEW DOES NOT SHOW ANY NON-CONFORMITY. THE PRODUCTS ARE IN CONFORMITY WITH THE SPECIFICATIONS. ALL CONTROLS ARE COMPLIANT, INCLUDING THE 100% TESTS CARRIED OUT ON EACH CATHETER DURING THE MANUFACTURING PROCESS. THESE TESTS ARE AS FOLLOWS: THE BLUNTNESS OF THE CATHETER TIP, MARKING THE CATHETERS, TIGHTNESS AND NON-OBSTRUCTION. THIS SERIOUS ADVERSE EVENT IS NOT TRACED TO A MALFUNCTION OR DEFECT OF OUR DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22293874&pc=DSQ | 3007042319-2025-01301 | 2025-06-16 04:00 | Death | HEARTWARE, INC. | 2025-06-19 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE (VAD) WAS MADE COMFORT CARE ONLY AND PASSED AWAY DUE TO HEART FAILURE. THE VAD REMAINS IMPLANTED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS HEART FAILURE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, HEART FAILURE AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22293955&pc=NPT | 2025587-2025-04354 | 2025-06-12 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-19 04:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Positioning; Perivalvular Leak; Malposition of Device; Insufficient Information; Activation Failure; Patient Device Interaction Problem | Myocardial Infarction; Obstruction/Occlusion; Heart Block; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE, ULTRASOUND WAS USED TO GUIDE ACCESS FOR THE LEFT FEMORAL ARTERY (LFA). THE PRECLOSE VASCULAR CLOSURE DEVICE WOULD NOT ENGAGE, AND MULTIPLE UNSUCCESSFUL ATTEMPTS AT VASCULAR CLOSURE WERE MADE. THE SITE WAS CLOSED WITH ANGIOSEAL, AND LFA ACCESS WAS REGAINED ABOVE THE INITIAL SITE, AT WHICH POINT AN OCCLUSION WAS DISCOVERED BELOW THE FIRST ACCESS SITE. THE PROCEDURE CONTINUED WITH A PLAN TO REPAIR THE LFA AFTERWARDS. DILATATION OF THE LFA WAS PERFORMED WITH A 14FR SENTRANT, AND A 14FR DCS WAS INSERTED INLINE OVER A NON-MEDTRONIC GUIDEWIRE. THE FIRST DEPLOYMENT OF THE VALVE WAS TO 80% DEEP ON THE LEFT CORONARY CUSP (LCC) SIDE AND WAS RECAPTURED. THE SECOND DEPLOYMENT WAS ALSO TO 80% AND STILL APPEARED DEEP; WHILE PLANNING TO RECAPTURE, THE VALVE WAS ACCIDENTALLY DEPLOYED DUE TO THE DCS HANDLE BEING QUICKLY ROTATED IN THE WRONG DIRECTION. THE FINAL DEPLOYMENT DEPTH WAS ESTIMATED AT 4 MILLIMETERS ON THE NON-CORONARY CUSP (NCC) AND 7 MILLIMETERS ON THE LCC, THOUGH THE VIEW HAD PARALLAX. THE VALVE APPEARED CONSTRICTED IN THE RAO VIEW, AND MODERATE PARAVALVULAR LEAK (PVL) WAS OBSERVED. FOLLOWING VALVE DEPLOYMENT, THE PATIENT EXPERIENCED INTERMITTENT COMPLETE HEART BLOCK WITH JUNCTIONAL ESCAPE, NECESSITATING THE INSERTION OF A TEMPORARY PACEMAKER. POST-DILATATION WAS PERFORMED WITH AN 18 MILLIMETER NON-MEDTRONIC BALLOON, BUT THE VALVE REMAINED CONSTRICTED WITH PVL. FURTHER POST-DILATATION WITH A 20 MILLIMETER NON-MEDTRONIC BALLOON WAS PERFORMED, RESULTING IN TRACE-MILD PVL, NO AORTIC INSUFFICIENCY, AND A MEAN GRADIENT OF 1 MMHG. THE PATIENT RETURNED TO SINUS RHYTHM. THE CARDIAC SURGEON PROCEEDED WITH LFA REPAIR UNDER GENERAL ANESTHESIA. AFTER LFA REPAIR, TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) SHOWED MILD TO MODERATE AORTIC REGURGITATION, THE VALVE DEPTH MEASURED 12 MILLIMETERS BELOW THE ANNULUS, AND THE MITRAL LEAFLET WAS OBSERVED TO BE INTERACTING WITH THE BASE OF THE EVOLUT VALVE. THE MITRAL VALVE HAD NO REGURGITATION AND A MEAN GRADIENT OF 4 MMHG. DUE TO CONCERNS ABOUT POTENTIAL FUTURE MITRAL LEAFLET DAMAGE FROM ONGOING INTERACTION WITH THE EVOLUT VALVE, THE DECISION WAS MADE TO SURGICALLY EXPLANT THE EVOLUT VALVE, REPLACE T HE AORTIC VALVE WITH A SURGICAL VALVE, AND REPLACE THE ASCENDING AORTA DUE TO A PORCELAIN AORTA. LIFE SUPPORT WAS WITHDRAWN TWO DAYS LATER, AND THE PATIENT DIED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED STATING THE MILD-MODERATE AORTIC REGURGITATION WAS PARAVALVULAR LEAK (PVL). PER THE PHYSICIAN, THE VALVE BEING RELEASED TOO DEEP IN THE LEFT VENTRICULAR OUTFLOW TRACT (LVOT) CAUSED THE EMERGENCY CARDIAC SURGERY AND SUBSEQUENTLY DURING THE POST CARDIAC SURGERY, THE PATIENT WAS IN REFRACTORY SHOCK DUE TO THE OCCLUSION OF THE RIGHT COMMON ARTERY AND CONTRIBUTING THE PATIENT DEATH. THE DEATH WAS DUE TO REFRACTORY CARDIOGENIC SHOCK SECONDARY TO THE RIGHT VENTRICLE (RV) DYSFUNCTION DUE TO ACUTE MYOCARDIAL INFARCTION POST CARDIAC SURGERY. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012617232); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE ; EXPLANT DATE SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT FX PLUS VALVE; PRODUCT ID EVFXPLUS-29 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE (B)(6) 2025 SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW: TWO MEDIA FILES WERE PROVIDED FOR REVIEW AND THE PATIENT¿S EXECUTIVE SUMMARY WAS PROVIDED FOR ANATOMICAL REVIEW. THE PATIENT¿S AORTIC VALVE AND SINOTUBULAR JUNCTION APPEARED TO BE SIGNIFICANTLY CALCIFIED WITH AN ANNULUS PERIMETER THAT MEASURED 74.3MM WITH A PERIMETER DERIVED DIAMETER OF 23.6MM SUGGESTING A 29MM VALVE. THE COMPUTED TOMOGRAPHY (CT) MEASUREMENTS SHOWED BILATERAL ILIOFEMORAL ARTERIES WITH MINIMUM DIAMETERS <(><<)> 5MM IN DIFFERENT SEGMENTS. AS STATED IN THE INSTRUCTIONS FOR USE (IFU), PATIENTS MUST PRESENT WITH TRANSARTERIAL ACCESS VESSELS WITH DIAMETERS THAT ARE =5.0 MM WHEN USING THE DEVICE. IN THIS CASE, TRANSFEMORAL ACCESS WAS PROHIBITED. IT WAS REPORTED THAT THE LEFT FEMORAL ARTERY WAS USED AS THE PRIMARY ACCESS AND CONSEQUENTLY VASCULAR COMPLICATIONS WERE REPORTED. THE MEDIA FILES SHOW AN IMPLANTED VALVE THAT APPEARS TO BE UNDER EXPANDED. IT WAS REPORTED THAT THE VALVE WAS INADVERTENTLY RELEASED DEEPER THAN DESIRED, AND A CONDUCTION DISTURBANCE OCCURRED. AS STATED IN THE IFU, POTENTIAL RISKS ASSOCIATED WITH THE IMPLANTATION OF THE BIOPROSTHESIS MAY INCLUDE BUT ARE NOT LIMITED TO CONDUCTION SYSTEM DISTURBANCES (FOR EXAMPLE, ATRIOVENTRICULAR NODE BLOCK, LEFT-BUNDLE BRANCH BLOCK, ASYSTOLE), WHICH MAY REQUIRE A PERMANENT PACEMAKER. FURTHERMORE, BIOPROSTHESIS IMPLANT DEPTH >5 MM MAY CONTRIBUTE TO AN INCREASED RISK OF CONDUCTION DISTURBANCES, WHICH MAY REQUIRE A PERMANENT PACEMAKER. A POST IMPLANT DILATATION WAS PERFORMED, AND FINAL ANGIOGRAM SHOWS AN EXPANDED VALVE THAT APPEARS DEEP (> 5MM) WITHOUT SIGNIFICANT EVIDENCE OF AORTIC REGURGITATION/PARAVALVULAR LEAK. IT WAS REPORTED THAT AFTER A VASCULAR REPAIR, AN ECHOCARDIOGRAM REVEALED MILD TO MODERATION AORTIC REGURGITATION AND MITRAL INTERACTION, AND THE PATIENT WAS CONVERTED TO SURGERY TO EXPLANT THE VALVE. IMAGES OF THE ECHOCARDIOGRAM WERE NOT PROVIDED FOR REVIEW. IT WAS REPORTED THAT THE PATIENT DIED TWO DAYS AFTER SURGERY, BUT THE IMAGES PROVIDED DO NOT VALIDATE THE CAUSE OF DEATH. UPDATED: D4, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED H.6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED B.5 AND H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22294089&pc=OZD | 1220648-2025-29594 | 2025-05-26 04:00 | Death | ABIOMED, INC. | 2025-06-19 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Display or Visual Feedback Problem; Decrease in Pressure | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT AN AIR IN PURGE ALARM APPEARED ON THE AUTOMATED IMPELLA CONTROLLER (AIC) DURING PATIENT SUPPORT. WHILE ATTEMPTING TO DE-AIR THE SYSTEM, THE SCREEN ON THE AIC FROZE. A PURGE CASSETTE EXCHANGE WAS PERFORMED, BUT THE ISSUE PERSISTED. THE AIC WAS SUBSEQUENTLY SWAPPED TO RESOLVE THE FAILURES. SUPPORT CONTINUED FOLLOWING THE EXCHANGE. HOWEVER, AFTER 11 DAYS OF SUPPORT, THE PATIENT EVENTUALLY EXPIRED. Manufacturer Narrative: THE CONSOLE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON INVESTIGATION CLOSURE, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22294185&pc=MKJ | 3004123209-2025-00169 | 2025-05-24 04:00 | Death | HEARTSINE TECHNOLOGIES LTD | 2025-06-19 04:00 | MKJ | PACKAGE,450P,PP01,EN,450-STR-US-10 | Failure to Analyze Signal | Cardiac Arrest | P160008 |  | 1 | Event Description: THE CUSTOMER CONTACTED HEARTSINE TO REPORT THAT THEIR DEVICE DID NOT ANALYZE WHEN CONNECTED TO A PATIENT. IN THIS STATE, DEFIBRILLATION THERAPY MAY BE DELAYED OR PREVENTED FROM BEING DELIVERED TO THE PATIENT IF NEEDED. THE PATIENT DID NOT SURVIVE. Manufacturer Narrative: HEARTSINE CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. HEARTSINE REQUESTS PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. THE CUSTOMER PROVIDED HEARTSINE WITH THE AVAILABLE PATIENT INFORMATION. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK. HEARTSINE HAS REQUESTED RETURN OF THE DEVICE FOR INVESTIGATION. UPON COMPLETION, THE CONCLUSIONS WILL BE SUBMITTED IN A FOLLOW-UP REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22294249&pc=NJE | 3008114965-2025-00481 | 2024-01-11 05:00 | Death | MEDOS INTERNATIONAL SARL | 2025-06-19 04:00 | NJE | CERENOVUS ENTERPRISE | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage |  |  | 1 | Event Description: THE EVENT WAS REPORTED VIA THE ICAD STUDY IN CHINA. ON (B)(6) 2024, THE 61-YEAR-OLD MALE PATIENT (SUBJECT # (B)(6) WAS ¿ADMITTED DUE TO "LEFT LIMB WEAKNESS FOR 4 MONTHS." THE PATIENT WAS DIAGNOSED WITH "CEREBROVASCULAR STENOSIS, SEQUELAE OF CEREBRAL INFARCTION." PHYSICAL EXAMINATION: CLEAR CONSCIOUSNESS, CLEAR SPEECH, APPROPRIATE RESPONSES, NORMAL LIMB MOVEMENTS, AND NORMAL MUSCULAR STRENGTH AND TONE. CRANIAL MRI (B)(6) 2023) SHOWED RIGHT MIDDLE CEREBRAL ARTERY STENOSIS. PAST MEDICAL HISTORY: CEREBRAL INFARCTION, HYPERTENSION (CURRENTLY TAKING "NIFEDIPINE CONTROLLED RELEASED TABLETS 30 MG Q.D", WITH WELL-CONTROLLED BLOOD PRESSURE). AFTER ADMISSION, ROUTINE BLOOD TESTS, D-DIMER + COAGULATION PANEL WERE BASICALLY NORMAL. THROMBOELASTOGRAPHY RESULTS: AA INHIBITION RATE 94.3% AND ADP INHIBITION RATE 62.1%; CYP2C19 INTERMEDIATE METABOLIZER, MODERATE RISK OF CLOPIDOGREL RESISTANCE (DOSE INCREASE OR SWITCHING TO STANDARD-DOSE TICAGRELOR WAS RECOMMENDED). MEDICATIONS: ASPIRIN 100 MG Q.D, CLOPIDOGREL 75 MG Q.D, ATORVASTATIN CALCIUM TABLETS 20 MG QN, NIFEDIPINE CONTROLLED RELEASED TABLETS 30 MG Q.D. THE PATIENT UNDERWENT A VASCULAR STENT PLACEMENT PROCEDURE ON (B)(6) 2024 USING AN UNKNOWN ENTERPRISE 2 VASCULAR RECONSTRUCTION DEVICE (PRODUCT CODE / LOT # UNKNOWN). POST-DILATION ANGIOGRAPHY SHOWED IMPROVEMENT IN STENOSIS. THE BALLOON WAS WITHDRAWN WHILE THE MICROGUIDEWIRE WAS RETAINED. THE SELECT PLUS STENT MICROCATHETER WAS ADVANCED ALONG THE MICROGUIDEWIRE TO THE DISTAL RIGHT MIDDLE CEREBRAL ARTERY. AFTER SATISFACTORY POSITIONING ACROSS THE STENOSIS SEGMENT, A 4.0MM*16MM ENTERPRISE 2 HEADLESS STENT WAS RELEASED. FOLLOW-UP ANGIOGRAPHY SHOWED SIGNIFICANT IMPROVEMENT IN STENOSIS, WITH A RESIDUAL STENOSIS OF ABOUT 28%, AND THE STENT HAD COMPLETELY COVERED THE LESION. NO OBVIOUS HEMORRHAGE (POSTOPERATIVE XPER-CT IMAGES WERE NOT SAVED BY THE PHYSICIAN). INTRAOPERATIVE DIASTOLIC BLOOD PRESSURE FLUCTUATED BETWEEN 58-70 MMHG, AND THE SYSTOLIC BLOOD PRESSURE RANGED FROM 105-145 MMHG. AT 12:52 ON (B)(6) 2024, THE PATIENT RETURNED FROM THE INTERVENTIONAL OPERATING ROOM. PHYSICAL EXAMINATION: CLEAR CONSCIOUSNESS, SPEECH DEFICIENCY, BILATERAL PUPILS EQUAL IN SIZE AND ROUND, NORMAL LIMB MUSCULAR TONE AFTER SURGERY, THE PATIENT CAME BACK TO WARD AND COMPLAINED OF HEAD DISCOMFORT, SUDDEN SPEECH DIFFICULTY, LEFT UPPER LIMB MUSCLE STRENGTH GRADE I, LEFT LOWER LIMB MUSCLE STRENGTH GRADE II. RESCUE MEASURES: URGENT 64-SLICE CT SHOWED: 1. POST CEREBRAL ARTERY BALLOON DILATATION AND STENT ANGIOPLASTY. 2. MASSIVE CEREBRAL HEMORRHAGE IN THE RIGHT TEMPORAL LOBE AND RIGHT THALAMO-BASAL GANGLIA AND RUPTURE INTO THE VENTRICULAR SYSTEM. 3. SUBARACHNOID HEMORRHAGE. 4. PONTINE LACUNAR INFARCTION CONSIDERED. 5. BILATERAL ETHMOID SINUSITIS AND BILATERAL MAXILLARY SINUSITIS WERE SHOWN (DICOM). THE PATIENT WAS TRANSFERRED TO THE GENERAL ICU AT 14:31 AND WAS COMATOSE AND GIVEN BEDSIDE VENTILATOR-ASSISTED BREATHING WITH UNEQUAL PUPILS BILATERALLY, 3.5 MM ON THE LEFT SIDE AND 4 MM ON THE RIGHT SIDE, WITH LIGHT REFLEX LOST, NORMAL MUSCULAR TONE IN ALL FOUR LIMBS, UNCOOPERATIVE MUSCLE STRENGTH TESTS, AND POSITIVE BABINSKI SIGN BILATERALLY. GENERAL ICU DIAGNOSIS AND TREATMENT PLAN: TRACHEAL INTUBATION AND RESPIRATORY CIRCULATION SUPPORT THERAPY WERE TEMPORARILY GIVEN; DUE TO EXCESSIVE INTRACRANIAL HEMORRHAGE, CAUSING MASS EFFECT AND RISK OF CEREBRAL HERNIA, EVACUATION OF INTRACRANIAL HEMATOMA + DECOMPRESSIVE CRANIECTOMY WAS PLANNED TO BE PERFORMED. CURRENT SITUATION OF THE PATIENT: SEDATION, TRACHEAL INTUBATION AND OXYGEN INHALATION, BILATERAL PUPILS EQUAL SIZE, DIAMETER 4 MM, LIGHT REFLEX LOST, LIMB MUSCULAR TONE WAS NORMAL, UNCOOPERATIVE MUSCLE STRENGTH EXAMINATION, BILATERAL BABINSKI SIGN WAS POSITIVE. AT 19:26 ON (B)(6) 2024, THE PATIENT UNDERWENT EVACUATION OF INTRACRANIAL HEMATOMA, DECOMPRESSIVE CRANIECTOMY AND INTRACRANIAL PRESSURE MONITORING SENSOR PLACEMENT IN THE GENERAL INTENSIVE CARE UNIT. IN THE STATE OF TRACHEAL INTUBATION, AFTER SUCCESSFUL GENERAL ANESTHESIA, THE PATIENT WAS PLACED IN THE SUPINE POSITION WITH THE RIGHT SHOULDER ELEVATED AND THE HEAD TURNED TO THE LEFT. A HORSESHOE-SHAPED INCISION WAS MADE IN THE RIGHT FRONTOTEMPORAL REGION. THE SURGICAL FIELD WAS ROUTINELY DISINFECTED AND DRAPED. THE SCALP WAS INCISED, AND HEMOSTASIS WAS ACHIEVED USING SCALP CLIPS. THE SKIN-MUSCLE FLAP WAS REFLECTED TOWARD THE TEMPORAL BASE TO EXPOSE THE RIGHT FRONTOTEMPORAL BONE. A BURR HOLE WAS MADE, AND A BONE FLAP WAS DRILLED OPEN. DURING THE OPERATION, APPROXIMATELY 60 ML OF OLD BLOOD CLOTS WERE EXPELLED. THE CLOTS WERE FRAGMENTED AND ASPIRATED. THE LATERAL VENTRICLE WAS NOTED TO BE OPEN. DIFFUSE EXTRAVASATION BLOOD WAS OBSERVED AT THE WOUND SURFACE, WHICH WAS MANAGED WITH METICULOUS HEMOSTASIS AND REPEATED IRRIGATION. NO OBVIOUS ACTIVE HEMORRHAGE WAS NOTED WITHIN THE SURGICAL CAVITY. DUE TO INCREASED INTRACRANIAL PRESSURE OBSERVED PRE-OP AND INTRA-OP, THE BONE FLAP WAS DISCARDED. THE SURGERY WAS COMPLETED SUCCESSFULLY WITH SATISFACTORY ANESTHESIA. AFTER EMERGENCY EVACUATION OF INTRACRANIAL HEMATOMA, THE INTRACRANIAL PRESSURE CONTINUED TO INCREASE, THE BILATERAL PUPILS WERE 4 MM IN DIAMETER, WITH LIGHT REFLEX LOST, AND THERE WAS NO SPONTANEOUS RESPIRATION. TREATMENT MEASURES SUCH AS DEHYDRATION TO REDUCE INTRACRANIAL PRESSURE, SEDATION AND ANALGESIA, HORMONE ANTI-INFLAMMATION, AND KEEPING THE SUBCUTANEOUS AND INTRACRANIAL HEMATOMA CAVITY DRAINAGE TUBES UNOBSTRUCTED WERE ADMINISTERED. THE REMAINING TREATMENTS INCLUDING FLUID REPLACEMENT, PRESSOR, STOMACH PROTECTION, AND MAINTENANCE OF WATER-ELECTROLYTE ACID-BASE BALANCE WERE PROVIDED AS SYMPTOMATIC TREATMENT. THE PATIENT HAD MASSIVE HEMORRHAGE, WITHOUT SPONTANEOUS RESPIRATION CURRENTLY, WITH EXTREMELY HIGH INTRACRANIAL PRESSURE, VERY POOR PROGNOSIS AND LONG HOSPITALIZATION PERIOD. AFTER COMMUNICATING WITH THE FAMILY MEMBERS, THE FAMILY MEMBERS STRONGLY REQUESTED DISCHARGE. THE FAMILY MEMBERS WERE REPEATEDLY INFORMED THE RISK AFTER DISCHARGE AND THEY SIGNED THE NOTICE FOR AUTOMATIC DISCHARGE, AND THEN THE DISCHARGE PROCEDURES WERE HANDLED. A SAFETY INCIDENT ASSESSMENT AND A ONE-MONTH FOLLOW-UP WERE CONDUCTED BY THE INVESTIGATOR ON THE SUBJECT ON (B)(6) 2024. IT WAS INFORMED THAT THE SUBJECT DIED AFTER BEING DISCHARGED AND RETURNING HOME ON (B)(6) 2024. THE PATIENT DIED ON: (B)(6) 2024 (DAY 2 POST SURGERY).¿ ADDITIONAL INFORMATION WAS RECEIVED ON 11-JAN-2024. PER THE ADDITIONAL INFORMATION, THE PATIENT EXPERIENCED A CEREBRAL HEMORRHAGE ON (B)(6) 2024, WHICH BECAME KNOWN TO THE SITE AND SPONSOR ON (B)(6) 2024. THE PRINCIPAL INVESTIGATOR (PI) ASSESSED THE EVENT AS A SERIOUS, SEVERE ADVERSE EVENT, THAT WAS POSSIBLY UNRELATED TO THE STUDY DEVICE AND POSSIBLY UNRELATED TO THE PRIMARY SURGICAL PROCEDURE. THIS EVENT REQUIRED IN PATIENT HOSPITALIZATION AND A SURGICAL INTERVENTION. THE EVENT WAS CLASSIFIED AS A SYMPTOMATIC CEREBRAL HEMORRHAGE. THE OUTCOME IS RECORDED AS ¿SYMPTOM IS ONGOING,¿ WITH NO END DATE LISTED. THERE WAS NO DEVICE DEFICIENCY. ADDITIONAL / MODIFIED INFORMATION WAS RECEIVED ON 11-JAN-2024 RELATED TO THE CEREBRAL HEMORRHAGE EVENT. THE ANSWER FIELD FOR WHETHER THE EVENT REQUIRED INPATIENT HOSPITALIZATION, WAS UPDATED FROM ¿YES¿ TO ¿NO.¿ THE OUTCOME WAS UPDATED FROM ¿SYMPTOM IS ONGOING¿ TO ¿DEATH¿ WITH AN END DATE OF ¿(B)(6) 2024.¿ Manufacturer Narrative: MANUFACTURER¿S REF. (B)(4). INFORMATION REGARDING PATIENT DATE OF BIRTH, GENDER, WEIGHT, RACE, AND ETHNICITY WERE NOT PROVIDED. SECTION D.4: THE PRODUCT CATALOG AND LOT NUMBERS ARE NOT AVAILABLE / NOT REPORTED. THE FULL UDI IS NOT AVAILABLE WITHOUT THE EXPIRATION DATE. SECTION E.1: THE INITIAL REPORTER PHONE AND EMAIL ADDRESS ARE NOT AVAILABLE / REPORTED. SECTION H.4: THE DEVICE MANUFACTURE DATE IS NOT KNOWN AS THE DEVICE LOT NUMBER IS NOT AVAILABLE / NOT REPORTED. BASED ON COMPLAINT INFORMATION, THE DEVICE REMAINS IMPLANTED AND IS THUS NOT AVAILABLE FOR EVALUATION. THE DEVICE LOT NUMBER WAS NOT AVAILABLE. THE MANUFACTURING DOCUMENTATION REVIEW COULD NOT BE PERFORMED WITHOUT THE LOT NUMBER. THERE WAS NO DEVICE PERFORMANCE ISSUE OR DEVICE MALFUNCTION REPORTED DURING THE PROCEDURE. PER THE ADDITIONAL INFORMATION RECEIVED ON 11-JAN-2024, IT WAS REPORTED THAT THE EVENT OF ¿CEREBRAL HAEMORRHAGE¿ WHICH RESULTED IN DEATH, WAS POSSIBLY UNRELATED TO THE ENTERPRISE2 STUDY DEVICE AND POSSIBLY UNRELATED TO THE PRIMARY PROCEDURE. THE INVESTIGATOR DETERMINED THAT THE CAUSE OF THE PATIENT'S BLEEDING WAS UNKNOWN BUT WAS CONSIDERED TO BE CAUSED BY A REPERFUSION INJURY. ADDITIONALLY, THE PATIENT¿S MEDICAL HISTORY OF HYPERTENSION, PREVIOUS CEREBRAL INFARCTION, AND ANTICOAGULATION THERAPY MAY HAVE CONTRIBUTED TO THE EVENT. HOWEVER, THE LOCATION OF THE HEMORRHAGE AND TEMPORAL GAP BETWEEN PROCEDURE TO THE EVENT, ARE BOTH RELEVANT. THEREFORE, THE CORRELATING RELATIONSHIP BETWEEN EVENT TO THE USED DEVICE/PROCEDURE AS A CONTRIBUTING FACTOR CANNOT BE RULED OUT ENTIRELY. BASED ON THIS INFORMATION, THE EVENT OF ¿CEREBRAL HAEMORRHAGE¿ AND THE OUTCOME OF ¿DEATH¿ WILL BE CONSERVATIVELY REPORTED TO THE US FDA REPORTING UNDER CRITERIA 21 CFR 803 WITH A CLASSIFICATION OF ¿DEATH.¿ THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. AS PART OF THE POST MARKET SURVEILLANCE PROGRAM, INFORMATION FROM THIS COMPLAINT IS TRENDED FOR STATISTICAL SIGNALS AND CORRECTIVE / PREVENTIVE ACTION MAY BE TRIGGERED AT A LATER TIME. SINCE THERE WAS NO EVIDENCE TO SUGGEST THE EVENT WAS RELATED TO A MANUFACTURING OR DESIGN ISSUE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS BLANK; THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION. THE MANUFACTURER WILL SUBMIT A SUPPLEMENTAL REPORT IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT. ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22294277&pc=OUT | 2029214-2025-01427 | 2022-02-13 05:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-06-19 04:00 | OUT | PIPELINE VANTAGE WITH SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Bacterial Infection; Inflammation; Paralysis; Sepsis; Movement Disorder | P100018 |  | 1 | Event Description: MEDTRONIC RECEIVED REPORT THAT ON 9-FEB-2022, A PATIENT UNDERWENT PIPELINE VANTAGE IMPLANTATION VIA RADIAL ACCESS WITH A RIST RADIAL ACCESS SHEATH CATHETER TO TREAT A LEFT VERTEBRAL ARTERY FUSIFORM ANEURYSM. THE ANEURYSM MAX DIAMETER WAS 32MM AND THE NECK WAS 10MM. THERE WERE NO REPORTED INTRA-OPERATIVE ISSUE OR DEVICE MALFUNCTION. COMPLETE ANEURYSM NECK COVERAGE WAS ACHIEVED AT THE END OF THE PROCEDURE. POST-OPERATIVELY ON (B)(6) 2022, THE PATIENT DEVELOPED PULMONARY EMBOLISM OF A SUBSEGMENTAL BRANCH INTENDED FOR THE MIDDLE LOBE THAT WAS ISOLATED, WITHOUT PULMONARY INFARCTION OPPOSITE EITHER ACUTE OR COR PULMONALE. DIAGNOSTIC LAB ON (B)(6) 2022 SHOWED THE PATIENT HAD DEVELOPED BILATERAL PNEUMOPATHY, PREDOMINANT ON THE LEFT WITH GRAM-POSITIVE COCCI IN DIPLOCOCCI AND CHAINS. THE EVENTS WERE NO CONSIDERED LIFE THREATENING BUT DID RESULT IN ADDITIONAL PERCUTANEOUS TREATMENT, PHYSICAL THERAPY, AND PROLONGED THE PATIENT'S HOSPITALIZATION. THE EVENTS WERE CONSIDERED TO BE UNRESOLVED AT THE TIME THE REPORT WAS RECEIVED. SITE ASSESSMENT DETERMINED THE EVENTS WERE NOT RELATED TO THE PIPELINE OR RIST DEVICE. THE EVENT OF PULMONARY EMBOLISM WAS PROBABLY RELATED TO THE PROCEDURE AND PNEUMOPATHY WAS NOT RELATED. MEDTRONIC STUDY SPONSOR ASSESSMENT DETERMINED THE EVENTS WERE NOT RELATED TO THE STUDY DEVICE BUT POSSIBLY RELATED TO THE PROCEDURE AND NOTED THE PULMONARY EMBOLISM COULD POSSIBLE BE RELATED TO THE RIST CATHETER OR DUAL ANTIPLATELET TREATMENT (DAPT) BUT THE EVENT OF PNEUMOPATHY WAS NOT RELATED TO EITHER DEVICE OR DAPT. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THE EVENT DID RESULT IN MEDICAL INTERVENTION NEEDED. THE PATIENT'S HOSPITAL END DATE WAS (B)(6) 2022. THE EVENT WAS RESOLVED ON (B)(6) 2022. THE MAXIMUM ANEURYSM DIAMETER WAS UPDATED TO 12MM AND THE ANEURYSM NECK SIZE WAS UPDATED TO 4MM. THERE WAS SIGNIFICANT STASIS AT THE END OF THE PROCEDURE. THE ANEURYSM OCCLUSION STATUS AT THE END OF THE PROCEDURE WAS RAYMOND <(>&<)> ROY CLASS 3. IN ADDITION, THE SITE REPORTED A POSSIBLE SECOND FLOW DIVERTER IMPLANTED AT THE INDEX PROCEDURE ON (B)(6) 2022, NOT DUE TO A DEVICE DEFICIENCY BUT RATHER DUE TO ANEURYSM MORPHOLOGY. ADDITIONAL INFORMATION RECEIVED REPORTED SITE HAD UPDATED OUTCOME TO RECOVERED/ RESOLVED WITH AE END DATE: (B)(6) 2022 FOR THE REPORTED PULMONARY EMBOLISM. SITE HAS DETERMINED PNEUMOPATHY WAS NOT A REPORTABLE EVENT AND HAVE DELETED THE EVENT. PATIENT WAS TREATED WITH A PERCUTANEOUS INTERVENTION AND PHYSICAL THERAPY FOR PULMONARY EMBOLISM. SITE HAS REPORTED THAT AT THE INDEX PROCEDURE ON (B)(6) 2022, RIST 079 RADIAL ACCESS GUIDE CATHETER (B)(6) WAS NOT SUCCESSFULLY INTRODUCED, ADJUNCT CATHETERS WERE NOT SUCCESSFULLY NAVIGATED THROUGH RIST, DID NOT PROVIDE ADEQUATE SUPPORT AND NAVIGABILITY FOR REASON "FAILURE OF CATHETERISM WITH RIST, SHIFTED TO ENVOY MPC 6F". NO DEVICE DEFICIENCY WAS SUBMITTED AND SITE HAS BEEN QUERIED TO DO SO, IF THERE WAS A DEFICIENCY. SITE HAS CONFIRMED THAT A SECOND FLOW DIVERTER (SILK VISTA) WAS IMPLANTED DUE TO ANEURYSM MORPHOLOGY. ADDITIONAL INFORMATION RECEIVED REPORTED DEVICE TWISTING RELATED TO THE RIST CATHETER. THE RIST CATHETER WAS PLACED IN THE VERTEBRAL ARTERY BUT DID NOT PROVIDE ENOUGH SUPPORT. THE PIPELINE WAS SUCCESSFULLY DEPLOYED WHEN THE SWITCH WAS MADE TO THE ENVOY MPC 6F CATHETER. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THE TWISTING ISSUE WAS CONFIRMED TO BE RELATED TO THE RIST CATHETER. THE PATIENT WAS HOSPITALIZED FROM (B)(6) 2022 TO (B)(6) 2022, AS THE PATIENT HAD A RIGHT MOTOR DEFICIT. THERE WAS CONCOMITANT/ADDITIONAL TREATMENT GIVEN FOR THE ADVERSE EVENT. THE PATIENT DIED ON (B)(6) 2022. THE EVENT WAS POSSIBLY RELATED TO THE STUDY DEVICE AND POSSIBLY RELATED TO THE STUDY PROCEDURE. THE PATIENT WAS UNDERGOING SURGERY FOR TREATMENT OF A FUSIFORM ANEURYSM WITH A MAX DIAMETER OF 12MM AND A 4MM NECK DIAMETER. THERE WAS SIGNIFICANT STASIS AT THE END OF THE PROCEDURE. THERE WAS COMPLETE NECK COVERAGE AT THE END OF THE PROCEDURE. THE ANEURYSM OCCLUSION STATUS WAS RAYMOND AND ROY CLASS 3 AT THE END OF THE PROCEDURE. ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE SITE UPDATED THEIR ASSESSMENT OF THE RIGHT MOTOR DEFICIT AND DEATH WAS NOT RELATED TO THE DEVICE OR PROCEDURE. THE SPONSOR ASSESSED THIS AS POSSIBLY RELATED TO THE DEVICE, PROCEDURE, AND DISEASE UNDER STUDY. ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE PATIENT WAS DIAGNOSED WITH PNEUMONITIS AND ONSET OF SEPSIS. IT WAS ALSO REPORTED THAT THE PATIENT HAD LEFT SIDE HEMIPLEGIA. IT WAS UNCLEAR IF THIS WAS NEW OR ASSOCIATED WITH PRE-EXISTING DYSTONIA OR OTHER CONDITION. ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE PATIENT WAS DISCHARGED FROM THE HOSPITAL (B)(6) 2022. ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE SITE HAS UPDATED PI RELATEDNESS TO NOT RELATED TO PROCEDURE OR DEVICE, HOWEVER, SPONSOR ALIGNS W/CEC ASSESSMENT AS POSSIBLE TO PROCEDURE, VANTAGE, AND NOT RELATED TO RIST OR DAPT. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE SITE ASSESSED AS POSSIBLY RELATED TO THE DEVICE PIPELINE VANTAGE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22294388&pc=QBJ | 2954323-2025-22898 | 2025-06-04 04:00 | Death | ABBOTT DIABETES CARE LTD | 2025-06-19 04:00 | QBJ | FREESTYLE LIBRE 3 | Low Readings | Fatigue; Hyperglycemia; Nausea; Vomiting | K223435 |  | 1 | Event Description: A LOW READINGS ISSUE WITH THE ABBOTT DIABETES CARE (ADC) DEVICE WAS REPORTED. A CALLER REPORTED THAT THE CUSTOMER RECEIVED SCANS THAT WERE BETWEEN "103 MG/DL AND 150 MG/DL" COMPARED TO HOW THE CUSTOMER FELT. THE CUSTOMER NOTED A HORIZONTAL TREND ARROW WHICH INDICATES GLUCOSE IS CHANGING SLOWLY (LESS THAN 1 MG/DL PER MINUTE). IT WAS FURTHER REPORTED THAT THERE WERE NO COMPARISON READING OBTAINED AS THE READINGS PROVIDED SCANS WERE ¿WITHIN THE PATIENT'S NORMAL RANGE". THE CUSTOMER EXPERIENCED SYMPTOMS DESCRIBED AS "BEING UNWELL, INCREASINGLY WEAK", NAUSEA, AND REPEATED VOMITING. THE CUSTOMER TO SEEN AT A HOSPITAL WHERE THEY WERE HOSPITALIZED, AND A LABORATORY RESULT OF 1000 MG/DL WAS OBTAINED. THE CUSTOMER WAS ADMINISTERED NOVORAPID FOR THE DIAGNOSIS OF HYPERGLYCEMIA. A POST-TREATMENT GLUCOSE RESULT OF 600 MG/DL WAS OBTAINED BUT ADDITIONAL TREATMENT WAS PROVIDED. THERE WAS NO REPORT OF DEATH OR PERMANENT INJURY ASSOCIATED WITH THIS EVENT.Event Description: A LOW READINGS ISSUE WAS REPORTED WITH THE ABBOTT DIABETES CARE (ADC) DEVICE. THE CUSTOMER'S CAREGIVER REPORTED THAT SCAN RESULTS FROM THE ADC DEVICE RANGED BETWEEN 103 MG/DL AND 150 MG/DL, WHICH APPEARED WITHIN THE CUSTOMER'S NORMAL RANGE AND SHOWED A HORIZONTAL TREND ARROW, INDICATING SLOW GLUCOSE CHANGE (LESS THAN 1 MG/DL PER MINUTE). DESPITE THESE READINGS, THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS WEAKNESS, NAUSEA, AND VOMITING. NO CAPILLARY BLOOD GLUCOSE COMPARISONS WERE MADE AT THE TIME. EMERGENCY SERVICES WERE CALLED, AND DURING TRANSPORT, A HEALTHCARE PROFESSIONAL MEASURED A BLOOD GLUCOSE LEVEL OVER 400 MG/DL. UPON HOSPITAL ADMISSION, FURTHER READINGS EXCEEDED 600 MG/DL, AND A LABORATORY RESULT CONFIRMED A GLUCOSE LEVEL OVER 1000 MG/DL. THE PATIENT WAS TREATED WITH INSULIN AND MONITORED CONTINUOUSLY. WHILE HOSPITALIZED, THE PATIENT DEVELOPED PNEUMONIA AND WAS TREATED WITH ANTIBIOTICS. NO GLUCOSE READINGS WERE TAKEN FROM AN ADC DEVICE DURING HOSPITALIZATION, AS ONLY BLOOD GLUCOSE VALUES WERE USED. THE SENSOR WAS REMOVED SHORTLY AFTER ADMISSION. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE CAREGIVER COULD NOT CONFIRM AN OFFICIAL CAUSE OF DEATH, NOTING THAT ALL DEVICES WERE TURNED OFF FOLLOWING THE ONSET OF PNEUMONIA. NO DEATH CERTIFICATE DETAILS WERE AVAILABLE AT THE TIME OF REPORTING. GIVEN THE INFORMATION PRESENTED AT THIS TIME, IT IS INCONCLUSIVE THAT THE ADC DEVICE HAS LED TO OR CONTRIBUTED TO THE REPORTED DEATH.Event Description: A LOW READINGS ISSUE WAS REPORTED WITH THE ABBOTT DIABETES CARE (ADC) DEVICE. THE CUSTOMER'S CAREGIVER REPORTED THAT SCAN RESULTS FROM THE ADC DEVICE RANGED BETWEEN 103 MG/DL AND 150 MG/DL, WHICH APPEARED WITHIN THE CUSTOMER'S NORMAL RANGE AND SHOWED A HORIZONTAL TREND ARROW, INDICATING SLOW GLUCOSE CHANGE (LESS THAN 1 MG/DL PER MINUTE). DESPITE THESE READINGS, THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS WEAKNESS, NAUSEA, AND VOMITING. NO CAPILLARY BLOOD GLUCOSE COMPARISONS WERE MADE AT THE TIME. EMERGENCY SERVICES WERE CALLED, AND DURING TRANSPORT, A HEALTHCARE PROFESSIONAL MEASURED A BLOOD GLUCOSE LEVEL OVER 400 MG/DL. UPON HOSPITAL ADMISSION, FURTHER READINGS EXCEEDED 600 MG/DL, AND A LABORATORY RESULT CONFIRMED A GLUCOSE LEVEL OVER 1000 MG/DL. THE PATIENT WAS TREATED WITH INSULIN AND MONITORED CONTINUOUSLY. WHILE HOSPITALIZED, THE PATIENT DEVELOPED PNEUMONIA AND WAS TREATED WITH ANTIBIOTICS. NO GLUCOSE READINGS WERE TAKEN FROM AN ADC DEVICE DURING HOSPITALIZATION, AS ONLY BLOOD GLUCOSE VALUES WERE USED. THE SENSOR WAS REMOVED SHORTLY AFTER ADMISSION. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE CAREGIVER COULD NOT CONFIRM AN OFFICIAL CAUSE OF DEATH, NOTING THAT ALL DEVICES WERE TURNED OFF FOLLOWING THE ONSET OF PNEUMONIA. NO DEATH CERTIFICATE DETAILS WERE AVAILABLE AT THE TIME OF REPORTING. GIVEN THE INFORMATION PRESENTED AT THIS TIME, IT IS INCONCLUSIVE THAT THE ADC DEVICE HAS LED TO OR CONTRIBUTED TO THE REPORTED DEATH. Manufacturer Narrative: SENSOR (B)(6) HAS BEEN RETURNED AND INVESTIGATED. VISUAL INSPECTION PERFORMED ON THE RETURNED SENSOR PATCH AND NO ISSUE WAS OBSERVED. WATERMARK WAS OBSERVED AT THE BASE OF THE SENSOR TAIL INDICATING SENSOR TAIL WAS PROPERLY INSERTED. DATA WAS EXTRACTED USING APPROVED SOFTWARE, AND EXTRACTION WAS SUCCESSFUL. THE RETURNED SENSOR WAS FOUND TO BE IN SENSOR STATE 7 WITH EVENT LOG 11 AND SENSOR STATE 8 WITH EVENT LOG 9 ARE AN INDICATION THAT THE SENSOR HAD TERMINATED DUE TO AN ERROR RECOGNIZED BY THE SENSOR. THIS IS A PART OF SOFTWARE DESIGN AND IS NOT AN INDICATION OF PRODUCT NON-CONFORMANCE. FUNCTIONALITY TESTING WAS PERFORMED ON THE SENSOR TO VERIFY IF SENSOR IS FUNCTIONING AS INTENDED. THE SENSOR WAS DE-CASED AND VISUAL INSPECTION WAS PERFORMED ON THE PRINTED CIRCUIT BOARD ASSEMBLY (PCBA), AND NO ISSUES WERE OBSERVED. THE SENSOR¿S BATTERY MEASURED TO BE WITHIN SPECIFIC RANGE. PERFORMED A SOURCE MEASURE UNIT (SMU) TO ENSURE THE SENSOR'S ELECTRONICS WERE FUNCTIONING CORRECTLY. POISE VOLTAGE AND SENSOR THERMISTOR TESTING WERE PERFORMED AND BOTH WERE WITHIN SPECIFICATION, INDICATING THE SENSOR WAS PROVIDING ACCURATE GLUCOSE READINGS. THE PASSING OF FUNCTIONALITY TESTING IS AN INDICATION THAT THERE WERE NO ISSUES WITH SENSOR FUNCTIONALITY AND ELECTRONICS, THEREFORE THIS ISSUE IS NOT CONFIRMED. AN EXTENDED INVESTIGATION HAS BEEN PERFORMED FOR THE REPORTED COMPLAINT AND THERE WAS NO INDICATION THAT THE PRODUCT DID NOT MEET SPECIFICATION. THE DEVICE HISTORY REVIEW (DHRS) FOR THE PRODUCT WAS REVIEWED AND THE DHRS SHOWED THE PRODUCT MET SPECIFICATIONS PRIOR TO RELEASE TO DISTRIBUTION. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED. Manufacturer Narrative: AN EXTENDED MANUFACTURING REVIEW IS PENDING AT THIS TIME. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE COMPLETED. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED. Manufacturer Narrative: THE PRODUCT HAS BEEN REQUESTED BACK FOR AN INVESTIGATION AND THE CUSTOMER AGREED TO RETURN THE DEVICE; HOWEVER, AT THIS TIME PRODUCT HAS NOT YET BEEN RECEIVED. AN INVESTIGATION HAS BEEN PERFORMED FOR THE REPORTED COMPLAINT, AND IT HAS BEEN DETERMINED THAT THERE WAS NO INDICATION THAT THE PRODUCT DID NOT MEET SPECIFICATION. THE DHRS (DEVICE HISTORY RECORD) FOR THE PRODUCT WERE REVIEWED, AND THE DHRS SHOWED THE PRODUCT MET SPECIFICATIONS PRIOR TO RELEASE TO DISTRIBUTION. THE MOST PROBABLE ROOT CAUSES ASSOCIATED WITH THIS FAILURE MODE ARE DISCONNECTED, FAULTY OR DAMAGED COMPONENTS, SOFTWARE/DATA CORRUPTION, OR MISUSE. HOWEVER, MITIGATIONS ARE IN PLACE TO REDUCE AND PREVENT SUCH ISSUES. ALL VITAL MANUFACTURING STEPS ARE VALIDATED, MONITORED, AND VERIFIED DURING MANUFACTURING TO ENSURE THE SYSTEM IS IN CONFORMANCE WITH THE VERIFIED DESIGN. ALL VITAL FUNCTIONS ARE MONITORED BY THE SYSTEM AND, WHEN NECESSARY, FUNCTION IS SUSPENDED TO SAFEGUARD AGAINST INACCURATE RESULTS. LABELING IS PROVIDED TO INSTRUCT THE USER ON THE INTENDED USE OF ALL VITAL PARTS OF THE SYSTEM TO MINIMIZE MISUSE. ALL COMPLAINTS AND COMPLAINT TRENDS ARE INVESTIGATED TO DETERMINE IF THERE IS A PRODUCT DEFECT/DEFICIENCY. IF A PRODUCT DEFECT/DEFICIENCY IS IDENTIFIED, A RISK EVALUATION IS COMPLETED AND COMPARED TO THE RISK MANAGEMENT REPORT, TO ENSURE THE RISK PROFILE HAS NOT CHANGED. ADDITIONALLY, AS A PART OF ABBOTT'S POST-MARKET SURVEILLANCE PROCESS, ALL RISK EVALUATIONS WITH ASSOCIATED COMPLAINT DATA ARE REVIEWED ANNUALLY TO DETERMINE IF THE RISK PROFILES HAVE CHANGED AS COMPARED TO THE PRODUCT RISK MANAGEMENT REPORTS. THESE MONITORING PROCESSES ENSURE THAT ALL PRODUCT RISK PROFILES REMAIN ACCEPTABLE AND HAVE A POSITIVE BENEFIT/RISK RATIO. IF PRODUCT IS RETURNED, A PHYSICAL INVESTIGATION WILL BE PERFORMED PER ADC'S ESTABLISHED PROCESSES AND PROCEDURES AND A REPORT WILL BE SUBMITTED UPON COMPLETION OF INVESTIGATION. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22294495&pc=QJP | 2029214-2025-01430 | 2025-06-17 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-06-19 04:00 | QJP | REACT CATHETER | Difficult to Remove; Difficult to Advance; Material Deformation; Adverse Event Without Identified Device or Use Problem | Nervous System Injury | K182097 |  | 1 | Event Description: MEDTRONIC RECEIVED REPORT THAT ISSUES OCCURRED WHEN USING A REACT-71 ASPIRATION CATHETER DURING A THROMBECTOMY PROCEDURE. IT WAS REPORTED THAT THE REACT-71 AND ALL ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). CATHETER F LUSHED WAS ADMINISTERED AS INDICATED IN THE IFU. THE REACT-71 COULD NOT PUSH UP WHEN TURNING AND THEN ALSO COULD NOT BE PULLED OUT. THE SURGEON PUSHED THE GAMX TO GO UPPER AND WAS ABLE TO PULL OUT THE REACT-71. WHEN REMOVED, IT WAS FOUND THAT THE DISTAL END OF THE REACT-71 WAS KINKED SO IT WAS REPLACED WITH A FIRMER CATHETER TO CONTINUE THE PROCEDURE. IT WAS NOTED THAT A REBAR MICROCATHETER WAS ALSO USED INT HE PROCEDURE THOUGH NOT DEVICE ISSUE WAS REPORTED REGARDING THE REPORT. IT WAS NOTED THAT THERE WAS NO PATIENT HARM OR INJURY ASSOCIATED WITH THE DEVICE ISSUES. HOWEVER, THE PATIENT DIED THE SAME DAY, (B)(6) 2025.Event Description: ADDITIONAL INFORMATION RECEIVED REPORTED CONDITION BEING TREATED WAS ACUTE INTRACRANIAL LARGE ARTERY OCCLUSION. AUTOPSY NOT AVAILABLE. CAUSE OF DEATH WAS NOT DETERMINED. DEATH NOT THOUGHT TO BE RELATED TO REACT-71 CATHETER. UNKNOWN IF DEATH IS RELATED TO PROCEDURE OR PATIENT BASELINE CONDITION. Manufacturer Narrative: PRODUCT ANALYSIS: AS FOUND CONDITION: THE REACT-71 CATHETER WAS RETURNED FOR ANALYSIS WITHIN A SHIPPING BOX, WITHIN A SEALED PLASTIC BIOHAZARD POUCH AND WITHIN AN OPENED REACT-71 OUTER CARTON. NO OTHER DEVICES INVOLVED IN THE EVENT WERE RETURNED. VISUAL INSPECTION/DAMAGE LOCATION DETAILS: NO DAMAGES WERE FOUND WITH THE REACT HUB. DRIED BLOOD WAS FOUND WITHIN THE HUB. THE MICRO CATHETER BODY WAS FOUND KINKED AT ~51.5CM AND ~78.5CM FROM THE PROXIMAL END; AND FOUND ACCORDIONED BETWEEN ~16.6CM AND ~12.4CM FROM THE DISTAL END. THE DISTAL TIP WAS FOUND CRUSHED. TESTING/ANALYSIS: THE REACT CATHETER TOTAL LENGTH WAS MEASURED TO BE ~141.6CM AND THE USABLE LENGTH WAS MEASURED TO BE ~134.4CM, WHICH IS WITHIN SPECIFICATION (TOTAL: 141CM ±. 3CM; USABLE: 132CM +3CM/-0CM). A 0.067¿ MANDREL WAS INSERTED INTO THE CATHETER HUB, THROUGH THE CATHETER BODY AND BECAME STUCK AT THE PROXIMAL KINKED LOCATION (~51.5CM). THE OUTER DIAMETER OF THE REACT-71 DISTAL END AT AN UNDAMAGED SECTION WAS MEASURED TO BE ~0.0798¿ WHICH IS WITHIN SPECIFICATION (SPECIFICATION: MAX OD = 0.0855¿). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID UNK-NV-REBAR (UNKNOWN); PRODUCT TYPE: ; IMPLANT DATE N/A; EXPLANT DATE N/A. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22294514&pc=CAW | 3013095415-2025-00537 | 2025-03-13 04:00 | Death | INVACARE SUZHOU | 2025-06-19 04:00 | CAW | INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | Insufficient Information | Unspecified Respiratory Problem | K020386 |  | 1 | Event Description: THE DEVICE'S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THE FOLLOWING EVENT: "ON (B)(6) 2025 [REDACTED], A HEALTH CARE SERVICE HAS BEEN INFORMED BY A CARER FROM THE 24/7 TEAM, THAT THE PATIENT WAS STRUGGLING TO GET OXYGEN THROUGH HIS TUBING ON THE CONCENTRATOR AND WAS STRUGGLING TO BREATH, SO THE PARAMEDIC HAD BEEN CALLED. THE PATIENT WAS WITH THE PARAMEDICS WHEN THE CARER CALLED IN AND THEY WERE ASSESSING HIM TO BE TAKEN TO HOSPITAL. THE CARER ADVISED THE PATIENT WAS GOING INTO HOSPITAL AND THEY WERE UNSURE WHEN HE WOULD BE RELEASED, AND NO ONE WOULD BE HOME TO ALLOW THE TECHNICIAN TO ATTEND. ON 30/04/2025, VIVISOL WAS INFORMED THAT THE PATIENT PASSED AWAY IN HOSPITAL. PATIENTS PRESCRIPTION ON THE CONCENTRATOR WAS 2-3LPM 20HPD. THIS CONCENTRATOR WAS INSTALLED WITH THE PATIENT ON (B)6) 2024." NO FURTHER DETAILS WERE PROVIDED. Manufacturer Narrative: H6: THIS DEVICE WAS MANUFACTURED AND PLACED INTO COMMERCIAL DISTRIBUTION BACK IN MAY 2011, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE ¿NONE.¿ THE REPORTER DID NOT PROVIDE ANY FURTHER INFORMATION ON THE CIRCUMSTANCES SURROUNDING THE INDIVIDUAL'S DEATH, NOR DID THEY PROVIDE THE DATE THAT THE PATIENT DIED. AS A RESULT, SECTION B2, DATE OF DEATH, SHALL BE LEFT BLANK. ALSO, SECTION D4 MODEL # AND D4 CATALOG # ARE "UNKNOWN" AS THEY WERE NOT REPORTED BY INVACARE. THE PREVIOUS DEVICE MANUFACTURER, INVACARE, IS INVESTIGATING THE REPORTED ISSUE AND HAS PROVIDED THEIR PRELIMINARY FINDINGS: "WE WOULD LIKE TO POINT OUT THAT "THE INVACARE PERFECTO2 CONCENTRATOR IS USED BY PATIENTS WITH RESPIRATORY DISORDERS WHO REQUIRE SUPPLEMENTAL OXYGEN. THE DEVICE IS NOT INTENDED TO SUSTAIN OR SUPPORT LIFE. FURTHERMORE, THE FOLLOWING SAFETY-RELEVANT INFORMATION IS STATED IN THE USER MANUAL OF THE PERFECTO2: "INVACARE RECOMMENDS AN ALTERNATE SOURCE OF SUPPLEMENTAL OXYGEN IN THE EVENT OF A POWER OUTAGE, ALARM CONDITION OR MECHANICAL FAILURE. CONSULT YOUR PHYSICIAN OR EQUIPMENT PROVIDER FOR THE TYPE OF RESERVE SYSTEM REQUIRED." THE INFORMATION WE HAVE AT THE TIME OF THIS REPORT IS VERY LIMITED. WE HAVE REQUESTED FURTHER INFORMATION AND THE PRODUCT FOR OUR INVESTIGATION. WHEN THE INVESTIGATION HAS BEEN CONCLUDED A FINAL REPORT WILL BE PROVIDED." INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22263261&pc=OZP | 2032227-2025-192646 | 2025-05-18 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-18 04:00 | OZP | 670G INSULIN PUMP MMT-1780KL | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Unspecified Infection; Loss of consciousness | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. THE EVENT INVOLVED PRODUCTS UNOMEDICAL, UNK_RESERVOIR, AND MMT-1780KL. TROUBLESHOOTING WAS PERFORMED AND FOUND THAT THE CAUSE OF DEATH WAS DUE TO SEPTIC AND A HEART ATTACK. THE CUSTOMER WAS HOSPITALIZED ON (B)(6) 2025, AND THE BLOOD GLUCOSE AT THE TIME OF THE HOSPITALIZATION WAS 400 MG/DL. IT WAS UNKNOWN WHETHER THE CUSTOMER WAS USING THE INSULIN PUMP SYSTEM WITHIN 48 HOURS OF THE REPORTED EVENT AND THE AUTO MODE/SMART GUARD FEATURE AT THE TIME OF THE EVENT. THE BLOOD GLUCOSE LEVEL AT THE TIME OF DEATH WAS UNKNOWN. THE PUMP WAS NOT WORN AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR UNOMEDICAL, UNK_RESERVOIR, AND MMT-1780KL. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22264781&pc=DTQ | 9611109-2025-90183 | 2025-05-07 04:00 | Death | LIVANOVA DEUTSCHLAND | 2025-06-18 04:00 | DTQ | SCPC | Pumping Stopped | Cardiac Arrest | K020571 |  | 1 | Event Description: LIVANOVA DEUTSCHLAND RECEIVED A REPORT THAT A SCPC, USED DURING ECMO TO GIVE LIFE SUPPORT TO A PATIENT DURING RESUSCITATION AND DURING OPERATION, SUDDENLY STOPPED ROTATING, CAUSING THE PATIENT'S CARDIAC ARREST. THEN, MEDICAL TEAM GAVE CARDIOPULMONARY RESUSCITATION (CPR), DEFIBRILLATION, MASSIVE REHYDRATION AND HIGH-DOSE VASOACTIVE DRUGS TO PATIENT. IN ADDITION, MANUAL HAND-CRANKING OF THE CENTRIFUGAL PUMP FOR BLOOD SUPPLY WAS PERFORMED AND AFTER SWITCHING OFF AND BACK ON OFF FOR THREE (3) TIMES, THE CENTRIFUGAL PUMP RETURNED TO NORMAL FUNCTION AND CONTINUED TO OPERATE. THE OUTCOME COMMUNICATED INITIALLY WAS THAT THE SEVERITY OF THE INJURY WAS SERIOUS. THROUGH FOLLOW UP COMMUNICATION, LIVANOVA BECAME AWARE THAT THE PATIENT DECEASED.Event Description: SEE INITIAL REPORT. Manufacturer Narrative: H11: LIVANOVA DEUTSCHLAND MANUFACTURES THE SCPC. THE INCIDENT OCCURRED IN CHINA. THROUGH FOLLOW-UP COMMUNICATION, LIVANOVA LEARNED THE FOLLOWING ADDITIONAL INFORMATION: AFTER THE INCIDENT, A LIVANOVA AUTHORIZED ENGINEER HAS CHECKED AND REPAIRED THE DEVICE. NO ERROR MESSAGE WAS DISPLAYED DURING THE PROCEDURE. INVOLVED PUMP LOG FILES WERE NOT PROVIDED. THE SYSTEM WAS USED WITH A REVOLUTION DISPOSABLE. INVESTIGATION IS ON-GOING. IF ANY ADDITIONAL INFORMATION PERTINENT TO THE REPORTED EVENT IS RECEIVED, IT WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. Manufacturer Narrative: HERE BELOW A SUMMARY OF THE INFORMATION COLLECTED REGARDING THE REPORTED EVENT: LIVANOVA DEUTSCHLAND RECEIVED A REPORT THAT A SCPC, USED DURING ECMO TO GIVE LIFE SUPPORT TO A PATIENT DURING RESUSCITATION AND DURING OPERATION, SUDDENLY STOPPED ROTATING, CAUSING THE PATIENT'S CARDIAC ARREST. THEN, MEDICAL TEAM GAVE CARDIOPULMONARY RESUSCITATION (CPR), DEFIBRILLATION, MASSIVE REHYDRATION AND HIGH-DOSE VASOACTIVE DRUGS TO PATIENT. IN ADDITION, MANUAL HAND-CRANKING OF THE CENTRIFUGAL PUMP FOR BLOOD SUPPLY WAS PERFORMED AND AFTER SWITCHING OFF AND BACK ON FOR THREE (3) TIMES, THE CENTRIFUGAL PUMP RETURNED TO NORMAL FUNCTION AND CONTINUED TO OPERATE. THE OUTCOME COMMUNICATED INITIALLY WAS THAT THE SEVERITY OF THE INJURY WAS SERIOUS. THROUGH FOLLOW UP COMMUNICATION, LIVANOVA BECAME AWARE THAT THE PATIENT DECEASED. INITIALLY, IT WAS REPORTED THAT NO ERROR WAS DISPLAYED DURING THE PROCEDURE. THROUGH FURTHER FOLLOW UP COMMUNICATION, IT WAS FOUND OUT THAT: THE SCPC WAS NOT EQUIPPED WITH BUBBLE/PRESSURE SENSORS. THE PATIENT WAS IN ECMO SUPPORT FOR A LONG TIME, DUE TO A MYOCARDIAL INFARCTION. WHEN THE ISSUE OCCURRED, THE MEDICAL STAFF WAS HANDING OVER THE PATIENT'S CONDITION TO THE EMERGENCY NURSE. SUDDENLY, AN ALARM FROM THE DEVICE COULD BE HEARD. THE HOSPITAL MEDICAL TEAM FOUND THAT THE BLOOD FLOW RATE HAD DECREASED, AND THE PUMP WAS NOT ROTATING. THE EXACT PUMP STOP DURATION IS UNKNOWN. WHILE THE PUMP WAS STOPPED, THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION AND DEFIBRILLATION WAS PERFORMED TWICE. MEDICAL TEAM THEN INCREASED THE BLOOD FLOW BY HAND-CRANK, AND THE TARGET BLOOD FLOW COULD NOT BE REACHED ALSO IN THIS CASE. AFTER HAND-CRANKING FOR 3 MINUTES, THE MEDICAL TEAM RE-INSTALLED THE PUMP TWICE (FIRST ATTEMPT TO REINSTALL THE PUMP WAS NOT SUCCESSFUL) AND THE PUMP RETURNED TO NORMAL FUNCTION. BASED ON INFORMATION COLLECTED UP TO DATE, THE REPORTED CONDITION IS DETERMINED TO BE AS A CENTRIFUGAL PUMP THAT STOPPED RUNNING AND TARGET FLOW COULD NOT BE REACHED WITH HAND-CRANK. ANALYSIS OF THE COLLECTED INFORMATION AND A REVIEW OF THE DHR DID NOT IDENTIFY ANY DEVIATIONS OR NONCONFORMITIES RELEVANT TO THE ISSUE. NO MANUFACTURING ISSUE WAS IDENTIFIED AT THE TIME OF PRODUCT RELEASE AND THERE IS NO RECORD OF A PRODUCT DEFECT AT THE TIME OF INSTALLATION OR DURING USE OF THE DEVICE, INCLUDING IN THE INFORMATION RELATED TO THIS EVENT. THEREFORE, A MANUFACTURING ISSUE WAS EXCLUDED AS A CAUSE. BASED ON COLLECTED INFORMATION, AFTER THE PUMP STOP, THE HAND-CRANK WAS PERFORMED, AND THE TARGET FLOW COULD NOT BE ACHIEVED. AFTER 3 MINUTES OF HAND CRANKING AND ATTEMPTS OF THE MEDICAL STAFF TO REPOSITION THE DISPOSABLE PUMP ON THE DRIVE UNIT, THE PUMP NORMAL FUNCTION COULD BE EVENTUALLY RESTORED. THIS IS IN LINE WITH THE FINDINGS OF THE ON-SITE TECHNICAL INSPECTION AND THE FUNCTIONAL TESTS THAT COULD NEITHER CONFIRM NOR REPRODUCE ANY MALFUNCTION OF SCPC EQUIPMENT. ALL THESE ELEMENTS SUGGEST THAT THE EXPERIENCED PROBLEM WAS LIKELY INDEPENDENT FROM THE DRIVE UNIT ITSELF. THEREFORE, LIVANOVA IS INVESTIGATING EXTERNAL FACTORS AS POTENTIAL CAUSES OF THE EVENT (I.E. PRESENCE OF KINK OR CLOTTING IN THE CIRCUITS OR IN THE CANNULA THAT MAY HAVE LED TO NO FLOW WITH CONSEQUENT ALARM AND PUMP STOP THAT WAS SUBSEQUENTLY RESOLVED BY CIRCUIT/TUBING/PUMP MANIPULATION DURING TROUBLESHOOTING). FURTHER FOLLOW UP COMMUNICATION WITH THE CUSTOMER IS ONGOING TO CLARIFY THE EXACT SEQUENCE OF EVENTS, EXTERNAL FACTORS THAT MAY HAVE LED TO THE REPORTED EVENT AND THE ALARM CODE THAT WAS DISPLAYED BY THE DEVICE WHEN THE ISSUE OCCURRED. BASED ON INFORMATION COLLECTED UP TO DATE AND RESULTS OF FUNCTIONAL TESTS CONDUCTED ON THE DEVICE, LIVANOVA HAS NOT IDENTIFIED ANY FAILURE OR MALFUNCTION OF THE SCPC SYSTEM. IF ANY ADDITIONAL INFORMATION PERTINENT TO THE REPORTED EVENT IS RECEIVED, IT WILL BE PROVIDED IN A SUPPLEMENTAL REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22273925&pc=MIH | 3002808486-2025-00135 | 2025-03-03 05:00 | Death | WILLIAM COOK EUROPE | 2025-06-18 04:00 | MIH | ZENITH® TX2 DISSECTION ENDOVASCULAR GRAFT TAPERED COMPONENT | Patient Device Interaction Problem | Hemothorax |  |  | 1 | Event Description: DESCRIPTION OF EVENT ACCORDING TO STUDY (WCE-1616: ZENITH TX2 DISSECTION W/PRO-FORM AND Z-TRAK PLUS (ZDEG)): ON (B)(6) 2025, 1112 DAYS POST PROCEDURE THE SITE HAS REPORTED DEATH OF PATIENT. PER THE SITE: ¿PATIENT SUFFERING FROM A LEAK FROM AN ENDOVASCULAR AORTIC PROTHESIS CAUSING HAEMOTHROAX AND CARDIAC COMPRESSION ASSOCIATED WITH BILATERAL PULMONARY EDEMA. COLLEGIAL DECISION TO PROVIDE COMFORT CARE IN THE ABSENCE OF ANY REALISTIC THERAPEUTIC OPTION. NO OTHER INFORMATION IS AVAILABLE AND THERE ARE QUERIES PENDING REQUESTING THIS EVENT BE REPORTED AS AN ADVERSE EVENT AND FOR THE SITE TO EVALUATE RELATIONSHIP OF DEVICE. PATIENT OUTCOME: DUE TO THE LEAK FROM THE ENDOVASCULAR AORTIC PROTHESIS, THIS CAUSED HEMOTHORAX AND CARDIAC COMPRESSION ASSOCIATED WITH BILATERAL PULMONARY EDEMA. NO REALISTIC TREATMENT OPTION AND PATIENT PROVIDED COMFORT CARE UNTIL DEATH.Event Description: ADDITIONAL INFORMATION RECEIVED ON 16JUL2025: THE SITE HAS RELATED THE DEATH TO PREEXISTING CONDITION PRIOR TO PROCEDURE AND PRIMARY DISEASE PROGRESSION. THE SITE HAS STATED THE CAUSE OF DEATH: "HAEMOTHORAX AND CARDIAC COMPRESSION FOLLOWING A LEAK FROM AN ENDOVASCULAR AORTIC PROTHESIS." Manufacturer Narrative: MANUFACTURERS REF # (B)(4). G4) SIMILAR TO DEVICE MARKETED UNDER 510(K)/PMA: P180001. INVESTIGATION IS STILL IN PROGRESS. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED. Manufacturer Narrative: MANUFACTURER¿S REF# (B)(4). H10) RELATED REPORT NUMBERS: (B)(4). DUE TO ADDITIONAL INFORMATION RECEIVED ON 18JUL2025 THE EVENT IS NO LONGER CONSIDERED REPORTABLE TO FDA AS NOTHING INDICATES DEVICE DEFICIENCY OR MALFUNCTION. COOK MEDICAL WILL CONTINUE TO MONITOR FOR SIMILAR EVENTS. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22278650&pc=DTB | 2649622-2025-16451 | 2025-06-03 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-06-18 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Pacing Problem; Under-Sensing | Cardiac Arrest; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC ARREST. IT WAS ALSO REPORTED THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING WHICH TRIGGERED DEVICE CLASSIFIED FALSE TERMINATION OF ATRIAL TACHYCARDIA/ATRIAL FIBRILLATION (AT/AF) EVENTS AND ALSO RESULTED IN INAPPROPRIATE ATRIAL PACING. THE RA LEAD REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT WAS ADMINISTERED EPINEPHRINE AFTER THEY CODED. RETURN OF SPONTANEOUS CIRCULATION WAS OBSERVED. AFTER ARRIVAL AT THE HEALTHCARE FACILITY, THE PATIENT RECEIVED FIFTEEN MINUTES OF COMPRESSIONS. THE PATIENT DECIDED THEY DID NOT WANT ANY 'HEROIC' MEASURES FROM THE HEALTHCARE PROFESSIONALS. THE PATIENT IS DECEASED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22279504&pc=DWF | 2184009-2025-00775 | 2020-12-01 05:00 | Death | PERFUSION SYSTEMS | 2025-06-18 04:00 | DWF | DLP AORTIC ROOT ANTEGRADE ANTEGRADE CANNULA | Obstruction of Flow; Air/Gas in Device | Cardiac Arrest | K790565 |  | 1 | Event Description: MEDTRONIC RECEIVED INFORMATION THAT AFTER THE USE OF A DLP AORTIC ROOT CANNULA, IT WAS REPORTED THAT DURING MICS (MINIMALLY INVASIVE CARDIAC SURGERY) SURGERY FOR MITRAL INSUFFICIENCY WAS PERFORMED. MITRAL VALVULOPLASTY WAS SCHEDULED AND STARTED, HOWEVER, VALVE REPLACEMENT SURGERY WAS PERFORMED MIDWAY AND THE CARDIAC ARREST TIME WAS 4 HOURS 51 MINUTES. EXTERNAL CIRCULATION SEPARATION WAS ATTEMPTED, BUT THE HEART DID NOT MOVE AND MYOCARDIAL INFARCTION OCCURRED DURING THE PROCEDURE. THE PATIENT WAS TRANSFERRED TO THE ICU OF ANOTHER HOSPITAL WITH IABP AND/OR ECMO ATTACHED, THE PATIENT'S LIFESPAN WAS EXTENDED FOR 2 MONTHS, BUT THE PATIENT DEVELOPED DIC AND DIED OF MULTI-ORGAN FAILURE. THE PRODUCT USED FOR VALVE REPLACEMENT IS UNKNOWN. THERE WAS A REPORT STATING THAT THE PRODUCT WAS USED DURING MICS (MINIMALLY INVASIVE CARDIAC SURGERY) SURGERY, BUT NO PRODUCT MALFUNCTIONS HAVE BEEN REPORTED. THERE WERE NO REPORTS FROM MEDICAL INSTITUTIONS REGARDING THIS CASE. IN THE ATTACHED SURGICAL FIELD VIDEO, THE AIR IN THE AORTIC BASE INCREASED INSIDE THE CANNULA, AND WITHOUT PERFORMING THE AIR REMOVAL OPERATION, THE CANNULA WAS REPLACED WITH AIR. THE IMAGES OF AIR BEING PUSHED INTO THE AORTIC ROOT SEVERAL TIMES WERE CLEARLY CONFIRMED. THE MEDICAL ACCIDENT INVESTIGATION AND SUPPORT CENTER CONDUCTED AN INVESTIGATION TO FIND OUT THAT THE CAUSE OF THE ACCIDENT WAS CAUSED BY A LARGE AMOUNT OF AIR BEING PUSHED INTO THE AORTIC ROOT AS A RESULT OF FAILURE TO PERFORM THE AIR REMOVAL OPERATIONS THAT SHOULD BE CARRIED OUT BY THE PHYSICIAN. IT WAS ADJUDGED THAT AN AIR BLOCK HAD OCCURRED IN THE RIGHT CORONARY ARTERY, CAUSING MYOCARDIAL INFARCTION. DEATH OCCURRED 2 MONTHS AFTER THE PROCEDURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22282436&pc=OZD | 1220648-2025-29415 | 2025-05-19 04:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. THE PATIENT HAD A CARDIAC HISTORY AND HAD COME IN FOR THE ROSS (AORTIC ROOT REPLACEMENT) PROCEDURE AND WHEN CONDITION DETERIORATED AN EMERGENT CORONARY ARTERY BYPASS GRAFT. THE PATIENT WAS UNABLE TO COME OFF BYPASS AFTER REPORTED SIX ATTEMPTS. EXTRACORPOREAL MEMBRANE OXYGENATION WAS CONSIDERED PRIOR TO PLACING INTRA -AORTIC BALLOON PUMP (IABP). THE PATIENT FAILED IABP/INOTROPIC SUPPORT AND THE IMPELLA TEAM WAS NOTIFIED OF NEED. DURING IMPELLA SUPPORT, THE PATIENT HAD EXTENSIVE SURGICAL/CHEST BLEEDING AND PROTAMINE WAS ADMINISTERED DUE TO SURGICAL OOZING. ADDITIONALLY, IT WAS REPORTED THAT THE IMPELLA 5.5 WAS UNABLE TO FLOW OPTIMALLY DUE TO LOW VOLUME. THE TEAM ATTEMPTED TO CORRECT BLEEDING SURGICALLY BEFORE CHEST WAS CLOSED. THE PATIENT CONTINUED TO BLEED POST CLOSURE. REPLACEMENT CLOTTING FACTORS WERE ADMINISTERED AT 1.5 LITERS CELL SAVER ALONG WITH OTHER BLOOD PRODUCTS. THE TEAM DECIDED THAT SURGICAL BLEED COULD NOT BE CORRECTED AND THE DECISION WAS MADE TO TRANSFER THE PATIENT TO THE UNIT. THE PATIENT REMAINED ON SUPPORT TO ALLOW THE FAMILY TIME TO DECIDE ON WHETHER TO WITHDRAW CARE OR ESCALATE CARE THROUGH THE ADDITION OF AN IMPELLA RP FLEX. THE PATIENT WAS ESCALATED DURING CARE. AS THE PATIENT CONTINUED SUPPORT, HEMODYNAMICS WERE DECLINING DESPITE MULTIPLE PRESSORS AND BICARBONATE PUSHES. ULTIMATELY, THE PATIENT CODED AND WAS NOT ABLE TO ACHIEVE RETURN OF SPONTANEOUS CIRCULATION. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW AND ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. THE DATA LOG SHOWS SEVERAL SUCTION ALARMS THROUGHOUT THE CASE RUN, WITH A TRENDING PLACEMENT SIGNAL. THERE WAS NO MOTOR CURRENT SPIKE OR POSITIONING-RELATED ALARM NOTED. THE CAUSE OF THE LOW PUMP FLOW ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED, BASED ON DATA LOG REVIEW AND GIVEN CLINICAL DETAILS. THE CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: MEDICAL SAFETY REVIEW OF THE EVENT NOTED THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA 5.5 DEVICE USED AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME GIVEN RELATED EVENTS OF BLEEDING AND LOW PUMP FLOW. THIS IS ONE OF TWO RELATED MEDICAL DEVICE REPORTS ASSOCIATED WITH THIS PATIENTS CARE. THE ASSOCIATED REPORT WILL REFLECT THE IMPELLA RP FLEX. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22282601&pc=OZD | 1220648-2025-29417 | 2020-01-04 05:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT ENDURED HEMOLYSIS DURING IMPELLA CP SUPPORT. FOLLOWING IMPLANT, THE PATIENT BEGAN TO DEVELOP HEMATURIA. THE SUPPORT LEVEL WAS REDUCED TO P-2 AND THE DEVICE WAS REPOSITIONED TO ADDRESS THE CONDITION. THE FOLLOWING DAY, THE DECISION WAS MADE TO EXPLANT THE PUMP AS THE HEMOLYSIS HAD NOT YET RESOLVED. THE PATIENT PASSED AWAY THE NEXT DAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22283114&pc=OZD | 1220648-2025-29427 | 2025-06-06 04:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Aortic Valve Insufficiency/ Regurgitation | P140003 |  | 1 | Event Description: IT WAS REPORTED FROM INFORMATION OBTAINED FROM A LITERATURE REVIEW THAT A PATIENT HAD EXTENSIVE MYOCARDIAL ISCHEMIA ACCOMPANIED BY ST ELEVATION. ON THE FIFTH DAY OF HOSPITALIZATION, THE INTRA-AORTIC BALLOON PUMP (IABP) WAS CHANGED TO AN IMPELLA CP IN ORDER TO IMPROVE PULMONARY HYPERTENSION BY REDUCING THE LEFT VENTRICULAR LOAD. ON THE EIGHTH DAY OF HOSPITALIZATION, THE IMPELLA CP WAS REMOVED DUE TO WORSENING AORTIC REGURGITATION, AND PERIPHERAL VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) WAS CHANGED TO CENTRAL VA ECMO TO ENSURE ANTEGRADE ARTERIAL BLOOD FLOW. HOWEVER, DESPITE INTENSIVE CARE, LEFT VENTRICULAR WALL MOTION DID NOT RECOVER, AND MECHANICAL CIRCULATORY SUPPORT WITH CENTRAL VA ECMO WAS CONTINUED. ON THE 23RD DAY OF HOSPITALIZATION, THE PATIENT DEVELOPED A HEMORRHAGIC CEREBRAL INFARCTION IN THE RIGHT CEREBRAL HEMISPHERE, AND DIED ON THE 26TH DAY OF HOSPITALIZATION. THE RELATIONSHIP BETWEEN THE IMPELLA AND CAUSE OF DEATH IS UNKNOWN Manufacturer Narrative: A.1, A.4 AND A.5 ARE UNKNOWN. D.4 CATALOG NUMBER, SERIAL NUMBER, LOT NUMBER, EXPIRATION DATE AND UNIQUE IDENTIFIER (UDI) # ARE UNKNOWN. D.6A AND D.6B IMPLANT AND EXPLANT DATES ARE UNKNOWN. E.1 FIRST NAME AND LAST NAME ARE UNKNOWN. H.4 DEVICE MANUFACTURE DATE IS UNKNOWN. THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: IMPELLA CP WITH SMART ASSIST CATHETER INSERTION, SECTION: AXILLARY INSERTION OF THE IMPELLA CP WITH SMART ASSIST CATHETER, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. "UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ SECTION: IMPELLA CP WITH SMARTASSIST CATHETER INSERTION: ¿IMPELLA CATHETER PERFORMANCE WILL BE COMPROMISED IF CORRECT PLACEMENT CANNOT BE CONFIRMED. WHILE OTHER IMAGING TECHNIQUES, SUCH AS TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE), PORTABLE C-ARM FLUOROSCOPY, OR CHEST X-RAY CAN HELP CONFIRM THE POSITION OF THE IMPELLA CATHETER AFTER PLACEMENT, THESE METHODS DO NOT ALLOW VISUALIZATION OF THE ENTIRE CATHETER ASSEMBLY AND ARE INADEQUATE FOR RELIABLY PLACING THE IMPELLA CATHETER ACROSS THE AORTIC VALVE¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER, IMPELLA CATHETER INLET TO THE AORTA: ¿IF THE INLET AREA OF THE IMPELLA CATHETER IS TOO CLOSE TO OR ENTANGLED IN THE PAPILLARY MUSCLE AND/OR SUBANNULAR STRUCTURES SURROUNDING THE MITRAL VALVE, IT CAN AFFECT MITRAL VALVE FUNCTION AND NEGATIVELY IMPACT CATHETER FLOW.¿ SECTION: UNDERSTANDING AND MANAGING IMPELLA CATHETER POSITION ALARMS ¿IF THE IMPELLA CATHETER IS COMPLETELY OUT OF THE VENTRICLE, DO NOT ATTEMPT TO REPOSITION THE CATHETER ACROSS THE VALVE WITHOUT A GUIDEWIRE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22283209&pc=MKJ | 0003015876-2025-01240 | 2025-05-19 04:00 | Death | PHYSIO-CONTROL, INC. - 3015876 | 2025-06-18 04:00 | MKJ | LIFEPAK® CR2 DEFIBRILLATOR | Failure to Power Up; Detachment of Device or Device Component | No Clinical Signs, Symptoms or Conditions | P170018 |  | 1 | Event Description: A DISTRIBUTOR CONTACTED STRYKER TO REPORT THEIR DEVICE WOULD NOT POWER ON AND THAT THE MAGNET WAS MISSING IN THE DEVICE'S LID. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THE PATIENT INVOLVED IS DECEASED.Event Description: A DISTRIBUTOR CONTACTED STRYKER TO REPORT THEIR DEVICE WOULD NOT POWER ON AND THAT THE MAGNET WAS MISSING IN THE DEVICE'S LID. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THE PATIENT INVOLVED IS DECEASED. Manufacturer Narrative: THE DEVICE WAS RETURNED TO STRYKER FOR EVALUATION. THE REPORTED ISSUE WAS DUPLICATED AND VERIFIED. REVIEW OF THE SOFTWARE LOGS INDICATED THE DEVICE WENT TO "REPLACE BATTERY" STATUS ON (B)(6) 2025 AND THEN TO "SHUTDOWN" STATUS ON (B)(6) 2025. THE CAUSE OF THE REPORTED ISSUE WAS DETERMINED TO BE A MISSING MAGNET IN THE LID. THE CUSTOMER RECEIVED A REPLACEMENT DEVICE AND THIS DEVICE WAS SCRAPPED BY STRYKER. Manufacturer Narrative: STRYKER CONTINUES TO INVESTIGATE THE REPORTED FAILURE AND WILL SUBMIT A SUPPLEMENTAL REPORT ON THIS EVENT TO THE FDA AS PROVIDED BY 21 CFR 803.56. A CLINICAL REVIEW WAS PERFORMED AND THE DEVICE CONTRIBUTION IS UNKNOWN DUE TO INSUFFICIENT INFORMATION. STRYKER CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. STRYKER REQUESTS PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. NO RESPONSE HAS BEEN RECEIVED FROM THE CUSTOMER. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22283511&pc=OZO | 2032227-2025-193284 | 2023-02-13 05:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-18 04:00 | OZO | 630G INSULIN PUMP MMT-1715KR 630G | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P150001 |  | 1 | Event Description: INFORMATION RECEIVED BY MEDTRONIC INDICATED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2023, AND THE CAUSE OF DEATH WAS CARDIAC ARREST, DIABETES COMPLICATIONS. THE BLOOD GLUCOSE VALUE UNKNOWN. THE EVENT INVOLVED PRODUCTS MMT-1715KR, MMT-243A, AND MMT-332A. TROUBLESHOOTING WAS PERFORMED FOR DECEASED REPORTING. THE CUSTOMER WAS WORN THE PUMP AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-1715KR. NO PRODUCT RETURN IS REQUIRED FOR MMT-243A. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22283580&pc=OZD | 1220648-2025-29432 | 2025-05-25 04:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Ventricular Fibrillation | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT WHILE IMPLANTED WITH AN IMPELLA 5.5 THE PATIENT EXPERIENCED VENTRICULAR FIBRILLATION. THE PATIENT WAS DEFIBRILLATED OVER TEN TIMES AND AMIODARONE AND LIDOCAINE WERE ADMINISTERED. DUE TO LACK OF BRAIN FUNCTION THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22283711&pc=DSQ | 2916596-2025-03834 | 2023-01-01 05:00 | Death | THORATEC CORPORATION | 2025-06-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE LITERATURE ARTICLE ¿HEARTMATE III POST-IMPLANT HEMODYNAMICS AS EARLY PREDICTOR OF OUTCOMES¿ THAT THE HEARTMATE 3 (HM3) MAY BE ASSOCIATED WITH INCREASED HOSPITALIZATIONS, BLEEDING, AND ULTIMATELY, WORSENING HEART FAILURE LEADING TO DEATH. THE STUDY AIMED TO IDENTIFY POST-IMPLANT HEMODYNAMIC FACTORS THAT MAY SERVE AS PROGNOSTIC MARKERS OF OUTCOMES. IT WAS HYPOTHESIZED THAT PATIENTS WITH CENTRAL VENOUS PRESSURE (CVP) LESS THAN 10MMHG, PULMONARY CAPILLARY WEDGE PRESSURE (PCWP) LESS THAN 15MMHG AND INCREASED CARDIAC OUTPUT AND INDEX WILL HAVE SIGNIFICANTLY FEWER GASTROINTESTINAL (GI) BLEEDS, LOWER RATE OF PROGRESSION OF CHRONIC KIDNEY DISEASE (CKD) AND DECREASED OVERALL HOSPITALIZATIONS. A SINGLE CENTER, RETROSPECTIVE COHORT STUDY OF ABOUT 200 PATIENTS WHO RECEIVED A HM3 IMPLANT FROM 2019 THROUGH 2023 WAS CONDUCTED. HEMODYNAMIC NUMBERS WERE COLLECTED PRE-IMPLANT AND TWO MONTHS POST-IMPLANT. PRIMARY OUTCOME MEASURES WERE TOTAL HOSPITAL ADMISSIONS, ADMISSIONS DUE TO HEART FAILURE AND DEATH. SECONDARY OUTCOME MEASURES WERE DRIVELINE OR SYSTEMIC INFECTIONS, GI BLEED, STROKE, SEIZURES, THROMBOEMBOLIC EVENTS AND GLOMERULAR FILTRATION RATE (GFR) PROGRESSION. PATIENTS WERE FOLLOWED UNTIL THEY EITHER UNDERWENT HEART TRANSPLANTATION OR UNTIL DEATH, WHICHEVER CAME FIRST. CURRENTLY DATA COLLECTION REMAINED ONGOING. HOWEVER, PRELIMINARY RESULTS WITH 25% OF DATA SHOW NOTABLE TRAJECTORIES. IN PATIENTS WITH POST-IMPLANT CVP LESS THAN 6 MMHG, THE AVERAGE NUMBER OF HOSPITALIZATIONS WAS 2.7 AS COMPARED TO 6.1 IN THOSE WITH CVP BETWEEN 6 AND 12MMHG. SIMILARLY, IN THE LOW CVP GROUP, THE AVERAGE NUMBER OF GI BLEEDS WAS 0.8 WHEREAS IT WAS 1 IN THE HIGHER CVP GROUP. IN PATIENTS WITH POST-IMPLANT PCWP LESS THAN 10MMHG, THE AVERAGE HOSPITALIZATIONS WAS 4, WHEREAS FOR THOSE WITH PCWP BETWEEN 10 AND 20MMHG, THE AVERAGE WAS 7. THERE WAS A MARKED DECREASE IN GI BLEEDS INPATIENTS WITH CARDIAC INDEX (CI) GREATER THAN 2.2L/MIN/M2 (0.5VS.1.17). IN PATIENTS WHO DIED PRIOR TO OR AFTER TRANSPLANT, THE AVERAGE POST-IMPLANT CVP WAS HIGHER (9.4VS7). CURRENT DATA ANALYSIS SHOWED CLINICAL OUTCOME DIFFERENCES THAT HAVE NOT YET REACHED SIGNIFICANCE GIVEN THE LOW SAMPLE SIZE AND OUTLIERS. HOWEVER, IT DEMONSTRATED THAT CERTAIN TARGETS IN HEMODYNAMIC PROFILE MAY CHANGE THE PROGNOSIS OF COMMON LEFT VENTRICULAR ASSIST DEVICE (LVAD) COMPLICATIONS AND MORTALITY. FURTHERMORE, THIS STUDY MAY INFORM HOW HEMODYNAMIC PROFILES CAN BE OPTIMIZED TO CHANGE PATIENT OUTCOMES AFTER LVAD IMPLANT. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBERS ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS THE DATE OF DATA COLLECTION (JAN 01, 2023). AUTHOR INFORMATION: PATNAIK, S., & MOSS, N. (2025). HEARTMATE III POST-IMPLANT HEMODYNAMICS AS EARLY PREDICTOR OF OUTCOMES. JOURNAL OF CARDIAC FAILURE, 31(1), 330. DOI: 10.1016/J.CARDFAIL.2024.10.382. (B)(6). MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RESEARCH ARTICLE TITLED "HEARTMATE III POST-IMPLANT HEMODYNAMICS AS EARLY PREDICTOR OF OUTCOMES", WAS RECEIVED. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE HEARTMATE 3 DEVICE SERIAL NUMBERS AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22283768&pc=OZD | 1220648-2025-29435 | 2025-05-28 04:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT DEVELOPED BLEEDING AT THE IMPELLA CP ACCESS SITE. PRESSURE WAS HELD AT THE SITE, HEPARIN WAS STOPPED, THE DRESSING WAS CHANGED, AND QUICK CLOT WAS APPLIED TO MANAGE THE CONDITION. TWO UNITS OF BLOOD WERE GIVEN TO THE PATIENT IN RESPONSE TO THE BLEED. IT WAS ADDITIONALLY REPORTED THAT THE PATIENT'S LABS HAD HEMOLYZED. THE SUPPORT LEVEL WAS REDUCED, AND TWO UNITS OF BLOOD WERE GIVEN TO TREAT THE CONDITION. LATER SUPPORT WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS AND ACCESS SITE BLEEDING HAS BEEN COMPLETED. DATA REVIEW: THERE IS NO DATA LOG FROM 05/27 - 5/28 RETURNED FOR REVIEW. PRODUCT ANALYSIS: PUMP WAS RETURNED FOR ANALYSIS. VISUAL INSPECTION FOUND NO ISSUES ON THE PUMP. PUMP WAS CONDUCTED HEMOLYSIS TEST TO OBSERVE IF ANY DEFECT ON THE PUMP. PUMP PASSED HEMOLYSIS TEST. CONCLUSION: THE ROOT CAUSE OF HEMOLYSIS WAS UNABLE TO BE DETERMINED AS LIMITED CLINICAL DETAILS WERE PROVIDED AND NO ISSUES WERE FOUND ON THE PUMP. THE ROOT CAUSE OF ACCESS SITE BLEEDING WAS MOST LIKELY USE-RELATED SINCE ANGLE MATCHING RESOLVED THE BLEED. B2 INTERVENTION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29435. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. IMPELLA CP WITH SMART ASSIST SYSTEM INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22283787&pc=NPT | 2025587-2025-04299 | 2024-08-01 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-18 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 3 MONTHS AND 15 DAYS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO AN UNKNOWN REASON. Manufacturer Narrative: UPDATED DATA: D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22283834&pc=MCM | 9710014-2025-00539 | 2016-03-05 05:00 | Death | MED-EL ELEKTROMEDIZINISCHE GERAETE GMBH | 2025-06-18 04:00 | MCM | CONCERTO | Adverse Event Without Identified Device or Use Problem | Meningitis | P000025 |  | 1 | Event Description: MENINGITIS AND VENTRICULITIS WAS REPORTED AS COMPLICATION FROM SURGERY TO REMOVE THE ACOUSTIC TUMOR (SCHWANNOMA). Manufacturer Narrative: THE DEVICE HAS BEEN EXPLANTED AND SHOULD BE RETURNED TO THE MANUFACTURER FOR EVALUATION. WHEN AVAILABLE, A DEVICE FAILURE ANALYSIS WILL BE SUBMITTED AS A FOLLOW UP REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22283977&pc=NPT | 2025587-2025-04302 | 2025-04-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-18 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 10 MONTHS FOLLOWING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED. THE CAUSE OF DEATH WAS UNKNOWN. PER THE PHYSICIAN, THE DEATH WAS NOT RELATED TO THE VALVE OR VALVE IMPLANT PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22284604&pc=NGV | 2124215-2025-38397 | 2025-01-11 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-18 04:00 | NGV | WATCHMAN LAA CLOSURE DEVICE & DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: REPORTED VIA JOURNAL ARTICLE. IT WAS REPORTED DEATH OCCURRED. THIS STUDY IS DESIGNED TO BE A NON-RANDOMIZED, RETROSPECTIVE COHORT STUDY. ALL NONVALVULAR AF PATIENTS WHO HAD UNDERGONE LAAC PERFORMED USING A WATCHMAN OR NON-BOSTON SCIENTIFIC (BSC) DEVICE BETWEEN 2016 AND 2022 AT CHANGHAI HOSPITAL WERE INCLUDED IN OUR RETROSPECTIVE STUDY. FROM JUNE 2016 TO FEBRUARY 2022, A TOTAL OF 522 CONSECUTIVE PATIENTS WHO HAD UNDERGONE LAAC WERE ENROLLED IN OUR CURRENT STUDY. AFTER THE 1:1 PROPENSITY-SCORE MATCHING, THE CURRENT ANALYSIS INCLUDED 184 WATCHMAN AND 184 NON-BSC DEVICE PATIENTS RESULTS: PROCEDURAL RELATED EVENTS INCLUDED 1 PATIENT WHO EXPERIENCED A PERICARDIAL EFFUSION WITH CARDIAC TAMPONADE. THIS CASE WAS THE RESULT OF CORONARY SINUS PERFORATION IN THE TRANSSEPTAL PUNCTURE PROCESS. 1 PATIENT HAD A DEVICE EMBOLIZATION, AND 1 PATIENT HAD AN ESOPHAGEAL PLEURAL FISTULA. 1 PERI-PROCEDURAL DEATH OCCURRED THE PATIENT DEVELOPED SUDDEN DEATH RESULTING FROM FAST AF, INDICATING ACUTE LEFT HEART FAILURE. FOLLOW UP OUTCOMES INCLUDED 15 ALL CAUSE DEATH PATIENTS. 16 PATIENTS THAT EXPERIENCED A CEREBRAL VASCULAR ACCIDENT AND 1 PATIENT WHO EXPERIENCED A TRANSIENT ISCHEMIC ATTACK. 4 PATIENTS HAD DEVICE RELATED THROMBUS PRESENT. 16 PATIENTS HAD A PERI-DEVICE LEAK THAT WAS NOTED. 6 PATIENTS WERE REPORTED TO HAVE EXPERIENCED A MAJOR BLEEDING EVENT. Manufacturer Narrative: B2 DATE OF DEATH - ARTICLE PUBLISH DATE USED AS DEATH DATE IS UNKNOWN. B3 DATE OF EVENT - ARTICLE PUBLISH DATE USED AS EVENT DATE IS UNKNOWN. D3 MODEL NUMBER, D3 LOT NUMBER, D4 EXPIRATION DATE, D4 UNIQUE IDENTIFIER (UDI) # ARE UNKNOWN AS THE INFORMATION WAS NOT PROVIDED IN THE LITERATURE ARTICLE. D6A IMPLANT DATE WAS REPORTED AS JUNE 2016 TO FEBRUARY 2022. LITERATURE CITATION: ZHANG, S., XIONG, S., ZHANG, S., CHEN, K., WANG, H., LI, K., ... & BAI, Y. (2025). WATCHMAN VERSUS LACBES DEVICE FOR PERCUTANEOUS LEFT ATRIAL APPENDAGE CLOSURE: A SINGLE-CENTER, PROPENSITY-MATCHED STUDY. BMC CARDIOVASCULAR DISORDERS, 25(1), 1-11.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22284753&pc=DSQ | 2916596-2025-03631 | 2025-05-12 04:00 | Death | THORATEC CORPORATION | 2025-06-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Dyspnea; Hypovolemia; Hypervolemia; Heart Failure/Congestive Heart Failure; Swelling/ Edema | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THEY WERE RECENTLY ADMITTED ON (B)(6) 2025 FOR FLUID OVERLOAD AND RIGHT HEART FAILURE WHILE NOT RESPONDING TO MILRINONE. THE PUMP OPERATED AS INTENDED, WAS NOT EXPLANTED, AND WOULD NOT BE RETURNED. THE CAUSE OF DEATH WAS RIGHT VENTRICLE FAILURE UNRESPONSIVE TO MILRINONE INITIATION. THE PATIENT REQUESTED TO GO TO COMFORT CARE RATHER THAN ESCALATE FURTHER. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. RIGHT HEART FAILURE DID NOT EXIST PRE-LVAD IMPLANT, AS IT WAS DIAGNOSED ON (B)(6) 2022. A DEVICE RELATED ISSUE DID NOT CONTRIBUTE TO RIGHT HEART FAILURE. THE PATIENT HAD SHORTNESS OF BREATH, BLE SWELLING, AND ABDOMINAL SWELLING. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING RIGHT HEART FAILURE, AND DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22284930&pc=DSQ | 3007042319-2025-01295 | 2024-12-27 05:00 | Death | HEARTWARE, INC. | 2025-06-18 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP/DRIVELINE | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Intracranial Hemorrhage; Unspecified Infection; Hemorrhagic Stroke; Ischemia Stroke; Ruptured Aneurysm | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING INTRACRANIAL HEMORRHAGE (ICH) IN LEFT VENTRICULAR ASSIST DEVICE (VAD) PATIENTS. THE AUTHORS DESCRIBED PATIENT OUTCOMES WHO PRESENTED WITH ICH. ALL PATIENTS WERE ADMITTED TO INTENSIVE CARE UNITS (CARDIAC OR NEUROLOGICAL) WITH AN IMMEDIATE HOLD OF ANTICOAGULANT THERAPY AND REVERSAL OF TREATMENT PER STANDARD REVERSAL PROTOCOLS. ALL PATIENTS UNDERWENT VASCULAR IMAGING AND SERIAL HEAD IMAGING. THE ETIOLOGIES OF ICH INCLUDED SPONTANEOUS HEMORRHAGE, HEMORRHAGIC STROKE CONVERSION OF ISCHEMIC STROKE, TRAUMATIC ICH, AND RUPTURED MYCOTIC ANEURYSM. THE TYPES OF HEMORRHAGE INCLUDED INTRAPARENCHYMAL HEMORRHAGE (IPH), SUBARACHNOID HEMORRHAGE (SAH), INTRAPARENCHYMAL HEMORRHAGE (IPH), AND INTRAVENTRICULAR HEMORRHAGE (IVH). SOME PATIENTS HAD MULTICOMPARTMENTAL HEMORRHAGE, OVER FIVE MM OF MIDLINE SHIFT, UNCAL HERNIATION, TRANS TENTORIAL, OR TONSILLAR HERNIATION AND SOME REQUIRED URGENT OR EMERGENT NEUROSURGICAL INTERVENTION. DECOMPRESSIVE SURGERY VIA CRANIOTOMY OR CRANIOSTOMY, VENTRICULOSTOMY, OR DIAGNOSTIC CEREBRAL ANGIOGRAPHY. UNKNOWN SURGICAL COMPLICATIONS OCCURRED IN A FEW PATIENTS, WITH SOME THAT REQUIRED REOPERATION FOR EVACUATION. THERE WERE PATIENTS WITH AN ACUTE DRIVELINE INFECTION AT THE TIME OF ICH. THERE WERE INPATIENT DEATHS AND MORTALITY AT 90-DAY AND ONE-YEAR; HOWEVER, THE SPECIFIC CAUSES OF DEATH WERE UNKNOWN. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/53 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: INCIDENCE, PRESENTATION, AND OUTCOMES OF INTRACRANIAL HEMORRHAGE IN LEFT VENTRICULAR ASSIST DEVICE PATIENTS. JOURNAL OF NEUROSURGERY. 2025. 142:1387¿1396. DOI: 10.3171/2024.7. JNS241497. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22285317&pc=NKM | 2135147-2025-03272 |  | Death | ABBOTT MEDICAL | 2025-06-18 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100009 |  | 1 | Event Description: THE ARTICLE "SURGERY AND TRANSCATHETER INTERVENTION FOR DEGENERATIVE MITRAL REGURGITATION IN THE UNITED STATES" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE PRACTICE TRENDS AND 3-YEAR OUTCOMES OF TRANSCATHETER EDGE-TO-EDGE REPAIR (TEER) AND SURGICAL REPAIR FOR DEGENERATIVE MITRAL REGURGITATION IN THE UNITED STATES.. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT AMONG MEDICARE BENEFICIARIES WITH DEGENERATIVE MITRAL REGURGITATION, AN INCREASE IN TEER UTILIZATION WAS ASSOCIATED WITH WORSE SURVIVAL, INCREASED HEART FAILURE READMISSIONS, AND MORE MITRAL REINTERVENTIONS.. [THE PRIMARY AUTHOR AND CORRESPONDING AUTHOR WAS JOANNA CHIKWE, AT CEDARS SINAI MEDICAL CENTER, LOS ANGELES, CALIFORNIA WITH CORRESPONDING EMAIL *JOANNA.CHIKWE@CSHS.ORG*. THE TIME FRAME OF THE STUDY WAS 2012 TO 2019. A TOTAL OF 27,170 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 7,755 UNDERWENT TRANSCATHETER EDGE TO EDGE REPAIR (TEER). AVERAGE AGE WAS 81 FOR TEER AND 71 FOR SURGICAL REPAIR, MAJORITY GENDER WAS FEMALE FOR TEER AND MALE FOR SURGICAL INTERVENTION. COMORBIDITIES INCLUDED DIABETES, HYPERTENSION, HYPERLIPIDEMIA, PERIPHERAL VASCULAR DISEASE, PRIOR STROKE, CEREBROVASCULAR DISEASE, PRIOR CAROTID SURGERY, CHRONIC KIDNEY DISEASE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, LIVER DISEASE, ATRIAL FIBRILLATION, CONGESTIVE HEART FAILURE, CORONARY ARTERY DISEASE, COAGULATION DISORDER, DRUG ABUSE, SMOKING, ALCOHOL ABUSE, CANCER. PERI AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, STROKE, RE-INTERVENTION, SURGICAL INTERVENTION, HEART FAILURE HOSPITALIZATION, ATRIAL FIBRILLATION. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING DIABETES, HYPERTENSION, HYPERLIPIDEMIA, PERIPHERAL VASCULAR DISEASE, PRIOR STROKE, CEREBROVASCULAR DISEASE, PRIOR CAROTID SURGERY, CHRONIC KIDNEY DISEASE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, LIVER DISEASE, ATRIAL FIBRILLATION, CONGESTIVE HEART FAILURE, CORONARY ARTERY DISEASE, COAGULATION DISORDER, DRUG ABUSE, SMOKING, ALCOHOL ABUSE, CANCER. COMPLICATIONS REPORTED INCLUDED DEATH, STROKE, RE-INTERVENTION, SURGICAL INTERVENTION, HEART FAILURE HOSPITALIZATION, ATRIAL FIBRILLATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: AS THIS IS FROM A LITERATURE REVIEW, THE DEATH DATE IS ESTIMATED. B3: AS THIS IS FROM A LITERATURE REVIEW, THE EVENT DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22285344&pc=OZD | 1220648-2025-29459 | 2025-05-25 04:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: A COMPLAINANT REPORTED THAT THE IMPELLA CP WAS PLACED ON A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK. IT WAS REPORTED THAT DURING IMPLANTATION, THE PURGE CASSETTE FELL TO THE FLOOR AND BECAME UNSTERILE. A NEW PURGE CASSETTE WAS USED. ADDITIONALLY, THERE WAS A CALL FROM THE PHYSICIAN THAT REPORTED THE HAD A SUCTION ALARM AND PUMP WAS REPOSITIONED AND PULLED BACK A LITTLE. IN ADDITION TO PLACEMENT SIGNAL ALARM ON THE AUTOMATED IMPELLA CONTROLLER (AIC) AND CARDIOPULMONARY RESUSCITATION AFTER ELECTRICAL STORM. HOWEVER, AFTER OPTIMIZING THE POSITION, THE AIC VALUES WERE IN RANGE. IT WAS NOTED THAT CARE WAS WITHDRAWN DUE TO TOTAL HEART FAILURE AND MULTI ORGAN FAILURE. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22285434&pc=DSQ | 2916596-2025-04015 | 2025-05-27 04:00 | Death | THORATEC CORPORATION | 2025-06-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hematoma; Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO A SUBDURAL HEMATOMA WITH MIDLINE SHIFT. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: IT WAS FURTHER REPORTED THAT IT WAS UNKNOWN WHETHER THE DEATH WAS DEVICE RELATED. THERE WERE NO CHANGES TO PATIENT CONDITION OR ANTICOAGULATION STATUS THAT MAY HAVE CONTRIBUTED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22285523&pc=DTB | 2649622-2025-16528 | 2025-05-21 04:00 | Death | MPRI | 2025-06-18 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Over-Sensing | Insufficient Information | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED OVERSENSING THAT RESULTED IN FALSE ATRIAL TACHYCARDIA/ATRIAL FIBRILLATION (AT/AF) DETECTIONS. IT WAS NOTED THAT THE DEVICE WAS SET TO ONLY MONITOR AT/AF, SO NO INAPPROPRIATE THERAPY WAS DELIVERED FOR THE FALSE DETECTION. HOWEVER, IT WAS REPORTED THAT ELEVEN DAYS LATER, THE PATIENT DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22285699&pc=OZD | 1220648-2025-29467 | 2024-08-09 04:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED HEMOLYSIS WITH A "DEFINITE" RELATIONSHIP TO THE IMPELLA DEVICE USED: ¿ON (B)(6) 2024, HGB 12.5 ON ADMISSION. ON (B)(6) 2024, HGB 11.4 BEFORE CARDIAC SURGERY. ON (B)(6) 2024, HGB DROPPED TO 7.5.¿ IT WAS FURTHER REPORTED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA USED AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER. ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER. DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART. FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22285708&pc=PYX | 1220648-2025-29468 | 2025-05-26 04:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | PYX | IMPELLA RP FLEX | Adverse Event Without Identified Device or Use Problem | Ischemia | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA RP FLEX DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE LIMB HAD A FASCIOTOMY PERFORMED FOR CONCERNS OF LIMB ISCHEMIA WHICH MAY HAVE BEEN DUE TO BP ISSUES RATHER THAN TRUE PUMP CAUSED ISCHEMIA. THE FAMILY MADE THE CHOICE TO WITHDRAW CARE AFTER FOUR (4) DAYS OF SUPPORT AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA RP SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER. SECTION: CONTRAINDICATIONS (UNITED STATES). ¿THE IMPELLA RP FLEX WITH SMARTASSIST SYSTEM IS CONTRAINDICATED FOR USE WITH PATIENTS EXPERIENCING ANY OF THE FOLLOWING CONDITIONS: DISORDERS OF THE PULMONARY ARTERY WALL THAT WOULD PRECLUDE PLACEMENT OR CORRECT POSITIONING OF THE IMPELLA RP FLEX WITH SMARTASSIST DEVICE; MECHANICAL VALVES, SEVERE VALVULAR STENOSIS OR VALVULAR REGURGITATION OF THE TRICUSPID OR PULMONARY VALVE; MURAL THROMBUS OF THE RIGHT ATRIUM OR VENA CAVA; ANATOMIC CONDITIONS PRECLUDING INSERTION OF THE PUMP; PRESENCE OF A VENA CAVA FILTER OR CAVAL INTERRUPTION DEVICE, UNLESS THERE IS CLEAR ACCESS FROM THE INTERNAL JUGULAR VEIN TO THE RIGHT ATRIUM THAT IS LARGE ENOUGH TO ACCOMMODATE A 22 FR CATHETER.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22285835&pc=OUT | 3008881809-2025-00344 | 2025-05-28 04:00 | Death | STRYKER NEUROVASCULAR CORK | 2025-06-18 04:00 | OUT | UNKNOWN_NEUROVASCULAR_PRODUCT | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P170024 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER THE SUBJECT STENT IMPLANTATION PROCEDURE THE PATIENT PRESENTED A SEVERE THROMBOSIS AND DIED THE DAY AFTER PROCEDURE. NO FURTHER INFORMATION IS AVAILABLE.Event Description: IT WAS REPORTED THAT AFTER THE SUBJECT STENT IMPLANTATION PROCEDURE THE PATIENT PRESENTED A SEVERE THROMBOSIS AND DIED THE DAY AFTER PROCEDURE. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: DUE TO THE AUTOMATED MANUFACTURING EXECUTION SYSTEM (MES) THERE ARE CONTROLS IN THE MANUFACTURING PROCESS TO ENSURE THE PRODUCT MET SPECIFICATIONS UPON RELEASE. THE SUBJECT DEVICE IS NOT AVAILABLE; THEREFORE, VISUAL TESTING AS WELL AS FUNCTIONAL TESTING CANNOT BE PERFORMED. THE REPORTED EVENT IS COVERED IN THE DEVICE DIRECTIONS FOR USE (DFU). AS WELL, THE RISK OF THE REPORTED EVENT IS DOCUMENTED IN THE RISK DOCUMENTATION AND THERE ARE CURRENT CONTROLS TO MITIGATE THE RISK OF THE AS REPORTED EVENT. THE REPORTED COMPLAINT COULD NOT BE CONFIRMED, AND IT COULD NOT BE DEFINITIVELY DETERMINED IF THE DEVICE FAILED TO MEET SPECIFICATIONS BECAUSE THE PRODUCT WAS NOT RETURNED. ADDITIONAL INFORMATION PROVIDED BY THE CUSTOMER INDICATED THE EXACT DATE IS NOT AVAILABLE. THE PROCEDURE HAPPENED BETWEEN (B)(6) 2025. PATIENT DEATH WAS THE DAY AFTER THE PROCEDURE. THE PRIMARY CAUSE OF PATIENT DEATH IS NOT AVAILABLE, BUT IT WAS REPORTED THAT THE PATIENT PRESENTED A SEVERE THROMBOSIS. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WAS NO RESISTANCE ENCOUNTERED AT ANY POINT DURING THE PROCEDURE. ACCORDING TO THE DOCTOR'S REPORT, THE PROCEDURE WENT PERFECTLY UNTIL THE END AND THE PATIENT LEFT THE CATHOLOGY ROOM IN GOOD HEALTH. THERE IS NO ALLEGATION AGAINST THE SUBJECT EVOLVE DEVICE. THE PHYSICIAN DOES NOT ALLEGE ANY MALFUNCTION OR NONCONFORMANCE AGAINST THE DEVICE. THE EVOLVE DEVICE PERFORMED AS INTENDED AND THE DEVICE WAS IMPLANTED PROPERLY. NO DIFFICULTIES WERE ENCOUNTERED DURING PREPARATION/TRANSFERRING/ADVANCEMENT /WITHDRAWAL OF THE SUBJECT DEVICE, THAT COULD HAVE CONTRIBUTED TO PATIENT DEATH. BASED UPON MEDICAL REVIEW, THE HARM OBSERVED IN THIS COMPLAINT IS ANTICIPATED IN NATURE AS PER THE DEVICE RISK ASSESSMENT. AN ASSIGNABLE CAUSE OF ANTICIPATED PROCEDURAL COMPLICATION WILL BE ASSIGNED TO THE REPORTED EVENT OF 'PATIENT DEATH¿ AND ¿PATIENT VESSEL THROMBOSIS' AS A PRODUCT RELATED ROOT CAUSE DOES NOT APPLY AND THE ISSUE IS DUE TO A KNOWN PHYSIOLOGICAL EFFECT OF THE PROCEDURE AND/OR PATIENT CONDITION NOTED WITH THE DIRECTIONS FOR USE, PRODUCT LABELING AND/OR RISK DOCUMENTATION FILES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22286021&pc=DSQ | 2916596-2025-03792 | 2025-05-17 04:00 | Death | THORATEC CORPORATION | 2025-06-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure; Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS DETERMINED TO BE RIGHT VENTRICULAR FAILURE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS RIGHT HEART FAILURE AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22286150&pc=NPT | 2015691-2025-04975 | 2025-03-06 05:00 | Death | EDWARDS LIFESCIENCES | 2025-06-18 04:00 | NPT | US SAPIEN 3 ULTRA VALVE 23MM | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 |  | 1 | Event Description: AS REPORTED FROM PATIENT SUPPORT, A CALL WAS RECEIVED FROM A CARE ADVOCATE WHO REPORTED A PATIENT UNDERWENT A TRANSFEMORAL TAVR PROCEDURE WITH A 23MM SAPIEN 3 ULTRA VALVE IN AORTIC POSITION AND PASSED AWAY "9 DAYS POST-TAVR." THE CARE ADVOCATE STATED THAT THE PATIENT "DIED AS A RESULT OF THE TAVR VALVE." WHEN ASKED FOR MORE INFORMATION ABOUT CAUSE OF DEATH, THE ADVOCATE ADDED THAT THE PATIENT "SUFFERED A STROKE" AND THAT THE DEATH CERTIFICATE HAD "A BUNCH OF THINGS LISTED BUT THE STROKE WAS THE FIRST CAUSE LISTED."Event Description: AS REPORTED FROM PATIENT SUPPORT, A CALL WAS RECEIVED FROM A CARE ADVOCATE WHO REPORTED A PATIENT UNDERWENT A TRANSFEMORAL TAVR PROCEDURE WITH A 23MM SAPIEN 3 ULTRA VALVE IN AORTIC POSITION AND PASSED AWAY AFTER TAVR. THE CARE ADVOCATE STATED THAT THE PATIENT "DIED AS A RESULT OF THE TAVR VALVE." WHEN ASKED FOR MORE INFORMATION ABOUT CAUSE OF DEATH, THE ADVOCATE ADDED THAT THE PATIENT SUFFERED A STROKE AND THAT THE DEATH CERTIFICATE HAD "A BUNCH OF THINGS LISTED" BUT THE STROKE WAS THE FIRST CAUSE LISTED. THE ADVOCATE DID NOT ASK TO SPEAK WITH EW PHYSICIANS. PER MEDICAL RECORDS, THE PATIENT WAS NOTED TO HAVE SPEECH DIFFICULTIES, AND RIGHT FACIAL DROOP SHORTLY AFTER TAVR. A STROKE CODE WAS CALLED. CT HEAD WAS NEGATIVE FOR ACUTE HEMORRHAGE. CTA HEAD/ NECK CONFIRMED NO LARGE VESSEL OCCLUSION. TNKASE WAS NOT ADMINISTERED, GIVEN INCREASED RISK FOR BLEEDING WITH AGE. SUBSEQUENT MRI SHOWED MULTIPLE SUBACUTE INFARCTS SCATTERED THROUGHOUT THE LMCA DISTRIBUTION OF L FRONTAL, PARIETAL, POSTERIOR MARGIN L PUTAMEN. NEURO RECOMMENDED RESTARTING ELIQUIS TO ALLOW PERMISSIVE HYPERTENSION. ON POST OP DAY 5, THE PATIENT HAD A HYPOTENSIVE EPISODE WITH WORSENING RIGHT-SIDED WEAKNESS. BP IMPROVED WITH IV FLUIDS AND CT HEAD REMAINED UNCHANGED. AFTER MULTIPLE BEDSIDE DISCUSSIONS WITH THE PATIENT AND FAMILY; THE PATIENT ELECTED TO PURSUE COMFORT MEASURES AND PEACEFULLY PASSED AWAY ON POST OP DAY 9. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS B4, B5, B7, G3, G6, H2, H6 TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS, AND H11 HAVE BEEN UPDATED. THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. THE COMPLAINT FOR STROKE WAS CONFIRMED BASED ON THE PROVIDED MEDICAL RECORDS. A REVIEW OF THE DHR, LOT HISTORY AND COMPLAINT HISTORY DID NOT PROVIDE ANY INDICATION THAT A MANUFACTURING NON-CONFORMANCE CONTRIBUTED TO THE COMPLAINT EVENT. A REVIEW OF THE IFU REVEALED NO DEFICIENCIES. PER THE INSTRUCTIONS FOR USE (IFU), PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS SUCH AS TRANSIENT ISCHEMIC ATTACK (TIA) AND STROKE ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND THE USE OF THE EDWARDS TRANSCATHETER HEART VALVE (THV) DEVICES. ACCORDING TO THE LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EW, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING THV. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING THV ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING THV DEMONSTRATED THAT MOST PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO SIGNIFICANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN THV AND AVR PATIENTS. AFTER THV, THERE APPEARS TO BE A MORE CONSIDERABLE PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT THV PATIENTS WITH MULTIPLE CO-MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. PER MEDICAL RECORDS, "THE PATIENT WAS NOTED TO HAVE SPEECH DIFFICULTIES, AND RIGHT FACIAL DROOP SHORTLY AFTER TAVR. A STROKE CODE WAS CALLED. CT HEAD WAS NEGATIVE FOR ACUTE HEMORRHAGE. CTA HEAD/ NECK CONFIRMED NO LARGE VESSEL OCCLUSION. TNKASE WAS NOT ADMINISTERED, GIVEN INCREASED RISK FOR BLEEDING WITH AGE." IN THIS CASE, LIMITED INFORMATION WAS PROVIDED; THEREFORE, A DEFINITIVE ROOT CAUSE IS UNABLE TO BE DETERMINED AT THIS TIME BUT MAY HAVE BEEN DUE TO PATIENT FACTORS (FEMALE SEX, AGE >75 YEARS, MULTIPLE RELEVANT CO-MORBIDITIES). COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22286154&pc=OZD | 1220648-2025-29474 | 2025-05-21 04:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemorrhage/Bleeding; Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT HAD MULTIPLE EPISODES FOR VENTRICULAR TACHYCARDIA (VT) PRIOR AND POST IMPELLA CP INSERTION. ONCE THE CP WAS INSERTED MEAN ARTERIAL PRESSURES WERE STABLE DURING PERSISTENT VT. ANTI ARRHYTHMIC WERE ADMINISTERED TO CONTROL VT AND SUBSIDED ONCE DRUGS WERE ADMINISTERED AND CIRCULATED. IT WAS FURTHER REPORTED THAT WHEN THE PATIENT WAS TRANSFERRED TO ANOTHER FACILITY, THE PATIENT ARRIVED BLEEDING FROM THE RIGHT GROIN. FEMSTOP WAS PLACED AND HEPARIN WAS HELD. IT WAS FURTHER REPORTED THAT DURING EXPLANT OF THE IMPELLA CP THAT THE PATIENT HAD NO PEDAL PULSES IN THE RIGHT LEG. A RIGHT THROMBECTOMY AND PATCH PERFORMED SUCCESSFULLY. ULTIMATELY, THE PATIENT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMART ASSIST SYSTEM INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22286169&pc=OZD | 1220648-2025-29475 | 2025-05-28 04:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | OZD | IMPELLA CP | Positioning Failure; Restricted Flow rate | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE IMPELLA CP PUMP'S FLOWS WERE LESS THAN EXPECTED AND ITS POSITIONING WAS BELIEVED TO BE BEHIND THE PAPILLARY MUSCLE. AS THE PATIENT WAS STABILIZING, THE STAFF ATTEMPTED TO REPOSITION THE PUMP. WHEN UNSUCCESSFUL, THE DECISION WAS MADE TO REMOVE THE DEVICE AND REWIRE THE LEFT VENTRICLE FOR BETTER POSITIONING. A STEELCORE .018 WIRE WAS UTILIZED TO REIMPLANT THE PUMP, HOWEVER, UPON ATTEMPTED REMOVAL, THE WIRE BECAME STUCK IN THE IMPELLA. THE PATIENT BEGAN TO RAPIDLY DECOMPENSATE, AND THE DECISION WAS MADE TO REPLACE THE PUMP. DESPITE THE NEW PUMP IMPLANT, THE PATIENT ARRESTED MULTIPLE TIMES WITH A TEMPORARY ROSC ACHIEVED AFTER SEVERAL ROUNDS OF CPR. WHEN THE PATIENT ARRESTED AGAIN SHORTLY AFTER, THEY COULD NOT BE STABILIZED AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES AND LOW PUMP FLOW HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED FOR INVESTIGATION. THE DATA LOG SHOWS THAT THE PUMP WAS STARTED IN AUTO MODE P9 BUT WAS UNABLE TO ACHIEVE FLOW MORE THAN 3 L/MIN. THERE WAS ONE SUCTION ALARM AT THE END OF THE CASE. NO POSITIONING ISSUES OR MOTOR CURRENT SPIKES WERE REPORTED DURING THE CASE RUN. ACCORDING TO THE CLINICAL NOTES, THE DOCTOR BELIEVED THE PUMP WAS POSITIONED BEHIND THE PAPILLARY MUSCLE, AND A FAILED ATTEMPT WAS MADE TO REPOSITION IT. THE DOCTOR ULTIMATELY REMOVED THE PUMP AND REWIRED IT AGAIN WITH A STEELCORE 0.018 WIRE. THE CAUSE OF THE LOW PUMP FLOW ISSUE WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND DATA LOG WAS NOT CONCLUSIVE TO DERIVE ROOT CAUSE. THE CAUSE OF THE POSITIONING ISSUE WAS NOT DETERMINED SINCE SUFFICIENT CLINICAL INFORMATION WAS NOT PROVIDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22286197&pc=OZD | 1220648-2025-29477 | 2025-05-26 04:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT ENDURED RECURRENT EPISODES OF VENTRICULAR TACHYCARDIA (VT). ANTIARRHYTHMIC MEDICATIONS WERE INITIALLY INCREASED, AND THE PATIENT WAS SEDATED IN RESPONSE TO THE CONDITION. WHEN THE VT RETURNED, THE PATIENT WAS SHOCKED 4-5 TIMES AND RECEIVED TWO ROUNDS OF CPR UNTIL STABLE. THE PATIENT'S CONDITION CONTINUED TO WORSEN, AND THEIR FAMILY CHANGED THEIR STATUS TO DO NOT RESUSCITATE (DNR). THE PATIENT PASSED AWAY AFTER TWO DAYS OF SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22286249&pc=OZD | 1220648-2025-29478 | 2025-05-28 04:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENT WAS ON IMPELLA CP FOR SUPPORT AND DURING SUPPORT, THE PATIENT HAD A COMPUTED TOMOGRAPHY (CT) THAT SHOWED A DIFFUSE ANOXIC INJURY AND MULTIPLE FOCI IN INTRAPARENCHYMAL HEMORRHAGE. NEUROSURGERY CONSULTED AND RECOMMENDED COMFORT MEASURES. THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22286326&pc=OZD | 1220648-2025-29480 | 2025-05-26 04:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA (VT) AND WAS SHOCKED TWO TIMES, WAS GIVEN FOUR MINUTES OF CPR, AND THE PUMP REMAINED ON FOR SUPPORT. PUMP WAS STILL SUPPORTING WHILE PLANNING FOR TRANSFER TO TERTIARY CENTER FOR CARE ESCALATION. AFTER JUST OVER FOUR DAYS THE PATIENT EXPIRED WHEN CARE WAS WITHDRAWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22286405&pc=DSQ | 2916596-2025-03842 | 2025-06-02 04:00 | Death | THORATEC CORPORATION | 2025-06-18 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPSIS. THIS WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. THE DOCTOR SAID THAT THE DEVICE WAS NOT THE CAUSE OF DEATH AS THE PUMP WAS FUNCTIONING AS INTENDED AND THE PATIENT SEPSIS WAS ASSOCIATED TO MULTIORGAN FAILURE. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. THE CLINICAL SYMPTOMS/CULTURES WERE UNKNOWN. THE SOURCE OF THE INFECTION/SEPSIS WAS UNKNOWN. THIS WAS TREATED WITH ANTIBIOTICS. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22286514&pc=OZD | 1220648-2025-29485 | 2025-05-26 04:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | OZD | IMPELLA CP | Malposition of Device | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT AN IMPELLA CP IMPLANTED IN A PATIENT WAS PULLED 1 CM OUT OF THE BODY AND THE DIASTOLIC PLACEMENT SIGNAL WAS 45 INSTEAD OF 27. NO PLACEMENT SIGNAL ALARMED. THE PATIENT EXPIRED ON SUPPORT THOUGH THE MEDICAL STAFF STATED THE IMPELLA DID NOT CAUSE THE DEATH OF THE PATIENT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22286752&pc=DSQ | 2916596-2025-03936 | 2025-06-02 04:00 | Death | THORATEC CORPORATION | 2025-06-18 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Sepsis; Respiratory Acidosis | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEVERE RESPIRATORY ACIDOSIS AND SEPSIS. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22286809&pc=KRA | 3014162263-2025-00040 | 2025-04-20 04:00 | Death | BALT EXTRUSION | 2025-06-18 04:00 | KRA | MAGIC INFUSION CATHETER | Material Rupture | Thromboembolism | K202366 |  | 1 | Event Description: IT WAS REPORTED THAT: "THE PHYSICIAN PLANNED TO PERFORM AN PICA ANEURYSM FOR EMBOLIZATION.. THE PHYSICIAN GETTING ACCESS TO THE VERTEBRAL ARTERY TOOK MAGIC 1.2FM AFTER THE PREPARATION OF THE CATHETER NAVIGATED TO THE PICA. THE PHYSICIAN PREPPED THE GLUE INJECTED THROUGH THE MAGIC CATHETER. MAGIC 1.2 FM MICRO CATHETER UNEXPECTEDLY RUPTURED DURING THE INJECTION OF GLUE RESULTING IN UNINTENDED EMBOLIZATION OF NORMAL VESSELS. IT IS A CRITICAL EVENT AND THAT LED TO CASCADE OF EVENTS RESULTING IN DEATH OF THE PATIENT. THE PHYSICIAN RETURN THE PRODUCT FOR EVALUATION OF THE MAGIC CATHETER." Manufacturer Narrative: BALT USA REFERENCE # (B)(4). CONCLUSION OF INVESTIGATION PENDING ANALYSIS FROM LEGAL MANUFACTURER, BALT EXTRUSION. THIS REPORTED COMPLAINT CONCERNS A BALT EXTRUSION DEVICE. AS LEGAL MANUFACTURER, BALT EXTRUSION IS RESPONSIBLE FOR THE ALL-POST-MARKET ACTIVITIES INCLUDING THE INVESTIGATION, ROOT CAUSE, AND CUSTOMER FOLLOW-UP RELATED TO THIS COMPLAINT. BALT USA COMMERCIALIZES A SIMILAR DEVICE IN THE US MARKET, WHICH IS ALSO MANUFACTURED BY BALT EXTRUSION AND HAS INITIATED THIS COMPLAINT FOR THE SOLE PURPOSES OF EVALUATION FOR POTENTIAL REPORTABILITY UNDER UNITED STATES MDR REQUIREMENTS. THE SUBMISSION OF THIS REPORT OR RELATED INFORMATION TO THE FDA, AND ITS RELEASE BY FDA, DOES NOT REFLECT A CONCLUSION BY THE PARTY SUBMITTING THIS REPORT OF THE FDA THAT THE REPORT OR RELATED INFORMATION IS AN ADMISSION THAT THE MANUFACTURER, THEIR EMPLOYEES, OR THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTABLE EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22287138&pc=OZD | 1220648-2025-29501 | 2025-05-31 04:00 | Death | ABIOMED, INC. | 2025-06-18 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ON IMPELLA 5.5 SUPPORT HAD A HEAD COMPUTED TOMOGRAPHY (CT) SCAN, AND A LARGE STROKE WAS OBSERVED. ADDITIONALLY, ANTICOAGULANT WAS HELD DUE TO OOZING AT THE SURGICAL SITE. THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22287438&pc=DSQ | 2916596-2025-03765 | 2023-07-15 04:00 | Death | THORATEC CORPORATION | 2025-06-18 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO STROKE AND RIGHT VENTRICULAR FAILURE (RVF). WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENTS; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING STROKE, RIGHT HEART FAILURE, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, "PATIENT CARE AND MANAGEMENT" (UNDER "RIGHT HEART FAILURE"), DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22287832&pc=BZD | 2518422-2025-106732 | 2021-01-01 05:00 | Death | RESPIRONICS, INC. | 2025-06-18 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THE USER OF A DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY. THE PATIENT'S SON IN LAW REPORTED THAT THE PATIENT WAS IN THE HOSPITAL AND THEN PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22287960&pc=FKX | 0008030665-2025-01397 | 2025-05-31 04:00 | Death | ERIKA DE REYNOSA, S.A. DE C.V. | 2025-06-18 04:00 | FKX | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Heart Failure/Congestive Heart Failure | K181108 |  | 1 | Event Description: A REGISTERED NURSE REPORTED TO FRESENIUS CUSTOMER SERVICE THIS PERITONEAL DIALYSIS (PD) PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED WHILE CONNECTED TO THE CYCLER. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE, IT WAS REPORTED THIS PATIENT EXPIRED ON (B)(6) 2025 AT HOME DUE TO COMPLICATIONS FROM A PRIOR CEREBRAL VASCULAR ACCIDENT (CVA), CARDIAC INSUFFICIENCY AND MULTIMORBIDITY. IT WAS AFFIRMED THAT THE PATIENT WAS CONNECTED TO HER LIBERTY SELECT CYCLER AT THE TIME OF DEATH. THE PATIENT¿S FAMILY WAS ALERTED TO AN ALARM FROM A GLUCOSE MONITORING SYSTEM (NOT A FRESENIUS PRODUCT) THAT STATED THERE WAS NO BLOOD FLOW FOR TESTING. THE PATIENT WAS FOUND PULSELESS AND UNRESPONSIVE WHEN EMERGENCY MEDICAL SERVICES (EMS) WERE ACTIVATED. THE PATIENT WAS GIVEN CARDIOPULMONARY RESUSCITATION AND ADVANCED CARDIAC LIFE SUPPORT BY EMS PERSONNEL; HOWEVER, THE PATIENT SUCCUMBED TO HER ILLNESS, AND SHE WAS PRONOUNCED DECEASED IN THE FIELD WITH NO TRANSPORT TO HOSPITAL. IT WAS CONFIRMED THAT THE PATIENT¿S PREVIOUS CVA AND CARDIAC INSUFFICIENCY WERE UNRELATED TO PD THERAPY. ADDITIONALLY, IT WAS CONFIRMED THE PATIENT¿S MULTIMORBIDITY AND HER SUBSEQUENT DEATH WERE NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S).Event Description: A REGISTERED NURSE REPORTED TO FRESENIUS CUSTOMER SERVICE THIS PERITONEAL DIALYSIS (PD) PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED WHILE CONNECTED TO THE CYCLER. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE, IT WAS REPORTED THIS PATIENT EXPIRED ON (B)(6) 2025 AT HOME DUE TO COMPLICATIONS FROM A PRIOR CEREBRAL VASCULAR ACCIDENT (CVA), CARDIAC INSUFFICIENCY AND MULTIMORBIDITY. IT WAS AFFIRMED THAT THE PATIENT WAS CONNECTED TO HER LIBERTY SELECT CYCLER AT THE TIME OF DEATH. THE PATIENT¿S FAMILY WAS ALERTED TO AN ALARM FROM A GLUCOSE MONITORING SYSTEM (NOT A FRESENIUS PRODUCT) THAT STATED THERE WAS NO BLOOD FLOW FOR TESTING. THE PATIENT WAS FOUND PULSELESS AND UNRESPONSIVE WHEN EMERGENCY MEDICAL SERVICES (EMS) WERE ACTIVATED. THE PATIENT WAS GIVEN CARDIOPULMONARY RESUSCITATION AND ADVANCED CARDIAC LIFE SUPPORT BY EMS PERSONNEL; HOWEVER, THE PATIENT SUCCUMBED TO HER ILLNESS, AND SHE WAS PRONOUNCED DECEASED IN THE FIELD WITH NO TRANSPORT TO HOSPITAL. IT WAS CONFIRMED THAT THE PATIENT¿S PREVIOUS CVA AND CARDIAC INSUFFICIENCY WERE UNRELATED TO PD THERAPY. ADDITIONALLY, IT WAS CONFIRMED THE PATIENT¿S MULTIMORBIDITY AND HER SUBSEQUENT DEATH WERE NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). Manufacturer Narrative: ADDITIONAL INFORMATION: D.10., H.3. PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL ANALYSIS. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NOT BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. VALVE ACTUATION TEST PASSED. SYSTEM AIR LEAK TEST PASSED. MUSHROOM HEAD CHECK PASSED. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S). UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED. Manufacturer Narrative: CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. THE CAUSE OF THIS PATIENT¿S DEATH CAN BE ATTRIBUTED TO COMPLICATIONS FROM A PRIOR CEREBRAL VASCULAR ACCIDENT (CVA), CARDIAC INSUFFICIENCY AND MULTIMORBIDITY AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION, PARTICULARLY IN THE ENVIRONMENT OF CARDIOPULMONARY DISEASE. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22288673&pc=DTB | 2649622-2025-16565 | 2025-05-25 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-06-18 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Pacing Problem; Under-Sensing | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THERE WAS UNDERSENSING ON THE RIGHT VENTRICULAR (RV) LEAD DURING POSSIBLE VENTRICULAR ARRYTHMIA. THERE WAS ALSO A RV LEAD INTEGRITY ALERT (LIA) FOR HIGH RATE NON-SUSTAINED EPISODES AND HIGH SENSING INTEGRITY COUNTER (SIC). IT WAS ALSO NOTED THAT THERE WAS UNDERSENSING ON THE RIGHT ATRIAL (RA) LEAD THAT APPEARED TO RESULT IN INAPPROPRIATE ATRIAL PACING. THE RV LEAD AND RA LEAD REMAIN IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT APPROXIMATELY FIVE DAYS LATER THE PATIENT EXPIRED. ADDITIONAL INFORMATION WAS REQUESTED BUT YIELD NO RESULTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: DTPA2D1 CRT-D, IMPLANTED: (B)(6) 2024. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22216822&pc=EOQ | 2955842-2025-25581 | 2025-05-21 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-06-17 04:00 | EOQ | ION | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Bradycardia; Hemorrhage/Bleeding | K182188 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE A PATIENT EXPIRED. THE BIOPSY WAS PERFORMED ON A METASTATIC MASS IN THE LEFT LUNG (PROBABLY LEFT UPPER LOBE). THE PROCEDURE WAS COMPLETED WITH NO MALFUNCTIONS OR ISSUES; HOWEVER, AT THE END OF THE PROCEDURE, WHEN THE PHYSICIAN WAS UNDOCKING THE ROBOT, BLEEDING WAS NOTICED FROM THE LUNG. THE PATIENT¿S HEART RATE THEN DROPPED TO 38BPM, ¿AND ESSENTIALLY THE PATIENT CRASHED AND SUBSEQUENTLY PASSED AWAY.¿ NO FURTHER DETAILS WERE PROVIDED. THE PHYSICIAN SPECULATED THAT THE PATIENT MAY HAVE EXPERIENCED A PULMONARY EMBOLUS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION WERE MADE; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. Manufacturer Narrative: A SYSTEM LOG REVIEW CANNOT BE PERFORMED BECAUSE THE SYSTEM LOGS ARE NOT AVAILABLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT UNDERWENT AN ION LUNG BIOPSY FOR METASTATIC DISEASE TO THE LUNG. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE. AS THE ION SYSTEM WAS UNDOCKED BLEEDING WAS NOTED ASSOCIATED WITH BRADYCARDIA INTO THE 30S. THE PATIENT DETERIORATED AND SUBSEQUENTLY DIED. THERE WAS NO MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. BASED ON THE AVAILABLE DATA THE EVENTS WERE LIKELY PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50 ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF ANY BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF MORE SEVERE CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22217312&pc=BZD | 2518422-2025-047472 | 2023-03-15 04:00 | Death | RESPIRONICS, INC. | 2025-06-17 04:00 | BZD | DREAMSTATION AUTO BIPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO BIPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22225778&pc=OZD | 1220648-2025-29359 | 2025-05-31 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP PRE PERCUTANEOUS CORONARY INTERVENTION FOR MECHANICAL CIRCULATORY SUPPORT AND DEVELOPED AN ACCESS SITE HEMATOMA AND SUBSEQUENTLY EXPIRED. THE ACCESS SITE WAS THE RIGHT FEMORAL ARTERY USING MICRO PUNCTURE AND ULTRASOUND FOR ACCESS. THE IMPELLA CP WAS IMPLANTED. AFTER HEPARIN, ACTIVATED CLOTTING TIME (ACT) WAS 290 SECONDS. THE PATIENT HAD 99% OCCLUDED LEFT MAIN AND OSTIAL LEFT ANTERIOR DESCENDING ARTERIES THAT WERE DIRECT STENTED. UPON COMPLETION, THE 14 FRENCH PEEL AWAY WAS REMOVED AND THE DEVICE WAS SECURED IN THE RIGHT GROIN. PRIOR TO TRANSPORT TO THE UNIT, A SMALL HEMATOMA WAS NOTED AT THE ACCESS SITE. THE PATIENT ARRIVED AT THE UNIT AND FEMSTOP WAS PLACED. HEMATOMA AND OOZING AROUND SITE WERE NOTED. THE LEG WAS COOL AND DISCOLORED BUT DOPPLER POPLITEAL AND POSTERIOR TIBIAL PULSES WERE PRESENT. FEMSTOP WAS REMOVED THE NEXT MORNING AND OOZING HAD STOPPED WITH THE LEG LOOKING PINKER AND WARMER. FOLEY PATENT WITH YELLOW URINE. ACT WAS 129 SECONDS. THERE WAS NO ANTICOAGULATION DUE TO THE HEMATOMA. HEMOGLOBIN WAS AT 7.8, HEMATOCRIT 23.2. AND PACKED RED BLOOD CELLS WERE ORDERED. THE IMPELLA CP WAS AT PERFORMANCE LEVEL P-8 FLOWING 3.6 LITERS. THIS SAME DAY PER THE FAMILY, CARE WAS WITHDRAWN AND THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: GIVEN THE IMPELLA-RELATED ACCESS SITE COMPLICATION WITH TRANSFUSION, THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR WITH THE DEMISE OUTCOME. THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER. ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22229288&pc=PYX | 1220648-2025-29367 | 2025-05-22 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | PYX | IMPELLA RP FLEX | Restricted Flow rate | Tachycardia | P170011 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT ENDURED AN EPISODE OF VENTRICULAR TACHYCARDIA DURING IMPELLA RP FLEX SUPPORT. THE PATIENT WAS CARDIOVERTED AND AMIODARONE WAS PUSHED TO TREAT THE CONDITION. IT WAS NOTED THAT THE MEASURED FLOW RATE AROUND THIS TIME WAS BELOW THE EXPECTED RANGE FOR THE SET P LEVEL. THE PATIENT DECOMPENSATED INTO LIVER FAILURE LATER THAT DAY AND PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. LOW PUMP FLOW: THE DATA LOGS SHOW ON (B)(6) A MOTOR CURRENT SPIKE WHICH CAUSED A SHIFT IN PUMP PARAMETERS: THERE WAS A SHARP DECREASE IN PURGE FLOWS WITH A RISE IN PURGE PRESSURE. THE PUMP FLOWS ALSO DECREASED AT THAT INSTANT AND THERE WAS A PLACEMENT SIGNAL SPIKE. THE FLOWS REMAIN LOWER AFTER THE MOTOR CURRENT SPIKE. PLACEMENT SIGNAL TRENDS UP FOR THE REMAINDER OF SUPPORT WHILE PUMP FLOWS TREND DOWN. THE ROOT CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY BIOMATERIAL AS EVIDENCED BY THE MOTOR CURRENT SPIKE AND ITS EFFECTS ON PUMP PARAMETERS IN THE DATA LOGS. VT: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. THROMBUS FAILURE MODE ADDED DUE TO LOG ANALYSIS REVEALING INGESTION. THROMBUS: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE THROMBUS WAS NOT DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER INSTRUCTIONS FOR USE & CLINICAL REFERENCE MANUAL SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ARRHYTHMIA, ATRIAL FIBRILLATION, BLEEDING, CARDIAC TAMPONADE, CARDIOGENIC SHOCK, DEATH, DEVICE MALFUNCTION, HEMOLYSIS, HEPATIC FAILURE, INSERTION SITE INFECTION, PERFORATION, PHLEGMASIA CERULEA DOLENS (A SEVERE FORM OF DEEP VENOUS THROMBOSIS), PULMONARY VALVE INSUFFICIENCY, RESPIRATORY DYSFUNCTION, SEPSIS, THROMBOCYTOPENIA, THROMBOTIC VASCULAR (NON-CENTRAL NERVOUS SYSTEM COMPLICATION, TRICUSPID VALVE INJURY, CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION), VENOUS THROMBOSIS, VENTRICULAR FIBRILLATION AND/OR TACHYCARDIA.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22229465&pc=NGV | 2124215-2025-38951 | 2025-05-25 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-17 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: REPORTED VIA CLINICAL STUDY. IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED, AND A 27MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS IMPLANTED ON (B)(6) 2024. THE PATIENT WAS DISCHARGED. THE PATIENT WAS CALLED FOR THE ROUTINE SIX (6) MONTH FOLLOW UP, AND THE PERSON THAT ANSWERED THE PHONE SAID THE PATIENT PASSED AWAY. A SECOND CALL WAS MADE TO TRY AND GET MORE INFORMATION DUE TO NO RECORDS IN THE CHART, BUT THERE WAS NO ANSWER AND NO VOICEMAIL. ON (B)(6) 2025 THE DEATH CERTIFICATE WAS RECEIVED, AND THE DIAGNOSIS WAS A CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD), AND THE PATIENT DATE OF DEATH WAS (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22230555&pc=HGM | 9610816-2025-000463 | 2025-05-14 04:00 | Death | PHILIPS MEDIZIN SYSTEME BÖBLINGEN GMBH | 2025-06-17 04:00 | HGM | AVALON FM50 FETAL MONITOR | Unable to Obtain Readings | Encephalopathy; Hypoxia in Utero | K140535 |  | 1 | Event Description: PHILIPS RECEIVED A COMPLAINT ON AN AVALON FM50 FETAL MONITOR INDICATING THAT THE FETAL HEART RATE (FHR) WAS NOT ABLE TO BE MEASURED AND THE BABY PASSED AWAY AFTER BIRTH. THE MOTHER AT 37 WEEKS PLUS 4 DAYS GESTATION PRESENTED WITH AN INTERMITTENT CATEGORY 2 TRACING (NEITHER REASSURING NOR NON-REASSURING TRACE) WITH IMPROVEMENT OVER 24 HOURS WITH POSITION CHANGES. THE PATIENT THEN HAD A QUICK CHANGE TO A CATEGORY 3 TRACING (PROLONGED DECELERATIONS) AND WAS TRANSFERRED EMERGENTLY TO THE OR, WHERE THE FETAL HEART RATE IMPROVED SO A VACUUM-ASSISTED DELIVERY WAS ATTEMPTED, WHICH WAS UNSUCCESSFUL, AND THE MOTHER WAS RELUCTANT TO GO TO C-SECTION. DURING THESE ATTEMPTS THE FHR WAS LOST FOR MORE THAN 8 MINUTES SO AN EMERGENT C-SECTION WAS PERFORMED. THE BABY WAS CYANOTIC AND UNRESPONSIVE AND REQUIRED RESUSCITATION. EMERGENT INTUBATION AND TRANSFER TO NICU AND ACTIVE COOLING WAS ATTEMPTED. CARE WAS SUBSEQUENTLY WITHDRAWN AND DEATH WAS DUE TO SEVERE HYPOXIC-ISCHEMIC ENCEPHALOPATHY. THE MATERNAL HEART RATE (MHR) WAS MONITORED AT THE TIME, AND IT WAS THOUGHT BECAUSE THE MHR AND FHR WERE BEING PICKED UP INTERCHANGEABLY, THERE WAS A DELAY IN CARE, WHICH AFFECTED THE OUTCOME. BASED ON THIS INFORMATION, THE CUSTOMER BELIEVES THE DEVICE MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. A PHILIPS TECHNICAL CONSULTANT (TC) WENT TO THE CUSTOMER SITE, AND WAS UNABLE TO DUPLICATE THE ISSUE THAT THE AUDIBLE HEART RATE WAS HALF THE RATE THAT SHOWED ON THE MONITOR DISPLAY. THE TC NOTED THAT THE CLINICIANS HAD DIFFICULTY DISCERNING THE MATERNAL AND FETAL HR. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22231366&pc=JOH | 8020889-2025-00294 | 2025-05-10 04:00 | Death | MALLINCKRODT MEDICAL | 2025-06-17 04:00 | JOH | SHILEY | Gas/Air Leak | Unintended Extubation; Insufficient Information | K150844 |  | 1 | Event Description: IT WAS REPORTED THAT AT A HOSPITAL, THE 8.5MM AND 7.0MM TRACHEOSTOMY TUBES HAD A LEAKING CUFF DURING THE PROCEDURE BASED ON THE VIDEO PROVIDED. THERE WAS NO REPORTED PATIENT OUTCOME.Event Description: IT WAS REPORTED THAT AT A HOSPITAL, THE 8.5MM AND 7.0MM TRACHEOSTOMY TUBES HAD A LEAKING CUFF DURING THE PROCEDURE BASED ON THE VIDEO PROVIDED. THE PATIENT WAS RE-INTUABATED TWICE, BUT EVENTUALLY DIED. Manufacturer Narrative: D10 CONCOMITANT PRODUCT: 8CN85R , 8CN85R 8.5MM ADT FLEX TRACH W TG CUFF X1, LOT# 202405168X MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22233889&pc=OZD | 1220648-2025-29377 | 2025-05-25 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 AND EXPERIENCED ATRIAL FIBRILLATION (AF). AMIODARONE WAS ADMINISTERED AND SUPPORT CONTINUED. ADDITIONALLY, THE PATIENT HAD A GASTROINTESTINAL BLEED. HEPARIN WAS WITHHELD AND TWO UNITS OF PACKED RED BLOOD CELLS WERE TRANSFUSED. THE PATIENT'S FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22233994&pc=NPT | 2025587-2025-04270 | 2025-05-22 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-17 04:00 | NPT | EVOLUT FX PLUS VALVE | Perivalvular Leak; Patient Device Interaction Problem | Cardiac Arrest; Non specific EKG/ECG Changes; Perforation; Ventricular Fibrillation; Heart Block; Valvular Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED DUE TO CALCIFICATION. FOLLOWING THE IMPLANT PROCEDURE, THE PATIENT WENT INTO AN UNSPECIFIED ATRIO-VENTRICULAR (AV) BLOCK WITH CARDIAC ARREST, AND A MEMBRANOUS SEPTUM PERFORATION WAS OBSERVED. SUBSEQUENTLY, CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED AND THE PATIENT'S HEART RHYTHM WENT BACK TO NORMAL; HOWEVER, VENTRICULAR FIBRILLATION WAS OBSERVED AND CPR WAS PERFORMED AGAIN WITH MULTIPLE DEFIBRILLATOR SHOCKS. AFTER APPROXIMATELY 30 MINUTES OF CPR, THE PATIENT DIED DUE TO HEART BLOCK/CARDIAC ARREST. PER THE PHYSICIAN, THE IMPLANT PROCEDURE AND THE PATIENT'S CALCIFIED ANATOMY CONTRIBUTED TO THE DEATH.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT COMPLETE HEART BLOCK (CHB) OCCURRED. THE MEMBRANOUS SEPTUM PERFORATION WAS OBSERVED FOLLOWING THE POST-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV). PER THE PHYSICIAN, THE CAUSE OF THE PERFORATION WAS DUE TO THE POST-IMPLANT BAV, AND THE DELIVERY CATHETER SYSTEM (DCS) AND GUIDEWIRE DID NOT CONTRIBUTE TO THE PERFORATION. ADDITIONALLY, THE VALVE COULD NOT BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE DEATH, BUT THE DCS COULD BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE DEATH.Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF A TRANSCATHETER VALVE, A PRE-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED DUE TO CALCIFICATION. FOLLOWING THE IMPLANT PROCEDURE, THE PATIENT WENT INTO AN UNSPECIFIED ATRIO-VENTRICULAR (AV) BLOCK WITH CARDIAC ARREST, AND A MEMBRANOUS SEPTUM PERFORATION WAS OBSERVED. SUBSEQUENTLY, CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED AND THE PATIENT'S HEART RHYTHM WENT BACK TO NORMAL; HOWEVER, VENTRICULAR FIBRILLATION WAS OBSERVED AND CPR WAS PERFORMED AGAIN WITH MULTIPLE DEFIBRILLATOR SHOCKS. AFTER APPROXIMATELY 30 MINUTES OF CPR, THE PATIENT DIED DUE TO HEART BLOCK/CARDIAC ARREST. PER THE PHYSICIAN, THE IMPLANT PROCEDURE AND THE PATIENT'S CALCIFIED ANATOMY CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED THAT COMPLETE HEART BLOCK (CHB) OCCURRED. THE MEMBRANOUS SEPTUM PERFORATION WAS OBSERVED FOLLOWING THE POST-IMPLANT BALLOON AORTIC VALVULOPLASTY (BAV). PER THE PHYSICIAN, THE CAUSE OF THE PERFORATION WAS DUE TO THE POST-IMPLANT BAV, AND THE DELIVERY CATHETER SYSTEM (DCS) AND GUIDEWIRE DID NOT CONTRIBUTE TO THE PERFORATION. ADDITIONALLY, THE VALVE COULD NOT BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE DEATH, BUT THE DCS COULD BE EXCLUDED AS A CONTRIBUTING FACTOR TO THE DEATH. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE POST-IMPLANT BAV WAS PERFORMED DUE TO MILD PARAVALVULAR LEAK. Manufacturer Narrative: UPDATED DATA: ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE D-EVOLUTFX-2329 IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE D-EVOLUTFX-2329 IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803 FOR THAT DEVICE. B5. D4. D10. CONTINUATION OF D10: PRODUCT ID {POST IMPLANT BAV}; PRODUCT LOT/SERIAL NUMBER {UNKNOWN}; PRODUCT TYPE: {BALLOON AORTIC VALVULOPLASTY}; IMPLANT DATE {N/A}; EXPLANT DATE {N/A}. H6. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. ADDED THIRD PARAGRAPH. H6. ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: DELIV SYS D-EVOLUTFX-2329, PRODUCT ID: D-EVOLUTFX-2329, SERIAL/LOT: UNKNOWN, USE BY DATE: UNKNOWN, UDI: UNKNOWN. THE CODES PRESENT IN SECTION H6 CORRESPOND TO MULTIPLE COMPONENTS/PRODUCTS THAT COMPRISE THIS REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22234083&pc=OZD | 1220648-2025-29378 | 2025-05-23 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemolysis | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP HAD HEMOLYZED LABS AND WAS PRODUCING RED URINE. BECAUSE THE PATIENT HAD NO NEURO FUNCTION, CARE WAS DISCONTINUED AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)." SECTION: HEMOLYSIS: "HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS." "PATIENT CONDITIONS: INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES; MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22234834&pc=OZD | 1220648-2025-29380 | 2025-05-23 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Stroke/CVA; Perforation of Vessels | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, INTERVENTION WAS COMPLETED TO CHRONIC TOTAL OCCLUSION (CTO) STENT TO LEFT ANTERIOR DESCENDING (LAD) AND DIAG. WHEN GOING TO PULL THE WIRES AFTER THE INTERVENTION THE DIAG WIRE CURLED UP AND BROKE OFF CAUSING A PERFORATION IN THE DIAG AND SHUTTING DOWN THE LAD. ECHO IN THE CARDIAC CARE LAB (CCL) SHOWING AN EFFUSION. SHORTLY AFTER THE PATIENT HAD A CARDIAC ARREST REQUIRING CPR. A PERICARDIAL DRAIN WAS PLACED. IMPELLA CP WAS PLACED EMERGENTLY DURING CPR. ONCE IT STARTED, THE PATIENT OBTAINED ROSC. BALLOON TAMPONADE WAS STILL UP ON DIAG. THE PHYSICIAN WAS CONCERNED FOR RV AND DIDN'T THINK CP WAS ENOUGH SUPPORT. PATIENT CANNULATED AND IMPELLA DROPPED TO P4. INTERVENTION TO LAD UTILIZING PENUMBRA FROM PROX TO DISTAL LAD. PTCA TO OM. BLEEDING HAD SLOWED DOWN FROM PERICARDIAL DRAIN. SWAN PLACED BUT HAD DIFFICULTY TO GET A CROSSED, SO ECMO FLOWS WERE DECREASED AND THEN IMPELLA SUCTION ALARMS PRESENT BLOOD WAS GIVEN AND THEN INCREASED FLOWS. DR STATED HE HAD CONCERNS THAT IMPELLA PERF THROUGH LEFT VENTRICLE (LV). HE SAID HE WOULD OPEN THE PATIENT IN THE ICU TO ACCESS. IT WAS NOTED THAT THE DIAG WAS STILL BLEEDING. THERE WAS NO LV PERF. IT WAS ALSO REPORTED THAT THE PATIENT HAD A STROKE. THE CARE WAS ULTIMATELY WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST SYSTEM INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿ IMPELLA CP WITH SMARTASSIST SYSTEM AND IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. Manufacturer Narrative: THE INVESTIGATION OF VASCULAR DAMAGE - PERIPHERAL VESSEL, STROKE/CVA, VENTRICLE PERFORATION, AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. VASCULAR DAMAGE - PERIPHERAL VESSEL: THE CAUSE OF THE VASCULAR DAMAGE - PERIPHERAL VESSEL WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS. VENTRICLE PERFORATION: THERE WAS NO ISSUE FOUND DURING THE CASE. CLINICAL DETAILS STATE ALTHOUGH LEFT VENTRICLE (LV) PERFORATION WAS FEARED, LV PERFORATION DID NOT OCCUR. STROKE/CVA: AS CLINICAL INFORMATION WAS NOT PROVIDED, THE TRUE ROOT CAUSE OF THE STROKE/CVA COULD NOT BE DETERMINED. LOW PUMP FLOW: DATA LOG REVIEW SHOWED SUCTION EVENTS AND LOW FLOWS FOLLOWING IMPLANT PROCEDURE UNTIL IMPELLA WAS TURNED DOWN TO P-1, THEN SUCTION RESOLVED. THE CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY PATIENT CONDITION RELATED DUE TO PT ON ECMO, SUCTION EVENTS, AND LOW VOLUME OF THE PATIENT. FLOWS WERE INCREASED TO P-2 AFTER CLINICAL DETAILS STATED BLOOD VOLUME WAS GIVEN AND THE SUCTION ALARMS CEASED AND FLOW RETURNED TO NOMINAL. HEALTH EFFECT - IMPACT CODE 4643 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29380.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22234996&pc=PYX | 1220648-2025-29382 | 2025-05-23 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | PYX | IMPELLA RP FLEX | Device Sensing Problem | Hemorrhage/Bleeding | P170011 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA RP FLEX DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THERE WERE PLACEMENT SIGNAL ISSUES. THE RP FLEX APPEARED TO STILL BE OPERATING WITHIN NORMAL LIMITS AND SUPPORT CONTINUED. ADDITIONALLY, THE PATIENT HAD BLEEDING OF UNKNOWN ETIOLOGY. EIGHT (8) UNITS OF BLOOD PRODUCTS WERE GIVEN TO THE PATIENT. HEPARIN WAS USED IN THE PURGE PRIOR TO THE PATIENT BLEEDING. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA RP SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER. SECTION: WARNINGS & CAUTIONS: CAUTIONS. ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ Manufacturer Narrative: VASCULAR DAMAGE - PERIPHERAL VESSEL: THE PATIENT HAD BLEEDING OF UNKNOWN ETIOLOGY. 8 UNITS OF BLOOD PRODUCTS WERE GIVEN TO THE PATIENT. THE ROOT CAUSE OF VASCULAR DAMAGE - PERIPHERAL VESSEL WAS NOT DETERMINED AS INSUFFICIENT CLINICAL INFORMATION WAS PROVIDED OPTICAL SIGNAL ISSUE: NO PRODUCT WAS RETURNED FOR ANALYSIS. THE DATA LOGS CONFIRM PSNR ALARM AND SHOW UPON INSERTION, A GRADUAL DECREASE IN RAW OPTICAL VALUES WHILE SNR AND CONTRAST REMAINED STABLE. CVP IS BELOW -100, TRIGGERING ALARMS. MOTOR CURRENT ALSO TRENDS DOWN AS RAW OPTICAL DOES. THE PUMP MAY HAVE BEEN EXPERIENCING SUCTION CONDITION AS LOW MOTOR CURRENT INDICATES LIMITED VOLUME. SINCE THE OPTICAL SENSOR IS TRENDING MORE NEGATIVE WITH THE MOTOR CURRENT ALSO TRENDING DOWN, AND OPTICAL IMPROVES (GOES POSITIVE) AS MOTOR CURRENT INCREASES, IT IS POSSIBLE THAT IT IS A SUCTION CONDITION CAUSING PSNR. CLINICAL MENTION THAT ISSUE RESOLVED AFTER DEVICE WAS SLIGHTLY REPOSITIONED. THE ROOT CAUSE OF THE OPTICAL SIGNAL ISSUE WAS UNABLE TO BE DETERMINED AS NO OPTICAL FAILURE PATTERN WAS IDENTIFIED IN THE LOGS AND NO PRODUCT WAS RETURNED TO CONFIRM POSSIBLE ROOT CAUSE WITH LIMITED CLINICAL INFORMATION. Manufacturer Narrative: B1 PRODUCT PROBLEM WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29382. H6 MEDICAL DEVICE PROBLEM CODE 2917 WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29382. H10 INVESTIGATION RESULTS FOR OPTICAL SIGNAL ISSUE WAS ERRONEOUSLY ENTERED ON MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29382-1.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22235001&pc=NVN | 2124215-2025-39784 | 2025-05-15 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-17 04:00 | NVN | INGEVITY+ | Adverse Event Without Identified Device or Use Problem | Pain; Perforation; Cardiac Tamponade |  |  | 1 | Event Description: IT WAS REPORTED THAT DURING A PACEMAKER IMPLANT PROCEDURE, THIS LEAD WAS IMPLANTED WITH NO ISSUES OBSERVED. ABOUT TWO WEEKS LATER, THE PATIENT WAS IN THE HOSPITAL AND DIED. THE OFFICIAL CAUSE OF DEATH WAS CARDIAC TAMPONADE. THE PHYSICIAN SUSPECTED A PERFORATION HAD OCCURRED DURING THE IMPLANT PROCEDURE. THE IMPLANTED LEAD WAS NOT EXPLANTED POST-MORTEM.Event Description: IT WAS REPORTED THAT DURING A PACEMAKER IMPLANT PROCEDURE, THIS LEAD WAS IMPLANTED WITH NO ISSUES OBSERVED. ABOUT TWO WEEKS LATER, THE PATIENT WAS IN THE HOSPITAL AND DIED. THE OFFICIAL CAUSE OF DEATH WAS CARDIAC TAMPONADE. THE PHYSICIAN SUSPECTED A PERFORATION HAD OCCURRED DURING THE IMPLANT PROCEDURE. THE IMPLANTED LEAD WAS NOT EXPLANTED POST-MORTEM. ADDITIONAL DETAILS WERE PROVIDED. IT WAS REPORTED THAT THE PATIENT PRESENTED BACK TO THE HOSPITAL THREE DAYS AFTER THE IMPLANT PROCEDURE, COMPLAINING OF RIGHT SHOULDER PAIN. IT WAS NOTED THE PATIENT HAD UNDERGONE AN ECHOCARDIOGRAM AFTER THE IMPLANT PROCEDURE, AND ANOTHER ONE WAS PERFORMED WHEN THEY RETURNED TO THE HOSPITAL. NO ADDITIONAL MEDICAL OR SURGICAL INTERVENTIONS, OR OTHER IMAGING, WERE REPORTED. THE PATIENT REMAINED IN THE HOSPITAL UNTIL THEIR DEATH. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED. THIS REPORT WILL BE UPDATED WHEN ADDITIONAL INFORMATION IS RECEIVED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS REQUESTED. THIS REPORT WILL BE UPDATED WHEN ADDITIONAL INFORMATION IS RECEIVED. IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22235131&pc=OZD | 1220648-2025-29384 | 2025-05-23 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | OZD | IMPELLA CP | Increase in Pressure | Hemolysis | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE TEAM PLACED THE CP FOR THE SUPPORT WITH DIFFICULTY DUE TO FEMORAL CALCIFICATION. THERE WAS RISING PP THAT PROMPTED THE TEAM TO ADD TISSUE PLASMINOGEN ACTIVATOR (TPA) TO THE PURGE. THE PATIENT ALSO HAD SIGNS OF HEMOLYSIS. THE PUMP REMAINED ON DESPITE THIS FOR A TOTAL OF FIVE (5) DAYS UNTIL THE FAMILY MADE THE CHOICE TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA CP WITH SMARTASSIST SYSTEM INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22235269&pc=NIK | 2124215-2025-39804 | 2025-03-29 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-17 04:00 | NIK | RESONATE? X4 CRT-D | Defective Device | Unspecified Heart Problem | P010012 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) PASSED AWAY WHILE IN (B)(6). THE CAUSE OF DEATH WAS NOT PROVIDED. HOWEVER, THE PATIENT ADVOCATE REPORTED THAT THE DEVICE ACTUALLY DAMAGED THE PATIENT'S HEART. THE DEVICE WAS NOT RETURNED AND WAS LIKELY NOT EXPLANTED POST-MORTEM. THIS REPORT WILL BE UPDATED IF ADDITIONAL INFORMATION IS RECEIVED. Manufacturer Narrative: THIS REPORT WILL BE UPDATED IF ADDITIONAL INFORMATION IS RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22235316&pc=DSQ | 2916596-2025-03607 | 2025-05-22 04:00 | Death | THORATEC CORPORATION | 2025-06-17 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22235330&pc=DQY | 2124215-2025-39378 | 2025-06-11 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-17 04:00 | DQY | WATCHMAN FXD CURVE? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Air Embolism; Bradycardia; Cardiac Arrest | K212228 |  | 1 | Event Description: IT WAS REPORTED AIR EMBOLISM AND PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE CLOSURE (LAA) PROCEDURE WAS BEING PERFORMED UNDER MONITORED ANESTHESIA CARE (MAC) AND INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING. A VERSACROSS CONNECT (VCC) LAAC KIT WAS SELECTED FOR USE. THE PHYSICIAN PERFORMED A SUCCESSFUL TRANSSEPTAL PUNCTURE (TSP) USING THE WATCHMAN FXD CURVE ACCESS SYSTEM (WAS). WHEN REMOVING THE VCC DILATOR, IT WAS NOTED THAT THERE WAS NO BLEED-BACK. THE END OF THE WAS WAS LIFTED UP AND HEARD A VACUUM/SUCKING SOUND, SO THE PHYSICIAN QUICKLY PLACED THE DILATOR BACK DOWN ON THE TABLE, LIFTED END OF THE WAS UP AGAIN AND HEARD THE SAME SUCKING SOUND AGAIN, QUICKLY LOWERED THE WAS DOWN. THIS IS WHEN IT WAS NOTED AIR BUBBLES ON THE ICE IMAGING, AND THE PATIENT'S HEART RATED DROPPED. THE PROCEDURE WAS ABORTED. TRANSCUTANEOUS PACING WAS INITIATED, PATIENT WAS INTUBATED, FLUIDS GIVEN AND RESUSCITATION EFFORTS STARTED ONCE PATIENT CONTINUED TO DECLINE. THE PATIENT WAS UNABLE TO BE SUCCESSFULLY RESUSCITATED. THE PATIENT WAS PRONOUNCED DEAD APPROXIMATELY AT 9:30 AM. IN THE PHYSICIAN'S OPINION, IT WAS UNCERTAIN IF THE PATIENT POSSIBLY HAD A VERY LOW OR NEGATIVE LEFT ATRIUM PRESSURE OR POSSIBLY TOOK A BREATH DURING MAC SEDATION CAUSING NEGATIVE PRESSURE RESULTING IN NO BLEED-BACK FROM THE WAS AFTER THE VCC DILATOR WAS REMOVED. THE PHYSICIANS WERE CERTAIN THE WAS/VCC SYSTEM WERE FLUSHED PROPERLY PRIOR TO INSERTION. THE WAS WAS INSPECTED AFTER THE EVENT AND APPEARED TO BE INTACT AND WORKING PROPERLY. AIR ENTRY BELIEVED TO BE AT THE PORT OF THE WAS RESULTING IN AIR IN THE LEFT SIDE OF THE HEART. THE VALVE WAS CLOSED AT THE END OF THE WAS AND ASPIRATION WAS FROM THE SIDE PORT. AIR INGRESS SPECIFICALLY TO THE LEFT SIDE OF THE HEART MAY HAVE OCCURRED WHEN THE VCC DILATOR WAS REMOVED, ULTIMATELY CAUSING CARDIAC ARREST. THE OFFICIAL CAUSE OF DEATH WAS AIR EMBOLISM RESULTING IN CARDIAC ARREST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22235505&pc=OZD | 1220648-2025-29383 | 2024-12-28 05:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | OZD | IMPELLA CP | Difficult or Delayed Positioning | Arrhythmia; Hemolysis | P140003 |  | 1 | Event Description: A NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED HEMOLYSIS WITH A "PROBABLE" RELATIONSHIP TO THE IMPELLA CP. LABS WERE HEMOLYZED. IT WAS FURTHER REPORTED THAT THE PATIENT CODED REQUIRING COMPRESSIONS. RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED AND THE IMPELLA CP WAS REPOSITIONED. THE PATIENT HAD AORTIC STENOSIS MAKING PROPER POSITIONING CHALLENGING. THE PATIENTS PROGNOSIS WAS POOR, AND CARE WAS WITHDRAWN. THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP DEVICE AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS, POSITIONING ISSUES, AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. HEMOLYSIS: THE CAUSE OF THE HEMOLYSIS WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL INFORMATION FOR IF THE HEMOLYSIS RESOLVED. POSITIONING ISSUES: THE CAUSE OF THE POSITIONING ISSUE WAS PATIENT CONDITION RELATED DUE TO THE PATIENTS NOTED AORTIC STENOSIS (AS) AND SMALL LV THAT MADE PROPER POSITIONING CHALLENGING. VF/VT/AF/AT: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22235544&pc=MHX | 1218950-2025-000358 | 2025-05-14 04:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-06-17 04:00 | MHX | PATIENT INFORMATION CENTER IX | Defective Alarm | Aspiration/Inhalation; Insufficient Information | K153702 |  | 1 | Event Description: IT WAS REPORTED THE CUSTOMER REQUESTED ASSISTANCE IN OBTAINING CLINICAL AUDIT LOGS FROM THE CENTRAL STATION FOLLOWING A PATIENT DEATH. THE CUSTOMER WANTED TO DETERMINE WHETHER THE DEVICE ALARMED APPROPRIATELY. THE CUSTOMER INDICATED THE PATIENT HAD ASPIRATED AT SOME POINT, WHICH LED TO DEATH. A LOG REVIEW WAS PERFORMED BY THE CUSTOMER, REVEALING THE ALARMS WERE GENERATED AS EXPECTED. THERE WERE MULTIPLE ALARMS WHICH WERE LATCHED BUT NOT ACKNOWLEDGED, WHICH MASKED LOWER-LEVEL ALARMS. THE LOG REVIEW PROVIDED ANSWERS TO THE CUSTOMER REQUEST. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. Manufacturer Narrative: A PHILIPS PRODUCT SUPPORT ENGINEERS (PSE) REVIEWED THE LOGS AND DETERMINED THAT THEY REFLECT THAT THE ALARMS FOR THE BED IN QUESTION DURING THE TIME PERIOD OF INTEREST WERE NOTED. A PHILIPS CLINICAL APPLICATIONS SPECIALIST (CAS) WENT ON SITE CONFIRMED THAT THE UNIT IS WORKING AS EXPECTED, THE CLINICAL AUDIT LOGS WERE ATTACHED, AND ALARMS WERE PRESENT AS EXPECTED WITH LATCHED AND NON-LATCHED ALARMS. BASED ON THIS INFORMATION, THERE WAS NO MALFUNCTION OF THE DEVICE. IN ADDITION, THERE WAS NO ALLEGATION THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. Manufacturer Narrative: E1: REPORTING INSTITUTION PHONE # (B)(6). E1: REPORTER PHONE # (B)(6). PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22235682&pc=DSQ | 2916596-2025-03866 | 2025-05-21 04:00 | Death | THORATEC CORPORATION | 2025-06-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Aneurysm; Coagulation Disorder; Hemorrhage/Bleeding; Shock; Heart Failure/Congestive Heart Failure; Mitral Valve Insufficiency/ Regurgitation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO A LARGE VOLUME OF HEMORRHAGIC SHOCK AND EVENTUAL WITHDRAWAL OF CARE.Event Description: IT WAS ADDITIONALLY REPORTED THAT LEFT VENTRICULAR ASSIST DEVICE (LVAD) AND CENTRAL RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT WAS COMPLICATED BY LARGE VOLUME HEMORRHAGIC STROKE. THERE WERE NO CHANGES TO PATIENT CONDITION OR ANTICOAGULATION STATUS THAT MAY HAVE CONTRIBUTED TO THE DEATH. ON (B)(6) 2025, A MITRAL VALVE REPAIR (MVR) AND LEFT VENTRICLE (LV) ANEURYSM REPAIR WERE PERFORMED, THE PATIENT SUFFERED A HEMORRHAGIC SHOCK, AND THE CHEST WAS LEFT OPEN. ON (B)(6) 2025, THE PATIENT WAS TAKEN BACK FOR A CHEST WASH OUT AND RVAD CANNULATION RECONFIGURATION TO ALLOW FOR PULMONARY BYPASS DUE TO SEVERELY ELEVATED PULSE VOLUME RECORDING (PVR) AND DIFFICULTY WITH ADEQUATE FLOWS ON RVAD AND LVAD. THE PATIENT DEVELOPED DIFFUSE COAGULOPATHY, AND THE CHEST WAS LEFT OPEN. THE DEVICE OPERATED AS EXPECTED AND WAS NOT CONSIDERED TO BE RELATED TO THE DEATH. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE PATIENT'S OUTCOME WAS NOT CONSIDERED TO BE DEVICE RELATED, AND THE DEVICE OPERATED AS EXPECTED. THE PUMP WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING BLEEDING AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 5 OF THE IFU, ¿SURGICAL PROCEDURES¿, EXPLAINS THAT MODERATE TO SEVERE AORTIC INSUFFICIENCY MUST BE CORRECTED AT THE TIME OF DEVICE IMPLANT. THIS SECTION ALSO STATES THAT IF THE AORTIC INSUFFICIENCY IS NOT ADDRESSED, THE DEVICE WILL NOT BE ABLE TO PROVIDE THE INTENDED FLOW. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THERE IS A RISK OF BLEEDING. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22235746&pc=DSQ | 2916596-2025-03868 | 2025-05-28 04:00 | Death | THORATEC CORPORATION | 2025-06-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH PATIENT OUTCOME THAT THE PATIENT PASSED AWAY DUE TO RESPIRATORY FAILURE RELATED TO AMIODARONE TOXICITY. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE ACCOUNT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE AND ARRHYTHMIA, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22236058&pc=OZD | 1220648-2025-29390 | 2025-05-22 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | OZD | IMPELLA 5.5 | Positioning Failure | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA (VT) AND WAS DEFIBRILLATED FIVE TIMES. THE FOLLOWING DAY, VT OCCURRED AGAIN. ECHOCARDIOGRAM AT BEDSIDE SHOWED THE IMPELLA 5.5 HAD NOT MOVED AT 4.8 CENTIMETERS BUT WAS UP AGAINST THE VENTRICLE WALL. THE IMPELLA 5.5 WAS REPOSITIONED TO 4.2 CENTIMETERS TO BE MOVED FROM THE VENTRICLE WALL. IT WAS FURTHER REPORTED THAT PATIENT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT AND POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. VT: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. ¿POSITIONING: PUMP REPOSITIONED TO 4.2CM TO HOPEFULLY HELP GET OFF THE WALL SOME. MYOCARDIUM EXTREMELY SENSITIVE AND EXTREMELY SMALL LEFT VENTRICLE (LV) CAVITY. THE CAUSE OF THE POSITIONING ISSUES IS MOST LIKELY PATIENT CONDITION RELATED DUE TO PATIENT'S EXTREMELY SMALL LV. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22236203&pc=OZD | 1220648-2025-29391 | 2025-05-23 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT CAME INTO THE HOSPITAL AS A NON-ST-SEGMENT ELEVATION MYOCARDIAL INFARCTION AND WITH CHEST PAIN. THE DOCTOR BROUGHT THE PATIENT TO THE CARDIAC CATHETERIZATION LAB. A 6FR SHEATH WAS PLACED IN RADIAL AND THE PATIENT WAS FOUND TO HAVE 99% LYMPHATIC MALFORMATION (LM) / LEFT ANTERIOR DESCENDING (LAD) ARTERY / CIRCUMFLEX ARTERY (CX)/ RIGHT CORONARY ARTERY (RCA) / WITH A END DIASTOLIC PRESSURE OF 30 WITH A EJECTION FRACTION OF 20%. CARDIOTHORACIC SURGERY WAS CONSULTED AND DR. (B)(6) WAS GOING TO TAKE HIM TO THE OPERATING ROOM TOMORROW FOR CORONARY ARTERY BYPASS GRAFTING. INTRA-AORTIC BALLOON PUMP WAS GOING TO BE PLACED BUT THE PATIENT DECOMPOSED IN THE CHRONIC LYMPHOCYTIC LEUKEMIA. DECISION BY THE DOCTOR TO PLACE IMPELLA CP FOR SUPPORT AND THEN TO FIX ALL THE VESSELS WITH STENTS. RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED VERY QUICKLY BY STAFF. THE DOCTORS ALL ASSISTED WITH THE INSERTION OF IMPELLA AND STENTS. PATIENTS LM / LAD / CX / RCA WAS ALL STENTED WHILE ON SUPPORT WITH A GOOD RESULT. PATIENT WAS INTUBATED. SHEATH WAS REMOVED FROM THE GROIN AND THE REPO WAS PLACED. NO BLEEDING WAS NOTED AT THE SITE. PLAN IS TO REST THE PATIENTS HEART FOR A FEW DAYS AND HOPE FOR HEART RECOVERY. THE FOLLOWING DAY THE PATIENT DEVELOPED FIXED AND DILATED PUPILS LAST NIGHT. A COMPUTED TOMOGRAPHY SCAN SHOWED RIGHT CEREBRAL HEMORRHAGE WITH INTRAVENTRICULAR AND SUBARACHNOID EXTENSION. ANTICOAGULATION IS ON HOLD DUE TO BRAIN BLEED. RIGHT, GROIN ACCESS SITE REMAINS CLEAN, DRY AND INTACT. IT WAS FURTHER REPORTED THAT THE PATIENT EXPIRED WHILE ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22236416&pc=FED | 2124215-2025-39274 | 2023-07-17 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-17 04:00 | FED | NAVIGATOR HD | Peeled/Delaminated; Material Split, Cut or Torn | Abdominal Pain; Atrial Fibrillation; Low Blood Pressure/ Hypotension; Ischemia; Sepsis; Septic Shock; Urinary Retention; Peritonitis; Constipation | K140323 |  | 1 | Event Description: IT WAS REPORTED THAT A NAVIGATOR ACCESS SHEATH DEVICE WAS USED DURING A URETEROSCOPY PROCEDURE. DURING THE PROCEDURE, THE NAVIGATOR ACCESS SHEATH "BROKE APART WHILE INSIDE THE PATIENT." REPORTEDLY, THE FRAGMENTS WERE REMOVED USING A GRASPER. NINE DAYS LATER, THE PATIENT "CALLED 911 COMPLAINING OF SEVERE ABDOMINAL PAIN, NAUSEA, INCLUDING ABDOMINAL DISTENSION AND FEELINGS OF CONSTIPATION." REPORTEDLY, THESE FEELINGS HAD BEEN "PRESENT SINCE HER DISCHARGE" FOLLOWING THE URETEROSCOPY PROCEDURE. UPON ARRIVAL AT THE EMERGENCY ROOM, IT WAS NOTED THE PATIENT WAS EXPERIENCING HYPOTENSION AND ATRIAL FIBRILLATION. OVER THE NEXT FIVE DAYS THE PATIENT'S "ABDOMINAL COMPLICATIONS WORSENED RAPIDLY." IT WAS NOTED THAT THE PATIENT WAS HAVING "PERITONITIS AND LIKELY AN ISCHEMIC OR INFECTIOUS BOWEL." IT WAS ALSO NOTED THAT THE URINE INPUT WAS SEVERELY LIMITED AND WAS ULTIMATELY DIAGNOSED WITH SEPSIS, SEPTIC SHOCK, AND LACTIC ACIDOSIS. AFTER SEVERAL DAYS, THE PATIENT WAS "UNRESPONSIVE IN HER HOSPITAL BED AND PRONOUNCED DEAD." IT WAS FURTHER REPORTED THAT AN AUTOPSY NOTED "HYPERTENSIVE AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE AND "COMPLICATIONS [FROM] URETEROSCOPY." FURTHER, IT WAS REPORTED THAT "THE PROCEDURE WAS COMPLICATED BY DISINTEGRATION OF THE URETEROSCOPE COVER AND SEVERE POSTOPERATIVE HYPOTENSION."Event Description: IT WAS REPORTED THAT A NAVIGATOR ACCESS SHEATH DEVICE WAS USED DURING A URETEROSCOPY PROCEDURE. DURING THE PROCEDURE, THE NAVIGATOR ACCESS SHEATH "BROKE APART WHILE INSIDE THE PATIENT." REPORTEDLY, THE FRAGMENTS WERE REMOVED USING A GRASPER. NINE DAYS LATER, THE PATIENT "CALLED 911 COMPLAINING OF SEVERE ABDOMINAL PAIN, NAUSEA, INCLUDING ABDOMINAL DISTENSION AND FEELINGS OF CONSTIPATION." REPORTEDLY, THESE FEELINGS HAD BEEN "PRESENT SINCE HER DISCHARGE" FOLLOWING THE URETEROSCOPY PROCEDURE. UPON ARRIVAL AT THE EMERGENCY ROOM, IT WAS NOTED THE PATIENT WAS EXPERIENCING HYPOTENSION AND ATRIAL FIBRILLATION. OVER THE NEXT FIVE DAYS THE PATIENT'S "ABDOMINAL COMPLICATIONS WORSENED RAPIDLY." IT WAS NOTED THAT THE PATIENT WAS HAVING "PERITONITIS AND LIKELY AN ISCHEMIC OR INFECTIOUS BOWEL." IT WAS ALSO NOTED THAT THE URINE INPUT WAS SEVERELY LIMITED AND WAS ULTIMATELY DIAGNOSED WITH SEPSIS, SEPTIC SHOCK, AND LACTIC ACIDOSIS. AFTER SEVERAL DAYS, THE PATIENT WAS "UNRESPONSIVE IN HER HOSPITAL BED AND PRONOUNCED DEAD." IT WAS FURTHER REPORTED THAT AN AUTOPSY NOTED "HYPERTENSIVE AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE AND "COMPLICATIONS [FROM] URETEROSCOPY." FURTHER, IT WAS REPORTED THAT "THE PROCEDURE WAS COMPLICATED BY DISINTEGRATION OF THE URETEROSCOPE COVER AND SEVERE POSTOPERATIVE HYPOTENSION." Manufacturer Narrative: BLOCKS D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI)# AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK H6: DEVICE CODE A040506 CAPTURES THE REPORTABLE EVENT OF SHEATH PEELED. PATIENT CODE E1002 CAPTURES THE REPORTABLE EVENT OF ABDOMINAL PAIN. PATIENT CODE E1007 CAPTURES THE REPORTABLE EVENT OF CONSTIPATION. PATIENT CODE E060102 CAPTURES THE REPORTABLE EVENT OF ATRIAL FIBRILLATION. PATIENT CODE E1024 CAPTURES THE REPORTABLE EVENT OF PERITONITIS. PATIENT CODE E0509 CAPTURES THE REPORTABLE EVENT OF ISCHEMIA. PATIENT CODE E1309 CAPTURES THE REPORTABLE EVENT OF URINARY RETENTION. PATIENT CODE E0306 CAPTURES THE REPORTABLE EVENT OF SEPSIS. IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF UNEXPECTED MEDICAL INTERVENTION. IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF PATIENT DEATH. IMPACT CODE F1908 CAPTURES THE REPORTABLE EVENT OF PROLONGED SURGERY. Manufacturer Narrative: BLOCKS D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK H2: ADDITIONAL INFORMATION: BLOCK H6 (PATIENT CODES, DEVICE CODES) BLOCK A6 (RACE). BLOCK H6: DEVICE CODE A0414 CAPTURES THE REPORTABLE EVENT OF SHEATH TORN. PATIENT CODE E1002 CAPTURES THE REPORTABLE EVENT OF ABDOMINAL PAIN. PATIENT CODE E1007 CAPTURES THE REPORTABLE EVENT OF CONSTIPATION. PATIENT CODE E060102 CAPTURES THE REPORTABLE EVENT OF ATRIAL FIBRILLATION. PATIENT CODE E1024 CAPTURES THE REPORTABLE EVENT OF PERITONITIS. PATIENT CODE E0509 CAPTURES THE REPORTABLE EVENT OF ISCHEMIA. PATIENT CODE E1309 CAPTURES THE REPORTABLE EVENT OF URINARY RETENTION. PATIENT CODE E0306 CAPTURES THE REPORTABLE EVENT OF SEPSIS. PATIENT CODE E233605 CAPTURES THE REPORTABLE EVENT OF SEPSIS SHOCK. PATIENT CODE E2321 CAPTURES THE REPORTABLE EVENT OF HYPOTENSION. IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF UNEXPECTED MEDICAL INTERVENTION. IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF PATIENT DEATH. IMPACT CODE F1908 CAPTURES THE REPORTABLE EVENT OF PROLONGED SURGERY. BLOCK H11: CORRECTION TO FIELD D6A (IMPLANT DATE).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22236467&pc=DTB | 2649622-2025-16398 | 2025-06-03 04:00 | Death | MPRI | 2025-06-17 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Over-Sensing; Under-Sensing | Ventricular Fibrillation | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT UPON REVIEW OF AN INTERROGATION REPORT OBTAINED AT THE EMERGENCY DEPARTMENT, RIGHT VENTRICULAR (RV) AND RIGHT ATRIAL (RA) LEAD UNDERSENSING WAS NOTED DURING VENTRICULAR TACHYCARDIA (VT) EVENTS, AND RA LEAD FAR FIELD R-WAVE (FFRW) OVERSENSING WAS NOTED ON THE CURRENT ELECTROGRAM (EGM). THE FOLLOWING DAY IT WAS INDICATED THAT OVERNIGHT, CARDIOLOGY HAD ORDERED THE DEVICE PROGRAMMED OFF/MAGNET PLACED OVER THE DEVICE. IT WAS INDICATED THAT THE PATIENT HAD BEEN STABLE; HOWEVER, THEY STARTED COUGHING, WENT INTO VENTRICULAR FIBRILLATION (VF) ARREST, AND PASSED AWAY. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22236481&pc=LWS | 2649622-2025-16399 | 2025-06-03 04:00 | Death | MPRI | 2025-06-17 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Under-Sensing | Ventricular Fibrillation | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT UPON REVIEW OF AN INTERROGATION REPORT OBTAINED AT THE EMERGENCY DEPARTMENT, RIGHT VENTRICULAR (RV) AND RIGHT ATRIAL (RA) LEAD UNDERSENSING WAS NOTED DURING VENTRICULAR TACHYCARDIA (VT) EVENTS, AND RA LEAD FAR FIELD R-WAVE (FFRW) OVERSENSING WAS NOTED ON THE CURRENT ELECTROGRAM (EGM). THE FOLLOWING DAY IT WAS INDICATED THAT OVERNIGHT, CARDIOLOGY HAD ORDERED THE DEVICE PROGRAMMED OFF/MAGNET PLACED OVER THE DEVICE. IT WAS INDICATED THAT THE PATIENT HAD BEEN STABLE; HOWEVER, THEY STARTED COUGHING, WENT INTO VENTRICULAR FIBRILLATION (VF) ARREST, AND PASSED AWAY. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22236497&pc=OZD | 1220648-2025-29397 | 2025-04-22 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemolysis | P140003 |  | 1 | Event Description: ABIOMED RECEIVED A LOQI STUDY COMPLAINT THAT REPORTED AN ADVERSE EVENT FOR WORSENING VENTRICULAR TACHYCARDIA WITH RELATIONSHIP TO IMPELLA CP LISTED AS POSSIBLY. "LOQI STATES: VTACH HAPPENED AFTER IMPELLA PLACEMENT. LEVOPHED STARTED. PATIENT WAS DEFIBRILLATED WITH 200J. PATIENT RETURNED BACK TO V PACED RHYTHM." LATER THE PATIENT ARRESTED AND PASSED AWAY.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT HAD DEVELOPED HEMOLYSIS WITH A RELATIONSHIP TO THE PUMP THAT WAS TREATED BY BLOOD PRODUCTS. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. B5 UPDATED WITH ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE H6 UPDATED WITH THE ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ B2 REQUIRED INTERVENTION SELECTION WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29397.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22236695&pc=GAW | 2210968-2025-06971 | 2024-01-25 05:00 | Death | ETHICON INC. | 2025-06-17 04:00 | GAW | PROLENE POLYPROPYLENE SUTURE UNKNOWN PRODUCT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K133356 |  | 1 | Event Description: TITLE: EFFICACY OF PROLENE VS VICRYL IN TRACHEAL RESECTION AND ANASTOMOSIS. THE AIM OF THIS STUDY IS TO COMPARE THE INCIDENCE OF MORBIDITY/MORTALITY DUE TO LEAKS AND LONG-TERM RECURRENCE OF TRACHEAL STENOSIS. BETWEEN JANUARY 2022 AND JUNE 2023, THERE WAS A TOTAL OF 66 CASES THAT REQUIRED TRACHEAL RESECTION AND ANASTOMOSIS. AMONG THEM, 32 PATIENTS UNDERWENT TRACHEAL RESECTION WITH PROLENE (GROUP P), AND THE REMAINING 34 UNDERWENT TRACHEAL RESECTION WITH VICRYL (GROUP V). REPORTED COMPLICATIONS ARE: GROUP P: N=6; LEAK FOLLOWING TRA (TRACHEAL RESECTION AND ANASTOMOSIS), TREATMENT: NOT REPORTED. GROUP V: N=1; LEAK FOLLOWING TRA (TRACHEAL RESECTION AND ANASTOMOSIS), TREATMENT: NOT REPORTED. N=3; MORTALITY DUE TO TRACHEA-RELATED COMPLICATIONS, TREATMENT: NONE. IN CONCLUSION, WE DEDUCED THAT VICRYL AND PROLENE PERFORM SIMILARLY IN REDUCING MORBIDITY (LEAK), BUT BOTH GROUPS' MORTALITY RATES REMAIN THE SAME. THEREFORE, TO ALLEVIATE THESE PROBLEMS, WE RECOMMEND USING PROLENE WHEN UNDERGOING TRACHEAL RESECTION AND ANASTOMOSIS. Manufacturer Narrative: PRODUCT COMPLAINT#: (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE CAUSE OF DEATH? DOES THE SURGEON BELIEVE THAT ANY OF THE ETHICON PRODUCTS INVOLVED CAUSED AND/OR CONTRIBUTED TO THE POST-OPERATIVE COMPLICATIONS/DEATH DESCRIBED IN THE ARTICLE? WHICH SPECIFIC ETHICON PRODUCTS HAVE BEEN USED DURING THE PROCEDURES (PRODUCT CODE, LOT NUMBER)? DOES THE SURGEON BELIEVE THERE WAS ANY DEFICIENCY WITH ANY OF THE ETHICON PRODUCTS USED IN THIS PROCEDURE? IF SO, PLEASE PROVIDE DETAILS. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE A COMPLAINT REFERENCE NUMBER. PATIENT DEMOGRAPHICS? WOULD THE JOURNAL AUTHOR / PHYSICIAN LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. H6: COMPONENT CODE: G07002 ¿ DEVICE NOT RETURNED. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. CITATION: DOI: 10.47009/JAMP.2024.6.1.330.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22236724&pc=CBK | 2518422-2025-106695 | 2025-06-06 04:00 | Death | RESPIRONICS, INC. | 2025-06-17 04:00 | CBK | TRILOGY 100 | Degraded; Electrical /Electronic Property Problem; Failure to Calibrate | Insufficient Information | K083526 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION CONCERNING A PRESERVATION OF EVIDENCE LETTER. THE COMPLAINT/ALLEGED ISSUE IS UNKNOWN, ONLY THAT A PRESERVATION OF EVIDENCE LETTER WAS RECEIVED WITH NO DETAILS. IT IS BELIEVED THAT THERE WAS A PATIENT DEATH, BUT NO DETAILS ON THE CAUSE OF DEATH WERE PROVIDED. THE MANUFACTURER IS TREATING THIS REPORT ARE BEING CONCERNED WITH THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. BASED ON THE INFORMATION CURRENTLY PROVIDED AND AVAILABLE, DEVICE CAUSE AND/OR CONTRIBUTION TO THE REPORTED EVENT OF DEATH CANNOT CURRENTLY BE CONFIRMED NOR REFUTED BASED ON THE EVIDENCE PRESENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION CONCERNING A PRESERVATION OF EVIDENCE LETTER. THE COMPLAINT/ALLEGED ISSUE IS UNKNOWN, ONLY THAT A PRESERVATION OF EVIDENCE LETTER WAS RECEIVED WITH NO DETAILS. IT IS BELIEVED THAT THERE WAS A PATIENT DEATH, BUT NO DETAILS ON THE CAUSE OF DEATH WERE PROVIDED. THE MANUFACTURER IS TREATING THIS REPORT ARE BEING CONCERNED WITH THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. BASED ON THE INFORMATION CURRENTLY PROVIDED AND AVAILABLE, DEVICE CAUSE AND/OR CONTRIBUTION TO THE REPORTED EVENT OF DEATH CANNOT CURRENTLY BE CONFIRMED NOR REFUTED BASED ON THE EVIDENCE PRESENT. TRILOGY 100 VENTILATOR, U.S.A. BT (P/N 1054260B S/N (B)(6)) RECEIVED AT MANUFACTURER'S PRODUCT INVESTIGATION LAB (PIL) FOR INVESTIGATION. UPON INITIAL EXTERNAL VISUAL INSPECTION OF THE SUSPECT TRILOGY 100 VENTILATOR, MANUFACTURER TECH DID NOTE SOME FOREIGN MATERIALS INDICATIVE OF CONTAMINATION AT VARIOUS SITES OF THE UNIT, INCLUDING WITHIN AND AROUND THE AIRFLOW OUTLET PORT. TECH RECORDED AND REVIEWED THE UNIT¿S EVENT LOG AND SETTINGS AND PROCEEDED TO OPERATE THE UNIT ON THE PRIMARY SETTINGS WITH WHICH IT WAS RECEIVED AT PIL AFTER CONNECTING IT TO A TEST LUNG. THE UNIT OPERATED WITHOUT ISSUES OR ALARMS. THE UNIT WAS CONNECTED TO A MULTI-FUNCTION TEST STATION WHERE TECH VERIFIED FAILURES FOR POS FLOW VERIFY, ZERO FLOW VERIFY AND CONTROL PRESSURE AND MONITOR PRESSURE. TECH DISASSEMBLED THE SUSPECT UNIT FOR INTERNAL INSPECTION. TECH FOUND NO SIGNS OF DAMAGE OR EVIDENCE OF DEFECTIVE OPERATION. TECH DID NOT DETECT ANY PINCHED OR BLOCKED TUBING. TECH ALSO FOUND THE AIRPATH FOAM TO BE FIRMLY SECURED, WITHOUT ANY SIGNS OF DELAMINATION OR DEGRADATION. TECH COULD NOT CONFIRM BLACK FOAM PARTICULATES OR FOAM DEGRADATION INSIDE THE SUSPECT UNIT AFTER DISASSEMBLY. AFTER REVIEW, TECH HAS DETERMINED THAT TRILOGY 100 VENTILATOR 1054260B/(B)(6) OPERATES AS DESIGNED AND FUNCTIONS AS EXPECTED IN REGARD TO THE AIRPATH FOAM BEING FOUND TO BE FIRMLY SECURED AND THE UNIT OPERATING WITHOUT ISSUES OR ALARMS. PER THE TEST FAILURES, IT WAS DETERMINED THAT THE UNIT DRIFTED OUT OF SPEC DUE TO SEPARATE CONTAMINATION ISSUE ON THE SENSOR BOARD AND REPLACEMENT OF THE SENSOR BOARD WOULD HAVE BEEN NEEDED TO ADDRESS THE ISSUE. CORRECTION HAS BEEN MADE TO SECTION H, TYPE OF REPORTED COMPLAINT FIELD FROM PRODUCT PROBLEM TO DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22237243&pc=OZD | 1220648-2025-29401 | 2025-05-24 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | OZD | IMPELLA CP | Difficult to Insert | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT DURING REMOVAL OF THE PEEL-AWAY SHEATH, THE PHYSICIAN INADVERTENTLY BEGAN TO CRACK AND PEEL THE COMPONENT AWAY DESPITE THE DISTAL END STILL BEING PARTLY WITHIN THE PATIENT'S BODY. UPON NOTICING THAT THE SHEATH WAS NOT FULLY OUTSIDE OF THE BODY, THE PHYSICIAN PROCEEDED WITH THEIR ATTEMPT. A HEMATOMA DEVELOPED IMMEDIATELY AND BOTH MANUAL PRESSURE, SUTURES, AND FEMOSTOP WERE PLACED TO MANAGE THE CONDITION. BLOOD WAS GIVEN TO COUNTERACT THE CONDITION. SUPPORT CONTINUED AS PLANNED WITH THE NEWLY IMPLANTED IMPELLA CP PUMP. THE PATIENT WAS EVENTUALLY TRANSITIONED TO COMFORT CARE AND PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND DIFFICULTY INSERTING/REMOVING INTRODUCER HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ACCESS SITE BLEEDING - MAJOR: THE CAUSE OF THE ACCESS SITE BLEEDING - MAJOR WAS USE RELATED DUE TO THE PHYSICIAN BREAKING THE PEEL AWAY SHEATH WITH STILL INSIDE THE PATIENT, CAUSING A HEMATOMA. DIFFICULTY INSERTING/REMOVING INTRODUCER: THE CAUSE OF THE DIFFICULTY INSERTING/REMOVING INTRODUCER WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS. B1 PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29401. H6 COMPONENT CODE 4742 WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29401. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT. USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER. ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22237466&pc=OZD | 1220648-2025-29404 | 2025-05-26 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | OZD | IMPELLA CP | Positioning Failure | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD A PATIENT ARRIVED BY EMERGENCY MEDICAL SERVICES AND WAS SEEN TO BE IN NEED OF IMPELLA CP AND EXTRACORPOREAL MEMBRANE OXYGENATION SUPPORT. WHILE ON SUPPORT, THE IMPELLA CP WAS ALARMING IMPELLA IN THE AORTA AND SUCTION. AN ECHOCARDIOGRAM BY THE PHYSICIAN ASSISTANT WAS QUESTIONABLE AT BEST DUE TO THE PATIENT'S PREVIOUS CHEST COMPRESSIONS AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE. THE IMPELLA CP WAS POSITIONED JUST ACROSS THE VALVE AND ADVANCED WHICH IMPROVED MOTOR CURRENT AND FLOWS. THE FOLLOWING DAY THERE WAS A SHORT EPISODE OF VENTRICULAR TACHYCARDIA WITH MOUTH CARE. FOUR DAYS LATER, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22237679&pc=OZD | 1220648-2025-29406 | 2025-05-24 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | OZD | IMPELLA CP | Positioning Failure | Hemolysis; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A CP PLACED FOR THE SUPPORT OF A PATIENT ADMITTED IN WITH A CARDIAC HISTORY AND STEMI. THE PATIENT HAD BEEN SHOCKED REPEATEDLY PRIOR TO THE PUMP PLACEMENT AND WAS INTUBATED AND SEDATED IN POOR CONDITION. THE CP WAS PLACED, THE PUMP WAS SEEN TO BE IN POOR POSITION AT THE VALVE AND THERE WAS SIGNS OF HEMOLYSIS. THE TEAM DID NOT REPOSITION BUT INSTEAD WHEN PATIENT MADE A DNR THE PATIENT EXPIRED AFTER JUST 39 HOURS OF SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION. ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS, ACCESS SITE BLEEDING-MINOR, AND POSITIONING ISSUES HAS BEEN COMPLETED. DATA REVIEW: DATA LOGS SHOWED PUMP RAN WITHOUT ISSUE DURING SUPPORT. THERE WERE MULTIPLE SUCTION ALARMS TRIGGERED DURING THE CASE. FAILURE MODE CHANGED FROM ACCESS SITE BLEEDING - MINOR TO ACCESS SITE BLEEDING - MAJOR BECAUSE BLOOD PRODUCT WERE GIVEN IN ADDITION TO MANUAL PRESSURE TO HELP MANAGED THE OOZING. THE CAUSE OF THE BLEEDING WAS MOST LIKELY PATIENT CONDITION RELATED SINCE PATIENT BLEEDING FROM ALL ORIFICES. THE ROOT CAUSE OF HEMOLYSIS WAS MOST LIKELY DUE TO PUMP WAS NOT IN THE PROPER POSITION SINCE THE PHYSICIAN CONFIRMED PUMP MIGHT BE INTERFERING WITH THE MITRAL VALVE (MV). THE ROOT CAUSE OF POSITIONING ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED BECAUSE PUMP WAS NOTED TO BE POSSIBLY INTERACTING WITH MV BUT WHEN REPOSITIONED, TRIGGERED ECTOPY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22237751&pc=OUT | 2029214-2025-01405 | 2021-06-13 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-06-17 04:00 | OUT | PIPELINE VANTAGE WITH SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P100018 |  | 1 | Event Description: MEDTRONIC RECEIVED A REPORT THAT THERE WAS TWISTING OF THE VANTAGE DEVICES DURING THE PROCEDURE. IT WAS DETERMINED TO BE A TEMPORARY TWISTING AND CONSIDERED A TECHNICAL ISSUE RATHER THAN A DEVICE MALFUNCTION. THE PHYSICIAN STATED THE DEVICE PERFORMED AS EXPECTED AND THE PROCEDURE WAS COMPLETED SUCCESSFULLY. THEREFORE, NO FURTHER INVESTIGATION WAS CARRIED OUT BY THE SITE. ADDITIONAL INFORMATION WAS RECEIVED AND IT WAS REPORTED DUAL ANTIPLATELET TREATMENT (DAPT) WAS ADMINISTERED. 5 DAYS PRIOR TO INDEX PROCEDURE, ASPIRIN 300 MG LOADING AND THEN 75 MG/DAILY FOR MAINTENANCE DOSE AND PRASUGREL 60MG LOADING AND 5 MG/DAILY MAINTENANCE DOSE. ANEURYSM LOCATION LEFT ICA-C6. ANEURYSM DIMENSIONS 32.5 X30 MM. MORPHOLOGY FUSIFORM. IT WAS CONCLUDED THAT NO DEVICE COMPLAINT TO BE REPORTED AS DEVICE WAS IMPLANTED SUCCESSFULLY AND PHYSICIAN DID NOT SEE ANY DEVICE ISSUE. NO DEVICE DAMAGE ¿ DEVICES WERE IMPLANTED SUCCESSFULLY. ADDITIONAL INFORMATION RECEIVED REPORTING THAT THE EVENT INVOLVED 2 PIPELINE VANTAGE DEVICES THAT WERE IMPLANTED FOR TREATMENT OF A LEFT C6 ICA FUSIFORM ANEURYSM MEASURING 32.5MM WITH NECK 30 MM. INCOMPLETE NECK COVERAGE WAS NOTED. IT WAS UNKNOWN IF THE EVENT WAS RELATED TO A DEVICE DEFICIENCY OR WHY 2 PIPELINE VANTAGE STENTS WERE IMPLANTED. NO PATIENT SYMPTOMS WERE REPORTED. ADDITIONAL INFORMATION RECEIVED REPORTED THE SECOND PIPELINE VANTAGE WAS IMPLANTED TO ADDRESS INCOMPLETE NECK COVERAGE BY THE FIRST PIPELINE VANTAGE. IT WAS NOT KNOWN AT THE TIME THE INFORMATION WAS RECEIVED WHETHER OR NOT THE INCOMPLETE COVERAGE WAS DUE TO ANY DEVICE MALFUNCTION. NECK COVERAGE WAS ACHIEVED AFTER IMPLANTATION OF THE SECOND PIPELINE VANTAGE. ADDITIONAL INFORMATION RECEIVED REPORTED THIS CASE WAS CONSIDERED A TECHNICAL ISSUE RATHER DEVICE MALFUNCTION. THE PHYSICIAN ALSO STATED THAT THE DEVICE PERFORMED AS EXPECTED AND PROCEDURE COMPLETED SUCCESSFULLY. A CONSTRUCT WITH 2 VANTAGE DEVICES WAS MADE DUE TO ANEURYSM SIZE AND MORPHOLOGY (GIANT FUSIFORM ANEURYSM). ADDITIONAL INFORMATION WAS RECEIVED INDICATING THE PATIENT HAD AN INTRACRANIAL HEMORRHAGE WITH DEATH. THE ADVERSE EVENT DID RESULT IN HOSPITALIZATION PRIOR TO THE DEATH. THE ADVERSE EVENT RESULTED AFTER POST-PROCEDURE AND WAS NOTED AS PROBABLY RELATED TO THE DISEASE UNDER STUDY. THE RESULT WAS A NEW OR WORSENING OF THE EXISTING NEUROLOGICAL DEFICITS. THE ADVERSE EVENT WAS POSSIBLY RELATED TO THE ANTIPLATELET MEDICINE, POSSIBLY RELATED TO THE DEVICE, AND POSSIBLY RELATED TO THE PROCEDURE. ADDITIONAL INFORMATION REPORTED THE HOSPITAL END DATE WAS ON (B)(6) 2021. ADDITIONAL INFORMATION RECEIVED CORRECTED THAT THE HOSPITAL END DATE WAS ON (B)(6) 2021. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE SITE RELATED ASSESSMENT WAS UPDATED TO NOT RELATED TO THE STUDY DEVICE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22237812&pc=OZD | 1220648-2025-29407 | 2025-05-20 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | OZD | IMPELLA 5.5 | Positioning Failure; Restricted Flow rate | Arrhythmia; Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT THE PATIENT ON IMPELLA 5.5 SUPPORT, THERE WAS SUCTION ON P8 WITH HYPOTENSION WHICH WAS RESOLVED BY LOWERING TO P7. THE IMPELLA WAS SHOWING THE INLET CLOSE TO THE SEPTUM AND REPOSITIONING WAS ATTEMPTED, BUT ECHO KEPT SHOWING THE INLET STILL CLOSE TO THE SEPTAL WALL. THE PATIENT WAS SHOCKED FOUR TIMES AND WAS RE-INTUBATED. IT WAS FURTHER REPORTED THAT THE PATIENT WAS MADE DNR AND EXPIRED ON SUPPORT. Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. THE ROOT CAUSE OF THE LOW PUMP FLOW AND POSITIONING ISSUES WAS NOT DETERMINED SINCE THERE WAS A LACK OF CLINICAL DETAILS, NO PRODUCT RETURNED, AND INSUFFICIENT DATA LOGS. THE PATIENT HAD A HISTORY OF VENTRICULAR TACHYCARDIA (VT) WITH AN AUTOMATIC IMPLANTABLE CARDIOVERTER DEFIBRILLATOR PRIOR TO IMPELLA SUPPORT. THE PATIENT WENT INTO VT STORM. THE MEDICAL TEAM DID NOT ATTRIBUTE VT TO THE IMPELLA, THUS H6 WAS REVISED TO REMOVE 1721 AND CHANGED TO 4445. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22237845&pc=NIK | 9614453-2025-02323 | 2024-11-01 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-06-17 04:00 | NIK | COBALT¿ HF QUAD CRT-D MRI SURESCAN¿ | Adverse Event Without Identified Device or Use Problem | Endocarditis; Drug Resistant Bacterial Infection | P010031 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) PATIENT PRESENTED TO THE HOSPITAL WITH A METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) BACTEREMIA INFECTION AND PNEUMONIA WITH PERSISTENTLY POSITIVE BLOOD CULTURES IN ADDITION TO AORTIC VALVE ENDOCARDITIS. THE PATIENT HAD A MRSA ABSCESS OF THE SPINE FROM THE C4 VERTEBRA THROUGH THE C6 VERTEBRA AND THE NEXT DAY HAD AN EPIDURAL DRAINAGE OF THE ABSCESS. THE PATIENT DIED TWO DAYS LATER. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. THE DEVICE WAS RETURNED AND SCREENING ANALYSIS WAS PERFORMED, BUT NO ISSUE WAS IDENTIFIED REQUIRING FULL ANALYSIS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22237863&pc=LWS | 2649622-2025-16413 | 2024-11-01 04:00 | Death | MPRI | 2025-06-17 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Endocarditis; Drug Resistant Bacterial Infection | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) PATIENT PRESENTED TO THE HOSPITAL WITH A METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) BACTEREMIA INFECTION AND PNEUMONIA WITH PERSISTENTLY POSITIVE BLOOD CULTURES IN ADDITION TO AORTIC VALVE ENDOCARDITIS. THE PATIENT HAD A MRSA ABSCESS OF THE SPINE FROM THE C4 VERTEBRA THROUGH THE C6 VERTEBRA AND THE NEXT DAY HAD AN EPIDURAL DRAINAGE OF THE ABSCESS. THE PATIENT DIED TWO DAYS LATER. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22237873&pc=OZD | 1220648-2025-29408 | 2025-05-30 04:00 | Death | ABIOMED, INC. | 2025-06-17 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Hernia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT ON IMPELLA 5.5 SUPPORT DEVELOPED OOZING IN THE POCKET. HEPARIN WAS HELD. IT WAS FURTHER REPORTED THAT THE PATIENT EXPERIENCED BRAIN BLEED AND BRAIN HERNIATION. THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22237920&pc=DTB | 2649622-2025-16415 | 2024-11-01 04:00 | Death | MPRI | 2025-06-17 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Endocarditis; Drug Resistant Bacterial Infection | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) PATIENT PRESENTED TO THE HOSPITAL WITH A METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) BACTEREMIA INFECTION AND PNEUMONIA WITH PERSISTENTLY POSITIVE BLOOD CULTURES IN ADDITION TO AORTIC VALVE ENDOCARDITIS. THE PATIENT HAD A MRSA ABSCESS OF THE SPINE FROM THE C4 VERTEBRA THROUGH THE C6 VERTEBRA AND THE NEXT DAY HAD AN EPIDURAL DRAINAGE OF THE ABSCESS. THE PATIENT DIED TWO DAYS LATER. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22237971&pc=OJX | 2649622-2025-16418 | 2024-11-01 04:00 | Death | MPRI | 2025-06-17 04:00 | OJX | ATTAIN PERFORMA MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Endocarditis; Drug Resistant Bacterial Infection | P080006 |  | 1 | Event Description: IT WAS REPORTED THAT THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D) PATIENT PRESENTED TO THE HOSPITAL WITH A METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) BACTEREMIA INFECTION AND PNEUMONIA WITH PERSISTENTLY POSITIVE BLOOD CULTURES IN ADDITION TO AORTIC VALVE ENDOCARDITIS. THE PATIENT HAD A MRSA ABSCESS OF THE SPINE FROM THE C4 VERTEBRA THROUGH THE C6 VERTEBRA AND THE NEXT DAY HAD AN EPIDURAL DRAINAGE OF THE ABSCESS. THE PATIENT DIED TWO DAYS LATER. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE PROXIMAL PORTION OF THE LEAD WAS RETURNED, ANALYZED, AND NO ANOMALIES WERE FOUND. VISUAL ANALYSIS OF THE LEAD INDICATED APPARENT EXPLANT DAMAGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22237991&pc=MKJ | 1220908-2025-02268 | 2025-05-30 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-06-17 04:00 | MKJ | R SERIES | Incorrect Interpretation of Signal | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ANALYZING THE HEART RHYTHM OF A PATIENT (AGE AND GENDER UNKNOWN), THE DEVICE ISSUED A "SHOCK ADVISED" PROMPT FOR A HEART RHYTHM THEY BELIEVE WAS NON-SHOCKABLE. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED; HOWEVER, IT WAS NOT A RESULT OF THE REPORTED MALFUNCTION. Manufacturer Narrative: THE ALS DEVICE WAS NOT RETURNED TO ZOLL MEDICAL CORPORATION, INSTEAD A STRIP FROM THE REPORTED EVENT WAS RECEIVED. REVIEW OF THE STRIP DOES NOT PROVIDE ENOUGH DATA TO EVALUATE THE ALGORTHM DECISION, BUT IT APPEARS THE OPERATOR ELECTED NOT TO SHOCK BASED ON THEIR ADVANCED LIFESUPPORT (ALS) TRAINING. ANALYSIS OF REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND. Manufacturer Narrative: JUSTIFICATION FOR NO UDI: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22238206&pc=DSQ | 3007042319-2025-01283 | 2025-03-18 04:00 | Death | HEARTWARE, INC. | 2025-06-17 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Pumping Problem | Arrhythmia; Intracranial Hemorrhage; Hypoxia; Sepsis; Tachycardia; Dizziness; Brain Injury; Syncope/Fainting; Heart Failure/Congestive Heart Failure | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING VENTRICULAR TACHYCARDIA (VT) ABLATION IN PATIENTS WITH A LEFT VENTRICULAR ASSIST DEVICE (VAD). PATIENTS WITH A LEFT VAD UNDERWENT VT ABLATION AFTER PRESENTING TO THE EMERGENCY DEPARTMENT WITH A SHOCK FROM THEIR IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD), A VAD ALARM, SYNCOPE, LIGHTHEADEDNESS, PALPITATIONS, OR WORSENING HEART FAILURE. THE AUTHORS DESCRIBED SOME PATIENT DEATHS WITHIN THIRTY DAYS OF THE ABLATION PROCEDURES. ONE PATIENT DIED FROM PROCEDURE-ASSOCIATED HYPOXIC BRAIN DAMAGE AS A RESULT OF POSTPROCEDURAL HYPOXEMIA. THE OTHER DEATHS OCCURRED IN THE FOLLOW-UP PERIOD; THE CAUSES OF DEATH WERE HEART FAILURE, UNKNOWN VAD FAILURE, SEPSIS, INTRACRANIAL BLEEDING, OR OTHER UNKNOWN CAUSES. ELECTROMAGNETIC INTERFERENCE OCCURRED DURING MAPPING WHICH RESULTED IN INCOMPLETE MAP ACQUISITION, TRANSIENT CATHETER VISUALIZATION LOSS, AND ONE CASE LED TO PROCEDURE ABORTION. PROCEDURE-RELATED COMPLICATIONS INCLUDED ACCESS SITE¿RELATED COMPLICATIONS, ONE PERICARDIAL EFFUSION OCCURRED, AND ONE IATROGENIC AORTIC DISSECTION BETWEEN THE RIGHT ILIAC ARTERY AND THE RENAL BRANCHES. THERE WERE PATIENTS WHO EXPERIENCED RECURRENCE OF VT OR ELECTRICAL STORM, AND THREE PATIENTS HAD A REPEAT ABLATION PROCEDURE. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/61 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: CHARACTERIZATION OF VENTRICULAR TACHYCARDIA ABLATION IN END-STAGE HEART FAILURE PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICE (CHANNELED REGISTRY). EUROPACE. 2025. 27, EUAF054. DOI: 10.1093/EUROPACE/EUAF054. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22238237&pc=NPT | 2025587-2025-04285 | 2025-04-11 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-17 04:00 | NPT | EVOLUT FX VALVE | Insufficient Information | Unspecified Infection; Septic Shock | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH FOLLOWING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE, THE PATIENT DIED DUE TO A N ON-CARDIOVASCULAR INFECTION AND SEPTIC SHOCK. PER THE PHYSICIAN, THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND DEVICE OR IMPLANT PROCEDURE. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22239547&pc=DSQ | 2916596-2025-03663 | 2025-05-12 04:00 | Death | THORATEC CORPORATION | 2025-06-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure; Thromboembolism; Multiple Organ Failure; Gastrointestinal Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTISYSTEM ORGAN FAILURE (MSOF). THE PATIENT'S BILIRUBIN WAS UP PRE-IMPLANT. THE PATIENT UNDERWENT A MIDLINE STERNOTOMY, AND A CUT AND SEW IN ORDER TO REPAIR THEIR MITRAL VALVE. PREOPERATIVELY, THE PATIENT'S HEART RATE WAS 80 BEATS PER MINUTE (BPM), BLOOD PRESSURE WAS 83/70, PULMONARY ARTERY PRESSURE WAS 64/28, AND CENTRAL VENOUS PRESSURE WAS 13. THEY WERE ALSO ON IMPELLA PREOPERATIVELY. AFTER THE OPERATION, THE PATIENT'S HEART RATE WAS 80, THEIR BLOOD PRESSURE WAS 98/78, PA 62.38, CENTRAL VENOUS PRESSURE WAS 17. THEIR VAD SPEED WAS 5200, FLOW WAS 4.3 LITERS PER MINUTE, PULSATILITY INDEX WAS 2.5, AND POWER WAS 3.6 W. THE RIGHT VENTRICULAR ASSIST DEVICE (RVAD) WAS 4.4 LPM, AND SPEED WAS 2190. THE PATIENT HAD GASTROINTESTINAL BLEEDING ON (B)(6) 2025 WITH MULTIPLE PRODUCTS, EMBOLISM IN RIGHT COLIC ARTERY. ON (B)(6) 2025, THE PATIENT HAD A WHITE OUT LEFT LUNG RELATED TO MUCUS PLUG, THEY WERE REINTUBATED/BRONCHOSCOPY. THEY HAD DUODENAL ULCERS, ISCHEMIC COLITIS. THEY HAD A TRACHEOSTOMY ON (B)(6) 2025. THEY ALSO RECEIVED CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT). RVAD WAS IMPLEMENTED DUE TO WORSENING RIGHT VENTRICLE FUNCTION. THEY HAD ALVEOLAR HEMORRHAGE, BRONCHOSCOPY, AND CLOT EVACUATION. DUE TO MSOF, THEY WERE TRANSITIONED TO COMFORT CARE. IT WAS REPORTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: ADDITIONAL H6 HEALTH EFFECT CIDES: RESPIRATORY SYSTEM E07: RESPIRATORY FAILURE E0742 HEPATIC AND BILIARY SYSTEM E11: HYPERBILIRUBINEMIA E1103 GASTROINTESTINAL SYSTEM E10: COLITIS E1036 VASCULAR SYSTEM E05: HEMORRHAGE/BLOOD LOSS/BLEEDING E0506. GENERALIZED DISORDERS E23: ULCERS E2339. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION E1: REPORTER EMAIL WAS NOT AVAILABLE. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6), AND THE REPORTED EVENTS AND PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. IT WAS REPORTED THAT THE PATIENT¿S OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WILL NOT BE EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED THROUGH THE MANUFACTURING EXECUTION SYSTEM (MES) AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS DEATH, BLEEDING, VENOUS THROMBOEMBOLISM, AND ARTERIAL NON-CENTRAL NERVOUS SYSTEM (CNS) THROMBOEMBOLISM AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. THE IFU ALSO LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RIGHT HEART FAILURE, RESPIRATORY FAILURE, RENAL DYSFUNCTION, AND HEPATIC DYSFUNCTION) AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION THERAPY AND INR (INTERNATIONAL NORMALIZED RATIO) RANGE, AS WELL AS CONSIDERATIONS FOR WHEN THERE IS A RISK OF BLEEDING. THIS SECTION ALSO OUTLINES INDICATIONS OF THROMBUS AS WELL AS HOW TO RESPOND TO SUCH EVENTS. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22241521&pc=DSQ | 3007042319-2025-01287 | 2022-01-01 05:00 | Death | HEARTWARE, INC. | 2025-06-17 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100047 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE IMPACT OF DIGOXIN USE FOLLOWING LEFT VENTRICULAR ASSIST DEVICE (VAD) IMPLANTATION. THE AUTHORS DESCRIBED DEATHS IN PATIENTS ON DIGOXIN AND THOSE NOT ON DIGOXIN; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN. THERE WERE PATIENTS WHO EXPERIENCED WORSENING RENAL FUNCTION, HYPERKALEMIA, GASTROINTESTINAL BLEED (GIB), AND RIGHT VENTRICULAR (RV) FAILURE. GIB WAS DEFINED WITH CLINICAL EVIDENCE OF BLEEDING (POSITIVE FECAL OCCULT BLOOD TESTING, MELENA, HEMATOCHEZIA, HEMATEMESIS AND/OR THE PRESENCE OF BLOOD IN THE GASTROINTESTINAL TRACT ON ENDOSCOPIC EVALUATION, VIDEO CAPSULE ENDOSCOPY, OR BLEEDING SCAN) ACCOMPANIED BY A DECREASE IN HEMOGLOBIN. RV FAILURE WAS DEFINED AS EVIDENCE OF ELEVATED CENTRAL VENOUS PRESSURE (CVP) ALONG WITH CLINICAL MANIFESTATION SUCH AS PERIPHERAL EDEMA, ASCITES, WORSENING RENAL, OR HEPATIC DYSFUNCTION. THE STATUS OF THE VADS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. OF NOTE, MULTIPLE PATIENTS AND MULTIPLE MANUFACTURERS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE BASELINE GENDER/AGE CHARACTERISTICS IS MALE/57 YEARS OLD. THE MODEL LISTED IN THE REPORT IS A REPRESENTATIVE OF THE MODEL FAMILY, AS THERE IS NO SPECIFIC MODEL LISTED. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: IMPACT OF DIGOXIN UTILIZATION ON CLINICAL OUTCOMES FOLLOWING LEFT VENTRICULAR ASSIST DEVICE IMPLANTATION. THE INTERNATIONAL JOURNAL OF ARTIFICIAL ORGANS. 2022. VOL. 45(11) 919¿926. DOI: 10.1177/03913988221112684. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22242485&pc=QJP | 2029214-2025-01410 | 2025-05-15 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-06-17 04:00 | QJP | PHENOM CATHETER | Physical Resistance/Sticking | Hemorrhage/Bleeding; Intracranial Hemorrhage; High Blood Pressure/ Hypertension; Rupture | K151638 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN INITIAL IMPLANT PROCEDURE USING THE PIPELINE EMBOLIZATION DEVICE 4.00MM X 30MM FOR MULTIPLE UNRUPTURED, SACCULAR ANEURYSMS LOCATED IN THE POSTERIOR COMMUNICATING SEGMENT OF THE LEFT INTERNAL CAROTID ARTERY (ICA), SEVERAL COMPLICATIONS OCCURRED. THE ANEURYSMS INCLUDED AN UPPER POSTERIOR COMMUNICATING ARTERY ANEURYSM MEASURING 7.1MM IN MAXIMUM DIAMETER WITH A 5.8MM NECK, AND ANOTHER ANEURYSM APPROXIMATELY 5MM IN SIZE BELOW. THE DISTAL LANDING ZONE WAS 3.3MM AND THE PROXIMAL LANDING ZONE WAS 3.9MM. THE VESSEL EXHIBITED SEVERE TORTUOSITY. FOLLOWING FEMORAL ARTERY PUNCTURE, A 6F 90CM LONG SHEATH, 6F 115CM TETHYS GUIDE CATHETER, PHENOM 27 MICROCATHETER, AND A 0.014 INCH 200CM GUIDEWIRE WERE USED TO ACCESS THE TARGET SITE. THE DEVICES WERE PREPARED ACCORDING TO THE INSTRUCTIONS FOR USE (IFU). 3D ROTATIONAL ANGIOGRAPHY WAS PERFORMED TO GUIDE DEVICE SELECTION AND PLACEMENT. THE STENT WAS HYDRATED AND DELIVERED, WITH THE TIP OPENED AT THE M1 SEGMENT OF THE MIDDLE CEREBRAL ARTERY. DURING PULL-BACK POSITIONING, THE STENT COULD NOT BE ANCHORED DUE TO A SHORT ANCHORING DISTANCE, RESULTING IN STENT RETRACTION. WHEN ATTEMPTING TO RETRIEVE THE STENT FOR REPOSITIONING, EXCESSIVE RESISTANCE WAS ENCOUNTERED. SUBSEQUENTLY, A RAPID RISE IN BLOOD PRESSURE WAS OBSERVED, AND ANGIOGRAPHY REVEALED RUPTURE AND BLEEDING OF THE INTERNAL CAROTID ARTERY WITH SUBARACHNOID HEMORRHAGE. PUPILLARY DILATION WAS NOTED ON THE OPERATING TABLE. CRANIOTOMY WAS PERFORMED AS A RESCUE INTERVENTION, BUT THE RESCUE WAS UNSUCCESSFUL AND THE PATIENT DIED. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE PROBLEM OCCURRED THAT DAY AND THE PATIENT DIED. THE PATIENT PASSED AWAY AT (B)(6) HOSPITAL, DOCTOR INFORMED THAT THE PATIENT DIED IN THE HOSPITAL. NO AUTOPSY RECORDS WILL BE AVAILABLE. THE PATIENT'S DEATH MAY BE RELATED TO THE PIPELINE SHIELD, AND IT MAY ALSO BE RELATED TO THE OPERATOR'S OPERATION. THERE WAS INTRACRANIAL ANEURYSM LEADING UP TO THE PATIENT'S DEATH. THE CAUSE OF THE RESISTANCE DURING RETRIEVAL MAY BE VASCULAR TORTUOSITY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22242870&pc=DSQ | 2916596-2025-03782 | 2024-12-24 05:00 | Death | THORATEC CORPORATION | 2025-06-17 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO STROKE. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. THE DEVICE WOULD NOT BE EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22243146&pc=DYB | 2015691-2025-04557 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hematoma | K200258 | E2016006 | 3 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 3 AORTIC BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 65 - 79 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 3 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 73.67 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4) AND (B)(4), THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED. Manufacturer Narrative: AFTER ADDITIONAL REVIEW, A CORRECTION IS REQUIRED TO THE "TIME TO EVENT" STATEMENT IN PREVIOUS H11. THE CORRECTED STATEMENT IS AS FOLLOWS: THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0 DAY(S).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22243148&pc=NPT | 2015691-2025-04553 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 AORTIC ANNULAR RUPTURE DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 74 - 74 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: AFTER ADDITIONAL REVIEW, A CORRECTION IS REQUIRED TO THE "TIME TO EVENT" STATEMENT IN PREVIOUS H11. THE CORRECTED STATEMENT IS AS FOLLOWS: THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0 DAY(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 ANNULAR RUPTURE DEATH EVENT(S) FOR THE SAPIEN 3 IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 74 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU) CARDIOVASCULAR INJURIES, SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES, AND HEMATOMA ARE A KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO THE THV TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, HEMATOMA, OR ANNULAR RUPTURE DURING THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE INCLUDE SIGNIFICANT THV OVERSIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION AND NARROW CALCIFIED STJ. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN VALVE RELIES ON NATIVE VALVE CALCIUM TO SECURELY ANCHOR TO THE ANNULUS. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22243154&pc=NPT | 2015691-2025-04560 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 AORTIC CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 74 - 74 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: AFTER ADDITIONAL REVIEW, A CORRECTION IS REQUIRED TO THE "TIME TO EVENT" STATEMENT IN PREVIOUS H11. THE CORRECTED STATEMENT IS AS FOLLOWS: THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0 DAY(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 74 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, CARDIOVASCULAR INJURIES SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR PERFORATION DURING A TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22243211&pc=NPT | 2015691-2025-04552 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 2 AORTIC CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 74 - 79 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: AFTER ADDITIONAL REVIEW, A CORRECTION IS REQUIRED TO THE "TIME TO EVENT" STATEMENT IN PREVIOUS H11. THE CORRECTED STATEMENT IS AS FOLLOWS: THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0 DAY(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 76.5 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST REOPERATION, INCLUDING EMERGENCY CARDIAC SURGERY, AS ADDITIONAL POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THIS EVENT TYPE (CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED) IS DEFINED AS: THE PATIENT SUBSEQUENTLY UNDERWENT OTHER TYPES OF CARDIAC SURGERY OR CATH LAB INTERVENTIONS (NOT OTHERWISE SPECIFIED) THAT WAS UNPLANNED. THIS COULD INCLUDE REOPERATION FOR DAMAGE TO ADJACENT STRUCTURES, SUCH AS THE MITRAL APPARATUS, OR POTENTIALLY VALVE EXPLANT RELATED TO EMBOLIZATION, OR OTHER EVENT TYPES. CARDIOVASCULAR INJURY, INCLUDING PERFORATION OR DISSECTION OF VALVULAR STRUCTURES (E.G., INJURY TO THE MITRAL VALVE/ APPARATUS), AND VALVE EMBOLIZATION REQUIRING INTERVENTION ARE KNOWN POTENTIAL COMPLICATIONS LISTED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU). CHORDAE TENDINEA RUPTURE IN TRANSCATHETER VALVE REPLACEMENT CAN OCCUR DURING ADVANCEMENT OF THE GUIDEWIRE, BAV CATHETER, OR DELIVERY SYSTEM AND IS MOST LIKELY TO OCCUR WITH ANTEGRADE APPROACHES, HOWEVER, CAN OCCUR WITH A RETROGRADE APPROACH. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVES (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, AND CAREFUL MANIPULATION OF DEVICES. PER THE THV TA TRAINING MANUALS, AFTER GAINING LV ACCESS WITH THE 0.035" SOFT GUIDEWIRE, THE WIRE SHOULD BE JIGGLED UNDER TEE IMAGING OF THE MITRAL VALVE. AN INCREASE IN MR SUGGESTS WIRE ENTANGLEMENT IN THE MITRAL SUB-VALVULAR APPARATUS. IF ENTANGLEMENT IS SUSPECTED, THE GUIDEWIRE SHOULD BE COMPLETELY REMOVED FROM THE VENTRICLE; THE OPERATOR SHOULD CHANGE DIRECTION OF THE NEEDLE AND REINSERT THE GUIDEWIRE INTO THE VENTRICLE, CHECKING AGAIN FOR WIRE ENTANGLEMENT. THE TF TRAINING MANUAL ALSO PROVIDES GUIDANCE FOR VALVE CROSSING AND WIRE EXCHANGE: EXCHANGE 0.035" AMPLATZ EXTRA-STIFF WIRE WITH PRE-SHAPED DISTAL END IN LEFT VENTRICLE. ENSURE GUIDE CATHETER IS ADVANCED UPON WIRE EXCHANGE TO ENSURE EXCHANGE WIRE DOES NOT GET CAUGHT IN THE MITRAL CHORDAE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO VENTRICULAR EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE/DELIVERY SYSTEM, A NARROW, CALCIFIED SINOTUBULAR JUNCTION, MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR VENTRICULAR EMBOLIZATION (I.E., SMALL, CALCIFIED STJ, MINIMAL LEAFLET CALCIFICATION), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. I IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22243220&pc=NPT | 2015691-2025-04549 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 8 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 8 AORTIC READMISSION (VALVE RELATED) DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 46 - 90 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 8 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 8 READMISSION (VALVE RELATED) DEATH EVENT(S) FOR THE SAPIEN 3 IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 73.38 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). VALVE RELATED RE-ADMISSION IN THE FOLLOW-UP PERIOD IS LIKELY DUE TO REOCCURRENCE OF SYMPTOMS. THIS MAY RESULT FROM REGURGITATION (CENTRAL OR PVL) OR PROGRESSION OF THE PRE-EXISTING DISEASE PROCESS AND MAY RESULT IN HEART FAILURE. CAUSES OF HEART FAILURE CAN BE MULTI-FACTORIAL. PER THE INSTRUCTION FOR USE (IFU), HEART FAILURE, VALVE REGURGITATION, AND VALVE DEGENERATION INCLUDING STENOSIS ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH AORTIC VALVE REPLACEMENT AND BIOPROSTHETIC HEART VALVES. CONGESTIVE HEART FAILURE CAN HAVE MULTIPLE ETIOLOGIES AND IS OFTEN DUE TO THE PROGRESSION OF UNDERLYING DISEASE PROCESSES, INCLUDING CORONARY ARTERY DISEASE, UNCONTROLLED HYPERTENSION, MYOCARDIAL INFARCTION, CARDIOMYOPATHY, FLUID OVERLOAD, OR VALVULAR DYSFUNCTION. RISK FACTORS THAT INCREASE A PATIENT'S RISK OF SUFFERING FROM CHF INCLUDE OBESITY, ADVANCED AGE, AND A HISTORY OF SMOKING. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO A NUMBER OF ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. VALVE STENOSIS MAY RESULT IN SYMPTOMS SUCH AS SOB AND DECREASED EXERCISE TOLERANCE, WHICH MAY BE ACCOMPANIED BY AN INCREASED GRADIENT ACROSS THE VALVE. THIS COULD BE DUE TO EARLY CALCIFICATION OF THE LEAFLETS, HOST TISSUE OVERGROWTH, OR IN RARE CASES A NON-FUNCTIONING LEAFLET. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED. Manufacturer Narrative: AFTER ADDITIONAL REVIEW, A CORRECTION IS REQUIRED TO THE "TIME TO EVENT" STATEMENT IN PREVIOUS H11. THE CORRECTED STATEMENT IS AS FOLLOWS: THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 152.88 DAY(S).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22243243&pc=NPT | 2015691-2025-04550 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P140031 | E2016006 | 5 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 5 AORTIC STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 72 - 79 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 5 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 5 STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 75.8 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS INCLUDING STROKE ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH INVASIVE CARDIAC PROCEDURE INCLUDING TRANSCATHETER HEART VALVE PROCEDURES AND THE USE OF THE EDWARDS THV DEVICES. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING TRANSCATHETER VALVE REPLACEMENT PROCEDURE. RISK FACTORS CORRELATING WITH A NUMBER OF PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE DEMONSTRATED THAT THE MAJORITY OF PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO IMPORTANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND AVR PATIENTS. AFTER TRANSCATHETER VALVE REPLACEMENT PROCEDURE, THERE APPEARS TO BE A MORE SIGNIFICANT PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT TRANSCATHETER VALVE REPLACEMENT PROCEDURE PATIENTS WITH MULTIPLE CO MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE HOWEVER, THE EVENT MAY BE RELATED TO THE MECHANISM DESCRIBED ABOVE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED. Manufacturer Narrative: AFTER ADDITIONAL REVIEW, A CORRECTION IS REQUIRED TO THE "TIME TO EVENT" STATEMENT IN PREVIOUS H11. THE CORRECTED STATEMENT IS AS FOLLOWS: THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 4.8 DAY(S).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22243259&pc=NPT | 2015691-2025-04555 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 AORTIC STROKE - UNDETERMINED DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 70 - 70 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 1 STROKE - UNDETERMINED DEATH EVENT(S) FOR THE SAPIEN 3 IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 70 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS INCLUDING STROKE ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH INVASIVE CARDIAC PROCEDURE INCLUDING TRANSCATHETER HEART VALVE PROCEDURES AND THE USE OF THE EDWARDS THV DEVICES. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING TRANSCATHETER VALVE REPLACEMENT PROCEDURE. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING TRANSCATHETER VALVE REPLACEMENT DEMONSTRATED THAT THE MAJORITY OF PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO IMPORTANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN TRANSCATHETER VALVE REPLACEMENT AND AVR PATIENTS. AFTER TAVR, THERE APPEARS TO BE A MORE SIGNIFICANT PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT TRANSCATHETER VALVE REPLACEMENT PROCEDURE PATIENTS WITH MULTIPLE CO MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THE EVENT IS LISTED AS UNDETERMINED STROKE. IT IS NOT KNOWN IF THE MECHANISM OF THE STROKE IS IN QUESTION, OR IF THE STROKE ITSELF IS IN QUESTION. THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. IT IS POSSIBLE THAT THE MECHANISM DESCRIBED ABOVE MAY HAVE CONTRIBUTED TO THE EVENT. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED. Manufacturer Narrative: AFTER ADDITIONAL REVIEW, A CORRECTION IS REQUIRED TO THE "TIME TO EVENT" STATEMENT IN PREVIOUS H11. THE CORRECTED STATEMENT IS AS FOLLOWS: THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 14 DAY(S).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22243302&pc=DYB | 2015691-2025-04562 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | K200258 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 MITRAL VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 73 - 73 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: AFTER ADDITIONAL REVIEW, A CORRECTION IS REQUIRED TO THE "TIME TO EVENT" STATEMENT IN PREVIOUS H11. THE CORRECTED STATEMENT IS AS FOLLOWS: THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 1 DAY(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 73 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22243303&pc=NPT | 2015691-2025-04554 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Insufficient Information | Aortic Valve Insufficiency/ Regurgitation | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 AORTIC REINTERVENTION - AORTIC VALVE DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 62 - 62 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: AFTER ADDITIONAL REVIEW, A CORRECTION IS REQUIRED TO THE "TIME TO EVENT" STATEMENT IN PREVIOUS H11. THE CORRECTED STATEMENT IS AS FOLLOWS: THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 255 DAY(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 REINTERVENTION - AORTIC VALVE DEATH EVENT(S) FOR THE SAPIEN 3 IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 62 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). INTRA-PROCEDURAL AORTIC VALVE RE-INTERVENTION WILL TYPICALLY RESULT FROM VALVE MALPOSITION OR REGURGITATION (PVL OR CENTRAL, INCLUDING LEAFLET RESTRICTION). THESE CONDITIONS ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO AORTIC MALPOSITION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, PRESERVED EJECTION FRACTION, SIGNIFICANT MITRAL ANNULAR CALCIFICATION (MAC), LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. DEPLOYMENT OF THE SAPIEN VALVE TOO AORTIC HAS THE POTENTIAL TO CONTRIBUTE TO SUBOPTIMAL COAPTATION OF THE SAPIEN VALVE LEAFLETS AND CAUSE CENTRAL AORTIC INSUFFICIENCY; IT CAN OBSTRUCT THE CORONARY OSTIA; AND LEAD TO EMBOLIZATION OF THE PROSTHESIS INTO THE ASCENDING AORTA. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR AORTIC MALPOSITION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. AORTIC VALVE RE-INTERVENTION IN THE FOLLOW-UP PERIOD (IN THE ABSENCE OF PROSTHETIC CARDIAC VALVE THROMBOSIS) WILL TYPICALLY RESULT FROM ON-GOING OR WORSENING REGURGITATION, VALVE DEGENERATION RELATED TO THE FORMATION OF CALCIFICATION OR PANNUS, OR VALVE MIGRATION. THESE EVENTS ARE IDENTIFIED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU) AS POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO A NUMBER OF ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. CALCIFICATION IS A WELL-RECOGNIZED FAILURE MODE OF BIOPROSTHETIC VALVES. THE MECHANISM OF CALCIFICATION OF BIOMATERIALS IS NOT COMPLETELY UNDERSTOOD BUT IS PROBABLY RELATED TO AN INABILITY OF THE NON-VIABLE CELLS TO MAINTAIN THEIR NORMALLY LOW INTRACELLULAR CONCENTRATION OF CALCIUM. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. ACCORDING TO LITERATURE REVIEW, VALVE MIGRATION RESULTS WHEN FORCES ACTING ON THE TRANSCATHETER HEART VALVE (THV) OVERCOME THE STRENGTH OF ATTACHMENT OF THE VALVE TO THE AORTIC WALL. STENT VALVES ARE SUBJECTED TO ANTEGRADE EJECTION FORCES DURING SYSTOLE. LESS-THAN-SEVERE AND NON-UNIFORMLY DISTRIBUTED CALCIFICATION OF THE NATIVE LEAFLETS, INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION CAN CONTRIBUTE TO VALVE MIGRATION. ADDITIONALLY, RESIDUAL OVERHANGING LEAFLETS CAN EXERT DOWNWARDS FORCE DURING DIASTOLE, CAUSING MIGRATION OF THE THV TOWARDS THE LEFT VENTRICLE IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22243306&pc=NPU | 2015691-2025-04561 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPU | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 MITRAL STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 67 - 67 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: AFTER ADDITIONAL REVIEW, A CORRECTION IS REQUIRED TO THE "TIME TO EVENT" STATEMENT IN PREVIOUS H11. THE CORRECTED STATEMENT IS AS FOLLOWS: THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 9 DAY(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 67 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS INCLUDING STROKE ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND THE USE OF THV DEVICES. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY BY EDWARDS LIFESCIENCES, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING TRANSCATHETER VALVE REPLACEMENT. RISK FACTORS CORRELATING WITH A NUMBER OF PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURES ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING TRANSCATHETER VALVE REPLACEMENT DEMONSTRATED THAT THE MAJORITY OF PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO IMPORTANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN TRANSCATHETER VALVE REPLACEMENT AND AVR PATIENTS. AFTER TRANSCATHETER VALVE REPLACEMENT, THERE APPEARS TO BE A MORE SIGNIFICANT PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT TRANSCATHETER VALVE REPLACEMENT PATIENTS WITH MULTIPLE CO MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22243332&pc=DYB | 2015691-2025-04556 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | K200258 | E2016006 | 8 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 8 AORTIC VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 62 - 90 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 7 MALE(S). Manufacturer Narrative: AFTER ADDITIONAL REVIEW, A CORRECTION IS REQUIRED TO THE "TIME TO EVENT" STATEMENT IN PREVIOUS H11. THE CORRECTED STATEMENT IS AS FOLLOWS: THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 64.38 DAY(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 8 VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 79 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22243613&pc=DYB | 2015691-2025-04558 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K200258 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 2 AORTIC BLEEDING - ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 79 - 86 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: AFTER ADDITIONAL REVIEW, A CORRECTION IS REQUIRED TO THE "TIME TO EVENT" STATEMENT IN PREVIOUS H11. THE CORRECTED STATEMENT IS AS FOLLOWS: THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0 DAY(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 BLEEDING - ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 82.5 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22243704&pc=NPT | 2015691-2025-04551 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Heart Block | P140031 | E2016006 | 3 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 3 AORTIC PERMANENT PACEMAKER DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 82 - 88 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: AFTER ADDITIONAL REVIEW, A CORRECTION IS REQUIRED TO THE "TIME TO EVENT" STATEMENT IN PREVIOUS H11. THE CORRECTED STATEMENT IS AS FOLLOWS: THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 2.33 DAY(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 PERMANENT PACEMAKER DEATH EVENT(S) FOR THE SAPIEN 3 IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 84.67 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, CONDUCTION SYSTEM DEFECTS (HEART BLOCK) ARRHYTHMIAS AND CONDUCTION SYSTEM DEFECTS THAT MAY REQUIRE A PERMANENT PACEMAKER ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH BALLOON AORTIC VALVULOPLASTY, DEPLOYMENT OF THE PROSTHETIC VALVE, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC) GUIDELINES, THE CLOSE ANATOMICAL RELATIONSHIP BETWEEN THE AORTIC VALVE COMPLEX AND THE BRANCHING ATRIOVENTRICULAR BUNDLE MAY PROVIDE AN EXPLANATION FOR THESE COMPLICATIONS OF THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY -CONDUCTION DISTURBANCES/ HEART BLOCK WRITTEN BY EDWARDS LIFESCIENCES, ATRIOVENTRICULAR CONDUCTION DISTURBANCES AFTER THV PLACEMENT ARE ASSOCIATED WITH MANY PATIENT RELATED AND PROCEDURAL RELATED FACTORS, INCLUDING PRE-OPERATIVE CO-MORBID STATUS, THE DEGREE AND BULKINESS OF AORTIC VALVE AND ANNULAR CALCIFICATION, INTER-VENTRICULAR SEPTAL THICKNESS, PRE-EXISTING ELECTROCARDIOGRAM ABNORMALITIES, THE DEPTH OF PROSTHESIS IMPLANTATION, AND THE PROFILE OF THE IMPLANTED PROSTHESIS. UNLIKE CONVENTIONAL AVR, WHERE THERE MAY BE LOCALIZED TRAUMA DUE TO DECALCIFICATION OF THE ANNULUS AND/OR SUTURE PLACEMENT IN THE PROXIMITY OF THE AV NODE OR THE BUNDLES, TRANSCATHETER VALVE REPLACEMENT MAY CAUSE CONDUCTION ABNORMALITIES THROUGH MECHANICAL IMPINGEMENT OF THE CONDUCTION SYSTEM BY THE PROSTHESIS. THE MECHANISMS OF THE DEVELOPMENT OF HEART BLOCK AFTER THV PLACEMENT ARE WELL DOCUMENTED AND DESCRIBED IN THE LITERATURE. IT IS ALSO DOCUMENTED THAT PRE-EXISTING HEART BLOCK IS COMMON IN PATIENTS UNDERGOING THV PLACEMENT OR SURGICAL AVR AND ANOTHER 4-6 % WILL DEVELOP POSTOPERATIVE HEART BLOCK, POTENTIALLY REQUIRING A PERMANENT PACEMAKER. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE, HOWEVER THE CONDUCTION DISTURBANCE IS LIKELY RELATED TO THE MECHANISM DESCRIBED ABOVE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22243790&pc=DYB | 2015691-2025-04559 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Insufficient Information | Unspecified Vascular Problem | K200258 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 2 AORTIC VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 77 - 79 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 78 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED. Manufacturer Narrative: AFTER ADDITIONAL REVIEW, A CORRECTION IS REQUIRED TO THE "TIME TO EVENT" STATEMENT IN PREVIOUS H11. THE CORRECTED STATEMENT IS AS FOLLOWS: THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0.5 DAY(S).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22246846&pc=PNJ | 2017865-2025-87539 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-17 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT ATRIUM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150035 |  | 1 | Event Description: IT WAS REPORTED FROM A STUDY ARTICLE THAT A CERTAIN GROUP OF PATIENTS FROM THE STUDY EXPERIENCED COMPLICATIONS AFTER OBTAINING AN ATRIAL LEADLESS PACEMAKER IMPLANT. THE COMPLICATIONS INCLUDED VENOUS THROMBOEMBOLISM, PUNCTURE SITE EVENTS, PERFORATION, DISLODGEMENT, INFECTION AND PERICARDITIS. THE PATIENT'S CONDITIONS WERE UNKNOWN. IT WAS ALSO REPORTED THAT A PERCENTAGE OF PATIENT'S PASSED AWAY DURING THE STUDY PERIOD AFTER IMPLANT. IT IS UNKNOWN IF THE PATIENT'S DEATHS WERE RELATED TO THE ATRIAL LEADLESS PACEMAKER AND/OR PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22247751&pc=BRY | 2016493-2025-87757 | 2025-05-22 04:00 | Death | CAREFUSION 303, INC. | 2025-06-17 04:00 | BRY | BD PYXIS¿ MEDSTATION ES AUXILIARY | Failure to Deliver | Cardiac Arrest | EXEMPT |  | 1 | Event Description: IT WAS REPORTED THAT DURING A POWER/INTERNET OUTAGE AT THE WILLIAM BEAUMONT ARMY MEDICAL CENTER, A NURSE AND A PHARMACIST COULD NOT LOG IN TO A BD PYXIS¿ MEDSTATION ES AND OPEN A DRAWER. THIS OCCURRED DURING A CODE BLUE IN THE EMERGENCY DEPARTMENT. IT WAS REPORTED THAT THE PATIENT UNFORTUNATELY DIED. MULTIPLE ATTEMPTS HAVE BEEN MADE (6 TOTAL) TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER FOR THIS EVENT. TO DATE, BD HAS NOT RECEIVED ANY ADDITIONAL INFORMATION. IF ADDITIONAL INFORMATION IS RECEIVED, THIS CASE WILL BE RE-EVALUATED. OUT OF AN ABUNDANCE OF CAUTION, THIS EVENT HAS BEEN DETERMINED TO BE REPORTABLE FOR DEATH.Event Description: IT WAS REPORTED THAT DURING A POWER/INTERNET OUTAGE AT THE (B)(6), A NURSE AND A PHARMACIST COULD NOT LOG IN TO A BD PYXIS¿ MEDSTATION ES AND OPEN A DRAWER. THIS OCCURRED DURING A CODE BLUE IN THE EMERGENCY DEPARTMENT. IT WAS REPORTED THAT THE PATIENT UNFORTUNATELY DIED. MULTIPLE ATTEMPTS HAVE BEEN MADE (6 TOTAL) TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER FOR THIS EVENT. TO DATE, BD HAS NOT RECEIVED ANY ADDITIONAL INFORMATION. IF ADDITIONAL INFORMATION IS RECEIVED, THIS CASE WILL BE RE-EVALUATED. OUT OF AN ABUNDANCE OF CAUTION, THIS EVENT HAS BEEN DETERMINED TO BE REPORTABLE FOR DEATH. Manufacturer Narrative: A DEVICE EVALUATION IS ANTICIPATED, BUT HAS NOT YET BEGUN. UPON COMPLETION OF THE INVESTIGATION, A SUPPLEMENTAL REPORT WILL BE FILED. Manufacturer Narrative: A REVIEW OF THE COMPLAINT HISTORY FOR SN: (B)(6) WAS PERFORMED IN SALESFORCE WHICH DID NOT LOCATE SIMILAR COMPLAINT(S) WITH THE SAME FAILURE MODE FOR THIS SERIAL NUMBER. A REVIEW OF THE DEVICE HISTORY RECORD FOR SN: (B)(6) WAS PERFORMED FROM THE DATE OF MANUFACTURE, 11-JAN-2022 AND CONFIRMED THAT THIS DEVICE WAS NOT PREVIOUSLY RETURNED FOR SERVICING AND THERE WERE NO PRODUCTION FAILURES WHICH CORRELATES TO THE CUSTOMER REPORTED ISSUE. UPON INVESTIGATION OF THE ACTUAL DEVICE USED IN THIS INCIDENT, IT WAS DETERMINED THAT POWER OUTAGE AND NETWORK FAILURE PREVENTED SOME STATIONS FROM ENTERING CRITICAL OVERRIDE. AS A RESULT, USERS WERE UNABLE TO LOG IN OR ACCESS MEDICATIONS, AND THERE WAS A CODE BLUE SITUATION, AND THE STAFF WERE UNABLE TO ACCESS THE STATION OR PULL MEDS. THE CUSTOMER TRIED ANOTHER TWO STATIONS NEARBY, BUT BOTH ACTED THE SAME. THE CUSTOMER ATTEMPTED TO OBTAIN THE MEDICATION FROM THE PHARMACY, BUT UNFORTUNATELY, THE PATIENT PASSED AWAY. AFTER THE PATIENT HAD PASSED AWAY, THEY ATTEMPTED TO LOG ON AGAIN AND TIMED HOW LONG IT TOOK¿TWENTY-FOUR SECONDS TO LOG IN AND NEARLY THIRTY SECONDS TO BRING UP THE PATIENT PROFILE. A FIELD SERVICE ENGINEER ARRIVED ON SITE AND WAS INFORMED OF A PATIENT DEATH THAT OCCURRED DURING AN UNSCHEDULED DOWNTIME. VERIFIED THE SERIAL NUMBER AND NO SERVICE WAS PERFORMED ON THE STATION. THE ENGINEER REMAINED ON SITE AWAITING FURTHER INSTRUCTIONS AND DEPARTED ONLY AFTER RECEIVING AUTHORIZATION. THE WORK ORDER WAS CLOSED WITH NO REPAIRS MADE. THE CUSTOMER REPORTED THAT THE STATION WAS FULLY FUNCTIONAL, AND NO FURTHER UPDATES WERE RECEIVED FROM CUSTOMER ADVOCACY. THE SYSTEM FUNCTIONED AS INTENDED AFTER THE ISSUE GOT RESOLVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22248079&pc=LOX | 2124215-2025-37666 | 2025-05-22 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-17 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT WAS ADMITTED FOR AN ELECTIVE CORONARY ANGIOGRAM FOLLOWING A POSITIVE CT CORONARY ANGIOGRAM (CTCA). THE PROCEDURE WAS INITIATED VIA RIGHT RADIAL ARTERY ACCESS, AND A 7F NON-BOSTON SCIENTIFIC RADIAL SHEATH WAS INSERTED. ANGIOGRAPHY REVEALED SEVERELY CALCIFIED LEFT ANTERIOR DESCENDING (LAD) ARTERY LESIONS, WHICH WAS ALSO NOTED TO BE MODERATELY TORTUOUS. A 6F GUIDEZILLA II GUIDE EXTENSION CATHETER WAS ENGAGED TO ASSIST DEVICE ADVANCEMENT. ROTATIONAL ATHERECTOMY WAS PERFORMED USING A 1.50 MM ROTAPRO DEVICE. INTRAVASCULAR ULTRASOUND (IVUS) WAS THEN PERFORMED FOLLOWED BY MULTIPLE BALLOON INFLATIONS WITH 3.00MM X 15MM, 2.75MM X 12MM AND 2.50MM X 12MM NC EMERGE BALLOON CATHETERS. THE BALLOON INFLATIONS WERE CARRIED OUT IN THE PROXIMAL AND MID-LAD SEGMENTS. DURING THIS PROCESS, HYPOTENSIVE EPISODES WERE OBSERVED AND MANAGED WITH MULTIPLE ARAMINE BOLUSES. A 2.50 X 32 MM SYNERGY XD DRUG-ELUTING STENT WAS THEN ADVANCED DOWN THE LAD. WHILE OPTIMAL POSITIONING WAS PERFORMED, THE STENT SHAFT WAS SHEARED, NOTED POSSIBLY DUE TO CONTACT WITH HEAVILY CALCIFIED PLAQUE. A PORTION OF THE SHAFT, INCLUDING THE UNDEPLOYED STENT AND STENT BALLOON, BECAME DETACHED AND LODGED WITHIN THE LAD, RESULTING IN VESSEL OCCLUSION. THE PATIENTS CONDITION BEGAN TO DETERIORATE, AND A CODE BLUE WAS ACTIVATED. A CARDIOTHORACIC SURGEON WAS URGENTLY CONTACTED, AND ADVANCED LIFE SUPPORT (ALS) MEASURES WERE INITIATED. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF) AND WAS DEFIBRILLATED WITH 200 JOULES. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, AND SINUS RHYTHM WAS RESTORED AFTER THE INITIAL SHOCK. PREPARATION BEGAN TO TRANSFER THE PATIENT TO CARDIAC SURGERY. THE PATIENT WAS INTUBATED AND THEN DETERIORATED BACK INTO VF, WHICH REQUIRED FOUR ADDITIONAL DEFIBRILLATION SHOCKS, THREE 1 MG DOSES OF ADRENALINE, AND 300 MG OF AMIODARONE, ADMINISTERED ACCORDING TO ALS PROTOCOL. A LUCAS MECHANICAL CPR DEVICE WAS APPLIED, AND RESUSCITATIVE EFFORTS WERE CONTINUED DURING TRANSFER TO THE CARDIAC SURGERY ROOM FOR URGENT CORONARY ARTERY BYPASS GRAFTING (CABG). AN EMERGENCY STERNOTOMY WAS PERFORMED, AND CARDIOPULMONARY BYPASS WAS INITIATED VIA FEMORAL CANNULATION. THE PATIENT UNDERWENT CABG TO RIGHT CORONARY ARTERY (RCA) X 2, LAD, AND DISTAL POSTERIOR DESCENDING ARTERY (PDA), WITH LEFT ATRIAL APPENDAGE (LAA) LIGATION ALSO COMPLETED. THE STENT, BALLOON, AND SHEARED SHAFT WERE SUCCESSFULLY RETRIEVED DURING SURGERY. THE PROCEDURE WAS COMPLETED. THE RETRIEVED STENT STRUT WAS ALSO NOTED TO BE DEFORMED. THE PATIENT WAS SUBSEQUENTLY TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). HOWEVER, DESPITE ALL INTERVENTIONS, THE PATIENT PASSED AWAY THE FOLLOWING DAY. Manufacturer Narrative: GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE PRODUCT STATUS AND ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22248125&pc=MVK | 3008642652-2025-05932 | 2025-04-14 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-17 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Signal Artifact/Noise | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE DEVICE WAS STARTED UP AT 09:34:39 ON (B)(6) 2025. AT 13:44:12, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH VARYING RATES/AMPLITUDES AND MOTION ARTIFACT. THE DEVICE PROPERLY DETECTED VF. VARYING RATES/AMPLITUDES AND THE LIFEVEST BEING SHUTDOWN PREVENTED THE LIFEVEST FROM TREATING THE PATIENT. THE DEVICE WAS SHUT DOWN AT 13:44:15 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22248162&pc=FGE | 3005099803-2025-02684 | 2025-02-23 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-17 04:00 | FGE | WALLFLEX BILIARY RX STENT SYSTEM | Use of Device Problem; Human-Device Interface Problem | Pancreatitis | K240464 |  | 1 | Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT A WALLFLEX FULLY COVERED BILIARY STENT RMV WAS IMPLANTED IN THE COMMON BILE DUCT OF THE LIVER TO TREAT A BILE LEAK DURING AN ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY PROCEDURE (ERCP) PERFORMED ON AN UNKNOWN DATE. POST STENT PLACEMENT, THE PATIENT DEVELOPED PANCREATITIS, AND THE PATIENT PASSED AWAY. IN THE PHYSICIAN'S ASSESSMENT, THE PATIENT DEVELOPED PANCREATITIS AS A POSSIBLE RESULT OF OCCLUSION TO THE PANCREATIC DUCT FROM THE WALLFLEX BILIARY STENT. NOTE: ACCORDING TO THE COMPLAINANT, THE WALLFLEX BILIARY FULLY COVERED STENT WAS USED TO TREAT A BILE LEAK. HOWEVER, PER THE WALLFLEX BILIARY RX FULLY COVERED STENT SYSTEM DIRECTIONS FOR USE, THE STENT IS INDICATED FOR USE IN THE PALLIATIVE TREATMENT OF BILIARY STRICTURES PRODUCED BY MALIGNANT NEOPLASMS, RELIEF OF MALIGNANT BILIARY OBSTRUCTION PRIOR TO SURGERY AND FOR TREATMENT OF BENIGN BILIARY STRICTURES. Manufacturer Narrative: BLOCK D4, H4: THE COMPLAINANT WAS UNABLE TO PROVIDE THE SUSPECT DEVICE LOT NUMBER. THEREFORE, THE MANUFACTURE AND EXPIRATION DATES ARE UNKNOWN. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK H6: IMDRF DEVICE CODE A22 CAPTURES THE REPORTABLE EVENT OF STENT OVERGROWTH. IMDRF PATIENT CODE E1021 CAPTURES THE REPORTABLE EVENT OF PANCREATITIS. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF THE PATIENT PASSED AWAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22248199&pc=MVK | 3008642652-2025-05931 | 2025-05-12 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-17 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 10:42:08 ON (B)(6) 2025. AT 17:50:21, AN ARRHYTHMIA WAS DETECTED. ECG WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 17:50:56 THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS SINUS RHYTHM @ 60 BPM WITH PAC¿S AND CPR/MOTION ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 18:22:22 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22248204&pc=NIQ | 2124215-2025-37627 | 2025-05-22 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-17 04:00 | NIQ | SYNERGY XD | Break; Material Deformation | Cardiac Arrest; Ventricular Fibrillation; Obstruction/Occlusion |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT WAS ADMITTED FOR AN ELECTIVE CORONARY ANGIOGRAM FOLLOWING A POSITIVE CT CORONARY ANGIOGRAM (CTCA). THE PROCEDURE WAS INITIATED VIA RIGHT RADIAL ARTERY ACCESS, AND A 7F NON-BOSTON SCIENTIFIC RADIAL SHEATH WAS INSERTED. ANGIOGRAPHY REVEALED SEVERELY CALCIFIED LEFT ANTERIOR DESCENDING (LAD) ARTERY LESIONS, WHICH WAS ALSO NOTED TO BE MODERATELY TORTUOUS. A 6F GUIDEZILLA II GUIDE EXTENSION CATHETER WAS ENGAGED TO ASSIST DEVICE ADVANCEMENT. ROTATIONAL ATHERECTOMY WAS PERFORMED USING A 1.50 MM ROTAPRO DEVICE. INTRAVASCULAR ULTRASOUND (IVUS) WAS THEN PERFORMED FOLLOWED BY MULTIPLE BALLOON INFLATIONS WITH 3.00MM X 15MM, 2.75MM X 12MM AND 2.50MM X 12MM NC EMERGE BALLOON CATHETERS. THE BALLOON INFLATIONS WERE CARRIED OUT IN THE PROXIMAL AND MID-LAD SEGMENTS. DURING THIS PROCESS, HYPOTENSIVE EPISODES WERE OBSERVED AND MANAGED WITH MULTIPLE ARAMINE BOLUSES. A 2.50 X 32 MM SYNERGY XD DRUG-ELUTING STENT WAS THEN ADVANCED DOWN THE LAD. WHILE OPTIMAL POSITIONING WAS PERFORMED, THE STENT SHAFT WAS SHEARED, NOTED POSSIBLY DUE TO CONTACT WITH HEAVILY CALCIFIED PLAQUE. A PORTION OF THE SHAFT, INCLUDING THE UNDEPLOYED STENT AND STENT BALLOON, BECAME DETACHED AND LODGED WITHIN THE LAD, RESULTING IN VESSEL OCCLUSION. THE PATIENTS CONDITION BEGAN TO DETERIORATE, AND A CODE BLUE WAS ACTIVATED. A CARDIOTHORACIC SURGEON WAS URGENTLY CONTACTED, AND ADVANCED LIFE SUPPORT (ALS) MEASURES WERE INITIATED. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF) AND WAS DEFIBRILLATED WITH 200 JOULES. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, AND SINUS RHYTHM WAS RESTORED AFTER THE INITIAL SHOCK. PREPARATION BEGAN TO TRANSFER THE PATIENT TO CARDIAC SURGERY. THE PATIENT WAS INTUBATED AND THEN DETERIORATED BACK INTO VF, WHICH REQUIRED FOUR ADDITIONAL DEFIBRILLATION SHOCKS, THREE 1 MG DOSES OF ADRENALINE, AND 300 MG OF AMIODARONE, ADMINISTERED ACCORDING TO ALS PROTOCOL. A LUCAS MECHANICAL CPR DEVICE WAS APPLIED, AND RESUSCITATIVE EFFORTS WERE CONTINUED DURING TRANSFER TO THE CARDIAC SURGERY ROOM FOR URGENT CORONARY ARTERY BYPASS GRAFTING (CABG). AN EMERGENCY STERNOTOMY WAS PERFORMED, AND CARDIOPULMONARY BYPASS WAS INITIATED VIA FEMORAL CANNULATION. THE PATIENT UNDERWENT CABG TO RIGHT CORONARY ARTERY (RCA) X 2, LAD, AND DISTAL POSTERIOR DESCENDING ARTERY (PDA), WITH LEFT ATRIAL APPENDAGE (LAA) LIGATION ALSO COMPLETED. THE STENT, BALLOON, AND SHEARED SHAFT WERE SUCCESSFULLY RETRIEVED DURING SURGERY. THE PROCEDURE WAS COMPLETED. THE RETRIEVED STENT STRUT WAS ALSO NOTED TO BE DEFORMED. THE PATIENT WAS SUBSEQUENTLY TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). HOWEVER, DESPITE ALL INTERVENTIONS, THE PATIENT PASSED AWAY THE FOLLOWING DAY. Manufacturer Narrative: GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22248214&pc=LOX | 2124215-2025-37684 | 2025-05-22 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-17 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT WAS ADMITTED FOR AN ELECTIVE CORONARY ANGIOGRAM FOLLOWING A POSITIVE CT CORONARY ANGIOGRAM (CTCA). THE PROCEDURE WAS INITIATED VIA RIGHT RADIAL ARTERY ACCESS, AND A 7F NON-BOSTON SCIENTIFIC RADIAL SHEATH WAS INSERTED. ANGIOGRAPHY REVEALED SEVERELY CALCIFIED LEFT ANTERIOR DESCENDING (LAD) ARTERY LESIONS, WHICH WAS ALSO NOTED TO BE MODERATELY TORTUOUS. A 6F GUIDEZILLA II GUIDE EXTENSION CATHETER WAS ENGAGED TO ASSIST DEVICE ADVANCEMENT. ROTATIONAL ATHERECTOMY WAS PERFORMED USING A 1.50 MM ROTAPRO DEVICE. INTRAVASCULAR ULTRASOUND (IVUS) WAS THEN PERFORMED FOLLOWED BY MULTIPLE BALLOON INFLATIONS WITH 3.00MM X 15MM, 2.75MM X 12MM AND 2.50MM X 12MM NC EMERGE BALLOON CATHETERS. THE BALLOON INFLATIONS WERE CARRIED OUT IN THE PROXIMAL AND MID-LAD SEGMENTS. DURING THIS PROCESS, HYPOTENSIVE EPISODES WERE OBSERVED AND MANAGED WITH MULTIPLE ARAMINE BOLUSES. A 2.50 X 32 MM SYNERGY XD DRUG-ELUTING STENT WAS THEN ADVANCED DOWN THE LAD. WHILE OPTIMAL POSITIONING WAS PERFORMED, THE STENT SHAFT WAS SHEARED, NOTED POSSIBLY DUE TO CONTACT WITH HEAVILY CALCIFIED PLAQUE. A PORTION OF THE SHAFT, INCLUDING THE UNDEPLOYED STENT AND STENT BALLOON, BECAME DETACHED AND LODGED WITHIN THE LAD, RESULTING IN VESSEL OCCLUSION. THE PATIENTS CONDITION BEGAN TO DETERIORATE, AND A CODE BLUE WAS ACTIVATED. A CARDIOTHORACIC SURGEON WAS URGENTLY CONTACTED, AND ADVANCED LIFE SUPPORT (ALS) MEASURES WERE INITIATED. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF) AND WAS DEFIBRILLATED WITH 200 JOULES. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, AND SINUS RHYTHM WAS RESTORED AFTER THE INITIAL SHOCK. PREPARATION BEGAN TO TRANSFER THE PATIENT TO CARDIAC SURGERY. THE PATIENT WAS INTUBATED AND THEN DETERIORATED BACK INTO VF, WHICH REQUIRED FOUR ADDITIONAL DEFIBRILLATION SHOCKS, THREE 1 MG DOSES OF ADRENALINE, AND 300 MG OF AMIODARONE, ADMINISTERED ACCORDING TO ALS PROTOCOL. A LUCAS MECHANICAL CPR DEVICE WAS APPLIED, AND RESUSCITATIVE EFFORTS WERE CONTINUED DURING TRANSFER TO THE CARDIAC SURGERY ROOM FOR URGENT CORONARY ARTERY BYPASS GRAFTING (CABG). AN EMERGENCY STERNOTOMY WAS PERFORMED, AND CARDIOPULMONARY BYPASS WAS INITIATED VIA FEMORAL CANNULATION. THE PATIENT UNDERWENT CABG TO RIGHT CORONARY ARTERY (RCA) X 2, LAD, AND DISTAL POSTERIOR DESCENDING ARTERY (PDA), WITH LEFT ATRIAL APPENDAGE (LAA) LIGATION ALSO COMPLETED. THE STENT, BALLOON, AND SHEARED SHAFT WERE SUCCESSFULLY RETRIEVED DURING SURGERY. THE PROCEDURE WAS COMPLETED. THE RETRIEVED STENT STRUT WAS ALSO NOTED TO BE DEFORMED. THE PATIENT WAS SUBSEQUENTLY TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). HOWEVER, DESPITE ALL INTERVENTIONS, THE PATIENT PASSED AWAY THE FOLLOWING DAY. Manufacturer Narrative: GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE THE PRODUCT STATUS AND ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22248283&pc=LOX | 2124215-2025-37656 | 2025-05-22 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-17 04:00 | LOX | NC EMERGE | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension |  |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT WAS ADMITTED FOR AN ELECTIVE CORONARY ANGIOGRAM FOLLOWING A POSITIVE CT CORONARY ANGIOGRAM (CTCA). THE PROCEDURE WAS INITIATED VIA RIGHT RADIAL ARTERY ACCESS, AND A 7F NON-BOSTON SCIENTIFIC RADIAL SHEATH WAS INSERTED. ANGIOGRAPHY REVEALED SEVERELY CALCIFIED LEFT ANTERIOR DESCENDING (LAD) ARTERY LESIONS, WHICH WAS ALSO NOTED TO BE MODERATELY TORTUOUS. A 6F GUIDEZILLA II GUIDE EXTENSION CATHETER WAS ENGAGED TO ASSIST DEVICE ADVANCEMENT. ROTATIONAL ATHERECTOMY WAS PERFORMED USING A 1.50 MM ROTAPRO DEVICE. INTRAVASCULAR ULTRASOUND (IVUS) WAS THEN PERFORMED FOLLOWED BY MULTIPLE BALLOON INFLATIONS WITH 3.00MM X 15MM, 2.75MM X 12MM AND 2.50MM X 12MM NC EMERGE BALLOON CATHETERS. THE BALLOON INFLATIONS WERE CARRIED OUT IN THE PROXIMAL AND MID-LAD SEGMENTS. DURING THIS PROCESS, HYPOTENSIVE EPISODES WERE OBSERVED AND MANAGED WITH MULTIPLE ARAMINE BOLUSES. A 2.50 X 32 MM SYNERGY XD DRUG-ELUTING STENT WAS THEN ADVANCED DOWN THE LAD. WHILE OPTIMAL POSITIONING WAS PERFORMED, THE STENT SHAFT WAS SHEARED, NOTED POSSIBLY DUE TO CONTACT WITH HEAVILY CALCIFIED PLAQUE. A PORTION OF THE SHAFT, INCLUDING THE UNDEPLOYED STENT AND STENT BALLOON, BECAME DETACHED AND LODGED WITHIN THE LAD, RESULTING IN VESSEL OCCLUSION. THE PATIENTS CONDITION BEGAN TO DETERIORATE, AND A CODE BLUE WAS ACTIVATED. A CARDIOTHORACIC SURGEON WAS URGENTLY CONTACTED, AND ADVANCED LIFE SUPPORT (ALS) MEASURES WERE INITIATED. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF) AND WAS DEFIBRILLATED WITH 200 JOULES. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, AND SINUS RHYTHM WAS RESTORED AFTER THE INITIAL SHOCK. PREPARATION BEGAN TO TRANSFER THE PATIENT TO CARDIAC SURGERY. THE PATIENT WAS INTUBATED AND THEN DETERIORATED BACK INTO VF, WHICH REQUIRED FOUR ADDITIONAL DEFIBRILLATION SHOCKS, THREE 1 MG DOSES OF ADRENALINE, AND 300 MG OF AMIODARONE, ADMINISTERED ACCORDING TO ALS PROTOCOL. A LUCAS MECHANICAL CPR DEVICE WAS APPLIED, AND RESUSCITATIVE EFFORTS WERE CONTINUED DURING TRANSFER TO THE CARDIAC SURGERY ROOM FOR URGENT CORONARY ARTERY BYPASS GRAFTING (CABG). AN EMERGENCY STERNOTOMY WAS PERFORMED, AND CARDIOPULMONARY BYPASS WAS INITIATED VIA FEMORAL CANNULATION. THE PATIENT UNDERWENT CABG TO RIGHT CORONARY ARTERY (RCA) X 2, LAD, AND DISTAL POSTERIOR DESCENDING ARTERY (PDA), WITH LEFT ATRIAL APPENDAGE (LAA) LIGATION ALSO COMPLETED. THE STENT, BALLOON, AND SHEARED SHAFT WERE SUCCESSFULLY RETRIEVED DURING SURGERY. THE PROCEDURE WAS COMPLETED. THE RETRIEVED STENT STRUT WAS ALSO NOTED TO BE DEFORMED. THE PATIENT WAS SUBSEQUENTLY TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). HOWEVER, DESPITE ALL INTERVENTIONS, THE PATIENT PASSED AWAY THE FOLLOWING DAY. Manufacturer Narrative: GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22248440&pc=DQY | 2124215-2025-37639 | 2025-05-22 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-17 04:00 | DQY | GUIDEZILLA? II | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Ventricular Fibrillation | K163314 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT WAS ADMITTED FOR AN ELECTIVE CORONARY ANGIOGRAM FOLLOWING A POSITIVE CT CORONARY ANGIOGRAM (CTCA). THE PROCEDURE WAS INITIATED VIA RIGHT RADIAL ARTERY ACCESS, AND A 7F NON-BOSTON SCIENTIFIC RADIAL SHEATH WAS INSERTED. ANGIOGRAPHY REVEALED SEVERELY CALCIFIED LEFT ANTERIOR DESCENDING (LAD) ARTERY LESIONS, WHICH WAS ALSO NOTED TO BE MODERATELY TORTUOUS. A 6F GUIDEZILLA II GUIDE EXTENSION CATHETER WAS ENGAGED TO ASSIST DEVICE ADVANCEMENT. ROTATIONAL ATHERECTOMY WAS PERFORMED USING A 1.50 MM ROTAPRO DEVICE. INTRAVASCULAR ULTRASOUND (IVUS) WAS THEN PERFORMED FOLLOWED BY MULTIPLE BALLOON INFLATIONS WITH 3.00MM X 15MM, 2.75MM X 12MM AND 2.50MM X 12MM NC EMERGE BALLOON CATHETERS. THE BALLOON INFLATIONS WERE CARRIED OUT IN THE PROXIMAL AND MID-LAD SEGMENTS. DURING THIS PROCESS, HYPOTENSIVE EPISODES WERE OBSERVED AND MANAGED WITH MULTIPLE ARAMINE BOLUSES. A 2.50 X 32 MM SYNERGY XD DRUG-ELUTING STENT WAS THEN ADVANCED DOWN THE LAD. WHILE OPTIMAL POSITIONING WAS PERFORMED, THE STENT SHAFT WAS SHEARED, NOTED POSSIBLY DUE TO CONTACT WITH HEAVILY CALCIFIED PLAQUE. A PORTION OF THE SHAFT, INCLUDING THE UNDEPLOYED STENT AND STENT BALLOON, BECAME DETACHED AND LODGED WITHIN THE LAD, RESULTING IN VESSEL OCCLUSION. THE PATIENTS CONDITION BEGAN TO DETERIORATE, AND A CODE BLUE WAS ACTIVATED. A CARDIOTHORACIC SURGEON WAS URGENTLY CONTACTED, AND ADVANCED LIFE SUPPORT (ALS) MEASURES WERE INITIATED. THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF) AND WAS DEFIBRILLATED WITH 200 JOULES. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, AND SINUS RHYTHM WAS RESTORED AFTER THE INITIAL SHOCK. PREPARATION BEGAN TO TRANSFER THE PATIENT TO CARDIAC SURGERY. THE PATIENT WAS INTUBATED AND THEN DETERIORATED BACK INTO VF, WHICH REQUIRED FOUR ADDITIONAL DEFIBRILLATION SHOCKS, THREE 1 MG DOSES OF ADRENALINE, AND 300 MG OF AMIODARONE, ADMINISTERED ACCORDING TO ALS PROTOCOL. A LUCAS MECHANICAL CPR DEVICE WAS APPLIED, AND RESUSCITATIVE EFFORTS WERE CONTINUED DURING TRANSFER TO THE CARDIAC SURGERY ROOM FOR URGENT CORONARY ARTERY BYPASS GRAFTING (CABG). AN EMERGENCY STERNOTOMY WAS PERFORMED, AND CARDIOPULMONARY BYPASS WAS INITIATED VIA FEMORAL CANNULATION. THE PATIENT UNDERWENT CABG TO RIGHT CORONARY ARTERY (RCA) X 2, LAD, AND DISTAL POSTERIOR DESCENDING ARTERY (PDA), WITH LEFT ATRIAL APPENDAGE (LAA) LIGATION ALSO COMPLETED. THE STENT, BALLOON, AND SHEARED SHAFT WERE SUCCESSFULLY RETRIEVED DURING SURGERY. THE PROCEDURE WAS COMPLETED. THE RETRIEVED STENT STRUT WAS ALSO NOTED TO BE DEFORMED. THE PATIENT WAS SUBSEQUENTLY TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). HOWEVER, DESPITE ALL INTERVENTIONS, THE PATIENT PASSED AWAY THE FOLLOWING DAY. Manufacturer Narrative: GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256904&pc=NPT | 2015691-2025-04569 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 2 AORTIC ANNULAR RUPTURE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 81 - 93 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 ANNULAR RUPTURE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU) CARDIOVASCULAR INJURIES, SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES, AND HEMATOMA ARE A KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO THE THV TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, HEMATOMA, OR ANNULAR RUPTURE DURING THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE INCLUDE SIGNIFICANT THV OVERSIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION AND NARROW CALCIFIED STJ. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN VALVE RELIES ON NATIVE VALVE CALCIUM TO SECURELY ANCHOR TO THE ANNULUS. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256905&pc=DYB | 2015691-2025-04576 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Retroperitoneal Hemorrhage | K200258 | E2016006 | 3 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 3 AORTIC BLEEDING - RETROPERITONEAL DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 63 - 78 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 BLEEDING - RETROPERITONEAL DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0.33 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256906&pc=DYB | 2015691-2025-04577 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K200258 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 1 AORTIC BLEEDING - ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 82 - 82 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 15 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256907&pc=DYB | 2015691-2025-04575 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hematoma | K200258 | E2016006 | 3 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 3 AORTIC BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 63 - 82 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 1 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256908&pc=NPT | 2015691-2025-04578 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140031 | E2016006 | 5 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 5 AORTIC CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 81 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 5 CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0.2 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE:(B)(4). PER THE INSTRUCTIONS FOR USE, CARDIOVASCULAR INJURIES SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR PERFORATION DURING A TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256909&pc=NPT | 2015691-2025-04568 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem |  | P140031 | E2016006 | 3 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 3 AORTIC CORONARY ARTERY COMPRESSION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 78 - 85 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 CORONARY ARTERY COMPRESSION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 1 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT THE SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS IN CLOSE PROXIMITY TO CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OCCLUSION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PCI). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE THV PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, DEPLOYMENT OF THE BIOPROSTHETIC HEART VALVE TOO AORTIC, AND SIGNIFICANT VALVE OVER SIZING COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE EDWARDS THV TRAINING MANUALS ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, ANNULUS TO CORONARY OSTIA DISTANCE, LENGTH OF THE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BAV, PATENCY OF CORONARIES DURING BAV, AND EXPANDED HEIGHT OF THE INTENDED THV. THE TRAINING MANUALS ALSO PROVIDE THE FOLLOWING TIPS FOR DETECTING RISK FOR LEFT MAIN OCCLUSION: (1) AORTOGRAM OR TEE PRIOR TO THV IMPLANTATION TO REVEAL BULKY CALCIFIED LEAFLETS; (2) DURING PRE-DILATATION, NOTE BULKY CALCIFICATION ON VALVE MOVING TOWARDS OSTIUM ON LEFT MAIN; AND (3) CONSIDER AORTOGRAM DURING VALVULOPLASTY TO ASSESS CORONARY FLOW. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256914&pc=NPT | 2015691-2025-04567 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 4 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 4 AORTIC PERCUTANEOUS CORONARY INTERVENTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 69 - 78 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 3 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 4 PERCUTANEOUS CORONARY INTERVENTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT THE SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED IN PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS IN CLOSE PROXIMITY TO THE CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PCI). THERE ARE MULTIPLE PATIENT FACTORS THAT COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, DEPLOYMENT OF THE BIOPROSTHETIC HEART VALVE TOO AORTIC, AND SIGNIFICANT VALVE OVER SIZING COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE EDWARDS THV MANUALS ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256915&pc=NPT | 2015691-2025-04563 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 17 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 17 AORTIC CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 63 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 11 FEMALE(S) AND 6 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 17 CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 1.94 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST REOPERATION, INCLUDING EMERGENCY CARDIAC SURGERY, AS ADDITIONAL POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THIS EVENT TYPE (CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED) IS DEFINED AS: THE PATIENT SUBSEQUENTLY UNDERWENT OTHER TYPES OF CARDIAC SURGERY OR CATH LAB INTERVENTIONS (NOT OTHERWISE SPECIFIED) THAT WAS UNPLANNED. THIS COULD INCLUDE REOPERATION FOR DAMAGE TO ADJACENT STRUCTURES, SUCH AS THE MITRAL APPARATUS, OR POTENTIALLY VALVE EXPLANT RELATED TO EMBOLIZATION, OR OTHER EVENT TYPES. CARDIOVASCULAR INJURY, INCLUDING PERFORATION OR DISSECTION OF VALVULAR STRUCTURES (E.G., INJURY TO THE MITRAL VALVE/ APPARATUS), AND VALVE EMBOLIZATION REQUIRING INTERVENTION ARE KNOWN POTENTIAL COMPLICATIONS LISTED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU). CHORDAE TENDINEA RUPTURE IN TRANSCATHETER VALVE REPLACEMENT CAN OCCUR DURING ADVANCEMENT OF THE GUIDEWIRE, BAV CATHETER, OR DELIVERY SYSTEM AND IS MOST LIKELY TO OCCUR WITH ANTEGRADE APPROACHES, HOWEVER, CAN OCCUR WITH A RETROGRADE APPROACH. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVES (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, AND CAREFUL MANIPULATION OF DEVICES. PER THE THV TA TRAINING MANUALS, AFTER GAINING LV ACCESS WITH THE 0.035" SOFT GUIDEWIRE, THE WIRE SHOULD BE JIGGLED UNDER TEE IMAGING OF THE MITRAL VALVE. AN INCREASE IN MR SUGGESTS WIRE ENTANGLEMENT IN THE MITRAL SUB-VALVULAR APPARATUS. IF ENTANGLEMENT IS SUSPECTED, THE GUIDEWIRE SHOULD BE COMPLETELY REMOVED FROM THE VENTRICLE; THE OPERATOR SHOULD CHANGE DIRECTION OF THE NEEDLE AND REINSERT THE GUIDEWIRE INTO THE VENTRICLE, CHECKING AGAIN FOR WIRE ENTANGLEMENT. THE TF TRAINING MANUAL ALSO PROVIDES GUIDANCE FOR VALVE CROSSING AND WIRE EXCHANGE: EXCHANGE 0.035" AMPLATZ EXTRA-STIFF WIRE WITH PRE-SHAPED DISTAL END IN LEFT VENTRICLE. ENSURE GUIDE CATHETER IS ADVANCED UPON WIRE EXCHANGE TO ENSURE EXCHANGE WIRE DOES NOT GET CAUGHT IN THE MITRAL CHORDAE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO VENTRICULAR EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE/DELIVERY SYSTEM, A NARROW, CALCIFIED SINOTUBULAR JUNCTION, MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR VENTRICULAR EMBOLIZATION (I.E., SMALL, CALCIFIED STJ, MINIMAL LEAFLET CALCIFICATION), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. I IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256916&pc=NPT | 2015691-2025-04571 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 1 AORTIC DEVICE EMBOLIZATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 DEVICE EMBOLIZATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, , MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, PRESERVED EJECTION FRACTION, LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE ,AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E. MINIMAL LEAFLET CALCIFICATION), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256917&pc=DYB | 2015691-2025-04599 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Retroperitoneal Hemorrhage | K200258 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 AORTIC BLEEDING - RETROPERITONEAL DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 27 - 27 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - RETROPERITONEAL DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; ; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256918&pc=NPT | 2015691-2025-04566 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Heart Block | P140031 | E2016006 | 8 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 8 AORTIC PERMANENT PACEMAKER DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 63 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALE(S) AND 4 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 8 PERMANENT PACEMAKER DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 5 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, CONDUCTION SYSTEM DEFECTS (HEART BLOCK) ARRHYTHMIAS AND CONDUCTION SYSTEM DEFECTS THAT MAY REQUIRE A PERMANENT PACEMAKER ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH BALLOON AORTIC VALVULOPLASTY, DEPLOYMENT OF THE PROSTHETIC VALVE, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC) GUIDELINES, THE CLOSE ANATOMICAL RELATIONSHIP BETWEEN THE AORTIC VALVE COMPLEX AND THE BRANCHING ATRIOVENTRICULAR BUNDLE MAY PROVIDE AN EXPLANATION FOR THESE COMPLICATIONS OF THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY -CONDUCTION DISTURBANCES/ HEART BLOCK WRITTEN BY EDWARDS LIFESCIENCES, ATRIOVENTRICULAR CONDUCTION DISTURBANCES AFTER THV PLACEMENT ARE ASSOCIATED WITH MANY PATIENT RELATED AND PROCEDURAL RELATED FACTORS, INCLUDING PRE-OPERATIVE CO-MORBID STATUS, THE DEGREE AND BULKINESS OF AORTIC VALVE AND ANNULAR CALCIFICATION, INTER-VENTRICULAR SEPTAL THICKNESS, PRE-EXISTING ELECTROCARDIOGRAM ABNORMALITIES, THE DEPTH OF PROSTHESIS IMPLANTATION, AND THE PROFILE OF THE IMPLANTED PROSTHESIS. UNLIKE CONVENTIONAL AVR, WHERE THERE MAY BE LOCALIZED TRAUMA DUE TO DECALCIFICATION OF THE ANNULUS AND/OR SUTURE PLACEMENT IN THE PROXIMITY OF THE AV NODE OR THE BUNDLES, TRANSCATHETER VALVE REPLACEMENT MAY CAUSE CONDUCTION ABNORMALITIES THROUGH MECHANICAL IMPINGEMENT OF THE CONDUCTION SYSTEM BY THE PROSTHESIS. THE MECHANISMS OF THE DEVELOPMENT OF HEART BLOCK AFTER THV PLACEMENT ARE WELL DOCUMENTED AND DESCRIBED IN THE LITERATURE. IT IS ALSO DOCUMENTED THAT PRE-EXISTING HEART BLOCK IS COMMON IN PATIENTS UNDERGOING THV PLACEMENT OR SURGICAL AVR AND ANOTHER 4-6 % WILL DEVELOP POSTOPERATIVE HEART BLOCK, POTENTIALLY REQUIRING A PERMANENT PACEMAKER. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE, HOWEVER THE CONDUCTION DISTURBANCE IS LIKELY RELATED TO THE MECHANISM DESCRIBED ABOVE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256919&pc=DYB | 2015691-2025-04597 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K200258 | E2016006 | 5 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 5 AORTIC BLEEDING - ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 27 - 88 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 3 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 5 BLEEDING - ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0.2 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256920&pc=NPT | 2015691-2025-04572 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 1 AORTIC STROKE - UNDETERMINED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 78 - 78 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 STROKE - UNDETERMINED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 1 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS INCLUDING STROKE ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH INVASIVE CARDIAC PROCEDURE INCLUDING TRANSCATHETER HEART VALVE PROCEDURES AND THE USE OF THE EDWARDS THV DEVICES. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING TRANSCATHETER VALVE REPLACEMENT PROCEDURE. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING TRANSCATHETER VALVE REPLACEMENT DEMONSTRATED THAT THE MAJORITY OF PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO IMPORTANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN TRANSCATHETER VALVE REPLACEMENT AND AVR PATIENTS. AFTER TAVR, THERE APPEARS TO BE A MORE SIGNIFICANT PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT TRANSCATHETER VALVE REPLACEMENT PROCEDURE PATIENTS WITH MULTIPLE CO MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THE EVENT IS LISTED AS UNDETERMINED STROKE. IT IS NOT KNOWN IF THE MECHANISM OF THE STROKE IS IN QUESTION, OR IF THE STROKE ITSELF IS IN QUESTION. THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. IT IS POSSIBLE THAT THE MECHANISM DESCRIBED ABOVE MAY HAVE CONTRIBUTED TO THE EVENT. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256921&pc=NPT | 2015691-2025-04564 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 13 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 13 AORTIC READMISSION (VALVE RELATED) DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 44 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 6 FEMALE(S) AND 7 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 13 READMISSION (VALVE RELATED) DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 168.38 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). VALVE RELATED RE-ADMISSION IN THE FOLLOW-UP PERIOD IS LIKELY DUE TO REOCCURRENCE OF SYMPTOMS. THIS MAY RESULT FROM REGURGITATION (CENTRAL OR PVL) OR PROGRESSION OF THE PRE-EXISTING DISEASE PROCESS AND MAY RESULT IN HEART FAILURE. CAUSES OF HEART FAILURE CAN BE MULTI-FACTORIAL. PER THE INSTRUCTION FOR USE (IFU), HEART FAILURE, VALVE REGURGITATION, AND VALVE DEGENERATION INCLUDING STENOSIS ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH AORTIC VALVE REPLACEMENT AND BIOPROSTHETIC HEART VALVES. CONGESTIVE HEART FAILURE CAN HAVE MULTIPLE ETIOLOGIES AND IS OFTEN DUE TO THE PROGRESSION OF UNDERLYING DISEASE PROCESSES, INCLUDING CORONARY ARTERY DISEASE, UNCONTROLLED HYPERTENSION, MYOCARDIAL INFARCTION, CARDIOMYOPATHY, FLUID OVERLOAD, OR VALVULAR DYSFUNCTION. RISK FACTORS THAT INCREASE A PATIENT'S RISK OF SUFFERING FROM CHF INCLUDE OBESITY, ADVANCED AGE, AND A HISTORY OF SMOKING. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO A NUMBER OF ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. VALVE STENOSIS MAY RESULT IN SYMPTOMS SUCH AS SOB AND DECREASED EXERCISE TOLERANCE, WHICH MAY BE ACCOMPANIED BY AN INCREASED GRADIENT ACROSS THE VALVE. THIS COULD BE DUE TO EARLY CALCIFICATION OF THE LEAFLETS, HOST TISSUE OVERGROWTH, OR IN RARE CASES A NON-FUNCTIONING LEAFLET. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256922&pc=NPT | 2015691-2025-04565 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P140031 | E2016006 | 13 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 13 AORTIC STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 48 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 8 FEMALE(S) AND 5 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 13 STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 6.62 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS INCLUDING STROKE ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH INVASIVE CARDIAC PROCEDURE INCLUDING TRANSCATHETER HEART VALVE PROCEDURES AND THE USE OF THE EDWARDS THV DEVICES. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING TRANSCATHETER VALVE REPLACEMENT PROCEDURE. RISK FACTORS CORRELATING WITH A NUMBER OF PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE DEMONSTRATED THAT THE MAJORITY OF PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO IMPORTANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND AVR PATIENTS. AFTER TRANSCATHETER VALVE REPLACEMENT PROCEDURE, THERE APPEARS TO BE A MORE SIGNIFICANT PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT TRANSCATHETER VALVE REPLACEMENT PROCEDURE PATIENTS WITH MULTIPLE CO MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE HOWEVER, THE EVENT MAY BE RELATED TO THE MECHANISM DESCRIBED ABOVE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256923&pc=NPT | 2015691-2025-04570 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Insufficient Information | Aortic Valve Insufficiency/ Regurgitation | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 2 AORTIC REINTERVENTION - AORTIC VALVE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 73 - 87 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 REINTERVENTION - AORTIC VALVE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 121 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). INTRA-PROCEDURAL AORTIC VALVE RE-INTERVENTION WILL TYPICALLY RESULT FROM VALVE MALPOSITION OR REGURGITATION (PVL OR CENTRAL, INCLUDING LEAFLET RESTRICTION). THESE CONDITIONS ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO AORTIC MALPOSITION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, PRESERVED EJECTION FRACTION, SIGNIFICANT MITRAL ANNULAR CALCIFICATION (MAC), LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. DEPLOYMENT OF THE SAPIEN VALVE TOO AORTIC HAS THE POTENTIAL TO CONTRIBUTE TO SUBOPTIMAL COAPTATION OF THE SAPIEN VALVE LEAFLETS AND CAUSE CENTRAL AORTIC INSUFFICIENCY; IT CAN OBSTRUCT THE CORONARY OSTIA; AND LEAD TO EMBOLIZATION OF THE PROSTHESIS INTO THE ASCENDING AORTA. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR AORTIC MALPOSITION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. AORTIC VALVE RE-INTERVENTION IN THE FOLLOW-UP PERIOD (IN THE ABSENCE OF PROSTHETIC CARDIAC VALVE THROMBOSIS) WILL TYPICALLY RESULT FROM ON-GOING OR WORSENING REGURGITATION, VALVE DEGENERATION RELATED TO THE FORMATION OF CALCIFICATION OR PANNUS, OR VALVE MIGRATION. THESE EVENTS ARE IDENTIFIED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU) AS POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO A NUMBER OF ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. CALCIFICATION IS A WELL-RECOGNIZED FAILURE MODE OF BIOPROSTHETIC VALVES. THE MECHANISM OF CALCIFICATION OF BIOMATERIALS IS NOT COMPLETELY UNDERSTOOD BUT IS PROBABLY RELATED TO AN INABILITY OF THE NON-VIABLE CELLS TO MAINTAIN THEIR NORMALLY LOW INTRACELLULAR CONCENTRATION OF CALCIUM. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. ACCORDING TO LITERATURE REVIEW, VALVE MIGRATION RESULTS WHEN FORCES ACTING ON THE TRANSCATHETER HEART VALVE (THV) OVERCOME THE STRENGTH OF ATTACHMENT OF THE VALVE TO THE AORTIC WALL. STENT VALVES ARE SUBJECTED TO ANTEGRADE EJECTION FORCES DURING SYSTOLE. LESS-THAN-SEVERE AND NON-UNIFORMLY DISTRIBUTED CALCIFICATION OF THE NATIVE LEAFLETS, INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION CAN CONTRIBUTE TO VALVE MIGRATION. ADDITIONALLY, RESIDUAL OVERHANGING LEAFLETS CAN EXERT DOWNWARDS FORCE DURING DIASTOLE, CAUSING MIGRATION OF THE THV TOWARDS THE LEFT VENTRICLE IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256924&pc=DYB | 2015691-2025-04573 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | K200258 | E2016006 | 25 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 25 AORTIC VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 63 - 94 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 13 FEMALE(S) AND 12 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 25 VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 31.8 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256926&pc=DYB | 2015691-2025-04598 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Hematoma | K200258 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 AORTIC BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 86 - 86 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256929&pc=NPT | 2015691-2025-04601 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem |  | P140031 | E2016006 | 4 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 4 AORTIC AORTIC DISSECTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 63 - 91 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 4 AORTIC DISSECTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0.75 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU) CARDIOVASCULAR INJURY, SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES, IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ASCENDING AORTIC DISSECTION MAY OCCUR WHEN MULTIPLE ATTEMPTS ARE MADE TO CROSS THE STENOTIC NATIVE VALVE, AND/OR WHEN EXCESSIVE FORCE IS USED. AORTIC DISSECTION IN A DISTAL SECTION OF THE AORTA CAN OCCUR WITH DEVICE MANIPULATION AND NAVIGATION THROUGH ANATOMIC TORTUOSITY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). THE THV TRAINING MANUALS PROVIDE GUIDANCE TO FACILITATE SAFE CROSSING OF THE NATIVE VALVE, INCLUDING CAMERA PROJECTIONS, HANDLING DURING ADVANCEMENT, AND TROUBLESHOOTING TECHNIQUES IF DIFFICULTY IS ENCOUNTERED. AS STATED, EXCESSIVE FORCE SHOULD NOT BE USED WHEN THE DEVICE HAS DIFFICULTY CROSSING THE STENOTIC VALVE. ADDING TENSION TO THE WIRE, PULLING BACK THE SYSTEM TO RE-ORIENT THE VALVE, AS NEEDED, AND TORQUEING OF THE FLEX CATHETER MAY BE HELPFUL IN SOLVING THE PROBLEM. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256930&pc=NPT | 2015691-2025-04588 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 | E2016006 | 8 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 8 AORTIC ANNULAR RUPTURE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 73 - 89 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 6 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 8 ANNULAR RUPTURE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU) CARDIOVASCULAR INJURIES, SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES, AND HEMATOMA ARE A KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO THE THV TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, HEMATOMA, OR ANNULAR RUPTURE DURING THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE INCLUDE SIGNIFICANT THV OVERSIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION AND NARROW CALCIFIED STJ. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN VALVE RELIES ON NATIVE VALVE CALCIUM TO SECURELY ANCHOR TO THE ANNULUS. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256931&pc=NPU | 2015691-2025-04582 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPU | EDWARDS COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 1 MITRAL ASD DEFECT CLOSURE DUE TO TRANSSEPTAL CATHETERIZATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 85 - 85 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 ASD DEFECT CLOSURE DUE TO TRANSSEPTAL CATHETERIZATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 2 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU) CARDIOVASCULAR INJURY, SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES, IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH TRANSCATHETER HEART VALVES. TRANSEPTAL APPROACHES ALLOW TRANSFEMORAL (VENOUS) ACCESS TO THE MITRAL VALVE BY PUNCTURING THE ATRIAL SEPTUM AND ADVANCING THE DELIVERY SYSTEM THROUGH THE OPENING. USUALLY, THE SEPTOSTOMY CLOSES ON ITS OWN OVER TIME AND DOES NOT REQUIRE TREATMENT. IN SOME CASES, THE INTERVENTIONALIST WILL ELECT TO CLOSE THE CREATED ATRIAL SEPTAL DEFECT (ASD) DURING THE INITIAL PROCEDURE PERCUTANEOUSLY WITH A SEPTAL OCCLUDER. PERSISTENT IATROGENIC ATRIAL SEPTAL DEFECTS (IASD) ARE NOT UNCOMMON AND MAY BE ASSOCIATED WITH THE USE OF LARGER SHEATHS. THESE MAY REQUIRE PLACEMENT OF A CLOSURE DEVICE IN A REPEAT PROCEDURE AND ARE CONSIDERED SERIOUS INJURIES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256932&pc=NPT | 2015691-2025-04590 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 4 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 4 AORTIC DEVICE EMBOLIZATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 27 - 86 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 2 MALE(S).. Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 4 DEVICE EMBOLIZATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 91.75 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, VALVE EMBOLIZATION IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, , MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, PRESERVED EJECTION FRACTION, LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE ,AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR EMBOLIZATION (I.E. MINIMAL LEAFLET CALCIFICATION), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256933&pc=NPT | 2015691-2025-04592 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Myocardial Infarction | P140031 | E2016006 | 3 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 3 AORTIC MYOCARDIAL INFARCTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 51 - 81 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 3 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 MYOCARDIAL INFARCTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 1.67 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), THROMBUS FORMATION, PLAQUE DISLODGEMENT, AND EMBOLIZATION THAT MAY RESULT IN MYOCARDIAL INFARCTION (MI) ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT THE SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS IN CLOSE PROXIMITY TO THE CORONARY OSTIA. IN ADDITION, CAUTION SHOULD BE EXERCISED WHEN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. MI RELATED TO THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND THE VALVE IMPLANT WILL MANIFEST INTRA-PROCEDURALLY OR IN THE IMMEDIATE POST-OPERATIVE PERIOD AND IS TYPICALLY DUE TO A COMBINATION OF PATIENT AND/OR PROCEDURAL FACTORS. AS DEFINED IN THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC) PUBLICATION ON TRANSCATHETER VALVE REPLACEMENT COMPLICATIONS, THE PERI-PROCEDURAL INTERVAL IS INCLUSIVE OF ALL EVENTS THAT BEGIN WITHIN 72 HRS. OF THE INDEX PROCEDURE. MI'S THAT OCCUR AFTER THE PERI-PROCEDURAL PERIOD (>72 HRS.) ARE TYPICALLY RELATED TO THE PATIENT'S UNDERLYING CORONARY DISEASE. THERE ARE MULTIPLE PATIENT FACTORS THAT COULD PUT THE PATIENT AT RISK FOR MI DURING THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE, AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. THE EDWARDS THV MANUALS ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256934&pc=NPT | 2015691-2025-04600 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140031 | E2016006 | 13 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 13 AORTIC CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 62 - 92 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 8 FEMALE(S) AND 5 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 13 CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, CARDIOVASCULAR INJURIES SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR PERFORATION DURING A TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. PER THE PROCEDURE DIDACTIC, PATIENTS WITH SMALL VENTRICLES ARE AT PARTICULARLY HIGH RISK FOR VENTRICULAR PERFORATION. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256935&pc=NPU | 2015691-2025-04580 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPU | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 1 MITRAL CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 85 - 85 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 2 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST REOPERATION, INCLUDING EMERGENCY CARDIAC SURGERY, AS ADDITIONAL POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THIS EVENT TYPE (CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED E031) IS DEFINED AS: THE PATIENT SUBSEQUENTLY UNDERWENT OTHER TYPES OF CARDIAC SURGERY OR CATH LAB INTERVENTIONS (NOT OTHERWISE SPECIFIED) THAT WAS UNPLANNED. THIS COULD INCLUDE REOPERATION FOR DAMAGE TO ADJACENT STRUCTURES, OR POTENTIALLY VALVE EXPLANT RELATED TO EMBOLIZATION, OR OTHER EVENT TYPES. CARDIOVASCULAR INJURY, INCLUDING PERFORATION OR DISSECTION OF VALVULAR STRUCTURES AND VALVE EMBOLIZATION REQUIRING INTERVENTION ARE KNOWN POTENTIAL COMPLICATIONS LISTED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU). PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVES (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, AND CAREFUL MANIPULATION OF DEVICES. PER THE THV TA TRAINING MANUALS, AFTER GAINING LV ACCESS WITH THE 0.035" SOFT GUIDEWIRE, THE WIRE SHOULD BE JIGGLED UNDER TEE IMAGING OF THE MITRAL VALVE. AN INCREASE IN MR SUGGESTS WIRE ENTANGLEMENT IN THE MITRAL SUB-VALVULAR APPARATUS. IF ENTANGLEMENT IS SUSPECTED, THE GUIDEWIRE SHOULD BE COMPLETELY REMOVED FROM THE VENTRICLE; THE OPERATOR SHOULD CHANGE DIRECTION OF THE NEEDLE AND REINSERT THE GUIDEWIRE INTO THE VENTRICLE, CHECKING AGAIN FOR WIRE ENTANGLEMENT. THE TF TRAINING MANUAL ALSO PROVIDES GUIDANCE FOR VALVE CROSSING AND WIRE EXCHANGE: EXCHANGE 0.035" AMPLATZ EXTRA-STIFF WIRE WITH PRE-SHAPED DISTAL END IN LEFT VENTRICLE. ENSURE THE GUIDE CATHETER IS ADVANCED UPON WIRE EXCHANGE TO ENSURE EXCHANGE WIRE DOES NOT GET CAUGHT IN THE MITRAL CHORDAE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO VENTRICULAR EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE/DELIVERY SYSTEM, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256936&pc=NPT | 2015691-2025-04584 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 18 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 18 AORTIC CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 27 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 7 FEMALE(S) AND 11 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 18 CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S) FOR THIS EVENT IS 1.39 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST REOPERATION, INCLUDING EMERGENCY CARDIAC SURGERY, AS ADDITIONAL POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THIS EVENT TYPE (CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED) IS DEFINED AS: THE PATIENT SUBSEQUENTLY UNDERWENT OTHER TYPES OF CARDIAC SURGERY OR CATH LAB INTERVENTIONS (NOT OTHERWISE SPECIFIED) THAT WAS UNPLANNED. THIS COULD INCLUDE REOPERATION FOR DAMAGE TO ADJACENT STRUCTURES, SUCH AS THE MITRAL APPARATUS, OR POTENTIALLY VALVE EXPLANT RELATED TO EMBOLIZATION, OR OTHER EVENT TYPES. CARDIOVASCULAR INJURY, INCLUDING PERFORATION OR DISSECTION OF VALVULAR STRUCTURES (E.G., INJURY TO THE MITRAL VALVE/ APPARATUS), AND VALVE EMBOLIZATION REQUIRING INTERVENTION ARE KNOWN POTENTIAL COMPLICATIONS LISTED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU). CHORDAE TENDINEA RUPTURE IN TRANSCATHETER VALVE REPLACEMENT CAN OCCUR DURING ADVANCEMENT OF THE GUIDEWIRE, BAV CATHETER, OR DELIVERY SYSTEM AND IS MOST LIKELY TO OCCUR WITH ANTEGRADE APPROACHES, HOWEVER, CAN OCCUR WITH A RETROGRADE APPROACH. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVES (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, AND CAREFUL MANIPULATION OF DEVICES. PER THE THV TA TRAINING MANUALS, AFTER GAINING LV ACCESS WITH THE 0.035'' SOFT GUIDEWIRE, THE WIRE SHOULD BE JIGGLED UNDER TEE IMAGING OF THE MITRAL VALVE. AN INCREASE IN MR SUGGESTS WIRE ENTANGLEMENT IN THE MITRAL SUB-VALVULAR APPARATUS. IF ENTANGLEMENT IS SUSPECTED, THE GUIDEWIRE SHOULD BE COMPLETELY REMOVED FROM THE VENTRICLE; THE OPERATOR SHOULD CHANGE DIRECTION OF THE NEEDLE AND REINSERT THE GUIDEWIRE INTO THE VENTRICLE, CHECKING AGAIN FOR WIRE ENTANGLEMENT. THE TF TRAINING MANUAL ALSO PROVIDES GUIDANCE FOR VALVE CROSSING AND WIRE EXCHANGE: EXCHANGE 0.035'' AMPLATZ EXTRA-STIFF WIRE WITH PRE-SHAPED DISTAL END IN LEFT VENTRICLE. ENSURE GUIDE CATHETER IS ADVANCED UPON WIRE EXCHANGE TO ENSURE EXCHANGE WIRE DOES NOT GET CAUGHT IN THE MITRAL CHORDAE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO VENTRICULAR EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE/DELIVERY SYSTEM, A NARROW, CALCIFIED SINOTUBULAR JUNCTION, MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR VENTRICULAR EMBOLIZATION (I.E., SMALL, CALCIFIED STJ, MINIMAL LEAFLET CALCIFICATION), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. I IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256937&pc=NPT | 2015691-2025-04594 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem |  | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 2 AORTIC CORONARY ARTERY COMPRESSION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 67 - 74 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 2 CORONARY ARTERY COMPRESSION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 2.5 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT THE SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED FOR PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS IN CLOSE PROXIMITY TO CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OCCLUSION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PCI). MULTIPLE PATIENT FACTORS COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE THV PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, DEPLOYMENT OF THE BIOPROSTHETIC HEART VALVE TOO AORTIC, AND SIGNIFICANT VALVE OVER SIZING COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE EDWARDS THV TRAINING MANUALS ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE PHYSICIAN IS INSTRUCTED TO EVALUATE THIS RISK EARLY IN THE PATIENT SCREENING PROCESS IN ALL PATIENTS. THE FOLLOWING FACTORS SHOULD BE CONSIDERED: DEGREE OF CALCIFICATION ON LEAFLETS, ANNULUS TO CORONARY OSTIA DISTANCE, LENGTH OF THE VALVE LEAFLET, WIDTH OF THE VALSALVA SINUSES, MOVEMENT OF THE LEAFLETS DURING BAV, PATENCY OF CORONARIES DURING BAV, AND EXPANDED HEIGHT OF THE INTENDED THV. THE TRAINING MANUALS ALSO PROVIDE THE FOLLOWING TIPS FOR DETECTING RISK FOR LEFT MAIN OCCLUSION: (1) AORTOGRAM OR TEE PRIOR TO THV IMPLANTATION TO REVEAL BULKY CALCIFIED LEAFLETS; (2) DURING PRE-DILATATION, NOTE BULKY CALCIFICATION ON VALVE MOVING TOWARDS OSTIUM ON LEFT MAIN; AND (3) CONSIDER AORTOGRAM DURING VALVULOPLASTY TO ASSESS CORONARY FLOW. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256938&pc=NPT | 2015691-2025-04585 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 18 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 18 AORTIC READMISSION (VALVE RELATED) DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 56 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 6 FEMALE(S) AND 12 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 18 READMISSION (VALVE RELATED) DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 134.61 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). VALVE RELATED RE-ADMISSION IN THE FOLLOW-UP PERIOD IS LIKELY DUE TO REOCCURRENCE OF SYMPTOMS. THIS MAY RESULT FROM REGURGITATION (CENTRAL OR PVL) OR PROGRESSION OF THE PRE-EXISTING DISEASE PROCESS AND MAY RESULT IN HEART FAILURE. CAUSES OF HEART FAILURE CAN BE MULTI-FACTORIAL. PER THE INSTRUCTION FOR USE (IFU), HEART FAILURE, VALVE REGURGITATION, AND VALVE DEGENERATION INCLUDING STENOSIS ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH AORTIC VALVE REPLACEMENT AND BIOPROSTHETIC HEART VALVES. CONGESTIVE HEART FAILURE CAN HAVE MULTIPLE ETIOLOGIES AND IS OFTEN DUE TO THE PROGRESSION OF UNDERLYING DISEASE PROCESSES, INCLUDING CORONARY ARTERY DISEASE, UNCONTROLLED HYPERTENSION, MYOCARDIAL INFARCTION, CARDIOMYOPATHY, FLUID OVERLOAD, OR VALVULAR DYSFUNCTION. RISK FACTORS THAT INCREASE A PATIENT'S RISK OF SUFFERING FROM CHF INCLUDE OBESITY, ADVANCED AGE, AND A HISTORY OF SMOKING. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO A NUMBER OF ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. VALVE STENOSIS MAY RESULT IN SYMPTOMS SUCH AS SOB AND DECREASED EXERCISE TOLERANCE, WHICH MAY BE ACCOMPANIED BY AN INCREASED GRADIENT ACROSS THE VALVE. THIS COULD BE DUE TO EARLY CALCIFICATION OF THE LEAFLETS, HOST TISSUE OVERGROWTH, OR IN RARE CASES A NON-FUNCTIONING LEAFLET. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256939&pc=NPT | 2015691-2025-04591 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Insufficient Information | Aortic Valve Insufficiency/ Regurgitation | P140031 | E2016006 | 4 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 4 AORTIC REINTERVENTION - AORTIC VALVE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 56 - 80 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 4 REINTERVENTION - AORTIC VALVE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 181.25 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). INTRA-PROCEDURAL AORTIC VALVE RE-INTERVENTION WILL TYPICALLY RESULT FROM VALVE MALPOSITION OR REGURGITATION (PVL OR CENTRAL, INCLUDING LEAFLET RESTRICTION). THESE CONDITIONS ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO AORTIC MALPOSITION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, SEVERE SEPTAL HYPERTROPHY, MINIMALLY OR BULKY/SEVERELY CALCIFIED AORTIC LEAFLETS, PRESERVED EJECTION FRACTION, SIGNIFICANT MITRAL ANNULAR CALCIFICATION (MAC), LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. DEPLOYMENT OF THE SAPIEN VALVE TOO AORTIC HAS THE POTENTIAL TO CONTRIBUTE TO SUBOPTIMAL COAPTATION OF THE SAPIEN VALVE LEAFLETS AND CAUSE CENTRAL AORTIC INSUFFICIENCY; IT CAN OBSTRUCT THE CORONARY OSTIA; AND LEAD TO EMBOLIZATION OF THE PROSTHESIS INTO THE ASCENDING AORTA. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN PATIENTS WITH HIGH-RISK ANATOMICAL FEATURES FOR AORTIC MALPOSITION (I.E., MINIMAL LEAFLET CALCIFICATION, SEVERE SEPTAL HYPERTROPHY), BAV MAY PROVIDE INDICATION OF POTENTIAL BALLOON MOVEMENT DURING VALVE DEPLOYMENT. AORTIC VALVE RE-INTERVENTION IN THE FOLLOW-UP PERIOD (IN THE ABSENCE OF PROSTHETIC CARDIAC VALVE THROMBOSIS) WILL TYPICALLY RESULT FROM ON-GOING OR WORSENING REGURGITATION, VALVE DEGENERATION RELATED TO THE FORMATION OF CALCIFICATION OR PANNUS, OR VALVE MIGRATION. THESE EVENTS ARE IDENTIFIED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU) AS POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO A NUMBER OF ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. CALCIFICATION IS A WELL-RECOGNIZED FAILURE MODE OF BIOPROSTHETIC VALVES. THE MECHANISM OF CALCIFICATION OF BIOMATERIALS IS NOT COMPLETELY UNDERSTOOD BUT IS PROBABLY RELATED TO AN INABILITY OF THE NON-VIABLE CELLS TO MAINTAIN THEIR NORMALLY LOW INTRACELLULAR CONCENTRATION OF CALCIUM. THE MECHANISM BEHIND WORSENING OR LATE PVL IS NOT WELL UNDERSTOOD BUT MAY BE RELATED TO CARDIAC REMODELING. ACCORDING TO LITERATURE REVIEW, VALVE MIGRATION RESULTS WHEN FORCES ACTING ON THE TRANSCATHETER HEART VALVE (THV) OVERCOME THE STRENGTH OF ATTACHMENT OF THE VALVE TO THE AORTIC WALL. STENT VALVES ARE SUBJECTED TO ANTEGRADE EJECTION FORCES DURING SYSTOLE. LESS-THAN-SEVERE AND NON-UNIFORMLY DISTRIBUTED CALCIFICATION OF THE NATIVE LEAFLETS, INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION CAN CONTRIBUTE TO VALVE MIGRATION. ADDITIONALLY, RESIDUAL OVERHANGING LEAFLETS CAN EXERT DOWNWARDS FORCE DURING DIASTOLE, CAUSING MIGRATION OF THE THV TOWARDS THE LEFT VENTRICLE IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256940&pc=NPT | 2015691-2025-04589 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 7 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 7 AORTIC PERCUTANEOUS CORONARY INTERVENTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 63 - 74 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 5 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 7 PERCUTANEOUS CORONARY INTERVENTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0.43 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), CORONARY FLOW OBSTRUCTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT THE SAFETY AND EFFECTIVENESS HAVE NOT BEEN ESTABLISHED IN PATIENTS WITH BULKY CALCIFIED AORTIC VALVE LEAFLETS IN CLOSE PROXIMITY TO THE CORONARY OSTIA. IN ADDITION, IT WARNS THAT CAUTION SHOULD BE EXERCISED IN IMPLANTING A BIOPROSTHESIS IN PATIENTS WITH CLINICALLY SIGNIFICANT CORONARY ARTERY DISEASE. CORONARY OBSTRUCTION CAN RESULT IN MYOCARDIAL ISCHEMIA OR INFARCTION DUE TO OBSTRUCTION OF THE CORONARY BLOOD FLOW AND MAY REQUIRE INTERVENTION (E.G., PCI). THERE ARE MULTIPLE PATIENT FACTORS THAT COULD PUT THE PATIENT AT RISK FOR CORONARY FLOW OBSTRUCTION DURING THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING SIGNIFICANT UNDERLYING CORONARY ARTERY DISEASE AND BULKY CALCIFICATION OF THE NATIVE LEAFLETS AND ROOT. DISPLACEMENT OF CALCIUM DEPOSITS WITH EMBOLIZATION OF DEBRIS INTO ONE OF THE ARTERIES, OR AORTIC DISSECTION WITH CONTINUITY OF THE RUPTURE INTO THE INTIMA OF ONE OF THE CORONARY OSTIA, CAN RESULT IN THIS COMPLICATION. ADDITIONALLY, MINIMAL DISTANCE BETWEEN THE NATIVE ANNULUS AND THE CORONARY OSTIA, BULKY CALCIFICATION, LONG NATIVE LEAFLETS, AND OBLITERATED CORONARY SINUSES. PROCEDURAL FACTORS SUCH AS TORN NATIVE LEAFLET DURING BALLOON VALVULOPLASTY, PLAQUE SHIFT, DEPLOYMENT OF THE BIOPROSTHETIC HEART VALVE TOO AORTIC, AND SIGNIFICANT VALVE OVER SIZING COULD ALSO CONTRIBUTE TO THIS COMPLICATION. THE EDWARDS THV MANUALS ADVISE THE OPERATOR ON PRE-PROCEDURE ASSESSMENT OF THE AORTIC VALVE, ROOT, AND CORONARY ANATOMY. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256941&pc=DYB | 2015691-2025-04574 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Insufficient Information | Unspecified Vascular Problem | K200258 | E2016006 | 9 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 9 AORTIC VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 63 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALE(S) AND 5 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 9 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 1.78 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256943&pc=NPT | 2015691-2025-04587 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Heart Block | P140031 | E2016006 | 11 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 11 AORTIC PERMANENT PACEMAKER DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 56 - 97 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 6 FEMALE(S) AND 5 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 11 PERMANENT PACEMAKER DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S) FOR THIS EVENT IS 10.18 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, CONDUCTION SYSTEM DEFECTS (HEART BLOCK) ARRHYTHMIAS AND CONDUCTION SYSTEM DEFECTS THAT MAY REQUIRE A PERMANENT PACEMAKER ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH BALLOON AORTIC VALVULOPLASTY, DEPLOYMENT OF THE PROSTHETIC VALVE, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC) GUIDELINES, THE CLOSE ANATOMICAL RELATIONSHIP BETWEEN THE AORTIC VALVE COMPLEX AND THE BRANCHING ATRIOVENTRICULAR BUNDLE MAY PROVIDE AN EXPLANATION FOR THESE COMPLICATIONS OF THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY -CONDUCTION DISTURBANCES/ HEART BLOCK WRITTEN BY EDWARDS LIFESCIENCES, ATRIOVENTRICULAR CONDUCTION DISTURBANCES AFTER THV PLACEMENT ARE ASSOCIATED WITH MANY PATIENT RELATED AND PROCEDURAL RELATED FACTORS, INCLUDING PRE-OPERATIVE CO-MORBID STATUS, THE DEGREE AND BULKINESS OF AORTIC VALVE AND ANNULAR CALCIFICATION, INTER-VENTRICULAR SEPTAL THICKNESS, PRE-EXISTING ELECTROCARDIOGRAM ABNORMALITIES, THE DEPTH OF PROSTHESIS IMPLANTATION, AND THE PROFILE OF THE IMPLANTED PROSTHESIS. UNLIKE CONVENTIONAL AVR, WHERE THERE MAY BE LOCALIZED TRAUMA DUE TO DECALCIFICATION OF THE ANNULUS AND/OR SUTURE PLACEMENT IN THE PROXIMITY OF THE AV NODE OR THE BUNDLES, TRANSCATHETER VALVE REPLACEMENT MAY CAUSE CONDUCTION ABNORMALITIES THROUGH MECHANICAL IMPINGEMENT OF THE CONDUCTION SYSTEM BY THE PROSTHESIS. THE MECHANISMS OF THE DEVELOPMENT OF HEART BLOCK AFTER THV PLACEMENT ARE WELL DOCUMENTED AND DESCRIBED IN THE LITERATURE. IT IS ALSO DOCUMENTED THAT PRE-EXISTING HEART BLOCK IS COMMON IN PATIENTS UNDERGOING THV PLACEMENT OR SURGICAL AVR AND ANOTHER 4-6 % WILL DEVELOP POSTOPERATIVE HEART BLOCK, POTENTIALLY REQUIRING A PERMANENT PACEMAKER. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE, HOWEVER THE CONDUCTION DISTURBANCE IS LIKELY RELATED TO THE MECHANISM DESCRIBED ABOVE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256944&pc=NPU | 2015691-2025-04579 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPU | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Obstruction/Occlusion | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 2 MITRAL LEFT VENTRICULAR OUTFLOW TRACT OBSTRUCTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 82 - 85 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 LEFT VENTRICULAR OUTFLOW TRACT OBSTRUCTION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 1 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE ARE: (B)(4).
PER THE INSTRUCTIONS FOR USE, VALVE MALPOSITION REQUIRING INTERVENTION IS A POTENTIAL ADVERSE EVENT ASSOCIATED WITH THE USE OF TRANSCATHETER HEART VALVES. OBSTRUCTION OF THE LEFT VENTRICULAR OUTFLOW TRACT CAN BE CAUSED BY PATIENT FACTORS (ANTERIOR MITRAL LEAFLET PROTRUDING INTO THE LVOT, SEPTAL BULGE) OR PROCEDURAL FACTORS (POSITIONING OF THE VALVE FRAME WITHIN THE ANNULUS). DEPENDING ON THE DEGREE OF OBSTRUCTION, CARDIAC OUTPUT AND HEMODYNAMIC STABILITY MAY BE AFFECTED. INACCURATE DEPLOYMENT IS GENERALLY A RESULT OF USE ERROR OR A COMBINATION OR PATIENT AND PROCEDURAL FACTORS. IN SOME CASES, LVOT OBSTRUCTION COULD RESULT IN CLINICALLY SIGNIFICANT HEMODYNAMIC COMPROMISE THAT MAY REQUIRE EXPLANTATION OF THE THV WITH SURGICAL CORRECTION. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256946&pc=NPU | 2015691-2025-04583 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPU | EDWARDS COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 1 MITRAL TRANSSEPTAL COMPLICATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 85 - 85 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 0 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 TRANSSEPTAL COMPLICATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 2 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU) CARDIOVASCULAR INJURY, SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES, IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH TRANSCATHETER HEART VALVES. TRANSEPTAL APPROACHES ALLOW TRANSFEMORAL (VENOUS) ACCESS TO THE MITRAL VALVE BY PUNCTURING THE ATRIAL SEPTUM AND ADVANCING THE DELIVERY SYSTEM THROUGH THE OPENING. USUALLY, THE SEPTOSTOMY CLOSES ON ITS OWN OVER TIME AND DOES NOT REQUIRE TREATMENT. IN SOME CASES, THE INTERVENTIONALIST WILL ELECT TO CLOSE THE CREATED ATRIAL SEPTAL DEFECT (ASD) DURING THE INITIAL PROCEDURE PERCUTANEOUSLY WITH A SEPTAL OCCLUDER. PERSISTENT IATROGENIC ATRIAL SEPTAL DEFECTS (IASD) ARE NOT UNCOMMON AND MAY BE ASSOCIATED WITH THE USE OF LARGER SHEATHS. THESE MAY REQUIRE PLACEMENT OF A CLOSURE DEVICE IN A REPEAT PROCEDURE AND ARE CONSIDERED SERIOUS INJURIES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256948&pc=NPT | 2015691-2025-04586 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P140031 | E2016006 | 14 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 14 AORTIC STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 61 - 97 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALE(S) AND 10 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 14 STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 8 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS INCLUDING STROKE ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH INVASIVE CARDIAC PROCEDURE INCLUDING TRANSCATHETER HEART VALVE PROCEDURES AND THE USE OF THE EDWARDS THV DEVICES. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING TRANSCATHETER VALVE REPLACEMENT PROCEDURE. RISK FACTORS CORRELATING WITH A NUMBER OF PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE DEMONSTRATED THAT THE MAJORITY OF PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO IMPORTANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND AVR PATIENTS. AFTER TRANSCATHETER VALVE REPLACEMENT PROCEDURE, THERE APPEARS TO BE A MORE SIGNIFICANT PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT TRANSCATHETER VALVE REPLACEMENT PROCEDURE PATIENTS WITH MULTIPLE CO MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE HOWEVER, THE EVENT MAY BE RELATED TO THE MECHANISM DESCRIBED ABOVE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256949&pc=NPT | 2015691-2025-04593 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P140031 | E2016006 | 3 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 3 AORTIC STROKE - UNDETERMINED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 84 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 2 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 STROKE - UNDETERMINED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 5.67 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4).
PER THE INSTRUCTIONS FOR USE, PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS INCLUDING STROKE ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH INVASIVE CARDIAC PROCEDURE INCLUDING TRANSCATHETER HEART VALVE PROCEDURES AND THE USE OF THE EDWARDS THV DEVICES. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING TRANSCATHETER VALVE REPLACEMENT PROCEDURE. RISK FACTORS CORRELATING WITH SEVERAL PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURE ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING TRANSCATHETER VALVE REPLACEMENT DEMONSTRATED THAT THE MAJORITY OF PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE AORTIC VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO IMPORTANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN TRANSCATHETER VALVE REPLACEMENT AND AVR PATIENTS. AFTER TAVR, THERE APPEARS TO BE A MORE SIGNIFICANT PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT TRANSCATHETER VALVE REPLACEMENT PROCEDURE PATIENTS WITH MULTIPLE CO MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THE EVENT IS LISTED AS UNDETERMINED STROKE. IT IS NOT KNOWN IF THE MECHANISM OF THE STROKE IS IN QUESTION, OR IF THE STROKE ITSELF IS IN QUESTION. THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. IT IS POSSIBLE THAT THE MECHANISM DESCRIBED ABOVE MAY HAVE CONTRIBUTED TO THE EVENT. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256950&pc=DYB | 2015691-2025-04581 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | K200258 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 2 MITRAL VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 76 - 82 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 1 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 2 VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 84.5 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256951&pc=NPU | 2015691-2025-04607 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPU | EDWARDS COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 2 MITRAL ASD DEFECT CLOSURE DUE TO TRANSSEPTAL CATHETERIZATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 80 - 81 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 ASD DEFECT CLOSURE DUE TO TRANSSEPTAL CATHETERIZATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 194 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU) CARDIOVASCULAR INJURY, SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES, IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH TRANSCATHETER HEART VALVES. TRANSEPTAL APPROACHES ALLOW TRANSFEMORAL (VENOUS) ACCESS TO THE MITRAL VALVE BY PUNCTURING THE ATRIAL SEPTUM AND ADVANCING THE DELIVERY SYSTEM THROUGH THE OPENING. USUALLY, THE SEPTOSTOMY CLOSES ON ITS OWN OVER TIME AND DOES NOT REQUIRE TREATMENT. IN SOME CASES, THE INTERVENTIONALIST WILL ELECT TO CLOSE THE CREATED ATRIAL SEPTAL DEFECT (ASD) DURING THE INITIAL PROCEDURE PERCUTANEOUSLY WITH A SEPTAL OCCLUDER. PERSISTENT IATROGENIC ATRIAL SEPTAL DEFECTS (IASD) ARE NOT UNCOMMON AND MAY BE ASSOCIATED WITH THE USE OF LARGER SHEATHS. THESE MAY REQUIRE PLACEMENT OF A CLOSURE DEVICE IN A REPEAT PROCEDURE AND ARE CONSIDERED SERIOUS INJURIES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256953&pc=DYB | 2015691-2025-04596 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Insufficient Information | Unspecified Vascular Problem | K200258 | E2016006 | 16 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 16 AORTIC VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 27 - 86 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 8 FEMALE(S) AND 8 MALE(S) Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 16 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 14.69 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4).
ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256954&pc=DYB | 2015691-2025-04595 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | K200258 | E2016006 | 23 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 23 AORTIC VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 27 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 5 FEMALE(S) AND 18 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 23 VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 29.17 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256955&pc=NPU | 2015691-2025-04608 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPU | EDWARDS COMMANDER DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 2 MITRAL CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 58 - 101 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 2 CARDIAC PERFORATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 11 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS COMMANDER DELIVERY SYSTEM ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, CARDIOVASCULAR INJURIES SUCH AS PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, MYOCARDIAL OR VALVULAR STRUCTURES ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND MAY REQUIRE INTERVENTION. THERE ARE SEVERAL POTENTIAL ETIOLOGIES FOR CARDIAC PERFORATION DURING A TRANSCATHETER VALVE REPLACEMENT PROCEDURE INCLUDING PERFORATION BY THE GUIDEWIRE, THE DELIVERY SYSTEM, OR THE TRANSVENOUS PACER (TVP) LEAD. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE TRANSCATHETER HEART VALVES. TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, FIXATION OF THE TVP TO PREVENT VENTRICLE PERFORATION, AND CAREFUL MANIPULATION OF DEVICES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256956&pc=DYB | 2015691-2025-04606 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | DYB | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | Adverse Event Without Identified Device or Use Problem | Unspecified Vascular Problem | K200258 | E2016006 | 3 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 3 MITRAL VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 58 - 80 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 3 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 VASCULAR COMPLICATION - MAJOR DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 92.33 DAY(S). DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256957&pc=NPU | 2015691-2025-04605 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPU | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 MITRAL STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 83 - 83 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 1 STROKE - ISCHEMIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 4 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE, PERMANENT OR TRANSIENT NEUROLOGICAL EVENTS INCLUDING STROKE ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSCATHETER VALVE REPLACEMENT PROCEDURE AND THE USE OF THV DEVICES. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY BY EDWARDS LIFESCIENCES, STROKE IS RECOGNIZED IN THE LITERATURE AS A WELL-KNOWN COMPLICATION IN A SMALL NUMBER OF PATIENTS UNDERGOING TRANSCATHETER VALVE REPLACEMENT. RISK FACTORS CORRELATING WITH A NUMBER OF PATIENT CO-MORBIDITIES HAVE BEEN IDENTIFIED. ALTHOUGH IN MANY CASES THE ROOT CAUSE OF THE EVENT IS UNABLE TO BE DETERMINED, STROKES DURING TRANSCATHETER VALVE REPLACEMENT PROCEDURES ARE UNDOUBTEDLY MULTIFACTORIAL, THE DOMINANT ETIOLOGY LIKELY BEING INTRA-PROCEDURE EMBOLIC EVENTS. A TRANSCRANIAL DOPPLER STUDY DURING TRANSCATHETER VALVE REPLACEMENT DEMONSTRATED THAT THE MAJORITY OF PROCEDURAL EMBOLIC EVENTS OCCURRED DURING BALLOON VALVULOPLASTY, MANIPULATION OF CATHETERS ACROSS THE VALVE, AND VALVE IMPLANTATION. AN ANALYSIS IN PATIENTS UNDERGOING VALVE SURGERY REVEALED FOUR BASELINE CHARACTERISTICS AND TWO PROCEDURAL EVENTS THAT WERE ASSOCIATED WITH EARLY POST-PROCEDURE STROKE: FEMALE SEX, EF < 30%, DIABETES, AGE OLDER THAN 70 YEARS, BYPASS PROCEDURE TIME> 120 MIN, AND CALCIFICATION OF THE ASCENDING AORTA. PREDICTORS OF LATE STROKE HAVE INCLUDED FEMALE SEX, AGE OLDER THAN 75 YEARS, ATRIAL FIBRILLATION, AND A HISTORY OF OR CURRENT SMOKING. THERE WERE NO IMPORTANT DIFFERENCES IN THE FREQUENCY OF LATE STROKES BETWEEN TRANSCATHETER VALVE REPLACEMENT AND AVR PATIENTS. AFTER TRANSCATHETER VALVE REPLACEMENT, THERE APPEARS TO BE A MORE SIGNIFICANT PROPORTION OF EARLY STROKES OCCURRING < 24 H POST-PROCEDURE, BUT TRANSCATHETER VALVE REPLACEMENT PATIENTS WITH MULTIPLE CO MORBIDITIES ARE PROBABLY AT HIGHER RISK OF BOTH EARLY AND LATE STROKES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256958&pc=NPU | 2015691-2025-04603 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPU | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 MITRAL DEVICE EMBOLIZATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS (B)(6). YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 DEVICE EMBOLIZATION DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 0 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), VALVE EMBOLIZATION IS A KNOWN POTENTIAL COMPLICATION ASSOCIATED WITH THE USE OF TRANSCATHETER HEART VALVES. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE/DELIVERY SYSTEM, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, INACCURATE MEASUREMENT OF THE LANDING ZONE, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. LATE VALVE EMBOLIZATIONS ARE LIKELY RELATED TO VALVE UNDER SIZING, MALPOSITION, OR INCOMPLETE EXPANSION, OR LACK OR ANCHORING MECHANISM (CHARACTERISTICS OF THE PRE-EXISTING VALVE OR RING). IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256959&pc=NPU | 2015691-2025-04602 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPU | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Insufficient Information | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P140031 | E2016006 | 2 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 2 MITRAL CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 58 - 101 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER: E2016006 AND SUMMARIZES 2 CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 11 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST REOPERATION, INCLUDING EMERGENCY CARDIAC SURGERY, AS ADDITIONAL POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THIS EVENT TYPE (CARDIAC SURGERY OR INTERVENTION - OTHER UNPLANNED E031) IS DEFINED AS: THE PATIENT SUBSEQUENTLY UNDERWENT OTHER TYPES OF CARDIAC SURGERY OR CATH LAB INTERVENTIONS (NOT OTHERWISE SPECIFIED) THAT WAS UNPLANNED. THIS COULD INCLUDE REOPERATION FOR DAMAGE TO ADJACENT STRUCTURES, OR POTENTIALLY VALVE EXPLANT RELATED TO EMBOLIZATION, OR OTHER EVENT TYPES. CARDIOVASCULAR INJURY, INCLUDING PERFORATION OR DISSECTION OF VALVULAR STRUCTURES AND VALVE EMBOLIZATION REQUIRING INTERVENTION ARE KNOWN POTENTIAL COMPLICATIONS LISTED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU). PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVES (THV). TRAINING INCLUDES PROPER GUIDEWIRE POSITIONING, AND CAREFUL MANIPULATION OF DEVICES. PER THE THV TA TRAINING MANUALS, AFTER GAINING LV ACCESS WITH THE 0.035" SOFT GUIDEWIRE, THE WIRE SHOULD BE JIGGLED UNDER TEE IMAGING OF THE MITRAL VALVE. AN INCREASE IN MR SUGGESTS WIRE ENTANGLEMENT IN THE MITRAL SUB-VALVULAR APPARATUS. IF ENTANGLEMENT IS SUSPECTED, THE GUIDEWIRE SHOULD BE COMPLETELY REMOVED FROM THE VENTRICLE; THE OPERATOR SHOULD CHANGE DIRECTION OF THE NEEDLE AND REINSERT THE GUIDEWIRE INTO THE VENTRICLE, CHECKING AGAIN FOR WIRE ENTANGLEMENT. THE TF TRAINING MANUAL ALSO PROVIDES GUIDANCE FOR VALVE CROSSING AND WIRE EXCHANGE: EXCHANGE 0.035" AMPLATZ EXTRA-STIFF WIRE WITH PRE-SHAPED DISTAL END IN LEFT VENTRICLE. ENSURE THE GUIDE CATHETER IS ADVANCED UPON WIRE EXCHANGE TO ENSURE EXCHANGE WIRE DOES NOT GET CAUGHT IN THE MITRAL CHORDAE. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO VENTRICULAR EMBOLIZATION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE/DELIVERY SYSTEM, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. THE THV TRAINING MANUALS INSTRUCT THE OPERATOR ON PROPER POSITIONING AND DEPLOYMENT OF THE VALVE, INCLUDING ALL PROCEDURAL AND ANATOMICAL CONSIDERATIONS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22256960&pc=NPU | 2015691-2025-04604 | 2025-04-01 04:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-17 04:00 | NPU | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Mitral Valve Insufficiency/ Regurgitation | P140031 | E2016006 | 1 | Event Description: THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 2 2025 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 MITRAL REINTERVENTION - MITRAL VALVE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 81 - 81 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 0 MALE(S). Manufacturer Narrative: THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 REINTERVENTION - MITRAL VALVE DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAY(S)) FOR THIS EVENT IS 68 DAY(S). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE ARE: (B)(4). INTRA-PROCEDURAL AND IN-HOSPITAL MITRAL VALVE RE-INTERVENTION WILL TYPICALLY RESULT FROM VALVE MALPOSITION OR REGURGITATION (PVL OR CENTRAL, INCLUDING LEAFLET RESTRICTION). THESE CONDITIONS ARE KNOWN POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV, DELIVERY SYSTEM, AND/OR ACCESSORIES. THERE ARE MULTIPLE PATIENT AND PROCEDURAL FACTORS THAT ALONE OR IN COMBINATION CAN CAUSE OR CONTRIBUTE TO VALVE MALPOSITION, INCLUDING IMPROPER POSITIONING PRIOR TO DEPLOYMENT, POOR IMAGE INTENSIFIER ANGLE, POOR COAXIAL ALIGNMENT OF THE VALVE AND DELIVERY SYSTEM, CHARACTERISTICS OF THE PRE-EXISTING VALVE OR RING, LOSS OF PACING CAPTURE, RAPID DEPLOYMENT, RELEASE OF STORED TENSION DURING DEPLOYMENT, AND MOVEMENT OF THE DELIVERY SYSTEM BY THE OPERATOR. VALVE MALPOSITION HAS THE POTENTIAL TO CONTRIBUTE TO SUBOPTIMAL COAPTATION OF THE THV VALVE LEAFLETS AND CAUSE CENTRAL INSUFFICIENCY; IT CAN LEAD TO MIGRATION OR EMBOLIZATION OF THE PROSTHESIS. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE THV. TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS AND PROCTORED PROCEDURES. THE CORRECT ALIGNMENT AND POSITIONING OF THE DEVICE AT THE POINT OF DEPLOYMENT IS EMPHASIZED AS A KEY FACTOR TO THE PLACEMENT AND FIXATION OF THE DEVICE. OPERATORS ARE ALSO INSTRUCTED TO USE FLUOROSCOPY AS THE PRIMARY METHOD OF VISUALIZATION FOR POSITIONING AND DEPLOYMENT. MITRAL VALVE RE-INTERVENTION IN THE FOLLOW-UP PERIOD (IN THE ABSENCE OF PROSTHETIC CARDIAC VALVE THROMBOSIS) WILL TYPICALLY RESULT FROM ON-GOING OR WORSENING REGURGITATION, VALVE DEGENERATION RELATED TO THE FORMATION OF CALCIFICATION OR PANNUS, OR VALVE MIGRATION. THESE EVENTS ARE IDENTIFIED IN THE PRODUCT INSTRUCTIONS FOR USE (IFU) AS POTENTIAL RISKS ASSOCIATED WITH THE USE OF THE THV. REGURGITATION WHICH DEVELOPS PROGRESSIVELY OVER TIME CAN BE DUE TO SEVERAL ISSUES INCLUDING PATIENT RELATED FACTORS OR STRUCTURAL VALVE DETERIORATION, INCLUDING CALCIFICATION, NON-CALCIFIC DEGENERATION, LEAFLET THICKENING OR FIBROSIS, OR A COMBINATION OF THESE. REGURGITATION MAY ALSO DEVELOP PROGRESSIVELY IF HOST FIBROTIC TISSUE, OR PANNUS, GROWS ONTO THE BIOPROSTHETIC VALVE. PANNUS, A CAUSE OF NONSTRUCTURAL DYSFUNCTION, MAY INTERFERE WITH FUNCTIONALITY OF THE DEVICE BY RESTRICTING THE LEAFLET MOTION LEADING TO ABNORMAL COAPTATION. CALCIFICATION IS A WELL-RECOGNIZED FAILURE MODE OF BIOPROSTHETIC VALVES. THE MECHANISM OF CALCIFICATION OF BIOMATERIALS IS NOT COMPLETELY UNDERSTOOD BUT IS PROBABLY RELATED TO AN INABILITY OF THE NON-VIABLE CELLS TO MAINTAIN THEIR NORMALLY LOW INTRACELLULAR CONCENTRATION OF CALCIUM. ANOTHER POTENTIAL CONDITION THAT MAY LEAD TO RE-INTERVENTION IS PATIENT PROSTHESIS MISMATCH AND/OR UNDER-EXPANDED VALVES DUE TO THV RESTRICTION FROM THE EXISTING VALVE FRAME. INCORRECT BIOPROSTHETIC VALVE SIZING, AND INCOMPLETE FRAME EXPANSION, CAN CONTRIBUTE TO VALVE MIGRATION IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22208983&pc=DSP | 2249723-2025-0002601 | 2025-05-27 04:00 | Death | DATASCOPE CORP. | 2025-06-16 04:00 | DSP | CARDIOSAVE HYBRID, TYPE B PLUG | Overheating of Device; Incorrect, Inadequate or Imprecise Result or Readings | Ischemic Heart Disease; Mitral Valve Insufficiency/ Regurgitation | K112372 |  | 1 | Event Description: IT WAS REPORTED BY THE DISTRIBUTOR THAT CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) EXPERIENCED OVERHEATING AND POWER ON TEST FAILURE 3 DURING USE ON PATIENT. THE DEVICE WAS TURNED OFF, WAITED 10 MINUTES, AND TURNED ON AGAIN. HOWEVER, THE FAULT PERSISTED. THE PATIENT DIED LATER THAN 4 HOURS AFTER FAILURE. IT HAS NOT BEEN PROVEN THAT THE EQUIPMENT WAS RESPONSIBLE FOR THE PATIENT'S DEATH. Manufacturer Narrative: DUE TO CHARACTER LIMITATION IN E1, FULL INITIAL REPORTER NAME IS (B)(6). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22209771&pc=PNJ | 9612164-2025-02936 | 2025-05-27 04:00 | Death | MEDTRONIC IRELAND | 2025-06-16 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P150033 |  | 1 | Event Description: THE PATIENT WAS REPORTED TO HAVE DIED FROM CARDIAC PERFORATION OF THE RIGHT VENTRICLE. AN ECHOCARDIOGRAM WAS PERFORMED TO CONFIRM THE PERFORATION, AFTER WHICH THE PATIENT CODED. PERICARDIOCENTESIS AND CARDIOPULMONARY RESUSCITATION WERE ALSO CONDUCTED. Manufacturer Narrative: CORRECTION OF SYSTEM REPORT: H3 SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA-DELSYS] PRODUCT ID MC2AVR1-DELSYS (SERIAL: (B)(6)). D9: YES, RETURN DATE: 06/02/2025> H3: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT ID MC2AVR1-DELSYS PRODUCT EVENT SUMMARY: THE FULL DELIVERY SYSTEM WAS RETURNED AND ANALYZED. THE LUMEN OF THE DELIVERY SYSTEM WAS TORN. THE DELIVERY SYSTEM INNER SHAFT WAS MECHANICALLY KINKED/BENT. THE ANALYST NOTED THE FULL DELIVERY SYSTEM WAS RETURNED WITH THE DEVICE INTACT IN THE DEVICE CUP. THERE WERE NO ANOMALIES FOUND WITH THE DISTAL EDGE OF THE DEVICE CUP. THE INNER SHAFT WAS KINKED AT 1.5 CM FROM THE DISTAL END OF THE RECAPTURE CONE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. CORRECTION: G3 FIELD FROM PREVIOUS FOLLOW-UP REPORT 001 WAS OMITTED. THE OMITTED DATE SHOULD HAVE BEEN 16-JUN-2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA-DELSYS] PRODUCT ID: MC2AVR1-DELSYS (SERIAL: (B)(6). D9: YES, RETURN DATE: 06/02/2025> H3: YES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22209978&pc=PNJ | 9612164-2025-02942 | 2025-05-27 04:00 | Death | MEDTRONIC IRELAND | 2025-06-16 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Chest Pain; Dyspnea; Sepsis; Discomfort; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT POST OPERATIVE OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT EXPERIENCED DISCOMFORT AND CHEST PRESSURE. AN ELECTROCARDIOGRAM (ECG) CONFIRMED A PERICARDIAL EFFUSION AND A PERICARDIOCENTESIS WAS PERFORMED. THE LEADLESS IPG REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT EXPERIENCED SEPSIS AND RENAL FAILURE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] SERIAL:(B) (6). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22210027&pc=DYB | 9612164-2025-02943 | 2025-05-27 04:00 | Death | MEDTRONIC IRELAND | 2025-06-16 04:00 | DYB | MICRA | Use of Device Problem | Cardiac Perforation | K132030 |  | 1 | Event Description: THE PATIENT WAS REPORTED TO HAVE DIED FROM CARDIAC PERFORATION OF THE RIGHT VENTRICLE. AN ECHOCARDIOGRAM WAS PERFORMED TO CONFIRM THE PERFORATION, AFTER WHICH THE PATIENT CODED. PERICARDIOCENTESIS AND CARDIOPULMONARY RESUSCITATION WERE ALSO CONDUCTED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22210148&pc=BZD | 2518422-2025-106642 | 2025-06-12 04:00 | Death | RESPIRONICS, INC. | 2025-06-16 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22210268&pc=BZD | 2518422-2025-106643 | 2025-06-12 04:00 | Death | RESPIRONICS, INC. | 2025-06-16 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. MULTIPLE GOOD FAITH EFFORTS HAVE BEEN MADE TO OBTAIN ADDITIONAL INFORMATION, WITHOUT CUSTOMER RESPONSE. A FINAL REPORT IS BEING SUBMITTED. IF NEW INFORMATION BECOMES AVAILABLE FOLLOWING THE COMPLETION OF THE DEVICE REPAIR THAT CHANGES THE OUTCOME OF THE INVESTIGATION AND ASSOCIATED MEDICAL DEVICE REPORTING A FOLLOW UP/SUPPLEMENTAL REPORT WILL BE COMPLETED. THE MANUFACTURER HAS UPDATED SECTION H IN THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22210311&pc=DSQ | 2916596-2025-03664 | 2025-05-26 04:00 | Death | THORATEC CORPORATION | 2025-06-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO REFRACTORY MULTIPLE ORGAN FAILURE (MOF) UNDER SEPSIS. PATIENT SUFFERED FROM A SEPTIC MOF WHICH WAS RELATED TO THE OUTCOME. NO DEVICE RELATED ISSUE, NEITHER DRIVELINE NOR DEVICE RELATED INFECTIONS WERE RELATED TO THE OUTCOME. THE ROOT CAUSE OF THE EVENT WAS NOT PROVIDED. THE LEFT VENTRICULAR ASSIST DEVICE WAS RUNNING AS EXPECTED. IT WAS NOTED THAT THE PATIENT'S OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: SECTION A: PATIENT INFORMATION WAS NOT PROVIDED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV C, AND THE HEARTMATE 3 PATIENT HANDBOOK, REV B, ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH, SEPSIS, AND MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION. ALTHOUGH ONLY SEPSIS WAS REPORTED BY THE ACCOUNT, SEVERAL SECTIONS OF THE HEARTMATE 3 LVAS IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22210456&pc=OZD | 1220648-2025-29303 | 2025-05-20 04:00 | Death | ABIOMED, INC. | 2025-06-16 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Thrombocytopenia | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A PATIENT DEVELOPED HEPARIN INDUCED THROMBOCYTOPENIA DURING IMPELLA 5.5 SUPPORT. TREATMENT WAS SWITCHED TO ARGATROBAN IN RESPONSE TO THE CONDITION. THE PATIENT'S UNDERLYING CONDITION FAILED TO IMPROVE FIVE DAYS LATER AND THE DECISION WAS MADE TO WITHDRAW CARE. THE PATIENT PASSED AWAY SHORTLY AFTER. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22210566&pc=OZD | 1220648-2025-29305 | 2025-05-20 04:00 | Death | ABIOMED, INC. | 2025-06-16 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Stroke/CVA; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT EXPERIENCED MULTIPLE COMPLICATIONS DURING IMPELLA 5.5 SUPPORT. ROUGHLY ONE DAY INTO SUPPORT, THE PATIENT EXPERIENCED PERSISTENT BLEEDING. THE PATIENT WAS TAKEN TO THE OPERATING ROOM FOR A WASHOUT AND SUPPORT CONTINUED WITH THE CURRENT SETUP. THE FOLLOWING DAY, THE PATIENT ENDURED A STROKE. THE PATIENT REMAINED SEDATED ON ECPELLA SUPPORT. ON THAT SAME DAY, THE PUMP WAS REPOSITIONED DUE TO SEVERAL RUNS OF VENTRICULAR TACHYCARDIA AND ECTOPY. THE P-LEVEL WAS LIMITED DUE TO THE PATIENT'S CARDIAC FUNCTION. THE DECISION WAS ULTIMATELY MADE TO WITHDRAW CARE AND THE PATIENT PASSED AWAY. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22210653&pc=BZD | 2518422-2025-047340 | 2023-12-19 05:00 | Death | RESPIRONICS, INC. | 2025-06-16 04:00 | BZD | DREAMSTATION AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT HER HUSBAND PASSED AWAY AND THAT SHE NEEDED TO KNOW WHEN HE RECEIVED HIS REPLACEMENT DEVICE FOR BILLING PURPOSES FOR A LOANER DEVICE HE HAD USED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22210829&pc=DSQ | 2916596-2025-03989 | 2025-05-20 04:00 | Death | THORATEC CORPORATION | 2025-06-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO INTRACRANIAL HEMORRHAGE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22210971&pc=MIH | 2247858-2025-00146 | 2025-05-18 04:00 | Death | BOLTON MEDICAL, INC. | 2025-06-16 04:00 | MIH | RELAY PRO THORACIC STENT-GRAFT SYSTEM | Sharp Edges | Vascular Dissection | P200045 |  | 1 | Event Description: "PATIENT DEVELOPED RETROGRADE TYPE A DISSECTION ON (B)(6) 2025 AND HAD SURGERY URGENT TO REPAIR IT. THE BARE STENT ON THE INFERIOR PORTION OF THE ARCH SEEMED TO BE THE CAUSE ACCORDING TO THE CT SURGEON WHO PERFORMED THE SURGERY. PATIENT EXPIRED THIS MORNING (B)(6) 2025". PATIENT OUTCOME: "DEATH". Manufacturer Narrative: BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22210982&pc=MIH | 9612164-2025-02948 | 2025-03-21 04:00 | Death | MEDTRONIC IRELAND | 2025-06-16 04:00 | MIH | VALIANT CAPTIVIA - FF | Leak/Splash; Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Rupture; Insufficient Information | P100040 |  | 1 | Event Description: A VALIANT CAPTIVIA STENT GRAFT WAS IMPLANTED IN THE ENDOVASCULAR TREATMENT OF A PAU. IT WAS REPORTED THAT THE PATIENT RETURNED FIVE DAYS LATER WITH CONCERNS FOR A POSSIBLE RUPTURE. IT WAS BELIEVED AFTER A CT SCAN WAS PERFORMED, THE AORTA MAY BE MYCOTIC AND WOULD NOW BE RELINED WITH A NON-MEDTRONIC STENT GRAFT. THE PATIENT THEN PASSED AWAY ON AN UNKNOWN DATE. THE CAUSE OF DEATH IS UNKNOWN BUT IT WAS SAID THE PATIENT WAS VERY SICK WITH OTHER CO MORBIDITIES AND HAD PRE-EXISTING ADVANCED UNKNOWN TYPE OF CANCER. A POSSIBLE RUPTURE AND POSSIBLE INFECTION WERE MENTIONED AS WELL AS CONCERNS OF POSSIBLE MYCOTIC INFECTION DETERIORATING THE VALIANT CAPTIVIA GRAFT. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5: IT WAS REPORTED THAT PER THE PHYSICIAN, THE VALIANT CAPTIVIA DID NOT CAUSE OR CONTRIBUTED TO THE PATIENT'S DEATH. THE VALIANT CAPTIVIA WAS RELINED WITH A NON-MEDTRONIC GRAFT DUE TO PTFE MATERIAL WANTED IN THE EVENT THAT THE VALIANT CAPTIVIA WERE POSSIBLY INFECTED. FILM ANALYSIS THE REPORTED ANEURYSM RUPTURE COULD NOT BE CONFIRMED, AS THE PROVIDED IMAGES WERE FROM BEFORE THE IMPLANTATION AND AFTER THE INTERVENTION PERFORMED TO TREAT THE RUPTURE. THE CAUSE OF THE PATIENT'S DEATH COULD NOT BE ASSESSED. THE RADIOLOGY REPORT FROM ONE WEEK POST-INDEX PROCEDURE STATED: "ON THE NON-CONTRAST STUDY, MODERATE TO SEVERE CALCIFIED ATHEROMATOUS PLAQUE WITHIN THE AORTA AND ITS BRANCHES IS NOTED. HIGH-DENSITY BLOOD IS NOTED SURROUNDING THE DESCENDING THORACIC AORTA AT THE SITE OF STENT GRAFT PLACEMENT. " "ON ARTERIAL PHASE IMAGING: THE AORTIC ROOT, ASCENDING AORTA, AND AORTIC ARCH APPEAR UNREMARKABLE. THE TEVAR WAS REINFORCED WITH A SECOND AORTIC STENT GRAFT AT THE SITE OF ENDOLEAK. AT PRESENT, THERE IS NO EVIDENCE OF ENDOLEAK. THERE IS NO EVIDENCE OF AORTIC STENOSIS. THE ABDOMINAL AORTA DEMONSTRATES MULTIFOCAL CALCIFIED PLAQUES." A SECOND RADIOLOGY REPORT FROM 10 DAYS POST-INDEX PROCEDURE NOTED: "THE AORTIC ROOT, ASCENDING AORTA, AND AORTIC ARCH ARE UNREMARKABLE. THERE IS A THORACIC ENDOVASCULAR AORTIC REPAIR STENT BEGINNING AT THE LEVEL OF THE AORTIC ARCH AND EXTENDING TO THE PROXIMAL THORACIC AORTA. THERE IS AN ADDITIONAL INNER SECOND ENDOVASCULAR STENT AT THE DISTAL END OF THE FIRST, SHOWING NO EVIDENCE OF ENDOVASCULAR LEAK. THERE IS A THORACIC FUSIFORM AORTIC ANEURYSM MEASURING APPROXIMATELY 4.7 X 6.1 CM, DEMONSTRATING THE AFOREMENTIONED ENDOVASCULAR STENT AND ANEURYSM LUMEN. WITHIN THE ANEURYSM LUMEN, THERE ARE MULTIPLE HYPERDENSE FOCI, WHICH ARE FAVORED TO REPRESENT COAGULATED BLOOD PRODUCTS. THESE FINDINGS ARE GROSSLY UNCHANGED FROM THE PRIOR CT EXAM. NO CONVINCING EVIDENCE OF ENDOLEAK INTO THE ANEURYSM LUMEN. THE THORACIC AORTA SHOWS NO DISSECTION OR NEW ANEURYSM. SEVERE ATHEROMATOUS DISEASE IS NOTED IN THE THORACIC AORTA." ANALYSIS OF THE REPORTED FILMS DID NOT REVEAL ANY OUT-OF-SPECIFICATION STENT GRAFT INTEGRITY ISSUES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: B5; ADDITIONAL INFORMATION RECEIVED ; IT WAS REPORTED THAT THE SUSPECTED RUPTURE WAS IDENTIFIED THROUGH A CT SCAN AND WAS BELIEVED TO BE ASSOCIATED WITH THE INFECTION. AN ENDOLEAK WAS PRESENT AND WAS SUSPECTED TO BE EITHER A TYPE IV POROSITY OR A TYPE III DUE TO THE INFECTION. IT WAS SAID THAT NEITHER COULD BE RULED OUT PRIOR TO THE PATIENT BEING EXPIRED. AN AUTOPSY WAS NOT PERFORMED. FILM EVALUATION SUMMARY: THE ENDOLEAK WAS CONFIRMED IN THE PROVIDED IMAGES, HOWEVER THE CAUSE AND SUBTYPE OF ENDOLEAK COULD NOT BE CONCLUSIVELY DETERMINED. AORTIC RUPTURE AND INFECTION COULD NOT BE CONFIRMED BASED ON THE AVAILABLE IMAGES. TYPE IA AND TYPE IB ENDOLEAKS ARE UNLIKELY, AS ADEQUATE PROXIMAL AND DISTAL SEALS WERE OBSERVED. A TYPE IIIA ENDOLEAK IS ALSO NOT CONSIDERED A FACTOR. WHILE A TYPE IIIB FABRIC ENDOLEAK CANNOT BE RULED OUT, THIS TYPE OF ENDOLEAK IS LESS LIKELY TO OCCUR SO SOON AFTER THE INDEX PROCEDURE. IT IS ALSO POSSIBLE THAT THIS MAY HAVE BEEN A TYPE IV ENDOLEAK CAUSED BY FABRIC POROSITY. THERE WAS NO EVIDENCE OF ENDOLEAK AFTER THE STENT GRAFT WAS RELINED, AND RELINING THE STENT GRAFT COULD HAVE RESOLVED EITHER TYPE OF ENDOLEAK. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22211713&pc=OZD | 1220648-2025-29322 | 2025-05-21 04:00 | Death | ABIOMED, INC. | 2025-06-16 04:00 | OZD | IMPELLA 5.5 | Pumping Stopped | Perforation of Vessels | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT STRUGGLED TO WEAN FROM THE BYPASS, NECESSITATING REVISION OF SOME GRAFTS AND A BENTALL PROCEDURE. DESPITE THESE EFFORTS, THE PATIENT WAS UNABLE TO COME OFF THE PUMP AND REQUIRED FIVE SHOCKS AND WAS PLACED ON CENTRAL VA ECMO FOR SUPPORT. THE DECISION WAS MADE TO PLACE AN IMPELLA 5.5 AS A VENT. THE TEAM ANTICIPATED DIFFICULTIES CROSSING THE AO VALVE DUE TO LACK OF PULSATILITY. AFTER MULTIPLE ATTEMPTS WITH VARIOUS CATHETERS, THE AL1 SUCCESSFULLY CROSSED THE VALVE, AND THE IMPELLA WAS PLACED WITHOUT FURTHER ISSUES. HOWEVER, UPON TURNING THE PUMP ON TO P2, FLOWS DISPLAYED 0.0/0.0, AND THE PUMP REVERTED TO P0. SEVERAL ATTEMPTS WERE MADE; HOWEVER, THE P-LEVEL CONSISTENTLY DEFAULTED BACK TO P0 BEFORE THE AUTOMATED IMPELLA CONTROLLER (AIC) TURNED OFF. THE AIC WAS RESTARTED PROMPTING IMPELLA DEFECTIVE" ALARM AND ALERT TO "REPLACE IMPELLA." A NEW AIC WAS USED, HOWEVER, THAT DID NOT RESOLVE THE ISSUE. AFTER SEVERAL ATTEMPTS THE DECISION WAS MADE TO EXPLANT THE DEFECTIVE PUMP AND IMPLANT A NEW 5.5 IMPELLA. AGAIN, CROSSING THE VALVE PROVED TO BE DIFFICULT. LIGHT CHEST COMPRESSIONS WERE PERFORMED TO HELP OPEN THE AORTIC VALVE; HOWEVER, EFFORTS WERE UNSUCCESSFUL. ECMO FLOWS WERE DECREASED, AND A SLIGHT INCREASE IN EPI WAS ADMINISTERED TO ASSIST WITH CROSSING THE VALVE, HOWEVER THIS ALSO FAILED. THE CHEST DRESSING WAS REMOVED, AND MANUAL HEART MASSAGE WAS PERFORMED TO FACILITATE OPENING OF THE AO VALVE, YET CROSSING REMAINED UNSUCCESSFUL. BLEEDING WAS OBSERVED DURING THE PROCEDURE, AND THE PATIENT RECEIVED 20 UNITS OF FFP, 19 UNITS OF RBC, 2 UNITS OF PLATELETS, AND 3 UNITS OF CRYOPRECIPITATE. THE PATIENT REMAINED NON-PULSATILE WITH COMPROMISED HEMODYNAMICS AND CONTINUED TO REQUIRE ONGOING TRANSFUSIONS OF BLOOD PRODUCTS. ULTIMATELY THE DECISION WAS MADE BY THE FAMILY TO WITHDRAW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: DIRECT AORTIC INSERTION: ¿MAINTAINING ACT AT OR ABOVE 250 SECONDS WILL HELP PREVENT A THROMBUS FROM ENTERING THE CATHETER AND CAUSING A SUDDEN STOP ON STARTUP.¿ SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER: ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PCI & IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22211765&pc=DSQ | 2916596-2025-03658 | 2024-11-14 05:00 | Death | THORATEC CORPORATION | 2025-06-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT KNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AN AUTOPSY WAS NOT PERFORMED. IT WAS UNKNOWN IF THE DEVICE WAS EXPLANTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22211926&pc=DSQ | 2916596-2025-03502 | 2025-05-20 04:00 | Death | THORATEC CORPORATION | 2025-06-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Arrhythmia; Aspiration/Inhalation; Cardiogenic Shock; Fluid Discharge | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED AN EMERGENT EPISODE OF DECOMPENSATION. THE SITE NOTED POSSIBLE ASPIRATION OF A GREEN SUBSTANCE. THE GREEN SUBSTANCE WAS VISUALLY SEEN EXITING THE DRIVELINE EXIT SITE. THE PATIENT WAS TAKEN FOR A COMPUTED TOMOGRAPHY (CT). AFTER THE CT WAS PERFORMED, THE PATIENT FURTHER DECOMPENSATED AND NEEDED CARDIOPULMONARY RESUSCITATION (CPR) WHILE BEING TRANSPORTED TO THE INTENSIVE CARE UNIT (ICU). THE PATIENT WAS SHOCKED DUE TO ARRHYTHMIA AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS PLACED AT BEDSIDE. THE PATIENT WAS THEN TAKEN TO THE OPERATING ROOM (OR) FOR INSERTION OF A RIGHT VENTRICULAR ASSIST DEVICE (RVAD) WITH OXYGENATOR CONFIGURATION, AND A LAPAROSCOPIC GASTROINTESTINAL GENERAL SURGERY EXPLORATION WAS PERFORMED. THE SITE REPORTED A LOW FLOW ALARM. THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) FLOWS ADJUSTED AND WERE MANAGED BY THE OR AND HEART FAILURE (HF) TEAMS TO SUPPORT THE PATIENT WITH OPTIMAL RVAD WITH OXYGENATOR AND LEFT VENTRICULAR ASSIST DEVICE (LVAD) FLOWS. LOG FILES WERE SENT FOR ANALYSIS ON (B)(6) 2025 AND REVEALED THAT THE EVENT LOG FILE DISPLAYED LOW FLOWS WHICH SEEMED TO BE PHYSIOLOGICAL IN NATURE WHERE THE LOW FLOW ESTIMATOR DROPPED BELOW 2.5 LITERS PER MINUTE (LPM). THE MCS EQUIPMENT APPEARED TO HAVE OPERATED AS EXPECTED. AS OF (B)(6) 2025, THE SITE REPORTED THAT THE PATIENT HAD PASSED AWAY DUE TO CARDIOGENIC SHOCK ON (B)(6) 2025. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. THE PUMP WAS NOT EXPLANTED. ADDITIONAL INFORMATION WAS UNKNOWN. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS OF DECOMPENSATION, ARRHYTHMIA, RIGHT HEART FAILURE, AND CARDIOGENIC SHOCK COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. ADDITIONALLY, A SPECIFIC CAUSE FOR THE REPORTED GREEN SUBSTANCE EXITING THE DRIVELINE EXIT SITE COULD NOT BE CONCLUSIVELY ESTABLISHED. FURTHERMORE, REVIEW OF THE SUBMITTED LOG FILES CONFIRMED MULTIPLE LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THESE ALARMS COULD NOT BE CONCLUSIVELY DETERMINED. THE SYSTEM CONTROLLER EVENT LOG FILE CONTAINED EVENTS FROM (B)(6) 2025 THROUGH (B)(6) 2025, PER THE TIMESTAMPS. EVENTS OCCURRING PRIOR TO THIS TIMEFRAME WERE OVERWRITTEN BY NEWER INCOMING EVENTS, PER DESIGN. MULTIPLE, TRANSIENT LOW FLOW HAZARD ALARMS WERE CAPTURED THROUGHOUT THE FILE. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED. THE PUMP OPERATED AT THE SET SPEED FOR THE ENTIRETY OF THE FILE. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATION. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) IFU AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. CHAPTER 1, "INTRODUCTION", LISTS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING CARDIAC ARRHYTHMIA AND RIGHT HEART FAILURE. THIS CHAPTER ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. CHAPTER 4 OF THE IFU DESCRIBES PUMP FLOW AND HAZARD ALARMS. THE IFU STATES THAT PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. THIS SECTION ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. CHAPTER 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, STATES THAT SUDDEN RIGHT VENTRICULAR FAILURE MAY DEVELOP DURING OR SHORTLY AFTER PUMP IMPLANTATION AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS. THIS SECTION ALSO PROVIDES INFORMATION ON CARING FOR THE DRIVELINE EXIT SITE. CHAPTER 7 OF THE IFU, ""ALARMS AND TROUBLESHOOTING"", AND SECTION 5 OF THE PATIENT HANDBOOK, ""ALARMS AND TROUBLESHOOTING"", ADDRESS ALL SYSTEM ALARM CONDITIONS, INCLUDING THE LOW FLOW HAZARD, AS WELL AS THE APPROPRIATE ACTIONS ASSOCIATED WITH EACH CONDITION. FURTHERMORE, SECTION 8 OF THE PATIENT HANDBOOK, ¿HANDLING EMERGENCIES¿, ALSO PROVIDES EXAMPLES OF EMERGENCIES AND THE PROPER ACTIONS TO TAKE IN THE EVENT AN EMERGENCY OCCURS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22211951&pc=MHX | 1218950-2025-000328 | 2025-04-22 04:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-06-16 04:00 | MHX | INTELLIVUE MX40 WLAN | Erratic Results | Insufficient Information | K113125 |  | 1 | Event Description: IT WAS REPORTED THE DEVICE DISPLAYED THE INCORRECT HEART RATE AND THE PATIENT PASSED AWAY. THE MX40 DEVICE MEASURED A HEART RATE OF 75BPM WHEN THE ACTUAL HEART RATE WAS 39BPM. RESUSCITATION WAS ATTEMPTED BUT THE PATIENT PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. Manufacturer Narrative: A PHILIPS REMOTE SERVICE ENGINEER (RSE) WORKED WITH THE CUSTOMER BIOMEDICAL ENGINEER (BIOMED). INFORMATION PROVIDED TO THE RSE INDICATED THAT THE PATIENT WAS NOT PACED BY AN EXTERNAL OR IMPLANTED PACEMAKER AT THE TIME OF THE EVENT. THE USERS WERE AWARE OF THE HEART RATE DISCREPANCY, BUT IT WAS ALSO INDICATED THIS DISCREPANCY DID IMPACT THE PATIENT OUTCOME. THE RSE WAS UNABLE TO DOWNLOAD THE CLINICAL AUDIT LOG, BUT WAS ABLE TO REVIEW REMOTELY. PER THE RSE, THE PIC IX LOGS SHOWED THAT THE MX40 DEVICE HAD ALARMED CORRECTLY DURING THE TIME OF THE ISSUE; THE LOG SPREADSHEET SHOWED THAT HEART RATE (HR) ALARMS WERE TRIGGERED IN THE 30S. THE RSE PROVIDED THE BIOMED WITH THIS INFORMATION. NO INFORMATION WAS PROVIDED, REGARDING THE REASON THE MX40 WOULD SHOW A HR OF 75BPM, WHEN THE HR WAS 39BPM. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED WE WERE UNABLE TO REPLICATE THE REPORTED PROBLEM. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. Manufacturer Narrative: PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22212003&pc=DSQ | 3007042319-2025-01275 | 2025-06-08 04:00 | Death | HEARTWARE, INC. | 2025-06-16 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS INTRACEREBRAL BLEEDING. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION AS IT REMAINS IMPLANTED. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION CONFIRMED THAT THE ASSOCIATED DEVICE MET ALL REQUIREMENTS FOR RELEASE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT'S CAUSE OF DEATH WAS INTRACEREBRAL BLEEDING. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, BLEEDING AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22212183&pc=BZD | 2518422-2025-106655 | 2025-05-30 04:00 | Death | RESPIRONICS, INC. | 2025-06-16 04:00 | BZD | REMSTAR PLUS W/C-FLEX | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM THE PATIENT'S SISTER STATING THAT THE PATIENT PASSED AWAY. SHE WANTED TO KNOW WHAT SHE COULD DO WITH THE REMSTAR PLUS W/ C-FLEX DEVICE. THERE WAS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE CALLER WAS DIRECTED TO A DME OR ELECTRONICS BIN TO DISPOSE OF THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22212287&pc=PNJ | 2017865-2025-87363 | 2025-05-25 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-16 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Positioning Failure | Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR AN IMPLANT PROCEDURE. DURING THE PROCEDURE, WHILE FIXATING THE VENTRICULAR LEADLESS PACEMAKER WITH THE DELIVERY CATHETER IT WAS NOTED THAT THE DEVICE'S TIP WAS UNSTABLE. SOON AFTER THE PATIENT'S BLOOD PRESSURE AND SATURATION DROPPED. AN ECHOCARDIOGRAM SHOWED THAT THERE WAS AN EPICARDIAL EFFUSION AND A CARDIAC PERFORATION WAS SUSPECTED. THE PATIENT WAS STABILIZED; THE DEVICE WAS RETRIEVED AND THE PROCEDURE ABANDONED. POST-PROCEDURE, THE PATIENT PASSED AWAY. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED, AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. REVIEW OF THE STERILIZATION RECORDS CONFIRMED NORMAL STERILIZATION CYCLES FOR THE PRODUCTS. THE PRODUCT WAS RETURNED, AND VISUAL INSPECTION WAS NORMAL. THE CAUSE OF INFECTION COULD NOT BE TRACED TO THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22212292&pc=PNJ | 2017865-2025-87366 | 2025-05-25 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-16 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR AN IMPLANT PROCEDURE. DURING THE PROCEDURE, WHILE FIXATING THE VENTRICULAR LEADLESS PACEMAKER WITH THE DELIVERY CATHETER IT WAS NOTED THAT THE DEVICE'S TIP WAS UNSTABLE. SOON AFTER THE PATIENT'S BLOOD PRESSURE AND SATURATION DROPPED. AN ECHOCARDIOGRAM SHOWED THAT THERE WAS AN EPICARDIAL EFFUSION AND A CARDIAC PERFORATION WAS SUSPECTED. THE PATIENT WAS STABILIZED; THE DEVICE WAS RETRIEVED AND THE PROCEDURE ABANDONED. POST-PROCEDURE, THE PATIENT PASSED AWAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22212448&pc=DWF | 2531527-2025-20002 | 2025-01-01 05:00 | Death | CARDIAC ASSIST INC. | 2025-06-16 04:00 | DWF | PROTEK DUO VENOUS CANNULA | Patient Device Interaction Problem | Hypovolemic Shock | K140999 |  | 1 | Event Description: LIVANOVA RECEIVED A MEDWATCH REPORT (B)(4) REPORT RELATED TO A PROTEK DUO CANNULA 29FR. PATIENT DEVELOPED A PROGRESSIVE DYSPNOEA FOLLOWING COVID INFECTION. A SURFACE ECHOCARDIOGRAM REVEALED PROGRESSION TO SEVERE MITRAL STENOSIS. TRANSOESOPHAGEAL ECHOCARDIOGRAM WAS CONFIRMATORY FOR SEVERE MITRAL STENOSIS WITH MEAN GRADIENT 12 MMHG WITH SEVERE THICKENING OF THE POSTERIOR GREATER THAN ANTERIOR MITRAL VALVE LEAFLETS AND MITRAL ANNULAR CALCIFICATION. CARDIAC CATHETERIZATION CONFIRMED MODERATE AORTIC AND SEVERE MITRAL STENOSIS WITH MEAN GRADIENT 12 AND 35 MMHG, LEFT MAIN DISEASE EXTENDING INTO THE PROXIMAL LAD (LEFT ANTERIOR DESCENDING). PATIENT MET IN CONSULTATION WITH DR. AND HAS ELECTED TO UNDERGO SURGICAL INTERVENTION (REPLACEMENT OF IMPELLA WITH PROTEK DEVICE IN THE RIGHT VENTRICLE OF THE HEART). PATIENT WAS ADMITTED PREOPERATIVELY FOR COMPLETION OF PREOPERATIVE TESTING. PATIENT UNDERWENT THE SURGERY FOR REPLACEMENT OF IMPELLA (NON LIVANOVA DEVICE) WITH PROTEK DUO CANNULA IN THE RIGHT HEART. APPROXIMATELY ONE HOUR AFTER RETURN TO THE CVICU (CARDIOVASCULAR INTENSIVE CARE UNIT) AFTER DEVICE PLACEMENT, THE CHEST TUBE OUTPUT INCREASED DRAMATICALLY, AND PRESSOR REQUIREMENTS ALSO INCREASED. THE PATIENT WAS THEN GIVEN MULTIPLE MASSIVE TRANSFUSION PACKS WHILE THE FAMILY AND SURGICAL TEAM DETERMINED THE BEST PATH FORWARD, AND ULTIMATELY ELECTED TO HAVE THE PATIENT PROCEED WITH COMFORT MEASURES. TRANSFUSIONS WERE STOPPED, AND THE PATIENT PROMPTLY PASSED SECONDARY TO HYPOVOLEMIC SHOCK. PATIENT DEAD.Event Description: SEE INITIAL REPORT. Manufacturer Narrative: D.4. LOT NUMBER IS UNKNOWN. THIS INFORMATION WILL BE PROVIDED IN A SUPPLEMENTAL REPORT IF MADE AVAILABLE. AS THE LOT NUMBER IS UNKNOWN, UNIQUE DEVICE IDENTIFIER (UDI) NUMBER AND EXPIRY DATE ARE NOT AVAILABLE H.4. AS THE LOT NUMBER IS UNKNOWN, THE DEVICE MANUFACTURE DATE COULD NOT BE DETERMINED. THIS INFORMATION WILL BE PROVIDED IN A SUPPLEMENTAL REPORT IF MADE AVAILABLE. H11: LIVANOVA MANUFACTURES THE PROTEK DUO CANNULA. THE INCIDENT OCCURRED IN (B)(6). AS ADDITIONAL INFORMATION REPORTED IN THE RECEIVED MEDTWATCH REPORT: DUE TO PATIENT INCREASED CHEST TUBE OUTPUT (FIRST PRIMARILY FROM LEFT PLEURAL TUBE AT APPROXIMATELY 200CC/HR AND SUBSEQUENTLY FROM THE MEDIASTINAL TUBES AS WELL) A MASSIVE BLOOD TRANSFUSION WITH EQUAL PRBC (PACKED RED BLOOD CELLS), FFP (FRESH FROZEN PLASMA) AND PLATELETS WERE STARTED AND ESCALATION OF PRESSORS. IN ADDITION, TEE (TRANSESOPHAGEAL ECHOCARDIOGRAM) WAS DONE WHICH SHOWED THAT THE NEWLY PLACED PROTEK DUO CANNULA WAS IN THE RV (RIGHT VENTRICLE) WITH WHAT LOOKED LIKE AN RV INJURY PRESENT. AFTER SPEAKING WITH PATIENT'S FAMILY IT WAS AGREED NOT TO PROCEED WITH A FURTHER SURGERY. THEREFORE, ANY COAGULOPATHY WAS CORRECTED AND DESPITE THIS, THE BLEEDING CONTINUED. THEN, IT WAS AGREED TO PROCEED WITH COMFORT CARE MEASURES. THE PATIENT'S FOLLOWING CONDITION NECESSITATING CRITICAL CARE MANAGEMENT: ACIDOSIS, ACUTE BLOOD LOSS ANAEMIA, ACUTE KIDNEY INJURY, ARRHYTHMIA, CARDIOGENIC SHOCK, COAGULOPATHY, ELECTROLYTE ABNORMALITIES, HYPOTENSION, HYPOVOLEMIC SHOCK, LIVER FAILURE, MULTISYSTEM ORGAN FAILURE, RENAL FAILURE, THROMBOCYTOPENIA. THE CRITICAL CARE TREATMENT AND MANAGEMENT INCLUDED ALKALI THERAPY, ANTIBIOTICS, BLOOD PRODUCT ADMINISTRATION, CONTINUOUS RENAL REPLACEMENT THERAPY, DISCUSSION OF PATIENT'S CONDITION AND THERAPEUTIC OPTIONS WITH MEDICAL DECISION MAKERS, ELECTROLYTE AND NUTRITION MANAGEMENT, FLUID ADMINISTRATION, INVASIVE HEMODYNAMIC MONITORING, MANAGEMENT OF MECHANICAL VENTILATION, REVIEW OF MEDICAL RECORDS, SEDATION MANAGEMENT, TITRATION OF INOTROPE INFUSIONS, TITRATION OF VASOPRESSOR INFUSIONS, DISCUSSION WITH CONSULTING SERVICES, TITRATION OF ANTICOAGULANTS. AS PER LIVANOVA MEDICAL EXPERTS' OPINION, THE INJURY (BLEEDING) IS AT THE CANNULA INSERTION SITE. THIS MAY BE RELATED TO THE USED INSERTION TECHNIQUE OR TO CANNULA MIGRATION. HOWEVER, AT THE MOMENT THERE IS NO EVIDENCE, AND NO DEFINITIVE CONCLUSION CAN BE REACHED AROUND THE RELATIONSHIP BETWEEN INJURED TISSUE AND THE DEVICE. MOREOVER, BASED ON INFORMATION PROVIDED, ALSO THE SUCTION OF THE IMPELLA (NON LIVANOVA DEVICE) GENERATING EXCESSIVE NEGATIVE PRESSURE MAY HAVE CAUSED THE INJURY TO THE RIGHT VENTRICLE AS THE PROCEDURE WAS FOR IMPELLA DEVICE REPLACEMENT. REGARDING THE ULTIMATE CAUSE OF PATIENT DEATH, MEDICAL ASSESSMENT OUTCOME IS THAT IT WAS LIKELY CAUSED BY ALL OTHER SERIOUS PATIENT CONDITIONS (E.G. COVID; MYCOBACTERIA INFECTION; CORONARY HEART DISEASE; DIABETES; HEPATIC/RENAL DYSFUNCTION; HYPERTENSION; IMMUNO-COMPROMISED; MORBIDLY OBESE; OTHER) AND THAT MOST LIKELY THE BLEEDING WAS NOT THE MAIN CAUSE OF THE DEATH. INVESTIGATION IS ONGOING. IF ANY ADDITIONAL INFORMATION PERTINENT TO THE REPORTED EVENT IS RECEIVED, IT WILL BE PROVIDED IN A SUPPLEMENTAL REPORT. Manufacturer Narrative: H10: RELATED MEDWATCH REPORT NUMBER HAS BEEN ADDED INTO THE DEDICATED SECTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22212589&pc=OZD | 1220648-2025-29329 | 2025-05-19 04:00 | Death | ABIOMED, INC. | 2025-06-16 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT THE DECISION FOR IMPELLA 5.5 WAS MADE ON A PATIENT WITH CARDIOGENETIC SHOCK. HOWEVER, IT WAS NOTED THAT DURING IMPLANTATION OF THE IMPELLA 5.5, IT WAS NOT POSSIBLE TO ADVANCE THE DEVICE THROUGH THE GRAFT. SEVERAL ATTEMPTS WERE MADE; HOWEVER, THEY WERE ALL UNSUCCESSFUL. INSTEAD, AN IMPELLA CP (B)(6) WAS SUCCESSFULLY IMPLANTED. DURING CP SUPPORT, THERE WAS INTERMITTENT SUCTION WHICH THEY INCREASED THE VOLUME. POSITION WAS VERIFIED BY ECHO. THE PATIENT WAS SCHEDULED FOR A BYPASS AND VENTRICULAR SEPTAL DEFECT CLOSURE SURGERY WITH A 5.5 ESCALATION. IT WAS NOTED THAT THE IMPELLA 5.5 WAS IN DIRECT AORTIC AND EXTRACORPOREAL LIFE SUPPORT CENTRAL IMPLANTED AFTER SUCCESSFUL HEART SURGERY. THE PATIENT WAS WITH OPEN THORAX ON WARD. ALSO NOTED WAS THAT AFTER SURGERY THAT LACTATE WAS DROPPING AND PARAMETERS WERE GETTING BETTER. HOWEVER, THE PATIENT ULTIMATELY EXPIRED ON SUPPORT. THIS CASE IS FOR THE CP (B)(6). THE MEDICAL SAFETY OFFICER REVIEWED AND DETERMINED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP PUMP 2 DEVICE USED AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME GIVEN THERE WAS INTERMITTENT SUCTION WITH SUPPORT. THE PATIENT EXPIRED THE FOLLOWING DAY AFTER ESCALATION IN THERAPY TO IMPELLA 5.5. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW HAS BEEN COMPLETED. THE PUMP WAS NOT RETURNED. DATA LOGS WERE NOT RECOVERED. THE CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY PATIENT CONDITION RELATED DUE TO ISSUE RESOLVING AROUND THE TIME ADDITIONAL VOLUME WAS GIVEN AND THE P-LEVEL WAS REDUCED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE MEDICAL SAFETY OFFICER REVIEWED AND DETERMINED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP PUMP 2 DEVICE USED AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME GIVEN THERE WAS INTERMITTENT SUCTION WITH SUPPORT. THE PATIENT EXPIRED THE FOLLOWING DAY AFTER ESCALATION IN THERAPY TO IMPELLA 5.5. THE IMPELLA 5.5 FOR DELIVERY ISSUE IS REPORTED ON MANUFACTURE REPORT NUMBER 1220648-2025-29319.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22212639&pc=KDJ | 3030306055-2025-00042 | 2025-05-24 04:00 | Death | VANTIVE US HEALTHCARE LLC | 2025-06-16 04:00 | KDJ | NI | Improper or Incorrect Procedure or Method | Peritonitis | NI |  | 1 | Event Description: IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT EXPERIENCED A BREACH IN ASEPTIC TECHNIQUE WHICH RESULTED IN PERITONITIS. THE BREACH IN ASEPTIC TECHNIQUE WAS NOT FURTHER DESCRIBED. THE PERITONITIS WAS MANIFESTED BY CLOUDY PD EFFLUENT. THE SAME DAY AS EVENT ONSET, THE PATIENT WAS HOSPITALIZED AND TREATED WITH VANCOMYCIN INJECTION (1GM, EVERY 96 HOURS, INTRAPERITONEAL, FOR SEVEN DAYS) AND MEROPENEM INJECTION (1GM, ONCE DAILY, INTRAPERITONEAL, FOR SEVEN DAYS) FOR PERITONITIS. EIGHT DAYS AFTER EVENT ONSET, THE PATIENT WAS TREATED WITH INJECTION OF MEROPENEM )1 GM EVERY 96 HOURS, INTRAVENOUS, ONGOING). TEN DAYS AFTER HOSPITAL ADMISSION, THE PATIENT WAS DISCHARGED. AT THE TIME OF THIS REPORT, THE PATIENT WAS NOT RECOVERED FROM THE PERITONITIS. SEVEN DAYS AFTER EVENT ONSET, PD THERAPY WAS STOPPED. THE PD CATHETER WAS REMOVED AND THE PATIENT WAS SHIFTED TO HEMODIALYSIS (ONGOING). IT WAS REPORTED THE PATIENT WAS RETRAINED ON THE PROPER ASEPTIC TECHNIQUE. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: THIS REPORT IS FOR A BREACH IN ASEPTIC TECHNIQUE WHICH RESULTED IN PERITONITIS. PER VANTIVE LABELING, USERS ARE INSTRUCTED TO USE ASEPTIC TECHNIQUE WHEN PERFORMING PERITONEAL DIALYSIS THERAPY. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: UPON FOLLOW UP, IT WAS REPORTED THAT MEROPENEM WAS DISCONTINUED ON AN UNKNOWN DATE. IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY PASSED AWAY IN THE HOSPITAL. THE CAUSE OF DEATH WAS REPORTED AS DUE TO LOW BLOOD PRESSURE. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. IT WAS REPORTED THAT THE PATIENT HAD NOT RECOVERED FROM THE PERITONITIS PRIOR TO DEATH. HD THERAPY WAS ONGOING. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22212851&pc=HIH | 1222780-2025-00331 | 2025-05-22 04:00 | Death | HOLOGIC, INC | 2025-06-16 04:00 | HIH | MYOSURE TISSUE REMOVAL DEVICE | Adverse Event Without Identified Device or Use Problem | Arrhythmia | K142029 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A PATIENT RECEIVED A MYOSURE XL PROCEDURE ON A 40 YEAR-OLD PATIENT, DURING THE PROCEDURE UP UNTIL 17 MINUTES EVERYTHING WAS DOING WELL AND TOTAL AMOUNT OF FLUID WAS 5 LITERS AND 3 LITERS SPILLED OUT OF THE PATIENT. THE HYSTEROSCOPE WAS REMOVED FROM THE PATIENT´S BODY AND THE PROCEDURE WAS ALMOST COMPLETE WHEN THE PATIENT SUDDENLY DEVELOPED ARRHYTHMIA AND AFTER 25 MINUTES OF RESUSCITATION FROM THE TEAM , THE PATIENT WAS TRANSFERRED TO THE ICU BUT DIED THE NEXT DAY. ECHOCARDIOGRAPHY WHICH DID NOT REVEAL ANYTHING. AN ABDOMINAL ULTRASOUND WAS PERFORMED AND THERE WERE NO PROBLEMS OBSERVED. PATIENT HAD A PREVIOUS HISTORY OF A MYOMECTOMY WITH RESECTOSCOPE A MONTH PRIOR TO THE PROCEDURE. THE PHYSICIAN BELIEVED THAT THE OPERATION WAS WELL PERFORMED AND THAT THE PATIENT´S DEATH HAD NOTHING TO DO WITH THE MYOSURE. THE FLUID MANAGEMENT SYSTEM WAS NOT A HOLOGIC PRODUCT. NO OTHER INFORMATION COULD BE OBTAINED. Manufacturer Narrative: HOLOGIC IS SUBMITTING THIS REPORT IN ACCORDANCE WITH 21. CFR PART 803. THE SUBMISSION IS BASED UPON THE CURRENTLY AVAILABLE INFORMATION. THE SUBMISSION OF THIS REPORT DOES NOT REPRESENT A CONFIRMATION OF DEVICE MALFUNCTION INVOLVING THE HOLOGIC PRODUCTS IN THE EVENT DESCRIBED. A DEVICE HISTORY RECORD (DHR) REVIEW WAS CONDUCTED FOR THE REPORTED LOT/SERIAL NUMBER. THE DEVICE WAS RELEASED MEETING ALL QA SPECIFICATIONS. THE DEVICE INVOLVED IN THIS EVENT WAS NOT RETURNED FOR EVALUATION PURPOSES THEREFORE VISUAL AND FUNCTIONAL ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED. WE ARE UNABLE TO CONFIRM A RELATIONSHIP BETWEEN THE DEVICE AND THE ISSUE REPORTED. THE INFORMATION OBTAINED DURING COMPLAINT INVESTIGATION WILL BE INCLUDED IN OUR GLOBAL COMPLAINT TRENDING AND PRODUCT SURVEILLANCE WILL CONTINUE TO MONITOR COMPLAINTS OF THIS TYPE FOR ADVERSE TRENDS. IF THE PRODUCT IS RECEIVED OR ADDITIONAL INFORMATION IS OBTAINED, THE INVESTIGATION WILL BE REOPENED ACCORDINGLY PER STANDARD OPERATING PROCEDURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22212890&pc=DQY | 2124215-2025-38698 | 2025-05-19 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-16 04:00 | DQY | WATCHMAN TRUSTEER? ACCESS SYSTEM | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Pericardial Effusion | K240018 |  | 1 | Event Description: IT WAS REPORTED A PERICARDIAL EFFUSION (PE) WAS NOTED. A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE WAS BEING PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 24MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS SELECTED FOR USE. A VERSACROSS CONNECT LAAC KIT WAS SELECTED FOR THE TRANSEPTAL PUNCTURE. IT WAS A CHALLENGING SEPTUM TO CANNULATE. THE VERSACROSS WIRE WAS SEEN TENTING SEPTUM, HOWEVER ONCE THE WAS SHEATH WAS BROUGHT UP, TENTING WAS LOST. AFTER 18 MINUTES OF ATTEMPTING, THE WIRE WAS ACROSS. THE WIRE WAS LOCATED IN THE LEFT ATRIUM HOVERING OVER AORTIC VALVE. THE PHYSICIAN WAS MADE AWARE, AND SEVERAL EFFUSION SCANS DID NOT NOTE ANY CHANGES. THE PHYSICIAN CONTINUED WITH PROCEDURE. THE PATIENT BLOOD PRESSURE WAS DECLINING. THE SCAN FOR PERICARDIAL EFFUSION (PE) WAS NEGATIVE, SO THE PROCEDURE CONTINUED. WHEN THE NON-BOSTON SCIENTIFIC PIGTAIL CATHETER WAS BEING INSERTED INTO LAA, THERE WAS A CHANGE IN THE ANATOMY OF LAA AND AN EFFUSION WAS NOTED IN TRANSVERSE SINUS. THE PATIENT WAS TAPPED TO TREAT THE EFFUSION. THE PATIENT STABILIZED AND THE WATCHMAN CLOSURE DEVICE WAS IMPLANTED. ADDITIONAL BLOOD WAS DRAINED AND THE PATIENT WAS TRANSFERRED TO INTENSIVE CARE UNIT (ICU) FOR OBSERVATION. THE PATIENT REMAINED STABLE AND WAS DISCHARGED 48 HOURS LATER.Event Description: IT WAS REPORTED A PERICARDIAL EFFUSION (PE) WAS NOTED. A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE WAS BEING PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 24MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS SELECTED FOR USE. A VERSACROSS CONNECT LAAC KIT WAS SELECTED FOR THE TRANSEPTAL PUNCTURE. IT WAS A CHALLENGING SEPTUM TO CANNULATE. THE VERSACROSS WIRE WAS SEEN TENTING SEPTUM, HOWEVER ONCE THE WAS SHEATH WAS BROUGHT UP, TENTING WAS LOST. AFTER 18 MINUTES OF ATTEMPTING, THE WIRE WAS ACROSS. THE WIRE WAS LOCATED IN THE LEFT ATRIUM HOVERING OVER AORTIC VALVE. THE PHYSICIAN WAS MADE AWARE, AND SEVERAL EFFUSION SCANS DID NOT NOTE ANY CHANGES. THE PHYSICIAN CONTINUED WITH PROCEDURE. THE PATIENT BLOOD PRESSURE WAS DECLINING. THE SCAN FOR PERICARDIAL EFFUSION (PE) WAS NEGATIVE, SO THE PROCEDURE CONTINUED. WHEN THE NON-BOSTON SCIENTIFIC PIGTAIL CATHETER WAS BEING INSERTED INTO LAA, THERE WAS A CHANGE IN THE ANATOMY OF LAA AND AN EFFUSION WAS NOTED IN TRANSVERSE SINUS. THE PATIENT WAS TAPPED TO TREAT THE EFFUSION. THE PATIENT STABILIZED AND THE WATCHMAN CLOSURE DEVICE WAS IMPLANTED. ADDITIONAL BLOOD WAS DRAINED AND THE PATIENT WAS TRANSFERRED TO INTENSIVE CARE UNIT (ICU) FOR OBSERVATION. THE PATIENT REMAINED STABLE AND WAS DISCHARGED 48 HOURS LATER. IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY APPROXIMATELY NINE (9) DAYS POST PROCEDURE, (B)(6) 2025. NO FURTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: B5 DESCRIBE EVENT OR PROBLEM: UPDATED WITH ADDITIONAL INFORMATION RECEIVED. H6: IMPACT CODE ADDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22212903&pc=OZD | 1220648-2025-29331 | 2025-05-28 04:00 | Death | ABIOMED, INC. | 2025-06-16 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT PLACED AN IMPELLA 5.5 PUMP TO SUPPORT A PATIENT WITH CARDIAC HISTORY AND IN NEED OF CARE ESCALATION FROM AN INTRA-AORTIC BALLOON PUMP (IABP) AND EXTRACORPOREAL MEMBRANE OXYGENATION. THE IMPELLA 5.5 WAS PLACED AXILLARY AND WAS FOUND TO HAVE A HEMATOMA AT THE IMPELLA INSERTION SITE AND OOZING AT OLD IABP INSERTION SITE TO THE RIGHT GROIN. HEPARIN CURRENTLY WAS ON HOLD AND RED BLOOD CELLS AND PLATELETS INFUSION WERE GIVEN TO THE PATIENT. LATER, THE FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. THE RELATIONSHIP BETWEEN THE IMPELLA 5.5 AND CAUSE OF DEATH IS UNKNOWN. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22213073&pc=NGV | 2124215-2025-33135 | 2025-03-28 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-16 04:00 | NGV | SYSTEM, APPENDAGE CLOSURE, LEFT ATRIAL | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: REPORTED VIA JOURNAL ARTICLE. IT WAS REPORTED THAT PATIENT DEATH OCCURRED. SOLO-CLOSE IS A SIMPLIFIED METHODOLOGY FOR LEFT ATRIAL APPENDAGE OCCLUSION (LAAO) WITH THE WATCHMAN DEVICE CONSISTING OF A SINGLE-OPERATOR, CONSCIOUS SEDATION, TEE-GUIDANCE, INTRAPROCEDURAL DEVICE SIZING, AND SAME-DAY DISCHARGE. A SINGLE-CENTER PROSPECTIVE ANALYSIS OF 208 CONSECUTIVE PATIENTS UNDERGOING SOLO-CLOSE BETWEEN DECEMBER 2020 AND JANUARY 2024 WAS PERFORMED. THE PRIMARY EFFICACY OUTCOME WAS THE RATE OF SIGNIFICANT PERI-DEVICE (GREATER THAN 5MM) LEAK AT 45-DAY TEE. THE PRIMARY SAFETY OUTCOME WAS THE COMPOSITE 3 POINT MAJOR ADVERSE CARDIOVASCULAR EVENTS OUTCOME (3P-MACE, COMPRISED OF CARDIOVASCULAR DEATH, NONFATAL MYOCARDIAL INFARCTION, OR NONFATAL STROKE) AT 45 DAYS. PROCEDURE TIMES, INCIDENTS OF ALL-CAUSE DEATH, DEVICE-RELATED THROMBUS, CEREBROVASCULAR ACCIDENT (CVA), DEVICE EMBOLIZATION, PERICARDIAL EFFUSION, AND MAJOR POST PROCEDURE BLEEDING WITHIN 45 DAYS OF THE PROCEDURE WAS ALSO COLLECTED. RESULTS AT 45 DAYS, THERE WERE NO SIGNIFICANT PERI-DEVICE LEAK (GREATER THAN 5MM). IMMEDIATE PROCEDURAL COMPLICATIONS INCLUDED ONE PARTIAL ESOPHAGEAL TEAR, ONE PULMONARY HEMORRHAGE, ONE TRANSIENT ISCHEMIC ATTACK, ONE ACCESS SITE-RELATED BLEEDING, AND ONE PERICARDIAL EFFUSION. THE INCIDENCES OF ALL-CAUSE MORTALITY, 3P-MACE, AND MAJOR BLEEDING WERE 3, 2, AND 5, RESPECTIVELY. THERE WERE TWO DEVICE RELATED THROMBUS. Manufacturer Narrative: B2 DATE OF DEATH - ARTICLE PUBLISH DATE USED AS DEATH DATE IS UNKNOWN. B3 DATE OF EVENT - ARTICLE PUBLISH DATE USED AS EVENT DATE IS UNKNOWN. D3 MODEL NUMBER, D3 LOT NUMBER, D4 EXPIRATION DATE, D4 UNIQUE IDENTIFIER (UDI) # ARE UNKNOWN AS THE INFORMATION WAS NOT PROVIDED IN THE LITERATURE ARTICLE. D6A IMPLANT DATE WAS REPORTED AS DECEMBER 2020 THROUGH JANUARY 2024. LITERATURE CITATION: GOLZARIAN, H., KNOUS, M., KLEMAN, A., HEMPFLING, G. L., & PATEL, S. M. (2025). CLINICAL OUTCOMES OF THE FIRST 200 CASES OF SINGLE-OPERATOR TEE-GUIDED LEFT ATRIAL APPENDAGE OCCLUSION (SOLO-CLOSE). JOURNAL OF THE AMERICAN COLLEGE OF CARDIOLOGY, 85(12_SUPPLEMENT), 929-929.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22213136&pc=OZD | 1220648-2025-29333 | 2025-05-21 04:00 | Death | ABIOMED, INC. | 2025-06-16 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT HAD ATRIAL FIBRILLATION (AF) AND INTRINSIC RHYTHM COMPETING WITH EPICARDIAL PACING WIRES. EPICARDIAL PACING WAS TURNED OFF TO ALLOW INTRINSIC RHYTHM AND SUPPORT CONTINUED. ADDITIONALLY, THE PATIENT'S ACCESS SITE WAS OOZING. THE SITE WAS REDRESSED, AND BLOOD PRODUCTS WERE ORDERED AND SUPPORT CONTINUED. THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: ACCESS SITE BLEEDING: THE ROOT CAUSE OF THE ACCESS SITE BLEEDING IS DETERMINED TO BE USE ISSUE BECAUSE THE PTT VALUES ARE HIGHER THAN THE THERAPEUTIC RANGE OF ABOUT 400 AS PER THE CLINICAL DETAILS AT THE TIME OF BLEEDING. VF/VT/AF/AT: "CARDIAC ARRYTHMIA CAN BE REPORTED DURING IMPELLA SUPPORT. TO MAKE A DETERMINATION AS TO THE CAUSE OF THE CARDIAC ARRYTHMIA, THE FOLLOWING CLINICAL INFORMATION MUST BE CONSIDERED: PATIENT'S MEDICAL HISTORY, INCLUDING ANY PRE-EXISTING CARDIAC CONDITIONS, PREVIOUS EPISODES OF ARRHYTHMIAS, OR STRUCTURAL HEART DISEASE. TIMING OF THE ARRYTHMIA EVENT IN RELATION TO IMPELLA SUPPORT (DURING OR POST-IMPLANT). ELECTROCARDIOGRAM (EKG) RECORDINGS OF THE ARRHYTHMIA EPISODES. EVALUATION OF ANY UNDERLYING ELECTRICAL DISTURBANCES, SUCH AS QT PROLONGATION OR ELECTROLYTE IMBALANCES. ASSESSMENT OF HEMODYNAMIC INSTABILITY DURING THE ARRHYTHMIA, INCLUDING BLOOD PRESSURE AND CARDIAC OUTPUT MEASUREMENTS. PRESENCE OF OTHER POTENTIALLY CONTRIBUTORY FACTORS, SUCH AS MEDICATION CHANGES, ISCHEMIA, OR ELECTROLYTE ABNORMALITIES. RESPONSE TO ANY ANTIARRHYTHMIC TREATMENTS OR INTERVENTIONS DURING THE EVENT. ANY DOCUMENTED DEVICE-RELATED ISSUES OR COMPLICATIONS DURING IMPELLA SUPPORT. EVALUATION OF POTENTIAL TRIGGERS OR PRECIPITATING FACTORS, SUCH AS CATHETER MANIPULATION, REPERFUSION INJURY, OR INCREASED MYOCARDIAL OXYGEN DEMAND. REVIEW OF ANY CONCOMITANT MEDICATIONS THAT MAY INFLUENCE CARDIAC ELECTROPHYSIOLOGY WITH A RETURNED IMPELLA, THE DEVICE CAN BE INSPECTED FOR ANY BURRS OR ROUGH EDGES THAT MAY CONTACT THE HEART WALL. TO DETERMINE THE TRUE ROOT CAUSE OF THE VT/VF, THE CLINICAL INFORMATION MENTIONED ABOVE WOULD NEED TO BE ASSESSED IN CONJUNCTION WITH EVIDENCE FROM THE RETURNED DEVICE. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED."" Manufacturer Narrative: ACCESS SITE BLEEDING: THE ROOT CAUSE OF THE ACCESS SITE BLEEDING IS DETERMINED TO BE USE ISSUE BECAUSE THE PTT VALUES ARE HIGHER THAN THE THERAPEUTIC RANGE OF ABOUT 400 AS PER THE CLINICAL DETAILS AT THE TIME OF BLEEDING. VF/VT/AF/AT: "CARDIAC ARRYTHMIA CAN BE REPORTED DURING IMPELLA SUPPORT. TO MAKE A DETERMINATION AS TO THE CAUSE OF THE CARDIAC ARRYTHMIA, THE FOLLOWING CLINICAL INFORMATION MUST BE CONSIDERED: PATIENT'S MEDICAL HISTORY, INCLUDING ANY PRE-EXISTING CARDIAC CONDITIONS, PREVIOUS EPISODES OF ARRHYTHMIAS, OR STRUCTURAL HEART DISEASE TIMING OF THE ARRYTHMIA EVENT IN RELATION TO IMPELLA SUPPORT (DURING OR POST-IMPLANT) ELECTROCARDIOGRAM (EKG) RECORDINGS OF THE ARRHYTHMIA EPISODES EVALUATION OF ANY UNDERLYING ELECTRICAL DISTURBANCES, SUCH AS QT PROLONGATION OR ELECTROLYTE IMBALANCES ASSESSMENT OF HEMODYNAMIC INSTABILITY DURING THE ARRHYTHMIA, INCLUDING BLOOD PRESSURE AND CARDIAC OUTPUT MEASUREMENTS PRESENCE OF OTHER POTENTIALLY CONTRIBUTORY FACTORS, SUCH AS MEDICATION CHANGES, ISCHEMIA, OR ELECTROLYTE ABNORMALITIES RESPONSE TO ANY ANTIARRHYTHMIC TREATMENTS OR INTERVENTIONS DURING THE EVENT ANY DOCUMENTED DEVICE-RELATED ISSUES OR COMPLICATIONS DURING IMPELLA SUPPORT EVALUATION OF POTENTIAL TRIGGERS OR PRECIPITATING FACTORS, SUCH AS CATHETER MANIPULATION, REPERFUSION INJURY, OR INCREASED MYOCARDIAL OXYGEN DEMAND REVIEW OF ANY CONCOMITANT MEDICATIONS THAT MAY INFLUENCE CARDIAC ELECTROPHYSIOLOGY WITH A RETURNED IMPELLA, THE DEVICE CAN BE INSPECTED FOR ANY BURRS OR ROUGH EDGES THAT MAY CONTACT THE HEART WALL. TO DETERMINE THE TRUE ROOT CAUSE OF THE VT/VF, THE CLINICAL INFORMATION MENTIONED ABOVE WOULD NEED TO BE ASSESSED IN CONJUNCTION WITH EVIDENCE FROM THE RETURNED DEVICE. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED." Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMART ASSIST SYSTEM INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: "ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA." SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: "REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE." SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): "ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22213139&pc=DSQ | 2916596-2025-03653 | 2025-05-22 04:00 | Death | THORATEC CORPORATION | 2025-06-16 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 DUE TO BACTEREMIA AND SEPSIS. THE VENTRICULAR ASSIST DEVICE (VAD) FUNCTIONED AS INTENDED. NO EQUIPMENT WOULD BE RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22213180&pc=DSQ | 2916596-2025-03550 | 2025-05-18 04:00 | Death | THORATEC CORPORATION | 2025-06-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. WHETHER THE DEATH WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. NO FURTHER INFORMATION WAS AVAILABLE BECAUSE THE PATIENT WAS ALONE AT HOME AND FOUND DEAD BY THE FAMILY. THE FAMILY DID NOT REQUEST FURTHER AUTOPSY AND WAS WAITING FOR HEALTH CARE PROVIDERS TO COMMUNICATE WITH THE FAMILY TO SEE IF THEY WERE WILLING TO PROVIDE MORE INFORMATION. INFORMATION FROM THE FAMILY SAID THE PRODUCT WAS WORKING NORMALLY WHEN THE PATIENT WAS FOUND. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV C, IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22213485&pc=OZD | 1220648-2025-29338 | 2025-05-22 04:00 | Death | ABIOMED, INC. | 2025-06-16 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Stroke/CVA | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT ON IMPELLA 5.5 SUPPORT HAD RUNS OF ATRIAL FIBRILLATION. ALSO, COMPUTED TOMOGRAPHY SCAN RESULT SHOWED A LARGE SUBACUTE RIGHT MIDDLE CEREBRAL ARTERY INFARCT. THE FAMILY ELECTED TO WITHDRAW CARE AND PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22213652&pc=OZD | 1220648-2025-29341 | 2025-05-20 04:00 | Death | ABIOMED, INC. | 2025-06-16 04:00 | OZD | IMPELLA 5.5 | Fluid/Blood Leak | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT PURGE SOLUTION WAS LEAKING FROM THE RED IMPELLA PLUG DURING IMPELLA 5.5 SUPPORT. THE DECISION WAS MADE TO CONTINUE WITH SUPPORT WHILE MONITORING THE LEAK. FOUR DAYS LATER, THE BEDSIDE NURSE REPORTED THAT THE PATIENT HAD A VENTRICULAR TACHYCARDIA EVENT EARLY IN THE MORNING REQUIRING CARDIOVERSION AT BEDSIDE. FOUR MORE DAYS LATER, INTEGRATED PURGE SIDE ARM WAS LEAKING PURGE SOLUTION. HOWEVER, AFTER 16 DAYS OF SUPPORT, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: CLEANING: ¿DO NOT CLEAN WITH OR EXPOSE ANY PART OF THE CLEAR SIDEARM OF THE IMPELLA CATHETER (E.G. INFUSION FILTER, PRESSURE RESERVOIR) TO ALCOHOL. ALCOHOL HAS BEEN SHOWN TO CAUSE CRACKS AND LEAKS IN THESE COMPONENTS. CAREFULLY READ LABELS ON COMMON SKIN PREPS AND LOTIONS TO AVOID USING ANY ALCOHOL-CONTAINING PRODUCTS IN THE AREA OF THE INFUSION FILTER OR PRESSURE RESERVOIR." ¿CLEAN THE CONNECTOR CABLE WITH 70% ISOPROPYL ALCOHOL.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22214144&pc=DTB | 2649622-2025-16286 | 2025-05-22 04:00 | Death | MPRI | 2025-06-16 04:00 | DTB | CAPSUREFIX NOVUS LEAD MRI SURESCAN | Pacing Problem; Under-Sensing | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P930039 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT ATRIAL (RA) LEAD EXHIBITED UNDERSENSING TRIGGERING DEVICE CLASSIFIED FALSE TERMINATION OF ATRIAL TACHYCARDIA/ATRIAL FIBRILLATION (AT/AF) EVENT(S) AT TIMES AND RESULTING IN INAPPROPRIATE ATRIAL PACING. THE RA LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: ADDITIONAL INFORMATION RECEIVED NOTED THAT THE PATIENT IS DECEASED.Event Description: ADDITIONAL INFORMATION RECEIVED NOTED THAT THE PATIENT CAUSE OF DEATH WAS DUE TO THE PATIENT BEING A VICTIM OF A TRAFFIC ACCIDENT IN WHICH A BUS COLLIDED WITH THE VEHICLE THE PATIENT WAS IN RESULTING IN THE DEATH OF THE PATIENT. THE CLINICIAN ATTESTS THAT THE DEVICE SYSTEM HAD NO CONTRIBUTION TO THE PATIENT¿S MANNER OF DEATH. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22214221&pc=NRY | 3007628272-2025-00028 |  | Death | CERENOVUS, INC. | 2025-06-16 04:00 | NRY | CEREGLIDE | Difficult to Advance; Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Thromboembolism; Vascular Dissection |  |  | 1 | Event Description: THE HEALTHCARE PROFESSIONAL REPORTED THAT DURING A MECHANICAL THROMBECTOMY PROCEDURE TO TREAT AN ACUTE ISCHEMIC STROKE TARGETING THE ANTERIOR CEREBRAL ARTERY (ACA), DEVICES WERE USED IN ACCORDANCE WITH THEIR RESPECTIVE INSTRUCTIONS FOR USE (IFUS). THE PROCEDURE WAS PERFORMED VIA THE COMBINED TECHNIQUE USING THE 132CM CEREGLIDE 71 CATHETER (NIC71132C / 31457321), AXS VECTA 46® INTERMEDIATE CATHETER (STRYKER), AND TREVO® STENT RETRIEVER (STRYKER). AFTER THE 8FR OPTIMO® BALLOON GUIDE CATHETER (BGC) (TOKAI MEDICAL) WAS PLACED IN THE INTERNAL CAROTID ARTERY, THE THROMBUS WAS RETRIEVED BY THE TREVO STENT RETRIEVER USING THE CEREGLIDE 71 CATHETER AND THE AXS VECTA 46. THE THROMBUS WAS REMOVED, BUT A VASCULAR DISSECTION OCCURRED. CONTINUOUS FLUSH WAS MAINTAINED DURING THE PROCEDURE. THE PATIENT¿S MOST CURRENT CONDITION WAS NOT KNOWN AT THE TIME OF THE COMPLAINT INITIATION. ON (B)(6) 2025, ADDITIONAL INFORMATION WAS RECEIVED. PER THE INFORMATION, RELATED TO THE FIRST PASS: THE TARGET LESION WAS THE LEFT INTERNAL CAROTID TOP OCCLUSION. AFTER THE 8FR OPTIMO® BGC WAS PLACED IN INTERNAL CAROTID ARTERY, THE THROMBUS WAS RETRIEVED BY A DIRECT ASPIRATION FIRST PASS TECHNIQUE (ADAPT) TECHNIQUE USING THE CEREGLIDE 71. ¿AFTER THE FIRST PASS, THE PHYSICIAN FOUND A CONTRALATERAL RIGHT A2 OCCLUSION, THE DEVICES APPROACHED THE LESION FROM THE SAME SIDE.¿ RELATED TO THE SECOND PASS, THE CEREGLIDE 71 WAS USED AS A DISTAL ACCESS CATHETER, AND IT WAS PLACED AT THE INTERNAL CAROTID (IC)-ANTERIOR CEREBRAL ARTERY (AC) BIFURCATION. THE AXS VECTA 46 WAS PLACED AT THE ORIGIN OF THE ACA. THE THROMBUS WAS THEN RETRIEVED BY THE COMBINED TECHNIQUE USING THE 3MM TREVO STENT RETRIEVER. POST-OPERATIVE CT REVEALED A VASCULAR DISSECTION AT THE ORIGIN OF THE A2. ADDITIONALLY, A HEMORRHAGE OF UNKNOWN CAUSE WAS OBSERVED EXTENSIVELY IN THE POSTERIOR AREA, LEADING TO COMPRESSION OF THE BRAINSTEM AND RESULTING IN RESPIRATORY DISORDER; HOWEVER, ADDITIONAL INTERVENTION COULD NOT BE PERFORMED, AND THE PROCEDURE WAS TERMINATED. THE PATIENT PASSED AWAY. THE INFORMATION INCLUDED THE PHYSICIAN¿S COMMENT: ¿VASCULAR DISSECTION AT THE ORIGIN OF THE ACA MAY HAVE OCCURRED DUE TO FORWARD ADVANCEMENT DURING STENT RETRIEVER TRACTION OR BECAUSE VECTA 46 WAS TOO ADVANCED. SINCE THE CEREGLIDE 71 WAS PLACED PROXIMAL TO THE ACA, I BELIEVED THERE WAS NO CAUSAL RELATIONSHIP WITH THE HEMORRHAGE. GENERALLY, IT IS UNLIKELY THAT HEMORRHAGE IN THE ANTERIOR CIRCULATION CAUSES HEMORRHAGE IN THE POSTERIOR AREA (BRAINSTEM), SO THE CAUSE REMAINS UNKNOWN.¿ ON (B)(6) 2025, ADDITIONAL INFORMATION WAS RECEIVED. PER THE INFORMATION, THE VASCULAR DISSECTION OCCURRED AT THE ORIGIN OF ACA RATHER THAN A2, AND HEMORRHAGE ALSO OCCURRED AT THE SAME LOCATION. HEMORRHAGE WAS ALSO CONFIRMED IN THE POSTERIOR AREA. THE ADDITIONAL INFORMATION ALSO CLARIFIED THE STATEMENT RELATED TO THE FIRST PASS. ON (B)(6) 2025, THE FIRST PASS TARGETED THE LEFT INTERNAL CAROTID TOP OCCLUSION. AFTER THE 8FR OPTIMO® BGC WAS PLACED IN INTERNAL CAROTID ARTERY, THE THROMBUS WAS RETRIEVED BY A DIRECT ASPIRATION FIRST PASS TECHNIQUE (ADAPT) TECHNIQUE USING THE CEREGLIDE 71. ¿AFTER THE FIRST PASS, THE PHYSICIAN FOUND A CONTRALATERAL RIGHT A2 OCCLUSION, THE DEVICES APPROACHED THE LESION FROM THE SAME SIDE. THE CLARIFICATION IS RELATED TO THE ¿SAME SIDE¿ THE SAME SIDE IN THIS CONTEXT IS ¿THE LEFT SIDE.¿ BASED ON THE ADDITIONAL INFORMATION RECEIVED ON 02-JUN-2025 THAT WAS FURTHER CLARIFIED ON 09-JUN-2025, THE EVENT HAS BEEN DEEMED USFDA REPORTABLE UNDER 21 CFR 803 WITH A CLASSIFICATION OF ¿DEATH.¿ Manufacturer Narrative: MANUFACTURER¿S REF. NO: (B)(4). INFORMATION REGARDING PATIENT IDENTIFIER, DATE OF BIRTH, AGE, GENDER, WEIGHT, RACE, AND ETHNICITY WERE NOT PROVIDED. SECTION B.2: DATE OF DEATH: THE DATE OF DEATH IS NOT KNOWN / REPORTED. SECTION B.3: DATE OF EVENT: THE DATE OF THE EVENT WAS NOT REPORTED / KNOWN. SECTION D.2B: PROCODE IS NRY/QJP. SECTION E.1: THE INITIAL REPORTER PHONE IS NOT AVAILABLE / REPORTED. BASED ON COMPLAINT INFORMATION, THE DEVICE IS NOT AVAILABLE TO BE RETURNED FOR ANALYSIS. A REVIEW OF MANUFACTURING DOCUMENTATION ASSOCIATED WITH THIS LOT (31457321) PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESSES RELATED TO THE REPORTED COMPLAINT. THERE WERE NO INTERNAL ACTIONS RELATED TO DEVICE MANUFACTURE OR INSPECTION. THERE WAS NO DEVICE PERFORMANCE ISSUE OR DEVICE MALFUNCTION REPORTED DURING THE PROCEDURE. BASED ON THE MANUFACTURING DOCUMENTATION REVIEW, THERE IS NO INDICATION THAT THE EVENT IS RELATED TO THE DEVICE MANUFACTURING PROCESS. BASED ON THE MANUFACTURING DOCUMENTATION REVIEW, THERE IS NO INDICATION THAT THE EVENT IS RELATED TO THE DEVICE MANUFACTURING PROCESS. THE EXACT CAUSE OF THE EVENT COULD NOT BE CONCLUSIVELY DETERMINED; HOWEVER, IT IS POSSIBLE THAT CIRCUMSTANCES OF THE PROCEDURE AND / OR DEVICE MANIPULATION / INTERACTION MAY HAVE CONTRIBUTED TO THE REPORTED FAILURE. VASCULAR DISSECTION IS A KNOWN COMPLICATION ASSOCIATED WITH THE USE OF THE CEREGLIDE71 CATHETER AND IS MENTIONED IN THE INSTRUCTIONS FOR USE (IFU) AS SUCH. THERE WAS NO ALLEGED QUALITY ISSUE RELATED TO THE USE OF THE DEVICE, AS THE DEVICE PERFORMED AS INTENDED. THERE ARE CLINICAL AND PROCEDURAL FACTORS INCLUDING CLOT BURDEN, VESSEL CHARACTERISTICS, AND MECHANICAL MANIPULATION OF DEVICES THAT MAY HAVE CONTRIBUTED TO THE EVENT RATHER THAN A DEVICE DESIGN OR MANUFACTURING ISSUE. A VASCULAR DISSECTION DURING A THROMBECTOMY IS LIKELY ATTRIBUTED TO THE STENT-RETRIEVER OR GUIDEWIRE. AS EXPLAINED IN THE REFERENCED MEDICAL LITERATURE, STENT-RETRIEVERS (SR) GENERATE A SUSTAINED RADIAL FORCE TO THE VESSEL WALL TO TRAP THE THROMBUS CLOT, WHICH MAY CAUSE INJURY TO THE ENDOTHELIUM. HOWEVER, SINCE THE EVENT WAS AN INTRAPROCEDURAL FINDING AND OCCURRING SHORTLY AFTER THE USE OF THE CEREGLIDE71, THE RELATIONSHIP OF THE DEVICE TO THE REPORTED EVENT CANNOT BE EXCLUDED. A VASCULAR DISSECTION IS CLASSIFIED AS A SERIOUS INJURY, AS IT INVOLVES DAMAGE TO THE VASCULAR STRUCTURE THAT CAN LEAD TO LIFE-THREATENING BLEEDING OR COMPROMISED BLOOD FLOW TO VITAL ORGANS. THE SEVERITY OF THIS EVENT IS CURRENTLY UNKNOWN. BASED ON THIS INFORMATION, THIS EVENT MEETS US FDA REPORTING CRITERIA UNDER 21 CFR 803 WITH A CLASSIFICATION OF ¿SERIOUS INJURY.¿ REFERENCE: LEE IH, HA SK, LIM DJ, CHOI JI. PREDICTORS OF INTRACRANIAL HEMORRHAGE AFTER MECHANICAL THROMBECTOMY USING A STENT-RETRIEVER FOR ANTERIOR CIRCULATION ISCHEMIC STROKE: A RETROSPECTIVE STUDY. MEDICINE (BALTIMORE). 2023 JAN 13;102(2):E32666. DOI: 10.1097/MD.0000000000032666. PMID: 36637951; PMCID: PMC9839270. PER THE ADDITIONAL INFORMATION RECEIVED ON 09-JUN-2025, IT WAS CLARIFIED THAT THE VASCULAR DISSECTION OCCURRED AT THE ORIGIN OF ACA RATHER THAN A2, AND A HEMORRHAGE ALSO OCCURRED AT THE SAME LOCATION. BASED ON THIS INFORMATION AND THE LOCATION OF THE CEREGLIDE71 DEVICE AT THE TIME OF THE EVENTS, THE CORRELATING RELATIONSHIP BETWEEN THE EVENTS OF ¿VASCULAR DISSECTION AND CEREBRAL HEMORRHAGE¿ TO THE USED CEREGLIDE71 DEVICE AS A CONTRIBUTING FACTOR CANNOT BE RULED OUT ENTIRELY. ADDITIONALLY, INFORMATION RECEIVED ON 02-JUN-2025 STATES A CONTRALATERAL RIGHT A2 OCCLUSION WAS FOUND AFTER THE FIRST PASS USING THE CEREGLIDE71 DEVICE; IT IS POSSIBLE THIS WAS A DISTAL EMBOLIZATION OF THE ORIGINAL THROMBUS, FOR WHICH THE CEREGLIDE71 DEVICE AS A CONTRIBUTING OR CAUSAL FACTOR CANNOT BE RULED OUT. WHILE THE CAUSE OF THE HEMORRHAGE IN THE POSTERIOR AREA REMAINS UNKNOWN, THE HEMORRHAGE IN THE ANTERIOR CIRCULATION MAY HAVE CONTRIBUTED TO COMPRESSION OF THE BRAINSTEM, WHICH ULTIMATELY RESULTED IN THE PATIENT¿S DEATH. THEREFORE, THE EVENTS OF ¿DISTAL EMBOLIZATION, VASCULAR DISSECTION, CEREBRAL HEMORRHAGE, AND DEATH¿ MEETS US FDA REPORTING CRITERIA UNDER 21 CFR 803 WITH A CLASSIFICATION OF ¿DEATH.¿ THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. AS PART OF THE POST MARKET SURVEILLANCE PROGRAM, INFORMATION FROM THIS COMPLAINT IS TRENDED FOR STATISTICAL SIGNALS AND CORRECTIVE / PREVENTIVE ACTION MAY BE TRIGGERED LATER. SINCE THERE WAS NO EVIDENCE TO SUGGEST THE EVENT WAS RELATED TO A MANUFACTURING OR DESIGN ISSUE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS BLANK; THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION. THE MANUFACTURER WILL SUBMIT A SUPPLEMENTAL REPORT IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT. ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. Manufacturer Narrative: MANUFACTURER¿S REF. NO: (B)(4). THE PURPOSE OF THIS MDR SUBMISSION IS TO INCLUDE THE ADDITIONAL EVENT INFORMATION RECEIVED ON 17-JUN-2025. [ADDITIONAL INFORMATION]: ON 17-JUN-2025, LIMITED ADDITIONAL INFORMATION WAS RECEIVED. THE ADDITIONAL INFORMATION INDICATED THAT THE HEMORRHAGE IN THE ANTERIOR CIRCULATION WAS CAUSED BY THE ACA DISSECTION, BUT THE CAUSE OF THE HEMORRHAGE IN THE POSTERIOR AREA (BRAINSTEM) WAS UNKNOWN. INFORMATION REGARDING THE HEMORRHAGE IN THE POSTERIOR AREA IS NOT AVAILABLE AT THIS TIME; WHEN INFORMATION IS AVAILABLE, IT WILL BE REPORTED / ADDED TO THE COMPLAINT FILE. UPDATED SECTIONS: B.4, G.3, G.6, H.2, AND H.11. THE MANUFACTURER WILL SUBMIT A SUPPLEMENTAL REPORT IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT. ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT. Manufacturer Narrative: MANUFACTURER¿S REF. NO: (B)(4). THE PURPOSE OF THIS MDR SUBMISSION IS TO INCLUDE THE ADDITIONAL EVENT INFORMATION RECEIVED ON 16-JUL-2025. [ADDITIONAL INFORMATION]: ON 16-JUL-2025, ADDITIONAL INFORMATION WAS RECEIVED. PER THE INFORMATION, THE PATIENT WAS AN 80-YEAR-OLD FEMALE. THE PATIENT EXPIRED THE SAME DAY AS THE PROCEDURE (THE PROCEDURE DATE WAS UNKNOWN / UNOBTAINABLE). THE OFFICIAL CAUSE OF DEATH WAS RESPIRATORY FAILURE DUE TO BRAINSTEM HERNIATION. THE INFORMATION INDICATED THAT THE VASCULAR TORTUOSITY WAS NOT SEVERE, ¿BUT THERE WAS A FEELING OF STUCK AT THE SITES OF TORTUOSITY FROM THE ICA TO THE ACA (WHICH WAS OF NORMAL TORTUOSITY DEGREE), MAKING IT DIFFICULT TO ADVANCE [THE] CEREGLIDE 71. THE CEREGLIDE 71 MADE A TOTAL OF TWO PASSES. THE HEMORRHAGE SUFFERED BY THE PATIENT WAS A SUBARACHNOID HEMORRHAGE. THE CT SCAN WAS DONE AFTER THE PLANNED PROCEDURE WAS COMPLETED. PER THE ADDITIONAL INFORMATION RECEIVED ON 16-JUL-2025, IT WAS DISCLOSED THAT THE HEMORRHAGE WAS CLASSIFIED AS A SUBARACHNOID HEMORRHAGE. ADDITIONALLY, IT WAS REPORTED THAT THERE WAS DIFFICULTY ADVANCING THE CEREGLIDE 71 CATHETER THROUGH SITES OF TORTUOSITY FROM THE ICA TO THE ACA. DIFFICULTY TRACKING A CATHETER THROUGH THE VASCULATURE IS A KNOWN PROCEDURAL OCCURRENCE. THE CONSEQUENCES OF TRACKING DIFFICULTY OCCURRING DURING CLINICAL USE OF THE DEVICE ARE USUALLY ADDRESSED BY MODIFICATION IN TECHNIQUE OR SUBSTITUTION WITH ANOTHER DEVICE. TRACKING DIFFICULTY IS MOST COMMONLY RELATED TO PATIENT ANATOMY, OPERATOR TECHNIQUE, AND APPROPRIATE DEVICE SELECTION. THEREFORE, THE POTENTIAL FOR PATIENT INJURY/DEATH OCCURRING AS A RESULT OF ANY TRACKING DIFFICULTY IS REMOTE. THERE IS NO INDICATION THAT THE TRACKING DIFFICULTY ENCOUNTERED AT TORTUOUS BENDS OF THE VESSEL CONTRIBUTED TO THE ADVERSE EVENTS REPORTED. THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. UPDATED SECTIONS: A.2, B.4, G.3, G.6, H.2, H.6, AND H.11. THE MANUFACTURER WILL SUBMIT A SUPPLEMENTAL REPORT IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT. ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22214260&pc=NIK | 9614453-2025-02312 | 2025-05-03 04:00 | Death | MEDTRONIC EUROPE SARL | 2025-06-16 04:00 | NIK | COBALT¿ XT HF QUAD CRT-D MRI SURESCAN¿ | Appropriate Term/Code Not Available | Cardiomyopathy | P010031 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENT¿S SPOUSE THAT THE PATIENT IS DECEASED AND NOTED THE CAUSE OF DEATH LISTED ON THE DEATH CERTIFICATE WAS NON-ISCHEMIC CARDIOMYOPATHY. THE SPOUSE INQUIRED IF THERE WAS ANYTHING THAT ¿WENT WRONG¿ WITH THE CARDIAC RESYNCHRONIZATION THERAPY DEFIBRILLATOR (CRT-D). ADDITIONAL INFORMATION RELATED TO THE CIRCUMSTANCES OF DEATH WERE REQUESTED AND NOT RECEIVED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22214487&pc=OZD | 1220648-2025-29346 | 2025-05-20 04:00 | Death | ABIOMED, INC. | 2025-06-16 04:00 | OZD | IMPELLA CP | Positioning Failure | Hemolysis; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THE PATIENTS URINE WAS RED. THE PUMP WAS DEEP THIS MORNING AND WAS PULLED BACK UNDER ECHOCARDIOGRAM. FLUID BOLUS WAS GIVEN. ADDITIONALLY, THERE WAS OOZING AT THE IMPELLA SITE AND SUTURES WERE ADDED. IT WAS FURTHER REPORTED THAT THE PATIENT¿S FAMILY DECIDED TO WITHDRAW CARE AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA USED AS AN ASSOCIATED FACTOR. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING, HEMOLYSIS, AND POSITIONING ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. ACCESS SITE BLEEDING - MAJOR: AFTER 1 DAY ON SUPPORT, OOZING WAS NOTED AT IMPELLA SITE REQUIRING THE PHYSICIAN TO PLACE SUTURES. THE PUMP WAS NOT RETURNED. THE CAUSE OF THE ACCESS SITE BLEEDING WAS MOST LIKELY USE RELATED DUE TO SUTURES REQUIRED THAT RESOLVED THE ISSUE. HEMOLYSIS: THE PUMP WAS NOT RETURNED. DATA LOG REVIEW SHOWED ADEQUATE FLOWS AND PLACEMENTS SIGNAL VALUES DURING DAY OF COMPLIANT, HOWEVER, 8 HOURS OF THE LOGS WERE UNACCOUNTED FOR. THE CAUSE OF THE HEMOLYSIS WAS MOST LIKELY IMPROPER POSITIONING OF THE DEVICE DUE TO CLINICAL DETAILS STATING THE PUMP WAS TOO DEEP IN THE LEFT VENTRICLE (LV). POSITIONING ISSUES: DATA LOG REVIEW SHOWED ADEQUATE FLOWS AND PLACEMENTS SIGNAL VALUES DURING DAY OF COMPLIANT, HOWEVER, 8 HOURS OF THE LOGS WERE UNACCOUNTED FOR. THE CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY USE RELATED DUE TO THE PUMP BEING INITIALLY PLACED TOO DEEP IN THE LV. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22214493&pc=OZD | 1220648-2025-29344 | 2025-05-20 04:00 | Death | ABIOMED, INC. | 2025-06-16 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN EXTENSIVE CARDIAC HISTORY/COMORBIDITIES WAS ON IMPELLA RP FLEX AND IMPELLA CP SUPPORT. THE PATIENT HAD BEEN FOUND DOWN AND UNRESPONSIVE IN THE COMMUNITY. CARDIOPULMONARY RESUSCITATION WAS GIVEN AND THE PATIENT WAS INTUBATED FOR PERCUTANEOUS CORONARY INTERVENTION. IT WAS NOTED THAT THE HEPARIN DRIP WAS TO BE STARTED WHEN LEFT FEMORAL OOZING WAS CONTROLLED. ULTIMATELY, THE PATIENT EXPIRED ON SUPPORT. Manufacturer Narrative: ABIOMED'S MEDICAL SAFETY OFFICER REVIEWED AND THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP DEVICE USED AS AN ASSOCIATED FACTOR TO THE PATIENT'S DEMISE OUTCOME GIVEN THE ACCESS SITE BLEEDING EVENT. HEPARIN IN PURGE MAY HAVE IMPACTED/EXACERBATED THE OUTCOME OF THE BLEED FOR THE PATIENT. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22215314&pc=LWS | 2649622-2025-16293 | 2025-05-22 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-06-16 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Defibrillation/Stimulation Problem; Under-Sensing | Ventricular Fibrillation; Insufficient Information | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD DISPLAYED UNDERSENSING AND THE PATIENT WENT INTO VENTRICULAR FIBRILLATION (VF) FOR AN EXTENDED PERIOD BEFORE BEING RESCUED INTERNALLY. THE PATIENT WAS INTUBATED AND PENDING COMFORT CARE. THE LEAD REMAINS IN USE. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY THE SAME DAY. THE CAUSE OF DEATH AND CONTRIBUTION OF THE DEVICE TO THE DEATH WAS UNKNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22215372&pc=NPT | 2015691-2025-04919 | 2025-05-27 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-16 04:00 | NPT | SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Difficult to Open or Close; Device Stenosis | Asystole | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS GERMANY AFFILIATE, APPROXIMATELY 9 MONTHS TAVR PROCEDURE WITH A 26MM SAPIEN 3 ULTRA VALVE, THE VALVE PRESENTED LEAFLET THICKENING AND REDUCED LEAFLET MOTION, WITH NO MOVEMENT OBSERVED IN THE NON-CORONARY CUSP, RIGHT CORONARY CUSP, AND MINIMAL MOVEMENT IN THE LEFT CORONARY CUSP WITH AN EOA OF 0.5 CM2. THE IMPLANTING HOSPITAL DISCHARGED THE PATIENT, INSTRUCTING THE PATIENT TO CONSULT WITH A LOCAL PHYSICIAN FOR ANTICOAGULATION THERAPY. HOWEVER, THE HOSPITAL DID NOT PROVIDE ANTICOAGULANT MEDICATION TO COVER THE PERIOD BETWEEN DISCHARGE AND THE SCHEDULED APPOINTMENT. THE APPOINTMENT WAS SET 4 DAYS LATER, BUT THE PATIENT PASSED AWAY THE DAY BEFORE ATTENDING THE CONSULTATION. THE PATIENT EXPERIENCED SYNCOPE AND WAS HOSPITALIZED. DURING THE NEXT EPISODE OF PHYSICAL EXERTION, THE PATIENT BECAME ASYSTOLE, AND RESUSCITATION EFFORTS WERE UNSUCCESSFUL. Manufacturer Narrative: CORRECTION TO E1. Manufacturer Narrative: THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THE INSTRUCTIONS FOR USE (IFU), AND TRAINING MANUALS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. IMAGERY WAS PROVIDED BY THE SITE AND EVALUATED BY AN EDWARDS CLINICAL SPECIALIST. THE VALVE PRESENTED WITH SEVERE AORTIC VALVE STENOSIS SHOWED IN INDEX ECHOCARDIOGRAM WITH AN AORTIC VALVE MEAN PEAK GRADIENT OF 51.71 MMHG, AND A HIGH CALCIUM SCORE. THE PROCEDURE WENT WELL, AND THE AORTIC VALVE MEAN GRADIENT OF 9.07 MMHG AFTER VALVE IMPLANTATION. APPROXIMATELY 9 MONTHS POST-IMPLANTATION, THE VALVE EXHIBITED SIGNS OF HALT (HYPO-ATTENUATED LEAFLET THICKENING) AND RLM (REDUCED LEAFLET MOTION) ON THE RIGHT CORONARY CUSP, LEFT CORONARY CUSP, AND NON-CORONARY CUSP AND TEE SHOWED AVA 0.472CM2 CONSISTENT WITH SEVERE AORTIC VALVE STENOSIS. THE REPORTED EVENT FOR LEAFLET THICKENED, LEAFLET MOTION RESTRICTED-IN PATIENT, AND STENOSIS WERE CONFIRMED BASED ON THE REVIEW OF THE PROVIDED IMAGERY. DURING THE MANUFACTURING PROCESS, ALL SAPIEN 3 ULTRA VALVES ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED FOR PROPER COAPTATION UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE COMPLAINT EVENT. HYPOATTENUATED LEAFLET THICKENING (HALT) MAY BE CAUSED BY SEVERAL PATIENT-RELATED FACTORS INCLUDING EARLY VALVE DETERIORATION (SVD) (E.G. STENOSIS/CALCIFICATION), NON-STRUCTURAL DYSFUNCTION (E.G. PANNUS), OR THROMBOSIS (FORMATION OF BLOOD CLOTS ON THE VALVE). IN THIS CASE, THE PATIENT WAS INSTRUCTED TO CONSULT THEIR LOCAL PHYSICIAN FOR ANTICOAGULATION THERAPY, HOWEVER, THE HOSPITAL DID NOT PROVIDE ANTICOAGULANT MEDICATION TO COVER THE PERIOD BETWEEN DISCHARGE AND THE SCHEDULED APPOINTMENT. THE APPOINTMENT WAS SET 4 DAYS LATER, BUT THE PATIENT PASSED AWAY ON THAT DAY BEFORE ATTENDING THE CONSULTATION. WHICH SUSPECTED THAT THE PATIENT WAS POTENTIALLY TO HAVE THROMBOSIS RESULTING IN THICKENED LEAFLET. VALVE THROMBOSIS IS THE FORMATION OF SIGNIFICANT BLOOD CLOTS FORMING ON THE VALVE. THESE CLOTS COULD SIGNIFICANTLY IMPACT THE FUNCTIONALITY OF THE VALVE RESULTING IN LEAFLET THICKENING. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (THROMBOSIS) MAY HAVE CONTRIBUTED TO THE COMPLAINT EVENT. PER THE INSTRUCTION FOR USE (IFU), VALVE STENOSIS IS A POTENTIAL RISK ASSOCIATED WITH BIOPROSTHETIC HEART VALVES. PER THE VALVE ACADEMIC RESEARCH CONSORTIUM (VARC), VALVE STENOSIS CAN RESULT FROM SEVERAL FACTORS, INCLUDING PANNUS, CALCIFICATION, SUPPORT STRUCTURE DEFORMATION (OUT-OF-ROUND CONFIGURATION), TRAUMA (CARDIA-PULMONARY RESUSCITATION, BLUNT CHEST TRAUMA), ENDOCARDITIS, PROSTHETIC VALVE THROMBOSIS, AND NATIVE LEAFLET PROLAPSE IMPEDING PROSTHETIC LEAFLET MOTION. STENOSIS OF AN IMPLANTED VALVE MAY BE A MANIFESTATION OF STRUCTURAL VALVE DETERIORATION (SVD). THIS TERM REFERS TO CHANGES INTRINSIC TO THE VALVE, AND CAN INCLUDE FAILURE MODES SUCH AS WEAR, CALCIFICATION, LEAFLET TEAR, STENT CREEP, LEAFLET DISRUPTION, OR LEAFLET RETRACTION. SVD MAY BE MILD AND NOT REQUIRE ANY INTERVENTION, OR IT MAY BE MODERATE TO SEVERE. IT CAN CAUSE THE HEART TO WORK HARDER TO EJECT BLOOD FROM THE VENTRICLE. DEPENDING ON SEVERITY IT COULD BE AN INDICATION FOR VALVE REPLACEMENT OR MEDICAL INTERVENTION. IN THIS CASE, THE PATIENT HAD LEAFLET THICKENING, WHICH COULD REDUCE THE EOA (EFFECTIVE ORIFICE AREA), RESULTING IN STENOSIS. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (THICKENED LEAFLETS) MAY HAVE CONTRIBUTED TO THE COMPLAINT EVENT. SINCE THE PATIENT WAS EXPERIENCING THE THICKENED LEAFLET, WHICH COULD AFFECT THE COAPTATION VALVE LEAFLETS, RESULTING IN LEAFLET MOTION RESTRICTED. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (THICKENED LEAFLET) MAY HAVE CONTRIBUTED TO THE COMPLAINT EVENT. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: INVESTIGATION IS STILL ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22215385&pc=DSQ | 2916596-2025-03574 | 2025-03-26 04:00 | Death | THORATEC CORPORATION | 2025-06-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Sepsis; Fungal Infection | P160054 |  | 1 | Event Description: EVENT SUMMARY: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO CANDIDA SEPSIS WITHOUT ANY DETAILS OF THE ROOT CAUSE. THE PATIENT SUFFERED FROM A SEVERE REFRACTORY CANDIDA SEPSIS WHICH WAS DIAGNOSED BY BLOOD CULTURE AND RELATED TO THE OUTCOME. NO DIRECT ROOT CAUSE COULD BE DETERMINED. DRIVELINE OR THE DEVICE AS A ROOT CAUSE WERE EXCLUDED. THE OUTCOME WAS CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22215496&pc=OZD | 1220648-2025-29355 | 2025-05-21 04:00 | Death | ABIOMED, INC. | 2025-06-16 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT. THREE DAYS INTO SUPPORT, THE PATIENT ENDURED SEVERAL EPISODES OF VENTRICULAR TACHYCARDIA STORM. THE PATIENT WAS SHOCKED 23 TIMES AND RECEIVED 20 SECONDS OF CHEST COMPRESSIONS TO COUNTERACT THE CONDITION. FIVE DAYS LATER, THE PATIENT CONTINUED TO DECOMPENSATE, AND CARE WAS WITHDRAWN. THE PATIENT EXPIRED UPON REMOVAL OF SUPPORT. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22215643&pc=OZD | 1220648-2025-29356 | 2025-05-21 04:00 | Death | ABIOMED, INC. | 2025-06-16 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Tachycardia; Pericardial Effusion | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PHYSICIAN WAS ABLE TO INDUCE VENTRICULAR TACHYCARDIA (VT) WITH SUCCESSFUL MAPPING. DURING THE END OF THE ABLATION, A PERICARDIAL EFFUSION WAS NOTED. THE PHYSICIAN PLACED A DRAIN, AND THE REMAINDER OF THE PROCEDURE WAS ABORTED DUE TO THE PATIENTS DECLINE. THE DECISION WAS MADE TO LEAVE THE IMPELLA CP IN AND SEND TO THE UNIT TO BE PLACED ON MULTIPLE ANTI ARRHYTHMIC AND CONTINUE IMPELLA CP SUPPORT. THE PLAN WAS TO STABILIZE THE PATIENT AND CORRECT VENTRICULAR TACHYCARDIA, LABS, AND HYPOTENSION DURING VENTRICULAR TACHYCARDIA EPISODES. IT WAS NOTED THAT DURING IMPELLA CP SUPPORT, THE PATIENT'S NEUROLOGY STATUS HAD IMPROVED. HOWEVER, VASOPRESSOR REQUIREMENTS WERE STILL VERY HIGH. THE DECISION WAS MADE TO PLACE THE PATIENT ON COMFORT MEASURES ONLY AND WITHDRAW CARE . THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22215688&pc=QBJ | 2954323-2025-22165 | 2025-05-30 04:00 | Death | ABBOTT DIABETES CARE LTD | 2025-06-16 04:00 | QBJ | FREESTYLE LIBRE 3 | Low Readings | Cardiac Arrest; Hyperglycemia; Loss of consciousness | K223435 |  | 1 | Event Description: ABBOTT DIABETES CARE RECEIVED A BFARM USER REPORT WHICH REPORTED THE FOLLOWING INFORMATION: A HEALTHCARE PROFESSIONAL (HCP) REPORTED ON BEHALF OF THE CUSTOMER THAT A LOW GLUCOSE READING ISSUE WAS OBSERVED WITH THE USE OF THE ABBOTT DIABETES CARE (ADC) DEVICE. THE CUSTOMER RECEIVED AN UNSPECIFIED LOW GLUCOSE READING AND IT IS UNKNOWN WHETHER THE CUSTOMER NOTED A TREND ARROW ON THE DEVICE. AS A RESULT, THE CUSTOMER PROVIDED SELF-TREATMENT WITH "GLUCOSE UPTAKE AND THE CESSATION OF INSULIN ADMINISTRATION." THE CUSTOMER EXPERIENCED SYMPTOMS OF HYPERGLYCEMIA, "FATAL METABOLIC DERANGEMENT," AND WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU) WHERE AN HCP OBTAINED AN UNSPECIFIED LABORATORY RESULT "CONFIRMING HYPERGLYCEMIA" AND PROVIDED UNSPECIFIED TREATMENT. THERE WERE NO ADDITIONAL DETAILS PROVIDED AS HCP COULD NOT BE REACHED TO TROUBLESHOOT FURTHER. THERE WAS NO REPORT OF DEATH OR PERMANENT INJURY ASSOCIATED WITH THIS EVENT.Event Description: ABBOTT DIABETES CARE RECEIVED A BFARM USER REPORT WHICH REPORTED THE FOLLOWING INFORMATION: A HEALTHCARE PROFESSIONAL (HCP) REPORTED ON BEHALF OF THE CUSTOMER THAT A LOW GLUCOSE READING ISSUE WAS OBSERVED WITH THE USE OF THE ABBOTT DIABETES CARE (ADC) DEVICE. THE CUSTOMER RECEIVED AN UNSPECIFIED LOW GLUCOSE READING AND IT IS UNKNOWN WHETHER THE CUSTOMER NOTED A TREND ARROW ON THE DEVICE. AS A RESULT, THE CUSTOMER PROVIDED SELF-TREATMENT WITH "GLUCOSE UPTAKE AND THE CESSATION OF INSULIN ADMINISTRATION." THE CUSTOMER EXPERIENCED SYMPTOMS OF HYPERGLYCEMIA, "FATAL METABOLIC DERANGEMENT," AND WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU) WHERE AN HCP OBTAINED AN UNSPECIFIED LABORATORY RESULT "CONFIRMING HYPERGLYCEMIA" AND PROVIDED UNSPECIFIED TREATMENT. THERE WERE NO ADDITIONAL DETAILS PROVIDED AS HCP COULD NOT BE REACHED TO TROUBLESHOOT FURTHER. THERE WAS NO REPORT OF DEATH OR PERMANENT INJURY ASSOCIATED WITH THIS EVENT.Event Description: ON 06-AUG-2025, ABBOTT DIABETES CARE (ADC) RECEIVED ADDITIONAL INFORMATION REGARDING THIS EVENT FROM THE CRIMINAL INVESTIGATION DEPARTMENT. THE CAREGIVER REPORTED THAT THE PATIENT HAD BEEN EXPERIENCING DIARRHEA FOR THE PAST FEW DAYS. THEY VISITED THE EMERGENCY ROOM THE DAY BEFORE AND WERE DISCHARGED TO GO HOME. ABOUT 30 MINUTES BEFORE THE ALARM WENT OFF, THE PATIENT EXPRESSED A DESIRE TO GO TO BED DUE TO FEELING UNWELL AGAIN. ADDITIONALLY, THEIR MOBILE PHONE WAS CONSTANTLY ALARMING WITH A BLOOD GLUCOSE (BG) READING OF "LO," WHICH LED THE PATIENT TO CONSUME MORE FOOD. MEASUREMENT LOGS EXTRACTED FROM THE APP INDICATE THAT, SINCE (B)(6) 2025, THE PATIENT'S BLOOD GLUCOSE LEVELS HAD PRIMARILY REMAINED IN THE NORMAL RANGE OR SLIGHTLY BELOW NORMAL. HOWEVER, ON THE MORNING OF (B)(6) 2025, THE PATIENT COLLAPSED. BLOOD GLUCOSE MEASUREMENTS TAKEN BY EMERGENCY SERVICES AND AN EMERGENCY DOCTOR AT 9:04 AND 9:15 A.M. BOTH REGISTERED AS "HIGH," INDICATING LEVELS WERE OUTSIDE THE MEASURABLE RANGE. EMERGENCY SERVICES HAD TO INTERVENE, AND THE PATIENT WAS IN A STATE OF RESUSCITATION UPON ARRIVAL AT THE HOSPITAL. DESPITE THIS, THE APP INCORRECTLY ALERTED THEM TO A HYPOGLYCEMIC EVENT AT THAT TIME. THE FIRST MEASURABLE BLOOD GLUCOSE LEVEL RECORDED DURING INTENSIVE CARE AT 12:14 P.M. WAS 1860 MG/DL. ACCORDING TO THE BLOOD GLUCOSE VALUES NOTED IN THE APP, THERE WERE NO INDICATIONS OF HYPERGLYCEMIA AT ANY TIME. THE TREATING PHYSICIANS DESCRIBED AN ALMOST PERMANENT HYPOGLYCEMIA ALARM OF THE APP WITH DEMONSTRABLY EXTREMELY HIGH HYPERGLYCEMIA. THE FOLLOWING MEDICATIONS WERE ADMINISTERED: A TOTAL OF 14 UNITS OF HUMAN INSULIN AND SODIUM BICARBONATE, ALONG WITH IV FLUIDS. ADDITIONALLY, EPINEPHRINE (ADRENALINE/SUPRARENIN) IV 2 MG, ASPIRIN IV 250 MG, HEPARIN IV 5000 IU, NABIC IV 4 G, MIDAZOLAM IV 20 MG, AND PREDNISOLONE IV 100 MG, AS WELL AS STEROFUNDIN IV 500 ML. THE CRIMINAL INVESTIGATION SUGGESTS THAT INACCURATELY LOW BLOOD GLUCOSE READINGS HAD PRESUMABLY BEEN COMMUNICATED VIA THE APP SINCE AT LEAST MONDAY, (B)(6) 2025. ON THE MORNING OF (B)(6) 2025, THE PATIENT¿S HYPERGLYCEMIC CONDITION ULTIMATELY LED TO CIRCULATORY ARREST, RESULTING IN THE PATIENT'S DEATH ON (B)(6) 2025. ADC HAS NOT YET RECEIVED THE AUTOPSY REPORT OR DEATH CERTIFICATE; THEREFORE, THE OFFICIAL CAUSE OF DEATH IS UNKNOWN AT THIS TIME. BASED ON THE CURRENTLY AVAILABLE INFORMATION, IT IS INCONCLUSIVE WHETHER THE ADC DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED DEATH. ADC IS IN THE PROCESS OF CONTACTING THE INVESTIGATING AUTHORITY, THE CRIMINAL INVESTIGATION DEPARTMENT 1, TO CONFIRM WHETHER THEY HAVE THE SENSOR IN THEIR POSSESSION. IF THEY DO, ADC WILL REQUEST THE PRODUCT BE RETURNED FOR FURTHER ANALYSIS. A FOLLOW-UP REPORT WILL BE PROVIDED IF ADC RECEIVES ANY ADDITIONAL INFORMATION REGARDING THIS INCIDENT. Manufacturer Narrative: ABBOTT DIABETES CARE (ADC) RECEIVED ADDITIONAL INFORMATION ON 06-AUG-2025. SECTION UPDATED AS FOLLOWS: B2 - OUTCOMES ATTRIBUTED TO AE. B2 - DATE OF DEATH . B5 - DESCRIBE EVENT OR PROBLEM IS UPDATED TO CAPTURE THIS INFORMATION. H6 - HEALTH EFFECT - CLINICAL CODE AND HEALTH EFFECT - IMPACT CODE. H6 - INVESTIGATION FINDINGS. H6 - INVESTIGATION CONCLUSIONS. THE DEVICE IN QUESTION WAS REQUESTED FOR RETURN. A FOLLOW-UP REPORT WILL BE SUBMITTED UPON COMPLETION OF INVESTIGATION ACTIVITIES OR IF ADDITIONAL INFORMATION IS OBTAINED. AN EXTENDED MANUFACTURING REVIEW WILL BE PERFORMED, AND A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. IF PRODUCT IS RETURNED, A PHYSICAL INVESTIGATION WILL BE PERFORMED PER ADC'S ESTABLISHED PROCESSES AND PROCEDURES AND A REPORT WILL BE SUBMITTED. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED. Manufacturer Narrative: REPEATED ATTEMPTS BY ADC TO RETRIEVE THE PRODUCT WERE UNSUCCESSFUL AND/OR THE CUSTOMER DISCARDED THE PRODUCT. THE MOST PROBABLE ROOT CAUSES ASSOCIATED WITH THIS FAILURE MODE ARE DISCONNECTED, FAULTY OR DAMAGED COMPONENTS, SOFTWARE/DATA CORRUPTION, OR MISUSE. HOWEVER, MITIGATIONS ARE IN PLACE TO REDUCE AND PREVENT SUCH ISSUES. ALL VITAL MANUFACTURING STEPS ARE VALIDATED, MONITORED, AND VERIFIED DURING MANUFACTURING TO ENSURE THE SYSTEM IS IN CONFORMANCE WITH THE VERIFIED DESIGN. ALL VITAL FUNCTIONS ARE MONITORED BY THE SYSTEM AND, WHEN NECESSARY, FUNCTION IS SUSPENDED TO SAFEGUARD AGAINST INACCURATE RESULTS. LABELING IS PROVIDED TO INSTRUCT THE USER ON THE INTENDED USE OF ALL VITAL PARTS OF THE SYSTEM TO MINIMIZE MISUSE. ALL COMPLAINTS AND COMPLAINT TRENDS ARE INVESTIGATED TO DETERMINE IF THERE IS A PRODUCT DEFECT/ DEFICIENCY. IF A PRODUCT DEFECT/ DEFICIENCY IS IDENTIFIED, A RISK EVALUATION IS COMPLETED AND COMPARED TO THE RISK MANAGEMENT REPORT, TO ENSURE THE RISK PROFILE HAS NOT CHANGED. ADDITIONALLY, AS A PART OF ABBOTT¿S POST-MARKET SURVEILLANCE PROCESS, ALL RISK EVALUATIONS WITH ASSOCIATED COMPLAINT DATA ARE REVIEWED ANNUALLY TO DETERMINE IF THE RISK PROFILES HAVE CHANGED AS COMPARED TO THE PRODUCT RISK MANAGEMENT REPORTS. THESE MONITORING PROCESSES ENSURE THAT ALL PRODUCT RISK PROFILES REMAIN ACCEPTABLE AND HAVE A POSITIVE BENEFIT/ RISK RATIO. AT THIS TIME, THE PRODUCT HAS NOT YET BEEN RETURNED FOR THIS COMPLAINT. AN EXTENDED INVESTIGATION HAS BEEN PERFORMED FOR THE REPORTED COMPLAINT AND THERE WAS NO INDICATION THAT THE PRODUCT DID NOT MEET SPECIFICATION. DHRS (DEVICE HISTORY REVIEW) FOR THE LIBRE SENSOR AND SENSOR KIT WAS REVIEWED AND THE DHRS SHOWED THE LIBRE SENSOR AND SENSOR KIT MEET SPECIFICATIONS PRIOR TO RELEASE TO DISTRIBUTION. IF PRODUCT IS RETURNED, A PHYSICAL INVESTIGATION WILL BE PERFORMED PER ADC'S ESTABLISHED PROCESSES AND PROCEDURES AND A REPORT WILL BE SUBMITTED UPON COMPLETION OF INVESTIGATION. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED. Manufacturer Narrative: THE PRODUCT HAS BEEN REQUESTED BACK FOR INVESTIGATION AND THE CUSTOMER AGREED TO RETURN THE DEVICE. A FOLLOW-UP REPORT WILL BE SUBMITTED UPON COMPLETION OF INVESTIGATION ACTIVITIES OR IF ADDITIONAL INFORMATION IS OBTAINED. THE DEVICE MANUFACTURING DATE IS UNKNOWN. THE DATE ENTERED IN SECTION H4 IS THE DATE ABBOTT DIABETES CARE BECAME AWARE OF THE EVENT. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22215926&pc=NVZ | 2017865-2025-87416 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-16 04:00 | NVZ | ELLIPSE VR | Failure to Deliver Shock/Stimulation; Under-Sensing | Arrhythmia | P910023 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY (B)(6) 2025 ALLEGEDLY DUE TO FAILURE OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) TO DETECT ARRHYTHMIA AND DELIVER SHOCK, WHICH REPORTEDLY COULD HAVE PREVENTED CARDIAC ARREST. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. Manufacturer Narrative: NOTE: THE HEALTHCARE PROVIDER/FACILITY IS LISTED AS - (B)(6). THE DEVICE HAS NOT YET BEEN RETURNED. FURTHER INFORMATION WAS REQUESTED BUT NO ADDITIONAL INFORMATION WAS AVAILABLE AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22215934&pc=LWS | 2017865-2025-87427 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-16 04:00 | LWS | DURATA STS OPTIM ACTIVE FIXATION | Failure to Deliver Shock/Stimulation; Under-Sensing | Arrhythmia | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY (B)(6) 2025 ALLEGEDLY DUE TO FAILURE OF THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) TO DETECT ARRHYTHMIA AND DELIVER SHOCK, WHICH REPORTEDLY COULD HAVE PREVENTED CARDIAC ARREST. THE ASSOCIATED RIGHT VENTRICULAR LEAD IS INCLUDED IN THIS REPORT. Manufacturer Narrative: A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTION STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES. Manufacturer Narrative: NOTE: THE HEALTHCARE PROVIDER/FACILITY IS LISTED AS - (B)(6). THE PRODUCT HAS NOT YET BEEN RETURNED. FURTHER INFORMATION WAS REQUESTED BUT NO ADDITIONAL INFORMATION WAS AVAILABLE AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22215953&pc=QFG | 3004464228-2025-26565 | 2025-05-20 04:00 | Death | INSULET CORPORATION | 2025-06-16 04:00 | QFG | OMNIPOD 5 CONTROLLER | Insufficient Information | Insufficient Information | K203768 |  | 1 | Event Description: THE PATIENT'S PHYSICIAN REPORTED THE PATIENT PASSED AWAY WHILE WEARING A POD. THE CAUSE OF DEATH WAS REPORTED TO BE UNKNOWN AND UNDER INVESTIGATION. ON A SUBSEQUENT COMMUNICATION WITH THE PATIENT'S MOTHER, IT WAS REPORTED THE PATIENT'S DEATH WAS GLUCOSE RELATED, AND SHE ALLEGED IT WAS DUE TO A DEVICE MALFUNCTION. NO ADDITIONAL INFORMATION WAS PROVIDED. IF FURTHER INFORMATION IS RECEIVED AT A LATER DATE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE PHYSICAL DEVICE WAS NOT RETURNED FOR INVESTIGATION. CLOUD DATA FROM THE USER FROM (B)(6) 2025 WAS REVIEWED. THE LAST PATIENT HISTORY BUFFER (PHB) DATA POINT WAS CREATED AT 06:38:41 ON (B)(6) 2025. THE DATA SHOWED THAT THE SYSTEM WAS USED PRIMARILY IN AUTOMATED MODE AND WHILE IN AUTOMATED MODE THE AUTOMATED ALGORITHM WAS ABLE TO APPROPRIATELY ADAPT THE MICROBOLUS AMOUNTS BASED ON THE ESTIMATED GLUCOSE VALUES RECEIVED BY THE POD FROM THE SENSOR AND THE USER INPUTS. THE LAST ESTIMATED GLUCOSE VALUE RECEIVED BY THE POD WAS 106 MG/DL AT 01:08:36 ON (B)(6) 2025. ABERRANT VALUES WERE GENERATED UNTIL 02:53:36 ON (B)(6) 2025 WHEN THE SENSOR BECAME INACTIVE. THE SYSTEM RESPONDED APPROPRIATELY, PUTTING THE SYSTEM INTO AUTOMATED MODE LIMITED AFTER FOUR MISSING CYCLES UNTIL THE SYSTEM MODE WAS SWITCHED TO MANUAL MODE AT 02:08:36 ON (B)(6) 2025. THE SYSTEM WAS USED IN MANUAL MODE WITH THE POD RECEIVING NO BLOOD GLUCOSE VALUES THROUGH THE LAST DATA POINT GENERATED AT 06:38:41 ON (B)(6) 2025. THE PHB DATA SHOWED THAT THE SYSTEM WAS CORRECTLY DELIVERING THE USER'S MANUAL BASAL PROGRAM WHILE THE SYSTEM WAS IN MANUAL MODE. NO EVIDENCE OF ISSUES WITH THE SYSTEM WERE OBSERVED IN THE AVAILABLE DATA. LOCKED DOWN SMARTPHONE: LOCKDOWN. OMNIPOD SOFTWARE APP VERSION: 1.2.4. SMARTPHONE OPERATING SYSTEM: N5004L-AM-Q-MV01602-06-01.06. SMARTPHONE HARDWARE: N5004L. CGM SENSOR TYPE: UNSPECIFIED. PLEASE NOTE, THAT SECTION D IS CAPTURING THE DEVICE IDENTIFIERS AS REPORTED BY THE COMPLAINANT. THIS MAY NOT ALIGN TO THE DEVICE CONFIGURATION REPORTED IN H11, AS THIS DATA IS PULLED FROM OUR CLOUD BASED ON THE REPORTED DATE OF EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22216106&pc=QMG | 3015053838-2025-00074 |  | Death | SHOCKWAVE MEDICAL, INC. | 2025-06-16 04:00 | QMG | SHOCKWAVE C2 CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | Material Split, Cut or Torn | Cardiac Arrest; Aspiration Pneumonitis | P200039 |  | 1 | Event Description: TWO SHOCKWAVE C2 CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETERS WERE USED TO TREAT A LESION IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD). THIS REPORT IS FOR THE FIRST IVL CATHETER. REFER TO MDR#3015053838-2025-00075 FOR THE SECOND IVL CATHETER. THE PATIENT WAS HOSPITALIZED FOR HEMODIALYSIS (HD). ECG ABNORMALITY WAS OBSERVED, AND THE PATIENT WAS THEN TRANSFERRED TO THE CARDIOVASCULAR INTERNAL MEDICAL DIVISION FOR CORONARY ANGIOGRAPHY (CAG) AND PERCUTANEOUS CORONARY INTERVENTION (PCI). IVL WAS PERFORMED AND AFTER 2 IVL PULSES WERE DELIVERED, THE BALLOON RUPTURED. IMMEDIATELY AFTER, CARDIAC ARREST OCCURRED. INTUBATION WAS PERFORMED AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS INTRODUCED TO STABILIZE THE GENERAL CONDITION OF THE PATIENT. THE PROCEDURE WAS COMPLETED USING ANOTHER IVL CATHETER. THE PATIENT WAS THEN TRANSFERRED TO THE URGENT CARE CENTER (UCC). THE NEXT DAY, THE PATIENT WAS CONFIRMED TO BE STABILIZED AND EXTUBATED. THE ECMO WAS THEN DISCONTINUED. AFTER 3 WEEKS, THE PATIENT WAS THEN TRANSFERRED TO THE NEPH-INTERNAL MEDICINE DIVISION. 1 WEEK AFTER THE TRANSFER, THE PATIENT WAS STABLE AND DISCHARGED FROM THE HOSPITAL. THE PATIENT DEVELOPED ASPIRATION PNEUMONITIS AND WAS THEN TRANSFERRED TO A NURSERY HOME WHERE THE PATIENT PASSED. THE PHYSICIAN STATES THAT THE PATIENT'S STATUS WAS EXTREMELY BAD. ALTHOUGH HEMODIALYSIS (HD) WAS INTRODUCED, IT COULD BE CONDUCTED ONLY TWICE DUE TO CARDIAC FAILURE. MYOCARDIAL INFARCTION WAS OBSERVED IN THE RIGHT CORONARY ARTERY (RCA) AND ISCHEMIA WAS OBSERVED IN THE LAD. THE PHYSICIAN SUGGESTS THAT THE IVL THERAPY MUST HAVE STRESSED THE HEART AS THE LAD WOULD BECOME ISCHEMIC. HE STATES THAT ASPIRATION PNEUMONITIS OCCURRED, AND THE GENERAL CONDITION OF THE PATIENT WAS BAD, THEREFORE CARDIAC ARREST COULD HAVE OCCURRED AT ANY MOMENT EVEN IF THE IVL THERAPY TRIGGERED IT. HOWEVER, HE STATES THAT IVL USAGE CANNOT BE EXCLUDED FROM THE POTENTIAL CAUSE OF CARDIAC ARREST AND THAT IT IS A POSSIBILITY. THE EXACT PROCEDURE DATE AND DEATH DATE ARE UNKNOWN. Manufacturer Narrative: THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION CANNOT BE PERFORMED. THE ROOT CAUSE OF THE BALLOON RUPTURE, CARDIAC ARREST, ASPIRATION PNEUMONITIS AND SUBSEQUENT DEATH COULD NOT BE DEFINITIVELY DETERMINED BASED ON THE INFORMATION AVAILABLE. THE PHYSICIAN STATED THAT THE PATIENT'S CONDITION PRIOR TO THE IVL PROCEDURE WAS EXTREMELY BAD. BASED ON REVIEW OF THE EVENT BY THE SHOCKWAVE MEDICAL SAFETY OFFICER, IT WAS DETERMINED THAT THE RELATIONSHIP OF THE DEVICE TO CARDIAC ARREST WAS PROBABLE AND THE RELATIONSHIP OF THE PROCEDURE TO CARDIAC ARREST WAS CAUSAL. BALLOON RUPTURES ARE A KNOWN FAILURE MODE WITH A LOW PROBABILITY OF OCCURRENCE. CONTRIBUTING FACTORS INCLUDE PATIENT CALCIUM, DAMAGE CAUSED WHEN THE BALLOON WALL COMES INTO CLOSE CONTACT WITH THE DEVICE EMITTERS, LIKELY CAUSES INCLUDE (A) AN IRREGULAR CALCIUM LESION CAPABLE OF COMPRESSING THE BALLOON WALL INTO CLOSE PROXIMITY WITH THE EMITTER(S) AND/OR (B) AN INCOMPLETELY INFLATED BALLOON. THE SHOCKWAVE INTRAVASCULAR LITHOTRIPSY (IVL) SYSTEM WITH THE SHOCKWAVE C2 CORONARY IVL CATHETER GENERATES ACOUSTIC PRESSURE PULSES WITHIN THE TARGET TREATMENT SITE, DISRUPTING CALCIUM WITHIN THE LESION ALLOWING SUBSEQUENT DILATATION OF A CORONARY ARTERY STENOSIS USING LOW BALLOON PRESSURE NOT TO EXCEED 4 ATM DURING IVL TREATMENT, AND 6 ATM FOR POST DILATATION. THE PRESSURE USED FOR THE IVL BALLOON CATHETER IS MUCH LOWER THAN OTHER BALLOON CATHETERS USED IN PERCUTANEOUS CORONARY INTERVENTION (PCI), THUS THE RISK ASSOCIATED WITH LOSS OF BALLOON PRESSURE DURING IVL CATHETER USE IS NEGLIGIBLE. THE ASSOCIATED PATIENT RISK FOR THIS FAILURE MODE IS LOW AND ADEQUATELY CAPTURED IN SWMI'S RISK DOCUMENTATION. THE LOT NUMBER OF THE DEVICE WAS NOT PROVIDED. THEREFORE, REVIEW OF THE DEVICE MANUFACTURING AND TEST DOCUMENTATION COULD NOT BE COMPLETED. HOWEVER, SHOCKWAVE MEDICAL INC. HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO DISTRIBUTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22216107&pc=QMG | 3015053838-2025-00075 |  | Death | SHOCKWAVE MEDICAL, INC. | 2025-06-16 04:00 | QMG | SHOCKWAVE C2 CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Aspiration Pneumonitis | P200039 |  | 1 | Event Description: TWO SHOCKWAVE C2 CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETERS WERE USED TO TREAT A LESION IN THE LEFT ANTERIOR DESCENDING ARTERY (LAD). THIS REPORT IS FOR THE SECOND IVL CATHETER. REFER TO MDR#: 3015053838-2025-00074 FOR THE FIRST IVL CATHETER. THE PATIENT WAS HOSPITALIZED FOR HEMODIALYSIS (HD). ECG ABNORMALITY WAS OBSERVED, AND THE PATIENT WAS THEN TRANSFERRED TO THE CARDIOVASCULAR INTERNAL MEDICAL DIVISION FOR CORONARY ANGIOGRAPHY (CAG) AND PERCUTANEOUS CORONARY INTERVENTION (PCI). IVL WAS PERFORMED AND AFTER 2 IVL PULSES WERE DELIVERED, THE BALLOON RUPTURED. IMMEDIATELY AFTER, CARDIAC ARREST OCCURRED. INTUBATION WAS PERFORMED AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) WAS INTRODUCED TO STABILIZE THE GENERAL CONDITION OF THE PATIENT. THE PROCEDURE WAS COMPLETED USING ANOTHER IVL CATHETER. THE PATIENT WAS THEN TRANSFERRED TO THE URGENT CARE CENTER (UCC). THE NEXT DAY, THE PATIENT WAS CONFIRMED TO BE STABILIZED AND EXTUBATED. THE ECMO WAS THEN DISCONTINUED. AFTER 3 WEEKS, THE PATIENT WAS THEN TRANSFERRED TO THE NEPH-INTERNAL MEDICINE DIVISION. 1 WEEK AFTER THE TRANSFER, THE PATIENT WAS STABLE AND DISCHARGED FROM THE HOSPITAL. THE PATIENT DEVELOPED ASPIRATION PNEUMONITIS AND WAS THEN TRANSFERRED TO A NURSERY HOME WHERE THE PATIENT PASSED. THE PHYSICIAN STATES THAT THE PATIENT'S STATUS WAS EXTREMELY BAD. ALTHOUGH HEMODIALYSIS (HD) WAS INTRODUCED, IT COULD NOT BE CONDUCTED TWICE DUE TO CARDIAC FAILURE. MYOCARDIAL INFARCTION WAS OBSERVED IN THE RIGHT CORONARY ARTERY (RCA) AND ISCHEMIA WAS OBSERVED IN THE LAD. THE PHYSICIAN SUGGESTS THAT THE IVL THERAPY MUST HAVE STRESSED THE HEART AS THE LAD WOULD BECOME ISCHEMIC. HE STATES THAT ASPIRATION PNEUMONITIS OCCURRED, AND THE GENERAL CONDITION OF THE PATIENT WAS BAD, THEREFORE CARDIAC ARREST COULD HAVE OCCURRED AT ANY MOMENT EVEN IF THE IVL THERAPY TRIGGERED IT. HOWEVER, HE STATES THAT IVL USAGE CANNOT BE EXCLUDED FROM THE POTENTIAL CAUSE OF CARDIAC ARREST AND THAT IT IS A POSSIBILITY. THE EXACT PROCEDURE DATE AND DEATH DATE ARE UNKNOWN. Manufacturer Narrative: THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION CANNOT BE PERFORMED. THE ROOT CAUSE OF THE CARDIAC ARREST, ASPIRATION PNEUMONITIS AND SUBSEQUENT DEATH COULD NOT BE DEFINITIVELY DETERMINED BASED ON THE INFORMATION AVAILABLE. THE PHYSICIAN STATED THAT THE PATIENT'S CONDITION PRIOR TO THE IVL PROCEDURE WAS EXTREMELY BAD. BASED ON REVIEW OF THE EVENT BY THE SHOCKWAVE MEDICAL SAFETY OFFICER, IT WAS DETERMINED THAT THE RELATIONSHIP OF THE DEVICE TO CARDIAC ARREST WAS PROBABLE AND THE RELATIONSHIP OF THE PROCEDURE TO CARDIAC ARREST WAS CAUSAL. THERE WAS NO MALFUNCTION REPORTED ON THE SECOND IVL CATHETER USED DURING THE PROCEDURE. FURTHERMORE, THE SECOND IVL CATHETER WAS INTRODUCED AFTER THE PATIENT WAS RESUSCITATED FOLLOWING THE CARDIAC ARREST. THIS REPORT IS BEING SUBMITTED OUT OF AN ABUNDANCE OF CAUTION SINCE THE CATHETER WAS USED DURING THE SAME PROCEDURE. THE LOT NUMBER OF THE DEVICE WAS NOT PROVIDED. THEREFORE, REVIEW OF THE DEVICE MANUFACTURING AND TEST DOCUMENTATION COULD NOT BE COMPLETED. HOWEVER, SHOCKWAVE MEDICAL INC. HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO DISTRIBUTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22216148&pc=NIP | 2017233-2025-06316 | 2025-05-20 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-16 04:00 | NIP | GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE | Insufficient Information | Insufficient Information | P040037 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT TREATMENT FOR AORTOILIAC OCCLUSIVE DISEASE IN THE DISTAL AORTA, BILATERAL COMMON AND EXTERNAL ILIAC ARTERIES. GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) DEVICES WERE IMPLANTED IN THE EXTERNAL ILIAC ARTERY. THE PHYSICIAN BEGAN WITH FEMORAL ENDARTERECTOMY OF THE COMMON FEMORAL ARTERIES. AFTER COMPLETION OF THE ENDARTERECTOMY, BILATERAL ACCESS WAS GAINED AND CORDIS 8 FR BRITE TIP® INTRODUCER SHEATHS WERE INSERTED. THE LESIONS WERE CROSSED; HOWEVER, THE PHYSICIAN ENCOUNTERED A SUBINTIMAL POSITION AT THE DISTAL AORTA. INITIALLY, A COOK MEDICAL BEBACK CROSSING CATHETER WAS ATTEMPTED TO BE USED. A PHILIPS PIONEER PLUS IVUS-GUIDED CATHETER WAS SELECTED FOR RE-ENTRY INTO THE DISTAL AORTA. TWO 8X79 VBX DEVICES WERE DEPLOYED BILATERALLY INTO THE DISTAL AORTA AND COMMON ILIAC ARTERIES USING A KISSING STENT TECHNIQUE. TO EXCLUDE THE AORTIC DISEASE, THE VBX DEVICES WERE SLIGHTLY RAISED IN THE BIFURCATION. TO PRESERVE THE HYPOGASTRIC ARTERY, AN 8X60 BOSTON SCIENTIFIC INNOVA¿ VASCULAR SELF-EXPANDING STENT SYSTEM (INNOVA STENT) WAS DEPLOYED INTO THE 8X79 VBX DEVICE THAT WAS IMPLANTED ON THE RIGHT SIDE. THE INNOVA STENT WAS OVERLAPPED WITH AN 8X59 VBX DEVICE THAT EXTENDED INTO THE EXTERNAL ILIAC ARTERY. THE 8X59 VBX BALLOON WAS INFLATED AGAIN TO BALLOON THE OVERLAP OF THE INNOVA STENT AND VBX DEVICE. AS REPORTED, THERE WERE NO ISSUES WITH BALLOON INFLATION. HOWEVER, WHEN ATTEMPTING TO WITHDRAW THE BALLOON CATHETER, THE PHYSICIAN NOTICED THE BALLOON CATHETER WAS STUCK TO THE GUIDEWIRE, PREVENTING ADVANCEMENT OR RETRACTION. SEVERAL ATTEMPTS WERE MADE TO DISENGAGE THE CATHETER BY PULLING NEGATIVE ON THE BALLOON, BUT ALL WERE UNSUCCESSFUL. IT IS UNKNOWN WHY THE BALLOON CATHETER BECAME STUCK ON THE GUIDEWIRE. THE BALLOON CATHETER, GUIDEWIRE, AND SHEATH WERE ALL REMOVED AS ONE FROM THE PATIENT. AFTER REMOVAL OF THE BALLOON CATHETER, GUIDEWIRE, AND SHEATH, THE PATIENT STARTED BLEEDING HEAVILY AND EXPERIENCED A DROP IN BLOOD PRESSURE, AND SUBSEQUENTLY CODED. A COOK MEDICAL CODA® BALLOON CATHETER WAS INSERTED FROM THE CONTRALATERAL SIDE TO CONTROL THE BLEEDING. AN 8 X 15 GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE WAS SUCCESSFULLY DEPLOYED ACROSS THE ILIAC ARTERY PERFORATION. CPR WAS PERFORMED FOR A SIGNIFICANT DURATION, BUT THE PATIENT EXPIRED. THE PATIENT'S DROP IN BLOOD PRESSURE WAS CAUSED BY AN ARTERIAL RUPTURE AND INTERNAL BLEEDING. HOWEVER, THE EXACT CAUSE OF DEATH REMAINS UNKNOWN. THE PHYSICIAN STATED THERE WAS SPECULATION THAT THE CATHETER, WHERE THE BALLOON WAS MOUNTED, MAY HAVE OVALIZED AND CAUSED THE CATHETER TO GET STUCK ON THE WIRE. THIS COULD NOT BE CONFIRMED AS THE BALLOON CATHETER AND GUIDEWIRE WERE IMMEDIATELY THROWN AWAY. Manufacturer Narrative: SECTION H6 UPDATED TO REFLECT COMPLETION OF INVESTIGATION. A REVIEW OF MANUFACTURING RECORDS VERIFIED THAT THE LOT INVOLVED IN THIS COMPLAINT MET ALL PRE-RELEASE SPECIFICATIONS. THE DEVICE REMAINS IMPLANTED AND WAS, THEREFORE, NOT AVAILABLE FOR ANALYSIS. NO CLINICAL IMAGES ENABLING DIRECT ASSESSMENT OF PRODUCT PERFORMANCE WERE RETURNED FOR EVALUATION. CAUSE OF THE REPORTED EVENT CANNOT BE ESTABLISHED BASED ON THE INFORMATION REPORTED TO GORE. FURTHER INFORMATION REGARDING THIS EVENT WAS REQUESTED BY GORE, BUT NO FURTHER INFORMATION HAS BEEN REPORTED, THEREFORE THIS INVESTIGATION IS CONSIDERED COMPLETE. Manufacturer Narrative: CBAS® HEPARIN SURFACE INCORPORATES CARMEDA HEPARIN MANUFACTURED FROM HEPARIN SODIUM API, WHICH IS COVALENTLY BOUND TO THE DEVICE SURFACE AND IS ESSENTIALLY NON-ELUTING. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22216206&pc=NPT | 2025587-2025-04258 | 2025-05-05 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-16 04:00 | NPT | MEDTRONIC TRANSCATHETER VALVE | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING THE PROGNOSTIC VALUE OF TRICUSPID ANNULAR DIMENSIONS IN TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI). THE STUDY POPULATION INCLUDED 522 PATIENTS FROM A SINGLE CENTER WHO WERE PREDOMINANTLY FEMALE WITH A MEAN AGE OF 82.1 YEARS OLD. ALL PATIENTS WERE IMPLANTED WITH A MEDTRONIC EVOLUT R OR EVOLUT PRO OR BIOPROSTHETIC VALVE. THREE PROCEDURAL DEATHS OCCURRED IN THE STUDY POPULATION. NO FURTHER DETAILS WERE PROVIDED ON THESE DEATHS. SEVEN DEATHS OCCURRED WITHIN 30 DAYS OF IMPLANT; HOWEVER, THERE WAS NO STATEMENT ESTABLISHING A CAUSAL OR CONTRIBUTORY RELATIONSHIP BETWEEN MEDTRONIC PRODUCT AND THE DEATHS. AMONG ALL PATIENTS, CLINICAL OBSERVATIONS INCLUDED: STROKE, ARRHYTHMIA REQUIRING PERMANENT PACEMAKER IMPLANT, AND MAJOR VASCULAR COMPLICATION. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. Manufacturer Narrative: CITATION: SCHÖRGHOFER ET AL. THE PROGNOSTIC VALUE OF TRICUSPID ANNULAR DIMENSIONS IN TAVI PATIENTS: A CT-BASED RETROSPECTIVE ANALYSIS OF RISK STRATIFICATION AND LONG-TERM OUTCOMES. J CLIN MED. 2025 MAY 5;14(9):3191. DOI: 10.3390/JCM14093191. EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22216253&pc=IYN | 3023245-2025-00051 | 2025-05-09 04:00 | Death | SIEMENS HEALTHINEERS | 2025-06-16 04:00 | IYN | ACUSON SEQUOIA | Use of Incorrect Control/Treatment Settings; Adverse Event Without Identified Device or Use Problem | Fetal Distress |  |  | 1 | Event Description: ON 12-JUN-2025, THE DESIGNATED COMPLAINT UNIT (DCU) RECEIVED A COMPLAINT (B)(4). PER THE SUBMITTER, ON 10-JUN-2025, A PATIENT SENT A LETTER TO THE CUSTOMER ADMINISTRATION (B)(6), STRATHPINE, AUSTRALIA REGARDING "USE OF A THERMAL INDEX FOR CRANIAL BONE (TIC) VALUE OF 3.03 AND 3.16 DURING TWO EARLY PREGNANCY ULTRASOUNDS". THE LETTER STATED THAT THE PATIENT HAD UNDERGONE TWO EARLY PREGNANCY ULTRASOUNDS, PERFORMED ON (B)(6) 2025 AND (B)(6) 2025 AT THE SITE USING THE ACUSON SEQUOIA SYSTEM WITH A THERMAL INDEX FOR CRANIAL BONE (TIC) VALUE OF 3.03 AND 3.16, RESPECTIVELY. THEN ON (B)(6) 2025, THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT AT (B)(6) HOSPITAL DUE TO BROWN DISCHARGE. AN ULTRASOUND WAS PERFORMED AT THE EMERGENCY ROOM, AND IT WAS FOUND THAT THE BABY HAD NO HEARTBEAT, AND ITS GROWTH CORRESPONDED TO 8 WEEKS AND 3/4 DAYS. THE PATIENT FURTHER ADDED THAT 2ND ULTRASOUND ON (B)(6) 2025 WITH (B)(6) HAD SHOWN A STRONG HEARTBEAT OF 130BPM WITH 8 WEEKS 4 DAYS GROWTH WHICH CONFIRMS THAT THE BABY DIED IN THE HOURS FOLLOWING THE ULTRASOUND.
THE PATIENT STATES THAT DURING BOTH HER ULTRASOUND SCANS AT (B)(6) WERE PERFORMED AT THERMAL INDEX FOR CRANIAL BONE (TIC) VALUE OF 3.03 AND 3.16. PER THE LETTER, "ACCORDING TO INTERNATIONALLY ACCEPTED GUIDELINES, TI AND MI VALUES DURING EARLY PREGNANCY SHOULD BE KEPT BELOW 1.0, PREFERABLY UNDER 0.7 - PARTICULARLY IN THE FIRST TRIMESTER DUE TO THE INCREASED SENSITIVITY OF DEVELOPING FETAL TISSUES TO THERMAL EFFECTS (SOURCE: AIUM.ORG: PRUDENT USE AND SAFETY OF DIAGNOSTIC ULTRASOUND IN PREGNANCY)." THE PATIENT STATES THAT THERE WAS NO ADVISE PROVIDED RELATED TO THE ELEVATED THERMAL OR MECHANICAL OUTPUT OR ITS POTENTIAL IMPLICATIONS, NOR WAS INFORMED ABOUT CONSENT OBTAINED FOR THIS LEVEL OF EXPOSURE BY THE CUSTOMER AT (B)(6). Manufacturer Narrative: A REVIEW OF THE AIUM GUIDELINES (SOURCE: AIUM.ORG: PRUDENT USE AND SAFETY OF DIAGNOSTIC ULTRASOUND IN PREGNANCY) FOR EARLY OB SCANS (<10 WEEKS FROM LMP) WAS DONE AND THE RECOMMENDATION IS TO MONITOR TIS (THERMAL INDEX FOR SOFT TISSUES) VALUES. THE ASUM, ISUOG GUIDELINES FOR USAGE OF ULTRASOUND DEVICES FOR FIRST TRIMESTER IMAGING STATES "WHEN PERFORMING DOPPLER ULTRASOUND, THE DISPLAYED THERMAL INDEX (TI) SHOULD BE LESS THAN OR EQUAL TO 1.0 AND EXPOSURE TIME SHOULD BE KEPT AS SHORT AS POSSIBLE (USUALLY NO LONGER THAN 5-10 MIN)." TIC (THERMAL INDEX FOR CRANIAL BONE) VALUES ARE NOT RELEVANT FOR EARLY OB SCANS, AS TIC IS ONLY FOR SKULL (CRANIAL) IMAGING OR SIMILAR EXAMS, WITH WIDE, THICK BONE IMMEDIATELY BELOW THE SKIN. IN EARLY PREGNANCY, THAT THICK SURFACE BONE IS ABSENT IN THE FETUS. REVIEW OF THE PATIENT IMAGES FROM ON (B)(6) 2025 FOUND THAT THE RELEVANT THERMAL INDEX FOR THIS EXAM IS TIS. IN THE FIRST SCAN (ON (B)(6) 2025) TIS WAS BELOW 0.7 AND WAS NEVER ABOVE 1.0 FOR MORE THAN ONE MINUTE. IN THE SECOND SCAN (ON (B)(6) 2025), TIS WAS ALWAYS BELOW 0.7. THERE WAS NO DEVIATION FROM THE STANDARD OF CARE FROM THE ASUM GUIDELINES. IN THIS CASE, THE CUSTOMER PRESET ON THE SYSTEM WAS SET TO DISPLAY THE TIC VALUES DURING THE SCANS. THE SYSTEM DISPLAYED TIC VALUE >3 AS THE THERMAL INDEX FOR CRANIAL BONE IS CALCULATED AS IF THICK SURFACE BONE IS PRESENT. THE SYSTEM CALCULATION IS NOT AWARE WHETHER SUCH A BONE IS PRESENT OR NOT DURING THIS EXAM. WHEN THE USERS KNOW THAT SUCH BONE IS NOT PRESENT, SUCH AS EARLY PREGNANCY CASES (<10 WEEKS), THEN THE USERS KNOW THAT VALUE IS NOT APPLICABLE. THUS,THE OUTPUT LEVELS USED IN THE SCANS FROM THE CUSTOMER USING THE ACUSON SEQUOIA SYSTEM ARE WITHIN THE AIUM/ASUM GUIDELINES FOR POWER. THEREFORE, IT IS UNLIKELY THAT THE DEVICE MALFUNCTION CAUSED OR CONTRIBUTED TO THE DEATH. REFERENCE: (B)(4). Manufacturer Narrative: BASED ON PRELIMINARY INVESTIGATION, IT IS UNLIKELY THAT THE SIEMENS DEVICE MALFUNCTION MAY HAVE CAUSED OR CONTRIBUTED TO THE DEATH. SIEMENS DEVICES FOLLOW THE AIUM GUIDELINES FOR POWER AND PROVIDES WARNINGS & GUIDELINES TO USERS VIA INSTRUCTIONS FOR USE ON SAFE USAGE OF THE DEVICES DURING SCANS. SIEMENS IS PERFORMING A DETAILED INVESTIGATION OF THE EVENT AND WILL SUPPLEMENT THE REPORT AFTER COMPLETION. REFERENCE: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22216289&pc=DSQ | 2916596-2025-03554 | 2025-01-28 05:00 | Death | THORATEC CORPORATION | 2025-06-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MALIGNANT, REFRACTORY RHYTHM DISTURBANCES. THE PATIENT SUFFERED FROM REFRACTORY, MALIGNANT RHYTHM DISTURBANCES WHICH LEAD TO THE OUTCOME AT LEAST. THE PATIENT DID RECEIVE MEDICAL OPTIMIZATION, CARDIOVERSION, AND DEFIBRILLATION AS TREATMENT. NO DEVICE RELATED ISSUES WERE MENTIONED AND/OR WERE RELATED TO THE DISTURBANCES. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED.Event Description: THE PATIENT EXPIRED IN AN EXTERNAL HOSPITAL SO THE DETAILS RELATED TO THE PATIENT'S PASSING WERE LESS DETAILS. SYMPTOMS WERE NOT PROVIDED, BUT THE PATIENT DID RECEIVE MEDICAL OPTIMIZATION, CARDIOVERSION, AND DEFIBRILLATION AS TREATMENT. THE TYPE OF ARRHYTHMIA WAS NOT PROVIDED, AND IT WAS UNKNOWN IF THE PATIENT HAD A HISTORY OF ARRHYTHMIA. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. B IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING CARDIAC ARRHYTHMIA AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿, LISTS ARRHYTHMIA AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22216410&pc=LWP | 2124215-2025-39703 | 2025-05-20 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-16 04:00 | LWP | ACCOLADE MRI DR | Incorrect, Inadequate or Imprecise Result or Readings; Telemetry Discrepancy | No Clinical Signs, Symptoms or Conditions | P150012/S000 |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER RETURNED TO STORAGE MODE AND COULD NOT BE INTERROGATED. IT WAS NOTED THAT THE PATIENT'S LAST FOLLOW-UP OCCURRED APPROXIMATELY FOUR YEARS AGO. AT THIS FOLLOW-UP, THERE WAS A PROJECTED LONGEVITY OF FOUR YEARS. THE PATIENT DID NOT ATTEND ANY CONSULTATION SUBSEQUENTLY. THE PHYSICIAN HAS ALERTED THE PATIENT, BUT THE PATIENT HAS NOT BEEN ABLE TO BE REACHED. THE DEVICE WILL NOT BE CHANGED OUT AS THE PATIENT'S CONDITION DOES NOT ALLOW THEM TO HAVE INVASIVE PROCEDURES. RECENTLY, THE PATIENT PASSED AWAY DUE TO A DIGESTIVE OBSTRUCTION. HOWEVER, THE PHYSICIAN WAS UNABLE TO CONCLUSIVELY DETERMINE WHETHER THE DEVICE PERFORMANCE AND BATTERY DEPLETION CONTRIBUTED TO THE PATIENT'S DEATH. IT IS CURRENTLY UNKNOWN IF THE DEVICE WILL BE RETURNED FOR ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22216448&pc=DSQ | 2916596-2025-03555 | 2023-01-02 05:00 | Death | THORATEC CORPORATION | 2025-06-16 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Fever; Pain; Sepsis; Post Traumatic Wound Infection; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A STAPHYLOCOCCUS AUREUS INFECTION WAS. THEY WERE ADMITTED SEVERAL TIMES FOR THE INFECTION. THE FIRST ADMISSION WAS ON (B)(6) 2023 FOR SEPSIS AND STAPHYLOCOCCUS AUREUS. THE SECOND ADMISSION WAS IN (B)(6) 2023, ALSO FOR STAPHYLOCOCCUS AUREUS. THE PATIENT WAS ADMITTED IN (B)(6) 2024 FOR VACUUM-ASSISTED CLOSURE (VAC) THERAPY, AND THEN AGAIN ON (B)(6) 2024 FOR STAPHYLOCOCCUS AUREUS OF THE STERNUM WHICH ENDED IN SEPSIS. THEY RECEIVED CHRONIC ANTIBIOTIC THERAPY WITH AN ADDITIONAL TWO TIMES OF VAC THERAPY TO CLEAN THE WOUND. THE PATIENT EXPERIENCED WOUND PAIN AND A FEVER. DRIVELINE/DEVICE INFECTION WAS ALSO PRESENT BUT THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS RUNNING AS EXPECTED. THE PATIENT PASSED AWAY ON 05DEC2024 DUE TO SEPSIS DUE TO CHRONIC STERNUM INFECTION. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION D6A: DATE OF IMPLANT CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22216482&pc=MVK | 3008642652-2025-05872 | 2025-06-03 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-16 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT AND AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 14:47:17 ON (B)(6) 2025. ( AT 18:19:48 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. AT 18:20:36, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS FINE VF WITH CPR/MOTION ARTIFACT. AT 18:21:26, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. RHYTHM AT THE TIME OF TREATMENT WAS FINE VF WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT. THE DEVICE WAS SHUT DOWN AT 19:01:43 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22216483&pc=MVK | 3008642652-2025-05873 | 2025-05-14 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-16 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 23:14:57 ON (B)(6) 2025. AT 06:20:29 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 06:21:19, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL, CPR/MOTION ARTIFACT, AND FLUTTER WAVES CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. AT 06:25:02, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE WITH CPR/MOTION ARTIFACT. AT 06:26:08, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL, CPR/MOTION ARTIFACT, AND FLUTTER WAVES CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. THE DEVICE WAS SHUT DOWN AT 06:26:39 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22200292&pc=FGE | 3005099803-2025-02729 | 2025-05-10 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-13 04:00 | FGE | WALLFLEX BILIARY RX STENT SYSTEM | Migration | Abdominal Pain; Fever; Perforation; Vomiting; Hernia; Peritonitis; Swelling/ Edema |  |  | 1 | Event Description: IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT A WALLFLEX BILIARY STENT WAS IMPLANTED IN THE BILE DUCT TO TREAT SEVERAL CHOLANGITIS EPISODES DURING AN ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY (ERCP) PROCEDURE PERFORMED ON (B)(6) 2024. MONTHS POST-PROCEDURE, THE PATIENT EXPERIENCED PAIN IN THE STOMACH AREA AND ON (B)(6) 2025, A CT SCAN WAS PERFORMED. THE RESULTS SHOWED THE STENT MIGRATED INTO THE SMALL BOWEL. THE CT SCAN SHOWED PARASTOMAL HERNIA BUT THERE WERE NO SIGNS OF INCARCERATION OR PERFORATION. THE PATIENT DEVELOPED INCREASED STOMACH, FEVER, INCREASED C-REACTIVE PROTEIN (CRP) LEVEL (10MG/L) AND WHITE CELLS. ANOTHER CT SCAN WAS PERFORMED ON (B)(6) 2025, THE PATIENT EXPERIENCED VOMITING AND NO GAS OR FECES EMPTYING FROM THE STOMA, AND A PERFORATION WAS NOTICED. THIS NEW CT SCAN SHOWED SMALL AMOUNTS OF EXTRALUMINAL GAS AND AN EDEMA AROUND A SHORT PART OF THE SMALL BOWEL JUST ORALLY TO THE DISLOCATED STENT. ON (B)(6) 2025, THE C-REACTIVE PROTEIN LEVEL OF THE PATIENT WAS 221 MG/L. THE PATIENT HAD A DISLOCATION OF THE STENT FROM THE MAIN BILIARY DUCT, AND BECAUSE OF A HERNIA THE STENT COULD NOT PASS THROUGH THE SMALL BOWEL AND THE STENT ERODED AND FINALLY PERFORATED THE BOWEL WALL. AFTER THE PERFORATION WAS DIAGNOSED, PERITONITIS FOLLOWED THE ADVERSE EVENT, AND AT THIS POINT IT WAS DECIDED THAT THE PATIENT SHOULD JUST RECEIVE PALLIATIVE CARE. THIS MEANS PAINKILLERS, SEDATIVE, STEROIDS ETC. THE PATIENT DIED DUE TO COMPLICATIONS RELATED TO THE PERFORATION ON (B)(6) 2025 Manufacturer Narrative: BLOCK D4, H4: THE COMPLAINANT WAS UNABLE TO PROVIDE THE COMPLAINT DEVICE LOT NUMBER. THEREFORE, THE MANUFACTURE AND EXPIRATION DATES ARE UNKNOWN; AND WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK H6: IMDRF DEVICE CODE A010402 CAPTURES THE REPORTABLE EVENT OF STENT MIGRATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22200406&pc=CAW | 3014777423-2025-00005 | 2025-05-20 04:00 | Death | QINGDAO KINGON MEDICAL SCIENCE AND TECHNOLOGY CO., | 2025-06-13 04:00 | CAW | REACT POC2000 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K223379 |  | 1 | Event Description: REACT HEALTH RECEIVED A COMPLAINT FROM A DURABLE MEDICAL EQUIPMENT PROVIDER OF A PATIENT PASSING AWAY WHILE USING THE DEVICE. THE DEVICE WAS RETURNED TO REACT HEATH AND EVALUATED ON 28-MAY-2025, WHERE IT WAS DETERMINED TO BE OPERATING ACCORDING TO THE MANUFACTURER'S SPECIFICATIONS. THE MANUFACTURER HAS NOT TAKEN POSSESSION OF THE DEVICE AS OF THIS REPORT. Manufacturer Narrative: THE INITIAL REPORT DID NOT STATE THAT A MALFUNCTION HAD OCCURRED. THE INVESTIGATION BY OUR DISTRIBUTOR IN THE USA CONCLUDED THAT THE DEVICE WAS PERFORMING AND WORKING AS INTENDED. THERE IS NO INDICATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT DEATH AND APPEARS TO HAVE SIMPLY BEEN A CONCOMITANT DEVICE. THE INVESTIGATION SAYS, "THE DEVICE WAS TESTED FOR 25 MINUTES. THE CLINICAL MENU SHOWS O2 AT 95%. THE DEVICE O2 LEVEL WAS ALSO CHECKED WITH AN OXYGEN ANALYZER, READING AT 94% OF O2 PURITY. THERE WERE NO SHUTDOWN ALARMS OR ALERTS LOGGED. THE DEVICE DOES NOT SHOW ANY SIGNS OF COSMETIC DAMAGE. THE RESULT OF THE INVESTIGATION IS THAT THE DEVICE IS WORKING AS INTENDED. NO ADDITIONAL INFORMATION IS AVAILABLE."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22200530&pc=DRA | 2124215-2025-38604 | 2025-04-18 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-13 04:00 | DRA | FARADRIVE STEERABLE SHEATH | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER THE PATIENT PASSED AWAY. A CONCOMITANT LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED AND A 31MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED FOLLOWING A PFA ABLATION TREATMENT USING A FARAWAVE PFA CATHETER. IT WAS REPORTED THAT THERE WERE NO COMPLICATIONS OR DIFFICULTIES DURING EITHER THE WATCHMAN NOR THE PFA PORTION OF THE PROCEDURE. SIX (6) DAYS POST INDEX PROCEDURE, THE PATIENT WAS DISCHARGED TO HOSPICE AND THE FOLLOWING DAY, THE PATIENT DIED. THE CAUSE OF DEATH REMAINS UNKNOWN, AND NO FURTHER DETAILS WERE AVAILABLE PER THE ACCOUNT. THE CATHETER WAS DISPOSED BY THE FACILITY AND WILL NOT BE RETURNED AS THE DEVICE PERFORMED AS EXPECTED DURING THE PROCEDURE AND THERE WERE NO REPORTS OF ANY PERFORMANCE CONCERNS. Manufacturer Narrative: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION BUT FURTHER INFORMATION ABOUT THE EVENT, INCLUDING SPECIFIC PRODUCT INFORMATION, WAS NOT AVAILABLE PER THE FACILITY. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22200967&pc=PNJ | 9612164-2025-02903 | 2025-05-09 04:00 | Death | MEDTRONIC IRELAND | 2025-06-13 04:00 | PNJ | MICRA¿ VR2 | Adverse Event Without Identified Device or Use Problem | Shock | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THREE DAYS POST IMPLANT OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE PATIENT EXPERIENCED AN INFECTION AND THE PATIENT PASSED.Event Description: IT WAS NOTED POST LEADLESS IPG PLACEMENT THAT THE PATIENT DEVELOPED MULTIORGAN FAILURE/SHOCK WITH LEFT VENTRICULAR EJECTION FRACTION DROPPING TO TEN PERCENT. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2VR01-DELSYS] (SERIAL: [(B)(6)]). D9: NO MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22201291&pc=PNJ | 9612164-2025-02908 | 2025-06-11 04:00 | Death | MEDTRONIC IRELAND | 2025-06-13 04:00 | PNJ | MICRA | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT DURING ATTEMPTED IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE PATIENT EXPERIENCED A CARDIAC PERFORATION RESULTING IN TAMPONADE AND PERICARDIAL EFFUSION. PERICARDIOCENTESIS WAS PERFORMED AND THE PATIENT WAS HOSPITALISED. THE LEADLESS IPG, DELIVERY SYSTEM AND INTRODUCER WERE ATTEMPTED/NOT USED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT DIED THREE DAYS AFTER UNSUCCESSFUL IMPLANT DUE TO INFECTION.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENTS CAUSE OF DEATH WAS SEPSIS AND THE ORIGIN OF THE SEPSIS WAS UNKNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA VR] PRODUCT ID [MC1VR01-DELSYS] (SERIAL: [(B)(6)). D9: NO MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED: B1, B2. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22201295&pc=DYB | 9612164-2025-02909 | 2025-06-11 04:00 | Death | MEDTRONIC IRELAND | 2025-06-13 04:00 | DYB | MICRA | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT DURING ATTEMPTED IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT EXPERIENCED A CARDIAC PERFORATION RESULTING IN TAMPONADE AND PERICARDIAL EFFUSION. PERICARDIOCENTESIS WAS PERFORMED AND THE PATIENT WAS HOSPITALISED. THE LEADLESS IPG, DELIVERY SYSTEM AND INTRODUCER WERE ATTEMPTED/NOT USED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT DIED THREE DAYS AFTER UNSUCCESSFUL IMPLANT DUE TO INFECTION.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENTS CAUSE OF DEATH WAS SEPSIS AND THE ORIGIN OF THE SEPSIS WAS UNKNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22201504&pc=NIP | 2183870-2025-00303 | 2025-05-20 04:00 | Death | COVIDIEN | 2025-06-13 04:00 | NIP | EVERFLEX SELF-EXPANDING PERIPHERAL STENT WITH ENTRUST DELIVERY SYSTEM LONG | Activation, Positioning or Separation Problem; Physical Resistance/Sticking | Hemorrhage/Bleeding; Perforation | P110023 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROCEDURE TO TREAT ILIAC ARTERY DISEASE UTILIZING THE DRUG COATED BALLOON 018 IPU06015013P D6.0 L150 UL0130 AND THE EVERFLEX ENTRUST 6X150X120 STENT. DURING THE PROCEDURE, MODERATE RESISTANCE WAS ENCOUNTERED WHEN ADVANCING BOTH DEVICES THROUGH THE SEVERELY CALCIFIED AND TORTUOUS RIGHT AND LEFT COMMON ILIAC ARTERIES. EXCESSIVE FORCE WAS USED DURING ADVANCEMENT OF THE DRUG-COATED BALLOON. A 7FR NON-MEDTRONIC SHEATH AND A NON-MEDTRONIC GUIDEWIRE WERE USED. A NON-MEDTRONIC INTRAVASCULAR LITHOTRIPSY CATHETER WAS USED FOR VESSEL PRE-DILATION. AN INFLATION DEVICE AND 50/50 CONTRAST INFLATION FLUID WERE USED FOR BALLOON INFLATION. THE DRUG-COATED BALLOON EXPERIENCED DEFLATION DIFFICULTIES DUE TO THE AMOUNT OF CALCIUM PRESENT, RESULTING IN THE BALLOON NOT FULLY DEFLATING AND BECOMING STUCK WITHIN THE VESSEL. NEGATIVE PRESSURE WAS APPLIED IN AN ATTEMPT TO DEFLATE THE BALLOON, BUT IT REMAINED INCOMPLETELY DEFLATED. THE BALLOON SUBSEQUENTLY DETACHED IN THE VESSEL AND WAS REMOVED USING A SNARE. THE EVERFLEX ENT RUST STENT EXPERIENCED A THUMBWHEEL MALFUNCTION, LEADING TO PARTIAL DEPLOYMENT AND THE CATHETER HANDLE WAS DELIBERATELY CRACKED TO AID STENT DEPLOYMENT. THE PHYSICIAN ATTEMPTED TO RE-SHEATH AND REMOVE THE PARTIALLY DEPLOYED STENT, BUT THE STENT DEPLOYED INSIDE THE SHEATH. MULTIPLE ATTEMPTS WERE MADE TO REMOVE THE STENT AND SHEATH USING A SNARE AND BALLOON AND SHEATH, BUT REMOVAL WAS UNSUCCESSFUL DUE TO SEVERE CALCIUM AND VESSEL TORTUOSITY. ULTIMATELY, THE NON-MEDTRONIC SNARE PERFORATED THE ILIAC ARTERY, RESULTING IN VESSEL DAMAGE AND INTERNAL BLEEDING. A VASCULAR SURGEON WAS CALLED TO ASSIST, AND ALL DEVICES WERE REMOVED FROM THE PATIENT. NON-MEDTRONIC COVERED STENTS WERE PLACED TO CONTROL THE BLEEDING. THE PATIENT WAS INTUBATED FOLLOWING THE COMPLICATION AND WAS INITIALLY REPORTED AS STABLE. HOWEVER, THE PATIENT LATER DIED DUE TO GROIN COMPLICATION BLEEDING. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: THERE WAS 10% LEFT TO DEFLATE PRIOR TO THE DETACHMENT IN THE VESSEL. NEITHER THE INPACT OR EVERFLEX WERE USED IN THE ILIAC ARTERIES, THE BALLOON AND STENT WERE BOTH USED IN THE SFA. THE INPACT ADMIRAL WAS NOT RELATED TO THE PATIENT DEATH, NOR WAS THE EVERFLEX STENT RELATED TO THE DEATH. BOTH THE STENT AND BALLOON WERE REMOVED WELL BEFORE THE PATIENT PASSED. THE RESULT OF DEATH WAS DUE TO GROIN COMPLICATION BLEEDING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22201641&pc=NVZ | 3004209178-2025-10148 | 2025-05-05 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-13 04:00 | NVZ | AZURE XT DR MRI SURESCAN | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED BY THE PATIENTS REPRESENTATIVE THAT THE PATIENT PASSED AWAY AND WAS HAVING HEART ISSUES. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22201807&pc=ONU | 9612164-2025-02913 | 2025-05-20 04:00 | Death | MEDTRONIC IRELAND | 2025-06-13 04:00 | ONU | IN.PACT 018 PACLITAXEL-COATED PERCUTANEOUS TRANSLUMINAL ANGIOPLASTY (PTA) BALLOO | Deflation Problem; Detachment of Device or Device Component | Hemorrhage/Bleeding; Perforation | P140010 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A PROCEDURE TO TREAT ILIAC ARTERY DISEASE UTILIZING THE DRUG COATED BALLOON 018 IPU06015013P D6.0 L150 UL0130 AND THE EVERFLEX ENTRUST 6X150X120 STENT. DURING THE PROCEDURE, MODERATE RESISTANCE WAS ENCOUNTERED WHEN ADVANCING BOTH DEVICES THROUGH THE SEVERELY CALCIFIED AND TORTUOUS RIGHT AND LEFT COMMON ILIAC ARTERIES. EXCESSIVE FORCE WAS USED DURING ADVANCEMENT OF THE DRUG-COATED BALLOON. A 7FR NON-MEDTRONIC SHEATH AND A NON-MEDTRONIC GUIDEWIRE WERE USED. A NON-MEDTRONIC INTRAVASCULAR LITHOTRIPSY CATHETER WAS USED FOR VESSEL PRE-DILATION. AN INFLATION DEVICE AND 50/50 CONTRAST INFLATION FLUID WERE USED FOR BALLOON INFLATION. THE DRUG-COATED BALLOON EXPERIENCED DEFLATION DIFFICULTIES DUE TO THE AMOUNT OF CALCIUM PRESENT, RESULTING IN THE BALLOON NOT FULLY DEFLATING AND BECOMING STUCK WITHIN THE VESSEL. NEGATIVE PRESSURE WAS APPLIED IN AN ATTEMPT TO DEFLATE THE BALLOON, BUT IT REMAINED INCOMPLETELY DEFLATED. THE BALLOON SUBSEQUENTLY DETACHED IN THE VESSEL AND WAS REMOVED USING A SNARE. THE EVERFLEX ENT RUST STENT EXPERIENCED A THUMBWHEEL MALFUNCTION, LEADING TO PARTIAL DEPLOYMENT AND THE CATHETER HANDLE WAS DELIBERATELY CRACKED TO AID STENT DEPLOYMENT. THE PHYSICIAN ATTEMPTED TO RE-SHEATH AND REMOVE THE PARTIALLY DEPLOYED STENT, BUT THE STENT DEPLOYED INSIDE THE SHEATH. MULTIPLE ATTEMPTS WERE MADE TO REMOVE THE STENT AND SHEATH USING A SNARE AND BALLOON AND SHEATH, BUT REMOVAL WAS UNSUCCESSFUL DUE TO SEVERE CALCIUM AND VESSEL TORTUOSITY. ULTIMATELY, THE NON-MEDTRONIC SNARE PERFORATED THE ILIAC ARTERY, RESULTING IN VESSEL DAMAGE AND INTERNAL BLEEDING. A VASCULAR SURGEON WAS CALLED TO ASSIST, AND ALL DEVICES WERE REMOVED FROM THE PATIENT. NON-MEDTRONIC COVERED STENTS WERE PLACED TO CONTROL THE BLEEDING. THE PATIENT WAS INTUBATED FOLLOWING THE COMPLICATION AND WAS INITIALLY REPORTED AS STABLE. HOWEVER, THE PATIENT LATER DIED DUE TO GROIN COMPLICATION BLEEDING. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: THERE WAS 10% LEFT TO DEFLATE PRIOR TO THE DETACHMENT IN THE VESSEL. NEITHER THE INPACT OR EVERFLEX WERE USED IN THE ILIAC ARTERIES, THE BALLOON AND STENT WERE BOTH USED IN THE SFA. THE INPACT ADMIRAL WAS NOT RELATED TO THE PATIENT DEATH, NOR WAS THE EVERFLEX STENT RELATED TO THE DEATH. BOTH THE STENT AND BALLOON WERE REMOVED WELL BEFORE THE PATIENT PASSED. THE RESULT OF DEATH WAS DUE TO GROIN COMPLICATION BLEEDING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22202154&pc=NPT | 2025587-2025-04219 | 2025-06-06 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-13 04:00 | NPT | EVOLUT PRO PLUS DCS | Failure to Advance; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING AN ATTEMPTED IMPLANTATION OF AN EVOLUT 26 PRO+ VALVE, THE VALVE WAS ADVANCED FROM THE RIGHT FEMORAL ARTERY ACCESS UNTIL THE AORTIC ARCH, WHERE DIFFICULTY WAS ENCOUNTERED IN ADVANCING FURTHER. A SUPPORTIVE GUIDE WAS INSERTED INTO THE CONTRALATERAL FEMORAL PIGTAIL WITHOUT IMPROVEMENT. A SNARE WAS THEN INSERTED FROM THE CONTRALATERAL FEMORAL ACCESS, ALLOWING THE GUIDE AND DELIVERY CATHETER SYSTEM TO BE PASSED. THE VALVE WAS AGAIN ADVANCED FROM THE FEMORAL ACCESS, AND IN THE ABDOMINAL AORTA, THE SNARE WAS CLOSED ON THE DISTAL SIDE OF THE CAPSULE. FURTHER ATTEMPTS TO PASS INTO THE DESCENDING AORTIC ARCH WHILE HOLDING THE SNARE IN TENSION WERE INEFFECTIVE. THE SNARE WAS REPOSITIONED ON THE NOSE CONE AND ANOTHER ATTEMPT WAS MADE, AT WHICH POINT A DROP IN BLOOD PRESSURE OCCURRED. IMMEDIATE ANGIOGRAPHY REVEALED A DISSECTION OF THE AORTIC ARCH BELOW THE SUBCLAVIAN ARTERY. SUBSEQUENTLY, THE PATIENT DIED THE SAME DAY. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH WAS THE RUPTURE OF THE AORTIC ARCH AND IT WAS UNKNOWN IF THE WALL OF THE AORTIC ARCH WAS ALREADY RUPTURED BY THE WIRE OR IF IT WAS CAUSED BY THE DELIVERY CATHETER SYSTEM (DCS).Event Description: IT WAS REPORTED THAT DURING AN ATTEMPTED IMPLANTATION OF AN EVOLUT 26 PRO+ VALVE, THE VALVE WAS ADVANCED FROM THE RIGHT FEMORAL ARTERY ACCESS UNTIL THE AORTIC ARCH, WHERE DIFFICULTY WAS ENCOUNTERED IN ADVANCING FURTHER. A SUPPORTIVE GUIDE WAS INSERTED INTO THE CONTRALATERAL FEMORAL PIGTAIL WITHOUT IMPROVEMENT. A SNARE WAS THEN INSERTED FROM THE CONTRALATERAL FEMORAL ACCESS, ALLOWING THE GUIDE AND DELIVERY CATHETER SYSTEM TO BE PASSED. THE VALVE WAS AGAIN ADVANCED FROM THE FEMORAL ACCESS, AND IN THE ABDOMINAL AORTA, THE SNARE WAS CLOSED ON THE DISTAL SIDE OF THE CAPSULE. FURTHER ATTEMPTS TO PASS INTO THE DESCENDING AORTIC ARCH WHILE HOLDING THE SNARE IN TENSION WERE INEFFECTIVE. THE SNARE WAS REPOSITIONED ON THE NOSE CONE AND ANOTHER ATTEMPT WAS MADE, AT WHICH POINT A DROP IN BLOOD PRESSURE OCCURRED. IMMEDIATE ANGIOGRAPHY REVEALED A DISSECTION OF THE AORTIC ARCH BELOW THE SUBCLAVIAN ARTERY. SUBSEQUENTLY, THE PATIENT DIED THE SAME DAY. ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH WAS THE RUPTURE OF THE AORTIC ARCH AND IT WAS UNKNOWN IF THE WALL OF THE AORTIC ARCH WAS ALREADY RUPTURED BY THE WIRE OR IF IT WAS CAUSED BY THE DELIVERY CATHETER SYSTEM (DCS). ADDITIONAL INFORMATION WAS RECEIVED WHICH CLARIFIED A NON-MEDTRONIC (SYMEDRIX INOWWI) GUIDEWIRE WAS USED FOR THE CASE.Event Description: IT WAS REPORTED THAT DURING AN ATTEMPTED IMPLANTATION OF AN EVOLUT 26 PRO+ VALVE, THE VALVE WAS ADVANCED FROM THE RIGHT FEMORAL ARTERY ACCESS UNTIL THE AORTIC ARCH, WHERE DIFFICULTY WAS ENCOUNTERED IN ADVANCING FURTHER. A SUPPORTIVE GUIDE WAS INSERTED INTO THE CONTRALATERAL FEMORAL PIGTAIL WITHOUT IMPROVEMENT. A SNARE WAS THEN INSERTED FROM THE CONTRALATERAL FEMORAL ACCESS, ALLOWING THE GUIDE AND DELIVERY CATHETER SYSTEM TO BE PASSED. THE VALVE WAS AGAIN ADVANCED FROM THE FEMORAL ACCESS, AND IN THE ABDOMINAL AORTA, THE SNARE WAS CLOSED ON THE DISTAL SIDE OF THE CAPSULE. FURTHER ATTEMPTS TO PASS INTO THE DESCENDING AORTIC ARCH WHILE HOLDING THE SNARE IN TENSION WERE INEFFECTIVE. THE SNARE WAS REPOSITIONED ON THE NOSE CONE AND ANOTHER ATTEMPT WAS MADE, AT WHICH POINT A DROP IN BLOOD PRESSURE OCCURRED. IMMEDIATE ANGIOGRAPHY REVEALED A DISSECTION OF THE AORTIC ARCH BELOW THE SUBCLAVIAN ARTERY. SUBSEQUENTLY, THE PATIENT DIED THE SAME DAY. Manufacturer Narrative: UPDATED: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID EVPROPLUS-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: N/A; EXPLANT DATE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A THIRD PARAGRAPH WAS ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: IMAGE REVIEW: PRE-IMPLANT COMPUTED TOMOGRAPHY (CT) IMAGING WAS PROVIDED WITH THE CHEST CT CONTAINING MOTION ARTIFACT/BLURRING, CONSIDER THE MEASUREMENTS ESTIMATES. ANNULUS PERIMETER AND PERIMETER DERIVED DIAMETER 67.0 MM/21.3 MM SUGGESTING A 26 MM VALVE. LEFT VENTRICULAR OUTFLOW TRACT (LVOT) PERIMETER AND PERIMETER DERIVED DIAMETER 64.4 MM/20.5 MM. THE AORTIC VALVE IS TRILEAFLET WITH MILD CALCIFICATION. ALL SINUS OF VALSALVA (SOV) DIAMETERS ARE BELOW THE RECOMMENDED RANGE OF 27 MM. THE PATIENT APPEARS NATIVE CORONARY ARTERY DEPENDENT. ALL SINUS HEIGHTS AND CORONARIES ARE WITHIN THE RECOMMENDED RANGES. ANNULAR ANGULATION 55°. DUE TO FEMORAL CT CONTAINING 1.2 MM SLICE THICKNESS WITH NOISE ARTIFACT/GRAINY IMAGING, CONSIDER MEASUREMENTS ESTIMATES. THE CASE REPORT WAS PROVIDED FOR REVIEW. ANNULUS TRACING IS NOT SMOOTH WHICH CAN ALTER THE MEASUREMENT. THE PERIMETER MEASURED 69.0 MM WITH A PERIMETER DERIVED DIAMETER OF 22.0 MM SUGGESTING A 26 MM VALVE. LVOT PERIMETER AND PERIMETER DERIVED DIAMETER 62.4 MM/19.9 MM. THE AORTIC VALVE IS TRILEAFLET WITH MILD CALCIFICATION. THE SOV MEASUREMENTS DO NOT MEET IN THE CENTER WHICH CAN ALTER THE MEASUREMENT AND APPEAR TO BE LABELED INCORRECTLY. LEFT SOV DIAMETER IS BELOW THE RECOMMENDED RANGE OF 27 MM WITH THE MEAN DIAMETER OF 27.1 MM. ALL SINUS HEIGHTS AND CORONARIES ARE WITHIN THE RECOMMENDED RANGES. ANNULAR ANGULATION 48°. INTRAPROCEDURAL FLUOROSCOPIC IMAGES WERE PROVIDED FOR REVIEW. A CORONARY ANGIOGRAM WAS PERFORMED PRIOR TO THE VALVE IMPLANT ATTEMPT. FLUOROSCOPIC VALVE LOAD INSPECTION WAS PERFORMED AND CONFIRMED A GOOD LOAD. IT WAS REPORTED THAT THERE WAS DIFFICULTY CROSSING THE AORTIC ARCH AND MULTIPLE MITIGATION ATTEMPTS FAILED. IT WAS REPORTED THAT A NON-MEDTRONIC (SYMEDRIX INOWWI) GUIDEWIRE WAS USED DURING THE PROCEDURE. OF NOTE, COMPATIBLE WIRES INCLUDE MEDTRONIC CONFIDA¿ GUIDEWIRE, BOSTON SCIENTIFIC SAFARI2¿* GUIDEWIRE, AND COOK LUNDERQUIST¿* EXTRA-STIFF GUIDEWIRE. IMAGES OF THE ATTEMPTED ADVANCEMENT WERE NOT PROVIDED FOR REVIEW. HOWEVER, A FINAL ANGIOGRAM REVEALED AN AORTIC DISSECTION/PERFORATION IN THE AORTIC ARCH. IT WAS REPORTED THAT THE PATIENT CONSEQUENTLY DIED. AS STATED IN THE INSTRUCTIONS FOR USE (IFU): CAUTION: THERE WILL BE SOME RESISTANCE WHEN THE CATHETER IS ADVANCED THROUGH THE VASCULATURE. IF THERE IS A SIGNIFICANT INCREASE IN RESISTANCE, STOP ADVANCEMENT AND INVESTIGATE THE CAUSE OF THE RESISTANCE (FOR EXAMPLE, MAGNIFY THE AREA OF RESISTANCE) BEFORE PROCEEDING. DO NOT FORCE PASSAGE. FORCING PASSAGE COULD INCREASE THE RISK OF VASCULAR COMPLICATIONS (FOR EXAMPLE, VESSEL DISSECTION OR RUPTURE). CAUTION: PERSISTENT FORCE ON THE CATHETER CAN CAUSE THE CATHETER TO KINK, WHICH COULD INCREASE THE RISK OF VASCULAR COMPLICATIONS (FOR EXAMPLE, VESSEL DISSECTION OR RUPTURE). UPDATED: D4, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22202221&pc=NVZ | 3004209178-2025-10160 | 2025-04-16 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-13 04:00 | NVZ | ADVISA MRI SURESCAN | Pacing Problem | Insufficient Information | P980035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY WHICH WAS POSSIBLY RELATED TO THE IMPLANTABLE PULSE GENERATOR (IPG). THE IPG WAS NOTED TO HAVE BEEN "FIRING ALL OVER THE PLACE".Event Description: IT WAS FURTHER REPORTED THAT NO PERFORMANCE ISSUE WITH THE IPG AND THE PATIENT'S DEATH WAS NOT BELIEVED TO BE LINKED TO THE PERFORMANCE OF THE IPG. Manufacturer Narrative: CORRECTED: B2. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION, B2. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22202288&pc=MAE | 2135147-2025-03156 | 2013-09-01 04:00 | Death | ABBOTT MEDICAL | 2025-06-13 04:00 | MAE | AMPLATZER DUCT OCCLUDER II, ADDITIONAL SIZES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P020024 |  | 1 | Event Description: THE ARTICLE, "PERCUTANEOUS CLOSURE OF PATENT DUCTUS ARTERIOSUS IN PREMATURE INFANTS: A FRENCH NATIONAL SURVEY", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, MULTICENTER STUDY TO DESCRIBE THE EVOLUTION OF THE TECHNICAL ASPECTS AS WELL AS OUTCOMES OF TRANSCATHETER PATENT DUCTUS ARTERIOSUS (PDA) IN PREMATURE INFANTS. DEVICES INCLUDED IN THE STUDY WERE AMPLATZER DUCT OCCLUDER II ADDITIONAL SIZES, MEDTRONIC MICROVASCULAR PLUG, AND COILS. THE ARTICLE CONCLUDED IN THIS LARGE SERIES OF PREMATURE INFANTS UNDERGOING TRANSCATHETER PDA CLOSURE, IT WAS DEMONSTRATED THAT THIS PROCEDURE CAN BE PERFORMED SUCCESSFULLY IN THE VAST MAJORITY OF PATIENTS WITH AN ACCEPTABLE COMPLICATION RATE. FUTURE EFFORTS SHOULD FOCUS ON MINIMIZING COMPLICATIONS, PARTICULARLY DEVICE-RELATED VASCULAR STENOSES. [(B)(6)]. THE TIME FRAME OF THE STUDY WAS FROM SEPTEMBER 2013 TO JUNE 2017. A TOTAL OF 102 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 91 (89.2%) RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 39 DAYS, THE AVERAGE WEIGHT WAS 1543G, AND GENDER MAJORITY WAS NOT REPORTED. COMORBIDITIES INCLUDED PRETERM BIRTH AND PATENT DUCTUS ARTERIOSUS. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF PERCUTANEOUS CLOSURE OF PATENT DUCTUS ARTERIOSUS IN PREMATURE INFANTS: A FRENCH NATIONAL SURVEY WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PRETERM BIRTH AND PATENT DUCTUS ARTERIOSUS. SOME OF THE COMPLICATIONS REPORTED WERE OBSTRUCTION/OCCLUSION AND DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: PERCUTANEOUS CLOSURE OF PATENT DUCTUS ARTERIOSUS IN PREMATURE INFANTS: A FRENCH NATIONAL SURVEY. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: PERCUTANEOUS CLOSURE OF PATENT DUCTUS ARTERIOSUS IN PREMATURE INFANTS: A FRENCH NATIONAL SURVEY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22202419&pc=PNJ | 9612164-2025-02919 | 2025-06-11 04:00 | Death | MEDTRONIC IRELAND | 2025-06-13 04:00 | PNJ | MICRA AV | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT POST IMPLANT PROCEDURE OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE PATIENT EXPERIENCED PERICARDIAL EFFUSION AND CARDIAC TAMPONADE TWO HOURS POST IMPLANT. REANIMATION WAS TRIED ON THIS PATIENT, ALREADY IN THE POOR STATE PRIOR TO IMPLANT PROCEDURE, BUT IT WAS NOT EFFICIENT. CONSEQUENTLY, THE PATIENT DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MC1AVR1] PRODUCT ID [MC1AVR1-DELSYS] (SERIAL: [(B)(6)]). D9: NO. DEV RTN TO MFR? NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22202425&pc=MKJ | 1220908-2025-02064 | 2025-05-23 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-06-13 04:00 | MKJ | X SERIES | Failure to Capture; Pacing Problem; Unable to Obtain Readings | Insufficient Information | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO PACE A PATIENT DURING A BRADYCARDIA EVENT (AGE 82 & GENDER FEMALE), THE DEVICE PROMPTED A "LEAD FAULT" MESSAGE AND THE PACER OUTPUT WAS OUT OF SPECIFICATION. THE DEVICE FAILED TO CAPTURE THE PATIENT'S HEART RHYTHM. THE COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: ZOLL MEDICAL CORPORATION HAS RECEIVED THE PRODUCT AND WILL BE PROVIDING A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS COMPLETED. Manufacturer Narrative: THE DEVICE WAS RETURNED TO ZOLL MEDICAL CORPORATION FOR EVALUATION. THE CUSTOMER'S REPORT WAS OBSERVED DURING REVIEW OF THE DEVICE DATA LOGS. HOWEVER, THE DEVICE WAS PUT THROUGH EXTENSIVE FUNCTIONAL TESTING INCLUDING STRESS TESTING, PACING ON DEMAND AND FIXED PACING UNDER VARIOUS CONDITIONS WITHOUT DUPLICATING THE REPORT. AN INTERNAL INSPECTION FOUND NO DISCREPANCIES. THE PACE/DEFIB ENGINE BOARD WAS REPLACED AS A PRECAUTION. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. ANALYSIS FOR REPORTS OF THIS TYPE HAS NOT IDENTIFIED AN INCREASE IN TREND.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22202469&pc=OZD | 1220648-2025-29280 | 2025-05-20 04:00 | Death | ABIOMED, INC. | 2025-06-13 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemolysis; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT ENDURED ATRIAL FIBRILLATION RVR AND VENTRICULAR TACHYCARDIA DURING IMPELLA CP SUPPORT. THE PUMP WAS REPOSITIONED IN RESPONSE TO THE CONDITION. IT WAS ADDITIONALLY NOTED THAT HEMOLYSIS HAD DEVELOPED. PLANS WERE MADE TO INITIATE CONTINUOUS RENAL REPLACEMENT THERAPY. HOWEVER, CARE WAS WITHDRAWN THE FOLLOWING DAY AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: THE INVESTIGATION OF HEMOLYSIS AND VF/VT/AF/AT HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. HEMOLYSIS: HEMOLYSIS WAS NOTED WITH NO ADDITIONAL INFORMATION. THE DATA LOGS SHOW A MOTOR CURRENT WITH SPEED DEVIATION ABOUT 2 HOURS AFTER SUPPORT STARTED. THERE WAS A MOTOR CURRENT SPIKE WITH SPEED DEVIATION AND A PUMP AUTO-REDUCTION IN P-LEVEL TO P1. THE PUMP WAS TURNED TO P6 AT THAT INSTANT THERE WAS ALSO A SUCTION ALARM WHICH RESOLVED LATER. CLINICAL REPORTED HEMOLYSIS SIGNS ABOUT 3 HOURS AFTER THIS PROBABLE INGESTION EVENT. THE ROOT CAUSE OF THE HEMOLYSIS WAS MOST LIKELY BIOMATERIAL AS EVIDENCED BY THE MOTOR CURRENT SPIKE WITH SPEED DEVIATION OCCURRING BEFORE CLINICAL REPORT OF HEMOLYSIS SIGNS VT/VF/AT/AF: AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE VT/VF WAS NOT DETERMINED. THROMBUS: BASED ON THE DATA LOGS AND CLINICAL DETAILS THROMBUS WAS ADDED AS A FAILURE MODE. AS THE NECESSARY INFORMATION WAS NOT PROVIDED, THE ROOT CAUSE OF THE THROMBUS WAS NOT DETERMINED B2 DATE OF DEATH WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29280. B2 LIFE THREATENING WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29280. G1 MANUFACTURING SITE NAME, ADDRESS, COUNTRY, AND FAX NUMBER WERE ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29280. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: HEMOLYSIS - SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER. ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION. ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22203011&pc=LWS | 3008973940-2025-04219 | 2025-06-06 04:00 | Death | MEDTRONIC SINGAPORE OPERATIONS | 2025-06-13 04:00 | LWS | SPRINT QUATTRO SECURE S MRI SURESCAN | Over-Sensing; Under-Sensing | Nausea; Tachycardia; Vomiting | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED UNDERSENSING AND T-WAVE OVERSENSING (TWOS). IT WAS ALSO REPORTED THAT THE PATIENT HAD BEEN ADMITTED TO THE HOSPITAL WITH HYPERKALEMIA, POTASSIUM ISSUES, ACUTE KIDNEY INJURY, NAUSEA, VOMITING, LOW HEMOGLOBIN, AND NOSE BLEEDS. THE PATIENT PASSED AWAY 3-4 HOURS AFTER BEING DISCHARGED. IT WAS NOTED THAT THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA (VT) BUT THERAPY WAS NOT DELIVERED DUE TO THE RV LEAD SENSING ISSUES. Manufacturer Narrative: CORRECTED: B2, B3, AND B7. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE LEAD WAS NOT RETURNED FOR ANALYSIS, HOWEVER, PERFORMANCE DATA COLLECTED FROM THE DEVICE WAS RECEIVED AND ANALYZED. ANALYSIS OF THE DEVICE MEMORY INDICATED T-WAVE OVERSENSING. ANALYSIS OF THE DEVICE MEMORY INDICATED RIGHT VENTRICULAR UNDERSENSING. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22203093&pc=NGV | 2135147-2025-03169 | 2016-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-06-13 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P200049 |  | 1 | Event Description: THE ARTICLE, "IMPACT OF CHA2DS2-VASC SCORE IN THE OUTCOMES OF PATIENTS UNDERGOING A SIMPLIFIED PATHWAY FOR PERCUTANEOUS LEFT ATRIAL APPENDAGE CLOSURE", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO INVESTIGATE LONG-TERM CLINICAL OUTCOMES, EFFECTIVENESS AND SAFETY OF PERFORMING LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURES WITHOUT PRE-PROCEDURAL IMAGING ASSESSMENT IN PATIENTS WITH HIGH AND VERY HIGH THROMBOEMBOLIC RISK BASED ON CHA2DS2-VASC SCORE. DEVICES INCLUDED IN THE STUDY WERE WATCHMAN OR WATCHMAN FLX (BOSTON SCIENTIFIC), AMULET (ST. JUDE MEDICAL), ULTRASEAL (CARDIA), OMEGA (ECLIPSE MEDICAL) AND LAMBRE (LIFETECH SCIENTIFIC CORPORATION) DEVICES. THE ARTICLE CONCLUDED THAT IN A CONSECUTIVE SERIES OF PATIENTS UNDERGOING LAAC, A SIMPLIFIED APPROACH WITHOUT PRE-OPERATIVE IMAGING ASSESSMENT SHOWED COMPARABLE RESULTS AT 2-YEAR FOR THE PRIMARY COMPOSITE ENDPOINT OF ALL-CAUSE DEATH, STROKE, SYSTEMIC EMBOLIZATION, OR BLEEDING EVENTS IN PATIENTS WITH CHA2DS2-VASC SCORE = 5 AND THOSE WITH CHA2DS2-VASC SCORE < 5. [(B)(6)]. THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2016 TO JANUARY 2023. A TOTAL OF 227 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH IT WAS NOT SPECIFIED HOW MANY RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 76 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED HEART FAILURE, HYPERTENSION, DIABETES MELLITUS, CHRONIC KIDNEY DISEASE, CORONARY ARTERY DISEASE, PRIOR MYOCARDIAL INFARCTION, PRIOR PERCUTANEOUS CORONARY INTERVENTION, PRIOR STROKE, PRIOR TRANSIENT ISCHEMIC ATTACK, PERIPHERAL ARTERY DISEASE, PRIOR INTRACRANIAL BLEEDING, PRIOR GASTROINTESTINAL BLEEDING, PRIOR SYMPTOMATIC BLEEDING, RECURRING EPISTAXIS, PLASMA CELL DYSCRASIAS, POOR ADHERENCE TO THERAPY, LABILE INTERNATIONAL NORMALIZED RATION, HIGH OCCUPATIONAL RISK, AND STROKE IN ANTICOAGULATION THERAPY. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF IMPACT OF CHA2DS2-VASC SCORE IN THE OUTCOMES OF PATIENTS UNDERGOING A SIMPLIFIED PATHWAY FOR PERCUTANEOUS LEFT ATRIAL APPENDAGE CLOSURE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HEART FAILURE, HYPERTENSION, DIABETES MELLITUS, CHRONIC KIDNEY DISEASE, CORONARY ARTERY DISEASE, PRIOR MYOCARDIAL INFARCTION, PRIOR PERCUTANEOUS CORONARY INTERVENTION, PRIOR STROKE, PRIOR TRANSIENT ISCHEMIC ATTACK, PERIPHERAL ARTERY DISEASE, PRIOR INTRACRANIAL BLEEDING, PRIOR GASTROINTESTINAL BLEEDING, PRIOR SYMPTOMATIC BLEEDING, RECURRING EPISTAXIS, PLASMA CELL DYSCRASIAS, POOR ADHERENCE TO THERAPY, LABILE INTERNATIONAL NORMALIZED RATION, HIGH OCCUPATIONAL RISK, AND STROKE IN ANTICOAGULATION THERAPY. SOME OF THE COMPLICATIONS REPORTED WERE STROKE, THROMBUS, PERICARDIAL EFFUSION, TRANSIENT ISCHEMIC ATTACK, HEMORRHAGE; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: IMPACT OF CHA2DS2-VASC SCORE IN THE OUTCOMES OF PATIENTS UNDERGOING A SIMPLIFIED PATHWAY FOR PERCUTANEOUS LEFT ATRIAL APPENDAGE CLOSURE. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: IMPACT OF CHA2DS2-VASC SCORE IN THE OUTCOMES OF PATIENTS UNDERGOING A SIMPLIFIED PATHWAY FOR PERCUTANEOUS LEFT ATRIAL APPENDAGE CLOSURE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22203494&pc=LYJ | 1644487-2025-10250 | 2023-05-03 04:00 | Death | CYBERONICS - HOUSTON | 2025-06-13 04:00 | LYJ | PULSE GEN MODEL 1000 | Adverse Event Without Identified Device or Use Problem | Convulsion, Clonic | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SUDEP. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS NOT RELATED TO VNS. NO ADDITIONAL RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22203601&pc=LWQ | 2135147-2025-03181 | 2025-05-24 04:00 | Death | ABBOTT MEDICAL | 2025-06-13 04:00 | LWQ | UNKNOWN MECHANICAL HEART VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P810002 |  | 1 | Event Description: THE ARTICLE, "VALVE-IN-VALVE TRANSCATHETER AORTIC VALVE REPLACEMENT VERSUS REDO SURGICAL AORTIC VALVE REPLACEMENT IN PATIENTS WITH DEGENERATED AORTIC BIOPROSTHESES: THREE-YEAR DEATH RATES AND HAEMODYNAMIC PERFORMANCE", WAS REVIEWED. THE ARTICLE PRESENTED A CASE STUDY OF A 79-YEAR-OLD FEMALE PATIENT WITH A PREVIOUSLY IMPLANTED BIOPROSTHETIC AORTIC VALVE. IT WAS REPORTED ON AN UNKNOWN DATE, AN UNKNOWN ABBOTT/ST. JUDE MEDICAL MECHANICAL VALVE WAS CHOSEN FOR IMPLANT FOR REDO-SURGICAL AORTIC VALVE REPLACEMENT. IT WAS REPORTED TEN DAYS POST-PROCEDURE, THE PATIENT PASSED AWAY FROM SUDDEN CARDIAC DEATH IN THE SERVICE. NO OTHER PATIENT SIGNS OR SYMPTOMS WERE REPORTED. [(B)(6)]. Manufacturer Narrative: AS REPORTED IN A RESEARCH ARTICLE, VALVE-IN-VALVE TRANSCATHETER AORTIC VALVE REPLACEMENT VERSUS REDO SURGICAL AORTIC VALVE REPLACEMENT IN PATIENTS WITH DEGENERATED AORTIC BIOPROSTHESES. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. B2: DEATH DATE WAS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: VALVE-IN-VALVE TRANSCATHETER AORTIC VALVE REPLACEMENT VERSUS REDO SURGICAL AORTIC VALVE REPLACEMENT IN PATIENTS WITH DEGENERATED AORTIC BIOPROSTHESES: THREE-YEAR DEATH RATES AND HAEMODYNAMIC PERFORMANCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22203612&pc=NPW | 2015691-2025-04834 | 2025-04-27 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-13 04:00 | NPW | EVOQUE TRICUSPID VALVE REPLACEMENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Heart Block | P230013 |  | 1 | Event Description: AS REPORTED BY THE EDWARDS LIFESCIENCES AFFILIATE IN GREECE, EDWARDS RECEIVED NOTIFICATION OF A 44M EVOQUE PROCEDURE IN TRICUSPID POSITION WHERE THE PATIENT WAS ON BETA BLOCKER AND DIGOXIN THAT WAS HELD FOR 24 HOURS POST PROCEDURAL AND REINSTATED AFTERWARDS. ADDITIONALLY, THE PATIENT WAS ON ATRIAL FIBRILLATION ON TELEMETRY FOR 48 HOURS POST PROCEDURE. THE PATIENT WAS DISCHARGED ON POST OPERATIVE DAY 2 AND SHORTLY AFTER ARRIVING HOME THEY EXPERIENCED DIZZINESS AND DESPITE EXTENSIVE CPR, THE PATIENT PASSED AWAY. AS PER MEDICAL OPINION, THE MOST LIKELY CASE OF DEATH WOULD BE COMPLETE HEART BLOCK ARRECT, FOLLOWED BY VENTRICULAR TACHYCARDIA/ VENTRICULAR FIBRILLATION (PATIENT HAD ELEVATED TROPONIN), PULMONARY EMBOLISM OR METALLIC VALVE THROMBOSIS. Manufacturer Narrative: THE FOLLOWING SECTIONS WERE UPDATED/CORRECTED/ADDED: B4, G3, G6, H1, H2, H6 AND H11. THE COMPLAINT FOR VALVE INTERACTS WITH CONDUCTION SYSTEM WAS CONFIRMED WITH OTHER EMPIRICAL EVIDENCE VIA INFORMATION RECEIVED BY EDWARDS. ULTIMATELY, THE PATIENT'S EARLY DISCHARGE PREVENTED THE MEDICAL TEAM FROM PROVIDING POTENTIAL TREATMENT SUCH AS A PACEMAKER IMPLANTATION. PER THE INSTRUCTIONS FOR USE (IFU), CONDUCTION SYSTEM DISTURBANCE AND/OR HEART BLOCK WHICH MAY REQUIRE TREATMENT (INCLUDING PERMANENT PACEMAKER) ARE POTENTIAL ADVERSE EVENTS. HEART BLOCK AND OTHER CONDUCTION DISTURBANCES ARE COMMON IN PATIENTS WITH UNDERLYING CARDIOVASCULAR DISEASE AND CAN BE EXACERBATED OR AGGRAVATED WITH STANDARD INTRA-OPERATIVE MEDICATIONS OR ANESTHESIA. SUCH EVENTS ARE RELATED TO THE PATIENT'S UNDERLYING CONDITION. OTHER MECHANISMS FOR CONDUCTION ABNORMALITIES COULD BE RELATED TO CORONARY OBSTRUCTION DUE TO AIR EMBOLISM, CORONARY SPASM AND/OR LOCAL EDEMA AFFECTING THE AV NODE. CONDUCTION DISTURBANCES HAVE ALSO BEEN DOCUMENTED DURING TRANSCATHETER CARDIAC PROCEDURES AS A RESULT OF DIRECT INTERACTION WITH CARDIAC ANATOMY, ESPECIALLY IN PATIENTS WITH UNDERLYING CONDUCTION ABNORMALITIES. OTHER POTENTIAL CAUSES INCLUDE TRAUMA BY A GUIDEWIRE OR DELIVERY SYSTEM. ADDITIONALLY, CARDIAC CONDUCTION SYSTEM COMPLICATIONS ARE KNOWN RISKS WITH TRICUSPID VALVE INTERVENTIONS DUE TO THE PROXIMITY OF THE ATRIOVENTRICULAR NODE (AV NODE) TO THE SEPTAL LEAFLET OF THE TRICUSPID VALVE. THESE CONDUCTION DISTURBANCES CAN LEAD TO POST-OPERATIVE HEART BLOCK REQUIRING PACEMAKER IMPLANTATION. SINCE NO MANUFACTURING NON-CONFORMANCES WERE CONFIRMED AND NO IFU/TRAINING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE/PREVENTATIVE ACTIONS ARE REQUIRED. Manufacturer Narrative: THE MANUFACTURER'S INVESTIGATION IS ONGOING AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22203665&pc=LWR | 2135147-2025-03183 | 2025-05-24 04:00 | Death | ABBOTT MEDICAL | 2025-06-13 04:00 | LWR | TRIFECTA¿ | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P100029 |  | 1 | Event Description: THE ARTICLE, "VALVE-IN-VALVE TRANSCATHETER AORTIC VALVE REPLACEMENT VERSUS REDO SURGICAL AORTIC VALVE REPLACEMENT IN PATIENTS WITH DEGENERATED AORTIC BIOPROSTHESES: THREE-YEAR DEATH RATES AND HAEMODYNAMIC PERFORMANCE", WAS REVIEWED. THE ARTICLE PRESENTED A CASE STUDY OF A 72-YEAR-OLD FEMALE PATIENT WITH A PREVIOUSLY IMPLANTED BIOPROSTHETIC AORTIC VALVE. IT WAS REPORTED ON AN UNKNOWN DATE, A 21MM TRIFECTA VALVE WAS CHOSEN FOR IMPLANT FOR REDO-SURGICAL AORTIC VALVE REPLACEMENT. DURING THE PROCEDURE, IT WAS REPORTED THE PATIENT UNEXPECTEDLY PASSED AWAY. NO OTHER PATIENT SIGNS OR SYMPTOMS WERE REPORTED. [(B)(6)]. Manufacturer Narrative: AS REPORTED IN A RESEARCH ARTICLE, VALVE-IN-VALVE TRANSCATHETER AORTIC VALVE REPLACEMENT VERSUS REDO SURGICAL AORTIC VALVE REPLACEMENT IN PATIENTS WITH DEGENERATED AORTIC BIOPROSTHESES: THREE-YEAR DEATH RATES AND HAEMODYNAMIC PERFORMANCE. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. B2: DATE OF DEATH WAS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: VALVE-IN-VALVE TRANSCATHETER AORTIC VALVE REPLACEMENT VERSUS REDO SURGICAL AORTIC VALVE REPLACEMENT IN PATIENTS WITH DEGENERATED AORTIC BIOPROSTHESES: THREE-YEAR DEATH RATES AND HAEMODYNAMIC PERFORMANCE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22204007&pc=MCX | 3004742232-2025-00169 | 2025-05-29 04:00 | Death | CARDIOVASCULAR SYSTEMS INCORPORATED | 2025-06-13 04:00 | MCX | DIAMONDBACK CORONARY | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Arrest; Obstruction/Occlusion | P130005 |  | 1 | Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A PROXIMAL AND MID LEFT ANTERIOR DESCENDING ARTERY (LAD). EIGHT LOW SPEED TREATMENTS WERE PERFORMED. AT THIS POINT, THE PATIENT HAD SLOW FLOW AND SUDDENLY HAD BRADYCARDIA. THE PATIENT WENT INTO CARDIAC ARREST. THE PHYSICIAN ATTEMPTED TO REVIVE THE PATIENT, BUT THE PATIENT EXPIRED. IN THE PHYSICIAN'S OPINION, ORBITAL ATHERECTOMY DID NOT CAUSE OR CONTRIBUTE TO PATIENT ADVERSE EVENTS. HOWEVER, THE PHYSICIAN COULD NOT DETERMINE WHAT LED TO THE PATIENT ADVERSE EVENTS. TRIGGER POINT INJECTION AND CARDIOPULMONARY RESUSCITATION WERE PERFORMED TO RESUSCITATE THE PATIENT. THE SLOW FLOW WAS TRANSIENT. IT WAS UNKNOWN WHAT EVENTS LED TO THE SLOW FLOW.Event Description: IT WAS REPORTED A DIAMONDBACK 360 CORONARY ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A PROXIMAL AND MID LEFT ANTERIOR DESCENDING ARTERY (LAD). EIGHT LOW SPEED TREATMENTS WERE PERFORMED. AT THIS POINT, THE PATIENT HAD SLOW FLOW AND SUDDENLY HAD BRADYCARDIA. THE PATIENT WENT INTO CARDIAC ARREST. THE PHYSICIAN ATTEMPTED TO REVIVE THE PATIENT, BUT THE PATIENT EXPIRED. IN THE PHYSICIAN'S OPINION, ORBITAL ATHERECTOMY DID NOT CAUSE OR CONTRIBUTE TO PATIENT ADVERSE EVENTS. HOWEVER, THE PHYSICIAN COULD NOT DETERMINE WHAT LED TO THE PATIENT ADVERSE EVENTS. TRIGGER POINT INJECTION AND CARDIOPULMONARY RESUSCITATION WERE PERFORMED TO RESUSCITATE THE PATIENT. THE SLOW FLOW WAS TRANSIENT. IT WAS UNKNOWN WHAT EVENTS LED TO THE SLOW FLOW. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. BASED ON THE INFORMATION RECEIVED, THE INVESTIGATION DETERMINED THAT THE REPORTED PATIENT HARMS INCLUDING CARDIAC ARREST, ARRHYTHMIA, OBSTRUCTION, UNEXPECTED MEDICAL INTERVENTION AND DEATH APPEAR TO BE RELATED TO OPERATIONAL CONTEXT. PER MEDICAL REVIEW, THERE IS VERY LIMITED INFORMATION PROVIDED HOWEVER, IT CAN BE STATED THAT THE OCCURRENCE OF SLOW FLOW THAT MOST LIKELY LED TO CARDIAC ARREST AND TO AN OUTCOME OF DEATH WAS RELATED TO THE PROCEDURE AND THAT THE OAD INDIRECTLY CONTRIBUTED/CAUSED IN THAT IT WAS USED IN A SEVERELY CALCIFIED LESION. BASED ON THE RESULTS OF THE COMPLAINT INVESTIGATION, THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. H6: CODES 4118, 3233, AND 11 NO LONGER APPLY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22204024&pc=NKM | 2135147-2025-03193 | 2016-01-01 05:00 | Death | ABBOTT MEDICAL | 2025-06-13 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "INCIDENCE, CORRELATES, AND OUTCOMES OF ECHOCARDIOGRAPHIC SMOKE-LIKE EFFECT AFTER TRANSCATHETER EDGE-TO-EDGE REPAIR OF MITRAL REGURGITATION WITH THE MITRACLIP DEVICE" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE MULTI CENTER STUDY, TO EVALUATE THE SMOKE-LIKE EFFECT (SE), THE SPONTANEOUS ECHOCARDIOGRAPHIC CONTRAST IN THE LEFT ATRIUM AT TRANSESOPHAGEAL ECHOCARDIOGRAPHY, HAS BEEN ANECDOTALLY REPORTED AFTER TRANSCATHETER EDGE-TO-EDGE REPAIR (TEER) FOR MITRAL REGURGITATION (MR), BUT UNCERTAINTY PERSISTS ON ITS IMPACT. THUS, THE AUTHORS AIMED AT APPRAISING THE INCIDENCE, CORRELATES, AND OUTCOMES OF SE AFTER TEER. DEVICES MENTIONED INCLUDE MITRACLIP. THE ARTICLE CONCLUDED THAT SE AFTER TEER IS NOT UNCOMMON AND IS TYPICALLY ASSOCIATED WITH A SIGNIFICANTLY WORSE CLINICAL PROFILE, PARTICULARLY PRIOR ATRIAL FIBRILLATION. IRRESPECTIVELY, SE IS NOT ASSOCIATED WITH ADVERSE OUTCOMES IN THE SHORT- OR LONG-TERM. ACCORDINGLY, IT SHOULD NOT BE CONSIDERED PER SE AS AN INDICATION FOR MORE AGGRESSIVE MEDICAL MANAGEMENT, WITH ANTITHROMBOTIC REGIMENS BEING INSTEAD INFORMED BY OTHER MORE ESTABLISHED INDICATIONS. [(B)(6)]. THE TIME FRAME OF THE STUDY WAS 2016 TO 2023. A TOTAL OF 2228 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH 2228 RECEIVED AN ABBOTT DEVICE. AS THIS EVENT IS FROM A LITERATURE REVIEW, THERE IS NO RELEVANT PATIENT INFORMATION (DATE OF BIRTH, AGE, GENDER, WEIGHT, AND MEDICAL HISTORY) TO REPORT. COMORBIDITIES INCLUDED HYPERTENSION, DIABETES, PRIOR MYOCARDIAL INFARCTION, PRIOR CORONARY REVASCULARIZATION, PRIOR MITRAL VALVE SURGERY, PRIOR CEREBROVASCULAR EVENT. (B)(6), UNK MITRACLIP PERI- AND POST-PROCEDURAL COMPLICATIONS INCLUDED DEATH, ATRIAL FIBRILLATION, TAMPONADE, SURGICAL INTERVENTION, BLEEDING, MEDICAL INTERVENTION, TRANSIENT ISCHEMIC ATTACK, MYOCARDIAL INFARCTION, PROLONGED HOSPITALIZATION, HEART FAILURE, STROKE, RECURRENT MR. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP DEVICE WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING HYPERTENSION, DIABETES, PRIOR MYOCARDIAL INFARCTION, PRIOR CORONARY REVASCULARIZATION, PRIOR MITRAL VALVE SURGERY, PRIOR CEREBROVASCULAR EVENT. COMPLICATIONS REPORTED INCLUDED DEATH, ATRIAL FIBRILLATION, TAMPONADE, SURGICAL INTERVENTION, BLEEDING, MEDICAL INTERVENTION, TRANSIENT ISCHEMIC ATTACK, MYOCARDIAL INFARCTION, PROLONGED HOSPITALIZATION, HEART FAILURE, STROKE, RECURRENT MR; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. AA MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. Manufacturer Narrative: IT IS UNKNOWN IF THE DEVICE IS RETURNING FOR ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22204138&pc=LYJ | 1644487-2025-10252 | 2025-05-20 04:00 | Death | HOUSTON | 2025-06-13 04:00 | LYJ | PULSE GEN MODEL 104 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY FOR AN UNKNOWN REASON. REVIEW OF THE PATIENT'S OBITUARY REVEALED THE PATIENT PASSED AWAY ON (B)(6) 2025. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22204278&pc=DSQ | 3007042319-2025-01272 | 2025-05-07 04:00 | Death | HEARTWARE, INC. | 2025-06-13 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A VENTRICULAR ASSIST DEVICE (VAD) PATIENT EXPERIENCED CARDIAC ARREST. THE PATIENT SUBSEQUENTLY PASSED AWAY. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED AFTER EXPERIENCING CARDIAC ARREST. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, CARDIOPULMONARY ARREST AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF CARDIOPULMONARY ARREST. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22204293&pc=NPT | 2015691-2025-04840 | 2025-02-16 05:00 | Death | EDWARDS LIFESCIENCES LLC | 2025-06-13 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P140031 |  | 1 | Event Description: THIS COMPLAINT WAS OPENED FROM THV/TVT REGISTRY ALTERNATIVE SUMMARY REPORTING ADVERSE EVENT SUBMISSION FOR EVENTS RECEIVED BY EW DURING QUARTER 2 OF 2025 FOR THE SAPIEN 3, SAPIEN 3 ULTRA AND SAPIEN 3 ULTRA RESILIA VALVE. MULTIPLE EVENTS WERE IDENTIFIED TO HAVE OCCURRED MORE THAN 1 YEAR POST-PROCEDURE OR WHERE THE THV WAS IMPLANTED THE TRICUSPID POSITION. THIS REPORT IS FOR AN VALVE-RELATED READMISSION DEATH EVENT 404 DAYS POST IMPLANT OF A 23MM EDWARDS SAPIEN 3 ULTRA RESILIA VALVE IN THE AORTIC POSITION FOR AN 81-YEAR-OLD MALE. Manufacturer Narrative: THE EVENT DOCUMENTED IN THIS COMPLAINT WAS RECEIVED BY EW FROM THE TRANSCATHETER VALVE THERAPY (TVT) REGISTRY DURING QUARTER 2 OF 2025 FOR THE SAPIEN XT, SAPIEN 3, SAPIEN 3 ULTRA AND SAPIEN 3 ULTRA RESILIA VALVE. THIS EVENT WAS IDENTIFIED TO HAVE OCCURRED MORE THAN 1-YEAR POST-INDEX PROCEDURE. THEREFORE, THE EVENT DOES NOT MEET FDA'S SUMMARY REPORTING EXEMPTION (E2016006) CRITERIA FOR TVT REGISTRY ADVERSE EVENTS. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR THE 23MM EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IS (B)(4). PER THE INSTRUCTIONS FOR USE (IFU), NON-EMERGENT REOPERATION, EMERGENCY CARDIAC SURGERY, REOPERATION, DEVICE MIGRATION OR MALPOSITION REQUIRING INTERVENTION AND DEVICE EXPLANT ARE LISTED AS POTENTIAL RISKS ASSOCIATED WITH THE DEVICE AND TRANSCATHETER VALVE REPLACEMENT PROCEDURE. THE IFU CAUTIONS THAT LONG-TERM DURABILITY HAS NOT BEEN ESTABLISHED FOR THE VALVE. REGULAR MEDICAL FOLLOW-UP IS ADVISED TO EVALUATE VALVE PERFORMANCE. ACCELERATED DETERIORATION OF THE VALVE DUE TO CALCIFIC DEGENERATION MAY OCCUR IN CHILDREN, ADOLESCENTS, OR YOUNG ADULTS AND IN PATIENTS WITH AN ALTERED CALCIUM METABOLISM. VALVE RECIPIENTS SHOULD BE MAINTAINED ON ANTICOAGULANT/ANTIPLATELET THERAPY, EXCEPT WHEN CONTRAINDICATED, AS DETERMINED BY THEIR PHYSICIAN. THIS DEVICE HAS NOT BEEN TESTED FOR USE WITHOUT ANTICOAGULATION. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. IT IS POSSIBLE THAT THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, HOWEVER GIVEN THE LIMITED INFORMATION AVAILABLE FURTHER ASSESSMENT OF THIS ADVERSE EVENT IS NOT POSSIBLE AT THIS TIME. VALVE RELATED READMISSIONS MAY BE DUE TO VALVE DYSFUNCTION (E.G., SEVERE, OR CLINICALLY SIGNIFICANT PARAVALVULAR LEAK, CENTRAL LEAK, SIGNIFICANT MALPOSITION, EARLY OR LATE ENDOCARDITIS, THROMBOSIS, STRUCTURAL VALVE DETERIORATION OR DEGENERATION) AND/OR OTHER REASONS UNRELATED TO THE EDWARDS DEVICES (E.G., TREATMENT OF OTHER PATIENT COMORBIDITIES). DURING THE MANUFACTURING PROCESS, ALL SAPIEN VALVES (ALL MODELS) ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. IN THIS CASE, READMISSION OCCURRED 404 DAYS POST-IMPLANT WITH NO INDICATION OF INTERVENTION. IT IS UNKNOWN IF THE EVENT WAS RELATED TO THE EDWARD'S DEVICES OR PATIENT CO-MORBIDITIES. NO ADDITIONAL INFORMATION WAS PROVIDED, AND NO ADDITIONAL INVESTIGATION IS POSSIBLE. DUE TO THE LIMITED INFORMATION AVAILABLE, THE DEFINITIVE CAUSE OR CONTRIBUTING FACTORS FOR THIS EVENT CANNOT BE DETERMINED. THE AVAILABLE INFORMATION DOES NOT REASONABLY SUGGEST THERE WAS A MALFUNCTION OF THE EDWARDS DEVICE OR THAT THE USE OR MISUSE OF THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT. THERE IS NO ALLEGATION OF DEFICIENCIES RELATED TO THE IDENTITY, QUALITY, DURABILITY, RELIABILITY, SAFETY, LABELING EFFECTIVENESS, OR PERFORMANCE OF THIS EDWARD DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: A SUPPLEMENTAL MANUFACTURER REPORT IS BEING SUBMITTED, FOLLOWING THE DISCOVERY OF AN ERROR DURING ADMINISTRATIVE REVIEW. B4, G3, AND G6 HAVE BEEN UPDATED. H6: HEALTH IMPACT HAS BEEN CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22204356&pc=NPT | 2015691-2025-04841 | 2025-05-21 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-13 04:00 | NPT | COMMANDER DELIVERY SYSTEM | Difficult to Remove; Device Slipped; Material Split, Cut or Torn | Insufficient Information | P140031 |  | 1 | Event Description: AS REPORTED BY AN EDWARDS LIFESCIENCES FIELD CLINICAL SPECIALIST, A 26MM SAPIEN 3 ULTRA RESILIA VALVE WAS TO BE IMPLANTED IN THE AORTIC POSITION VIA LEFT AXILLARY APPROACH UTILIZING A COMMANDER DELIVERY SYSTEM AND 14 ESHEATH OPTIMA. VALVE ALIGNMENT WAS PERFORMED IN A STRAIGHT SECTION OF THE ANATOMY. THERE WAS DIFFICULTY DURING FINE ADJUST OF THE DELIVERY SYSTEM. AFTER VALVE ALIGNMENT, THERE WAS SIGNIFICANT TENSION ON THE SYSTEM, MAKING VALVE POSITIONING DIFFICULT. UPON ATTEMPTING TO INFLATE THE BALLOON, THERE WAS NO INDICATION OF BALLOON INFLATION. WHEN PULLING BACK, THE ENDOFLATOR FILLED WITH BLOOD, INDICATING A BREAK IN THE LINE/POTENTIAL IC BOND FAILURE. THE DECISION WAS MADE TO ATTEMPT TRANSFEMORAL ACCESS, AND A NEW 26MM SAPIEN 3 ULTRA RESILIA VALVE WAS SUCCESSFULLY DEPLOYED. WHEN ATTEMPTING TO WITHDRAW THE SHEATH AND DELIVERY SYSTEM FROM THE LEFT AXILLARY AS A UNIT, THE VALVE AND DELIVERY SYSTEM APPEARED TO GET STUCK IN THE AXILLARY ARTERY. DELIVERY SYSTEM AND VALVE WERE UNABLE TO BE WITHDRAWN INTO THE SHEATH DUE TO ANGULATION AND BLOOD FILLING THE DELIVERY SYSTEM BALLOON. A CUTDOWN WAS PERFORMED, BUT THE OPERATOR WAS UNSUCCESSFUL IN ACCESSING THE VALVE, SO THE PATIENT WAS CONVERTED TO FULL STERNOTOMY. THIS WAS UNSUCCESSFUL AND THE PATIENT EXPIRED. UPON REMOVAL OF THE DEVICES, THE COMMANDER DELIVERY SYSTEM BALLOON APPEARED TORN. ADDITIONALLY, AFTER THE SHEATH WAS REMOVED, IT WAS FLUSHED BY THE FIELD CLINICAL SPECIALIST, AND A PINPOINT HOLE WAS FOUND IN THE SHEATH APPROXIMATELY 10CM FROM THE SIDE PORT.Event Description: ADDITIONAL CLARIFICATION WAS PROVIDED ABOUT FINE ADJUSTMENT STEP OF THE PROCEDURE. THERE WAS DIFFICULTY DURING FINE ADJUSTMENT OF THE DELIVERY SYSTEM. AFTER VALVE ALIGNMENT, THERE WAS SIGNIFICANT TENSION ON THE SYSTEM, MAKING VALVE POSITIONING DIFFICULT. WHEN THE PUSHER WAS PULLED BACK, IT SEEMED TO CHANGE THE ORIENTATION AND LOOKED LIKE IT WAS A 90-DEGREE ANGLE TO THE REST OF THE DELIVERY SYSTEM. WHEN THIS HAPPENED, THE VALVE WOULD THEN MOVE UP AND OUT OF POSITION OFF THE BALLOON. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR ADDITIONAL INFORMATION RECEIVED FROM THE SITE. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: ADDITIONAL INFORMATION ADDED TO B5, ADDITIONAL CODE ADDED TO H6 DEVICE CODE(S). Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR CORRECTION AND ADDITIONAL INFORMATION BASED ON THE ENGINEER EVALUATION. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: CORRECTED D9, ADDITIONAL CODE ADDED TO H6 COMPONENT CODES, ADDITIONAL CODES ADDED TO H6 TYPE OF INVESTIGATION, CORRECTED H6 INVESTIGATION FINDINGS, AND CORRECTED H6 INVESTIGATION CONCLUSIONS. THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED TO EDWARDS LIFESCIENCES FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. DEVICE-RELATED IMAGERY WAS PROVIDED, AND THE FOLLOWING WAS OBSERVED: CRIMPED VALVE ON COMMANDER INFLATION BALLOON WITH BALLOON TORN AT I/C BOND. PROXIMAL BALLOON WINGS WERE FLARED. PATIENT TISSUE WAS PARTIALLY COVERING THE CRIMPED VALVE, AND SUTURES INTERTWINED AROUND THE DELIVERY SYSTEM. THE REPORT OF THE VALVE MOVING ON THE BALLOON WORKING LENGTH DURING POSITIONING WAS UNABLE TO BE CONFIRMED. AVAILABLE INFORMATION SUGGESTS PROCEDURAL FACTORS (BUILT-UP TENSION) LIKELY CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT THERE WAS SIGNIFICANT TENSION ON THE SYSTEM AFTER VALVE ALIGNMENT, MAKING VALVE POSITIONING DIFFICULT. IN THIS CASE, THE RELEASE OF BUILT-UP TENSION WITHIN THE DELIVERY SYSTEM DURING THE PROCEDURE MAY HAVE RESULTED IN THE REPORTED VALVE MOVEMENT ON BALLOON DURING FLEX TIP RETRACTION. THE REPORT OF THE TORN DELIVERY SYSTEM BALLOON WAS CONFIRMED BASED ON PROVIDED IMAGERY. AVAILABLE INFORMATION SUGGESTS PROCEDURAL FACTORS (HIGH ALIGNMENT FORCES) LIKELY CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT THERE WAS NO INDICATION OF BALLOON INFLATION UPON ATTEMPTING TO INFLATE THE BALLOON. WHEN PULLING BACK, THE ENDOFLATOR FILLED WITH BLOOD, INDICATING A BREAK IN THE LINE/POTENTIAL IC BOND FAILURE. IN THIS CASE, HIGH FORCES ON THE SYSTEM DURING VALVE ALIGNMENT MAY RESULT IN CRIMP BALLOON TEARING PRIOR TO THV DEPLOYMENT. THE REPORTED DIFFICULTY EXPERIENCED REMOVING THE DELIVERY SYSTEM AND VALVE THROUGH THE VASCULATURE AND SUBSEQUENT PATIENT DEATH WAS CONFIRMED BASED ON PROVIDED IMAGERY. AVAILABLE INFORMATION SUGGESTS PATIENT FACTORS (TORTUOSITY) AND/OR PROCEDURAL FACTORS (NON-COAXIAL ALIGNMENT) LIKELY CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT THE DELIVERY SYSTEM AND VALVE WERE UNABLE TO BE WITHDRAWN INTO THE SHEATH DUE TO ANGULATION AND BLOOD FILLING THE DELIVERY SYSTEM BALLOON. IN THIS CASE, TORTUOUS PATIENT ANATOMY CAN CONTRIBUTE TO THE NON-COAXIAL WITHDRAWAL OF THE DELIVERY SYSTEM. IT IS ALSO POSSIBLE THAT NON-COAXIAL ALIGNMENT MAY HAVE RESULTED IN THE CRIMPED THV INTERACTING WITH PATIENT VASCULATURE OR THE SHEATH, CAUSING THE WITHDRAWAL DIFFICULTY, AND RESULTING IN POTENTIAL INJURIES AS REPORTED. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED ADDITIONAL INFORMATION AFTER ADMINISTRATIVE REVIEW. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: A2. Manufacturer Narrative: INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22204481&pc=DPS | 1218950-2025-000355 | 2025-04-08 04:00 | Death | PHILIPS MEDICAL SYSTEMS | 2025-06-13 04:00 | DPS | PAGEWRITER | Incorrect, Inadequate or Imprecise Result or Readings | Insufficient Information | K113144 |  | 1 | Event Description: IT WAS REPORTED THERE WAS A DISCREPANCY IN RHYTHM INTERPRETATION BETWEEN THE CARDIOGRAPH AND THE MONITOR ECG. INFORMATION RECEIVED FROM THE CUSTOMER INDICATED THE DEVICE HAD 'BEEN A CONTRIBUTING FACTOR TO PATIENT DEATHS AND PATIENT INCIDENTS'. THE CUSTOMER CLARIFIED THE ALLEGATION TO MEAN THAT THE CRITICAL VALUE STATEMENT (TEXT) IS NOT APPEARING IN THE CUSTOMER THIRD-PARTY CERNER ELECTRONIC MEDICAL RECORD (EMR) BUT IS SHOWING ON THE MX MONITOR AND PAGEWRITER TC CARTS. Manufacturer Narrative: THE ISSUE WAS INVESTIGATED BY A PHILIPS SENIOR PRODUCT SUPPORT ENGINEER, A PHILIPS INTEGRATION ENGINEER (IE) AND A PHILIPS SOLUTIONS ARCHITECT. THE IE PERFORMED AN ANALYSIS OF THE CONTENTS PHILIPSXML BEING SENT BY THE PIC IX C.03 SYSTEM AND WORKED WITH THE CUSTOMER¿S INTEGRATION ANALYST TO REVIEW WHAT WAS SEEN IN THE DATAMED OUTPUT. THE MONITOR DATA PATHWAY IS SUMMARIZED AS FOLLOWS: THE PHILIPS MX SERIES MONITOR SENDS THE CRITICAL VALUE STATEMENTS TO A PHILIPS PATIENT INFORMATION CENTER IX (PIC IX CENTRAL MONITORING) C.03 SYSTEM. THE PIC IX C.03 INCLUDES THESE CRITICAL VALUE STATEMENTS IN THE PHILIPS XML OF THE 12-LEAD ECG RESULT TO A THIRD-PARTY MIDDLEWARE PRODUCT 'DATAMEDFT', NOT SOLD OR DISTRIBUTED BY PHILIPS TO THIS CUSTOMER. THE CUSTOMERS PAGEWRITERTC CARTS ARE LICENSED AND CONFIGURED TO SEND DICOM FORMATTED ECG STUDIES DIRECTLY TO THE CUSTOMER'S CERNER EMR. THE INVESTIGATION DETERMINED THAT THE PHILIPS DICOM IS DESIGNED TO AND FUNCTIONS TO MEET THE UNIVERSAL DICOM STANDARDS. THIS STANDARD DOES NOT INCLUDE THE TRANSFER OF CRITICAL VALUE STATEMENTS. IN THIS CASE, THE CUSTOMER MISTAKENLY BELIEVED THAT THE CRITICAL VALUES WOULD TRANSFER TO THE CERNER EMR AND, BECAUSE OF THIS, THE USER WAS NOT ALERTED TO THE PATIENT CONDITION IN THE EXPECTED WAY. PHILIPS CARDIOGRAPH FOLLOWS THE UNIVERSAL DICOM STANDARDS AND THE PHILIPS DEVICE IS CONFORMING TO SPECIFICATIONS. AS A RESULT, THE CERNER ELECTRONIC MEDICAL RECORD DOES NOT DISPLAY THE CRITICAL VALUE STATEMENTS WHEN PAGEWRITERTC CARTS ARE LICENSED AND CONFIGURED TO SEND DICOM FORMATTED ECG STUDIES. IN SUMMARY, CERNER SOFTWARE IS NOT ABLE TO DISPLAY/PRINT CRITICAL VALUE. CERNER DOESN'T PUBLISH ITS DICOM CONFORMANCE STATEMENT FOR PHILIPS TO DO THE SYSTEM INTEGRATION. CERNER DOES NOT OPEN A CHANNEL FOR PHILIPS TO DO THE SYSTEM INTEGRATION. DICOM STANDARD DOESN'T DEFINE/RESERVE ¿DICOM TAG¿ FOR CRITICAL VALUE. ANY FURTHER CHANGES TO DICOM CAN ONLY BE PERFORMED IN COOPERATION WITH CERNER. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. Manufacturer Narrative: THE CUSTOMER CLARIFIED THE ALLEGATION TO MEAN THAT THE CRITICAL VALUE STATEMENT (TEXT) IS NOT APPEARING IN THE CUSTOMER THIRD-PARTY CERNER ELECTRONIC MEDICAL RECORD (EMR) BUT IS SHOWING ON THE MX MONITOR AND PAGEWRITER TC CARTS. THE ISSUE WAS INVESTIGATED BY A PHILIPS SENIOR PRODUCT SUPPORT ENGINEER, A PHILIPS INTEGRATION ENGINEER (IE) AND A PHILIPS SOLUTIONS ARCHITECT. THE IE PERFORMED AN ANALYSIS OF THE CONTENTS PHILIPSXML BEING SENT BY THE PIC IX C.03 SYSTEM AND WORKED WITH THE CUSTOMER¿S INTEGRATION ANALYST TO REVIEW WHAT WAS SEEN IN THE DATAMED OUTPUT. THE MONITOR DATA PATHWAY IS SUMMARIZED AS FOLLOWS: THE PHILIPS MX SERIES MONITOR SENDS THE CRITICAL VALUE STATEMENTS TO A PHILIPS PATIENT INFORMATION CENTER IX (PIC IX CENTRAL MONITORING) C.03 SYSTEM. THE PIC IX C.03 INCLUDES THESE CRITICAL VALUE STATEMENTS IN THE PHILIPS XML OF THE 12-LEAD ECG RESULT TO A THIRD-PARTY MIDDLEWARE PRODUCT 'DATAMEDFT', NOT SOLD OR DISTRIBUTED BY PHILIPS TO THIS CUSTOMER. THE CUSTOMERS PAGEWRITERTC CARTS ARE LICENSED AND CONFIGURED TO SEND DICOM FORMATTED ECG STUDIES DIRECTLY TO THE CUSTOMER'S CERNER EMR. THE INVESTIGATION DETERMINED THAT THE PHILIPS DICOM IS DESIGNED TO AND FUNCTIONS TO MEET THE UNIVERSAL DICOM STANDARDS. THIS STANDARD DOES NOT INCLUDE THE TRANSFER OF CRITICAL VALUE STATEMENTS. IN THIS CASE, THE CUSTOMER MISTAKENLY BELIEVED THAT THE CRITICAL VALUES WOULD TRANSFER TO THE CERNER EMR AND, BECAUSE OF THIS, THE USER WAS NOT ALERTED TO THE PATIENT CONDITION IN THE EXPECTED WAY. GOOD FAITH EFFORTS ARE BEING MADE TO OBTAIN SERIAL NUMBER UDI INFORMATION. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22204551&pc=NGV | 2135147-2025-03204 | 2025-05-21 04:00 | Death | ABBOTT MEDICAL | 2025-06-13 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Tachycardia; Ventricular Fibrillation; Pericardial Effusion | P200049 |  | 1 | Event Description: IT WAS REPORTED ON (B)(6) 2025 A 22MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS SELECTED FOR IMPLANT USING A NON-ABBOTT DELIVERY SYSTEM. THE PATIENT WAS IN SINUS RHYTHM. TRANSSEPTAL ACCESS WAS INFERIOR AND MID. THE PATIENT'S ACTIVATED CLOTTING TIME (ACT) LEVEL WAS BETWEEN 250 AND 350 SECONDS. DURING THE PROCEDURE THE OCCLUDER WAS PARTIALLY RE-CAPTURED ONCE. CLOSE CRITERIA WERE MET AND THE OCCLUDER HAD A TIRE-SHAPED COMPRESSION. THE OCCLUDER WAS IMPLANTED. APPROXIMATELY AN HOUR AND A HALF AFTER THE PROCEDURE THE PATIENT WAS FOUND UNRESPONSIVE IN RECOVERY. CARDIOPULMONARY RESUSCITATION (CPR) WAS STARTED AND THE PATIENT WENT INTO VENTRICULAR FIBRILLATION. AN X-RAY WAS PERFORMED AND THE OCCLUDER WAS IN PLACE. ADDITIONAL CPR WAS PERFORMED. A TRANSTHORACIC ECHOCARDIOGRAM (TTE) SHOWED A SMALL, CIRCUMFERENTIAL PERICARDIAL EFFUSION IN THE APEX OF THE HEART. PERICARDIOCENTESIS WAS PERFORMED. THE PATIENT WAS SHOCKED SEVERAL TIMES AND PRESENTED WITH BOUTS OF VENTRICULAR FIBRILLATION, VENTRICULAR TACHYCARDIA, AND ASYSTOLE. 1500 ML OF DARK LIQUID WAS DRAINED. ANOTHER PERICARDIAL ACCESS WAS DONE AND 200 ML OF LIQUID WAS DRAINED. THE PATIENT WAS SHOCKED SEVERAL MORE TIMES. AT THIS POINT CPR HAD BEEN PERFORMED FOR AN HOUR AND TWENTY MINUTES. NO EKG ACTIVITY WAS NOTED AND THE PATIENT WAS PRONOUNCED DEAD.Event Description: IT WAS REPORTED ON (B)(6) 2025 A 22MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS SELECTED FOR IMPLANT USING A NON-ABBOTT DELIVERY SYSTEM. THE PATIENT WAS IN SINUS RHYTHM. TRANSSEPTAL ACCESS WAS INFERIOR AND MID. THE PATIENT'S ACTIVATED CLOTTING TIME (ACT) LEVEL WAS BETWEEN 250 AND 350 SECONDS. DURING THE PROCEDURE THE OCCLUDER WAS PARTIALLY RE-CAPTURED ONCE. CLOSE CRITERIA WERE MET AND THE OCCLUDER HAD A TIRE-SHAPED COMPRESSION. THE OCCLUDER WAS IMPLANTED. APPROXIMATELY AN HOUR AND A HALF AFTER THE PROCEDURE THE PATIENT WAS FOUND UNRESPONSIVE IN RECOVERY. CARDIOPULMONARY RESUSCITATION (CPR) WAS STARTED AND THE PATIENT WENT INTO VENTRICULAR FIBRILLATION. AN X-RAY WAS PERFORMED AND THE OCCLUDER WAS IN PLACE. ADDITIONAL CPR WAS PERFORMED. A TRANSTHORACIC ECHOCARDIOGRAM (TTE) SHOWED A SMALL, CIRCUMFERENTIAL PERICARDIAL EFFUSION IN THE APEX OF THE HEART. PERICARDIOCENTESIS WAS PERFORMED. THE PATIENT WAS SHOCKED SEVERAL TIMES AND PRESENTED WITH BOUTS OF VENTRICULAR FIBRILLATION, VENTRICULAR TACHYCARDIA, AND ASYSTOLE. 1500 ML OF DARK LIQUID WAS DRAINED. ANOTHER PERICARDIAL ACCESS WAS DONE AND 200 ML OF LIQUID WAS DRAINED. THE PATIENT WAS SHOCKED SEVERAL MORE TIMES. AT THIS POINT CPR HAD BEEN PERFORMED FOR AN HOUR AND TWENTY MINUTES. NO EKG ACTIVITY WAS NOTED AND THE PATIENT WAS PRONOUNCED DEAD. THE CAUSE OF DEATH WAS SUSPECTED FROM THE PERICARDIAL EFFUSION OR A STROKE BUT THIS WAS UNCONFIRMED. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT FOR PATIENT EFFECTS OF PERICARDIAL EFFUSION, TACHYCARDIA, CARDIAC ARREST, VENTRICULAR FIBRILLATION AND PATIENT DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE REMAINED IMPLANTED AND WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL DEFINED MANUFACTURING SPECIFICATIONS. BASED ON THE INFORMATION RECEIVED AND PER EVENT DETAILS, THE CAUSE OF THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED. IT IS POSSIBLE DUE TO PROCEDURAL CONDITIONS, HOWEVER THIS CANNOT BE CONFIRMED. THERE WAS NO ALLEGATION OF MALFUNCTION AGAINST THE ABBOTT DEVICE OR PROCEDURE. THE REPORTED UNEXPECTED MEDICAL INTERVENTIONS WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES, AS PERICARDIOCENTESIS AND CPR WERE PERFORMED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. FROM MEDICAL REVIEW: "THE CAUSE OF DEATH IS PROCEDURE RELATED DUE TO A CARDIAC INJURY CAUSING AN ACUTE PERICARDIAL BLEEDING."
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22204707&pc=NIM | 2183870-2025-00305 | 2025-02-28 05:00 | Death | COVIDIEN | 2025-06-13 04:00 | NIM | PROTEGE RX CAROTID STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Stroke/CVA | P060001 |  | 1 | Event Description: OBJECTIVE INTRAPROCEDURAL ATHERO EMBOLIZATION DURING CAROTID ARTERY STENTING (CAS) CAN BE REDUCED THROUGH CAREFUL PATIENT SELECTION, CONSIDERATION OF VASCULAR ANATOMY AND LESION CHARACTERISTICS, OPERATOR AND INSTITUTIONAL EXPERIENCE, PERI-PROCEDURAL ANTITHROMBOTIC AND ANTIPLATELET THERAPY, AND USE OF EMBOLIC PROTECTION. HOWEVER, CAS CAN ALSO RESULT IN STROKE AS THE STENT IS DEPLOYED AND EMBOLIC PROTECTION WITHDRAWN. THE FREE-CELL AREA OF MOST CLOSED-CELL STENTS IS <(><<)>5 MM2, AND =5 MM2 FOR OPEN-CELL STENTS. THE LARGER AREA MAY PERMIT ESCAPE OF MORE ATHEROMATOUS DEBRIS. COMPARISONS OF CLINICAL OUTCOMES BETWEEN CLOSED-CELL AND OPEN-CELL STENTS HAVE BEEN INCONCLUSIVE. THE AIM OF THIS STUDY IS TO COMPARE CLINICAL OUTCOMES ASSOCIATED WITH CAS USING OPEN-CELL VS CLOSED-CELL STENTS. METHODS THE CREST-2-REGISTRY (C2R) ENROLLS ASYMPTOMATIC AND SYMPTOMATIC PATIENTS FOR WHOM CAS IS FAVORED BECAUSE OF HIGH RISK FOR SURGERY OR PATIENT PREFERENCE. C2R IMPLEMENTS OPERATOR- AND SITE-CREDENTIALING, CAREFUL LESION SELECTION, AND STANDARDIZED PROCEDURAL PROTOCOLS. PATIENT CHARACTERISTICS, PROCEDURAL DETAILS, AND OUTCOMES ARE RECORDED. INTERVENTIONISTS MAY USE UNITED STATES FOOD AND DRUG ADMINISTRATION-APPROVED DEVICES INCLUDING OPEN-CELL STENTS (RX ACCULINK [ABBOTT VASCULAR], PRECISE PRO RX [CORDIS-CARDINAL HEALTH], AND PROTÉGÉ RX [MEDTRONIC/COVIDIEN]), OR CLOSED-CELL STENTS (XACT [ABBOTT VASCULAR] AND WALLSTENT MONORAIL ENDOPROSTHESIS [BOSTON SCIENTIFIC]). MULTIVARIABLE LOGISTIC REGRESSION WAS USED TO ASSESS RELATE STENT CELL CONFIGURATION TO PERI-PROCEDURAL (30-DAY) STROKE OR DEATH (SD). RESULTS OF 5307 PROCEDURES PERFORMED BY 163 INTERVENTIONISTS ACROSS 101 CLINICAL CENTERS, 2054 (38.7%) RECEIVED OPEN-CELL STENTS, AND 3253 (61.3%) RECEIVED CLOSED-CELL STENTS. IN THE PERIPROCEDURAL PERIOD, 91 PATIENTS (1.7%) EXPERIENCED A STROKE (3 WERE FATAL), AND 16 PATIENTS DIED WITHOUT EXPERIENCING STROKES (0.4%). AFTER ADJUSTING FOR AGE, SEX, SYMPTOMATIC STATUS, AND CASE URGENCY, AND FOR EFFECT-MODIFICATION BY INDICATION, PERIPROCEDURAL SD WAS SIGNIFICANTLY HIGHER WHEN AN OPEN-CELL STENT WAS PLACED IN A PRIMARY LESION COMPARED WITH CLOSED-CELL STENTS (3.5 EVENTS PER 100 PROCEDURES USING OPEN-CELL STENTS [95% CONFIDENCE INTERVAL [CI], 2.6-4.7] VS 2.2% [95% CI, 1.6-3.0] USING CLOSED-CELL STENTS (ODDS RATIO, 1.59; 95% CI, 1.13-2.23; P <(><<)> .01). PERIPROCEDURAL SD WAS NOT SIGNIFICANTLY DIFFERENT BETWEEN STENT TYPES WHEN PLACED IN A RESTENOTIC LESION (1.2% [95% CI, 0.4-3.3]) USING OPEN-CELL STENTS VS 4.0% (95% CI, 2.2-7.2) USING CLOSED-CELL STENTS (ODDS RATIO, 0.31; 95% CI, 0.09-1.01; P = .052). CONCLUSIONS STENT DESIGN INFLUENCES PERIPROCEDURAL STROKE OR DEATH IN CAROTID STENTING. CLOSED-CELL STENTS ARE ASSOCIATED WITH A LOWER EVENT RATE WHEN TREATING PRIMARY ATHEROSCLEROSIS, BUT NOT IN THE SETTING OF RESTENOSIS. Manufacturer Narrative: CAROTID ARTERY STENTING WITH OPEN VS CLOSED STENT CELL CONFIGURATIONS IN THE CREST-2 REGISTRY 2025 THE AUTHOR(S). PUBLISHED BY ELSEVIER INC. ON BEHALF OF THE SOCIETY FOR VASCULAR SURGERY. THIS IS AN OPEN ACCESS ARTICLE UNDER THE CC BY LICENSE (HTTP://CREATIVECOMMONS.ORG/LICENSES/BY/4.0/). HTTPS://DOI.ORG/10.1016/J.JVS.2025.02.025 A2 AVERAGE AGE A3 MAJORITY GENDER B3 DATE OF PUBLICATION MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22205053&pc=OZD | 1220648-2025-29292 | 2025-05-14 04:00 | Death | ABIOMED, INC. | 2025-06-13 04:00 | OZD | IMPELLA CP | Pumping Stopped; Failure to Read Input Signal; Device Sensing Problem |  | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT (UNKNOWN RACE AND ETHNICITY) WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND DECOMPENSATED RIGHT AFTER IMPLANT. VASOPRESSIN, EPINEPHRINE, AND AMIODARONE DRIPS WERE STARTED. THE PATIENT HAD SUCTION ALARMS AND PERFORMANCE LEVEL P-6 WAS THE HIGHEST LEVEL ABLE TO BE ACHIEVED. ECHOCARDIOGRAM HAD POOR WINDOWS; HOWEVER, THE PHYSICIAN WAS CONTENT WITH THE IMPELLA CP POSITION. THE PATIENT CONTINUED TO HAVE POOR OXYGENATION AND POTENTIAL PLAN FOR EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). THE PATIENT CONTINUED TO BE LABILE. THE PATIENT WAS PLANNED FOR TRANSFER AND DURING TRANSFER THE FOLLOWING DAY, UPON AUTOMATED IMPELLA CONTROLLER (AIC) TO AIC TRANSFER, ¿IMPELLA DEFECTIVE¿ ALARM APPEARED. THE NEW AIC USED AND ¿PLACEMENT SIGNAL NOT RELIABLE¿ ALARMED. PER TRANSPORT PERSONNEL, THE IMPELLA SHUT OFF THREE TIMES AND RESTARTED. IT WAS NOTED THE IMPELLA CORD NOT COMPLETELY PLUGGED IN AS SMALL GAP WAS VISIBLE BETWEEN CONSOLE PLUG IN AND IMPELLA CORD. THE PATIENT'S MEAN ARTERIAL PRESSURE WAS IN THE 70S AND ECMO AT 3L. IMPELLA CORD WAS UNPLUGGED TO ASSESS PRONGS, WHICH WERE ALL STRAIGHT THEN REPLUGGED IN AND STILL NOT FITTING FLUSH INTO PLUG IN. THE IMPELLA CP DEVICE WAS AT P-6/ WITH FLOWS OF 2.4L AND PLACEMENT SIGNAL WAS NOT VISIBLE ON AIC. THE PHYSICIAN WAS OKAY WITH MUTING THE PLACEMENT SIGNAL UNRELIABLE ALARM WHICH WAS COMPLETED WITH GUIDANCE. AN ECHOCARDIOGRAM WAS COMPLETED AND THE IMPELLA CP MEASURED 3.46CM WITHIN MID LEFT VENTRICLE. THE PATIENT WAS NOTED TO ALSO HAS SEVERE MITRAL REGURGITATION AND THE PHYSICIAN FELT, HOWEVER, THE MITRAL REGURGITATION WAS CAUSING A FALSE NORMAL EJECTION FRACTION. THE PATIENT REMAINS INTUBATED AND SEDATED ON SUPPORT. HOWEVER, THE PATIENT WOULD EVENTUALLY EXPIRE. CARE WAS WITHDRAWN. Manufacturer Narrative: MEDICAL SAFETY REVIEWED THE EVENT AND DETERMINED THAT THE DEMISE OUTCOME IS ASSIGNED TO BOTH THE IMPELLA CP PUMP AND AUTOMATED IMPELLA CONTROLLER. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP PUMP AND AUTOMATED IMPELLA CONTROLLER USED AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME GIVEN EVENT DETAILS ON BOTH DEVICES. THE IMPELLA CP IS ONE OF TWO DEVICES ASSOCIATED WITH THE EVENT. REFER TO INITIAL MEDICAL DEVICE REPORT FOR THE AUTOMATED IMPELLA CONTROLLER SERIAL NUMBER (B)(6) EVENT DATE 14-MAY-2025 AND PATIENT IDENTIFIER END IN (B)(6). THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION FOR THE OPTICAL SIGNAL ISSUE, TEMPORARY PUMP STOP, AND THE PUMP NOT DETECTED ISSUE HAS BEEN COMPLETED. THE DATA LOGS SHOW THAT THE IMPELLA WAS RUNNING WITHOUT ISSUE 5/13 - 5/14 TILL 11:49 AM ON CONSOLE IC9283. THERE WERE NO IMPELLA DEFECTIVE ALARMS, PUMP STOPS OR PLACEMENT SIGNAL ISSUES. THERE IS A GAP IN TIME IN WHICH THE PUMP WAS PLUGGED INTO IC2043 IN WHICH THEY RECEIVED AN IMPELLA DEFECTIVE ALARM. THERE ARE NO DATA LOGS RETURNED FOR IC2043. THE ROOT CAUSE OF THE PUMP NOT DETECTED CANNOT BE DETERMINED DUE TO LACK OF DATA LOGS OR PRODUCT RETURNED. 5/14 AT 1:39PM THE PUMP IS PLUGGED INTO IC2045, THE PUMP IS BEGINNING TO RAMP UP P-LEVEL UNTIL IT HAS ALARM 115 AND 119 PUMP STOPS WITH PS NOT RELIABLE ALARMS. THE SNR IS 0, GAIN AND LIGHT ARE INCREASED, LAMP IS MAXED OUT AND CONTRAST IS VARIABLE. THERE ARE NO SPIKES IN MOTOR CURRENT (MC) OR INCREASES IN PURGE PRESSURE, BUT THE PURGE FLOW INCREASES FROM 5 TO ABOUT 17ML/HR. 5/14 AT 1:52 THE PUMP IS RUNNING AGAIN, MOTOR CURRENT WAS VERY STABLE, PS WAS MAXED OUT AT 3000 TRIGGERING PSNR ALARMS AND THERE WERE CONTROLLER ERROR ALARMS. THE SNR IS 0, GAIN AND LIGHT ARE INCREASED, LAMP IS MAXED OUT AND CONTRAST IS VARIABLE. THE PUMP RUNS UNTIL ABOUT 5/26 WHERE IS IT ULTIMATELY EXPLANTED. DUE TO LACK OF PRODUCT RETURNED, THE ROOT CAUSE OF THE TEMPORARY PUMP STOP CANNOT BE DETERMINED. DUE TO LACK OF PRODUCT RETURNED THE ROOT CAUSE OF THE OPTICAL SIGNAL ISSUE CANNOT BE DETERMINED. MSO REVIEW: THE PATIENT'S OUTCOME WAS REVIEWED BY ABIOMED'S MEDICAL SAFETY OFFICER (MSO) AND IT WAS DETERMINED THAT THE PATIENT'S DEATH IS ASSIGNED TO BOTH THE PUMP AND AIC (CONSOLE). THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP PUMP AND AUTOMATED IMPELLA CONTROLLER USED AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME GIVEN EVENT DETAILS ON BOTH DEVICES. ADDED H6 MED DEV PROB CODES 1503, 2917
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22205133&pc=DSQ | 2916596-2025-03894 | 2023-07-31 04:00 | Death | THORATEC CORPORATION | 2025-06-13 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLE ¿EVALUATION OF WARFARIN DOSE AND INR TIME IN THERAPEUTIC RANGE IN LEFT VENTRICULAR ASSIST DEVICE PATIENTS WITH SLEEVE GASTRECTOMY¿ THAT THE HEARTMATE 3 MAY BE ASSOCIATED WITH, BLEEDING, THROMBOEMBOLIC EVENTS, PUMP THROMBOSIS, TRANSPLANT, AND DEATH. A SINGLE-CENTER, OBSERVATIONAL, RETROSPECTIVE COHORT REVIEW WAS PERFORMED WITH THE INTENT OF IDENTIFYING TRENDS IN WARFARIN REQUIREMENTS BASED ON DOSE-RESPONSE. DATA WAS COLLECTED THROUGH REVIEW OF ELECTRONIC MEDICAL RECORDS. PATIENTS WERE IDENTIFIED FOR INCLUSION FROM A COMPREHENSIVE LIST OF PATIENTS IMPLANTED WITH A CONTINUOUS FLOW LEFT VENTRICULAR ASSIST DEVICE (HEARTWARE HVAD, HEARTMATE II, AND HEARTMATE 3) AT THE STUDY CENTER BETWEEN JANUARY 1, 2016, AND JULY 31, 2023. THOSE THAT RECEIVED A CONCOMITANT SLEEVE GASTRECTOMY (SG) DURING LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION WERE IDENTIFIED USING CURRENT PROCEDURAL TERMINOLOGY (CPT) CODES. PATIENTS WERE ELIGIBLE FOR INCLUSION IF THEY WERE GREATER THAN OR EQUAL TO 18 YEARS OF AGE AT TIME OF IMPLANTATION WITH AN ACTIVE WARFARIN PRESCRIPTION FOR AT LEAST 8 WEEKS POST DISCHARGE. PATIENTS WERE EXCLUDED IF THEY HAD AN ALLERGY, INTOLERANCE OR CONTRAINDICATION TO WARFARIN, WERE DECEASED WITHIN 30 DAYS OF IMPLANT OR PRIOR TO HOSPITAL DISCHARGE, HAD A SG PERFORMED IN A SEPARATE ADMISSION PRIOR TO LVAD IMPLANTATION, OR WERE NON-ADHERENT TO WARFARIN THERAPY. 335 PATIENTS WERE EVALUATED FOR INCLUSION. OF THOSE, 33 MET EXCLUSION CRITERIA, LEAVING 48 PATIENTS ELIGIBLE IN THE LVAD PLUS SG GROUP AND 254 IN THE LVAD-ONLY GROUP. AFTER PROPENSITY SCORE MATCHING, 48 LVAD PLUS SG PATIENTS WERE MATCHED, RESULTING IN 96 LVAD-ONLY PATIENTS. THEREFORE, A TOTAL OF 144 PATIENTS WERE INCLUDED IN THE FINAL ANALYSES. THE PRIMARY ENDPOINT WAS THE COMPARISON OF MEAN TOTAL WEEKLY DOSE REQUIREMENT OF WARFARIN AT BASELINE (DAY OF INITIAL HOSPITAL DISCHARGE) AND AT INTERVALS ASSESSED BIWEEKLY THEREAFTER FOR AT LEAST 8 WEEKS. SECONDARY END POINTS INCLUDED THE TIME IN THERAPEUTIC INR RANGE (TTR) CALCULATED USING THE ROSENDAAL METHOD, AND THE MEAN WEIGHT-BASED TOTAL WEEKLY DOSE REQUIREMENTS. SECONDARY ENDPOINTS WERE COLLECTED AT BASELINE, WEEK 4, AND WEEK 8. SAFETY ENDPOINTS INCLUDED THE INCIDENCE OF MAJOR BLEEDING, MINOR BLEEDING, AND THROMBOEMBOLIC EVENTS. MAJOR AND MINOR BLEEDING EVENTS WERE DEFINED ACCORDING TO THE INTERNATIONAL SOCIETY OF THROMBOSIS AND HEMOSTASIS (ISTH) BLEEDING DEFINITIONS OF MAJOR BLEEDING AND CLINICALLY RELEVANT NON-MAJOR BLEEDING, RESPECTIVELY. THROMBOEMBOLIC EVENTS WERE A COMPOSITE OF VENOUS THROMBOEMBOLISM, ARTERIAL THROMBOEMBOLISM AND PUMP THROMBOSIS. DATA BEYOND 8 WEEKS WAS COLLECTED RETROSPECTIVELY FOR ALL PRIMARY AND SECONDARY OUTCOMES AT MONTH 3, 6, AND 12 IF FOLLOW-UP DATA WAS AVAILABLE. DATA REPORTING CEASED AFTER TIME POINT OF PATIENT DEATH, HEART TRANSPLANTATION, NON-ADHERENCE TO WARFARIN THERAPY, OR AT 12 MONTHS POST-DISCHARGE. THE MEAN TOTAL WEEKLY DOSE OF WARFARIN AT BASELINE WAS NUMERICALLY LOWER IN THE LVAD PLUS SG GROUP (30.9 MG) COMPARED WITH THE LVAD-ONLY GROUP (34.9 MG), HOWEVER THIS WAS NOT STATISTICALLY SIGNIFICANT. THE TOTAL WEEKLY DOSE CONTINUED TO RISE THROUGH THE STUDY PERIOD IN THE LVAD PLUS SG GROUP WHEREBY MONTH 12, THIS GROUP HAD A HIGHER WEEKLY DOSE COMPARED WITH THE LVAD-ONLY GROUP (46.8 MG VS 38.1 MG, P < 0.05). WHEN COMPARING PATIENTS THAT HAD AN INR GOAL OF 2¿3 THROUGHOUT THE STUDY PERIOD (N = 40 IN LVAD PLUS SG, N = 82 IN LVAD-ONLY), SIMILAR RESULTS WERE OBSERVED EXCEPT FOR A SIGNIFICANTLY LOWER DOSE IN THE LVAD PLUS SG GROUP AT BASE LINE. THERE WERE NO DIFFERENCES IN CLINICAL EVENTS INCLUDING BLEEDING OR THROMBOEMBOLIC EVENTS OBSERVED BETWEEN GROUPS FOR THE DURATION OF THE STUDY PERIOD. FOR SAFETY, THE INCIDENCE OF MAJOR BLEEDING WAS 10% IN EACH GROUP, ALL DOCUMENTED AS GASTROINTESTINAL BLEEDS IN THE LVAD PLUS SG GROUP, AND MOST OF THE BLEEDING BEING INTRA-CRANIAL IN THE LVAD-ONLY GROUP. THERE WAS NO DIFFERENCE IN MAJOR OR MINOR BLEEDING EVENTS. A LOW RATE OF THROMBOEMBOLIC EVENTS WAS OBSERVED. IN CONCLUSION, WARFARIN REQUIREMENTS ARE GENERALLY LOWER AT BASELINE POST-SG, ESPECIALLY WHEN COMPARED IN A WEIGHT-BASED FASHION, BUT BY 12 MONTHS, THE LVAD PLUS SG GROUP REQUIRE HIGHER DOSES COMPARED WITH LVAD IMPLANTED PATIENTS THAT HAVE NOT UNDERGONE SG. Manufacturer Narrative: SECTION B: DATE OF EVENT IS APPROXIMATE AS THE DATA WERE COLLECTED BETWEEN 1JAN2016 AND 31JUL2023. SECTION D4: DEVICE SERIAL NUMBER AND LOT NUMBER WERE NOT PROVIDED, AND EXPIRATION DATE AND MANUFACTURE DATE (H4) CANNOT BE DETERMINED. THE PRIMARY UDI NUMBER IN D4 IS REFLECTIVE OF ALL AVAILABLE INFORMATION. AUTHOR CITATION: ROBERTSON, S., (2025). EVALUATION OF WARFARIN DOSE AND INR TIME IN THERAPEUTIC RANGE IN LEFT VENTRICULAR ASSIST DEVICE PATIENTS WITH SLEEVE GASTRECTOMY. THE INTERNATIONAL JOURNAL OF ARTIFICIAL ORGANS, 48(2), 77¿83. HTTPS://DOI.ORG/10.1177/03913988251317820. SECTION E1: CUSTOMER SITE WAS (B)(6) MEDICAL CENTER. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: SECTION E2 CORRECTION. SECTION E3 CORRECTION. MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. NO PRODUCT WAS EVALUATED UNDER THIS COMPLAINT. THE HEARTMATE 3 DEVICE SERIAL NUMBERS, AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. D AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22205532&pc=BZD | 2518422-2025-106611 | 2025-06-05 04:00 | Death | RESPIRONICS, INC. | 2025-06-13 04:00 | BZD | REP DREAMSTATION AUTO CPAP, DOM - RECRT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT REPORTEDLY PASSED AWAY BEFORE RECEIVING THE REPLACEMENT DREAMSTATION AUTO CPAP DEVICE, AND THE DEVICE WAS NO LONGER NEEDED. THE DAUGHTER WAS INSTRUCTED TO SEND THE DEVICE BACK WITH THE RETURN LABEL. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22205543&pc=OAE | 2029046-2025-01939 | 2025-05-29 04:00 | Death | BIOSENSE WEBSTER INC | 2025-06-13 04:00 | OAE | QDOT MICRO¿ CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Tamponade | P210027 |  | 1 | Event Description: THE REPORT INDICATED THAT DURING ISCHEMIC VT (VENTRICULAR TACHYCARDIA) ABLATION PROCEDURE WITH A QDOT MICRO¿ CATHETER, THE PATIENT EXPERIENCED BLOOD PRESSURE DROP, AND A PERICARDIAL EFFUSION WAS DISCOVERED WITH INTRACARDIAC ECHO TREATED WITH REMOVAL OF ¾ LITER OF BLOOD REMOVED AND DRAIN LEFT IN PLACE. THE PATIENT EXPIRED WITHIN A WEEK AFTER THE PROCEDURE. DURING ISCHEMIC VT ABLATION, AN IMPELLA DEVICE WAS PLACED AT THE BEGINNING OF THE PROCEDURE FOR CARDIAC SUPPORT. THE PHYSICIAN WAS ABLATING IN THE LV (LEFT VENTRICULAR) APEX WHEN IT WAS NOTED THAT THE PATIENT'S BLOOD PRESSURE WAS SLOWLY DROPPING. A PERICARDIAL EFFUSION WAS DISCOVERED USING INTRACARDIAC ECHO. THERE WAS NO SPIKE IN IMPEDANCE OR FORCE READINGS PRIOR TO FINDING THE EFFUSION. A PERICARDIOCENTESIS WAS PERFORMED AND APPROXIMATELY "3/4 LITER" OF BLOOD WAS REMOVED. THE PERICARDIAL DRAIN WAS LEFT IN. THE PATIENT WAS TRANSFERRED TO ICU WITH THE IMPELLA DEVICE IN PLACE FOR FURTHER MONITORING. THE PHYSICIAN STATED THAT THE PATIENT HAD DIED "WITHIN THE LAST WEEK", BUT THE DAY OF DEATH IS UNKNOWN. THE PHYSICIAN STATED THAT THE PATIENT HAD EXTENSIVE HEALTH ISSUES, AND THAT THE FAMILY DECIDED TO DISCONTINUE MEDICAL TREATMENT. DEVICES USED FOR ABLATION INCLUDED QDOT, 8FR SOUNDSTAR, LARGE CURVE VIZIGO SHEATH, SMALL CURVE VIZIGO SHEATH, CARTO 3, NGEN, AND JOSEPHEN QUAD CATHETER. ALL CATHETERS HAVE BEEN DISCARDED AND ARE NOT AVAILABLE FOR RETURN. NON BWI DEVICES USED BAYLESS TRANSEPTAL NEEDLE AND IMPELLA. Manufacturer Narrative: D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. BASED ON THE INFORMATION AVAILABLE, THERE IS NO INDICATION THAT A DESIGN OR MANUFACTURING ISSUE HAS CAUSED THE REPORTED COMPLAINT CONDITION. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC. OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22205706&pc=OZD | 1220648-2025-29295 | 2025-05-14 04:00 | Death | ABIOMED, INC. | 2025-06-13 04:00 | OZD | AUTOMATED IMPELLA CONTROLLER | Failure to Read Input Signal |  | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT (UNKNOWN RACE AND ETHNICITY) WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND DECOMPENSATED RIGHT AFTER IMPLANT. VASOPRESSIN, EPINEPHRINE, AND AMIODARONE DRIPS WERE STARTED. THE PATIENT HAD SUCTION ALARMS AND PERFORMANCE LEVEL P-6 WAS THE HIGHEST LEVEL ABLE TO BE ACHIEVED. ECHOCARDIOGRAM HAD POOR WINDOWS; HOWEVER, THE PHYSICIAN WAS CONTENT WITH THE IMPELLA CP POSITION. THE PATIENT CONTINUED TO HAVE POOR OXYGENATION AND POTENTIAL PLAN FOR EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). THE PATIENT CONTINUED TO BE LABILE. THE PATIENT WAS PLANNED FOR TRANSFER AND DURING TRANSFER THE FOLLOWING DAY, UPON AUTOMATED IMPELLA CONTROLLER (AIC) TO AIC TRANSFER, ¿IMPELLA DEFECTIVE¿ ALARM APPEARED. THE NEW AIC USED AND ¿PLACEMENT SIGNAL NOT RELIABLE¿ ALARMED. PER TRANSPORT PERSONNEL, THE IMPELLA SHUT OFF THREE TIMES AND RESTARTED. IT WAS NOTED THE IMPELLA CORD NOT COMPLETELY PLUGGED IN AS SMALL GAP WAS VISIBLE BETWEEN CONSOLE PLUG IN AND IMPELLA CORD. THE PATIENT'S MEAN ARTERIAL PRESSURE WAS IN THE 70S AND ECMO AT 3L. IMPELLA CORD WAS UNPLUGGED TO ASSESS PRONGS, WHICH WERE ALL STRAIGHT THEN REPLUGGED IN AND STILL NOT FITTING FLUSH INTO PLUG IN. THE IMPELLA CP DEVICE WAS AT P-6/ WITH FLOWS OF 2.4L AND PLACEMENT SIGNAL WAS NOT VISIBLE ON AIC. THE PHYSICIAN WAS OKAY WITH MUTING THE PLACEMENT SIGNAL UNRELIABLE ALARM WHICH WAS COMPLETED WITH GUIDANCE. AN ECHOCARDIOGRAM WAS COMPLETED AND THE IMPELLA CP MEASURED 3.46CM WITHIN MID LEFT VENTRICLE. THE PATIENT WAS NOTED TO ALSO HAS SEVERE MITRAL REGURGITATION AND THE PHYSICIAN FELT, HOWEVER, THE MITRAL REGURGITATION WAS CAUSING A FALSE NORMAL EJECTION FRACTION. THE PATIENT REMAINS INTUBATED AND SEDATED ON SUPPORT. HOWEVER, THE PATIENT WOULD EVENTUALLY EXPIRE. CARE WAS WITHDRAWN. Manufacturer Narrative: MEDICAL SAFETY REVIEWED THE EVENT AND DETERMINED THAT THE DEMISE OUTCOME IS ASSIGNED TO BOTH THE IMPELLA CP PUMP AND AUTOMATED IMPELLA CONTROLLER. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE THE IMPELLA CP PUMP AND AUTOMATED IMPELLA CONTROLLER USED AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME GIVEN EVENT DETAILS ON BOTH DEVICES. THE AUTOMATED IMPELLA CONTROLLER IS ONE OF TWO DEVICES ASSOCIATED WITH THE EVENT. REFER TO MANUFACTURING REPORT NUMBER 1220648-2025-29292 FOR THE REPORT ON THE IMPELLA CP. THE AUTOMATED IMPELLA CONTROLLER DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF PUMP NOT DETECTED HAS BEEN COMPLETED. THE CAUSE OF THE IMPELLA DEFECTIVE ALARM WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND LOGS WERE NOT AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22205826&pc=OZD | 1220648-2025-29297 | 2025-05-14 04:00 | Death | ABIOMED, INC. | 2025-06-13 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Ischemia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WITH AN ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT AND EXPERIENCED LIMB ISCHEMIA REQUIRING SURGICAL INTERVENTION AND THEREAFTER WOULD EXPIRE. AFTER IMPELLA CP DEVICE IMPLANT, THE PATIENT UNDERWENT EMERGENT PERCUTANEOUS CORONARY INTERVENTION WITH ONE DRUG-ELUTING STENT (DES) TO THE LEFT MAIN AND THREE DES TO THE LEFT ANTERIOR DESCENDING ARTERY WITH IMPELLA CP DEVICE IMPLANTED FOR CARDIOGENIC SHOCK MANAGEMENT. THE PATIENT WAS TRANSFERRED TO THE UNIT ON SUPPORT. LATER THAT DAY, THE PATIENT DEVELOPED LIMB ISCHEMIA AND UNDERWENT A FEM-FEM BYPASS. THE PATIENT CONTINUED ON IMPELLA CP MCS UNTIL HOURS LATER BEFORE BEING ESCALATED TO AN IMPELLA 5.5 DEVICE. THE FOLLOWING DAY AFTER ESCALATION, APPROXIMATELY 16 HOURS LATER, THE PATIENT WOULD EXPIRE. CARE WAS WITHDRAWN. Manufacturer Narrative: MEDICAL SAFETY REVIEWED THE EVENT AND DETERMINED THERE IS NOT ENOUGH INFORMATION TO EXCLUDE THE IMPELLA CP USED AS AN ASSOCIATED FACTOR TO THE PATIENT'S OUTCOME GIVEN THE DEVICE-RELATED ISCHEMIA EVENT REQUIRING SURGICAL INTERVENTION AND PATIENT EXPIRING SHORTLY THEREAFTER. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER; THEREFORE, AN EVALUATION/ANALYSIS OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM. SECTION: POTENTIAL ADVERSE EVENTS. ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY¿ (INCLUDING VENTRICULAR PERFORATION).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22205984&pc=BTI | 2518422-2025-047253 | 2023-06-07 04:00 | Death | RESPIRONICS, INC. | 2025-06-13 04:00 | BTI | INNOSPIRE ESSENCE | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K042655 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH AN ESSENCE RENTAL WITH SSD 230V50HZ DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED
DEATH. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A PATIENT WITH AN ESSENCE RENTAL WITH SSD 230V50HZ DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. AFTER ATTEMPTS MADE BY THE MANUFACTURER, THE DEVICE HAS NOT YET BEEN RETURNED FOR EVALUATION. THERE HAS BEEN NO RESPONSE FROM THE CUSTOMER ON THE RETURN OF THE DEVICE, THEREFORE THE DEVICE HAS BEEN MARKED LOST. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. BOX H HAS BEEN UPDATED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22206197&pc=OZD | 1220648-2025-29298 | 2025-05-15 04:00 | Death | ABIOMED, INC. | 2025-06-13 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT HAD A HIGH-RISK PERCUTANEOUS CORONARY INTERVENTION SUCCESSFULLY SUPPORTED BY THE IMPELLA CP. THERE WERE MULTIPLE BALLOON INFLATIONS AND SHOCKWAVE WERE NEEDED TO DELIVER 3 DES. THE PATIENT REMAINED STABLE THROUGHOUT THE CASE. AFTER 2 HOURS OF SUPPORT THE WAS IMPELLA WEANED AND EXPLANTED. THE ACCESS SITE CLOSED VIA TWO PERCLOSE AND A 6FR ANGIOSEAL WAS ADDED DUE TO TRACK OOZE. LATER IT WAS REPORTED THAT ABOUT 21 DAYS THE PATIENT WAS FOUND UNRESPONSIVE FOLLOWING A MOTOR VEHICLE ACCIDENT. A LARGE AMOUNT OF BLOOD WAS NOTED UNDER THE PATIENT THAT WAS COMING FROM RIGHT FEMORAL ACCESS SITE (IMPELLA CP ACCESS SITE). PARAMEDICS WERE UNABLE TO REVIVE THE PATIENT AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22206376&pc=CBK | 2518422-2025-106603 | 2025-06-11 04:00 | Death | RESPIRONICS, INC. | 2025-06-13 04:00 | CBK | TRILOGY EVO | No Apparent Adverse Event | Insufficient Information | K181166 |  | 1 | Event Description: THIS NOTIFICATION WAS OPENED DUE TO A REPORT FROM A RESPIRATORY THERAPIST (RT) WHO WAS TOLD BY A FAMILY MEMBER WHEN THEY CALLED THE PATIENT TO SCHEDULE AN APPOINTMENT TO COME OUT FOR A VENT CHECK, THAT THE PATIENT HAD PASSED AWAY. THE RT REPORTED THE EVENT OCCURRED (B)(6) 2025. THE PATIENT WAS FOUND AROUND 1-2PM AND 911 WAS CALLED. THE REPORTER STATED THERE WAS NO DEVICE MALFUNCTION AT THIS TIME AND THE PATIENT WAS ON A PASSIVE CIRCUIT. ACCORDING TO THE PATIENT'S BROTHER, HE FOUND THE PATIENT AND STATED THE VENTILATOR WAS OFF. BASED ON THE INFORMATION CURRENTLY PROVIDED AND AVAILABLE, DEVICE CAUSE AND/OR CONTRIBUTION TO THE REPORTED EVENT CANNOT BE CONFIRMED, NOR REFUTED AT THIS TIME BASED ON THE EVIDENCE PRESENT. FURTHER GOOD FAITH EFFORTS AND INFORMATION GATHERING ARE REQUIRED TO DISPOSITION DEVICE CAUSE AND/OR CONTRIBUTION TO THE PATIENT HARMS INCURRED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO DETERMINE ISSUES REGARDING THE FUNCTIONALITY OF THE PATIENT'S VENTILATOR. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE TRILOGY EVO VENTILATOR WAS RETURNED FOR DEVICE EVALUATION AND THE EVALUATION WAS COMPLETED ON 12 AUGUST 2025 WITH THE FOLLOWING FINDINGS: A TRILOGY EVO WAS RETURNED FOR INVESTIGATION DUE TO THE CUSTOMER REQUESTING IT BE CHECKED AFTER A PATIENT PASSED AWAY. THERE WAS NO COMPLAINT THAT THE DEVICE MALFUNCTIONED. THE PRODUCT INVESTIGATION LABORATORY (PIL) WAS UNABLE TO CONFIRM ANY MALFUNCTION OF THE TRILOGY EVO. THERE IS NO EVIDENCE THAT THE DEVICE MALFUNCTIONED IN A MANNER THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY IF IT WERE TO RECUR.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22206649&pc=CBK | 2518422-2025-106615 | 2025-06-02 04:00 | Death | RESPIRONICS, INC. | 2025-06-13 04:00 | CBK | TRILOGY EVO, JAPAN | Adverse Event Without Identified Device or Use Problem | Unspecified Respiratory Problem | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT "A PATIENT'S FAMILY HAS REQUESTED USAGE DATA, AND SUCH BE SUBMITTED BECAUSE THE PATIENT PASSED AWAY. THE CAUSE OF DEATH HAS BEEN REPORTED AS ASPHYXIATION. THE DEVICE KEPT OPERATING WHEN THE EVENT OCCURRED. NO PATIENT INFORMATION WAS DISCLOSED". THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT "A PATIENT'S FAMILY HAS REQUESTED USAGE DATA, AND SUCH BE SUBMITTED BECAUSE THE PATIENT PASSED AWAY. THE CAUSE OF DEATH HAS BEEN REPORTED AS ASPHYXIATION. THE DEVICE KEPT OPERATING WHEN THE EVENT OCCURRED. NO PATIENT INFORMATION WAS DISCLOSED". THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE WAS EVALUATED. THERE WERE NO ABNORMALITIES DURING AN OPERATIONAL CHECK. UPON CHECKING THE EVENT LOG, IT WAS OBSERVED THAT THE "LOW TIDAL VOLUME" ALARM AND "LOW MINUTE VENTILATION" ALARM OCCURRED FOR 1827 SECONDS, AND THE "CIRCUIT DISCONNECT" ALARM OCCURRED FOR 1828 SECONDS. HOWEVER, THERE WERE NO ERRORS INDICATING MALFUNCTIONS OF THE DEVICE. WHEN AN INTERNAL INSPECTION WAS PERFORMED TO CHECK AN ABNORMALITY OF THE DEVICE, NO ABNORMALITIES WERE OBSERVED. UPON CHECKING THE WAVEFORM DATA RECORDED IN THE DEVICE, IT WAS OBSERVED THAT THE LEAK INCREASED. BASED ON THE ABOVE, IT IS BELIEVED THAT THERE ARE NO ABNORMALITIES IN THE DEVICE, AND THE CIRCUIT HAD BEEN REMOVED WHILE USING THE DEVICE. PERFORMED FINAL TEST, BATTERY CHECK, VALVE CHECK, RUN IN TESTS AND CLEANING THEN CONFIRMED THE UNIT OPERATED PROPERLY. THE SOFTWARE VERSION WAS CHECKED. (1.06.10.00).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22206740&pc=NVN | 2124215-2025-39234 | 2024-12-24 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-13 04:00 | NVN | INGEVITY? MRI | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS PACEMAKER PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED TO BE SHOCK DUE TO SEPSIS. THE PACEMAKER AND LEAD SYSTEM WAS EXPLANTED, AND THE PRODUCTS WERE RETURNED FOR ANALYSIS ABOUT SIX MONTHS LATER.Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS PACEMAKER PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED TO BE SHOCK DUE TO SEPSIS. THE PACEMAKER AND LEAD SYSTEM WAS EXPLANTED, AND THE PRODUCTS WERE RETURNED FOR ANALYSIS ABOUT SIX MONTHS LATER.Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS PACEMAKER PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED TO BE SHOCK DUE TO SEPSIS. THE PACEMAKER AND LEAD SYSTEM WAS EXPLANTED, AND THE PRODUCTS WERE RETURNED FOR ANALYSIS ABOUT SIX MONTHS LATER. Manufacturer Narrative: BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN THE MISSING INITIAL REPORTER NAME. THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS. Manufacturer Narrative: BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN THE MISSING INITIAL REPORTER NAME. THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS. ANALYSIS OF THE RETURNED PRODUCT PROVIDED NO RELEVANT INFORMATION FOR THE INFECTION-RELATED ALLEGATION, AND REVIEW OF PRODUCTION RECORDS FOUND NO ISSUES. WE HAVE DETERMINED THAT THE CLINICAL OBSERVATION OF INFECTION IS A KNOWN INHERENT RISK WITH THE USE OF THIS PRODUCT. Manufacturer Narrative: BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN THE MISSING INITIAL REPORTER NAME. THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS. ANALYSIS OF THE RETURNED PRODUCT PROVIDED NO RELEVANT INFORMATION FOR THE INFECTION-RELATED ALLEGATION, AND REVIEW OF PRODUCTION RECORDS FOUND NO ISSUES. WE HAVE DETERMINED THAT THE CLINICAL OBSERVATION OF INFECTION IS A KNOWN INHERENT RISK WITH THE USE OF THIS PRODUCT. THIS REPORT IS BEING FILED TO CORRECT THE EVALUATION METHOD CODE IN H6.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22206745&pc=LWP | 2124215-2025-39231 | 2024-12-24 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-13 04:00 | LWP | ESSENTIO? MRI DR | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS PACEMAKER PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED TO BE SHOCK DUE TO SEPSIS. THE PACEMAKER AND LEAD SYSTEM WAS EXPLANTED, AND THE PRODUCTS WERE RETURNED FOR ANALYSIS ABOUT SIX MONTHS LATER.Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS PACEMAKER PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED TO BE SHOCK DUE TO SEPSIS. THE PACEMAKER AND LEAD SYSTEM WAS EXPLANTED, AND THE PRODUCTS WERE RETURNED FOR ANALYSIS ABOUT SIX MONTHS LATER. Manufacturer Narrative: BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN THE MISSING INITIAL REPORTER NAME. THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS. Manufacturer Narrative: BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN THE MISSING INITIAL REPORTER NAME. THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS. UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THE DEVICE WAS SUCCESSFULLY INTERROGATED AND THE DEVICE MEMORY WAS REVIEWED. VISUAL INSPECTION FOUND NO IRREGULARITIES. NO DEVICE ISSUES WERE NOTED THAT WOULD HAVE MADE INFECTION MORE LIKELY THAN WHAT IS DESCRIBED IN THE INSTRUCTIONS FOR USE AS A KNOWN INHERENT RISK OF THE USE OF THIS PRODUCT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22206751&pc=NVN | 2124215-2025-39237 | 2024-12-24 05:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-13 04:00 | NVN | INGEVITY? MRI | Adverse Event Without Identified Device or Use Problem | Unspecified Infection; Sepsis | P150012 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS PACEMAKER PASSED AWAY. THE CAUSE OF DEATH WAS REPORTED TO BE SHOCK DUE TO SEPSIS. THE PACEMAKER AND LEAD SYSTEM WAS EXPLANTED, AND THE PRODUCTS WERE RETURNED FOR ANALYSIS ABOUT SIX MONTHS LATER. Manufacturer Narrative: BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN THE MISSING INITIAL REPORTER NAME. THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22206872&pc=NGV | 2135147-2025-03220 | 2025-05-21 04:00 | Death | ABBOTT MEDICAL | 2025-06-13 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Pulmonary Embolism; Cardiac Arrest | P200049 |  | 1 | Event Description: IT WAS REPORTED ON (B)(6) 2025 A 22MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE OCCLUDER WAS SELECTED FOR IMPLANT. THE OCCLUDER WAS IMPLANTED. POST-PROCEDURE THE PATIENT WENT INTO CARDIAC ARREST. THE PATIENT WAS PROVIDED FLUID RESUSCITATION AND TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). PULSELESS ELECTRICAL ACTIVITY (PEA) OCCURRED AND THE PATIENT WAS INTUBATED. IT WAS SUSPECTED THAT THE PATIENT HAD A CARDIAC TAMPONADE. THE DECISION WAS MADE TO DRAIN THE EFFUSION; HOWEVER, IT WAS OBSERVED THAT THERE WAS NO PERICARDIAL EFFUSION. THE PATIENT CONTINUED TO CODE UNTIL THE PATIENT'S FAMILY REQUESTED ALL ATTEMPTS TO CEASE. THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS ATTRIBUTED TO A PULMONARY EMBOLISM. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT FOR PATIENT EFFECTS OF CARDIAC ARREST AND DEATH WAS REPORTED. A RETURNED DEVICE INSPECTION, TO RULE OUT ANY DEVICE-RELATED CAUSES, COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE INFORMATION RECEIVED, THE OFFICIAL CAUSE OF DEATH WAS THE CARDIAC ARREST. THE CAUSE OF CARDIAC ARREST COULD NOT BE CONCLUSIVELY DETERMINED. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES WHERE CPR/RESUSCITATION AND PERICARDIOCENTESIS WERE PERFORMED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING. FROM MEDICAL REVIEW: IN THE ABSENCE OF ADDITIONAL CLINICAL INFORMATION RELATED IT IS NOT POSSIBLE TO DETERMINE THE EXACT CAUSE FOR THE CARDIAC ARREST. THE CARDIAC ARREST IS MOST LIKELY PROCEDURE RELATED AND NOT RELATED TO A DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22207829&pc=MNS | 2518422-2025-106623 | 2025-06-11 04:00 | Death | RESPIRONICS, INC. | 2025-06-13 04:00 | MNS | DREAMSTATION BIPAP AVAPS | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION FROM A CLINIC STATING THAT A PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION THAT THE DREAMSTATION BIPAP AVAPS DEVICE CONTRIBUTED TO THE DEATH. IT IS UNKNOWN IF THE PATIENT WAS ON/USING THE DEVICE DURING THE TIME OF DEATH. THE CLINICAL ASSESSMENT TEAM STATES THAT THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USER ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED. BASED ON AVAILABLE INFORMATION AT THIS TIME, THE ALLEGED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. NO FURTHER ACTION IS REQUIRED. IF ADDITIONAL INFORMATION IS OBTAINED ABOUT THIS EVENT, PLEASE SEND IT TO THE PMS CLINICAL EXPERT FOR REVIEW. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM ANY MALFUNCTION OF THE DEVICE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22352603&pc=QJI | MW5172082 |  | Death | BETA BIONICS, INC. | 2025-06-30 04:00 | QJI | ILET DOSING DECISION SOFTWARE | Failure to Sense; Failure to Deliver | Insufficient Information |  |  | 1 | Event Description: OER RECEIVED A CALL FROM THE FDA EMERGENCY CALL CENTER REPORTING A DEATH IN (B)(6). THE CALLER, (B)(6), IS A DEATH INVESTIGATOR AT (B)(6) REPORTING THE DEATH OF A 56-YEAR-OLD MALE WITH TYPE 1 DIABETES. HIS INSULIN WAS RECENTLY CHANGED, AND HE RECEIVED A NEW PUMP 2-3 WEEKS AGO. THE DEVICE IS A BETA BIONICS ILET BIONIC PANCREAS PUMP. THE FAMILY NOTICED SOME PROBLEMS WITH THE DEVICE, BUT DID NOT PROVIDE ANY SPECIFICS. ON FEBRUARY (B)(6), THE INDIVIDUAL WENT TO BED AND WAS FOUND DECEASED THE NEXT MORNING. FROM THE BIOMETRICS APP OF THE DEVICE, IT WAS OBSERVED THAT THE GLUCOSE LEVEL DECLINED SIGNIFICANTLY AFTER MIDNIGHT ON FEB (B)(6) AND INSULIN DELIVERY EVENTUALLY STOPPED. THE CALLER MENTIONED SEEING INFORMATION ON A PREVIOUS FDA INVESTIGATION FROM 2023 THAT APPEARED SIMILAR. Event Description: ADDITIONAL INFORMATION RECEIVED FOR MW5172082 ON 8/14/2025 TO MAKE THE REPORTER UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22353028&pc=MIH | 2247858-2025-00158 | 2025-05-30 04:00 | Death | BOLTON MEDICAL, INC. | 2025-06-30 04:00 | MIH | RELAY PRO NBS THORACIC STENT-GRAFT SYSTEM | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Rupture; Insufficient Information | P200045 |  | 1 | Event Description: "DEATH: THE RELAY PRO STENT GRAFT WAS ATTEMPTED TO BE IMPLANTED USING A PULL-THROUGH TECHNIQUE. HOWEVER, DUE TO THE SEVERE BENDING OF THE DESCENDING AORTA, THE DESCENDING AORTIC ANEURYSM RUPTURED BEFORE THE TECHNIQUE COULD BE USED, WHEN THE DELIVERY SYSTEM WAS ADVANCED INTO THE ARTERY. THE PATIENT SUBSEQUENTLY DIED. THERE WERE NO PROBLEMS WITH THE STENT GRAFT ITSELF, AND THIS EVENT WOULD LIKELY HAVE OCCURRED NO MATTER WHICH DEVICE HAD BEEN USED. CAUSE OF DEATH: RUPTURE OF THE DESCENDING AORTIC ANEURYSM. PHYSICIAN'S COMMENT ON CAUSAL RELATIONSHIP: THE BLOOD VESSEL WAS WEAK, AND THIS EVENT WOULD LIKELY HAVE OCCURRED NO MATTER WHICH DEVICE HAD BEEN USED. OPERATION TYPE: STENT GRAFT IMPLANTATION BLOOD LOSS: UNKNOWN NO IMAGE AVAILABLE DUE TO HOSPITAL POLICY PRE-CASE PLAN AVAILABLE UPON REQUEST NO ADDITIONAL INFORMATION AVAILABLE DUE TO HOSPITAL POLICY (TC#(B)(4). PATIENT OUTCOME: "THE PATIENT DIED." Manufacturer Narrative: BOLTON MEDICAL, INC. (D/B/A TERUMO AORTIC), HEREIN KNOWN AS THE "COMPANY", IS SUBMITTING THIS REPORT PURSUANT TO 21 CFR PART 803, HAS MADE REASONABLE EFFORTS TO OBTAIN COMPLETE INFORMATION, AND HAS PROVIDED AS MUCH AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT IS BASED ON INFORMATION OBTAINED BY THE COMPANY, WHICH MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE REGULATIONS. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN, UNAVAILABLE, OR NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY ANYONE THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE CERTAIN REGULATIONS, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22353055&pc=OZD | 1220648-2025-29865 | 2025-06-05 04:00 | Death | ABIOMED, INC. | 2025-06-30 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Thrombocytopenia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA 5.5 WAS FOUND TO HAVE A SMALL SUBARACHNOID HEMORRHAGE ON COMPUTED TOMOGRAPHY (CT) OF THE HEAD WITHOUT CONTRAST. A CT ANGIOGRAM OF THE HEAD/NECK WERE NEGATIVE. ARGATROBAN WAS ON HOLD. NEUROSURGERY WAS FOLLOWING AND THE ATTENDING MEDICAL DOCTOR DOES NOT BELIEVE FINDINGS ARE TO BE IMPELLA RELATED. THE FOLLOWING DAY, THERE WERE CONCERNS FOR THROMBOCYTOPENIA. ARGATROBAN MIGHT BE RESTARTED WITH CONSERVATIVE PARTIAL THROMBOPLASTIN TIME GOAL. APPROXIMATELY A WEEK LATER, ADVANCED THERAPY WAS DENIED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22353129&pc=CBK | 2518422-2025-106917 | 2025-05-21 04:00 | Death | RESPIRONICS, INC. | 2025-06-30 04:00 | CBK | TRILOGY EVO | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K181166 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION THAT A LOANER TRILOGY EVO VENTILATOR WAS RETURNED TO THE MANUFACTURER DUE TO A PATIENT PASSING AWAY. IT WAS NOTED THAT THE PATIENT WAS NOT ON THE DEVICE AT THE TIME OF DEATH AND THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS EVALUATED BY THE MANUFACTURER AND NO ACTIONABLE ERROR CODES OR MALFUNCTIONS WERE OBSERVED. THE UNIT WAS CALIBRATED AND CONFIRMED TO BE WORKING CORRECTLY. OPERATION HRS = 0.8 / BLOWER HRS = 0.8. IN ADDITION, THE TECHNICIAN REPLACED THE AIR INLET FILTER & PARTICULATE FILTER. THE UNIT WAS DECONTAMINATED BEFORE RETURNING TO THE ENABLE ROOM. THE SOFTWARE VERSION 1.06.10.00. PHILIPS WAS NOTIFIED THAT THE PATIENT PASSED AWAY ON 21/5. THE DEVICE ARRIVED IN NR ON 23/5 WITHOUT THE POWER CORD OR BAG. HEALTHSHARE TO BE INFORMED THAT THESE ACCESSORIES WILL NEED TO BE REPLACED. THIS EVENT WAS REVIEWED BY THE MANUFACTURER CLINICAL ASSESSMENT TEAM AND WAS DETERMINED: THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USER ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. BASED ON AVAILABLE INFORMATION, THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. NO FURTHER ACTION IS REQUIRED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH" A REPORT WILL BE FILED WITH ALL APPLICABLE REGULATORY AGENCIES. AN ATTEMPT TO GATHER ADDITIONAL INFORMATION ON THE CAUSE OF DEATH IS BEING REQUESTED. IF ADDITIONAL INFORMATION IS PROVIDED A SUPPLEMENTAL REPORT WILL BE FILED WITH ALL APPLICABLE REGULATORY AGENCIES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22355785&pc=FRN | 2016493-2025-91345 | 2025-05-13 04:00 | Death | CAREFUSION SD | 2025-06-30 04:00 | FRN | ALARIS SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K133532 |  | 1 | Event Description: IT WAS REPORTED THE CUSTOMER IS REQUESTING A LOG REVIEW FOLLOWING AN EVENT WHERE THE PATIENT HAS DIED. ALTHOUGH DUE DILIGENCE ATTEMPTS HAVE BEEN MADE, AT THIS TIME, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: THE ACTUAL DATE OF DEATH IS UNKNOWN. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22355823&pc=FRN | 2016493-2025-91347 | 2025-05-13 04:00 | Death | CAREFUSION SD | 2025-06-30 04:00 | FRN | ALARIS SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K133532 |  | 1 | Event Description: IT WAS REPORTED THE CUSTOMER IS REQUESTING A LOG REVIEW FOLLOWING AN EVENT WHERE THE PATIENT HAS DIED. ALTHOUGH DUE DILIGENCE ATTEMPTS HAVE BEEN MADE, AT THIS TIME, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: THE ACTUAL DATE OF DEATH IS UNKNOWN. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22355909&pc=DTQ | 8010762-2025-0000275 | 2025-06-25 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-06-30 04:00 | DTQ | HEART LUNG MACHINE | Pumping Stopped | Insufficient Information | K943803 |  | 1 | Event Description: THE EVENT OCCURRED IN INDIA DURING PATIENT TREATMENT. IT WAS REPORTED THAT THE ERROR MESSAGE "HEAD" WAS DISPLAYED ON A HL20 PUMP WHICH CAUSED THE PUMP TO STOP. A GETINGE FIELD SERVICE TECHNICIAN WAS ONSITE FOR FURTHER INVESTIGATION AND FOUND THAT THE PUMP BELTS WERE BROKEN AND NEEDS TO BE REPLACED. PERFUSIONIST SHIFTED MAIN CIRCULATION CIRCUIT TO ANOTHER PUMP. NO HARM OF THE PATIENT HAS BEEN REPORTED DUE TO THE EXCHANGE OF THE PUMP. FURTHER INFORMATION PROVIDED BY THE SSU (SALES AND SERVICE UNIT) ON 2026-06-27 THAT AFTER THE SURGERY THE PATIENT WAS SHIFTED TO ICU AND THE NEXT DAY IN THE MORNING THE PATIENT PASSED AWAY. AT THAT TIME IN THE ICU WHEN THE PATIENT PASSED AWAY, THE PATIENT WAS NOT CONNECTED TO THE HL20. SINCE THE REPORTED ERROR MESSAGE: ¿HEAD¿ STOPPED THE PUMP AND THE PATIENT PASSED AWAY A REPORT IS REQUIRED. COMPLAINT ID: (B)(4).Event Description: COMPLAINT ID: (B)(4). Manufacturer Narrative: THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE INDIA MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, HL 20¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K943803. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HL 20 WITH CATALOG NUMBER 701043262. Manufacturer Narrative: THE EVENT OCCURRED IN INDIA DURING PATIENT TREATMENT. IT WAS REPORTED THAT THE ERROR MESSAGE "HEAD" WAS DISPLAYED ON A HL20 PUMP WHICH CAUSED THE PUMP TO STOP. A GETINGE FIELD SERVICE TECHNICIAN WAS ONSITE FOR FURTHER INVESTIGATION AND FOUND THAT THE PUMP BELTS WERE BROKEN AND NEEDS TO BE REPLACED. PERFUSIONIST SHIFTED MAIN CIRCULATION CIRCUIT TO ANOTHER PUMP. NO HARM OF THE PATIENT HAS BEEN REPORTED DUE TO THE EXCHANGE OF THE PUMP. FURTHER INFORMATION PROVIDED BY THE SSU (SALES AND SERVICE UNIT) ON 2026-06-27 THAT AFTER THE SURGERY THE PATIENT WAS SHIFTED TO ICU AND THE NEXT DAY IN THE MORNING THE PATIENT PASSED AWAY. AT THAT TIME IN THE ICU WHEN THE PATIENT PASSED AWAY, THE PATIENT WAS NOT CONNECTED TO THE HL20. SINCE THE REPORTED ERROR MESSAGE: ¿HEAD¿ STOPPED THE PUMP AND THE PATIENT PASSED AWAY A REPORT IS REQUIRED. A GETINGE FIELD SERVICE TECHNICIAN (FST) WAS ONSITE FOR INVESTIGATION OF THE DEVICE. THE FST REPLACED THE BELT KIT. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. AS CONFIRMED BY THE GETINGE FST, THE BELTS HAVE NEVER BEEN REPLACED ON THAT DEVICE AND WERE OVERDUE FOR REPLACEMENT. ACCORDING TO THE HL20 SERVICE MANUAL, BELTS WHICH ARE FOUND TWISTED, HAVE WEAR, CRACKING, OR SERVICE LIFE OVER 12 MONTHS MUST BE REPLACED. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED ON 2025-07-04 FOR THE PERIOD OF (B)(6) 2017 TO (B)(6) 2025. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED PRODUCT AND FAILURE. THERE IS NO INDICATION THAT MANUFACTURING ISSUES OCCURRED DURING THIS TIME, THUS PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE DEVICE WAS MANUFACTURED ON 2017-10-11. IN ADDITION A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. FOR THE AFFECTED SERIAL NUMBER WITHIN THE TIMEFRAME OF THE SEARCH NO ADDITIONAL COMPLAINT WAS FOUND FOR THE SAME ISSUE. THE INVESTIGATION IS ONGOING. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22355913&pc=BZD | 2518422-2025-047798 | 2023-01-20 05:00 | Death | RESPIRONICS, INC. | 2025-06-30 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO THIRD PARTY SERVICE CENTER AND EVALUATED. EVALUATION RESULT STATED THAT FOAM PARTICLES WERE NOT OBSERVED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22356587&pc=DSQ | 2916596-2025-04262 | 2024-05-24 04:00 | Death | THORATEC CORPORATION | 2025-06-30 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. THE DEVICE WAS NOT BEING RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO HOSPITAL POLICY, THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, REV. B, IS ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22356651&pc=DSQ | 2916596-2025-04258 | 2024-06-03 04:00 | Death | THORATEC CORPORATION | 2025-06-30 04:00 | DSQ | HEARTMATE II LEFT VENTRICULAR ASSIST DEVICE | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY, AND THE CAUSE WAS UNKNOWN. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO HOSPITAL POLICY, THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6), WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISIONS OF THE IFU CAN ALSO BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22356769&pc=PNJ | 2017865-2025-87827 | 2025-06-13 04:00 | Death | ABBOTT | 2025-06-30 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem | Hypersensitivity/Allergic reaction | P150035 |  | 1 | Event Description: A PATIENT WAS BEING IMPLANTED WITH A LEADLESS PACEMAKER SYSTEM. FOLLOWING AN INJECTION OF CONTRAST, THE PATIENT EXPERIENCED LOW OXYGEN LEVELS, WHICH RESULTED IN A NEED TO RESUSCITATE THE PATIENT. AFTER TWENTY MINUTES OF ATTEMPTS, RESUSCITATION WAS STOPPED AND THE PATIENT PASSED AWAY. THE PHYSICIAN ALLEGED THE CAUSE OF DEATH WAS DUE TO AN ALLERGIC REACTION TO THE CONTRAST WHICH RESULTED IN A SPASTIC LUNG. THE PHYSICIAN DOES NOT BELIEVE THE LEADLESS PACEMAKER CAUSE OR CONTRIBUTED TO THE DEATH OF THE PATIENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22357302&pc=DSQ | 2916596-2025-04084 | 2025-05-23 04:00 | Death | THORATEC CORPORATION | 2025-06-30 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Vasodilatation | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO POST OPERATIVE COMPLICATIONS. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED AND WAS DUE TO VASOPLEGIA HEMODYNAMIC INSTABILITY. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22357365&pc=DSQ | 2916596-2025-04329 | 2025-06-10 04:00 | Death | THORATEC CORPORATION | 2025-06-30 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD MULTISYSTEM ORGAN FAILURE AND THEIR FAMILY DECIDED TO WITHDRAW FROM CARE. THE PATIENT PASSED AWAY (B)(6) 2025.Event Description: IT WAS LATER REPORTED THAT THE MULTIORGAN FAILURE WAS LIKELY A NATURAL PROGRESSION OF THE PATIENT'S UNDERLYING DISEASE PROCESS FOR WHICH THEY RECEIVED THE LEFT VENTRICULAR ASSIST DEVICE. IT WAS ALSO REPORTED THAT DEVICE OPERATED AS INTENDED AND THAT THE DEVICE WAS FUNCTIONING APPROPRIATELY. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: IT WAS ORIGINALLY REPORTED THAT THE PATIENT HAD MULTISYSTEM ORGAN FAILURE (MSOF) AND THEIR FAMILY DECIDED TO WITHDRAW FROM CARE. THE PATIENT ULTIMATELY EXPIRED ON (B)(6) 2025. HOWEVER, FURTHER INFORMATION COMMUNICATED BY THE ACCOUNT REVEALED THAT THE MSOF WAS LIKELY A NATURAL PROGRESSION OF THE PATIENT'S UNDERLYING DISEASE PROCESS FOR WHICH THEY RECEIVED THE LEFT VENTRICULAR ASSIST DEVICE (LVAD). THE DEVICE OPERATED AS INTENDED, AND THE PATIENT'S OUTCOME WAS NOT CONSIDERED TO BE DEVICE RELATED. IT WAS COMMUNICATED THAT THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. ALTHOUGH NO DEVICE-RELATED ISSUES WERE REPORTED, SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: H6- 3261 - MULTIPLE ORGAN DYSFUNCTION SYNDROME. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22357551&pc=DTQ | 3013876692-2025-00052 | 2025-06-25 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-06-30 04:00 | DTQ | HEART LUNG MACHINE | Pumping Stopped | Insufficient Information | K943803 |  | 1 | Event Description: THE EVENT OCCURRED IN INDIA DURING PATIENT TREATMENT. IT WAS REPORTED THAT THE ERROR MESSAGE "HEAD" WAS DISPLAYED ON A HL20 PUMP WHICH CAUSED THE PUMP TO STOP. A GETINGE FIELD SERVICE TECHNICIAN WAS ONSITE FOR FURTHER INVESTIGATION AND FOUND THAT THE PUMP BELTS WERE BROKEN AND NEEDS TO BE REPLACED. PERFUSIONIST SHIFTED MAIN CIRCULATION CIRCUIT TO ANOTHER PUMP. NO HARM OF THE PATIENT HAS BEEN REPORTED DUE TO THE EXCHANGE OF THE PUMP. FURTHER INFORMATION PROVIDED BY THE SSU (SALES AND SERVICE UNIT) ON 2026-06-27 THAT AFTER THE SURGERY THE PATIENT WAS SHIFTED TO ICU AND THE NEXT DAY IN THE MORNING THE PATIENT PASSED AWAY. AT THAT TIME IN THE ICU WHEN THE PATIENT PASSED AWAY, THE PATIENT WAS NOT CONNECTED TO THE HL20. SINCE THE REPORTED ERROR MESSAGE: ¿HEAD¿ STOPPED THE PUMP AND THE PATIENT PASSED AWAY A REPORT IS REQUIRED. COMPLAINT ID: (B)(4).Event Description: COMPLAINT ID: (B)(4). Manufacturer Narrative: THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE INDIA MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, HL 20¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K943803. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HL 20 WITH CATALOG NUMBER 701043262. Manufacturer Narrative: THE EVENT OCCURRED IN INDIA DURING PATIENT TREATMENT. IT WAS REPORTED THAT THE ERROR MESSAGE "HEAD" WAS DISPLAYED ON A HL20 PUMP WHICH CAUSED THE PUMP TO STOP. A GETINGE FIELD SERVICE TECHNICIAN WAS ONSITE FOR FURTHER INVESTIGATION AND FOUND THAT THE PUMP BELTS WERE BROKEN AND NEEDS TO BE REPLACED. PERFUSIONIST SHIFTED MAIN CIRCULATION CIRCUIT TO ANOTHER PUMP. NO HARM OF THE PATIENT HAS BEEN REPORTED DUE TO THE EXCHANGE OF THE PUMP. FURTHER INFORMATION PROVIDED BY THE SSU (SALES AND SERVICE UNIT) ON 2026-06-27 THAT AFTER THE SURGERY THE PATIENT WAS SHIFTED TO ICU AND THE NEXT DAY IN THE MORNING THE PATIENT PASSED AWAY. AT THAT TIME IN THE ICU WHEN THE PATIENT PASSED AWAY, THE PATIENT WAS NOT CONNECTED TO THE HL20. SINCE THE REPORTED ERROR MESSAGE: ¿HEAD¿ STOPPED THE PUMP AND THE PATIENT PASSED AWAY A REPORT IS REQUIRED. A GETINGE FIELD SERVICE TECHNICIAN (FST) WAS ONSITE FOR INVESTIGATION OF THE DEVICE. THE FST REPLACED THE BELT KIT. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. AS CONFIRMED BY THE GETINGE FST, THE BELTS HAVE NEVER BEEN REPLACED ON THAT DEVICE AND WERE OVERDUE FOR REPLACEMENT. ACCORDING TO THE HL20 SERVICE MANUAL, BELTS WHICH ARE FOUND TWISTED, HAVE WEAR, CRACKING, OR SERVICE LIFE OVER 12 MONTHS MUST BE REPLACED. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED ON 2025-07-04 FOR THE PERIOD OF 2017-10-11 TO 2025-06-25. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED PRODUCT AND FAILURE. THERE IS NO INDICATION THAT MANUFACTURING ISSUES OCCURRED DURING THIS TIME, THUS PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE DEVICE WAS MANUFACTURED ON 2017-10-11. IN ADDITION A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. FOR THE AFFECTED SERIAL NUMBER WITHIN THE TIMEFRAME OF THE SEARCH NO ADDITIONAL COMPLAINT WAS FOUND FOR THE SAME ISSUE. THE INVESTIGATION IS ONGOING. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22357786&pc=OZD | 1220648-2025-29877 | 2025-06-10 04:00 | Death | ABIOMED, INC. | 2025-06-30 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemolysis; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA 5.5 HAD OOZING AT THE ACCESS SITE. ONE UNIT OF BLOOD WAS TRANSFUSED AND HEPARIN WAS WITHHELD. ADDITIONALLY, THE PATIENT'S LABS WERE HEMOLYZED THE PATIENT WAS TRANSFERRED TO PALLIATIVE CARE AND EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS.¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22358138&pc=OZD | 1220648-2025-29881 | 2025-06-11 04:00 | Death | ABIOMED, INC. | 2025-06-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THERE WAS COMPLETE LOSS OF PULSATILITY WITH EACH LESION. DUE TO LACTATE AND DEPENDENCY ON LEVOPHED, THE DECISION WAS MADE TO LEAVE THE IMPELLA OVERNIGHT TO COOL THE PATIENT OFF WITH POTENTIAL EXPLANT THE NEXT DAY. THE PEEL AWAY WAS REMOVED AND REPO WAS ADVANCED. THE DEVICE WAS SECURED IN PLACE. SOME BLEEDING WAS NOTED AROUND REPO. PERCLOSE WAS TIGHTENED AND SURGICEL WAS APPLIED TO THE SITE. GOOD PULSE PRESENT. DURING SUPPORT, IT WAS REPORTED THAT THE PATIENT WAS STABLE. HOWEVER, THE PATIENT HAD AN UNEXPECTED ARRYTHMIA AND WENT ASYSTOLIC AND EXPIRED ON SUPPORT. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22358276&pc=OZD | 1220648-2025-29879 | 2025-06-10 04:00 | Death | ABIOMED, INC. | 2025-06-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Ischemia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED A CEREBRAL VASCULAR ACCIDENT WHICH WAS CONFIRMED WITH A CT. SUBSEQUENTLY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION.¿ Manufacturer Narrative: THE INVESTIGATION OF CVA/STROKE AND LIMB ISCHEMIA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. LIMB ISCHEMIA BASED ON INVESTIGATION FINDINGS/REVIEW OF THE CLINICAL DATA.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22358417&pc=LWS | 2124215-2025-43129 | 2025-05-09 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-30 04:00 | LWS | INOGEN EL ICD DR | Pacing Problem; Inaccurate Synchronization; Under-Sensing | Arrhythmia | P960040/S306 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS REPORTED TO HAVE PASSED AWAY ON (B)(6) 2025. A REVIEW OF ICD DATA WAS REQUESTED BY THE FIELD TO DETERMINE IF THERE WAS ANY INFORMATION STORED IN THE ICD MEMORY AROUND THE TIME THE PATIENT PASSED. REVIEW OF DEVICE DATA BY BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) CONFIRMED ON THE DATE OF DEATH, UNDERSENSING WAS OBSERVED ON THE RIGHT VENTRICULAR (RV) CHANNEL, CAUSING INAPPROPRIATE PACING FROM THE ICD, DESPITE THERE STILL BEING INTRINSIC ACTIVITY. THIS SEEMED TO TRIGGER AN ARRYTHMIA IN THE RV, WHICH ACCELERATED SIGNIFICANTLY, AND INCLUDED VARYING AMPLITUDE MORPHOLOGY MEASUREMENTS. THERE WAS COMPETITION BETWEEN THE PACING BY THE ICD AND THE PATIENT'S INTRINSIC SIGNALS IN THE RV. THE DEVICE CONTINUED TO UNDERSENSE THE RV ACTIVITY, THUS CAUSING MORE INTERMITTENT PACING. AFTER ANTI-TACHYCARDIA PACING AND SHOCK DELIVERY, THE ICD CONTINUED TO PACE BECAUSE OF ARRHYTHMIA UNDERSENSING. AT THE SAME TIME, THE UNDERLYING VENTRICULAR ARRHYTHMIA ONGOING WAS ONGOING. BASED ON THE DATA AVAILABLE, TS CONCLUDED THE UNDERSENSING LIKELY TRIGGERED THE VENTRICULAR ARRYTHMIA, WHICH EVENTUALLY CAUSED THE PATIENT'S DEMISE. THE LEADS TRENDS PRIOR TO THE PATIENT'S DEATH WERE STABLE AND IN-RANGE. ALL EVIDENCE INDICATES THE ICD WAS EXPLANTED FROM THE BODY AND WILL NOT BE RETURNED FOR ANALYSIS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22359056&pc=PYX | 1220648-2025-29890 | 2025-06-12 04:00 | Death | ABIOMED, INC. | 2025-06-30 04:00 | PYX | IMPELLA RP FLEX | Device Sensing Problem | Cardiovascular Insufficiency; No Clinical Signs, Symptoms or Conditions | P170011 |  | 1 | Event Description: THE COMPLAINANT HAD AN IMPELLA RP FLEX BEING ATTACHED TO AN AUTOMATED IMPELLA CONTROLLER (AIC) FOR USE IN A PATIENT ADMITTED IN WITH ACUTE MYOCARDIAL INFARCTION/CARDIOGENIC SHOCK. WHEN THE PUMP WAS CONNECTED TO THE AIC, THE OPTICAL SIGNAL WAS NOT FUNCTIONING APPROPRIATELY SO THE PUMP WAS REPLACED. THE IMPELLA RP FLEX WAS NEVER INSERTED INTO THE BODY, JUST CONNECTED TO THE AIC. THE DEVICE WAS STILL IN PLASTIC PACKAGING WHEN CONNECTED AND HAD NOT BEEN HANDLED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: H.6 HEALTH EFFECT - CLINICAL CODE 4582 WAS INCORRECTLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29890. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22359793&pc=PNJ | 2017865-2025-90558 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-30 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC FOR LEADLESS PACEMAKER IMPLANT PROCEDURE. DURING THE IMPLANT, THE PATIENT SUSTAINED A BLOOD PRESSURE DROP AND TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) REVEALED PERICARDIAL EFFUSION. EMERGENT CARDIOPULMONARY RESUSCITATION WAS PERFORMED BUT WAS UNSUCCESSFUL. THE PATIENT WAS PRONOUNCED DECEASED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22359922&pc=BZD | 2518422-2025-106908 | 2025-06-17 04:00 | Death | RESPIRONICS, INC. | 2025-06-30 04:00 | BZD | PHILIPS HOME CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED A CALL FROM THE PATIENT'S SON STATING THAT THE PATIENT PASSED AWAY YESTERDAY, HE ALSO STATED THAT HE HAD AN AEROSOL DELIVERY SYSTEM, HE ALSO STATED THAT HIS FATHER DID NOT PASS AWAY DUE TO THE DEVICE, HE WAS ON HOSPICE AND NOT USING THE DEVICE. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22359996&pc=OZD | 1220648-2025-29893 | 2025-06-12 04:00 | Death | ABIOMED, INC. | 2025-06-30 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT EXPERIENCED SUCTION ALARMS DUE TO LACK OF VOLUME. 3L HAD ALREADY BEEN ADMINISTERED. IT WAS RECOMMENDED TO REDUCE THE P-LEVEL AS FAR AS NECESSARY, ADMINISTER VOLUME AND INCREASE THE P-LEVEL AGAIN. IN ADDITION, THE PUNCTURE SITE WAS BLEEDING HEAVILY, AND THE PATIENT RECEIVED SEVERAL BLOOD BAGS. THE VASCULAR SURGEON CIRCUMCISED THE PUNCTURE SITE AND THUS STOPPED THE BLEEDING. ADDITIONALLY, THE PATIENT HAD A PERICARDIAL EFFUSION WITH 100 ML/H ASPIRATE. THEY HAD ALREADY ADMINISTERED 5 BLOOD BAGS AND 5 FFP`S. ULTIMATELY THE PATIENT DETERIORATED CARDIOPULMONARY OVERNIGHT. THE PATIENT EXPERIENCED A SILENT ASPIRATION AND BECAME INCREASINGLY WORSE, WHICH IS WHY HE WAS INTUBATED. DESPITE THE IMPELLA CP AND CATECHOLAMINE ADMINISTRATION, HE CONTINUED TO DETERIORATE, AND THE DOCTORS DISCONTINUED THE TREATMENT, AND THE PATIENT PASSED AWAY. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMART ASSIST SYSTEM INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22360004&pc=PNJ | 2017865-2025-90557 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-30 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED IN CLINIC FOR LEADLESS PACEMAKER IMPLANT PROCEDURE. DURING THE IMPLANT, THE PATIENT SUSTAINED A BLOOD PRESSURE DROP AND TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) REVEALED PERICARDIAL EFFUSION. EMERGENT CARDIOPULMONARY RESUSCITATION WAS PERFORMED BUT WAS UNSUCCESSFUL. THE PATIENT WAS PRONOUNCED DECEASED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22360137&pc=OZD | 1220648-2025-29903 | 2025-06-11 04:00 | Death | ABIOMED, INC. | 2025-06-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Great Vessel Perforation | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT IMPLANTED WITH AN IMPELLA AND EXHIBITED A PERFORATION OF THE LEFT VENTRICLE AND A HEMOPNEUMOTHORAX. THE PATIENT EXPIRED FROM THE INJURY. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER AS IT WAS NOT EXPLANTED FROM THE DECEASED PATIENT'S BODY. THE PUMP WAS TURNED OFF BY THE MEDICAL DOCTOR. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST SYSTEM WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS & CAUTIONS: WARNINGS, SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22360167&pc=LWR | 2135147-2025-03554 | 2025-04-25 04:00 | Death | ABBOTT MEDICAL | 2025-06-30 04:00 | LWR | EPIC MAX VALVE (AORTIC) | Adverse Event Without Identified Device or Use Problem |  | P040021 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 19MM EPIC MAX VALVE WAS SELECTED FOR IMPLANT IN A PATIENT WHO HAD UNDERGONE A PRIOR TAVI PROCEDURE. POST-PROCEDURE, THE PATIENTS CORONARY ARTERY WAS NOTED TO HAVE BECOME OCCLUDED. THE PATIENT PASSED AWAY AS A RESULT. THE PHYSICIAN BELIEVES THAT THE ARTERY WAS BLOCKED DUE TO OVERSIZING OF THE BIOPROSTHETIC AND BECAUSE THE TAVI WAS REMOVED. IT WAS ALSO NOTED AS A POSSIBILITY THAT TISSUE FRAGMENTS MAY HAVE ENTERED THE CORONARY ARTERY DURING THE REMOVAL OF THE TAVI. THE PHYSICIAN BELIEVES THAT THE EPIC DEVICE FUNCTIONED AS INTENDED AND THAT THE DEATH WAS NOT RELATED TO A PRODUCT MALFUNCTION. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22360496&pc=BZD | 2518422-2025-106905 | 2025-06-23 04:00 | Death | RESPIRONICS, INC. | 2025-06-30 04:00 | BZD | PHILPS HOME CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DEATH OF A PATIENT. THE MANUFACTURER RECEIVED A CALL FROM THE PATIENT'S DAUGHTER STATING HER MOTHER PASSED AWAY IN 2017. THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION. THE PATIENT'S DAUGHTER WAS UNABLE TO PROVIDE THE SERIAL NUMBER OR ANY DEVICE INFORMATION. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22360610&pc=DSQ | 2916596-2025-04334 | 2025-05-26 04:00 | Death | THORATEC CORPORATION | 2025-06-30 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE PATIENT OUTCOME; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED BY THE ACCOUNT. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22360682&pc=LWS | 2649622-2025-17652 | 2025-04-14 04:00 | Death | MEDTRONIC PUERTO RICO VILLALBA | 2025-06-30 04:00 | LWS | SPRINT QUATTRO SECURE MRI SURESCAN | Impedance Problem | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD TRIGGERED AN ALERT FOR SUPERIOR VENA CAVA IMPEDANCE. THE RV LEAD REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE SUPERIOR VENA CAVA (SVC) WAS OFF. IT WAS ALSO NOTED THAT THE PATIENT PASSED AWAY APPROXIMATELY TWO AND A HALF MONTHS AFTER THE SVC IMPEDANCE WARNING, A CAUSE OF DEATH WAS REQUESTED BUT HAS NOT BEEN PROVIDED. Manufacturer Narrative: CONTINUATION OF D10: 419688 LEAD IMPLANTED: (B)(6) 2012; 5076-52 LEAD IMPLANTED: (B)(6) 2010. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22360804&pc=DSQ | 2916596-2025-04333 | 2025-06-10 04:00 | Death | THORATEC CORPORATION | 2025-06-30 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RENAL FAILURE. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED. WHETHER THE OUTCOME WAS DEVICE OR THERAPY RELATED WAS NOT PROVIDED BY THE CUSTOMER. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RENAL FAILURE AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22360815&pc=DSQ | 2916596-2025-04383 | 2025-06-12 04:00 | Death | THORATEC CORPORATION | 2025-06-30 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AT THE HOSPITAL DUE TO CARDIAC ARREST. DEATH WAS NOT SUSPECTED TO BE DEVICE RELATED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22360885&pc=OZD | 1220648-2025-29905 | 2025-06-12 04:00 | Death | ABIOMED, INC. | 2025-06-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Perforation of Vessels; Cardiac Perforation | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, DURING PCI PROCEDURE, THERE WAS A DISTAL WIRE PERFORATION IN THE POSTEROLATERAL BRANCH (PLB) OF THE HEART WHICH OCCURRED AFTER THROMBECTOMY OF THE DISTAL RIGHT CORONARY ARTERY (RCA). BALLOON TAMPONADE FOR SEVERAL MINUTES DID NOT CLOSE PERFORATION. TWO COILS PLACED TO CLOSE PERFORATION, PATIENT LOST POSTEROLATERAL ARTERY (PLA) AND POSTERIOR DESCENDING ARTERY (PDA) BRANCH. ANGIOGRAPHICALLY THE PLB PERFORATION WAS CLOSED. PERICARDIUM WAS FILLING WITH BLOOD. RIGHT VENTRICLE AND LEFT VENTRICLE ACCESSED FOR PERFORATIONS BUT NO OBVIOUS SOURCE. TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE), TRANSTHORACIC ECHOCARDIOGRAM (TTE) AND ANGIOGRAMS PERFORMED BUT WERE UNABLE TO LOCATE AND THE PATIENT EXPIRED IN THE CATH LAB. Manufacturer Narrative: THE INVESTIGATION OF VENTRICLE PERFORATION HAS BEEN COMPLETED. THE IMPELLA CP PUMP WAS RETURNED. NO PHYSICAL ABNORMALITIES OR DAMAGES FOUND WITH PUMP. THE ROOT CAUSE OF THE VENTRICLE PERFORATION ISSUE WAS NOT DETERMINED DUE TO INSUFFICIENT CLINICAL INFORMATION AND UNKNOWN RELATION TO IMPELLA. A.4 WEIGHT WAS ADDED AS IT WAS INADVERTENTLY OMITTED ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29905. B.7 INFORMATION WAS ADDED THAT WAS INADVERTENTLY OMITTED ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29905. H.6 HEALTH EFFECT - CLINICAL CODE 2135 WAS ENTERED INCORRECTLY ON INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29905. A NEW CODE WAS ADDED. Manufacturer Narrative: THE IMPELLA DEVICE WAS RETURNED, AND AN EVALUATION IS UNDERWAY. UPON COMPLETION OF OUR ANALYSIS, A SUPPLEMENTAL REPORT WILL BE FILED. INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: IMPELLA CP WITH SMARTASSIST SYSTEM AND IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE. SECTION: WARNINGS & CAUTIONS: WARNINGS SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES ¿TO REDUCE THE RISK OF CARDIAC INJURY (INCLUDING VENTRICULAR PERFORATION), PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY IN THE SETTINGS OF ACUTE INFARCTION WITH TISSUE NECROSIS.¿ ¿TO REDUCE THE RISK OF VASCULAR INJURY, PHYSICIANS SHOULD EXERCISE CAUTION WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX PERIPHERAL VASCULAR ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: UNREPAIRED ABDOMINAL AORTIC ANEURYSM, SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM, DISSECTION OF THE ASCENDING/ TRANSVERSE/DESCENDING AORTA, CHRONIC ANATOMICAL CHANGES IN THE RELATIONSHIP OF THE AORTA/AORTIC VALVE/VENTRICULAR ALIGNMENT, SIGNIFICANT MOBILE ATHEROMATOUS DISEASE IN THE THORACIC OR ABDOMINAL AORTA OR PERIPHERAL VESSELS.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING A RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY IN THE LEFT VENTRICLE AFTER CPR WITH ECHOCARDIOGRAPHY GUIDANCE. IMPELLA CP WITH SMARTASSIST SYSTEM AND IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & IMPELLA RP SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER & RP FLEX WITH SMARTASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: WARNINGS & CAUTIONS: CAUTIONS ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22361777&pc=OZD | 1220648-2025-29907 | 2025-06-12 04:00 | Death | ABIOMED, INC. | 2025-06-30 04:00 | OZD | IMPELLA 5.5 | Increase in Pressure; Infusion or Flow Problem | Unspecified Infection; Gastrointestinal Hemorrhage; No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THERE WAS PUMP SATURATION. THERE WAS A CHANGE IN THE WAVEFORMS AND EQUALIZATION IN MAX AND MIN FLOWS. IT WAS RECOMMENDED TO EXPLANT THE PUMP DUE TO POTENTIAL PUMP STOP. THE STAFF DECIDED TO CONTINUE SUPPORT, THE PUMP WAS TURNED TO P 9 AND FLOWS EQUALIZED DEVICE THE WAVEFORMS RETURNED TO NORMAL AND FLOWS BEGAN TO SETTLE TO AN AVERAGE OF 5.3 MAX 6.0 MIN 4.7. IT WAS STILL RECOMMENDED TO WEAN AND EXPLANT, HOWEVER SUPPORT CONTINUED.Event Description: ADDITIONALLY, THE PATIENT HAD AN INFECTION. CULTURES WERE POSITIVE AND ANTIBIOTICS WERE STARTED. IT IS UNKNOWN IF THE IMPELLA CONTRIBUTED TO THE INFECTION. FURTHERMORE, THE PATIENT HAD GASTROINTESTINAL BLEED. BLOOD PRODUCTS WERE GIVEN. MOREOVER, HIGH PURGE PRESSURE WAS NOTED. TISSUE PLASMA ACTIVATOR WAS ADDED TO THE PURGE. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION HAS BEEN COMPLETED. NO PRODUCT WAS RETURNED. CLINICAL DETAILS STATE THERE WAS A CHANGE IN WAVEFORMS AND EQUALIZATION IN MAX AND MIN FLOWS ON THE SECOND DAY OF SUPPORT. PUMP WAS VERIFIED TO BE IN CORRECT POSITION. AFTER ~2 HOURS, PUMP WAVEFORMS RETURNED TO NORMAL. DATA LOG REVIEW SHOWED A MOTOR CURRENT SPIKE WITH A STEP UP IN MOTOR CURRENT, SPEED DEVIATION, AND IMMEDIATE SATURATED FLOWS. NO PRODUCT WAS RETURNED. THE ROOT CAUSE OF THE SATURATED FLOW WAS MOST LIKELY BIOMATERIAL INGESTION SINCE DATA LOG REVIEW SHOWED A MOTOR CURRENT SPIKE AND INCREASE AS THE SATURATED FLOWS BEGAN. THE FAILURE MODE THROMBUS WAS ADDED SINCE DATA LOG REVIEW INDICATED BIOMATERIAL INGESTION RELATED TO THE SATURATED FLOW. THE ROOT CAUSE OF THE THROMBUS WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. CLINICAL DETAILS STATED THAT PURGE PRESSURE WAS IN THE 1100S MMHG ON (B)(6) 2025. THEY ADMINISTERED TPA AND THE PURGE PRESSURE DECREASED BACK TO THE 700S MMHG. DATA LOG REVIEW SHOWED A PRE-DISTURBANCE WHERE THE PURGE PRESSURE INCREASED FROM ~600 TO ~800 MMHG FOR ABOUT 11 HOURS BEFORE RESOLVING. AFTER THAT RESOLVED, THE PURGE PRESSURE GRADUALLY RISES FROM ~600 MMHG TO >1100 MMHG OVER ABOUT 24 HOURS. THE HIGH PURGE PRESSURE IS SUSTAINED FOR ~8 HOURS BEFORE GRADUALLY RESOLVING. THERE WERE MULTIPLE LONG BAG/CASSETTE CHANGES THROUGHOUT THIS TIME. ALTHOUGH NONE OF THESE OCCURRED AROUND THE TIMES THAT THE PURGE PRESSURE STARTED TO INCREASE, THEY COULD HAVE CONTRIBUTED TO A CONTINUED RISE. THE ROOT CAUSE OF THE HIGH PURGE PRESSURE WAS MOST LIKELY BIOMATERIAL BUILDUP DUE TO THE GRADUAL INCREASE OF PURGE PRESSURE OVER ~24 HOURS. CLINICAL DETAILS MENTION AN INFECTION TREATED BY ANTIBIOTICS DURING THE CASE. PRODUCT WAS NOT RETURNED AND LOGS ARE NOT APPLICABLE. THE ROOT CAUSE OF THE INFECTION/SEPSIS WAS UNABLE TO BE DETERMINED DUE TO INSUFFICIENT CLINICAL DETAILS. CLINICAL DETAILS MENTION A GI BLEED DURING THE CASE. THE PATIENT WAS ON SYSTEMIC HEPARIN FOR PART OF THE CASE. PRODUCT WAS NOT RETURNED AND LOGS ARE NOT APPLICABLE. THE CAUSE OF THE VASCULAR DAMAGE WAS NOT DETERMINED SINCE INSUFFICIENT CLINICAL DETAILS WERE PROVIDED AND NO PRODUCT WAS RETURNED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. Manufacturer Narrative: B1 REVISED TO REFLECT BOTH ADVERSE EVENT AND PRODUCT PROBLEM. B2 UPDATED TO REFLECT INTERVENTION, DEATH, AND DATE OF DEATH. B5 UPDATED WITH ADDITIONAL INFORMATION RECEIVED FROM THE CLINICAL SITE. D6B ADDED. H1 REVISED TO REFLECT DEATH. H6 UPDATED WITH ADDITIONAL FAILURE MODES. H10 INSTRUCTIONS FOR USE ADDED. THE IMPELLA PUMP WAS NOT RETURNED FOR INVESTIGATION. SHOULD NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. INSTRUCTIONS FOR USE. IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK & CP WITH SMARTASSIST SECTION: TABLE 3.2 IMPELLA CATHETER COMPONENTS ¿THE INFUSION FILTER PREVENTS BACTERIAL CONTAMINATION AND AIR FROM ENTERING THE PURGE LUMEN.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿ IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING. SECTION: WARNINGS & CAUTIONS: CAUTIONS. ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH KNOWN OR SUSPECTED UNREPAIRED ABDOMINAL AORTIC ANEURYSM OR SIGNIFICANT DESCENDING THORACIC AORTIC ANEURYSM OR DISSECTION OF THE ASCENDING, TRANSVERSE, OR DESCENDING AORTA.¿ CAUTION: ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS. SECTION: PRE-SUPPORT EVALUATION. SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER. SECTION: DIRECT AORTIC INSERTION. SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER IN PATIENTS WITH COMPLEX ANATOMY. THIS INCLUDES PATIENTS WITH KNOWN OR SUSPECTED: DECREASED VENTRICULAR CAVITY SIZE, VENTRICULAR ANEURYSMS, THIN-WALLED VENTRICLES DUE TO CHRONIC DILATION, CONGENITAL HEART DISEASE, OR COMPROMISED CARDIAC TISSUE QUALITY.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ IMPELLA 5.5 WITH SMARTASSIST. SECTION: USING THE AUTOMATED IMPELLA CONTROLLER WITH THE IMPELLA CATHETER ¿UNRESOLVED PURGE PRESSURE HIGH ALARM IF NOT RESOLVED BY THE RECOMMENDATIONS PROVIDED, HIGH PURGE PRESSURE¿WHICH TRIGGERS THE ¿PURGE PRESSURE HIGH¿ ALARM MESSAGE¿COULD BE AN INDICATION OF A KINK IN THE IMPELLA CATHETER. IN THIS CASE, THE MOTOR IS NO LONGER BEING PURGED AND MAY EVENTUALLY STOP. CLINICIANS SHOULD MONITOR MOTOR CURRENT AND CONSIDER REPLACING THE IMPELLA CATHETER WHENEVER A RISE IN MOTOR CURRENT IS SEEN.¿ E1 WAS ERRONEOUSLY SELECTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29907.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22361936&pc=NPT | 2025587-2025-04638 | 2025-06-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-30 04:00 | NPT | EVOLUT PRO PLUS VALVE | Patient Device Interaction Problem | Cardiac Arrest; Low Blood Pressure/ Hypotension | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE WITH THE EVOLUT 26 VALVE, PRE-DILATION WAS PERFORMED WITH A NON-COMPLIANT BALLOON, SIZE 18 X 40. FOLLOWING IMPLANTATION OF THE VALVE, THE PATIENT EXPERIENCED HEMODYNAMIC INSTABILITY WITH LOW BLOOD PRESSURE AND SUBSEQUENTLY SUFFERED CARDIORESPIRATORY ARREST. RESUSCITATION MANEUVERS WERE PERFORMED. ECHOCARDIOGRAPHY CONFIRMED LEFT CORONARY OCCLUSION, AND CORONARY CATHETERIZATION WAS CONDUCTED AND A CHIMNEY STENT WAS IMPLANTED. AN INTRA-AORTIC BALLOON WAS INSTALLED, BUT ALL MANEUVERS WERE UNSUCCESSFUL, AND DEATH WAS CONFIRMED IN THE OPERATING ROOM. THE CAUSE OF DEATH WAS REPORTED AS CORONARY OCCLUSION ORIGINATING FROM A NATIVE LEAFLET THAT DETACHED DURING PRE-DILATION WITH THE CRYSTAL BALLOON.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH WAS THE LEFT CORONARY OCCLUSION, AND THE VALVE CAN BE EXCLUDED AS A CONTRIBUTING CAUSE. Manufacturer Narrative: IMAGE REVIEW: PRE-IMPLANT COMPUTED TOMOGRAPHY (CT) IMAGING PROVIDED FROM (B)(6) 2025. CT CONTAINS MOTION ARTIFACT/BLURRING ¿ PLEASE CONSIDER MEASUREMENTS ESTIMATES. ANNULUS PERIMETER IS 65.3 MM AND PERIMETER DERIVED MEASURES 20.8 MM SUGGESTING A 26 MM EVOLUT PER SIZING MATRIX. AORTIC VALVE APPEARS MILDLY CALCIFIED. SINUS OF VALSALVA (SOV) MEAN DIAMETER IS 26 MM WITH RCC AND LCC TO COMMISSURES MEASURING 27 MM. SINUS HEIGHTS ARE WITHIN RECOMMENDED RANGE. AORTIC ROOT ANGLE IS 39°. REPORT REVIEW: CASE REPORT PROVIDED FROM (B)(6) 2025. ANNULUS PERIMETER IS 65.3 MM WITH A PERIMETER DERIVED DIAMETER OF 20.8 MM SUGGESTING 26 MM EVOLUT PER SIZING MATRIX. CALCIFICATION SEEN IN ALL LEAFLET BORDERS. SOV MEAN DIAMETER IS 25.7 MM WITH ALL SINUS TO COMMISSURES MEASURING 27 MM. SINUS HEIGHTS ARE WITHIN RECOMMENDED RANGE. AORTIC ROOT ANGLE IS 39°. UPDATED DATA: B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: H2 H3 H11 IMAGE REVIEW: INTRAPROCEDURAL FLUOROSCOPIC IMAGES WERE PROVIDED FOR REVIEW OF THE EVENT. FLUOROSCOPIC VALVE LOAD INSPECTION WAS NOT PERFORMED THUS IT IS NOT POSSIBLE TO VALIDATE A GOOD LOAD. A PRE BALLOON AORTIC VALVULOPLASTY WAS PERFORMED PRIOR TO THE VALVE IMPLANT. THE VALVE WAS DEPLOYED IN THE LEFT ANTERIOR OBLIQUE (LAO) VIEW AT A DEPTH OF APPROXIMATELY 6 MILLIMETER (MM) ON THE NON-CORONARY CUSP AND 8MM ON THE LEFT CORONARY CUSP WITH EVIDENCE OF FLOW TO THE LEFT CORONARY ARTERY. AS STATED IN THE INSTRUCTIONS FOR USE (IFU), THE RECOMMENDED TARGET DEPTH IS 3 MM RELATIVE TO THE VALVE ANNULUS. IF THE IMPLANT DEPTH IS 1 MM OR >5 MM, CONSIDER RECAPTURE. IN THIS CASE, THE VALVE WAS RELEASED, AND THE SUBSEQUENT ANGIOGRAM REVEALED THE ABSENCE OF FLOW TO THE LEFT CORONARY ARTERY. IT WAS REPORTED THAT CARDIOPULMONARY RESUSCITATION EFFORTS WERE INITIATED. CORONARY RE-ACCESS WAS ACHIEVED WITH A CATHETER AND A GUIDEWIRE, AND A CORONARY STENT WAS IMPLANTED. FINAL ANGIOGRAM SHOWED EVIDENCE THAT FLOW WAS RESTORED; HOWEVER, IT WAS REPORTED THAT THE PATIENT EVENTUALLY DIED. EVIDENCE SUPPORTS THAT CORONARY OCCLUSION WAS THE MOST LIKELY CAUSE OF DEATH. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: D-EVPROP23-29 (LOT: 0012490476); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID: L-EVPROP23-29 (LOT: 0012399243); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID: GWBC30 (LOT: 92106989); PRODUCT TYPE: 0193-GUIDEWIRE; IMPLANT DATE N/A; EXPLANT DATE N/A. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22363428&pc=DSQ | 2916596-2025-04089 | 2025-05-31 04:00 | Death | THORATEC CORPORATION | 2025-06-30 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS DECEASED ON (B)(6) 2025. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV D, IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22363476&pc=OZD | 1220648-2025-29912 | 2025-06-12 04:00 | Death | ABIOMED, INC. | 2025-06-30 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT THE PATIENT IMPLANTED WITH AN IMPELLA CP EXPERIENCED BLEEDING AT THE RIGHT FEMORAL INSERTION SITE. THE SITE WAS RE-DRESSED WITH ANGLE MATCH DRESSING AND SUTURE WAS APPLIED. THE FAMILY WITHDREW CARE, AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF THE ACCESS SITE BLEEDING ISSUE WAS MOST LIKELY USE RELATED, SITE RE-DRESSED WITH ANGLE MATCH DRESSING AND OOZING STOPPED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: THE INSTRUCTIONS FOR USE FOR THE IMPELLA CP WITH SMART ASSIST SYSTEM STATES THE FOLLOWING: SECTION: GENERAL PATIENT CARE CONSIDERATIONS ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22363847&pc=OZD | 1220648-2025-29913 | 2025-06-11 04:00 | Death | ABIOMED, INC. | 2025-06-30 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PUMP ACCESS SITE WAS OOZING AND WAS REDRESSED, SUTURES WERE TIGHTENED, AND BLOOD WAS GIVEN FOR DROPPED HEMOGLOBIN. IN ADDITION, THE PUMP HAD SUCTION DESPITE ALL TROUBLESHOOTING. THE FAMILY MADE THE CHOICE TO WITHDRAW CARE AS PROGNOSIS WAS IN QUESTION WITH POSSIBLE ENCEPHALOPATHY AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF ACCESS SITE BLEEDING AND LOW PUMP FLOW HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. LOW PUMP FLOW: DATA LOG REVIEWED: SUCTIONS ALARMS OCCURRED AFTER IMPELLA INSERTION ABOVE P-4, DROP IN FLOW OBSERVED AT P4 & 5. LOOKS FLOW RESOLVED AFTER BLOOD INFUSION, NO PLACEMENT SIGNAL OR POSITIONING ISSUE ALARMS SEEN. NO MC SPIKE OUTSIDE OF P-LEVEL CHANGES OCCURRED. THE CAUSE OF LOW PUMP FLOW MOST LIKELY WAS PATIENT CONDITION RELATED SINCE AFTER BLOOD INFUSION WAS ABLE TO INCREASE FLOWS. ACCESS SITE BLEEDING: THE CAUSE OF BLEEDING (OOZING) ISSUE MOST LIKELY WAS USE RELATED SINCE RETIGHTENED PERCLOSE SUTURES WITHIN SYRINGE/STOPCOCK AND SITE REDRESSED REQUIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: GENERAL PATIENT CARE CONSIDERATIONS: ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT, USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES), ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22364436&pc=NPT | 2015691-2025-05292 | 2025-06-10 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-30 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem | Rupture | P140031 |  | 1 | Event Description: AS REPORTED, THE CASE INVOLVED THE IMPLANTATION OF A 23MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION VIA A TRANSFEMORAL APPROACH. DURING THE PROCEDURE, DIFFICULTY WAS ENCOUNTERED IN CROSSING THE NATIVE ANNULUS DUE TO A HORIZONTAL AORTA (76 DEGREE). DESPITE THIS CHALLENGE, THE VALVE WAS SUCCESSFULLY DEPLOYED. HOWEVER, AN ANNULAR RUPTURE WAS IDENTIFIED. PERICARDIAL DRAINAGE WAS PERFORMED, AND THE PATIENT WAS CONVERTED TO SURGERY FOR A MINI-STERNOTOMY. UNFORTUNATELY, THE PATIENT EXPIRED. UPON REMOVAL, THE EDWARDS DEVICES SHOWED NO SIGNS OF DAMAGE. THE CAUSE OF DEATH WAS DETERMINED TO BE THE ANNULAR RUPTURE. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. AORTIC DISSECTION OR RUPTURE MAY OCCUR WHEN MULTIPLE ATTEMPTS ARE MADE TO CROSS THE STENOTIC NATIVE VALVE, AND/OR WHEN EXCESSIVE FORCE IS USED. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). THE THV TRAINING MANUALS GUIDE TO FACILITATE SAFE CROSSING OF THE NATIVE OR FAILING VALVE, INCLUDING CAMERA PROJECTIONS, HANDLING DURING ADVANCEMENT, AND TROUBLESHOOTING TECHNIQUES IF DIFFICULTY IS ENCOUNTERED. AS STATED, EXCESSIVE FORCE SHOULD NOT BE USED WHEN THE DEVICE HAS DIFFICULTY CROSSING THE STENOTIC VALVE. ADDING TENSION TO THE WIRE, PULLING BACK THE SYSTEM TO RE-ORIENT THE VALVE, AS NEEDED, AND TORQUING THE FLEX CATHETER MAY HELP SOLVE THE PROBLEM. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS INDICATE DURING THE PROCEDURE; DIFFICULTY WAS ENCOUNTERED IN CROSSING THE NATIVE ANNULUS DUE TO A HORIZONTAL AORTA (76 DEGREE). DESPITE THIS CHALLENGE, THE VALVE WAS SUCCESSFULLY DEPLOYED. HOWEVER, AN ANNULAR RUPTURE WAS IDENTIFIED AND THE PATIENT EXPIRED. PER MEDICAL OPINION, THE CAUSE OF DEATH WAS DETERMINED TO BE THE ANNULAR RUPTURE. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22364516&pc=DSQ | 3007042319-2025-01388 | 2025-06-10 04:00 | Death | HEARTWARE, INC. | 2025-06-30 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Respiratory Failure | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT THE VENTRICULAR ASSIST DEVICE (VAD) PATIENT WAS ADMITTED FOR RESPIRATORY FAILURE TREATMENT. THE PATIENT SUBSEQUENTLY EXPIRED, AND THE CAUSE OF DEATH WAS RESPIRATORY FAILURE. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT EXPIRED DUE TO RESPIRATORY FAILURE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, RESPIRATORY DYSFUNCTION AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE WAS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22364729&pc=MKJ | 1220908-2025-02319 | 2025-06-02 04:00 | Death | ZOLL MEDICAL CORPORATION | 2025-06-30 04:00 | MKJ | X SERIES | Failure to Charge | Ventricular Fibrillation | P160022 |  | 1 | Event Description: COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A 65-YEAR-OLD MALE PATIENT, THE DEVICE FAILED TO CHARGE. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. Manufacturer Narrative: ZOLL MEDICAL CORPORATION EVALUATED THE DEVICE AND THE CUSTOMER'S REPORT WAS NOT REPLICATED OR CONFIRMED. THE DEVICE WAS PUT THROUGH EXTENSIVE TESTING INCLUDING BENCH HANDLING AND DEFIB CYCLE TESTING WITH A KNOWN GOOD BATTERY WITHOUT DUPLICATING THE REPORT. THE BATTERY USED WAS NOT RETURNED AS PART OF THIS INVESTIGATION. THE DEVICE WAS RECERTIFIED AND RETURNED TO THE CUSTOMER. IT IS RECOMMENDED TO THE CUSTOMER TO REVIEW BATTERY MANAGEMENT INSTRUCTIONS AND TO RECALIBRATE INVENTORY OF BATTERIES, AS A PRECAUTION. NO TREND IS ASSOCIATED WITH REPORTS OF THIS TYPE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22365039&pc=DSQ | 2916596-2025-04097 | 2025-06-12 04:00 | Death | THORATEC CORPORATION | 2025-06-30 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS CONGESTIVE HEART FAILURE (CHF) AND MULTIPLE SYSTEM ORGAN FAILURE (MSOF). THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. THE DEVICE WOULD NOT BE RETURNED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISIONS OF THE IFU CAN ALSO BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22365288&pc=QZI | 2124215-2025-43298 | 2025-05-29 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-30 04:00 | QZI | FARAWAVE | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED POST PROCEDURE. A PULSED FIELD ABLATION (PFA) PROCEDURE WAS PERFORMED WITH A FARAWAVE PFA CATHETER, AND 24 HOURS AFTER THE PROCEDURE, THE PATIENT PASSED AWAY. NO ISSUES WERE OBSERVED DURING THE PROCEDURE AND THE PROCEDURE WAS COMPLETED WITHOUT PATIENT COMPLICATIONS. THE PATIENT WAS DISCHARGED HEMODYNAMICALLY STABLE WITH NO COMPLICATIONS POST PROCEDURE. THE OFFICIAL CAUSE OF DEATH WAS NOT CONFIRMED. NO FURTHER INFORMATION IS AVAILABLE. THE CATHETER IS NOT EXPECTED TO RETURN FOR ANALYSIS AS IT WAS DISPOSED POST PROCEDURE. Manufacturer Narrative: IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT THIS DETAIL WAS NOT AVAILABLE FOLLOWING GOOD FAITH EFFORTS. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22365375&pc=DSQ | 2916596-2025-04066 | 2018-03-01 05:00 | Death | THORATEC CORPORATION | 2025-06-30 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿POTENTIAL BENEFITS OF AORTIC VALVE OPENING IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES¿, THAT HEARTMATE 3 MAY BE ASSOCIATED WITH VALVE INSUFFICIENCY/REGURGITATION, HEMODYNAMIC INSTABILITY, TRANSPLANT, AND DEATH. AORTIC VALVE REGURGITATION (AR) IN PATIENTS WITH LEFT VENTRICULAR ASSIST DEVICES (LVADS) LEADS TO RETROGRADE FLOW FROM THE LVAD'S OUTFLOW BACK INTO THE LEFT VENTRICLE (LV). THIS CAUSES SYSTEMIC HYPOPERFUSION AND INCOMPLETE UNLOADING OF THE LV. SIGNIFICANT AR NEGATIVELY IMPACTS LVAD PATIENT OUTCOMES, RESULTING IN LOWER SURVIVAL RATES AND HIGHER REHOSPITALIZATION RATES. THE INCIDENCE OF AR INCREASES WITH THE DURATION OF MECHANICAL CIRCULATORY SUPPORT (MCS). FACTORS SUCH AS AGE, GENDER, BODY MASS INDEX (BMI), AND LVAD SETTINGS INFLUENCE THE DEVELOPMENT OF AR DURING MCS. WHEN THE AORTIC VALVE (AV) FAILS TO OPEN DURING THE CARDIAC CYCLE (NON- OPENING AV, NAVO), PERSISTENT PRESSURE ON THE VALVE THROUGHOUT SYSTOLE AND DIASTOLE MAY LEAD TO AV DEGENERATION AND SUBSEQUENT AR. 4 NAVO DURING MCS HAS BEEN ASSOCIATED WITH AR. AV OPENING OR NAVO IS INFLUENCED NOT ONLY BY LVAD SETTINGS BUT ALSO BY THE MYOCARDIAL AND VALVULAR FUNCTION OF THE HEART SUPPORTED BY THE LVAD. FOR INSTANCE, SEVERE MITRAL REGURGITATION MAY PREVENT AV OPENING REGARDLESS OF LVAD ADJUSTMENT. THE OBJECTIVE OF THIS STUDY WAS TO EVALUATE CLINICAL AND CARDIAC PARAMETERS ASSOCIATED WITH AORTIC VALVE OPENING IN PATIENTS WITH LVAD. IN THIS RETROSPECTIVE ANALYSIS, THE FILES OF ALL PATIENTS WITHOUT CONGENITAL HEART DISEASE WHO UNDERWENT CFLVAD IMPLANTATION BETWEEN (B)(6) 2018 AND (B)(6) 2020 WERE REVIEWED AND THE COLLECTED DATA WERE ENTERED IN A REDCAP DATABASE. THE SPEED OF THE DEVICE WAS CALIBRATED FOR OPTIMAL BLOOD CIRCULATION AND ORGAN FUNCTION AFTER IMPLANTATION, BEFORE THE PATIENT WAS DISCHARGED AND AT EACH SUBSEQUENT OUTPATIENT APPOINTMENT. RAMP TESTS AND RIGHT HEART CATHETERIZATION WERE NOT ROUTINELY PERFORMED ON ALL PATIENTS. THESE INVESTIGATIONS WERE CARRIED OUT ONLY FOR SPECIFIC REASONS, SUCH AS POTENTIALLY SIGNIFICANT AORTIC VALVE INSUFFICIENCY, PATIENTS WITH EXERTIONAL DYSPNEA, OR PATIENTS WITH RIGHT HEART FAILURE. FOR ROUTINE PATIENTS, THE DEVICE SETTINGS WERE ADJUSTED BASED ON DEVICE PARAMETERS, ECHOCARDIOGRAPHIC FINDINGS, AND CLINICAL ASSESSMENT. SEPTUM POSITION, THE SEPTUM SHOULD BE CENTRALLY POSITIONED. IF IT WAS NOT AND THERE WAS NO OBVIOUS CAUSE, SUCH AS A SEPTIC CONDITION OR VOLUME OVERLOAD, INVASIVE DIAGNOSTIC AND/OR EVALUATION IN THE RAMP TEST WERE CONDUCTED. AORTIC VALVE OPENING, INTERMITTENT OPENING OF THE AORTIC VALVE WAS CONSIDERED INDICATIVE OF GOOD LVAD SPEED ADJUSTMENT. IF THE AORTIC VALVE OPENED IN EVERY CYCLE, THE DEVICE SPEED WAS INCREASED. IF THE AORTIC VALVE DID NOT OPEN AT ALL, THE DEVICE SPEED WAS REDUCED IF POSSIBLE. THE SPEED REDUCTION WAS CARRIED OUT GRADUALLY. FROM THE 219 PATIENTS, 156 PATIENTS PRESENTED TO THE FOLLOW- UP, ALIVE AND ON THE LVAD AFTER SIX MONTHS. 56 PATIENTS DIED EARLIER, THREE WERE HEART TRANSPLANTED, ONE GOT WEANED FROM THE DEVICE, AND THREE WERE LOST TO FOLLOW- UP. IN TWO OF THE 156 REVISED ECHOCARDIOGRAMS, WE WERE NOT ABLE TO VISUALIZE THE AV. IN THE REMAINING 154, 99 SHOWED AV OPENING IN EVERY CYCLE OR INTERMITTENTLY (64%), IN CONTRAST TO THE OTHER 36% WITH CONTINUOUSLY CLOSED AV ON THE 6- MONTH FOLLOW- UP ECHOCARDIOGRAPHY. WHEN SUBDIVIDING THE PATIENTS WITH AVO BETWEEN CONTINUOUSLY VS. INTERMITTENTLY VALVE OPENING AND COM PARING THE TWO GROUPS WITH THE NAVO SUBJECTS, THE RESULTS DID NOT DIFFER SIGNIFICANTLY FROM THOSE MENTIONED. WHEN ANALYZING INDEPENDENT PREDICTORS OF AV OPENING IN THE FOLLOW- UP IN A MULTIVARIATE LOGISTIC REGRESSION A WORSE NYHA CLASS WAS ASSOCIATED WITH LESS AV OPENING (NYHA >II: OR 0.29, IC 95% 0.10¿0.83, P 0.021), AS IT DID A HIGHER TR GRADE (TR =2: 0R 0.40, IC 95% 0.17 0.95, P 0.04) FROM THE ECHOCARDIOGRAPHIC POINT OF VIEW. LARGER LEFT VENTRICULAR END- DIASTOLIC DIAMETERS (LVEDD) AND A BETTER LEFT VENTRICULAR EJECTION FRACTION (LVEF) APPEARED AS SIGNIFICANT PREDICTORS OF A REGULAR AV OPENING AT THE TIME OF FOLLOW- UP AFTER ADJUSTING FOR OTHER RISK FACTORS (LVEDD: OR 1.06, IC 95% 1.02¿1.11, P 0.006; LVEF: OR 1.06, IC 95% 1.01¿1.11, P 0.02). AMONG ALL THE RIGHT VENTRICULAR FUNCTIONAL PARAMETERS TESTED, TAPSE SHOWED THE STRONGER ASSOCIATION WITH AV OPENING, ALTHOUGH IT DID NOT REACH STATISTICAL SIGNIFICANCE (OR 1.13, IC 95% 0.99¿1.27, P 0.07). THE MEDIAN TIME FROM PRIMARY IMPLANTATION TO THE END OF THE FOLLOW- UP WAS 3.2 YEARS (IQR 2.1 YEARS). THE HEART FAILURE MEDICATION WAS MAXIMALLY OPTIMIZED IN ALL SUBJECTS, WITHOUT DIFFERENCES WHEN DIVIDING BY AV OPENING. MOST PATIENTS WERE ON BETA- BLOCKERS AND MRA (88% AND 84%, RESPECTIVELY). TWO- THIRDS OF THEM TOLERATED ARNI (66%), AND THE REST WERE HALF ON ACE- INHIBITORS (15%) AND HALF ON ARB (15%). 102 PATIENTS HAD A HEARTWARE® SYSTEM (66%) VS. 52 PATIENTS WITH A HEARTMATE 3¿ (34%). A WORSE NYHA CLASS WAS ASSOCIATED WITH LESS AV OPENING (NYHA >II: OR 0.29, IC 95% 0.10¿0.83, P 0.021), AS IT DID A HIGHER TR GRADE (TR =2: 0R 0.40, IC 95% 0.17 0.95, P 0.04) FROM THE ECHOCARDIOGRAPHIC POINT OF VIEW. LARGER LEFT VENTRICULAR END- DIASTOLIC DIAMETERS (LVEDD) AND A BETTER LEFT VENTRICULAR EJECTION FRACTION (LVEF) APPEARED AS SIGNIFICANT PREDICTORS OF A REGULAR AV OPENING AT THE TIME OF FOLLOW- UP AFTER ADJUSTING FOR OTHER RISK FACTORS (LVEDD: OR 1.06, IC 95% 1.02¿1.11, P 0.006; LVEF: OR 1.06, IC 95% 1.01¿1.11, P 0.02). AFTER A MEDIAN FOLLOW- UP OF 3.2 YEARS, 85 OF THESE 156 PATIENTS WERE STILL ALIVE AND ON THE DEVICE (55%), 39 DIED DURING FOLLOW- UP (25%), 20 UNDERWENT HEART TRANSPLANTATION (13%), 3 UNDERWENT RVAD- IMPLANTATION (2%), AND 8 UNDERWENT LVAD EXPLANTATION DUE TO MYOCARDIAL RECOVERY (5%). THE SURVIVAL RATE ONE AND THREE YEARS AFTER THE 6- MONTH FOLLOW- UP WAS 78% AND 72%, RESPECTIVELY, IN THE GROUP WHO SHOWED REGULARLY AV OPENING AT 6 MONTHS AND 88% AND 85% IN THOSE WHO DID NOT. THE DIFFERENCE OF SURVIVAL RATE BETWEEN THE TWO COHORTS WAS NOT STATISTICALLY SIGNIFICANT. IT WAS FOUND AVO TO BE ASSOCIATED WITH LOWER NYHA CLASS, BETTER RIGHT HEART FUNCTION, AND A LOWER RATE OF AR IN PATIENTS WITH CENTRIFUGAL CONTINUOUS FLOW LVAD. THE SPEED OF THE DEVICE ALONE DOES NOT SOLELY INFLUENCE AV OPENING. STRUCTURAL HEART DAMAGE, SUCH AS SEVERE MITRAL REGURGITATION OR A VERY POOR EJECTION FRACTION, MAY ALSO IMPACT LV FILLING PRESSURES AND AV OPENING. ALTHOUGH THE TWO GROUPS OF PATIENTS SHOWED NO DIFFERENCES IN TERMS OF EJECTION FRACTION, A BETTER LVEF APPEARED AS SIGNIFICANTLY ASSOCIATED WITH REGULAR AVO AT THE TIME OF FOLLOW- UP AFTER ADJUSTING FOR OTHER RISK FACTORS. THERE IS A THOUGHT THAT A HIGHER MEAN ARTERIAL PRESSURE CAN BE AN INDIRECT EXPRESSION OF BETTER LV FUNCTION. THE BETTER LEFT VENTRICULAR FUNCTION MAY LEAD TO THE OPENING OF THE AORTIC VALVE AND CONTRIBUTE TO THE GENERATION OF A HIGHER BLOOD PRESSURE. THIS STUDY FOCUSED ON STABLE PATIENTS SIX MONTHS AFTER IMPLANTATION WHO PRESENTED IN THE OUTPATIENT CLINIC. IN THESE PATIENTS, THE VASODILATOR AND HEART FAILURE DRUGS WERE RAPIDLY INCREASED UP TO THE MAXIMUM DOSE TOLERATED BY THE PATIENT WITHIN THE FIRST 6 MONTHS AFTER IMPLANTATION ACCORDING TO A STANDARDIZED PROTOCOL. A LARGE PROPORTION OF THE PATIENTS WERE ALREADY TREATED WITH BETA- BLOCKERS, MINERALOCORTICOID RECEPTOR ANTAGONISTS, AND INHIBITORS OF THE RENIN¿ANGIOTENSIN SYSTEM AT THE TIME OF THE SIX-MONTH FOLLOW- UP, WITH NO SIGNIFICANT DIFFERENCES IN MEDICAL TREATMENT REGARDING AORTIC VALVE OPENING. PATIENTS WITH NON- AORTIC VALVE OPENING (NAVO) HAD A SIGNIFICANTLY HIGHER PREVALENCE OF RELEVANT AORTIC VALVE INSUFFICIENCY. IN CONCLUSION, THE PRESENCE OF AORTIC VALVE OPENING WAS ASSOCIATED WITH ADVANTAGEOUS PARAMETERS SUCH AS BETTER RIGHT VENTRICULAR. THE PRESENCE OF AORTIC VALVE OPENING WAS ASSOCIATED WITH ADVANTAGEOUS PARAMETERS SUCH AS BETTER RIGHT VENTRICULAR. Manufacturer Narrative: B5:SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DORTA, E.R., MEYN, R., MÜLLER, M., HOERMANDINGER, C., SCHOENRATH, F., FALK,V., MEYER NICOLAS MERKE, A., POTAPOV, E., MULZER, J., KNIERIM, J. ARTIFICIAL ORGANS. 2025;49:441¿450. RECEIVED: 16 JUNE 2024, REVISED: 18 SEPTEMBER 2024, ACCEPTED: 3 OCTOBER 2024 DOI: 10.1111/AOR.14891. DEPARTMENT OF CARDIOLOGY, ANGIOLOGY AND INTENSIVE CARE, DEUTSCHES HERZZENTRUM DER CHARITÉ (DHZC), BERLIN, GERMANY CHARITÉ ¿ UNIVERSITÄTSMEDIZIN BERLIN, CORPORATE MEMBER OF FREIE UNIVERSITÄT BERLIN UND HUMBOLDT- UNIVERSITÄT ZU BERLIN, BERLIN, GERMANY DZHK (GERMAN CENTRE FOR CARDIOVASCULAR RESEARCH), PARTNER SITE BERLIN, BERLIN, GERMANY DEPARTMENT OF CARDIOTHORACIC AND VASCULAR SURGERY, DEUTSCHES HERZZENTRUM DER CHARITÉ (DHZC), BERLIN, GERMANY. DEPARTMENT OF INTERNAL MEDICINE AND CARDIOLOGY, SANA PAULINENKRANKENHAUS BERLIN, BERLIN, GERMANY DEPARTMENT OF HEALTH SCIENCES AND TECHNOLOGY, TRANSLATIONAL CARDIOVASCULAR TECHNOLOGY, EIDGENÖSSISCHE TECHNISCHE HOCHSCHULE ZÜRICH, ZURICH, SWITZERLAND. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22365466&pc=LNQ | 3027386225-2025-00078 | 2025-05-12 04:00 | Death | ENTERRA MEDICAL, INC. | 2025-06-30 04:00 | LNQ | ENTERRA II IPG | No Apparent Adverse Event | Insufficient Information | H990014 |  | 1 | Event Description: PRODUCT RETURNED TO MEDTRONIC FOR ANALYSIS: REPORTED IN MEDTRONIC DATABASE. RECEIVED FROM (B)(6). CODED AS C-500 NON-COMPLAINT AT MEDTRONIC AND CLOSED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22365531&pc=DSQ | 2916596-2025-04038 | 2025-04-29 04:00 | Death | THORATEC CORPORATION | 2025-06-30 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, "INTRODUCTION", LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22365659&pc=GDW | 2647580-2025-01992 | 2021-09-12 04:00 | Death | US SURGICAL PUERTO RICO | 2025-06-30 04:00 | GDW | TA | Mechanics Altered; Adverse Event Without Identified Device or Use Problem | Failure to Anastomose; Abscess; Adhesion(s); Anemia; Bacterial Infection; Dehydration; Purulent Discharge; Emotional Changes; Erythema; Foreign Body Reaction; Hematoma; Hemorrhage/Bleeding; Unspecified Infection; Inflammation; Muscle Weakness; Necrosis; Pain; Pleural Effusion; Rash; Shock; Electrolyte Imbalance; Hernia; Ulcer; Anxiety; Distress; Discomfort; Depression; Impaired Healing; Ambulation Difficulties; Ascites; Abdominal Distention; Blister; Unspecified Tissue Injury; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  |  | 1 | Event Description: THE PATIENT¿S ATTORNEY ALLEGED A DEFICIENCY AGAINST THE DEVICE. THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING A SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, LEUKOCYTOSIS, ASCITES, SEROSITIS, LOOPS OF DISTENDED AIR AND DILATED SMALL/LARGE BOWEL, GASEOUS DISTENTION, ILEUS, BILATERAL BASILAR SUBSEGMENTAL ATELECTASIS, FLUID COLLECTION AT THE SITE OF PRIMARY ANASTOMOSIS, PNEUMOPERITONEUM, LABS ABNORMAL FOR WBC; H & H, DEFECT AT ANASTOMOTIC SITE, ULCERATION, INFLAMMATION, NECROSIS, REACTIVE CHANGES, OPEN WOUNDS, HERNIA DEFECT, LOCULES OF AIR, ABSCESS, INFECTED FLUID COLLECTION, PLEURAL EFFUSION, HEMATOMAS, PSEUDOMONAS, ABNORMAL ELECTROLYTE COUNTS, BLOODY FLUID, PURULENT FLUID, HYPOKALEMIA, ARTHRODESIS STATUS, ANEMIA, MAJOR DEPRESSIVE DISORDER, ANXIETY DISORDER, GENERALIZED MUSCLE WEAKNESS, ABNORMALITIES OF GAIT & MOBILITY, LEAKING ILEOSTOMY BAG CAUSING BLISTERING RED RASH, AFRAID TO DRINK ANYTHING, DEHYDRATED WITH ORANGE URINE, UNABLE TO AMBULATE BY SELF, UNCOMFORTABLE, ADHESIONS, AMBULATING WITH ROLLING WALKER, ENTEROSTOMY COMPLICATIONS, HYPOMAGNESEMIA, & DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, DIVERTING LOOP ILEOSTOMY, CT SCAN, X-RAY, ABSCESS DRAINAGE, REHABILITATION CENTER STAY, PHYSICAL THERAPY, IV ANTIBIOTICS, PICC LINE PLACEMENT, INSERTION OF SVC INFUSION DEVICE, STUDIES DONE, NG TUBE PLACEMENT, PAIN MEDICATION, WOUND CARE, IV FLUIDS, OCCUPATIONAL THERAPY, COLOSTOMY CARE, PAIN MANAGEMENT, HOSPITALIZATION, REVERSAL OF LOOP ILEOSTOMY WITH RESECTION AND PRIMARY ANASTOMOSIS, WOUND PACKED, ELECTROLYTE SUPPLEMENTATION, TAKING LOVENOX FOR DVT PROPHYLAXIS, CPR, & 5 ROUNDS OF EPINEPHRINE. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED DUE TO A PULMONARY EMBOLISM. THE PATIENT¿S HISTORY OF MULTIPLE RECENT ABDOMINAL SURGERIES WITH A LONG HOSPITALIZATION COULD BE CONSIDERED AS A RISK FACTOR FOR THROMBOEMBOLISM. Manufacturer Narrative: ADDITIONAL INFO: B7. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: H6 IME E2402: BILATERAL BASILAR SUBSEGMENTAL ATELECTASIS, LABS ABNORMAL FOR WBC; H & H, LOCULES OF AIR LEUKOCYTOSIS, LOOPS OF DISTENDED AIR IN LARGE/SMALL BOWEL, GASEOUS DISTENTION, ILEUS ARTHRODESIS STATUS MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22365869&pc=MCX | 2124215-2025-41979 | 2025-06-09 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-30 04:00 | MCX | ROTAPRO? | Adverse Event Without Identified Device or Use Problem | Thrombosis/Thrombus | P900056 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT PRESENTED WITH CORONARY ARTERY DISEASE AND UNDERWENT PERCUTANEOUS CORONARY INTERVENTION. THE TARGET LESION WAS LOCATED IN THE MID LEFT ANTERIOR DESCENDING ARTERY (LAD). A 1.25MM ROTAPRO WAS SELECTED FOR USE. DURING THE PROCEDURE, AFTER STENT DEPLOYMENT, A THROMBOSIS OCCURRED IN THE LAD. POST DILATATION WITH BALLOON ANGIOPLASTY WAS PERFORMED; HOWEVER, THE PATIENT DIED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22366088&pc=LWS | 9614453-2025-02517 |  | Death | MEDTRONIC EUROPE SARL | 2025-06-30 04:00 | LWS | EVERA MRI S VR SURESCAN | Failure to Deliver Shock/Stimulation | Insufficient Information | P980016 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS DECEASED. IT WAS REPORTED THAT THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) DID NOT DELIVER A SHOCK AND THE PATIENT EXPIRED. THE DEVICE WAS EXPLANTED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22366415&pc=LNQ | 3027386225-2025-00079 | 2025-05-12 04:00 | Death | ENTERRA MEDICAL, INC. | 2025-06-30 04:00 | LNQ | ENTERRA II IPG | No Apparent Adverse Event | Insufficient Information | H990014 |  | 1 | Event Description: PRODUCT RETURNED TO MEDTRONIC FOR ANALYSIS - REPORTED IN MEDTRONIC DATABASE. RETURNED FROM (B)(6) FUNERAL HOME IN (B)(6). CODED AS C-500 NON-COMPLAINT AT MEDTRONIC AND CLOSED. ORIGINALLY IMPLANTED ON (B)(6) 2022.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22366606&pc=NPT | 2025587-2025-04648 | 2025-06-26 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-30 04:00 | NPT | EVOLUT PRO PLUS VALVE | Fluid/Blood Leak; Insufficient Information; Activation Failure; Patient Device Interaction Problem; Central Regurgitation | Cardiac Arrest; Valvular Insufficiency/ Regurgitation; Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THE TRANSCATHETER BIPROSTHETIC VALVE INTO A PATIENT WITH CALCIFICATION, THE VALVE HYPO-EX PANDED. A BALLOON DILATION WAS PERFORMED WITH A 28 MM BALLOON HOWEVER, DURING THE DILATION, THE BALLOON STOPPED DILATING DUE TO A STORM. THE PATIENT DIED. THE CAUSE OF DEATH IS UNKNOWN.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT CLARIFIED PREVIOUSLY REPORTED INFORMATION INDICATING THAT DURING THE PROCEDURE THE PATIENT CRASHED 10 TIMES.Event Description: IT WAS REPORTED THAT DURING THE IMPLANT OF THE TRANSCATHETER BIOPROSTHETIC VALVE INTO A PATIENT WITH CALCIFICATION, THE VALVE BECAME HYPO-EXPANDED WITH REFLUX. A 28 MILLIMETER (MM) BALLOON DILATION WAS PERFORMED, HOWEVER DURING THE BALLOON DILATION, AN ELECTRICAL STORM OCCURRED AND STOPPED THE DILATION. THE PATIENT DIED. ADDITIONAL INFORMATION WAS RECEIVED THAT THE VALVE DID NOT EXPAND AND SEVERE REGURGITATION OCCURRED. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS NOT KNOWN BUT THE VALVE AND DELIVERY CATHETER SYSTEM (DCS) COULD NOT BE EXCLUDED AS CONTRIBUTING FACTORS.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE TYPE OF SEVERE REGURGITATION WAS CENTRAL. Manufacturer Narrative: IMAGE REVIEW: INTRAPROCEDURAL FLUOROSCOPIC IMAGES WERE PROVIDED FOR REVIEW OF THE EVENT. PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. FLUOROSCOPIC VALVE LOAD INSPECTION WAS PERFORMED AND CONFIRMED A GOOD LOAD. A PRE BALLOON AORTIC VALVULOPLASTY WAS PERFORMED PRIOR TO THE VALVE IMPLANTATION. THE VALVE WAS DEPLOYED JUST PRIOR TO THE POINT OF NO RECAPTURE AND SIGNIFICANT UNDER EXPANSION WAS NOTED. PER MEDTRONIC BEST PRACTICES, IF A VALVE IS UNDER-EXPANDED: REMOVE SYSTEM, PERFORM PRE-DILATATION, AND IMPLANT NEW VALVE WITH NEW DELIVERY SYSTEM. HOWEVER, THE VALVE WAS RECAPTURED AND REDEPLOYED RESULTING IN AN INFOLD. THE INFOLD IS CHARACTERIZED BY AN INWARD FOLD OR CREASE IN THE VALVE, EXTENDING FROM THE INFLOW. INFOLDING COULD OCCUR DUE A MISLOADED VALVE, ANATOMICAL CHARACTERISTICS SUCH AS CALCIUM, AND RECAPTURES. IF AN INFOLD IS IDENTIFIED, THE VALVE SHOULD BE RECAPTURED, REMOVED FROM THE BODY, AND DISCARDED. NEW STERILE COMPONENTS MUST BE USED. DESPITE THE INFOLD, THE VALVE WAS RELEASED. THERE IS EVIDENCE THAT CARDIOPULMONARY RESUSCITATION EFFORTS WERE IN PROGRESS DURING THE POST IMPLANT DILATATION. IT WAS REPORTED THAT THE PATIENT SUBSEQUENTLY DIED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: PRO+ DELIVERY CATHETER SYSTEM (DCS), PRODUCT ID: D-EVPROP34, LOT: 0012659440, USE BY DATE: 2027-02-06, UDI: (B)(4). UPDATED: B5, H6, H11. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVPROP34 (LOT: 0012583313); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID D-EVPROP34 (LOT: 0012659440); PRODUCT TYPE: 0195-HEART VALVES; MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTED DATA: B5. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22366674&pc=MUD | 2015691-2025-05308 | 2025-06-05 04:00 | Death | EDWARDS LIFE SCIENCES | 2025-06-30 04:00 | MUD | FORE-SIGHT ELITE MONITOR | Adverse Event Without Identified Device or Use Problem | Low Oxygen Saturation | K143675 |  | 1 | Event Description: DURING USE OF A FORESIGHT ELITE MONITOR, A 15-YEAR-OLD MALE PATIENT WAS DECLARED BRAIN DEAD FOLLOWING AN OPEN-HEART SURGERY FOR REPLACEMENT OF PULMONARY VALVE REPLACEMENT AND PFO CLOSURE. PER THE LOGS, THE FORESIGHT ELITE MONITOR DISPLAYED SIGNIFICANT DROPS IN TISSUE OXYGEN SATURATION. IT IS SUSPECTED THAT THAT THE ALARM VOLUME WAS SET TO WHISPER DURING THE CASE. THE FACILITY STATED THAT THE SURGEON WAS UNAWARE OF THE LOW TISSUE OXYGEN SATURATION (STO2) VALUES DURING THE CASE. THE FACILITY DOES NOT SUSPECT A PRODUCT MALFUNCTION. Manufacturer Narrative: A PRODUCT EVALUATION WAS PERFORMED AT THE CUSTOMER SITE TO ASSESS THE FUNCTIONALITY AND INTEGRITY OF THE FORESIGHT ELITE MONITOR AND PRE-AMP CABLES. IN ADDITION, THE ANALYSIS ALSO INCLUDED A REVIEW OF THE MONITOR'S CASE DATA LOG FOR THE REPORTED EVENT. THE PRODUCT INVESTIGATION AND THE CASE DATA LOG ANALYSIS CONCLUDED THAT THE MONITOR AND THE TWO PRE-AMP CABLES USED IN THE COMPLAINT CASE FUNCTIONED AS INTENDED. THE MONITOR PASSED ALL TESTING INCLUDING AN ALARM VOLUME TEST. THE ALARM VOLUME TEST WAS PERFORMED MEASURING THE ALARM AUDIO VOLUME THREE TIMES FOR ALL ALARM LEVELS USING A CALIBRATED SOUND LEVEL METER. MONITOR CASE DATA LOG ANALYSIS: SYSTEM TRIGGERED LOW AND HIGH STO2 ALARMS ACCORDING TO THE USER SETTINGS. IT WAS CONFIRMED THAT VALUES WERE DISPLAYED FOR UP TO FOUR CHANNELS EVERY TWO SECONDS. FOUR BODY LOCATIONS WERE MONITORED: LEFT AND RIGHT CEREBRAL, LEFT AND RIGHT FLANK MUSCLES. AS DEFINED IN THE USER MANUAL, MULTIPLE INFORMATIONAL ALARMS, AND MEDIUM AND LOW PRIORITY ALARMS OCCURRED. THE STO2 ALARM LIMITS (%) WERE SET BY THE USER. IT WAS CONCLUDED THAT THERE WAS NO PRODUCT NONCONFORMANCE. THE DEVICE ALARMED AS INTENDED. Manufacturer Narrative: THE DEVICE EVALUATION IS ANTICIPATED. HOWEVER, THE COMPLAINT CANNOT BE CONFIRMED WITHOUT THE COMPLETION OF A PRODUCT EVALUATION. A SUPPLEMENTAL REPORT WILL BE FORTHCOMING WHEN THE INVESTIGATION IS COMPLETED. A DEVICE HISTORY RECORD REVIEW WAS PERFORMED, AND NO RELATED NON-CONFORMANCES WERE FOUND. NO ESCALATION IS REQUIRED. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS A PART OF THE MONTHLY REVIEW.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22366754&pc=LNQ | 3027386225-2025-00080 | 2025-05-12 04:00 | Death | ENTERRA MEDICAL, INC. | 2025-06-30 04:00 | LNQ | ENTERRA II IPG | No Apparent Adverse Event | Insufficient Information | H990014 |  | 1 | Event Description: PRODUCT RETURNED TO MEDTRONIC FOR ANALYSIS - REPORTED IN MEDTRONIC DATABASE. RECEIVED FROM (B)(6) FUNERAL HOME IN (B)(6). CODED AS C-500 NON-COMPLAINT AT MEDTRONIC AND CLOSED. ORIGINALLY IMPLANTED ON (B)(6) 2021.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22366922&pc=MVK | 3008642652-2025-06343 | 2025-04-18 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-30 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN INAPPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 11:12:26 ON (B)(6) 2025. AT 04:02:50 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS VF WITH CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. AT 04:03:29, THE PATIENT RECEIVED THE INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC, CPR/MOTION ARTIFACT, AND ELECTRODE LEAD FALL OFF CONTRIBUTED TO THE FALSE DETECTION. THE RHYTHM AT THE TIME OF TREATMENT WAS OBSCURED DUE TO CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. POST SHOCK RHYTHM WAS OBSCURED DUE TO CPR/MOTION ARTIFACT/ELECTRODE LEAD FALL OFF. THE ELECTRODE BELT WAS DISCONNECTED AT 04:53:26 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22366927&pc=MVK | 3008642652-2025-06344 | 2025-04-18 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-30 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED SIX APPROPRIATE TREATMENTS AND THREE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 23:07:19 ON (B)(6) 2025. AT 02:09:48 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS NSVT TRANSITIONING TO VT @ 230 BPM. AT 02:12:22, THE PATIENT RECEIVED THE FIRST APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VT @ 230 BPM. POST SHOCK RHYTHM WAS VF. AT 02:12:43, THE PATIENT RECEIVED THE SECOND APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 220 BPM. AT 02:13:17, THE PATIENT RECEIVED THE THIRD APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VT @ 240 BPM. POST SHOCK RHYTHM WAS VF. AT 02:13:48, THE PATIENT RECEIVED THE FOURTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VF. POST SHOCK RHYTHM WAS VT @ 210 BPM. AT 02:14:22, THE PATIENT RECEIVED THE FIFTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VT @ 240 BPM. POST SHOCK RHYTHM WAS NSVT TRANSITIONING TO VT FROM 240-250 BPM/VF/FINE VF. AT 02:37:04, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS FINE VF. AT 02:38:16, THE PATIENT RECEIVED THE SIXTH APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS FINE VF. POST SHOCK RHYTHM WAS ASYSTOLE WITH TACTILE ARTIFACT. POST-SHOCK ASYSTOLE IS A KNOWN AND POTENTIALLY ADVERSE OUTCOME OF DEFIBRILLATION THERAPY. AT 02:39:28, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 02:42:47, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. AT 02:43:15, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH CPR/MOTION ARTIFACT. AT 02:43:42, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF CARDIAC ACTIVITY AND CPR/MOTION ARTIFACT CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE WITH CPR/MOTION ARTIFACT. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM, WITH MOTION/TACTILE ARTIFACT. THE ELECTRODE BELT WAS DISCONNECTED AT 02:43:49 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22367107&pc=BZD | 2518422-2025-106963 | 2025-06-26 04:00 | Death | RESPIRONICS, INC. | 2025-06-30 04:00 | BZD | DREAMSTATION AUTO CPAP | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION A USER OF A DREAMSTATION AUTO CPAP PASSED AWAY. THE PATIENT'S SPOUSE REPORTED THE PATIENT'S DEATH AND DID NOT WANT TO ANSWER ANY QUESTIONS. THE CAUSE OF DEATH WAS NOT PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22368108&pc=NAY | 2955842-2025-27878 | 2024-11-20 05:00 | Death | INTUITIVE SURGICAL, INC | 2025-06-30 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure; Thrombosis/Thrombus | K131861 |  | 1 | Event Description: A LITERATURE ARTICLE DESCRIBED A SINGLE SITE, PROSPECTIVE STUDY FROM DECEMBER 2022 TO APRIL 2024 FOR THE SAFETY AND FEASIBILITY OF 3-DIMENSIONAL (3D) VIRTUAL LUNG RECONSTRUCTION IN ROBOTIC SEGMENTECTOMY. THE STUDY FOCUSED ON THE RATE OF CONVERSION FROM SEGMENTECTOMY TO LOBECTOMY WHEN PREOPERATIVE PLANNING IS PERFORMED WITH 3D RECONSTRUCTION IN PATIENTS WITH STAGE I NON¿SMALL CELL LUNG CANCER LESS THAN 3 CM. HIGH-RESOLUTION CT SCAN OF THE CHEST WITH CONTRAST IN THE PULMONARY ARTERY PHASE BEFORE THE OPERATION WAS UTILIZED FOR THE 3D MODELING. THE DA VINCI XI SYSTEM WAS USED WITH THE FIREFLY FLUORESCENCE IMAGING CAMERA AS A NIF LIGHT SOURCE AND WITH INDOCYANINE GREEN DYE. OF THE 67 PATIENTS THAT MET THE STUDY CRITERIA, THE ARTICLE NOTED ONE MORTALITY THAT WAS SECONDARY TO MULTIORGAN FAILURE IN THE SETTING OF MULTIPLE POSTOPERATIVE THROMBOTIC EVENTS OF UNCLEAR ETIOLOGY. NO ADDITIONAL DETAILS WERE INCLUDED. THERE WAS NO INDICATION OF ANY DA VINCI SYSTEM ISSUES IN THE ARTICLE. ADDITIONAL INFORMATION WAS REQUESTED FROM THE DESIGNATED AUTHOR CONTACT; NO RESPONSE WAS RECEIVED. Manufacturer Narrative: NO INVESTIGATION COULD BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE OR DA VINCI SYSTEM IDENTIFIERS. THERE WAS NO REPORT OF, OR INDICATION OF, ANY DA VINCI PRODUCT ISSUES, AND NO INDICATION THAT ANY DA VINCI PRODUCTS CONTRIBUTED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE MENTIONED PATIENT IN THIS LITERATURE ARTICLE DIED FROM MULTISYSTEM ORGAN FAILURE FOLLOWING A ROBOTIC PROCEDURE. ALTHOUGH THEY DEVELOPED MULTISYSTEM ORGAN FAILURE IN THE SETTING OF MULTIPLE POSTOPERATIVE THROMBOTIC EVENTS OF UNCLEAR ETIOLOGY, THE SPECIFIC REASON THEY DEVELOPED THIS COMPLICATION AND THE EVENTS THAT LED TO THESE PROBLEMS AND THE DEATH ARE NOT PROVIDED. NO ALLEGATION AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT IS MADE IN THE ARTICLE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SECTIONS A2 PATIENT AGE, AND A3 PATIENT GENDER: NO SPECIFIC PATIENT DEMOGRAPHICS WERE PROVIDED. THE STUDY MEDIAN PATIENT AGE WAS 70 (65-74 YEARS), AND THE GENDER RATIO FEMALE (39) TO MALE (28). SECTION D3 EVENT DATE: DUE TO THE LACK OF SPECIFIC INFORMATION REGARDING THE EVENT DATE ASSOCIATED WITH THE ADVERSE EVENT, THE ARTICLE PUBLISHED DATE WAS USED. CITATION: BROWNE, IKENNAH L., PATEL, YOGITA S., HANNA, NADER. M., HAIDER, ESHAN, HANNA, WAEL C. 3D VIRTUAL LUNG RECONSTRUCTION IN ROBOTIC SEGMENTECTOMY ¿ A SAFETY AND FEASIBILITY TRIAL. JTCVS TECHNIQUES. 2025;29:150-60. HTTPS://DOI.ORG/10.1016/J.XJTC.2024.10.024.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22343103&pc=DTZ | 9681834-2025-00117 | 2025-06-01 04:00 | Death | TERUMO CORPORATION, ASHITAKA | 2025-06-27 04:00 | DTZ | CAPIOX®FX | Fluid/Blood Leak | Low Blood Pressure/ Hypotension | K071494 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A SMALL BLOOD LEAK ARE DETECTED AT THE ARTERIAL CONNECTION PORT OUTLET DURING THE SETUP OF THE EXTRACORPOREAL CIRCUIT FOR THE PATIENT'S SURGERY. THE ISSUE WAS IDENTIFIED WHEN THE CIRCUIT WAS PRIMED. SINCE THE PATIENT WAS ALREADY CONNECTED TO THE CARDIOPULMONARY BYPASS (CPB) AND THERE WAS INSUFFICIENT TIME TO REPLACE THE CIRCUIT, THE LEAK WAS RESOLVED BY APPLYING A TEGADERM TO PREVENT A HIGH-RISK SITUATION. THE PATIENT EXPERIENCED HYPOTENSION AND AIRWAY BLEEDING, WHICH WAS THE CAUSE OF DEATH. Manufacturer Narrative: D6A: IMPLANTED DATE: DEVICE WAS NOT IMPLANTED D6B: EXPLANTED DATE: DEVICE WAS NOT EXPLANTED E3: OCCUPATION: PERFUSIONIST G4: 510K: K1300520. THE ACTUAL DEVICE IS NOT AVAILABLE FOR RETURN; THEREFORE, AN EVALUATION OF THE ACTUAL DEVICE WAS UNABLE TO BE CONDUCTED. THE MANUFACTURING RECORD AND THE SHIPPING INSPECTION RECORD OF THE ACTUAL DEVICE - NO ANOMALY WAS FOUND. PAST COMPLAINT FILE - NO OTHER SIMILAR REPORT OF THE PRODUCT WITH THE INVOLVED PRODUCT CODE/LOT NUMBER WAS FOUND. CAUSE OF OCCURRENCE/CONCLUSION BASED ON THE INVESTIGATION RESULT, NO ANOMALY WAS FOUND IN THE MANUFACTURING RECORDS. SINCE THE ACTUAL DEVICE COULD NOT BE INVESTIGATED, IT WAS NOT POSSIBLE TO CLARIFY THE CAUSE OF OCCURRENCE. RELEVANT IFU REFERENCE: "DO NOT USE AN OXYGENATOR AND RESERVOIR THAT LEAKS. REPLACE IT WITH ANOTHER CAPIOX FX25 OXYGENATOR AND RESERVOIR." TERUMO MEDICAL CORPORATION (TMC) (IMPORTER) REGISTRATION NO. 2243441 IS SUBMITTING THIS REPORT ON BEHALF OF ASHITAKA FACTORY OF TERUMO CORPORATION (MANUFACTURER) REGISTRATION NO. 9681834.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22343698&pc=DTQ | 8010762-2025-0000270 | 2025-05-15 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-06-27 04:00 | DTQ | CARDIOHELP | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Hemorrhage/Bleeding | K133598 |  | 1 | Event Description: MAQUET CARDIOPULMONARY GMBH WAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES ON 2025-06-25: THE VIGILANCE AREA OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT INVOLVING THE PRODUCT ¿CARDIOHELP SYSTEM¿, MANUFACTURED BY GETINGE AND DISTRIBUTED BY, OCCURRED ON 2025-05-15 WITH REPORT NUMBER (B)(4). THE INFORMATION WAS PROVIDED THAT THE PATIENT HAD THE INCIDENT "MAJOR BLEEDING: RETROPERITONEAL HEMORRHAGE" AND THE CONSEQUENCE FOR THE PATIENT WAS "PATIENT PASSED AWAY¿ NO ADDITIONAL INFORMATION WAS RECEIVED. FURTHER INFORMATION ABOUT THE EVENT AND DEVICE WAS REQUESTED BUT IS STILL PENDING. AS THE PATIENT PASSED AWAY, A REPORT IS REQUIRED. COMPLAINT ID (B)(4).Event Description: COMPLAINT ID#: (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE SPAIN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K133598. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR CARDIOHELP WITH CATALOG NUMBER 701048012. Manufacturer Narrative: MAQUET CARDIOPULMONARY GMBH WAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES ON 2025-06-25: THE VIGILANCE AREA OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT INVOLVING THE PRODUCT ¿CARDIOHELP SYSTEM¿, MANUFACTURED BY GETINGE AND DISTRIBUTED BY, OCCURRED ON 2025-05-15. THE INFORMATION WAS PROVIDED THAT THE PATIENT HAD THE INCIDENT "MAJOR BLEEDING: RETROPERITONEAL HEMORRHAGE" AND THE CONSEQUENCE FOR THE PATIENT WAS "PATIENT PASSED AWAY¿. NEW INFORMATION WAS RECEIVED ON 2025-06-27 THAT NEITHER THE CARDIOHELP DEVICE NOR THE DISPOSABLE WERE INVOLVED. THE CUSTOMER CONFIRMED THAT THE PATIENT PASSED AWAY FROM HEMORRHAGE DURING A COMPLEX INTERVENTIONAL PROCEDURE AND THAT NO GETINGE EQUIPMENT WAS INVOLVED. BASED ON THE AVAILABLE INFORMATION THE REPORTED EVENT "MAJOR BLEEDING RETROPERITONEAL HEMORRHAGE " COULD NOT BE LINKED TO A DEVICE DEFECT/FAILURE. THEREFORE THE COMPLAINT COULD NOT BE CONFIRMED. THE CONFIRMATION WAS RECEIVED THAT THERE WAS NO PRODUCT MALFUNCTION ON A GETINGE PRODUCT, THEREFORE THIS COMPLAINT IS NOT REPORTABLE. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22344370&pc=OZD | 1220648-2025-29838 | 2025-06-10 04:00 | Death | ABIOMED, INC. | 2025-06-27 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS ON IMPELLA CP SUPPORT AND EXHIBITED A BLEED WITH MIDLINE SHIFT. THE IMPELLA WAS EXPLANTED AND THE PATIENT WAS PLACED ON COMFORT CARE UNTIL HE EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM, SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22344524&pc=DSQ | 2916596-2025-03937 | 2025-06-05 04:00 | Death | THORATEC CORPORATION | 2025-06-27 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hematoma; Muscle Weakness; Convulsion/Seizure; Speech Disorder; Hemorrhagic Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT INITIALLY PRESENTED WITH SLIGHT SLURRED SPEECH AND LEFT ARM WEAKNESS THAT QUICKLY PROGRESSED TO SEIZURE-LIKE ACTIVITY AND GLASGOW COMA SCALE (GCS) OF 3. THE PATIENT THEN PASSED AWAY DUE TO CATASTROPHIC SUBDURAL HEMATOMA (SDH). THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. THE DEVICE WOULD NOT BE EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. THE DEVICE HAD OPERATED AS EXPECTED. THE PATHOLOGY WAS PENDING BUT IT APPEARED THAT THE CAUSE OF SDH MAY HAVE BEEN A NEOPLASM/GLIOBLASTOMA. THE CAUSE OF SDH WAS NOT DUE TO A POST MECHANICAL FALL OR ACCIDENT. PATHOLOGY REPORTED AN AGGREGATE OF TAN-BROWN, RAGGED, HEMORRHAGIC TISSUE FRAGMENTS MEASURING 5.0 X 4.0 X 1.5 CENTIMETERS. NO FURTHER INFORMATION WAS AVAILABLE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22344705&pc=OZD | 1220648-2025-29842 | 2025-06-12 04:00 | Death | ABIOMED, INC. | 2025-06-27 04:00 | OZD | IMPELLA CP | Patient-Device Incompatibility | Cardiovascular Insufficiency | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED AN IMPELLA CP WAS BEING PLACED RATHER THAN AN IMPELLA 5.5 THAT COULD NOT BE USED AS THE ARM/AXILLARY WAS TO BE USED FOR CANNULATION OF EXTRACORPOREAL MEMBRANE OXYGENATION. THE IMPELLA CP WAS PLACED FEMORALLY BUT THE DELIVERY WAS COMPLICATED AND THE .018 WIRE KINKED. THE TEAM REPLACED THE ABIOMED WIRE. FURTHER DETAILS PROVIDED STATED THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY USE RELATED SINCE NON ABIOMED PRODUCT WAS USED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH RISK PERCUTANEOUS CORONARY INTERVENTION SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: CAUTIONS: ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22344728&pc=DSQ | 3007042319-2025-01382 | 2025-06-21 04:00 | Death | HEARTWARE, INC. | 2025-06-27 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Pneumonia; Insufficient Information | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR ASSIST DEVICE DIED, WITH THE CAUSE OF DEATH UNKNOWN. IT WAS UNKNOWN IF THERE WERE ANY PATIENT SYMPTOMS OR COMPLICATIONS ASSOCIATED WITH THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT DEVELOPED PNEUMONIA AND REFUSED INTERVENTION TO BE INTUBATED. THE VAD WAS TURNED OFF, SUPPORT WAS WITHDRAWN AND THE PATIENT DIED. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: (B)(6) WAS NOT RETURNED FOR EVALUATION. THIS COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE DEVICE HISTORY RECORD WAS NOT PERFORMED AS THE REPORTED EVENT IS NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. LOG FILE ANALYSIS WAS NOT PERFORMED SINCE LOG FILES WERE NOT AVAILABLE. BASED ON THE AVAILABLE INFORMATION, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, RESPIRATORY DYSFUNCTION, INFECTION, AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. THERE IS NO EVIDENCE THAT THE PATIENT HAD A HISTORY OF SIMILAR ADVERSE EVENTS. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR ADDITIONAL EVENT DETAILS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22345039&pc=NPT | 2025587-2025-04590 | 2022-11-26 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-27 04:00 | NPT | MEDTRONIC TRANSCATHETER VALVE | Fluid/Blood Leak; Device Dislodged or Dislocated; Patient Device Interaction Problem; Migration; Central Regurgitation | Cardiac Arrest; Fall; Hematoma; Intracranial Hemorrhage; Pneumonia; Ventricular Fibrillation; Cardiogenic Shock; Foreign Body In Patient; Heart Failure/Congestive Heart Failure; Valvular Insufficiency/ Regurgitation; Mitral Valve Insufficiency/ Regurgitation | P130021 |  | 1 | Event Description: LITERATURE WAS REVIEWED REGARDING TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) FOR LEFT VENTRICULAR ASSIST DEVICE (LVAD) RELATED AORTIC REGURGITATION. THE STUDY POPULATION CONSISTED OF ELEVEN PATIENTS WITH LVADS WHO UNDERWENT TAVR FOR AORTIC REGURGITATION WITH A MEDTRONIC COREVALVE OR EVOLUT VALVE. ULTIMATELY, EIGHT OF THE ELEVEN PATIENTS DIED WITHIN FOUR YEARS OF TAVR. CAUSES OF DEATH LISTED BY THE AUTHORS: PROCEDURAL COMPLICATIONS (INITIAL VALVE DISLODGED, SEVERE AORTIC REGURGITATION, SECOND VALVE DISLODGED, VENTRICULAR FIBRILLATION, ATTEMPTED SNARING, AND CARDIOPULMONARY RESUSCITATION) THAT CULMINATED IN DEATH (ONE CASE); CARDIAC ARREST (ONE CASE); PNEUMONIA (ONE CASE); FALL AND INTRACRANIAL HEMORRHAGE (ONE CASE); RIGHT VENTRICLE FAILURE AFTER LVAD EXCHANGE (ONE CASE); RIGHT VENTRICLE FAILURE (ONE CASE); HEART FAILURE PROGRESSION (ONE CASE); AND CARDIOGENIC SHOCK AND SEVERE MITRAL REGURGITATION, WHICH WERE THOUGHT TO BE DUE TO THE PATIENT¿S UNDERLYING CONDITION, NOT THE LATE VALVE MIGRATION AND ISCHEMIA (ONE CASE). COMPLICATIONS ASSOCIATED WITH TAVR THAT DID NOT LEAD TO DEATH: VALVE DISLODGEMENT, VALVE SNARED/REPOSITIONED, ADDITIONAL VALVE REQUIRED, BALLOON DILATION TO FULLY EXPAND VALVE FRAME, TRACE TO SEVERE PARAVALVULAR REGURGITATION, VALVE THROMBUS, STROKE (SUBARACHNOID HEMORRHAGE), HEART TRANSPLANT, AND LEFT/RIGHT FEMORAL ARTERY HEMATOMAS REQUIRING BLOOD TRANSFUSIONS. Manufacturer Narrative: CITATION: GONDI KT, TAM MC, CHETCUTI SJ, ET AL. TRANSCATHETER AORTIC VALVE REPLACEMENT FOR LEFT VENTRICULAR ASSIST DEVICE-RELATED AORTIC REGURGITATION: THE MICHIGAN MEDICINE EXPERIENCE. J SOC CARDIOVASC ANGIOGR INTERV. 2022;2(1):100530. PUBLISHED 2022 NOV 26. DOI: 10.1016/J.JSCAI.2022.100530 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT AND DATE OF DEATH. EARLIEST APPROVED MEDTRONIC TAVR PRODUCT USED FOR PRODUCT CODE AND PMA#. NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT, NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22345279&pc=QZK | 2017233-2025-06362 | 2025-05-01 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-27 04:00 | QZK | GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension; Vascular Dissection | P230023 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT A TAMBE PROCEDURE TO TREAT THORACIC AND INFRARENAL ANEURYSM UTILIZING GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS. REPORTEDLY, ALL DEVICES WERE IMPLANTED SUCCESSFULLY AND PATIENT WAS DOING WELL POST PROCEDURE AS THE FINAL ANGIO RUN SHOWED EVERYTHING LOOKED GOOD. ON (B)(6) 2025, FIVE DAYS POST PROCEDURE, PATIENT PASSED AWAY. SOMETIME PRIOR TO THIS, PATIENT WAS IN ICU AS SHE HAD HYPOTENSION AS PATIENT'S BLOOD PRESSURE DROPPED, X-RAY WAS DONE, WHICH WAS ABNORMAL THEN A CT SCAN WAS DONE WHICH SHOWED A TYPE B DISSECTION AND PATIENT HAD GOTTEN BETTER BUT THEN QUICKLY CODED AND PASSED AWAY. ON (B)(6) 2025, THE PHYSICIAN NOTIFIED FIELD SALES ASSOCIATE OF PATIENT'S PASSING AND REPORTED PER THE INITIAL AUTOPSY REPORTS WHICH ARE UNREMARKABLE BUT THERE IS A TYPE B DISSECTION PROXIMAL TO THE TAMBE DEVICE. THIS DISSECTION WAS NOT PRESENT WHEN THEY DID THE FINAL ANGIO RUN DURING THE ORIGINAL PROCEDURE. IT COULD BE A PROCEDURAL ISSUE BUT EXACT CAUSE OF DEATH IS UNKNOWN AND A MORE COMPREHENSIVE AUTOPSY IS BEING CONDUCTED TO ASCERTAIN THE EXACT CAUSE OF DEATH. Manufacturer Narrative: H6: CODE C19 - A REVIEW OF THE MANUFACTURING RECORDS INDICATED THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE B20 -THE DEVICE REMAINS IMPLANTED AND WAS THEREFORE NOT AVAILABLE FOR ENGINEERING EVALUATION BY GORE. IMAGING EVALUATION SUMMARY: THE IMAGING EVALUATION PERFORMED BY A CLINICAL IMAGING SPECIALIST SHOWED THE FOLLOWING: ¿TWO TIME POINTS AVAILABLE FOR EVALUATION: PRE-IMPLANTATION CTAS DATED (B)(6) 2024 AND (B)(6) 2025. THE CTA IMAGE SET IS STITCHED TOGETHER FOR THE (B)(6) 2025 STUDY. ¿THE PRE-IMPLANTATION CTA DATED (B)(6) 2025 IS AN ABDOMEN/PELVIS CTA SCAN ONLY AND SHOWS: OTHIS TIME-POINT DOES NOT INCLUDE ENOUGH DESCENDING THORACIC AORTA (DTA) TO GIVE PROXIMAL DIAMETERS OF A TAMBE DEVICE CASE PLAN. OTHE DISTAL DTA, ABDOMINAL AORTA, AND COMMON ILIAC ARTERIES ARE HEAVILY CALCIFIED. ¿THE PRE-IMPLANTATION CTA DATED (B)(6) 2025 SHOWS: OTHIS CTA SCAN APPEARS TO BE 2 DIFFERENT SCANS STITCHED TOGETHER. OPROXIMAL DIAMETERS (BETWEEN 45MM ¿ 85MM PROXIMAL TO THE CELIAC ARTERY) APPEAR TO RANGE FROM ~23MM ¿ 25.5MM. OTHERE DOES NOT APPEAR TO BE AN AORTIC DISSECTION IN THE THORACIC OR ABDOMINAL AORTA. ¿IMAGES PROVIDED DO NOT ALLOW FOR EVALUATION IN RELATION TO THE REPORTED COMPLAINT. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22345331&pc=DYB | 3007284313-2025-04057 | 2025-05-30 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-27 04:00 | DYB | GORE® DRYSEAL FLEX INTRODUCER SHEATH | Device Handling Problem | Rupture | K160254 |  | 1 | Event Description: THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT A TEVAR PROCEDURE TO TREAT AN ANEURYSM UTILIZING A GORE® DRYSEAL FLEX INTRODUCER SHEATH. REPORTEDLY, WHEN GETTING READY TO PUT DEVICE IN THE 24FR SHEATH, PHYSICIAN TOOK THE DILATOR OUT AND IN DOING SO RUPTURED THE COMMON AND INTERNAL ILIAC ARTERY LEADING TO PATIENT DEATH. THERE WERE NO ANATOMICAL CHALLENGES WITH TORTUOSITY OR CALCIUM BUT ONE SIDE HAD SMALLER ACCESS SO PHYSICIAN USED OTHER SIDE ACCESS AND PATIENT DEATH IS DUE TO PROCEDURAL ISSUE WITH THE DILATOR. FURTHER INFORMATION IS NOT AVAILABLE. Manufacturer Narrative: C20- A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE CONDUCTED BECAUSE THE SERIAL/LOT NUMBER REMAINS UNKNOWN. IFU STATEMENTS: THE DRYSEAL SHEATH DEVICE INSTRUCTIONS FOR USE (IFU) WERE REVIEWED WITH RESPECT TO THE COMPLAINT DETAIL FOR THE APPLICABLE REGION AND TIME PERIOD, AND THE FOLLOWING IFU STATEMENTS WERE IDENTIFIED IN RELATION TO THE COMPLAINT. WARNINGS: ¿ ADEQUATE VESSEL ACCESS IS REQUIRED TO INTRODUCE THE SHEATH INTO THE VASCULATURE. CAREFUL EVALUATION OF VESSEL SIZE, ANATOMY, TORTUOSITY, AND DISEASE STATE (INCLUDING CALCIFICATION, PLAQUE, AND THROMBUS) IS REQUIRED TO ENSURE SUCCESSFUL SHEATH INTRODUCTION AND SUBSEQUENT WITHDRAWAL. IF VESSEL IS NOT ADEQUATE FOR ACCESS, MAJOR BLEEDING, VESSEL DAMAGE, OR SERIOUS INJURY TO THE PATIENT, INCLUDING DEATH, MAY RESULT. ¿ IF VESSEL SIZE IS SMALLER THAN THE NOMINAL BODY OD (TABLE 1), MAJOR BLEEDING, VESSEL DAMAGE, OR SERIOUS INJURY TO THE PATIENT, INCLUDING DEATH, MAY RESULT. INSTRUCTIONS FOR OPERATION OR USE: SHEATH PREPARATION - 2. VERIFY THE VESSEL IS OF ADEQUATE DIAMETER AND TORTUOSITY TO ACCOMMODATE THE INTRODUCER SHEATH. 8. INSERT THE DILATOR TIP THROUGH THE VALVE AND INTO THE SHEATH UNTIL THE LOCKING CAP ON THE DILATOR IS IN CONTACT WITH THE MATING SURFACE OF THE VALVE. LOCK THE DILATOR TO THE VALVE BODY BY TWISTING THE LOCKING CAP (CLOCKWISE) ON THE DILATOR UNTIL THE POINTER ON THE DILATOR CAP ALIGNS WITH THE ¿LOCK¿ ICON ON THE VALVE BODY. THIS ENSURES THAT THE TAPERED PORTION OF THE DILATOR IS BEYOND THE LEADING END OF THE INTRODUCER SHEATH TIP. THIS WILL OPTIMIZE THE FLEXIBILITY OF THE LEADING END OF THE DILATOR AND ENSURE THAT THE DILATOR STAYS IN PLACE WHILE ADVANCING THE INTRODUCER SHEATH WITH DILATOR INTO THE PATIENT¿S ACCESS VESSEL. 9. HYDROPHILIC COATING ACTIVATION: WET THE ENTIRE OUTER SURFACE OF THE SHEATH TUBE WITH EITHER STERILE SALINE OR WATER TO ACTIVATE THE HYDROPHILIC COATING. NOTE: IT IS IMPORTANT TO KEEP THE SHEATH TUBE OUTER SURFACE WET/SLIPPERY TO TOUCH THROUGHOUT THE PROCEDURE AND DURING INSERTION OF THE INTRODUCER SHEATH. ONCE THE HYDROPHILIC COATING IS ACTIVATED, THE SHEATH OUTER TUBE WILL BE WET AND SLIPPERY. ACTIVATE THE INTRODUCER SHEATH COATING IMMEDIATELY PRIOR TO INSERTING INTO THE VASCULATURE. FOR PROCEDURES OF EXTENDED DURATION, IT MAY BE NECESSARY TO REACTIVATE THE HYDROPHILIC COATING. THIS CAN BE ACHIEVED THROUGH MINOR ROTATIONAL MOVEMENT OF THE SHEATH TO ALLOW BLOOD TO REACTIVATE THE COATING. ENSURE THE DILATOR IS LOCKED TO THE INTRODUCER SHEATH VALVE PRIOR TO INSERTION INTO THE ACCESS VESSEL OR PRIOR TO ANY SHEATH ADVANCEMENT WHILE IN THE PATIENT¿S VASCULATURE. POTENTIAL DEVICE AND PROCEDURE-RELATED ADVERSE EVENTS: POSSIBLE ADVERSE EVENTS AND COMPLICATIONS THAT MAY OCCUR WITH THE USE OF GORE® DRYSEAL FLEX INTRODUCER SHEATH OR ANY INTRODUCER SHEATH OR IN ANY ENDOVASCULAR DEVICE INSERTION PROCEDURE AND WHICH MAY OR MAY NOT REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO (SHOWN IN ENGLISH ALPHABETICAL ORDER): ¿ BLOOD LOSS, BLEEDING, HEMATOMA. ¿ DEATH. ¿ VASCULAR TRAUMA (I.E., DISSECTION, RUPTURE, PERFORATION, TEAR, ETC.) W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22345367&pc=DTQ | 3013876692-2025-00051 | 2025-05-15 04:00 | Death | MAQUET CARDIOPULMONARY GMBH | 2025-06-27 04:00 | DTQ | CARDIOHELP | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Hemorrhage/Bleeding | K133598 |  | 1 | Event Description: MAQUET CARDIOPULMONARY GMBH WAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES ON 2025-06-25: THE VIGILANCE AREA OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT INVOLVING THE PRODUCT ¿CARDIOHELP SYSTEM¿, MANUFACTURED BY GETINGE AND DISTRIBUTED BY, OCCURRED ON 2025-05-15 WITH REPORT NUMBER (B)(4). THE INFORMATION WAS PROVIDED THAT THE PATIENT HAD THE INCIDENT "MAJOR BLEEDING: RETROPERITONEAL HEMORRHAGE" AND THE CONSEQUENCE FOR THE PATIENT WAS "PATIENT PASSED AWAY¿ NO ADDITIONAL INFORMATION WAS RECEIVED. FURTHER INFORMATION ABOUT THE EVENT AND DEVICE WAS REQUESTED BUT IS STILL PENDING. AS THE PATIENT PASSED AWAY, A REPORT IS REQUIRED. COMPLAINT ID (B)(4).Event Description: COMPLAINT ID#: (B)(4). Manufacturer Narrative: A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE SPAIN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K133598. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR CARDIOHELP WITH CATALOG NUMBER 701048012. Manufacturer Narrative: MAQUET CARDIOPULMONARY GMBH WAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES ON 2025-06-25: THE VIGILANCE AREA OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT INVOLVING THE PRODUCT ¿CARDIOHELP SYSTEM¿, MANUFACTURED BY GETINGE AND DISTRIBUTED BY, OCCURRED ON 2025-05-15. THE INFORMATION WAS PROVIDED THAT THE PATIENT HAD THE INCIDENT "MAJOR BLEEDING: RETROPERITONEAL HEMORRHAGE" AND THE CONSEQUENCE FOR THE PATIENT WAS "PATIENT PASSED AWAY¿. NEW INFORMATION WAS RECEIVED ON 2025-06-27 THAT NEITHER THE CARDIOHELP DEVICE NOR THE DISPOSABLE WERE INVOLVED. THE CUSTOMER CONFIRMED THAT THE PATIENT PASSED AWAY FROM HEMORRHAGE DURING A COMPLEX INTERVENTIONAL PROCEDURE AND THAT NO GETINGE EQUIPMENT WAS INVOLVED. BASED ON THE AVAILABLE INFORMATION THE REPORTED EVENT "MAJOR BLEEDING RETROPERITONEAL HEMORRHAGE " COULD NOT BE LINKED TO A DEVICE DEFECT/FAILURE. THEREFORE THE COMPLAINT COULD NOT BE CONFIRMED. THE CONFIRMATION WAS RECEIVED THAT THERE WAS NO PRODUCT MALFUNCTION ON A GETINGE PRODUCT, THEREFORE THIS COMPLAINT IS NOT REPORTABLE. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22345572&pc=NGV | 2135147-2025-03448 | 2025-05-29 04:00 | Death | ABBOTT MEDICAL | 2025-06-27 04:00 | NGV | AMPLATZER AMULET | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | P200049 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 20MM AMPLATZER AMULET LEFT ATRIAL APPENDAGE (LAA) OCCLUDER WAS CHOSEN FOR IMPLANTATION UTILIZING AN UNKNOWN DELIVERY SYSTEM. THE DEVICE WAS IMPLANTED WITHOUT REPORTED ISSUES. IT WAS REPORTED THE PATIENT DIED ON (B)(6) 2025. PATIENT WAS FOUND UNRESPONSIVE AND TRANSPORTED TO EMERGENCY ROOM WHERE THEY WERE DECLARED DECEASED. THE CAUSE OF DEATH WAS CARDIAC ARREST. THE CAUSE OF THE CARDIAC ARREST IS UNKNOWN. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. Manufacturer Narrative: AN EVENT OF CARDIAC ARREST AND SUBSEQUENT DEATH APPROXIMATELY 10 DAYS POST-PROCEDURE WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO SIMILAR COMPLAINTS REPORTED FROM THIS LOT. BASED ON THE AVAILABLE INFORMATION AND MEDICAL REVIEW, THE EXACT CAUSE OF THE CARDIAC ARREST AND DEATH COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22345806&pc=OZD | 1220648-2025-29848 | 2025-06-06 04:00 | Death | ABIOMED, INC. | 2025-06-27 04:00 | OZD | IMPELLA 5.5 | Restricted Flow rate; Malposition of Device | Hemolysis | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 AND EXPERIENCED LOW PUMP FLOW. THE ISSUE WAS RESOLVED BY INCREASING THE P LEVEL OF THE PUMP. THE PATIENT'S LABS WERE HEMOLYZED BUT THE PUMP WAS NOT REPOSITIONED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER. ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE. ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS. ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿. ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION. ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿ Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW, PUMP STOP, POSITIONING ISSUES, AND HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. DATA LOG REVIEW: THE DATA LOG SHOWS A TRENDING PLACEMENT SIGNAL WITHOUT ANY REPORTED MOTOR CURRENT SPIKES DURING THE LOW PUMP FLOW ISSUE WHILE RUNNING ON A STEADY P-LEVEL OF 3. "PUMP POSITION IN AORTA" ALARMS WERE PRESENT DURING THE TIME OF LOW PUMP FLOW. NO MAJOR SUCTION ALARMS WERE REPORTED DURING THE ENTIRE CASE RUN. ACCORDING TO THE CLINICAL NOTES, THE PUMP WAS LEFT SHALLOW DUE TO A SMALL LV, AND BLOOD WAS GIVEN DURING THE NOTED LOW PUMP FLOW; THE PUMP WAS ALSO PROVIDING ECMO SUPPORT. LOW PUMP FLOW: THE CAUSE OF THE LOW PUMP FLOW ISSUE WAS MOST LIKELY RELATED TO THE PATIENT'S CONDITION, BASED ON THE DATA LOG REVIEW AND THE PROVIDED CLINICAL INFORMATION ABOUT THE PATIENT'S CONDITION RELATED TO SMALL LEFT VENTRICLE (LV), LOW BLOOD VOLUME AND EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SUPPORT DURING PUMP USE. POSITIONING ISSUE: THE CAUSE OF THE POSITIONING ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED SMALL LV. HEMOLYSIS: THE CAUSE OF THE HEMOLYSIS ISSUE WAS MOST LIKELY RELATED TO THE PATIENT'S CONDITION, AS THE PATIENT EXPERIENCED LOW BLOOD VOLUME AND POSITIONING ISSUES ACCORDING TO THE CLINICAL DETAILS. THE PLACEMENT SIGNAL TREND ALSO CONFIRMS ASPECTS OF THE PATIENT¿S CONDITION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22345820&pc=OZP | 2032227-2025-199494 | 2025-05-29 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-27 04:00 | OZP | MINIMED 770G US SYSTEM BLE CONNECT 3.0 MG/DL | Adverse Event Without Identified Device or Use Problem | Hyperglycemia; Hypoglycemia; Cancer; Insufficient Information | P150001 |  | 1 | Event Description: IT WAS REPORTED TO MEDTRONIC MINIMED THAT THE CUSTOMER HAD PASSED AWAY ON (B)(6) 2025, AND THE CAUSE OF DEATH WAS MYELODYSPLASIA. THE CUSTOMER REPORTED A BLOOD GLUCOSE VALUE OF 400 MG/DL. THE CUSTOMER REPORTED THE HYPERGLYCEMIA AND HYPOGLYCEMIA. THE EVENT INVOLVED PRODUCTS MMT-332A, MMT-1880, AND MMT-381A. TROUBLESHOOTING WAS PERFORMED AND FOUND THAT THE DOCUMENT OF ANY HEALTH ISSUES OR ILLNESSES THAT MAY HAVE CONTRIBUTED TO OR LED UP TO PASSING WAS THE MYELODYSPLASIA AND DIABETES. THE CUSTOMER WAS STATED THAT THE PUMP WAS NOT WORN AT THE TIME OF PASSING. NO PRODUCT RETURN IS REQUIRED FOR MMT-332A, MMT-1880, AND MMT-381A. Manufacturer Narrative: THIS MDR RELATED TO THE PUERTO RICO MANUFACTURING SITE HAS BEEN ASSIGNED A MEDWATCH NUMBER FROM THE MEDTRONIC MINIMED NORTHRIDGE SITE, PER VARIANCE 5. CURRENTLY IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22345824&pc=DSQ | 2916596-2025-04041 | 2025-06-07 04:00 | Death | THORATEC CORPORATION | 2025-06-27 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTIORGAN DYSFUNCTION. THE MULTIORGAN FAILURE AND SUBSEQUENT PATIENT DEATH WAS NOT CONSIDERED TO BE A NATURAL PROGRESSION OF THE PATIENTS UNDERLYING DISEASE PROCESS FOR WHICH THEY WERE RECEIVING LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE HAD OPERATED AS EXPECTED. THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED EVENT AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22345992&pc=OZD | 1220648-2025-29849 | 2025-06-06 04:00 | Death | ABIOMED, INC. | 2025-06-27 04:00 | OZD | IMPELLA 5.5 | Failure to Advance | No Clinical Signs, Symptoms or Conditions | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT IMPLANTATION OF AN IMPELLA 5.5 FAILED AND COULD NOT BE ADVANCED INTO THE AXILLARY ARTERY. INSTEAD, AN IMPELLA CP WAS IMPLANTED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿. SECTION: WARNINGS & CAUTIONS: WARNINGS - ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION. [ADDITION FOR 5.0, 5.5 ONLY] IN INSTANCES WHERE THE IMPELLA PUMP HAS BEEN PLACED PRIOR TO PERFORMING CARDIAC SURGERY WITH AORTIC CROSS CLAMPING AND CARDIOPLEGIC ARREST, CARE SHOULD BE TAKEN WHEN MANIPULATING THE HEART WHEN THE PUMP POSITION IS FIXED WITH APPLICATION OF THE AORTIC CROSS CLAMP ACROSS THE PUMP CATHETER.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION) WHEN ADVANCING OR TORQUING THE IMPELLA, ADJUSTMENTS SHOULD BE PERFORMED UNDER IMAGING GUIDANCE.¿ SECTION: WARNINGS & CAUTIONS: CAUTIONS - ¿DILATORS AND CATHETERS SHOULD BE REMOVED SLOWLY FROM THE SHEATH. RAPID REMOVAL MAY DAMAGE THE VALVE, RESULTING IN BLOOD FLOW THROUGH THE VALVE.¿ SECTION: WARNINGS & CAUTIONS: WARNINGS - ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE.¿ Manufacturer Narrative: THE INVESTIGATION OF DELIVERY ISSUES HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. BASED ON THE CLINICAL DETAILS THE ROOT CAUSE OF THE DELIVERY ISSUE WAS MOST LIKELY PATIENT CONDITION RELATED, THE DOCTOR TRIED MULTIPLE TIMES TO ADVANCE THE 5.5 INTO THE AXILLARY ARTERY. WE TRIED MULTIPLE INSERTION TRICKS WITH NO SUCCESS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22345996&pc=MFA | 3007284006-2025-00120 | 2025-05-29 04:00 | Death | SPECTRANETICS CORPORATION | 2025-06-27 04:00 | MFA | GLIDELIGHT LASER SHEATH | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation | P960042 |  | 1 | Event Description: A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE TWO RIGHT VENTRICULAR (RV), A RIGHT ATRIAL (RA), AND A LEFT VENTRICULAR (LV) LEAD DUE TO CIED SYSTEM/POCKET INFECTION, REDUNDANT LEAD, AND OCCLUSION. SPECTRANETICS LLD LEAD LOCKING DEVICE (LLDS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 16F GLIDELIGHT LASER SHEATH, THE RA, LV, AND ONE RV LEAD WERE REMOVED. DURING EXTRACTION OF THE LAST RV LEAD, THE GLIDELIGHT DEVICE PERFORATED THE SUPERIOR VENA CAVA (SVC) AND RESCUE EFFORTS BEGAN, INCLUDING STERNOTOMY. THE PERFORATION WAS REPAIRED; HOWEVER, THE SVC WALL HAD BECOME VERY THIN, AND THERE WAS DIFFICULTY IN REPAIRING. THE PATIENT WAS TRANSFERRED TO ICU, BUT UNFORTUNATELY DID NOT SURVIVE THE PROCEDURE. THIS REPORT CAPTURES THE 16F GLIDELIGHT DEVICE IN USE WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. Manufacturer Narrative: A2) PATIENT DATE OF BIRTH - NOT AVAILABLE FROM FACILITY. A5) PATIENT ETHNICITY - NOT AVAILABLE FROM FACILITY. D4) DEVICE SERIAL NUMBER - NOT AVAILABLE FROM FACILITY. H3) THE GLIDELIGHT WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) GREAT VESSEL PERFORATION AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE GLIDELIGHT DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22346217&pc=DSQ | 2916596-2025-03938 | 2025-03-29 04:00 | Death | THORATEC CORPORATION | 2025-06-27 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Renal Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT WAS ADMITTED ON (B)(6) 2025 WITH ACUTE RENAL FAILURE AND CARDIOGENIC SHOCK SECONDARY TO BIVENTRICULAR FAILURE. THEY WERE STABILIZED WITH CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT) AND IMPELLA 5.5, AND THEN THEY WERE IMPLANTED WITH THE HEARTMATE 3 DEVICE. POST OPERATIVELY, THEY WERE UNABLE TO SUCCESSFULLY WEAN OFF INOTROPIC RIGHT VENTRICLE SUPPORT AND REMAINED ON AND OFF PRESSORS THROUGHOUT. THEY WERE ALSO UNABLE TO WEAN OFF OF CRRT DUE TO HEMODYNAMIC INSTABILITY. THEY WERE UNABLE TO WEAN OFF OF THE VENTILATOR AND THEY WERE TRACED. THEY BECAME BACTEREMIA ON (B)(6) 2025 AND HAD BEEN FOLLOWED CLOSELY BY ID. GIVEN INDWELLING DEVICES; THE SITE WAS NOT SURE THEY COULD CLEAR THE INFECTION. OVER THE LAST 24-48 HOURS, PRESSOR REQUIREMENTS CONTINUED TO ESCALATE. THE FAMILY DECIDED NOT TO PURSUE CHEST COMPRESSIONS. THE SITE WAS NOT ABLE TO MAINTAIN AN ADEQUATE PRESSOR ON MAX PRESSOR SUPPORT AND THE PATIENT WAS PRONOUNCED AT 0752 AM. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED, AND THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU, DOCUMENT REV. D AND THE HEARTMATE 3 PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, PUMP POCKET) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22346290&pc=DYB | 3005334138-2025-00399 |  | Death | ABBOTT MEDICAL | 2025-06-27 04:00 | DYB | SWARTZ¿ BRAIDED TRANSSEPTAL GUIDING INTRODUCER, SL1¿, 63 CM LENGTH, 8.5F | Gas/Air Leak | Air Embolism; Stroke/CVA | K052644 |  | 1 | Event Description: DURING A LEFT ATRIAL PROCEDURE, AN AIR EMBOLISM OCCURRED WHICH RESULTED IN A NEUROLOGICAL EVENT. THE VALVE OF THE SHEATH WAS REPORTEDLY LEAKING AND THE PATIENT EXPERIENCED ST-ELEVATION IN ALL LEADS FOLLOWED BY ELECTRO-MECHANICAL DISSOCIATION. REANIMATION MEASURES WERE INITIATED. AFTER THE PATIENT WAS STABILIZED AND ST-ELEVATION RETURNED TO NORMAL, A CORONARY ANGIOGRAM WAS PERFORMED WHICH EXCLUDED CAD AND THROMBOEMBOLISM. THE PATIENT WAS TRANSFERRED TO INTENSIVE CARE AND CURRENT NEUROLOGICAL STATUS IS NOT CLEAR.Event Description: DURING A LEFT ATRIAL PROCEDURE, AN AIR EMBOLISM OCCURRED WHICH RESULTED IN A NEUROLOGICAL EVENT AND THE PATIENT LATER EXPIRED. THE VALVE OF THE SHEATH WAS REPORTEDLY LEAKING AND THE PATIENT EXPERIENCED ST-ELEVATION IN ALL LEADS FOLLOWED BY ELECTRO-MECHANICAL DISSOCIATION. REANIMATION MEASURES WERE INITIATED. AFTER THE PATIENT WAS STABILIZED AND ST-ELEVATION RETURNED TO NORMAL, A CORONARY ANGIOGRAM WAS PERFORMED WHICH EXCLUDED CAD AND THROMBOEMBOLISM. THE PATIENT WAS TRANSFERRED TO INTENSIVE CARE AND CURRENT NEUROLOGICAL STATUS WAS NOT CLEAR. NOTIFICATION WAS RECEIVED ON (B)(6) 2025 INDICATING THAT THE PATIENT EXPIRED ON (B)(6) 2025. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT NO RESPONSE HAS BEEN RECEIVED TO DATE. Manufacturer Narrative: THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED AIR EMBOLISM AND STROKE REMAINS UNKNOWN. Manufacturer Narrative: CORRECTION: B2, H1, H6 (HEALTH IMPACT CODE). ADDITIONAL INFORMATION: B5, G3, H2, H3.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22346291&pc=NPV | 2015691-2025-05246 | 2025-06-04 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-27 04:00 | NPV | ALTERRA ADAPTIVE PRESTENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Great Vessel Perforation | P200015 |  | 1 | Event Description: AS REPORTED BY A FIELD CLINICAL SPECIALIST, THE PATIENT PRESENTED FOR AN ALTERRA CASE. THE PATIENT WITH A HISTORY OF PULMONARY HYPERTENSION WITH PRIOR EVALUATION FOR MITRAL STENOSIS, WAS NOTED TO HAVE SEVERE PULMONARY REGURGITATION AT WORKUP. BASELINE PRESSURES (AT START OF PROCEDURE) OF RIGHT VENTRICLE (RV) 78/11, PULMONARY ARTERY (PA) 77/27, PULMONARY CAPILLARY WEDGE PRESSURE (PCWP) 15, SYSTEMIC ARTERIAL PRESSURE 99/69. FOLLOWING HEMODYNAMICS, A NON-EDWARDS WIRE WAS PLACED TO THE DISTAL RIGHT PULMONARY ARTERY (DRPA) AND THE ALTERRA PRE-STENT WAS ADVANCED THROUGH THE NON-EDWARDS SHEATH, SLIGHTLY DISTAL TO THE ORIGIN OF THE RIGHT PULMONARY ARTERY (RPA) WITHOUT DIFFICULTY. THE PRE-STENT WAS EXPOSED TO APPROXIMATELY 30% AND MAIN PULMONARY ARTERY (MPA) ANGIOGRAPHY WAS PERFORMED. THE ALTERRA WAS THEN WITHDRAWN TO A LEVEL AT/NEAR THE NATIVE PULMONARY ANNULUS. THE ALTERRA WAS FURTHER EXPOSED TO 50% AT WHICH TIME IT WAS NOTED THAT FORWARD PRESSURE WAS APPLIED TO THE ALTERRA DELIVERY SYSTEM (DS); THE IMPLANTER WAS INSTRUCTED TO RELEASE THE OUTER SHAFT OF THE DS, AS THIS WAS CREATING DISTAL MEDIAL/INFERIOR "CANTING" OF THE PRE-STENT, AS WELL AS EXCESSIVE LEFT/LATER DEVICE BIAS. THE ALTERRA WAS FURTHER EXPOSED TO 65%, AT WHICH TIME IT WAS SUGGESTED TO RECAPTURE AND THE PRE-STENT AS DELIVERY OF THE DEVICE WOULD HAVE RESULTED IN CANTING (LACK OF COAXIALITY). ALTERRA WAS SUCCESSFULLY RECAPTURED, READVANCED AND APPEARED TO BE ALIGNED WITHIN THE ANATOMY, AT THE APPROPRIATE LEVEL, WITH THE INFLOW APICES SLIGHTLY INFERIOR TO THE ANNULUS. AT TIME OF RELEASE, IT WAS NOTED PER FLOUROSCOPY THAT FORWARD TENSION REMAINED ON THE SYSTEM, PARTIALLY INVERTING THE INFLOW APICES. THE STENT CONNECTED WAS THEN VERIFIED SEPARATED FROM PRE-STENT AND THE DS WAS REMOVED. THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE WITH DECLINING SYSTEMIC PRESSURES; THEREFORE, CPR WAS INITIATED. A TEE PROBE WAS PLACED, AND A LARGE PERICARDIAL EFFUSION WAS NOTED - THE PATIENT RECEIVED PROTAMINE AT THIS TIME. AN EMERGENCY PERICARDIOCENTESIS WAS PERFORMED, RESULTING IN MARGINAL SYSTEMIC PRESSURES, HOWEVER BLEEDING HAD SLOWED AS NOTED DURING MANUAL ASPIRATION. A MPA ANGIO WAS PERFORMED AND FRANK EXTRAVASATION WAS NOTED AT THE LEVEL OF THE MID/LEFT LATERAL MPA - IT WAS ALSO NOTED THAT THE ALTERRA PRE-STENT HAD MIGRATED DEEPER INTO THE RVOT (PRESUMABLY FROM CHEST COMPRESSIONS). CT SURGERY WAS PRESENT AT THIS TIME AND DESCRIBED OPEN REPAIR AS NOT AN OPTION, DUE TO AGE, COMORBIDITIES, AND COMPLEX NATURE OF A POTENTIAL OPERATIVE RESCUE. AS THE PRESSURE REMAINED LABILE, THE PATIENT WAS NEXT PLACED ON VA ECMO, REQUIRING FULL ANTICOAGULATION AND THE DEGREE OF BLEEDING INTO THE PERICARDIAL SPACE INCREASED DRAMATICALLY, REQUIRING AGGRESSIVE AUTO-TRANSFUSION. MULTIPLE VASCULAR ACCESS SITES WERE OBTAINED AT THIS TIME TO INCLUDE PLACEMENT OF AN ANTEGRADE RIGHT SUPERFICIAL FEMORAL ARTERY (RSFA) SHEATH FOR PERFUSION, DISTAL TO THE FEMORAL AORTIC CANNULA. THE PATIENT WAS THEN TRANSFERRED TO THE ICU IN CRITICAL CONDITION. PER ADDITIONAL INFORMATION THE PATIENT PASSED AWAY THE SAME DAY AS THE PROCEDURE. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22346508&pc=GEI | 9610614-2025-00034 | 2025-05-06 04:00 | Death | ERBE ELEKTROMEDIZIN GMBH | 2025-06-27 04:00 | GEI | ERBE VIO 3 | Appropriate Term/Code Not Available | Bowel Perforation | K190823 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT INCIDENT OCCURRED WITH THE ELECTROSURGICAL UNIT (ESU/GENERATOR) DURING A PARTIAL GASTRECTOMY AND HERNIA REPAIR OF THE ABDOMEN. THE ESU WAS USED WITH AN ERBE LAP BISECT [PART NUMBER (P/N) 20195-204, LOT NUMBER (L/N): INFORMATION NOT PROVIDED (NI) WHICH INCLUDED AN ERBE LAP BISECT INSERT (P/N 20195-206, L/N 2304793)] (NOTE: THE ERBE LAP BISECT AND INSERT ARE NOT CLEARED FOR DISTRIBUTION IN THE U.S. AND ARE NOT DISTRIBUTED IN THE U.S.). NO INFORMATION WAS PROVIDED REGARDING ANY OTHER ACCESSORY USED IN THE PROCEDURE. ALSO, THE SPECIFIC ESU SETTINGS USED WERE NOT PROVIDED. DURING THE OPERATION AND WHEN ACTIVATED IN A CUTTING MODE, THERMALLY INDUCED PERFORATIONS OCCURRED ON A BOWEL LOOP DUE TO TISSUE COMING INTO CONTACT WITH THE NON-INSULATED JOINT OF THE BISECT INSTRUMENT. SUBSEQUENTLY, THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT OF ANOTHER HOSPITAL FOR FURTHER SURGICAL TREATMENT OF THE PERFORATION SITES. HOWEVER, THE PATIENT'S CONDITION DETERIORATED, AND HE DIED ON (B)(6) 2025. Manufacturer Narrative: THE ESU AND LAP BISECT INSERT WERE EVALUATED. THE FINDINGS WERE AS FOLLOWS: ESU.THE GENERATOR WAS INSPECTED/TESTED. THE UNIT WAS FOUND TO BE FUNCTIONING AS INTENDED. THE EVALUATION INCLUDED AN ELECTRICAL SAFETY CHECK, A FUNCTIONAL CHECK OF EACH OF THE EQUIPMENT'S FEATURES AND A POWER OUTPUT CHECK. THE ESU WAS/IS WITHIN SPECIFICATIONS AND ALL FEATURES WERE/ARE FUNCTIONING PROPERLY. IN ADDITION, NO ANOMALIES WERE FOUND IN THE DEVICE HISTORY RECORD (DHR) OF THE ESU. LAP BISECT INSERT AN INSPECTION OF THE INSERT SHOWED SIGNS OF DAMAGE DURING USE. ITS BLACK PTFE INSULATION WAS BADLY DAMAGED IN THE JOINT AREA. TRACES OF MELTING CAN BE SEEN ON THE WHITE PTFE INSULATION IN THE REAR AREA. THE SHAFT END OF THE INSERT WAS HEAVILY SCALED. A MATERIAL/MANUFACTURING DEFECT WAS NOT EVIDENT. IN CONCLUSION, NO ERBE EQUIPMENT PROBLEM WAS FOUND THAT WOULD HAVE CAUSED OR CONTRIBUTED TO THE INCIDENT. THERE WERE MANY FACTORS INVOLVED IN THE EVENT. THE THERMAL LESIONS OF THE BOWEL WERE INITIALLY PRESUMABLY CAUSED BY (UNINTENTIONAL) OR INCORRECT CONTACT OF THE PROXIMAL NON-INSULATED BRANCH PARTS OF THE INSTRUMENT WITH THE BOWEL. THESE AREAS CAN ALSO BECOME VERY HOT, AND CARE MUST BE TAKEN TO ENSURE THAT THERE IS NO UNINTENTIONAL CONTACT WITH TISSUE IN THIS AREA. THE NOTES ON USE FOR THE INSTRUMENT EXPLICITLY STATES THAT THE PRODUCT, INCLUDING ALL INSULATION, MUST BE CHECKED FOR MECHANICAL DAMAGE BEFORE EACH USE AND THAT THE PRODUCT MUST NOT BE USED IN THE EVENT OF DAMAGE. THERE IS ALSO ANOTHER WARNING THAT UNINTENTIONAL TISSUE CONTACT WITH LIVE, NON-INSULATED INSTRUMENT PARTS CAN LEAD TO UNINTENTIONAL COAGULATION AND IRREVERSIBLE TISSUE DAMAGE. THIS ALSO APPLIES TO THE JOINT AREA BEHIND THE INSTRUMENT JAWS. THE INSTRUMENT WAS ALSO CLEANED INTRAOPERATIVELY WITH A SALINE SOLUTION, WHICH SHOULD ALSO BE AVOIDED ACCORDING TO THE ACCESSORY'S INSTRUCTIONS. IN SUMMARY PER THE REPORTED INFORMATION AND THE EVALUATION OF ERBE'S PRODUCTS, THE INCIDENT IS RELATED TO USER ERROR. NO TRENDS HAVE BEEN IDENTIFIED AND ERBE USA, INC. IS NOW CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22346631&pc=LOX | 2124215-2025-42642 | 2025-06-23 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-27 04:00 | LOX | NC EMERGE | Break; Difficult to Remove; Detachment of Device or Device Component | Sepsis; No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: IT WAS REPORTED THAT REMOVAL DIFFICULTY AND TIP DETACHMENT OCCURRED. THE PATIENT PRESENTED WITH ST-ELEVATION MYOCARDIAL INFARCTION AND UNDERWENT PERCUTANEOUS CORONARY INTERVENTION. A 5.0MM X20MM SYNERGY MEGATRON DRUG-ELUTING STENT WAS ADVANCED AND DEPLOYED IN THE RIGHT CORONARY ARTERY. THEN, A 5.00MM X 15MM NC EMERGE BALLOON CATHETER WAS ADVANCED TO POST-DILATE THE LESION. HOWEVER, DURING REMOVAL, THE DEVICE GOT STUCK AT THE SUBCUTANEOUS RIGHT GROIN AND COULD NOT BE REMOVED, AND THE BALLOON TIP FRACTURED. A VASCULAR SURGEON EXPLORED AND HELPED IN REMOVING THE DEVICE FROM THE RIGHT GROIN. ALL THE OTHER DEVICES WERE NON-BOSTON SCIENTIFIC AND WERE REMOVED OUTSIDE OF THE PATIENT. THE LAST IMAGE SHOWED THE STENT DEPLOYED AND POST DILATED, SO THE PHYSICIAN DECIDED TO STOP THE PROCEDURE AND TRANSFERRED THE PATIENT TO THE CARDIAC CARE UNIT. THERE WERE NO PATIENT COMPLICATIONS.Event Description: IT WAS REPORTED THAT REMOVAL DIFFICULTY AND TIP DETACHMENT OCCURRED. THE PATIENT PRESENTED WITH ST-ELEVATION MYOCARDIAL INFARCTION AND UNDERWENT PERCUTANEOUS CORONARY INTERVENTION. A 5.0MM X20MM SYNERGY MEGATRON DRUG-ELUTING STENT WAS ADVANCED AND DEPLOYED IN THE RIGHT CORONARY ARTERY. THEN, A 5.00MM X 15MM NC EMERGE BALLOON CATHETER WAS ADVANCED TO POST-DILATE THE LESION. HOWEVER, DURING REMOVAL, THE DEVICE GOT STUCK AT THE SUBCUTANEOUS RIGHT GROIN AND COULD NOT BE REMOVED, AND THE BALLOON TIP FRACTURED. A VASCULAR SURGEON EXPLORED AND HELPED IN REMOVING THE DEVICE FROM THE RIGHT GROIN. ALL THE OTHER DEVICES WERE NON-BOSTON SCIENTIFIC AND WERE REMOVED OUTSIDE OF THE PATIENT. THE LAST IMAGE SHOWED THE STENT DEPLOYED AND POST DILATED, SO THE PHYSICIAN DECIDED TO STOP THE PROCEDURE AND TRANSFERRED THE PATIENT TO THE CARDIAC CARE UNIT. THERE WERE NO PATIENT COMPLICATIONS. IT WAS FURTHER REPORTED THAT IT WAS THE PROXIMAL SHAFT OF THE BALLOON WHICH WAS DETACHED RATHER THAN THE TIP. FOLLOWING THE PROCEDURE, THE PHYSICIAN RECHECKED THE VESSEL AGAIN AND FOUND THE STENT WAS IN GOOD POSITION. THE PATIENT WAS ADMITTED WITH SEPSIS AND HAD A NECROTIZING FASCIITIS AND DEBRIDEMENT. TWO DAYS LATER, THE PATIENT PASSED AWAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22346915&pc=OZD | 1220648-2025-29852 | 2025-06-01 04:00 | Death | ABIOMED, INC. | 2025-06-27 04:00 | OZD | IMPELLA 5.5 | Positioning Failure | Hemorrhage/Bleeding; Tachycardia; Brain Injury; Ectopic Heartbeat | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT EXPERIENCED SUSTAINED VENTRICULAR TACHYCARDIA (VT). THE IMPELLA WAS REPOSITIONED, AND THE PATIENT WAS GIVEN 500 CUBIC CENTIMETERS OF ALBUMIN, WHICH RESOLVED THE VENTRICULAR TACHYCARDIA. IN ADDITION, THE P LEVEL WAS DECREASED DUE TO THE VT. THEY ALSO NOTED THAT HEPARIN WAS STOPPED DUE TO BLEEDING. THE PATIENT WAS NOT FOLLOWING COMMANDS. HOWEVER, THE PATIENT MOVED ALL FOUR EXTREMITIES. TWO UNITS OF RED BLOOD CELLS WERE ADMINISTERED AND HEPARIN WAS RESTARTED. THE IMPELLA 5.5 WAS PULLED BACK ONE CENTIMETER AFTER ECTOPY. IT WAS NOTED THAT THE PATIENT WAS OFF SEDATION. HOWEVER, THERE WAS NO NEUROLOGICAL RESPONSE. THERE WAS CONCERN FOR COAGULOPATHY AND HEPARIN WAS STOPPED AGAIN. A COMPUTED TOMOGRAPHY SCAN SHOWED ANOXIC BRAIN INJURY. THE IMPELLA 5.5 WAS REPOSITIONED UNDER TRANSESOPHAGEAL ECHOCARDIOGRAPHY DUE TO CONTACT WITH THE MITRAL APPARATUS. ULTIMATELY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22346928&pc=NPT | 2135147-2025-03495 | 2025-06-03 04:00 | Death | ABBOTT MEDICAL | 2025-06-27 04:00 | NPT | NAVITOR VISION | Perivalvular Leak | Low Blood Pressure/ Hypotension; Perforation; Pericardial Effusion | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 25MM NAVITOR VALVE WAS SELECTED FOR IMPLANT. DURING CASE REVIEW PLANNING, IT WAS DISCUSSED THE ANNULAR AND LVOT CALCIUM BURDEN. THE PATIENT'S PERIMETER DERIVED WAS 22.1MM AND THE MINIMUM ON THE ANNULUS MEASURED AT 19.8MM. DURING THE PROCEDURE, THE PATIENT WAS PROPERLY PREPPED, ACCESSED WAS OBTAINED, AND A SUCCESSFUL A PRE-BAV WAS CARRIED OUT WITH A 20MM NON-ABBOTT BALLOON. THE NAVITOR DEVICE WAS THEN INSERTED AND DEPLOYED TO 80% IN CUSP OVERLAP VIEW. THE DEPTH WAS 4MM ON THE NCC AND 5MM ON THE LCC. THE DECISION WAS MADE TO FULLY DEPLOY THE VALVE. THE DEPLOYMENT WAS SUCCESSFUL WITH ALL THREE RETAINER TABS DETACHING. THE FINAL DEPT OF IMPLANT WAS 4MM ON THE NCC AND 4.5MM ON THE LCC. IT WAS NOTED VIA AORTOGRAM THAT THE PATIENT HAD MODERATE PVL. THE PHYSICIAN BEGAN PREPPING A 22MM NON-ABBOTT, DURING WITH TIME THE PATIENTS BLOOD PRESSURE BEGAN DROPPING. THE PHYSICIAN TOOK ANOTHER AORTOGRAM TO ASSESS POTENTIAL CAUSES OF HYPOTENSION. IT WAS NOTED VIA TTE THAT THE PATIENT HAD PERICARDIAL EFFUSION. ANESTHESIA WAS SUPPORTING THE BLOOD PRESSURE WHILE SIMULTANEOUSLY CONVERTING FROM MAC SEDATION TO GENERAL. THE PHYSICIAN PERFORMED A PERICARDIOCENTESIS, DURING WHICH THE PATIENTS BLOOD PRESSURE WAS AROUND 40-50 SYSTOLIC. THE DRAIN WAS SUCCESSFULLY PLACED AND BLOOD WAS PULLED AS THE BLOOD AND ANESTHESIA WERE RAPIDLY REINFUSED INTO THE PATIENT. THE PATIENT STABILIZED AND A FURTHER TTE WAS PERFORMED TO FIND THE SOURCE OF THE BLEED. THE PATIENT WAS STABLE FOR SEVERAL MINUTES, THEN AGAIN BECAME HYPOTENSIVE. AGAIN, THE PATIENT STABILIZED; BUT ONCE MORE BECAME HYPOTENSIVE. DURING THIS ALL, THE GROINS WERE ASSESSED VIA DSA, ALONG WITH RE-ASSESSING THE AORTIC VALVE AND THE AORTIC ROOT. A PIGTAIL WAS PLACED IN THE LEFT VENTRICLE AND AN LV GRAM WAS PERFORMED. THE PHYSICIAN NOTED AN AREA OF CONTRAST JUST OUTSIDE OF THE LEFT CORONARY CUSP OF THE VALVE FRAME IN WHICH HE BELIEVED TO BE AN ANNULAR RUPTURE. IT WAS ALSO NOTED THAT THE VALVE LOOKED TO HAVE EXPANDED MORE, AS THE PREVIOUSLY NOTED MODERATE PVL WAS NO LONGER PRESENT. NO FURTHER CARE COULD BE PERFORMED TO TREAT THE ISSUE. THE PATIENT ENDED UP PASSING AWAY DUE TO ANNULAR RUPTURE. THE PHYSICIAN BELIEVES THAT IT WAS DUE TO THE PATIENTS ANATOMY AND THAT THE PRE-BAV BROKE THE CALCIUM AND THEN A MIX OF THE NAVITOR VALVE¿S NITINOL EXPANDING, AND HYPERTENSIVE EPISODES CAUSED BY ANESTHESIA MADE THE POTENTIAL SMALL TEAR OR RUPTURE WORSE.Event Description: IT WAS REPORTED THAT ON 03 JUN. 2025, A 25MM NAVITOR VALVE WAS SELECTED FOR IMPLANT. DURING CASE REVIEW PLANNING, IT WAS DISCUSSED THE ANNULAR AND LVOT CALCIUM BURDEN. THE PATIENT'S PERIMETER DERIVED WAS 22.1MM AND THE MINIMUM ON THE ANNULUS MEASURED AT 19.8MM. DURING THE PROCEDURE, THE PATIENT WAS PROPERLY PREPPED, ACCESSED WAS OBTAINED, AND A SUCCESSFUL A PRE-BAV WAS CARRIED OUT WITH A 20MM NON-ABBOTT BALLOON. THE NAVITOR DEVICE WAS THEN INSERTED AND DEPLOYED TO 80% IN CUSP OVERLAP VIEW. THE DEPTH WAS 4MM ON THE NCC AND 5MM ON THE LCC. THE DECISION WAS MADE TO FULLY DEPLOY THE VALVE. THE DEPLOYMENT WAS SUCCESSFUL WITH ALL THREE RETAINER TABS DETACHING. THE FINAL DEPT OF IMPLANT WAS 4MM ON THE NCC AND 4.5MM ON THE LCC. IT WAS NOTED VIA AORTOGRAM THAT THE PATIENT HAD MODERATE PVL. THE PHYSICIAN BEGAN PREPPING A 22MM NON-ABBOTT, DURING WITH TIME THE PATIENTS BLOOD PRESSURE BEGAN DROPPING. THE PHYSICIAN TOOK ANOTHER AORTOGRAM TO ASSESS POTENTIAL CAUSES OF HYPOTENSION. IT WAS NOTED VIA TTE THAT THE PATIENT HAD PERICARDIAL EFFUSION. ANESTHESIA WAS SUPPORTING THE BLOOD PRESSURE WHILE SIMULTANEOUSLY CONVERTING FROM MAC SEDATION TO GENERAL. THE PHYSICIAN PERFORMED A PERICARDIOCENTESIS, DURING WHICH THE THE PATIENTS BLOOD PRESSURE WAS AROUND 40-50 SYSTOLIC. THE DRAIN WAS SUCCESSFULLY PLACED AND BLOOD WAS PULLED AS THE BLOOD AND ANESTHESIA WERE RAPIDLY REINFUSED INTO THE PATIENT. THE PATIENT STABILIZED AND A FURTHER TTE WAS PERFORMED TO FIND THE SOURCE OF THE BLEED. THE PATIENT WAS STABLE FOR SEVERAL MINUTES, THEN AGAIN BECAME HYPOTENSIVE. AGAIN, THE PATIENT STABILIZED; BUT ONCE MORE BECAME HYPOTENSIVE. DURING THIS ALL, THE GROINS WERE ASSESSED VIA DSA, ALONG WITH RE-ASSESSING THE AORTIC VALVE AND THE AORTIC ROOT. A PIGTAIL WAS PLACED IN THE LEFT VENTRICLE AND AN LV GRAM WAS PERFORMED. THE PHYSICIAN NOTED AN AREA OF CONTRAST JUST OUTSIDE OF THE LEFT CORONARY CUSP OF THE VALVE FRAME IN WHICH HE BELIEVED TO BE AN ANNULAR RUPTURE. IT WAS ALSO NOTED THAT THE VALVE LOOKED TO HAVE EXPANDED MORE, AS THE PREVIOUSLY NOTED MODERATE PVL WAS NO LONGER PRESENT. NO FURTHER CARE COULD BE PERFORMED TO TREAT THE ISSUE. THE PATIENT ENDED UP PASSING AWAY DUE TO ANNULAR RUPTURE. THE PHYSICIAN BELIEVES THAT IT WAS DUE TO THE PATIENTS ANATOMY AND THAT THE PRE-BAV BROKE THE CALCIUM AND THEN A MIX OF THE NAVITOR VALVE¿S NITINOL EXPANDING, AND HYPERTENSIVE EPISODES CAUSED BY ANESTHESIA MADE THE POTENTIAL SMALL TEAR OR RUPTURE WORSE. Manufacturer Narrative: AN EVENT OF MODERATE PERIVALVULAR LEAK (PVL), PERICARDIAL EFFUSION, HYPOTENSION, AND ANNULAR RUPTURE LEADING TO PATIENT DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE AVAILABLE INFORMATION AND MEDICAL REVIEW, THE CAUSE OF THE REPORTED PVL COULD NOT BE CONCLUSIVELY DETERMINED. THE CAUSE OF THE HYPOTENSION IS LIKELY A CASCADING EFFECT OF THE REPORTED PERICARDIAL EFFUSION. THE CAUSE OF THE REPORTED PERICARDIAL EFFUSION COULD NOT BE CONCLUSIVELY DETERMINED. THE CAUSE OF THE REPORTED DEATH IS LIKELY RELATED TO ANNULAR RUPTURE. THE CAUSE OF THE ANNULAR RUPTURE LIKELY RELATED TO A COMBINATION OF PROCEDURAL (PRE-IMPLANT BALLOON VALVULOPLASTY) AND PATIENT (CALCIFICATION) CONDITIONS. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS A RESULT OF CASE-SPECIFIC CIRCUMSTANCES AS PERICARDIOCENTESIS WAS PERFORMED AND THE CONVERTING FROM MONITORED ANESTHESIA CARE (MAC) SEDATION TO GENERAL ANESTHESIA. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING DESIGN OR MANUFACTURING OF THE DEVICE. Manufacturer Narrative: THE DEVICE REMAINS IMPLANTED IN PATIENT. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22347010&pc=OAE | 2029046-2025-02106 | 2025-06-03 04:00 | Death | BIOSENSE WEBSTER INC | 2025-06-27 04:00 | OAE | QDOT MICRO | Adverse Event Without Identified Device or Use Problem | Low Blood Pressure/ Hypotension | P210027 |  | 1 | Event Description: IT WAS REPORTED DURING AND AFTER ATRIAL FIBRILLATION ABLATION WITH QDOT MICRO AND THE PATIENT EXPERIENCED HYPOTENSION DURING AND AFTER THE PROCEDURE, AND AFTER TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) AND THE PATIENT PASSED AWAY. IT WAS REPORTED THAT THE PATIENT PASSED AWAY WITHIN THE NIGHT AFTER UNDERGOING CARDIAC ABLATION. DURING THE PROCEDURE, THE PATIENT BECAME HYPOTENSIVE. THE PERICARDIAL SPACE WAS EVALUATED WITH THE INTRACARDIAC ECHO CATHETER (ICE). NO PERICARDIAL EFFUSION WAS SEEN ON THE ICE IMAGES. THE ANESTHESIA TEAM TREATED BLOOD PRESSURE ISSUES WITH MEDICATION AND THE PROCEDURE WAS CONTINUED. AT ONE POINT IN THE PROCEDURE, AIR WAS NOTICED WITHIN THE BOSTON SCIENTIFIC FARADRIVE SHEATH. THE PHYSICIAN REMOVED THE AIR AND THE PROCEDURE WAS CONTINUED. AFTER THE ABLATION, THE PATIENT WAS STABLE AND TRANSPORTED TO THE POST OPERATIVE AREA. IT WAS THERE THAT THE PATIENT EXHIBITED SIGNS OF EXTENSIVE HYPOTENSION. THERE WAS LIMITED INFORMATION REGARDING THE EVENT DETAILS INCLUDING MEDICAL INTERVENTIONS THAT WERE PROVIDED TO THE PATIENT. THE PATIENT WAS LATER TRANSPORTED TO THE INTENSIVE CARE UNIT AND LATER PASSED AWAY. THE PATIENT REQUIRED EXTENDED HOSPITALIZATION BECAUSE THE PATIENT WAS RE-INTUBATED AND SENT TO THE ICU. THE PHYSICIAN¿S OPINION OF THE CAUSE OF DEATH WAS UNKNOWN, POSSIBLY PATIENT RELATED. THE OTHER CATHETER USED FOR ABLATION ON THE LEFT ATRIAL SIDE WAS THE BS FARAPULSE CATHETER, AND THE QDOT FJ CURVE WAS USED ONLY FOR THE CTI LINE. FORCE SENSING ABLATION CATHETER WAS USED, AND FORCE VISUALIZATION FEATURES USED WERE GRAPH, DASHBOARD, VECTOR, AND VISITAG. VISITAG MODULE PARAMETERS FOR STABILITY USED WERE 3,3,3 SURPOINT RECOMMENDED SETTINGS. ADDITIONAL FILTER USED WITH THE VISITAG WAS RESPIRATION, AND COLOR OPTION USED PROSPECTIVELY WAS SURPOINT. Manufacturer Narrative: DEVICE INVESTIGATION DETAILS: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. NOTE: FOR FIELD D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REF. # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22347011&pc=LWS | 3004209178-2025-11051 | 2025-04-19 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-27 04:00 | LWS | EVERA MRI XT DR SURESCAN | Failure to Deliver Shock/Stimulation; Battery Problem | Tachycardia; Ventricular Fibrillation | P980016 |  | 1 | Event Description: IT WAS REPORTED, BY A REPRESENTATIVE FOR THE PATIENT, THAT THE PATIENT IS DECEASED. THE PATIENT WAS HOSPITALIZED POST SHOCK DELIVERY FROM THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND THE BATTERY LONGEVITY SHOWED FOUR YEARS. DURING THE HOSPITAL STAY, OVER A TWO-DAY DURATION, THE PATIENT EXPERIENCED VENTRICULAR TACHYCARDIA (VT)/VENTRICULAR FIBRILLATION (VF) EPISODES, AND THE ICD DELIVERED BETWEEN EIGHTY TO ONE HUNDRED SHOCKS; HOWEVER, THE SHOCKS WERE UNABLE TO GET THE PATIENT OUT OF VF. THE PATIENT WAS DISCHARGED EIGHT DAYS AFTER ADMISSION AND PASSED AWAY TWO DAYS LATER. THE CALLER STATED THAT NO ONE CHECKED THE BATTERY LIFE, AND THEY BELIEVED THERE WASN'T ENOUGH BATTERY LIFE FOR THE ICD TO SHOCK THE PATIENT. ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND NOT RECEIVED. Manufacturer Narrative: CONTINUATION OF D10: 407645 LEAD, IMPLANTED: (B)(6) 2006. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22347055&pc=NPT | 2025587-2025-04603 | 2025-06-04 04:00 | Death | MEDTRONIC IRELAND | 2025-06-27 04:00 | NPT | EVOLUT FX DCS | Insufficient Information; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Vascular Dissection | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS EVFXPLUS-29 IN A PATIENT PRESENTED WITH "EXTREMELY HOSTILE" ANATOMY AND THICK CALC IFICATION, A PRE-IMPLANT BALLOON VALVULOPLASTY WAS PERFORMED DUE TO CALCIFICATION. THERE WAS CONCERN THAT CHUNKS OF CALCIUM COULD POTENTIALLY CAUSE OCCLUSION FOLLOWING THE PROCEDURE. A CONTRAST INJECTION WAS USED TO ASSESS CORONARY PERFUSION. THE RIGHT CORONARY AR TERY DEMONSTRATED SLUGGISH AND DELAYED PERFUSION, PROMPTING THE TEAM TO TAKE PROTECTIVE MEASURES. AFTER VALVE DEPLOYMENT, THE PATIENT'S BLOOD PRESSURE DROPPED, AND A POST-DILATION WAS PERFORMED WITH A 22 MM NON-MEDTRONIC BALLOON. A SNORKEL STENT WAS DEPLOYED. DESPITE THESE INTERVENTIONS, THE BLOOD PRESSURE DID NOT IMPROVE, AND AN ASCENDING AORTA DISSECTION EXTENDING TO THE ROOT BY THE RIGHT AND NON-CORONARY CUSPS WAS IDENTIFIED. UPON FURTHER REVIEW, IT APPEARED THE DISSECTION OCCURRED AFTER THE VALVULOPLASTY BUT WAS NOT RECOGNIZED UNTIL AFTER FULL DEPLOYMENT. THE TEAM CONVERTED TO OPEN SURGERY TO REPAIR THE DISSECTION AND REPLACED THE VALVE WITH A NON-MEDTRONIC SURGICAL VALVE. A SURGICAL AORTIC VALVE REPLACEMENT WAS PERFORMED AND THE TRANSCATHETER VALVE WAS EXPLANTED. THE PATIENT REMAINED ON THE TABLE AWAITING WEANING FROM CARDIOPULMONARY BYPASS. THE PROCEDURE WAS DELAYED BY APPROXIMATELY 4-HOURS. THE PATIENT PASSED AWAY THE FOLLOWING DAY. THE PHYSICIAN ASSESSED THAT THE PROCEDURE CONTRIBUTED TO THE DISSECTION, BUT THE MEDTRONIC DEVICE DID NOT.Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS EVFXPLUS-29 IN A PATIENT PRESENTED WITH "EXTREMELY HOSTILE" ANATOMY AND THICK CALCIFICATION, A PRE-IMPLANT BALLOON VALVULOPLASTY WAS PERFORMED DUE TO CALCIFICATION. THERE WAS CONCERN THAT CHUNKS OF CALCIUM COULD POTENTIALLY CAUSE OCCLUSION FOLLOWING THE PROCEDURE. A CONTRAST INJECTION WAS USED TO ASSESS CORONARY PERFUSION. THE RIGHT CORONARY ARTERY DEMONSTRATED SLUGGISH AND DELAYED PERFUSION, PROMPTING THE TEAM TO TAKE PROTECTIVE MEASURES. AFTER VALVE DEPLOYMENT, THE PATIENT'S BLOOD PRESSURE DROPPED, AND A POST-DILATION WAS PERFORMED WITH A 22 MM NON-MEDTRONIC BALLOON. A SNORKEL STENT WAS DEPLOYED. DESPITE THESE INTERVENTIONS, THE BLOOD PRESSURE DID NOT IMPROVE, AND AN ASCENDING AORTA DISSECTION EXTENDING TO THE ROOT BY THE RIGHT AND NON-CORONARY CUSPS WAS IDENTIFIED. UPON FURTHER REVIEW, IT APPEARED THE DISSECTION OCCURRED AFTER THE VALVULOPLASTY BUT WAS NOT RECOGNIZED UNTIL AFTER FULL DEPLOYMENT. THE TEAM CONVERTED TO OPEN SURGERY TO REPAIR THE DISSECTION AND REPLACED THE VALVE WITH A NON-MEDTRONIC SURGICAL VALVE. A SURGICAL AORTIC VALVE REPLACEMENT WAS PERFORMED AND THE TRANSCATHETER VALVE WAS EXPLANTED. THE PATIENT REMAINED ON THE TABLE AWAITING WEANING FROM CARDIOPULMONARY BYPASS. THE PROCEDURE WAS DELAYED BY APPROXIMATELY 4-HOURS. THE PATIENT PASSED AWAY THE FOLLOWING DAY. THE PHYSICIAN ASSESSED THAT THE PROCEDURE CONTRIBUTED TO THE DISSECTION, BUT THE MEDTRONIC DEVICE DID NOT. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT PER THE PHYSICIAN THE VALVE, THE DCS, AND THE GUIDEWIRE DID NOT CONTRIBUTE TO THE VASCULAR INJURY. THE VALVE AND DCS CAN ALSO BE EXCLUDED AS CONTRIBUTING FACTORS TO THE PATIENT'S DEATH. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS A COMBINATION OF AN ASCENDING AORTIC DISSECTION AND A ROOT RUPTURE. OF NOTE, THE PATIENT HAD HEAVILY CALCIFIED ANATOMY THAT WAS DESCRIBED AS "HOSTILE" AND "RISKY". THE PATIENT'S ANATOMY MAY HAVE CONTRIBUTED TO THE CAUSE OF DEATH. Manufacturer Narrative: SECTION D REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER MEDICAL PRODUCTS IN USE DURING THE EVENT INCLUDE: BRAND NAME EVOLUT FX PLUS VALVE; PRODUCT ID EVFXPLUS-29 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. ADDED SECOND PARAGRAPH. D3. D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22347293&pc=DSQ | 2916596-2025-04023 | 2025-05-15 04:00 | Death | THORATEC CORPORATION | 2025-06-27 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY. THE PATIENT WAS FOUND DEAD ON A RUNNING PUMP. NO SYSTEM CONTROLLER WAS RETRIEVED. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE PATIENT¿S OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22347622&pc=NPT | 2025587-2025-04609 | 2021-06-17 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-27 04:00 | NPT | EVOLUT PRO PLUS VALVE | Insufficient Information | Insufficient Information | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER VALVE AORTIC REPLACEMENT PROCEDURE (TAVR), THE PATIENT DIED DUE TO AN UNKNOWN REASON. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: {{D-EVPROP2329US}}; PRODUCT TYPE: {{0195 HEART VALVES}}. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22347740&pc=OZD | 1220648-2025-29857 | 2025-06-06 04:00 | Death | ABIOMED, INC. | 2025-06-27 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Hematoma; Hemolysis; Tachycardia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND HAD A SMALL HEMATOMA AT THE ACCESS SITE AND HEMOLYSIS. A FEMSTOP WAS APPLIED, PRESSURE WAS HELD, AND HEPARIN WAS WITHHELD UNTIL THE SITE WAS STABLE AND THE P LEVEL OF THE PUMP WAS DECREASED. FURTHERMORE, THE PATIENT HAD VENTRICULAR TACHYCARDIA AND WAS DEFIBRILLATED. CARE WAS EVENTUALLY WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE,¿AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA CP WITH SMART ASSIST SYSTEM. SECTION: GENERAL PATIENT CARE CONSIDERATIONS. ¿ASSESS ACCESS SITE FOR BLEEDING AND HEMATOMA.¿ SECTION: ENTERING CARDIAC OUTPUT: USE OF THE REPOSITIONING SHEATH AND PEEL-AWAY INTRODUCER: ¿REMOVE THE PEEL-AWAY INTRODUCER COMPLETELY FROM THE ARTERY OVER THE CATHETER SHAFT TO PREVENT TRAUMA AND SIGNIFICANT BLEEDING AND APPLY MANUAL PRESSURE ABOVE THE PUNCTURE SITE.¿ SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER: ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿ SECTION: SUCTION: ¿SUCTION MAY OCCUR IF THE BLOOD VOLUME AVAILABLE FOR THE IMPELLA CATHETER IS INADEQUATE OR RESTRICTED. SUCTION LIMITS THE AMOUNT OF SUPPORT THAT THE IMPELLA CATHETER CAN PROVIDE TO THE PATIENT AND RESULTS IN A DECREASE IN ARTERIAL PRESSURE AND CARDIAC OUTPUT. IT CAN DAMAGE BLOOD CELLS, LEADING TO HEMOLYSIS. IT MAY ALSO BE AN INDICATOR OF RIGHT HEART FAILURE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22348107&pc=NIK | 2017865-2025-90461 | 2025-06-10 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-27 04:00 | NIK | QUADRA ASSURA MP ICD | Failure to Capture | Insufficient Information; No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PRESENTED REMOTELY VIA MERLIN.NET. IT WAS NOTED THAT IT SEEMED LIKE THERE COULD BE LOSS OF LEFT VENTRICULAR (LV) CAPTURE ON THE IMPLANTABLE CARDIOVERTER DEFIBRILLATOR. THIS WAS TO BE DOCUMENTED AND DISCUSSED WITH THE PHYSICIAN. NO OTHER ISSUES WERE REPORTED.Event Description: NEW INFORMATION RECEIVED NOTES THE PATIENT WAS ADMITTED IN HOSPITAL BETWEEN (B)(6) 2025. THE PATIENT HAD A CONSULT FOR VASCULAR ISSUES IN CLINIC PRIOR TO ADMISSION. THE PATIENT PASSED AWAY (B)(6) 2025. THE CAUSE OF DEATH WAS UNKNOWN. THERE IS NO KNOWN ALLEGATION FROM A HEALTH PROFESSIONAL THAT SUGGESTS THAT THE DEATH WAS DUE TO THE DEVICE/SYSTEM. FURTHER INFORMATION WAS REQUESTED FROM THE CLINICIAN BUT WAS NOT PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22348195&pc=CBK | 2518422-2025-106902 | 2025-06-03 04:00 | Death | RESPIRONICS, INC. | 2025-06-27 04:00 | CBK | TRILOGY 202 VENTILATOR, INTERNATIONAL | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K093905 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING ACUTE NIV (NON-INVASIVE VENTILATION) CIRCUIT HAD DISCONNECTED DUE TO PATIENT PULLING OFF AND STAFF NOTICED ALARMS WERE OFF ON ACUTE NIV MACHINE, THEREFORE DID NOT ALARM TO ALERT STAFF. PATIENT FOUND UNRESPONSIVE WITH VENTILATOR DISCONNECTED. NO ALARMS TO ALERT STAFF OF DISCONNECTION. MET (MEDICAL EMERGENCY TEAM) CALL PUT OUT, UNFORTUNATELY PATIENT PASSED AWAY ON MET TEAM ARRIVAL. VENTILATOR CHECKED WITH PHYSIO NEXT MORNING, PATIENT HAD CHANGED VENTILATION MODES DURING DAY, NO ALARMS HAD BEEN SET, THE PREVIOUS ALARMS DON'T FOLLOW THROUGH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION ALLEGING ACUTE NIV (NON-INVASIVE VENTILATION) CIRCUIT HAD DISCONNECTED DUE TO PATIENT PULLING OFF AND STAFF NOTICED ALARMS WERE OFF ON ACUTE NIV MACHINE, THEREFORE DID NOT ALARM TO ALERT STAFF. PATIENT FOUND UNRESPONSIVE WITH VENTILATOR DISCONNECTED. NO ALARMS TO ALERT STAFF OF DISCONNECTION. MET (MEDICAL EMERGENCY TEAM) CALL PUT OUT, UNFORTUNATELY PATIENT PASSED AWAY ON MET TEAM ARRIVAL. VENTILATOR CHECKED WITH PHYSIO NEXT MORNING, PATIENT HAD CHANGED VENTILATION MODES DURING DAY, NO ALARMS HAD BEEN SET, THE PREVIOUS ALARMS DON'T FOLLOW THROUGH. THE VENTILATOR WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. NO COSMETIC DAMAGE TO DEVICE WAS FOUND AND NO ACTIONABLE ERRORS WERE FOUND IN THE EVENT LOG THAT OCCURRED WHILE IN POSSESSION OF THE CUSTOMER. THE DEVICE WAS TESTED AND AN ERROR CODE RELATED TO A COMMUNICATION ISSUE WAS OBSERVED. ERR_DSP_LOSS_OF_RIPC_COMMUNICATION SIGNIFIES THAT A BRIEF SYSTEM RESET OCCURRED DUE TO DETECTION OF AN ISSUE. THIS IS A SYSTEM RELIABILITY FUNCTION TO KEEP THE SYSTEM WORKING THROUGH TRANSIENT UPSETS. THE SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. THIS EVENT DID NOT OCCUR WHILE IN USE BY THE CUSTOMER AND NO EVIDENCE WAS FOUND REGARDING THIS EVENT BEING RELATED TO THE REPORTED COMPLAINT. THE SUBJECT DEVICE WAS BENCH TESTED AND FUNCTIONED AS DESIGNED. THE FULL DATA EXPORT THAT WAS GIVEN TO PHILIPS RESPIRONICS BY THE CUSTOMER INCLUDED THE ENTIRETY OF THE EVENT LOG EVENTS FROM JULY 21, 2024 TO JULY 1, 2025. THIS DATA IS RECORDED TO THE VENTILATOR INTERNALLY AND DOES NOT REQUIRE AN SD CARD TO BE IN USE DURING EVENTS. DURING AN ANALYSIS OF THE EXPORTED EVENT LOG, IT WAS DETERMINED THAT NONE OF THE ALARM SETTINGS WERE CHANGED BETWEEN JUNE 3, 2025 AND JUNE 25, 2025. ALL USER-SETTABLE ALARMS WERE TURNED OFF (WITH THE EXCEPTION OF THE LOW SPO2 ALARM), INCLUDING THE CIRCUIT DISCONNECT ALARM. SYSTEM ALARMS EXIST ON THE TRILOGY 202 THAT ARE NOT ABLE TO BE TURNED OFF OR HAVE THEIR LIMITS ADJUSTED. IN THE PC MODE THAT WAS BEING USED DURING THE EVENT, ONE OF THESE ALARMS IS LOW PEAK INSPIRATORY PRESSURE (PIP), WHICH CAN OCCUR DURING LARGE LEAKS AND CIRCUIT DISCONNECTS. WHILE THE LOW PIP ALARM CAN BE SEEN IN THE EVENT LOG FOR PRIOR TIMES ON (B)(6) 2025 (THE LAST LOW PIP ALARM ON (B)(6) 2025 OCCURRED AT APPROXIMATELY 14:59 BST), INSTANCES OF THIS ALARM DO NOT APPEAR IN THE TIMELINE IMMEDIATELY SURROUNDING THE REPORTED EVENT. NO EVIDENCE WAS FOUND OF THE REPORTED CIRCUIT DISCONNECT IMMEDIATELY PRIOR TO THE ESTIMATED TIME THE PATIENT WAS FOUND UNRESPONSIVE (21:33 BST ON (B)(6) 2025). ACCORDING TO THE USER MANUAL FOR THE TRILOGY 202, WARNINGS AND CAUTIONS, PAGE 3 STATES: FOR VENTILATOR DEPENDENT PATIENTS, DO NOT RELY ON ANY SINGLE ALARM TO DETECT A CIRCUIT DISCONNECT CONDITION. THE LOW TIDAL VOLUME, LOW MINUTE VENTILATION, LOW RESPIRATORY RATE, AND APNEA ALARMS SHOULD BE USED IN CONJUNCTION WITH THE CIRCUIT DISCONNECT AND LOW PEAK INSPIRATORY PRESSURE ALARMS. TEST THE OPERATION OF THE CIRCUIT DISCONNECT FUNCTION DAILY AND WHENEVER A CHANGE IS MADE TO THE PATIENT CIRCUIT. AN INCREASE IN CIRCUIT RESISTANCE CAN PREVENT PROPER OPERATION OF SOME ALARMS. DO NOT SET THE LOW PEAK INSPIRATORY PRESSURE ALARM TOO LOW, OR THE SYSTEM MAY NOT DETECT LARGE CIRCUIT LEAKS, OR A PATIENT DISCONNECT. BASED ON THE EVALUATION OF THE DEVICE AND DOWNLOADED DEVICE EVENT DATA, PHILIPS RESPIRONICS CONCLUDES THAT THE DEVICE OPERATED AS DESIGNED. THE CUSTOMER REPORTED THAT THE SD CARD WAS NOT FULLY INSERTED INTO THE VENT RESULTING IN MISSING THERAPY DATA AROUND THE REPORTED EVENT TIME, THEREFORE DEVICE CONTRIBUTION TO A USE ERROR THAT MAY HAVE LED TO THE PATIENT DEATH CANNOT BE CONFIRMED OR REFUTED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22349222&pc=BZD | 2518422-2025-106903 | 2025-06-23 04:00 | Death | RESPIRONICS, INC. | 2025-06-27 04:00 | BZD | REMSTAR SYSTEM ONE 60 SERIES | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131982 |  | 1 | Event Description: THE MANUFACTURER ATTEMPTED TO CONTACT THE PATIENT IN REFERENCE TO A REMSTAR AUTO WITH HEATED HUMIDIFIER, SYSTEM ONE 60 SERIES DEVICE REPLACEMENT. THE PATIENT'S SPOUSE ANSWERED THE CALL AND STATED THE PATIENT WAS DECEASED. THE PATIENT'S SPOUSE ENDED THE CALL PRIOR TO ANY ADDITIONAL INFORMATION BEING OBTAINED. DETAILS SURROUNDING THE PATIENT'S DEATH WERE NOT PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22349386&pc=DSQ | 2916596-2025-04087 | 2025-06-13 04:00 | Death | THORATEC CORPORATION | 2025-06-27 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED AND THE DEVICE FUNCTIONED AS EXPECTED. AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22349480&pc=DQA | 3019388613-2025-00048 | 2022-02-01 05:00 | Death | MASIMO - 15750 ALTON PKWY | 2025-06-27 04:00 | DQA | RADICAL-7 | Application Program Freezes, Becomes Nonfunctional | Hypoxia | K212161 |  | 1 | Event Description: PER MEDWATCH MW5170850 THE CUSTOMER REPORTED "A SIX-MONTH-OLD INFANT DIED DURING PSG WHEN A HIDDEN FREEZE MASKED A DROP IN OXYGEN SATURATION." THERE WAS A PATIENT DEATH REPORTED AS A PART OF THIS COMPLAINT. Manufacturer Narrative: OTHER, OTHER TEXT: DEVICE DETAILS FOR THE RADICAL-7 WERE OMITTED FROM THE MEDWATCH REPORT, AND THE DEVICE HAS NOT BEEN RETURNED TO MASIMO FOR INVESTIGATION. AS A RESULT, THE UDI, DEVICE SERIAL NUMBER, AND MANUFACTURE DATE ARE UNKNOWN. AS THE INITIAL REPORTER REQUESTED TO REMAIN ANONYMOUS, MASIMO IS UNABLE TO GATHER ADDITIONAL INFORMATION FROM THE REPORTER. BASED ON THE INFORMATION PROVIDED IN THE MEDWATCH REPORT, MASIMO DOES NOT BELIEVE THAT THIS IS A PROBLEM WITH THE RADICAL-7 DEVICE AND ITS HARDWARE, BUT RATHER AN ISSUE WITH THE USE OF AN UNVALIDATED DIY RS-232 TO USB CABLE USED TO STREAM SPO2 VALUES INTO THE PSG SYSTEM. IT WAS REPORTED THAT DURING PSG TESTING, THE UNVALIDATED CABLE EXPERIENCED OVER 120 COMMUNICATION DROPOUTS. WHEN THESE DROPOUTS OCCURRED, THE PSG TRACE FROZE ON THE LAST VALUE RESULTING IN HIDDEN DESATURATION EVENTS. MASIMO HAS DETERMINED THAT THE USE OF AN UNVALIDATED DIY CABLE, COMBINED WITH INAPPROPRIATE DEVICE SETTINGS, IMPROPER SYSTEM INTEGRATION INTO THE PSG PLATFORM AND A LACK OF INVESTIGATION BY THE CUSTOMER DESPITE KNOWN AND DOCUMENTED ISSUES CONTRIBUTED TO THE REPORTED EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22349552&pc=DSQ | 2916596-2025-04010 | 2025-06-05 04:00 | Death | THORATEC CORPORATION | 2025-06-27 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hypoxia; Ischemia; Respiratory Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO ACUTE HYPOXEMIC RESPIRATORY FAILURE, BOWEL ISCHEMIA. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. AN AUTOPSY WOULD NOT BE PERFORMED. THE PUMP WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6) , AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV D, AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV A, ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH AND RESPIRATORY FAILURE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22349606&pc=DSQ | 2916596-2025-04128 | 2025-06-14 04:00 | Death | THORATEC CORPORATION | 2025-06-27 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT WAS ON HOSPICE. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WAS NOT RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. DUE TO HOSPITAL POLICY, THE CUSTOMER COMMUNICATED THAT NO ADDITIONAL INFORMATION WILL BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISIONS OF THE IFU CAN ALSO BE FOUND ON THE ELECTRONIC IFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22349655&pc=NIQ | 2024168-2025-06626 | 2021-12-01 05:00 | Death | ABBOTT VASCULAR | 2025-06-27 04:00 | NIQ | XIENCE PRIME EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P110019 |  | 1 | Event Description: IN A SUMMARY ARTICLE SITING THE ILUMIEN IV CLINICAL TRIAL IT WAS REPORTED THAT USE OF XIENCE EVEROLIMUS-ELUTING STENTS (EES) ARE SAFE AND EFFECTIVE FOR TREATING SINGLE-LAYER IN-STENT RESTENOSIS (ISR). THE ARTICLE REFERS TO THE RIBS IV CLINICAL TRIAL WHERE IMPLANTATION OF THE XIENCE PRIME STENT LED TO A SIGNIFICANTLY LARGER MINIMUM LUMEN DIAMETER AT 6 TO 9 MONTH FOLLOW-UP RESULTING IN A LOWER RATE OF CARDIAC DEATH, MYOCARDIAL INFARCTION AND TARGET-VESSEL REVASCULARIZATION AT 1 YEAR FOLLOW-UP. THE RESTENT--ISR CLINICAL TRIAL IS ALSO REFERENCED, NOTING IN THE 3 YEAR COMPOSITE POST IMPLANTATION OF THE XIENCE V STENT, DEATH, MYOCARDIAL INFARCTION, STENT THROMBOSIS AND TARGET-VESSEL REVASCULARIZATION OCCURRED. DETAILS ARE LISTED IN THE ARTICLE "SAFETY AND EFFICACY OF COBALT CHROMIUM EVEROLIMUS-ELUTING STENTS FOR TREATMENT OF IN-STENT RESTENOSIS: AN ILUMIEN IV SUBSTUDY". NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORDS FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PART NUMBER AND/OR LOT NUMBER WAS NOT REPORTED, AND THE DEVICE WAS NOT RETURNED. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT DEATH, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE REPORTED PATIENT DEATH IS LISTED IN THE XIENCE PRIME, XIENCE PRIME SMALL VESSEL (SV), AND XIENCE PRIME LONG LENGTH (LL), EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2 - DATE OF DEATH ESTIMATED. B3 - DATE OF EVENT ESTIMATED. D4- THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL DEVICE REPORTED IN B5 IS CAPTURED UNDER SEPARATE MEDWATCH REPORT. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE ARE CAPTURED UNDER SEPARATE MEDWATCH REPORT. ATTACHMENT ARTICLE TITLED: "SAFETY AND EFFICACY OF COBALT CHROMIUM EVEROLIMUS-ELUTING STENTS FOR TREATMENT OF IN-STENT RESTENOSIS: AN ILUMIEN IV SUBSTUDY".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22349724&pc=NGV | 2124215-2025-42595 | 2025-06-13 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-27 04:00 | NGV | WATCHMAN FLX? PRO | Adverse Event Without Identified Device or Use Problem | Cardiogenic Shock; Syncope/Fainting | P130013 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED, AND A 24 MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED ON (B)(6) 2025. THE PATIENT WAS DISCHARGED. ON (B)(6) 2025 THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT (ED) FOR SYNCOPE. ON (B)(6) 2025 THE PATIENT PASSED AWAY FROM A CARDIOGENIC SHOCK.Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED, AND A 24MM WATCHMAN FLX PRO CLOSURE DEVICE WAS IMPLANTED ON (B)(6) 2025. THE PATIENT WAS DISCHARGED. ON (B)(6) 2025 THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT (ED) FOR SYNCOPE. ON (B)(6) 2025 THE PATIENT PASSED AWAY FROM A CARDIOGENIC SHOCK. IT WAS FURTHER REPORTED THAT WHILE THE PATIENT WAS IN THE HOSPITAL, PRESENTED ALTERED MENTAL STATUS AND EXPERIENCED ATRIAL ARRHYTHMIAS DRIVEN BY UNDERLYING EMPHYSEMA, FOLLOWING A PRIOR ABLATION PROCEDURE. Manufacturer Narrative: B5 DESCRIBE EVENT OR PROBLEM: UPDATED WITH ADDITIONAL INFORMATION RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22349895&pc=NIQ | 2024168-2025-06633 | 2021-12-01 05:00 | Death | ABBOTT VASCULAR | 2025-06-27 04:00 | NIQ | XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEM | Adverse Event Without Identified Device or Use Problem | Unspecified Heart Problem | P070015 |  | 1 | Event Description: IN A SUMMARY ARTICLE SITING THE ILUMIEN IV CLINICAL TRIAL IT WAS REPORTED THAT USE OF XIENCE EVEROLIMUS-ELUTING STENTS (EES) ARE SAFE AND EFFECTIVE FOR TREATING SINGLE-LAYER IN-STENT RESTENOSIS (ISR). THE ARTICLE REFERS TO THE RIBS IV CLINICAL TRIAL WHERE IMPLANTATION OF THE XIENCE PRIME STENT LED TO A SIGNIFICANTLY LARGER MINIMUM LUMEN DIAMETER AT 6 TO 9 MONTH FOLLOW-UP RESULTING IN A LOWER RATE OF CARDIAC DEATH, MYOCARDIAL INFARCTION AND TARGET-VESSEL REVASCULARIZATION AT 1 YEAR FOLLOW-UP. THE RESTENT--ISR CLINICAL TRIAL IS ALSO REFERENCED, NOTING IN THE 3 YEAR COMPOSITE POST IMPLANTATION OF THE XIENCE V STENT, DEATH, MYOCARDIAL INFARCTION, STENT THROMBOSIS AND TARGET-VESSEL REVASCULARIZATION OCCURRED. DETAILS ARE LISTED IN THE ARTICLE "SAFETY AND EFFICACY OF COBALT CHROMIUM EVEROLIMUS-ELUTING STENTS FOR TREATMENT OF IN-STENT RESTENOSIS: AN ILUMIEN IV SUBSTUDY". NO OTHER INFORMATION Manufacturer Narrative: THE DEVICE WAS NOT RETURNED FOR ANALYSIS. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A QUERY OF THE COMPLAINT HANDLING DATABASE FOR THE REPORTED LOT REVEALED THERE IS NO INDICATION OF A LOT SPECIFIC ISSUE. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT(S), AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THE REPORTED PATIENT EFFECT(S) OF DEATH IS LISTED IN THE XIENCE V EVEROLIMUS ELUTING CORONARY STENT SYSTEMS INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT(S) OF CORONARY STENTING PROCEDURES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2 - DATE OF DEATH IS ESTIMATED. B3 - DATE OF EVENT ESTIMATED. D4- THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. D6A - IMPLANT DATE ESTIMATED. THE ADDITIONAL DEVICE REPORTED IN B5 IS CAPTURED UNDER SEPARATE MEDWATCH REPORT. THE ADDITIONAL PATIENT EFFECTS REPORTED IN THE ARTICLE IS CAPTURED UNDER SEPARATE MEDWATCH REPORT. ATTACHMENT ARTICLE TITLED: "SAFETY AND EFFICACY OF COBALT CHROMIUM EVEROLIMUS-ELUTING STENTS FOR TREATMENT OF IN-STENT RESTENOSIS: AN ILUMIEN IV SUBSTUDY".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22349907&pc=DSQ | 2916596-2025-04221 | 2023-06-15 04:00 | Death | THORATEC CORPORATION | 2025-06-27 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT CONSIDERED TO BE DEVICE OR THERAPY RELATED. NO AUTOPSY WAS PERFORMED. THE DEVICE WAS NOT EXPLANTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22350288&pc=OZD | 1220648-2025-29862 | 2025-06-02 04:00 | Death | ABIOMED, INC. | 2025-06-27 04:00 | OZD | IMPELLA CP | Restricted Flow rate | Tachycardia | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT HAD VENTRICULAR TACHYCARDIA (VT) DURING MANIPULATION AT THE AORTIC VALVE. THE PUMP POSITION WAS ASSESSED AND VERIFIED. THE PUMP ALSO HAD LOW FLOWS DUE TO MISSING PRELOAD. THE PATIENT EXPIRED WHEN PATIENT WENT ASYSTOLIC AS CIRCULATION WAS NOT ABLE TO BE MAINTAINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER; THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22350523&pc=NAY | 2955842-2025-27881 | 2025-05-26 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-06-27 04:00 | NAY | DAVINCI X | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Pain | K171294 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI-ASSISTED BILATERAL INGUINAL HERNIA REPAIR, THE PATIENT EXPERIENCED SUDDEN, UNEXPLAINED PAIN AND RETURNED TO THE HOSPITAL IN CARDIAC ARREST. THE SITE ROBOTICS COORDINATOR REPORTED THAT THE SURGICAL PROCEDURE WAS COMPLETED WITHOUT ANY INTRAOPERATIVE COMPLICATIONS. THE PATIENT WOKE UP WELL, WAS EXTUBATED, AND WAS TRANSFERRED TO THE ICU AS A PRECAUTION DUE TO COMORBIDITIES AND ADVANCED AGE. ON POSTOPERATIVE DAY (POD) 3, THE PATIENT WAS TRANSFERRED OUT OF THE ICU AND REMAINED HOSPITALIZED FOR AN ADDITIONAL DAY BEFORE BEING DISCHARGED HOME ON POD 4 IN STABLE CONDITION. ON POD 8, THE PATIENT EXPERIENCED A SUDDEN ONSET OF UNSPECIFIED PAIN AND RETURNED TO THE HOSPITAL IN CARDIAC ARREST. AFTER APPROXIMATELY 40 MINUTES OF RESUSCITATION EFFORTS, THE PATIENT EXPIRED. THE CAUSE OF DEATH IS SUSPECTED TO BE EITHER A HEART ATTACK OR A PULMONARY EMBOLISM. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, LARGE NEEDLE DRIVER, FENESTRATED BIPOLAR FORCEPS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED 8 DAYS FOLLOWING A REPORTEDLY UNEVENTFUL ROBOTIC HERNIA REPAIR. THE PATIENT RETURNED WITH PAIN BUT THE LOCATION AND TYPE OF PAIN IS NOT PROVIDED. A PULMONARY EMBOLISM OR CARDIAC EVENT WAS SPECULATED TO HAVE OCCURRED BUT NO SUPPORTING INFORMATION FOR THOSE POTENTIAL CAUSES WAS PROVIDED. NO ALLEGATION AGAINST ANY INTUITIVE PRODUCT OR INSTRUMENT WAS MADE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22350559&pc=MNS | 2518422-2025-106887 | 2025-06-24 04:00 | Death | RESPIRONICS, INC. | 2025-06-27 04:00 | MNS | RENTAL, DREAMSTATION ST30, GB | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: MANUFACTURER RECEIVED A CALL FROM THE CLINIC THAT THE PATIENT PASSED AWAY. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22350564&pc=BZD | 2518422-2025-106883 | 2025-06-24 04:00 | Death | RESPIRONICS, INC. | 2025-06-27 04:00 | BZD | PHILIPS CPAP | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K131982 |  | 1 | Event Description: MANUFACTURER RECEIVED A CALL FROM THE PATIENT'S SISTER STATING THAT HER BROTHER WAS ON THE DEVICE FOR YEARS BUT PASSED AWAY IN 2012 AND WAS REQUESTING FOR THE SERIAL NUMBER OF HER BROTHER'S DEVICE, THE CALLER WAS TRANSFERRED TO PATIENT RECALL SUPPORT. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22350607&pc=PNG | 9680904-2025-00017 | 2025-05-06 04:00 | Death | CAREFUSION, INC | 2025-06-27 04:00 | PNG | 1000ML PLEURX SPARE BOTTLE LATAM | Adverse Event Without Identified Device or Use Problem | Insufficient Information | NA |  | 1 | Event Description: THERE WAS NO REPORTED PATIENT IMPACT. HOWEVER, THERE WAS AN ABSENCE OF VACUUM, WHICH WAS DETECTED DURING INSTALLATION. DESPITE THE ISSUE, DRAINAGE WAS SUCCESSFULLY PERFORMED USING OTHER BOTTLES. THE CLAMP WAS PROPERLY CLOSED UNTIL AFTER THE PLUNGER WAS PRESSED. THE CATHETER WAS LOCATED IN THE ABDOMEN. IT IS UNCLEAR WHETHER THE EVENT DIRECTLY OR INDIRECTLY INVOLVED THE PATIENT OR USER. THE CATHETER WAS IMPLANTED ON (B)(6) 2025, AN ISSUE WITH THE BOTTLE OCCURRED ON (B)(6) 2025, FOLLOWED BY A COMPLAINT ON (B)(6) 2025, WHICH IS BEING RECORDS IN THIS COMPLAINT. THE PATIENT PASSED AWAY IN (B)(6) 2025. THE RELATIONSHIP BETWEEN THE PATIENT'S DEATH, THE DEVICE USE, AND THE PROCEDURE WAS NOT PROVIDED. Manufacturer Narrative: H10: MANUFACTURING REVIEW: A DEVICE HISTORY RECORD COULD NOT BE EVALUATED AS THE LOT NUMBER IS UNKNOWN. H10: INVESTIGATION SUMMARY: NO PHOTOS OR SAMPLES WERE RECEIVED BY OUR QUALITY TEAM FOR EVALUATION THEREFORE THE FAILURE MODE COULD NOT BE VERIFIED. THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE MODE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED FAILURE COULD NOT BE DETERMINED BASED UPON PROVIDED INFORMATION. IT IS UNKNOWN WHETHER PATIENT AND/OR PROCEDURAL ISSUES CONTRIBUTED TO THE REPORTED EVENT. H10: B3: DATE OF EVENT: THE DATE OF AWARENESS WAS ENTERED IN THE DATE OF EVENT FIELD AS THE ACTUAL DATE OF EVENT IS UNKNOWN. H10: B2: DATE OF DEATH: THE REPORTED DATE WAS (B)(6) 2025, THEREFORE, THE DATE ENTERED IS (B)(6) 2025. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22350755&pc=NPT | 2135147-2025-03526 | 2025-06-05 04:00 | Death | ABBOTT MEDICAL | 2025-06-27 04:00 | NPT | NAVITOR VISION | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Vascular Dissection | P190023 |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2025, A 27MM NAVITOR VALVE WAS SELECTED FOR IMPLANT USING A LARGE FLEXNAV DELIVERY SYSTEM. NO EFFUSION WAS OBSERVED AND THE PATIENT WAS IN STABLE CONDITION PRIOR TO PROCEDURE. PRE-DILATATION WAS PERFORMED WITH A 22MM BALLOON; "LESS THAN PERIMETER DERIVED AS PER IFU DUE TO SIGNIFICANT CALCIUM." DURING DEVICE PREPARATION, THE VALVE WAS PREPPED AND LOADED PER IFU. APPROXIMATELY 30 SECONDS POST-DEPLOYMENT, THE PATIENT LOST PRESSURE AND CIRCUMFERENTIAL PERICARDIAL EFFUSION WAS IDENTIFIED. CARDIOPULMONARY RESUSCITATION WAS COMMENCED AND PERICARDIOCENTESIS WAS PERFORMED. HOWEVER, THE PATIENT PASSED DUE TO ANNULAR RUPTURE. THE USER BELIEVES THAT THE "PRE-DILATATION BALLOON INITIATED THE ANNULAR RUPTURE FOLLOWED BY THE NAVITOR VALVE DEPLOYMENT." Manufacturer Narrative: INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22351719&pc=NAY | 2955842-2025-27887 | 2025-05-28 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-06-27 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Adhesion(s); Hemorrhage/Bleeding | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED RADICAL CYSTECTOMY AND URINARY DIVERSION PROCEDURE, THE PATIENT EXPIRED POST-OPERATIVELY DUE TO BLEEDING. THE PATIENT REPORTEDLY HAD SEVERE ADHESIONS, WAS UNDERGOING CHEMOTHERAPY, AND HAD FRAGILE BLOOD VESSELS. THE SURGEONS CONVERTED TO LAPAROTOMY, BUT COULD NOT IDENTIFY THE SOURCE OF THE BLEEDING. NO ADDITIONAL INFORMATION WAS PROVIDED. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION WERE MADE; NO RESPONSE WAS RECEIVED. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 8MM 30 DEGREE ENDOSCOPE PLUS, 0 DEGREE ENDOSCOPE, MONOPOLAR CURVED SCISSORS, FENESTRATED BIPOLAR FORCEPS, PROGRASP FORCEPS, LARGE SUTURECUT NEEDLE DRIVER, VESSEL SEALER EXTEND X2. A SITE HISTORY REVIEW FOUND THAT ALL MULTIPLE USE INSTRUMENTS WERE USED IN SUBSEQUENT PROCEDURES WITH NO REPORTED ISSUES THAT WOULD BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AFTER A RADICAL CYSTECTOMY. BLEEDING IS SUGGESTED AS THE CAUSE OF DEATH AND POTENTIAL CONTRIBUTING FACTORS ARE SUGGESTED BUT HOW AND WHEN THE BLEEDING OCCURRED IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22352191&pc=FMI | 9681413-2025-00003 | 2025-04-06 04:00 | Death | TERUMO EUROPE N.V | 2025-06-27 04:00 | FMI | K-PACK II | Delivered as Unsterile Product | Shock | K984576 |  | 1 | Event Description: THE USER FACILITY REPORTED AN INCIDENT TO THE HEALTH AUTHORITY OF POLAND FOLLOWING THE DEATH OF A CHILD AFTER A VACCINATION. TO ENSURE THE SAFETY OF PATIENTS, AN IMMEDIATE REVIEW OF THE BATCH RECORDS AND A THOROUGH INVESTIGATION FOCUSING ON THE STERILITY OF THE TWO NEEDLE LOTS IN QUESTION HAVE BEEN REQUESTED. THE EVENT OCCURRED POST-TREATMENT, AND THE FINAL PATIENT IMPACT WAS DEATH. Manufacturer Narrative: D4: LOT#: 2306025 OR 2311027. D6A: IMPLANTED DATE: DEVICE WAS NOT IMPLANTED. D6B: EXPLANTED DATE: DEVICE WAS NOT EXPLANTED. E3: OCCUPATION: SPECIALIST, SPU IM OPS BE, PRIMARY PACK. H4: DEVICE MANUFACTURE DATE: 2023-06-25 OR 11-27-2023. THE ACTUAL DEVICE IS NOT AVAILABLE FOR RETURN. THE BATCH RECORDS FOR LOT NUMBERS: 2306025 AND 2311027 (KN2516RB05) WERE REVIEWED. DURING NEEDLE ASSEMBLY, AS PART OF IN-PROCESS CONTROL, 40 ASSEMBLED NEEDLES (NEEDLE ASSEMBLED WITH CASE) WERE VISUALLY INSPECTED TWICE PER SHIFT (8-HOUR SHIFTS DURING THE WEEK AND 12-HOUR SHIFTS DURING THE WEEKEND) FOR FOREIGN MATTER AND DAMAGE TO THE CASE. ONE NON-CONFORMITY WAS RECORDED FOR CASE DAMAGE; HOWEVER, THE DAMAGE WAS SUPERFICIAL AND POSED NO RISK OF CRACKING THE CASE. DURING PACKAGING, 40 LABELED PRODUCTS WERE VISUALLY INSPECTED TWICE PER SHIFT (8-HOUR SHIFTS DURING THE WEEK AND 12-HOUR SHIFTS DURING THE WEEKEND) FOR FOREIGN MATTER AND DAMAGE TO THE CASE AND CAP. NO DEVIATIONS OR ABNORMALITIES RELATED TO THE REPORTED COMPLAINT WERE OBSERVED DURING PRODUCTION. DURING FINAL INSPECTION, WHICH IS PART OF RELEASE TESTING, 1,250 UNITS WERE EXAMINED AND NO DEFECTS WERE DETECTED. NO RELEVANT DEVIATIONS OR ABNORMALITIES WERE NOTED IN THE BATCH RECORDS OR DURING THE INSPECTIONS DESCRIBED ABOVE THAT WOULD RELATE TO THE REPORTED COMPLAINT. LIMULUS AMEBOCYTE LYSATE (LAL) TESTING AND STERILIZATION CYCLES FOR BOTH LOTS WERE PERFORMED ACCORDING TO SPECIFICATIONS, WITH NO OBSERVATIONS NOTED. THE 2023 BIOBURDEN MONITORING PROGRAM WAS REVIEWED, AND RESULTS WERE CONSISTENT WITH PREVIOUS YEARS. MONITORING SAMPLES FROM JUNE 2023 AND NOVEMBER 2023 (CLOSEST TO THE MANUFACTURING DATES) DID NOT SHOW ANY DEVIATIONS. ADDITIONALLY, NO MATERIALS OF ANIMAL ORIGIN WERE USED. ALL REMAINING RETENTION SAMPLES WERE VISUALLY INSPECTED FOR FOREIGN MATTER AND BROKEN OR DAMAGED CASES WITH THE NAKED EYE ON 06/05/2025. NO DEFECTS WERE FOUND. NO COMPLAINT SAMPLE WAS RECEIVED; THEREFORE, NO DIRECT INVESTIGATION OF A SAMPLE COULD BE PERFORMED. NON-CONFORMITIES RELATED TO THE FINISHED PRODUCT BATCHES AND SEMI-FINISHED NEEDLE ASSEMBLY BATCHES WERE REVIEWED. NO RELATED NON-CONFORMITIES WERE IDENTIFIED, EXCEPT FOR ONE: A NON-CONFORMITY WAS RECORDED FOR COSMETIC DAMAGE (NO CRACK) TO THE CASE DURING NEEDLE ASSEMBLY ON LINE NA7. THE AFFECTED PRODUCTION WAS CONTAINED AND SCRAPPED. NOTIFICATIONS RELATED TO THE FINISHED PRODUCT AND SEMI-FINISHED NEEDLE ASSEMBLY BATCHES WERE ALSO REVIEWED. NO RELATED NOTIFICATIONS WERE FOUND. THERE IS NO CORRECTIVE AND PREVENTIVE ACTION (CAPA) ASSOCIATED WITH ISSUES INVOLVING THE STERILE BARRIER SYSTEM OF THE PRODUCTS. BASED ON THE INVESTIGATION, INCLUDING BATCH RECORD REVIEW AND RETENTION SAMPLE INSPECTION, NO ROOT CAUSE RELATED TO THE PRODUCTION PROCESS WAS IDENTIFIED. FOR INFORMATION REGARDING THE IMPORTER REPORT FOR THIS EVENT, PLEASE REFER TO SECTION H10.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22352260&pc=LYJ | 1644487-2025-10300 | 2020-04-14 04:00 | Death | LIVANOVA USA, INC. | 2025-06-27 04:00 | LYJ | PULSE GEN MODEL 103 | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P970003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AFTER A GRAND MAL SEIZURE. NO OTHER RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS RECEIVING THERAPY AT THE TIME OF DEATH AND EXPERIENCING SEIZURE REDUCTION WITH VNS. THE PATIENT'S CAUSE OF DEATH WAS UNKNOWN. Manufacturer Narrative: D7A, CORRECTED DATA: INITIAL REPORT INADVERTENTLY SUBMITTED WITHOUT REPORTING THIS INFORMATION. E4, CORRECTED DATA: INITIAL REPORT INADVERTENTLY SUBMITTED WITHOUT REPORTING THIS INFORMATION. Manufacturer Narrative: LIVANOVA USA, INC. SUBMITS THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION, BASED ON INFORMATION THAT LIVANOVA HAS OBTAINED, BUT MAY NOT HAVE BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION, OR A CONCLUSION BY FDA OR ANYONE ELSE, THAT THE DEVICE, LIVANOVA, OR LIVANOVA'S EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS¿ OR "MALFUNCTIONS¿. THESE WORDS ARE INCORPORATED INTO THE FDA 3500A MEDWATCH FORM BY THE FDA, AND LIVANOVA OBJECTS TO THEIR USE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22352521&pc=QFG | 3019004087-2025-00614 | 2025-05-25 04:00 | Death | BETA BIONICS, INC | 2025-06-27 04:00 | QFG | ILET BIONIC PANCREAS | Insufficient Information | Insufficient Information | K231485 |  | 1 | Event Description: ON 27-MAY-2025, A CERTIFIED DIABETES CARE AND EDUCATION SPECIALIST (CDCES) NOTIFIED BETA BIONICS THAT THE USER HAD PASSED AWAY OVER THE WEEKEND. FOLLOW-UP ON (B)(6) 2025 CONFIRMED THE USER PASSED AWAY ON (B)(6) 2025. AT THE TIME OF DISCOVERY, THE USER WAS WEARING THE ILET, BUT THE INFUSION SET WAS FOUND TO BE PARTIALLY DISLODGED. THE REPORTED CAUSE OF DEATH WAS NATURAL CAUSES AND HIGH BLOOD PRESSURE. NO AUTOPSY WAS PERFORMED. REVIEW OF DEVICE DATA CONFIRMS THE ILET WAS IN USE THROUGH (B)(6) 2025. A DEVICE INVESTIGATION IS PENDING. A FOLLOW-UP REPORT WILL BE SUBMITTED ONCE EVALUATION IS COMPLETE. Manufacturer Narrative: NO PRODUCT WAS RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD (DHR) REVIEW WAS COMPLETED, AND THIS DEVICE PASSED ALL MANUFACTURING RELEASE CRITERIA FOR DISTRIBUTION. THERE WERE NO ISSUES IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. AT THE TIME OF THIS REPORT, THE ILET EVALUATION IS STILL IN PROGRESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22334236&pc=OAE | 2029046-2025-02087 | 2025-05-30 04:00 | Death | BIOSENSE WEBSTER INC | 2025-06-26 04:00 | OAE | QDOT MICRO | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P210027 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION ABLATION PROCEDURE WITH A QDOT MICRO AND THE PATIENT DIED. THE REPORTER WAS NOTIFIED BY STAFF MEMBERS IN THE LAB THAT THE PATIENT EXPIRED OVER THE WEEKEND. THE PATIENT HAD GROIN COMPLICATIONS WHICH OCCURRED ON THE DAY OF THE PROCEDURE. THE PATIENT COMPLAINED OF GROIN ISSUES TO THEIR PHYSICIAN WHILE IN POST-ANESTHESIA CARE UNIT (PACU). ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THAT THE DATE OF DEATH WAS (B)(6) 2025. THE CAUSE OF DEATH WAS REPORTED AS UNKNOWN. INTERVENTION INCLUDED VASOPRESSORS AND DIAGNOSTIC TESTING INCLUDING ULTRASOUND, ECHOCARDIOGRAM AND COMPUTED TOMOGRAPHY (CT). OTHER RELEVANT HISTORY / PRE-EXISTING CONDITION THAT THE PHYSICIAN STATED WAS THAT THE PATIENT WAS ON LOVENOX AND COUMADIN. Manufacturer Narrative: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22334884&pc=DYB | 9612164-2025-03161 | 2025-06-06 04:00 | Death | MEDTRONIC IRELAND | 2025-06-26 04:00 | DYB | MICRA | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion; Insufficient Information | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC PERFORATION, CARDIAC TAMPONADE, PERICARDIAL EFFUSION, CARDIAC ARREST AND BLOOD PRESSURE DROP DURING A PROCEDURE TO IMPLANT A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. WHEN THE DELIVERY SYSTEM WAS PRESSED, ANGIOGRAPHY SHOWED SLIGHT SLIPPAGE AND THE TIP HAD PERFORATED OUTSIDE THE RIGHT VENTRICLE, AND THE DELIVERY SYSTEM WAS REMOVED FROM THE BODY. PERFORATION WAS FOUND UPON FURTHER ANGIOGRAPHY AND PERICARDIAL EFFUSION AND CARDIAC TAMPONADE WERE CONFIRMED BY ECHOCARDIOGRAM. BLOOD PRESSURE DROP WAS NOTED. DRAINAGE MEASURES WERE TAKEN; BLOOD COULD NOT BE DRAWN VERY MUCH, SO THE PATIENT WAS TEMPORARILY STABILIZED, BUT A HEART MASSAGE WAS PERFORMED AFTER CARDIAC ARREST. MEDICATION WAS ADMINISTERED AND DRAINAGE WAS RESUMED. AT THIS POINT, RED BLOOD WAS CONFIRMED. THE BLEEDING DID NOT STOP, SO SURGERY WAS PERFORMED. THORACOTOMY WAS PERFORMED, AND A PERFORATION WAS OBSERVED IN THE LEFT VENTRICLE. THE MYOCARDIUM WAS FRAGILE, AND THE BLEEDING COULD NOT BE COMPLETELY STOPPED, SO THE OPERATION WAS COMPLETED. IT WAS REPORTED THAT THE PATIENT DIED. THE PHYSICIAN'S ASSESSED THAT AFTER THE DRAINAGE FOLLOWING THE RIGHT VENTRICULAR PERFORATION STABILIZATION, THE LEFT VENTRICLE PERFORATED FOR AN UNKNOWN REASON, WHICH LED TO INCREASED BLEEDING AND DEATH. THE CAUSE IS BELIEVED TO BE THE PATIENT'S MYOCARDIAL CONDITION. THE CARDIAC MUSCLE WAS FRAGILE. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE INTRODUCER WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. THE ANALYST NOTED THE INTRODUCER WAS RETURNED WITH THE DILATOR SEPARATED FROM THE SHEATH OF THE INTRODUCER. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22334910&pc=PNJ | 9612164-2025-03162 | 2025-06-06 04:00 | Death | MEDTRONIC IRELAND | 2025-06-26 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC PERFORATION, CARDIAC TAMPONADE, PERICARDIAL EFFUSION, CARDIAC ARREST AND BLOOD PRESSURE DROP DURING A PROCEDURE TO IMPLANT A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) SYSTEM. WHEN THE DELIVERY SYSTEM WAS PRESSED, ANGIOGRAPHY SHOWED SLIGHT SLIPPAGE AND THE TIP HAD PERFORATED OUTSIDE THE RIGHT VENTRICLE, AND THE DELIVERY SYSTEM WAS REMOVED FROM THE BODY. PERFORATION WAS FOUND UPON FURTHER ANGIOGRAPHY AND PERICARDIAL EFFUSION AND CARDIAC TAMPONADE WERE CONFIRMED BY ECHOCARDIOGRAM. BLOOD PRESSURE DROP WAS NOTED. DRAINAGE MEASURES WERE TAKEN; BLOOD COULD NOT BE DRAWN VERY MUCH, SO THE PATIENT WAS TEMPORARILY STABILIZED, BUT A HEART MASSAGE WAS PERFORMED AFTER CARDIAC ARREST. MEDICATION WAS ADMINISTERED AND DRAINAGE WAS RESUMED. AT THIS POINT, RED BLOOD WAS CONFIRMED. THE BLEEDING DID NOT STOP, SO SURGERY WAS PERFORMED. THORACOTOMY WAS PERFORMED, AND A PERFORATION WAS OBSERVED IN THE LEFT VENTRICLE. THE MYOCARDIUM WAS FRAGILE, AND THE BLEEDING COULD NOT BE COMPLETELY STOPPED, SO THE OPERATION WAS COMPLETED. IT WAS REPORTED THAT THE PATIENT DIED. THE PHYSICIAN'S ASSESSED THAT AFTER THE DRAINAGE FOLLOWING THE RIGHT VENTRICULAR PERFORATION STABILIZATION, THE LEFT VENTRICLE PERFORATED FOR AN UNKNOWN REASON, WHICH LED TO INCREASED BLEEDING AND DEATH. THE CAUSE IS BELIEVED TO BE THE PATIENT'S MYOCARDIAL CONDITION. THE CARDIAC MUSCLE WAS FRAGILE. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE FULL DELIVERY SYSTEM WAS RETURNED AND ANALYZED. THE LUMEN OF THE DELIVERY SYSTEM WAS TORN. THE ANALYST NOTED THE FULL DELIVERY SYSTEM WAS RETURNED WITH THE DEVICE INTACT IN THE DEVICE CUP. THERE WERE NO ANOMALIES FOUND WITH THE DISTAL EDGE OF THE DEVICE CUP. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [(B)(6)]). D9: YES, RETURN DATE: 12-JUN-2025 H3: YES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: PRODUCT EVENT SUMMARY: THE DEVICE WAS RETURNED AND ANALYZED. RETURNED PRODUCT ANALYSIS WAS PERFORMED AND NO ANOMALIES WERE FOUND. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID MI2355A NTRODUCER THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: (B)(6)), D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22334971&pc=NPT | 2015691-2025-05186 | 2025-03-27 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-26 04:00 | NPT | EDWARDS LIFESCIENCES SAPIEN 3 TRANSCATHETER HEART VALVE | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Insufficient Information | P140031 |  | 1 | Event Description: IT WAS REPORTED BY THE IMPLANT PATIENT REGISTRY THAT ON THE SAME DAY AS A TRANSFEMORAL TAVR WITH A 29MM SAPIEN 3 VALVE, THE VALVE WAS EXPLANTED. THERE IS NO ADDITIONAL INFORMATION.Event Description: IT WAS REPORTED BY THE IMPLANT PATIENT REGISTRY, THAT ON THE SAME DAY AS A TRANSFEMORAL TAVR WITH A 29 MM SAPIEN 3 VALVE, THE VALVE WAS EXPLANTED DUE TO UNKNOWN REASONS. PER MEDICAL RECORDS RECEIVED, THE PATIENT PRESENTED WITH SHORTNESS OF BREATH, MILD CHEST DISCOMFORT, INCREASING FATIGUE, AND DIZZINESS. WORKUP REVEALED SEVERE AORTIC STENOSIS, AORTIC VALVE AREA 0.5CM2 AND MEAN GRADIENT OF 68 MMHG. A CT PERFORMED AS PART OF THE TAVR PATIENT SCREENING PROTOCOL SHOWED HEAVILY CALCIFIED BICUSPID AORTIC VALVE. THE PATIENT UNDERWENT A TRANSFEMORAL TAVR PROCEDURE WITH A 29 MM SAPIEN 3 VALVE IN THE NATIVE AORTIC ANNULUS. A TYPE I AORTIC DISSECTION OCCURRED, AND THE PATIENT WAS TAKEN TO THE OR FOR A BENTALL SURGICAL PROCEDURE WITH 21 MM EDWARDS KONECT BIOPROSTHETIC AORTIC VALVE CONDUIT WITH REIMPLANTATION OF CORONARY ARTERIES AND REPLACEMENT OF THE ASCENDING AORTA AND AORTIC ROOT. POST OPERATIVE COURSE WAS COMPLICATED BY CARDIAC ARREST, RESPIRATORY FAILURE, ACUTE KIDNEY INJURY, ACUTE LIVER FAILURE, AND METABOLIC ACIDOSIS. THE PATIENT CONDITION CONTINUED TO DETERIORATE AND UPON DISCUSSION WITH THE FAMILY ON POST OPERATIVE DAY 2, IT WAS DECIDED TO WITHDRAW TO COMFORT CARE. THE PATIENT SHORTLY AFTER BEING EXTUBATED. THE PATHOLOGY REPORT NOTED THAT THE AORTIC VALVE LEAFLETS CONSIST OF 2 TAN IRREGULAR HEART VALVE FRAGMENTS, SHOWING EXTENSIVE CALCIFICATION. THERE WAS NO ALLEGATION OR INDICATION OF AN EDWARDS DEVICE DYSFUNCTION. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), NON-EMERGENT REOPERATION, EMERGENCY CARDIAC SURGERY, REOPERATION, DEVICE MIGRATION OR MALPOSITION REQUIRING INTERVENTION AND DEVICE EXPLANT ARE LISTED AS POTENTIAL RISKS ASSOCIATED WITH THE DEVICE AND TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. THE IFU CAUTIONS THAT LONG-TERM DURABILITY HAS NOT BEEN ESTABLISHED FOR THE VALVE. REGULAR MEDICAL FOLLOW-UP IS ADVISED TO EVALUATE VALVE PERFORMANCE. ACCELERATED DETERIORATION OF THE VALVE DUE TO CALCIFIC DEGENERATION MAY OCCUR IN CHILDREN, ADOLESCENTS, OR YOUNG ADULTS AND IN PATIENTS WITH AN ALTERED CALCIUM METABOLISM. VALVE RECIPIENTS SHOULD BE MAINTAINED ON ANTICOAGULANT/ANTIPLATELET THERAPY, EXCEPT WHEN CONTRAINDICATED, AS DETERMINED BY THEIR PHYSICIAN. THIS DEVICE HAS NOT BEEN TESTED FOR USE WITHOUT ANTICOAGULATION. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. VALVE REINTERVENTION MAY BE DUE TO VALVE DYSFUNCTION (E.G., SEVERE, OR CLINICALLY SIGNIFICANT PARAVALVULAR LEAK, CENTRAL LEAK, SIGNIFICANT MALPOSITION, EARLY/LATE ENDOCARDITIS, THROMBOSIS, STRUCTURAL VALVE DETERIORATION OR DEGENERATION) AND/OR OTHER REASONS UNRELATED TO THE EDWARDS DEVICES (E.G., TREATMENT OF OTHER PATIENT COMORBIDITIES). DURING THE MANUFACTURING PROCESS, ALL SAPIEN VALVES (ALL MODELS) ARE 100% VISUALLY INSPECTED FOR DEFECTS AND 100% TESTED UNDER PHYSIOLOGICAL BACKPRESSURE CONDITIONS PRIOR TO RELEASE FOR DISTRIBUTION. THEREFORE, IT IS HIGHLY UNLIKELY THAT A MANUFACTURING DEFECT OR DEVICE MALFUNCTION CONTRIBUTED TO THE EVENT. IN THIS CASE, DESPITE MULTIPLE INVESTIGATIONAL ATTEMPTS, IT WAS NOT POSSIBLE TO OBTAIN ADDITIONAL DETAILS REGARDING THIS EVENT. DUE TO LIMITED INFORMATION, A DEFINITIVE ROOT CAUSE WAS UNABLE TO BE DETERMINED AT THIS TIME. AT THE TIME OF THIS REPORT, THE INFORMATION AVAILABLE DOES NOT REASONABLY SUGGEST THERE WAS A MALFUNCTION OF THE EDWARDS DEVICE OR THAT THE USE OR MISS-USE OF THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT. THERE IS NO ALLEGATION OF DEFICIENCIES RELATED TO THE IDENTITY, QUALITY, DURABILITY, RELIABILITY, SAFETY, LABELING EFFECTIVENESS, OR PERFORMANCE OF THIS EDWARD DEVICE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, THE INFORMATION WILL BE REVIEWED, AND THE FILE UPDATED ACCORDINGLY. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME. Manufacturer Narrative: THIS IS BEING SUBMITTED TO CORRECT H6 HEALTH IMPACT CODE. Manufacturer Narrative: PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES ON CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST THAT PATIENT FACTORS (HEAVILY CALCIFIED BICUSPID AORTIC ANNULUS) MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22335144&pc=QIX | 2010441-2025-00001 | 2024-11-18 05:00 | Death | Unknown Manufacturer | 2025-06-26 04:00 | QIX | FENTANYL | Insufficient Information | Myocardial Infarction |  |  | 1 | Event Description: MFR REPORT #: US-ALVOGEN-(B)(4). #1: ACUTE MYOCARDIAL INFARCTION V28.0 (PATIENT DIED ON (B)(6) 2024, WEARING ONE OF THE FENTANYL PATCHES, PATIENT'S DEATH CERTIFICATE STATED ACUTE MYOCARDIAL INFARCTION): FROM TO - SERIOUS - FATAL. THIS SPONTANEOUS CASE REPORT WAS RECEIVED FROM A CONSUMER OR OTHER NON-HEALTH PROFESSIONAL (PATIENT'S WIFE) VIA MEDICAL INFORMATION CALL CENTER (MICC) (REFERENCE NUMBER: (B)(4)) FROM UNITED STATES ON 17-JUN-2025. THIS CASE CONCERNS A MALE PATIENT OF AN UNKNOWN AGE WHO WAS USING FENTANYL PATCH FOR AN UNKNOWN INDICATION AND DIED ON (B)(6) 2024, WEARING ONE OF THE FENTANYL PATCHES, AND DEATH CERTIFICATE STATED ACUTE MYOCARDIAL INFARCTION (PT: ACUTE MYOCARDIAL INFARCTION) (CLINICAL TERM: MYOCARDIAL INFARCTION; PROBLEM TERM: INSUFFICIENT DEVICE PROBLEM INFORMATION). PATIENT'S RELEVANT MEDICAL HISTORY, PAST DRUG HISTORY, AND FAMILY HISTORY WERE NOT REPORTED. CONCURRENT CONDITIONS, CONCOMITANT AND CO-SUSPECT MEDICATIONS WERE NOT REPORTED. ON AN UNKNOWN DATE, THE PATIENT STARTED TREATMENT WITH 25 MCG/HR FENTANYL (SYSTEM, TRANSDERMAL DELIVERY, CDER OR CBER LED; APPROPRIATE COMPONENT TERM/CODE NOT AVAILABLE) FOR AN UNKNOWN INDICATION (BATCH/LOT NUMBER: 108319, EXPIRATION DATE: APR-2027 AND DEVICE MANUFACTURER DATE: UNKNOWN). USAGE OF THE DEVICE WAS UNKNOWN, AND THE OPERATOR OF DEVICE WAS LAY USER/PATIENT. REPORTER STATED THAT HER HUSBAND DIED ON (B)(6) 2024 AND HE WAS WEARING ONE OF THESE PATCHES. SHE ALSO ASKED WHAT SHE COULD DO NOW AND REQUESTED FOR FURTHER ADVISE. ON 17-JUN-2025, REPORTER MENTIONED THAT SHE DID NOT KNOW IF THE PATCHES WERE STACKED. SHE MENTIONED THAT THE LAST BOX HE WAS USING WAS A RECALLED BATCH. SHE MENTIONED THAT DEATH CERTIFICATE MENTIONED ACUTE MYOCARDIAL INFARCTION AND WAS CREMATED PRIOR TO HEARING OF THE RECALL. SHE PROVIDED PICTURES, HOWEVER, DETAILS WERE NOT IDENTIFIABLE. REPORTED CAUSE OF DEATH WAS REPORTED AS ACUTE MYOCARDIAL INFARCTION. AUTOPSY WAS NOT PERFORMED. AS THE REPORTER DID NOT KNOW WHETHER THE PATCHES FROM THE RECALLED BATCH WERE STACKED, THERE WAS INSUFFICIENT INFORMATION. THEREFORE, THE TYPE OF INVESTIGATION, INCLUDING THE METHOD TERM, WAS CAPTURED AS 'INSUFFICIENT INFORMATION AVAILABLE,' THE RESULT TERM AS 'NO FINDINGS AVAILABLE,' AND THE CONCLUSION TERM AS 'APPROPRIATE INVESTIGATION CONCLUSION TERM/CODE NOT AVAILABLE.' ACCORDING TO THE REPORTER, THE PATCHES BELONGED TO THE AFFECTED LOT. BASED ON MEDICAL ASSESSMENT, THE RECALL PRESENTS A SIGNIFICANT SAFETY CONCERN, AS THE POTENTIAL RISK OF A PRODUCT CAN CAUSE SUBSTANTIAL AND UNREASONABLE HARM TO PUBLIC HEALTH. THE REPORTED EVENT ACUTE MYOCARDIAL INFARCTION, MEET THE FDA'S MDR REPORTABILITY CRITERIA AS DEFINED UNDER 21 CFR PART 803. THEREFORE, THIS QUALIFIES AS A 5-DAY REPORTABLE EVENT. THE REMEDIAL ACTION WAS INITIATED AS RECALL. NO LABORATORY TESTS AND PROCEDURES WERE REPORTED. THE CASE IS CONSIDERED AS SERIOUS (DEATH) (IMPACT TERM: DEATH, SERIOUS INJURY/ILLNESS/IMPAIRMENT). THE ACTION TAKEN WITH FENTANYL (SYSTEM, TRANSDERMAL DELIVERY, CDER OR CBER LED; APPROPRIATE COMPONENT TERM/CODE NOT AVAILABLE) WAS NOT APPLICABLE. THE OUTCOME OF THE EVENT WAS FATAL. THE REPORTER'S ASSESSMENT OF THE CAUSAL RELATIONSHIP OF THE EVENT WITH FENTANYL WAS PROVIDED AS POSSIBLE AT THE TIME OF THIS REPORT. NO FURTHER INFORMATION IS AVAILABLE. MDR COMMENT: - ACCORDING TO THE MEDICAL ASSESSMENT, A MALE PATIENT USING FENTANYL PATCH FOR AN UNKNOWN INDICATION DIED ON (B)(6) 2024, WEARING ONE OF THE FENTANYL PATCHES AND HIS DEATH CERTIFICATE STATED ACUTE MYOCARDIAL INFARCTION (ACUTE MYOCARDIAL INFARCTION). DEATH IS CONSIDERED POTENTIALLY DUE TO RECALL BATCH LOT 108319. HENCE, THERE IS A POTENTIAL SAFETY IMPACT WHICH CAUSED HARM TO PUBLIC HEALTH AND THE EVENT OF DEATH IS CONSIDERED RELATED TO THE DEVICE COMPONENT OF THE PRODUCT AND NECESSITATED REMEDIAL ACTION TO PREVENT AN UNREASONABLE RISK OF SUBSTANTIAL HARM TO THE PUBLIC HEALTH, THEREFORE, THE CASE QUALIFIES AS 5-DAY REPORTABLE, AND AN EXPEDITED REPORT WILL ALSO BE SUBMITTED TO FDA. SENDER COMMENT: THIS 5-DAY/15-DAY REPORTABLE CASE CONCERNS A MALE PATIENT OF AN UNKNOWN AGE WHO WAS USING FENTANYL PATCH FOR AN UNKNOWN INDICATION AND DIED ON (B)(6) 2024, WEARING ONE OF THE FENTANYL PATCHES AND PATIENT'S DEATH CERTIFICATE STATED ACUTE MYOCARDIAL INFARCTION (ACUTE MYOCARDIAL INFARCTION). THE EVENT OF ACUTE MYOCARDIAL INFARCTION WAS ASSESSED AS SERIOUS AND UNLISTED. THE CAUSAL ROLE OF FENTANYL FOR REPORTED EVENT OF ACUTE MYOCARDIAL INFARCTION IS ASSESSED AS POSSIBLY RELATED. HOWEVER, LACK OF INFORMATION REGARDING PATIENT¿S RELEVANT PAST MEDICAL HISTORY; PAST DRUG HISTORY; FAMILY HISTORY AND PRE-EXISTING RISK FACTORS PRECLUDES COMPREHENSIVE ASSESSMENT.Event Description: MFR REPORT#: (B)(4) / 2010441-2025-00001. #1: ACUTE MYOCARDIAL INFARCTION V28.0 (PATIENT DIED ON (B)(6) 2024, WEARING ONE OF THE FENTANYL PATCHES, PATIENT'S DEATH CERTIFICATE STATED ACUTE MYOCARDIAL INFARCTION): FROM TO - SERIOUS - FATAL. THIS SPONTANEOUS CASE REPORT WAS RECEIVED FROM A CONSUMER OR OTHER NON-HEALTH PROFESSIONAL (PATIENT'S WIFE) VIA MEDICAL INFORMATION CALL CENTER (MICC) (REFERENCE NUMBER: (B)(4) FROM UNITED STATES ON 17-JUN-2025. THIS CASE CONCERNS A MALE PATIENT OF AN UNKNOWN AGE WHO WAS USING FENTANYL PATCH FOR AN UNKNOWN INDICATION AND DIED ON (B)(6) 2024, WEARING ONE OF THE FENTANYL PATCHES, AND DEATH CERTIFICATE STATED ACUTE MYOCARDIAL INFARCTION (PT: ACUTE MYOCARDIAL INFARCTION) (CLINICAL TERM: MYOCARDIAL INFARCTION; PROBLEM TERM: INSUFFICIENT DEVICE PROBLEM INFORMATION). PATIENT'S RELEVANT MEDICAL HISTORY, PAST DRUG HISTORY, AND FAMILY HISTORY WERE NOT REPORTED. CONCURRENT CONDITIONS, CONCOMITANT AND CO-SUSPECT MEDICATIONS WERE NOT REPORTED. ON AN UNKNOWN DATE, THE PATIENT STARTED TREATMENT WITH 25 MCG/HR FENTANYL (SYSTEM, TRANSDERMAL DELIVERY, CDER OR CBER LED; APPROPRIATE COMPONENT TERM/CODE NOT AVAILABLE) FOR AN UNKNOWN INDICATION (BATCH/LOT NUMBER: 108319, EXPIRATION DATE: APR-2027 AND DEVICE MANUFACTURER DATE: UNKNOWN). USAGE OF THE DEVICE WAS UNKNOWN, AND THE OPERATOR OF DEVICE WAS LAY USER/PATIENT. REPORTER STATED THAT HER HUSBAND DIED ON (B)(6) 2024 AND HE WAS WEARING ONE OF THESE PATCHES. SHE ALSO ASKED WHAT SHE COULD DO NOW AND REQUESTED FOR FURTHER ADVISE. ON (B)(6) 2025, REPORTER MENTIONED THAT SHE DID NOT KNOW IF THE PATCHES WERE STACKED. SHE MENTIONED THAT THE LAST BOX HE WAS USING WAS A RECALLED BATCH. SHE MENTIONED THAT DEATH CERTIFICATE MENTIONED ACUTE MYOCARDIAL INFARCTION AND WAS CREMATED PRIOR TO HEARING OF THE RECALL. SHE PROVIDED PICTURES; HOWEVER, DETAILS WERE NOT IDENTIFIABLE. REPORTED CAUSE OF DEATH WAS REPORTED AS ACUTE MYOCARDIAL INFARCTION. AUTOPSY WAS NOT PERFORMED. AS THE REPORTER DID NOT KNOW WHETHER THE PATCHES FROM THE RECALLED BATCH WERE STACKED, THERE WAS INSUFFICIENT INFORMATION. THEREFORE, THE TYPE OF INVESTIGATION, INCLUDING THE METHOD TERM, WAS CAPTURED AS 'INSUFFICIENT INFORMATION AVAILABLE,' THE RESULT TERM AS 'NO FINDINGS AVAILABLE,' AND THE CONCLUSION TERM AS 'APPROPRIATE INVESTIGATION CONCLUSION TERM/CODE NOT AVAILABLE.' ACCORDING TO THE REPORTER, THE PATCHES BELONGED TO THE AFFECTED LOT. BASED ON MEDICAL ASSESSMENT, THE RECALL PRESENTS A SIGNIFICANT SAFETY CONCERN, AS THE POTENTIAL RISK OF A PRODUCT CAN CAUSE SUBSTANTIAL AND UNREASONABLE HARM TO PUBLIC HEALTH. THE REPORTED EVENT ACUTE MYOCARDIAL INFARCTION, MEET THE FDA'S MDR REPORTABILITY CRITERIA AS DEFINED UNDER 21 CFR PART 803. THEREFORE, THIS QUALIFIES AS A 5-DAY REPORTABLE EVENT. THE REMEDIAL ACTION WAS INITIATED AS RECALL. NO LABORATORY TESTS AND PROCEDURES WERE REPORTED. THE CASE IS CONSIDERED AS SERIOUS (DEATH) (IMPACT TERM: DEATH, SERIOUS INJURY/ILLNESS/IMPAIRMENT). THE ACTION TAKEN WITH FENTANYL (SYSTEM, TRANSDERMAL DELIVERY, CDER OR CBER LED; APPROPRIATE COMPONENT TERM/CODE NOT AVAILABLE) WAS NOT APPLICABLE. THE OUTCOME OF THE EVENT WAS FATAL. THE REPORTER'S ASSESSMENT OF THE CAUSAL RELATIONSHIP OF THE EVENT WITH FENTANYL WAS PROVIDED AS POSSIBLE AT THE TIME OF THIS REPORT. NO FURTHER INFORMATION IS AVAILABLE. AS PER AN EMAIL RECEIVED FROM FDA, A SUPPLEMENTAL REPORT NEEDS TO BE SUBMITTED AS INITIAL 5-DAY REPORT (CONSERVATIVELY PROCESSED AS 5-DAY) COULD NOT BE PROCESSED BECAUSE BOTH THE ¿5-DAY¿ AND ¿INITIAL¿ CHECKBOXES WERE SELECTED IN SECTION G6 (TYPE OF REPORT) FIELD OF MEDWATCH. INITIALLY, THIS REPORT WAS CONSERVATIVELY CLASSIFIED AND SUBMITTED AS A 5-DAY REPORT, SINCE THE CASE PERTAINED TO A PREVIOUSLY RECALLED LOT OF FENTANYL. HOWEVER, THE REMEDIAL ACTIONS FOR THAT LOT HAD ALREADY BEEN COMPLETED, AND THE RECALL WAS INITIATED BEFORE THE COMPANY BECAME AWARE OF THIS EVENT. AS NO ADDITIONAL CORRECTIVE ACTION WAS NECESSARY, THE EVENT HAS NOW BEEN RECLASSIFIED AS A 30-DAY REPORT. ACCORDINGLY, THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO INITIAL REPORT BY SELECTING THE ¿30-DAY" CHECKBOX IN SECTION G6 OF THE MEDWATCH FORM, IN ACCORDANCE WITH THE RECOMMENDATIONS PROVIDED BY THE FDA PROBLEM REPORT SPECIALIST. HENCE, AN AMENDMENT IS PERFORMED ON 09-JUL-2025 TO SUBMIT THE CASE AS A 30-REPORT. MDR COMMENT: ACCORDING TO THE MEDICAL ASSESSMENT, A MALE PATIENT USING FENTANYL PATCH FOR AN UNKNOWN INDICATION DIED ON (B)(6) 2024, WEARING ONE OF THE FENTANYL PATCHES (LOT NUMBER: 108319) AND HIS DEATH CERTIFICATE STATED ACUTE MYOCARDIAL INFARCTION (ACUTE MYOCARDIAL INFARCTION). THE AFFECTED LOT: (108319) WAS ALREADY RECALLED FROM THE MARKET AS THERE WAS A POTENTIAL THAT PATCHES COULD BE MULTI-STACKED, ADHERED ONE ON TOP OF THE OTHER, IN A SINGLE PRODUCT POUCH. ADDITIONALLY, AS THE COMPANY HAD ALREADY COMPLETED A REMEDIAL ACTION TO PREVENT AN UNREASONABLE RISK OF SUBSTANTIAL HARM TO THE PUBLIC HEALTH AND HAVE FILED 5-DAY REPORTS PREVIOUSLY FOR THE EVENTS THAT LED THE COMPANY TO RECOGNIZE THE NEED FOR THE REMEDIAL ACTION OF RECALL. THE RECALL WAS COMPLETED PRIOR TO THE COMPANY BECOMING AWARE OF THIS REPORT. THEREFORE, THE SUBSEQUENT ADDITIONAL REPORTABLE EVENT (INCLUDING THE CURRENT REPORT) ASSOCIATED WITH THE REMEDIAL ACTION (I.E., EVENTS THAT DO NOT NECESSITATE A NEW REMEDIAL ACTION) WOULD BE FILED AS 30-DAY REPORT. ALTHOUGH THERE IS INSUFFICIENT INFORMATION IF THE PATCHES USED BY THE PATIENT WERE STACKED OR NOT AS THE REPORTER WAS NOT SURE REGARDING SAME AND AS THE PATCH USED BY THE PATIENT BELONGS TO THE RECALLED LOT HENCE THE DEATH OF THE PATIENT IS CONSIDERED POTENTIALLY DUE TO THE OVER STACKING OF THE PATCHES REPORTED FOR THIS LOT, HENCE THIS CASE QUALIFIES TO BE A 30-DAY REPORTABLE CASE IN ACCORDANCE WITH REGULATORY GUIDELINES. SENDER COMMENT: THIS 15-DAY REPORTABLE CASE CONCERNS A MALE PATIENT OF AN UNKNOWN AGE WHO WAS USING FENTANYL PATCH FOR AN UNKNOWN INDICATION AND DIED ON (B)(6) 2024, WEARING ONE OF THE FENTANYL PATCHES AND PATIENT'S DEATH CERTIFICATES STATED ACUTE MYOCARDIAL INFARCTION (ACUTE MYOCARDIAL INFARCTION). THE EVENT OF ACUTE MYOCARDIAL INFARCTION WAS ASSESSED AS SERIOUS AND UNLISTED. THE CAUSAL ROLE OF FENTANYL FOR REPORTED EVENT OF ACUTE MYOCARDIAL INFARCTION IS ASSESSED AS POSSIBLY RELATED. HOWEVER, LACK OF INFORMATION REGARDING PATIENT¿S RELEVANT PAST MEDICAL HISTORY; PAST DRUG HISTORY; FAMILY HISTORY AND PRE-EXISTING RISK FACTORS PRECLUDES COMPREHENSIVE ASSESSMENT.Event Description: MFR REPORT #: (B)(4) / 2010441-2025-00001. #1: ACUTE MYOCARDIAL INFARCTION V28.0 (PATIENT DIED ON (B)(6) 2024, WEARING ONE OF THE FENTANYL PATCHES, PATIENT'S DEATH CERTIFICATE STATED ACUTE MYOCARDIAL INFARCTION): FROM (B)(6) 2024 TO - SERIOUS - FATAL. THIS SPONTANEOUS CASE REPORT WAS RECEIVED FROM A CONSUMER OR OTHER NON-HEALTH PROFESSIONAL (PATIENT'S WIFE) VIA MEDICAL INFORMATION CALL CENTER (MICC) (REFERENCE NUMBER: (B)(4)) FROM UNITED STATES ON 17-JUN-2025. THIS CASE CONCERNS A MALE PATIENT OF AN UNKNOWN AGE WHO WAS USING FENTANYL PATCH FOR CHRONIC PAIN AND DIED ON (B)(6) 2024, WEARING ONE OF THE FENTANYL PATCHES, AND DEATH CERTIFICATE STATED ACUTE MYOCARDIAL INFARCTION (PT: ACUTE MYOCARDIAL INFARCTION) (CLINICAL TERM: MYOCARDIAL INFARCTION; PROBLEM TERM: INSUFFICIENT DEVICE PROBLEM INFORMATION). PATIENT'S RELEVANT MEDICAL HISTORY AND FAMILY HISTORY WERE NOT REPORTED. CONCURRENT CONDITION INCLUDED CHRONIC PAIN. PAST DRUG HISTORY INCLUDED THE PATIENT WAS ON FENTANYL PATCHES FOR LIKE 20 YEARS APPROXIMATELY ON OR ABOUT 2000, FOR CHRONIC PAIN, WHEN ORAL MEDICATION (UNSPECIFIED) PROVED TO NOT BE SUFFICIENT FOR PAIN CONTROL OR RELIEF. CONCOMITANT AND CO-SUSPECT MEDICATIONS WERE NOT REPORTED. ON AN UNKNOWN DATE (SINCE YEARS), THE PATIENT WAS ON TREATMENT WITH FENTANYL TRANSDERMAL PATCHES (SYSTEM, TRANSDERMAL DELIVERY, CDER OR CBER LED; APPROPRIATE COMPONENT TERM/CODE NOT AVAILABLE) FOR CHRONIC PAIN (BATCH/LOT NUMBER: UNKNOWN, EXPIRATION DATE: UNKNOWN AND DEVICE MANUFACTURER DATE: UNKNOWN). HOWEVER, ON AN UNKNOWN DATE, THE PATIENT USED STARTED USING NEW FENTANYL TRANSDERMAL PATCH 25 MCG/HR (SYSTEM, TRANSDERMAL DELIVERY, CDER OR CBER LED; APPROPRIATE COMPONENT TERM/CODE NOT AVAILABLE) (BATCH/LOT NUMBER: 108319, EXPIRATION DATE: APR-2027 AND DEVICE MANUFACTURER DATE: UNKNOWN). USAGE OF THE DEVICE WAS UNKNOWN, AND THE OPERATOR OF DEVICE WAS LAY USER/PATIENT. REPORTER STATED THAT HER HUSBAND DIED ON (B)(6) 2024 AND HE WAS WEARING ONE OF THESE PATCHES. SHE ALSO ASKED WHAT SHE COULD DO NOW AND REQUESTED FOR FURTHER ADVISE. ON 17-JUN-2025, REPORTER MENTIONED THAT SHE DID NOT KNOW IF THE PATCHES WERE STACKED. SHE MENTIONED THAT THE LAST BOX HE WAS USING WAS A RECALLED BATCH. SHE MENTIONED THAT DEATH CERTIFICATE MENTIONED ACUTE MYOCARDIAL INFARCTION AND WAS CREMATED PRIOR TO HEARING OF THE RECALL. SHE PROVIDED PICTURES, HOWEVER, DETAILS WERE NOT IDENTIFIABLE. ON (B)(6) 2025, THE PATIENT'S WIFE SHARED THE PICTURE OF THE FENTANYL PATCHES AND ASKED IF COMPANY WANTED HER TO OPEN THE OTHER REMAINING RECALLED BOXES OF PATCHES. ON 30-JUN-2025, REPORTER WAS ADVISED NOT TO OPEN THE REMAINING PATCHES AND INFORMED FEW POINTS TO RETURN THE RECALLED LOT OF FENTANYL PATCHES. FURTHER, ON (B)(6) 2025, PATIENT'S WIFE MENTIONED THAT PATIENT SUFFERED ACUTE MYOCARDIAL INFARCTION ON (B)(6) 2024. ON (B)(6) 2025, PATIENT'S WIFE STATED THAT PATIENT HAD BEEN ON FENTANYL PATCHES SINCE YEARS AND MOST OF THE TIME THE USED COMPANY'S PATCHES. AS THE REPORTER DID NOT KNOW WHETHER THE PATCHES FROM THE RECALLED BATCH WERE STACKED, THERE WAS INSUFFICIENT INFORMATION. THEREFORE, THE TYPE OF INVESTIGATION, INCLUDING THE METHOD TERM, WAS CAPTURED AS 'INSUFFICIENT INFORMATION AVAILABLE,' THE RESULT TERM AS 'NO FINDINGS AVAILABLE,' AND THE CONCLUSION TERM AS 'APPROPRIATE INVESTIGATION CONCLUSION TERM/CODE NOT AVAILABLE.' ACCORDING TO THE REPORTER, THE PATCHES BELONGED TO THE AFFECTED LOT. BASED ON MEDICAL ASSESSMENT, THE RECALL PRESENTS A SIGNIFICANT SAFETY CONCERN, AS THE POTENTIAL RISK OF A PRODUCT CAN CAUSE SUBSTANTIAL AND UNREASONABLE HARM TO PUBLIC HEALTH. THE REPORTED EVENT ACUTE MYOCARDIAL INFARCTION, MEET THE FDA'S MDR REPORTABILITY CRITERIA AS DEFINED UNDER 21 CFR PART 803. THEREFORE, THIS QUALIFIES AS A 30-DAY REPORTABLE EVENT. THE REMEDIAL ACTION WAS INITIATED AS RECALL. REPORTED CAUSE OF DEATH WAS REPORTED AS ACUTE MYOCARDIAL INFARCTION. AUTOPSY WAS NOT PERFORMED. NO LABORATORY TESTS AND PROCEDURES WERE REPORTED. THE CASE IS CONSIDERED AS SERIOUS (DEATH) (IMPACT TERM: DEATH, SERIOUS INJURY/ILLNESS/IMPAIRMENT). THE ACTION TAKEN WITH FENTANYL (SYSTEM, TRANSDERMAL DELIVERY, CDER OR CBER LED; APPROPRIATE COMPONENT TERM/CODE NOT AVAILABLE) WAS NOT APPLICABLE. THE OUTCOME OF THE EVENT WAS FATAL. THE REPORTER'S ASSESSMENT OF THE CAUSAL RELATIONSHIP OF THE EVENT WITH FENTANYL WAS PROVIDED AS POSSIBLE AT THE TIME OF THIS REPORT. NO FURTHER INFORMATION IS AVAILABLE. AS PER AN EMAIL RECEIVED FROM FDA, A SUPPLEMENTAL REPORT NEEDS TO BE SUBMITTED AS INITIAL 5-DAY REPORT (CONSERVATIVELY PROCESSED AS 5-DAY) COULD NOT BE PROCESSED BECAUSE BOTH THE "5-DAY" AND "INITIAL" CHECKBOXES WERE SELECTED IN SECTION G6 (TYPE OF REPORT) FIELD OF MEDWATCH. INITIALLY, THIS REPORT WAS CONSERVATIVELY CLASSIFIED AND SUBMITTED AS A 5-DAY REPORT, SINCE THE CASE PERTAINED TO A PREVIOUSLY RECALLED LOT OF FENTANYL. HOWEVER, THE REMEDIAL ACTIONS FOR THAT LOT HAD ALREADY BEEN COMPLETED, AND THE RECALL WAS INITIATED BEFORE THE COMPANY BECAME AWARE OF THIS EVENT. AS NO ADDITIONAL CORRECTIVE ACTION WAS NECESSARY, THE EVENT HAS NOW BEEN RECLASSIFIED AS A 30-DAY REPORT. ACCORDINGLY, THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO INITIAL REPORT BY SELECTING THE "30-DAY" CHECKBOX IN SECTION G6 OF THE MEDWATCH FORM, IN ACCORDANCE WITH THE RECOMMENDATIONS PROVIDED BY THE FDA PROBLEM REPORT SPECIALIST. HENCE, AN AMENDMENT IS PERFORMED ON 09-JUL-2025 TO SUBMIT THE CASE AS A 30-REPORT. FOLLOW-UP INFORMATION OF THIS CASE WAS RECEIVED FROM A CONSUMER OR OTHER NON-HEALTH PROFESSIONAL (PATIENT'S WIFE) VIA MEDICAL INFORMATION CALL CENTER (MICC) (REFERENCE NUMBER: (B)(4)) FROM UNITED STATES ON (B)(6) 2025. CONCURRENT CONDITION (CHRONIC PAIN), PAST DRUG HISTORY, ONSET DATE OF EVENT, SUSPECT PRODUCT DETAIL (NDC NUMBER AND OTHER DETAILS), ADDITIONAL SUSPECT REGIMEN AND INDICATION OF SUSPECT PRODUCT WAS ADDED. DEVICE AVAILABLE FOR EVALUATION WAS UPDATED TO YES. NARRATIVE WAS UPDATED ACCORDINGLY. NO FURTHER INFORMATION IS AVAILABLE. MDR COMMENT: - ACCORDING TO THE MEDICAL ASSESSMENT, A MALE PATIENT USING FENTANYL PATCH FOR CHRONIC PAIN DIED ON (B)(6) 2024, WEARING ONE OF THE FENTANYL PATCHES (LOT NUMBER 108319) AND HIS DEATH CERTIFICATE STATED ACUTE MYOCARDIAL INFARCTION (ACUTE MYOCARDIAL INFARCTION). THE AFFECTED LOT (108319) WAS ALREADY RECALLED FROM THE MARKET AS THERE WAS A POTENTIAL THAT PATCHES COULD BE MULTI-STACKED, ADHERED ONE ON TOP OF THE OTHER, IN A SINGLE PRODUCT POUCH. ADDITIONALLY, AS THE COMPANY HAD ALREADY COMPLETED A REMEDIAL ACTION TO PREVENT AN UNREASONABLE RISK OF SUBSTANTIAL HARM TO THE PUBLIC HEALTH AND HAVE FILED 5-DAY REPORTS PREVIOUSLY FOR THE EVENTS THAT LED THE COMPANY TO RECOGNIZE THE NEED FOR THE REMEDIAL ACTION OF RECALL. THE RECALL WAS COMPLETED PRIOR TO THE COMPANY BECOMING AWARE OF THIS REPORT. THEREFORE, THE SUBSEQUENT ADDITIONAL REPORTABLE EVENT (INCLUDING THE CURRENT REPORT) ASSOCIATED WITH THE REMEDIAL ACTION (I.E., EVENTS THAT DO NOT NECESSITATE A NEW REMEDIAL ACTION) WOULD BE FILED AS 30-DAY REPORT. ALTHOUGH THERE IS INSUFFICIENT INFORMATION IF THE PATCHES USED BY THE PATIENT WERE STACKED OR NOT AS THE REPORTER WAS NOT SURE REGARDING SAME AND AS THE PATCH USED BY THE PATIENT BELONGS TO THE RECALLED LOT HENCE THE DEATH OF THE PATIENT IS CONSIDERED POTENTIALLY DUE TO THE OVER STACKING OF THE PATCHES REPORTED FOR THIS LOT; ADDITIONALLY FOLLOW-UP WAS RECEIVED; PROVIDING ADDITIONAL PATIENT DETAILS (CONCURRENT CONDITION (CHRONIC PAIN), PAST DRUG HISTORY, ONSET DATE OF EVENT, SUSPECT PRODUCT DETAIL, AND ADDITIONAL SUSPECT REGIMEN), HOWEVER, REPORTED INFORMATION DOES NOT ALTER THE INITIAL MEDICAL ASSESSMENT OF THE CASE. HENCE, THIS CASE QUALIFIES TO BE A 30-DAY REPORTABLE CASE IN ACCORDANCE WITH REGULATORY GUIDELINES. SENDER COMMENT: THIS 15-DAY REPORTABLE CASE CONCERNS A MALE PATIENT OF AN UNKNOWN AGE WHO WAS USING FENTANYL PATCH FOR PATCH FOR CHRONIC PAIN AND DIED ON (B)(6) 2024, WEARING ONE OF THE FENTANYL PATCHES AND PATIENT'S DEATH CERTIFICATE STATED ACUTE MYOCARDIAL INFARCTION (ACUTE MYOCARDIAL INFARCTION). THE EVENT OF ACUTE MYOCARDIAL INFARCTION WAS ASSESSED AS SERIOUS AND UNLISTED. THE CAUSAL ROLE OF FENTANYL FOR REPORTED EVENT OF ACUTE MYOCARDIAL INFARCTION IS ASSESSED AS POSSIBLY RELATED. HOWEVER, LACK OF INFORMATION REGARDING PATIENT¿S RELEVANT PAST MEDICAL HISTORY; PAST DRUG HISTORY; FAMILY HISTORY AND PRE-EXISTING RISK FACTORS PRECLUDES COMPREHENSIVE ASSESSMENT. Manufacturer Narrative: MANUFACTURER REPORT NUMBER: 2010441-2025-00001, WILL BE RE-SUBMITTED AS A 30-DAY REPORT VIA A SIGNIFICANT AMENDMENT. THE INITIAL 5-DAY REPORT COULD NOT BE PROCESSED BECAUSE BOTH THE "5-DAY" AND "INITIAL" CHECKBOXES WERE SELECTED IN SECTION G6 (TYPE OF REPORT) FIELD OF MEDWATCH. INITIALLY, THIS REPORT WAS CONSERVATIVELY CLASSIFIED AND SUBMITTED AS A 5-DAY REPORT, SINCE THE CASE PERTAINED TO A PREVIOUSLY RECALLED LOT OF FENTANYL. HOWEVER, THE REMEDIAL ACTIONS FOR THAT LOT HAD ALREADY BEEN COMPLETED, AND THE RECALL WAS INITIATED BEFORE THE COMPANY BECAME AWARE OF THIS EVENT. AS NO ADDITIONAL CORRECTIVE ACTION WAS NECESSARY, THE EVENT HAS NOW BEEN RECLASSIFIED AS A 30-DAY REPORT. ACCORDINGLY, THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO INITIAL REPORT BY SELECTING THE "30-DAY" CHECKBOX IN SECTION G6 OF THE MEDWATCH FORM, IN ACCORDANCE WITH THE RECOMMENDATIONS PROVIDED BY THE FDA PROBLEM REPORT SPECIALIST. Manufacturer Narrative: MANUFACTURER REPORT NUMBER: 2010441-2025-00001, WILL BE RE-SUBMITTED AS A 30-DAY REPORT VIA A SIGNIFICANT AMENDMENT. THE INITIAL 5-DAY REPORT COULD NOT BE PROCESSED BECAUSE BOTH THE ¿5-DAY¿ AND ¿INITIAL¿ CHECKBOXES WERE SELECTED IN SECTION G6 (TYPE OF REPORT) FIELD OF MEDWATCH. INITIALLY, THIS REPORT WAS CONSERVATIVELY CLASSIFIED AND SUBMITTED AS A 5-DAY REPORT, SINCE THE CASE PERTAINED TO A PREVIOUSLY RECALLED LOT OF FENTANYL. HOWEVER, THE REMEDIAL ACTIONS FOR THAT LOT HAD ALREADY BEEN COMPLETED, AND THE RECALL WAS INITIATED BEFORE THE COMPANY BECAME AWARE OF THIS EVENT. AS NO ADDITIONAL CORRECTIVE ACTION WAS NECESSARY, THE EVENT HAS NOW BEEN RECLASSIFIED AS A 30-DAY REPORT. ACCORDINGLY, THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO INITIAL REPORT BY SELECTING THE ¿30-DAY" CHECKBOX IN SECTION G6 OF THE MEDWATCH FORM, IN ACCORDANCE WITH THE RECOMMENDATIONS PROVIDED BY THE FDA PROBLEM REPORT SPECIALIST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22335375&pc=OZD | 1220648-2025-29801 | 2025-06-07 04:00 | Death | ABIOMED, INC. | 2025-06-26 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Atrial Fibrillation; Hemorrhagic Stroke | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THE PATIENT WENT INTO ATRIAL FIBRILLATION. THE PATIENT WAS UNABLE TO RECEIVE AMINO, SO SUPPORT CONTINUED, AND STAFF MONITORED. THE PATIENT ALSO HAD A HEMORRHAGIC STROKE WHICH WAS CONFIRMED BY COMPUTED TOMOGRAPHY. THE PATIENT REMAINED COMATOSE AND WAS MADE DO NOT RESUSCITATE. THE PATIENT EXPIRED. Manufacturer Narrative: THE INVESTIGATION OF VF/VT/AF/AT AND STROKE/CVA HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES). ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: WARNINGS, CAUTIONS & PRECAUTIONS: ¿TO REDUCE THE POSSIBILITY OF FIBERS BEING DRAWN INTO THE IMPELLA, CUSTOMERS SHOULD AVOID EXPOSING THE INLET AND CANNULA SECTION OF THE IMPELLA HEART PUMPS TO ANY SURFACES OR FLUID BATHS WHERE THE DEVICE CAN COME INTO CONTACT WITH LOOSE OR FLOATING FIBERS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22335492&pc=MAL | 2242352-2025-00737 | 2025-05-30 04:00 | Death | INTERVASCULAR SAS | 2025-06-26 04:00 | MAL | HEMASHIELD PLATINUM | Off-Label Use | Hemorrhage/Bleeding | K021213 |  | 1 | Event Description: A SURVEY WAS PERFORMED BY EXTRACORPOREAL LIFE SUPPORT ORGANIZATION (ELSO) REGARDING THE USE OF VASCULAR GRAFTS WITH GRAFT-FACILITATED ACCESS FOR VEINO ATERIAL-EXTRACOPOREAL MEMBRAN OXYGENATION (VA-ECMO) AND MECHANICAL CIRCULATORY SUPPORT (MCS). AS PART OF THESE TYPE OF PROCEDURES, ADVERSE EVENTS WERE REPORTED. THIS REPORT RELATES THE CASES OF DEATH MENTIONED IN THE SURVEY. BELOW IS A SUMMARY OF THE SURVEY. OBJECTIVE OF THE STUDY: THE FOCUS OF THIS SURVEY WAS ON THE USE OF SIDE GRAFTS ANASTOMOSED TO THE AXILLARY OR SUBCLAVIAN ARTERY AS PORT ACCESS FOR ARTERIAL PERFUSION IN V-A ECLS OR FOR IMPELLA OR FOR A COMBINATION OF SUCH DEVICES (REFERRED TO AS ECPELLA OR ECMELLA). EVENT DESCRIPTION: THE SURVEY DESCRIBED THAT MOST RESPONDENTS FOR BOTH GLOBALLY AND WITHIN THE US HAD EXPERIENCED AN ADVERSE EVENT AT THE GRAFT SITE. THIS REPRESENTED 53.2% GLOBALLY AND 64.3% IN THE US OF THE 62 PEOPLE WHO RESPOND TO THE QUESTIONS. GLOBALLY, 8.3% AND 15.8% IN US OF THE 36 RESPONDENTS, EXPERIENCED BLEEDING EVENTS WHICH LED TO THE DEATH OF THE PATIENT. GLOBALLY, IT REPRESENTS APPROXIMATELY 3 DEATHS. TO BE NOTED THAT DATA FROM THIS SURVEY ARE AGGREGATED DATA FROM A VOLUNTARY SURVEY, WITH NO PATIENT-INDIVIDUAL CLINICAL DATA, NO VERIFICATION OF FITTING COMPLIANCE, AND NO COMPARATIVE ANALYSIS BY PRODUCT TYPE. THERE IS NO CORRELATION BETWEEN A SPECIFIC GRAFT AND REPORTED ADVERSE EVENTS. COMPLAINT # (B)(4). THIS REPORT IS THE FIRST IN A SERIES OF 4. Manufacturer Narrative: (4111/3233) ATTEMPTS TO CONTACT THE SURVEY AUTHOR TO GET MORE INFORMATION OF THE ADVERSE EVENTS REPORTED AS PART OF THE SURVEY WILL BE PERFORMED. (11) THE INVESTIGATION IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE SENT UPON COMPLETION OF THE INVESTIGATION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22335646&pc=KRG | 2182208-2025-03093 | 2025-05-28 04:00 | Death | MEDTRONIC, INC. | 2025-06-26 04:00 | KRG | CARELINK PROGRAMMER | Unexpected Shutdown | Bradycardia; Cardiac Arrest; Asystole | P890003 |  | 1 | Event Description: IT WAS REPORTED THAT THE PROGRAMMER WAS TURNED OF SPONTANEOUSLY AND NEVER TURNED ON. THE PATIENT WAS DECEASED ON THIS EVENT BUT PROGRAMMER WAS NOT RELATED TO THE DECEASE OF THE PATIENT. TROUBLESHOOTING HAS BEEN DONE BY CHECKING THE POWER CORD WITH OTHER PROGRAMMER AND WORKED NORMALLY AND REPLACED THE PROGRAMMER IMMEDIATELY WITH THE MOBILE PROGRAMMER SO THAT THE IMPLANT PROCEDURE WAS NOT INTERRUPTED BY THE CHANGE OF PROGRAMMER. IT WAS NOTED THAT THE PATIENT EXPERIENCED ASYSTOLE, BRADYCARDIA AND PAUSES. IT WAS NOTED THAT THE HIGH THRESHOLDS AND LOSS OF CAPTURE WERE PATIENT CONDITION. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22335758&pc=PNJ | 9612164-2025-03174 | 2025-06-06 04:00 | Death | MEDTRONIC IRELAND | 2025-06-26 04:00 | PNJ | MICRA¿ AV2 | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion; Asystole | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY TEN TO FIFTEEN MINUTES POST IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE PATIENT EXPERIENCED ASYSTOLE, CARDIAC ARREST, CARDIAC TAMPONADE AND PERICARDIAL EFFUSION. CHEST COMPRESSIONS WERE ADMINISTERED AND, AFTER APPROXIMATELY TWENTY TO THIRTY MINUTES OF LIFE SAVING MEASURES, THE PATIENT WAS DECLARED DECEASED.Event Description: IT WAS FURTHER REPORTED THAT THE PATIENT EXPERIENCED SUSPECTED CARDIAC PERFORATION. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC2AVR1-DELSYS] (SERIAL: [(B)(6)]). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22335900&pc=NGV | 2124215-2025-42211 | 2025-05-15 04:00 | Death | BOSTON SCIENTIFIC CORPORATION | 2025-06-26 04:00 | NGV | WATCHMAN FLX? | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P130013 |  | 1 | Event Description: REPORTED VIA CLINICAL STUDY. IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED. THE PATIENT UNDERWENT SUCCESSFUL PLACEMENT OF 31MM WATCHMAN FLX DEVICE WITH DEPLOYED DEVICE DIAMETER OF 24.4 MM. THE PATIENT WAS DISCHARGED ON APIXABAN (10 MG/DAY), CLOPIDOGREL (75 MG/DAY) AND ASPIRIN (81 MG/DAY) ON (B)(6) 2023. ON (B)(6) 2025, 1220 DAYS POST INDEX PROCEDURE, THE PATIENT DIED, AND THE CAUSE OF DEATH WAS UNKNOWN. THE AUTOPSY PERFORMED IS UNKNOWN. AT THE TIME OF DEATH, THE PATIENT WAS ON ASPIRIN (81MG/DAY) AND RIVAROXABAN (20 MG/DAY). THERE WAS NO ACTION TAKEN AND DIAGNOSTICS PERFORMED IN RESPONSE TO THE EVENT.Event Description: REPORTED VIA CLINICAL STUDY IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS PERFORMED. THE PATIENT UNDERWENT SUCCESSFUL PLACEMENT OF 31MM WATCHMAN FLX DEVICE WITH DEPLOYED DEVICE DIAMETER OF 24.4 MM. THE PATIENT WAS DISCHARGED ON APIXABAN (10 MG/DAY), CLOPIDOGREL (75 MG/DAY) AND ASPIRIN (81 MG/DAY) ON (B)(6) 2023. ON (B)(6) 2025, 1220 DAYS POST INDEX PROCEDURE, THE PATIENT DIED, AND THE CAUSE OF DEATH WAS UNKNOWN. THE AUTOPSY PERFORMED IS UNKNOWN. AT THE TIME OF DEATH, THE PATIENT WAS ON ASPIRIN (81MG/DAY) AND RIVAROXABAN (20 MG/DAY). THERE WAS NO ACTION TAKEN AND DIAGNOSTICS PERFORMED IN RESPONSE TO THE EVENT. IT WAS FURTHER REPORTED THAT THE PATIENT DATE OF IMPLANT AND DATE OF DISCHARGED WAS ON (B)(6) 2022. Manufacturer Narrative: B5 DESCRIBE EVENT OR PROBLEM: UPDATED WITH ADDITIONAL INFORMATION RECEIVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22336263&pc=PNJ | 9612164-2025-03176 | 2025-06-05 04:00 | Death | MEDTRONIC IRELAND | 2025-06-26 04:00 | PNJ | MICRA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P150033 |  | 1 | Event Description: IT WAS REPORTED THAT POST IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT PASSED AWAY. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME [MICRA] PRODUCT ID [MC1VR01-DELSYS] (SERIAL: (B)(6). D9: NO. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22336266&pc=DYB | 9612164-2025-03177 | 2025-06-05 04:00 | Death | MEDTRONIC IRELAND | 2025-06-26 04:00 | DYB | MICRA | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K132030 |  | 1 | Event Description: IT WAS REPORTED THAT POST IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT PASSED AWAY. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22336338&pc=OZD | 1220648-2025-29805 | 2025-06-08 04:00 | Death | ABIOMED, INC. | 2025-06-26 04:00 | OZD | IMPELLA CP | Malposition of Device | Cardiac Arrest | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND SHOWED FLAT WAVEFORMS. THE PHYSICIAN REPOSITIONED THE IMPELLA CP BUT THE PATIENT CODED. DESPITE OBTAINING RETURN OF SPONTANEOUS CIRCULATION BY PROVIDING CARDIOPULMONARY RESUSCITATION, THE PATIENT LATER EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22336406&pc=MOF | 3012931345-2025-00107 | 2025-04-30 04:00 | Death | STRYKER NEUROVASCULAR-UTAH-SALT LAKE CITY | 2025-06-26 04:00 | MOF | SYNCHRO SELECT-14 SUPPORT ST 215CM | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding | K202522 |  | 1 | Event Description: IT WAS REPORTED THAT DURING ANTERIOR COMMUNICATING ARTERY (ACOM ARTERY) ANEURYSM CASE, PATIENT VESSEL WALL WAS PERFORATED BY SUBJECT GUIDEWIRE WHEN TRYING TO SECURE THE DISTAL PORTION OF THE VESSEL AT THE INTENDED PLACEMENT SITE FOR STENT PLACEMENT WHICH LED TO SUBARACHNOID HEMORRHAGE. ANOTHER BALLOON CATHETER WAS USED AS MEDICAL INTERVENTION TO CONTROL BLEEDING. THE PATIENT DID NOT RECOVER FROM THE DAMAGE CAUSED BY SUBARACHNOID HEMORRHAGE (SAH) AND DIED ON THE FOURTH POSTOPERATIVE DAY. NO FURTHER INFORMATION IS AVAILABLE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22336599&pc=QZI | 3002648230-2025-00573 | 2025-06-09 04:00 | Death | MEDTRONIC CRYOCATH LP | 2025-06-26 04:00 | QZI | PULSESELECT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P230017 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT'S DEATH WAS ASSOCIATED WITH THE USE OF THE PULSESELECT CATHETER ON THE CTI LINE. ALLEGATIONS INCLUDED A PRODUCT ISSUE AND OFF-LABEL USE OF PULSESELECT ON THE CTI LINE, WITH A REPORTED RISK OF DEATH WHEN USING THE DEVICE IN THIS MANNER. IT WAS STATED THAT CAUTION SHOULD BE EXERCISED IF CONSIDERING THIS USE. THE CASE OUTCOME IS UNKNOWN. THE PATIENT EXPIRED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT INCOMING INFORMATION INDICATES THAT THE MEDTRONIC PRODUCTS WERE NOT RELATED TO THE PATIENT'S DEATH. THE DEATH WAS RELATED TO NON-MEDTRONIC PRODUCTS. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22336630&pc=DSQ | 2916596-2025-03804 | 2024-03-01 05:00 | Death | THORATEC CORPORATION | 2025-06-26 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿INFLUENCE OF PREOPERATIVE ANATOMICAL CHARACTERISTICS ON THE DIRECTION OF INFLOW CANNULA IN HEARTMATE 3¿ THAT THE ANATOMICAL CHARACTERISTICS RELATED TO INFLOW CANNULA (IC) DIRECTION WAS INVESTIGATED THROUGH ANALYSIS OF COMPUTED TOPOGRAPHY (CT) IMAGES AND RELATION BETWEEN PROGNOSIS AND INFLOW CANNULA DIRECTION. 48 PATIENTS WERE REVIEWED WHO HAD UNDERWENT HM 3 IMPLANTATION, BETWEEN JUNE 2019 AND MARCH 2024, AND HAD CT IMAGES TAKE PRE- AND POST-IMPLANTATION. THE DIRECTION OF INFLOW CANNULA FOR GROUP A (N=18) WAS THE LEFT VENTRICULAR ANTERIOR/LATERAL WALL, GROUP B (N=21) WAS THE AORTIC/MITRAL VALVE, AND GROUP C (N=9) WAS LEFT VENTRICULAR POSTERIOR/INFERIOR WALL. THE MEAN AGE OF THE PATIENTS WAS 53 (44¿59) YEARS, AND 40 (83%) WERE MALE. ON PREOPERATIVE CT, GROUP A HAD A HIGHER MITRAL VALVE HEIGHT FROM THE LV APEX THAN GROUP B (68 ± 13 AND 52 ± 14 MM, RESPECTIVELY; P < 0.01). GROUP C HAD A LONGER DISTANCE BETWEEN THE LV APEX AND CHEST WALL THAN GROUP B (20 ± 9 AND 9 ± 6 MM, P < 0.01). GROUP C HAD A SHORTER THORACIC DEPTH FROM THE LV APEX THAN DID GROUP B (24 ± 9 AND 39 ± 11 MM, P < 0.01). THE IC DIRECTION IN HM3 WAS INFLUENCED BY THE PREOPERATIVE POSITION OF THE MITRAL VALVE, LV APEX, CHEST WALL, AND DIAPHRAGM. THE IC DIRECTION IN HM3 DID NOT SIGNIFICANTLY AFFECT SURVIVAL RATES. COMPLICATIONS AFTER HM3 IMPLANTATION INCLUDED CEREBRAL INFARCTION, CEREBRAL HEMORRHAGE, BLEEDING COMPLICATIONS INCLUDING GASTROINTESTINAL BLEEDING, HEART FAILURE, AND ARRHYTHMIA. CEREBROVASCULAR ACCIDENT (CVA) WAS EXPERIENCED BY 1 PATIENT IN GROUP A, 1 PATIENT IN GROUP B, AND 1 PATIENT IN GROUP C. HEART FAILURE WAS EXPERIENCED BY 2 PATIENTS IN GROUP A, 5 PATIENTS IN GROUP B, AND 1 PATIENT IN GROUP C. ARRYTHMIA WAS EXPERIENCED BY 2 PATIENTS IN GROUP A, 2 PATIENTS IN GROUP B, AND 3 PATIENTS IN GROUP C. BLEEDING COMPLICATIONS WERE EXPERIENCED BY 2 PATIENTS IN GROUP A, 3 PATIENTS IN GROUP B, AND 1 PATIENT IN GROUP C. LOW FLOW ALARMS WERE EXPERIENCED BY 1 PATIENT IN GROUP A, 2 PATIENTS IN GROUP B, AND 1 PATIENT IN GROUP C. THE IC DIRECTION WAS CONSIDERED TO HAVE LITTLE INFLUENCE ON LVAD DRIVELINE INFECTION. THUS, LVAD DRIVELINE INFECTION WAS EXCLUDED FROM COMPLICATIONS AFTER HM3 IMPLANTATION IN THIS STUDY. THE SURVIVAL RATES AFTER HM3 IMPLANTATION AT 3 YEARS IN GROUPS A, B, AND C WERE 88%, 90%, AND 100%, RESPECTIVELY (P = 0.67). THE RATES OF FREEDOM FROM COMPLICATIONS AFTER HM3 IMPLANTATION AT 3 YEARS IN GROUPS A, B, AND C WERE 50%, 43%, AND 20%, RESPECTIVELY (P = 0.70). Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED PATIENT OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RESEARCH ARTICLE TITLED, ¿INFLUENCE OF PREOPERATIVE ANATOMICAL CHARACTERISTICS ON THE DIRECTION OF INFLOW CANNULA IN HEARTMATE 3¿, WAS RECEIVED. THIS ARTICLE SOUGHT TO INVESTIGATE THE INFLUENCE OF ANATOMICAL CHARACTERISTICS ON INFLOW CANNULA (IC) DIRECTION IN PATIENTS WITH IMPLANTED WITH THE HEARTMATE 3 LVAS THROUGH ANALYSIS OF COMPUTED TOMOGRAPHY (CT) IMAGES AND ASSESSMENT OF IC DIRECTION AND PROGNOSIS. THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED AS THE HEARTMATE 3 DEVICE SERIAL NUMBER AS WELL AS OTHER SPECIFIC CASE/PATIENT INFORMATION, ARE NOT AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. A IS CURRENTLY AVAILABLE. THE IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT HAS BEEN ENTERED AS 01MAR2024, AS PATIENTS WERE IMPLANTED BETWEEN JUN2019 AND MAR2024. DEPARTMENT OF CARDIOVASCULAR SURGERY, OSAKA UNIVERSITY GRADUATE SCHOOL OF MEDICINE, OSAKA, JAPAN, DEPARTMENT OF THORACIC AND CARDIOVASCULAR, SURGERY, DOKKYO MEDICAL UNIVERSITY SAITAMA MEDICAL CENTER, SAITAMA, JAPAN IMAOKA, S., YOSHIOKA, D., SAITO, S., KAWAMURA, T., KAWAMURA, A., MATSUURA, R., MISUMI, Y., TODA, K., & MIYAGAWA, S. (2025). INFLUENCE OF PREOPERATIVE ANATOMICAL CHARACTERISTICS ON THE DIRECTION OF INFLOW CANNULA IN HEARTMATE 3. ARTIFICIAL ORGANS. HTTPS://DOI.ORG/10.1111/AOR.15018 NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22336954&pc=QZI | 3002648230-2025-00576 | 2025-06-09 04:00 | Death | MEDTRONIC CRYOCATH LP | 2025-06-26 04:00 | QZI | PULSESELECT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P230017 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT'S DEATH WAS ASSOCIATED WITH THE USE OF THE PULSESELECT CATHETER ON THE CTI LINE. ALLEGATIONS INCLUDED A PRODUCT ISSUE AND OFF-LABEL USE OF PULSESELECT ON THE CTI LINE, WITH A REPORTED RISK OF DEATH WHEN USING THE DEVICE IN THIS MANNER. IT WAS STATED THAT CAUTION SHOULD BE EXERCISED IF CONSIDERING THIS USE. THE CASE OUTCOME IS UNKNOWN. THE PATIENT EXPIRED. NO FURTHER PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.Event Description: IT WAS FURTHER REPORTED THAT INCOMING INFORMATION INDICATES THAT THE MEDTRONIC PRODUCTS WERE NOT RELATED TO THE PATIENT'S DEATH. THE DEATH WAS RELATED TO NON-MEDTRONIC PRODUCTS. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22337240&pc=OZD | 1220648-2025-29810 | 2025-06-09 04:00 | Death | ABIOMED, INC. | 2025-06-26 04:00 | OZD | IMPELLA CP | Restricted Flow rate; Deformation Due to Compressive Stress | Insufficient Information | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND EXPERIENCED LOW PUMP FLOW. THE IMPELLA CP POSITION WAS CONFIRMED TO BE CORRECT BUT A KINK WAS NOTED PROXIMAL TO THE OUTFLOW CAGE. THE PUMP WAS SUCCESSFULLY REPLACED WITH A NEW IMPELLA CP BUT THE PATIENT ULTIMATELY EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. Manufacturer Narrative: THE INVESTIGATION OF LOW PUMP FLOW AND PRODUCT DAMAGE (CANNULA KINK) HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. LOW FLOW - THE ROOT CAUSE OF THE LOW PUMP FLOW WAS MOST LIKELY A KINKED CANNULA SINCE A CANNULA KINK WAS OBSERVED UNDER X-RAY. CANNULA KINK - THE ROOT CAUSE OF THE PRODUCT DAMAGE (CANNULA KINK) WAS NOT DETERMINED SINCE PRODUCT WAS NOT RETURNED AND INSUFFICIENT CLINICAL DETAILS PROVIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22337800&pc=DSQ | 2916596-2025-03959 | 2022-04-07 04:00 | Death | THORATEC CORPORATION | 2025-06-26 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD A PSEUDOMONAS AERUGINOSA DRIVELINE EXIT SITE INFECTION THAT REQUIRED IV ANTIBIOTICS WITH AN ONSET OF 07APR2022. THE PATIENT WAS MOST RECENTLY ADMITTED ON (B)(6) 2025 FOR THE INFECTION. AFTER MULTIPLE ROUNDS OF TREATMENT WAS PERFORMED, THE PATIENT DECIDED TO END TREATMENT AND GO HOME ON HOSPICE. THE PATIENT WAS DISCHARGED ON (B)(6) 2025. THE PATIENT PASSED AWAY AT HOME WHILE ON HOSPICE CARE. THE PUMP WAS ON AT THE TIME OF DEATH AND WAS SAID TO HAVE BEEN WORKING AS INTENDED. Manufacturer Narrative: MANUFACTURER¿S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. INFECTION IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. INFECTION HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22337848&pc=NPT | 2015691-2025-05198 | 2025-06-05 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-26 04:00 | NPT | EDWARDS SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE | Material Split, Cut or Torn; Central Regurgitation |  | P140031 |  | 1 | Event Description: AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING A TRANSFEMORAL TAVR (TRANSCATHETER AORTIC VALVE REPLACEMENT) PROCEDURE WITH A 20MM SAPIEN 3 ULTRA (S3U) VALVE IN A PRE-EXISTING SURGICAL VALVE, THERE WAS "SIGNIFICANT WASTE" OF THE S3U VALVE. THE VALVE WAS POST DILATED WITH A 20MM TRU BALLOON, WHICH WAS INFLATED TO 14ATMS. WHILE THERE WAS SOME RELIEF OF THE VALVE WAIST, THE PATIENT'S BLOOD PRESSURE (BP) DID NOT RECOVER. ECHOCARDIOGRAM SHOWED WIDE OPEN, SEVERE AI (AORTIC REGURGITATION) UPON REMOVAL OF THE GUIDEWIRE. IT WAS SUSPECTED THAT THE S3U LEAFLETS TORE DURING BALLOONING. A SECOND S3U VALVE WAS IMPLANTED AND THE PATIENT BEGAN TO IMPROVE HEMODYNAMICALLY. APPROXIMATELY 3 MINUTES LATER, THE PATIENT DECOMPENSATED AND FULL RESUSCITATION WAS STARTED WITH CARDIOPULMONARY RESUSCITATION (CPR), INTUBATION AND MEDICATIONS, AFTER WHICH THE PATIENT'S BP STABILIZED. THE PATIENT WAS STILL DEPENDENT ON THE TEMPORARY PACEMAKER AND WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). THE PATIENT PASSED AWAY ON POSTOPERATIVE DAY 4. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS B4, G3, G6, H2, H6: COMPONENT CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. ADDITIONS TO SECTIONS H6: COMPONENT CODE AND TYPE OF INVESTIGATION. CORRECTIONS TO SECTIONS H6: INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS. THE EVENT(S) REPORTED IS/ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. IMAGING WAS PROVIDED AND REVEALED THE SURGICAL VALVE TRUE ID IS ONLY 17MM, WHICH LIMITS TRANSCATHETER HEART VALVE EXPANSION. IN THIS CASE, THE COMPLAINT OF LEAFLET TEAR WAS UNABLE TO BE CONFIRMED DUE TO UNAVAILABILITY OF RELEVANT IMAGERY AND/OR MEDICAL RECORDS. WHILE THE COMPLAINT OF CENTRAL REGURGITATION COULD ALSO NOT BE DEFINITIVELY CONFIRMED, THE INFORMATION REPORTED BY THE CLINICAL SPECIALIST SUPPORTS THE LIKELIHOOD OF SUCH AN EVENT. AS REPORTED, POST-DILATION WAS PERFORMED WITH A TRU BALLOON (NON-EDWARDS (EW) DEVICE) IN AN ATTEMPT TO RELIEVE THE WAIST OBSERVED ON THE THV, WHICH COULD HAVE OVER-EXPANDED AND/OR OVER-STRETCHED THE VALVE, RESULTING IN THE SUBSEQUENT SUSPECTED LEAFLET TEAR AND REPORTED CENTRAL REGURGITATION. IN ADDITION, IT WAS RECOMMENDED TO USE THE SAME DELIVERY SYSTEM TO PERFORM THE POST DILATION PER PROCEDURAL TRAINING MANUAL. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PROCEDURAL FACTORS (POST-DILATION WITH NON-EW BALLOON) MAY HAVE CONTRIBUTED TO THE COMPLAINT EVENT. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22337970&pc=OZD | 1220648-2025-29813 | 2025-06-09 04:00 | Death | ABIOMED, INC. | 2025-06-26 04:00 | OZD | IMPELLA 5.5 | Positioning Failure | Hemolysis; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED THAT A NSTEMI PATIENT WITH AN IMPELLA 5.5 WAS MADE. DURING SUPPORT, HEMOLYSIS WAS NOTED IN URINE. ECHO DONE AND PUMP WAS TURNED TOWARD MITRAL. PLAN TO REPOSITION AT BEDSIDE. ADDITIONALLY, LV THROMBUS 1CM X 1CM WAS NOTED ON TTE WITH DEFINITY. DECISION ON HOLDING OFF ON PUMP REPO AWAY FROM MV GIVEN THROMBUS CONCERN. ON ANOTHER, ECHO SHOWED PUMP TO BE SHALLOW AT 3.4 CM AND FACED TOWARDS THE MITRAL. LV THROMBUS NOTED. CONSULT SURGERY FOR POSSIBLE REPOSITIONING, URINE CONTINUED TO BE DARK. ADDITIONAL INFORMATION STATES THAT CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE IFU STATES: THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION, SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER, SECTION: DIRECT AORTIC INSERTION, SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES. "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH." . IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿ Manufacturer Narrative: THE INVESTIGATION OF POSITIONING ISSUES, THROMBUS, AND HEMOLYSIS HAS BEEN COMPLETED. THE IMPELLA DEVICE WAS NOT RETURNED FOR EVALUATION. WITHOUT ADDITIONAL CLINICAL DETAILS, THE ROOT CAUSE OF THE REPORTED ISSUES COULD NOT BE DETERMINED. B1 PRODUCT PROBLEM WAS INADVERTENTLY OMITTED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29813. D4 MODEL NUMBER WAS ERRONEOUSLY ENTERED ON THE INITIAL MANUFACTURER DEVICE REPORT NUMBER 1220648-2025-29813.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22337988&pc=PNJ | 2017865-2025-90375 | 2025-06-06 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-26 04:00 | PNJ | AVEIR¿ DELIVERY CATHETER | Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR AVEIR LEADLESS PACEMAKER (LP) IMPLANT PROCEDURE. DURING IMPLANT, THE RIGHT VENTRICULAR LP WAS NOTED TO INTERMITTENTLY CAPTURE. DUE TO SUDDEN MOVEMENT FROM THE PATIENT, THE LP WAS NOTED TO SHIFT ORIENTATION, UPON WHICH THE THRESHOLD INCREASED. FLUID FILLED THE PERICARDIAL SPACE FROM CARDIAC TAMPONADE AND THE PATIENT HAD GONE INTO CARDIAC ARREST. EMERGENCY PERICARDIOCENTESIS WAS PERFORMED. INTUBATION AND CARDIOPULMONARY RESUSCITATION WAS ATTEMPTED. THE PATIENT EXPIRED DURING PROCEDURE ON (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22338201&pc=PNJ | 2017865-2025-90374 | 2025-06-06 04:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-26 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Intermittent Capture | Arrhythmia; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150035 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PRESENTED FOR AVEIR LEADLESS PACEMAKER (LP) IMPLANT PROCEDURE. DURING IMPLANT, THE RIGHT VENTRICULAR LP WAS NOTED TO INTERMITTENTLY CAPTURE. DUE TO SUDDEN MOVEMENT FROM THE PATIENT, THE LP WAS NOTED TO SHIFT ORIENTATION, UPON WHICH THE THRESHOLD INCREASED. FLUID FILLED THE PERICARDIAL SPACE FROM CARDIAC TAMPONADE AND THE PATIENT HAD GONE INTO CARDIAC ARREST. EMERGENCY PERICARDIOCENTESIS WAS PERFORMED. INTUBATION AND CARDIOPULMONARY RESUSCITATION WAS ATTEMPTED. THE PATIENT EXPIRED DURING PROCEDURE ON (B)(6) 2025.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22338356&pc=MNS | 2518422-2025-047768 | 2021-03-26 04:00 | Death | RESPIRONICS, INC. | 2025-06-26 04:00 | MNS | DREAMSTATION ST30 | Degraded; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER HAS ALLEGED THAT THE PATIENT SADLY PASSED AWAY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN CONFIRMED NO FAULT FOUND AND PASSED ALL TESTS. Manufacturer Narrative: ON PREVIOUSLY SUBMITTED REPORT, SECTION "DESCRIBE EVENT OR PROBLEM" WAS INCORRECT. IT SHOULD BE - THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE DREAMSTATION ST30 DEVICE. THE REPORTER HAS ALLEGED THAT THE PATIENT SADLY PASSED AWAY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN CONFIRMED NO EVIDENCE OF FOAM PARTICLES IN THE AIR PATH DURING THE DEVICE EVALUATION. THE TECHNICIAN CONFIRMED NO FAULT FOUND AND PASSED ALL TESTS. THE DEVICE WAS SCRAPPED. SECTION "ADVERSE EVENT/PRODUCT PROBLEM" IN BOX B, " (DEVICE) PROBLEM CODE GRID (1)" IN BOX H WAS INCORRECT. IN THIS REPORT SECTION "ADVERSE EVENT/PRODUCT PROBLEM" IN BOX B, " (DEVICE) PROBLEM CODE GRID (1)" IN BOX H HAS BEEN UPDATED/CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22338536&pc=MNS | 2518422-2025-047770 | 2021-11-23 05:00 | Death | RESPIRONICS, INC. | 2025-06-26 04:00 | MNS | DREAMSTATION ST30 | Degraded; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | K102465 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER HAS ALLEGED THAT THE PATIENT HAS SADLY PASSED AWAY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: ON PREVIOUSLY SUBMITTED REPORT, SECTION "DESCRIBE EVENT OR PROBLEM" WAS INCORRECT. IT SHOULD BE - THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE DREAMSTATION ST30 DEVICE. THE REPORTER HAS ALLEGED THAT THE PATIENT HAS SADLY PASSED AWAY. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. SECTION "ADVERSE EVENT/PRODUCT PROBLEM" IN BOX B, " (DEVICE) PROBLEM CODE GRID (1)" IN BOX H WAS INCORRECT. IN THIS REPORT SECTION "ADVERSE EVENT/PRODUCT PROBLEM" IN BOX B, " (DEVICE) PROBLEM CODE GRID (1)" IN BOX H HAS BEEN UPDATED/CORRECTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22338537&pc=OUT | 2029214-2025-01468 | 2025-06-17 04:00 | Death | MICRO THERAPEUTICS, INC. DBA EV3 | 2025-06-26 04:00 | OUT | PIPELINE VANTAGE WITH SHIELD TECHNOLOGY | Adverse Event Without Identified Device or Use Problem | Ischemia; Muscle Weakness; Paresis; Dysphasia; Coma; Nervous System Injury; Speech Disorder; Thrombosis/Thrombus |  |  | 1 | Event Description: MEDTRONIC RECEIVED A REPORT THAT THE PATIENT WAS UNDERGOING OFF-LABEL EMERGENCY TREATMENT FOR GIANT FUSIFORM ANEURYSM OF THE PROXIMAL BASILAR TRUNK WITH BRAINSTEM COMPRESSION UNDER GENERAL ANESTHESIA WITH LEFT RADIAL ARTERY ACCESS INTO THE LEFT VERTEBRAL ARTERY. THE SURGEON AND PATIENT WERE BOTH AWARE THAT THE DEVICE USED WAS ASSOCIATED WITH A RECALL PRIOR TO THE PROCEDURE. ALTHOUGH GETTING ACCESS IN A VERY TORTUOUS LEFT VERTEBRAL ARTERY WAS DIFFICULT, THE SURGEON WAS ABLE TO EFFECTIVELY AND ACCURATELY DEPLOY THREE 6MM X 50MM VANTAGE 027 DEVICES AND ONE 6MM X 30MM VANTAGE 027 DEVICE ACROSS THE ANEURYSM IN THE VERTEBROBASILAR CIRCULATION. THE PATIENT CAME OUT FROM ANESTHESIA AND HAD A VERY MILD DRIFT IN THE RIGHT UPPER EXTREMITY AND A VERY MILD, TRANSIENTWEAKNESS IN THE RIGHT LOWER DIVISION OF THE FACE. THROUGH THE NIGHT, THE PATIENT DID VERY WELL AND WAS STABLE, DYSARTHRIA AND DYSPHAGIA. THE PATIENT WAS PLACED ON 325 MG OF ASPIRIN AND 75 MG OF PLAVIX FOR 7 DAYS PRIOR TO THE PROCEDURE. THIS WAS CONTINUED AFTER THE PROCEDURE AND WAS SWITCHED TO 81 MG OF ASPIRIN AND 90 MG OF BRILINTA TWICE A DAY. UNFORTUNATELY, ON POSTOPERATIVE DAY #1 AROUND 9 TO 9:30 AM, THE PATIENT DEVELOPED SIGNIFICANT LEFT-SIDED DENSE HEMIPARESIS, WHICH PROGRESSED TO SIGNIFICANT SWALLOWING PROBLEMS, ACCOMPANIED BY INCREASED TONE IN ALL FOUR EXTREMITIES. BY POSTOPERATIVE DAY #2, THIS HAD PROGRESSED TO FULL UNRESPONSIVENESS WITH INTUBATION, AND MAGNETIC RESONANCE IMAGING (MRI)) CONFIRMED THROMBOSIS OF THE ANEURYSM (AS ANTICIPATED) WITH SIGNIFICANT COMPRESSION OF THE PONS, BUT NOW WITH DIFFUSION-WEIGHTED ABNORMALITY IN THAT AREA INDICATIVE OF ACUTE INJURY. IN ACCORDANCE WITH THE PATIENT'S ADVANCED DIRECTIVE AND WITH THE AGREEMENT OF THE FAMILY, THE PATIENT UNDERWENT COMPASSIONATE EXTUBATION, CONTINUING COMFORT MEASURES ONLY FOR HIS CARE. THE PATIENT PASSED AWAY. IT WAS NOTED THAT NO DEVICE DEFICIENCIES OR PROBLEMS WITH DELIVERY OR COMPLICATIONS OCCURRED DURING THE PROCEDURE. THE REPORTED DEVICE AND ANY ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). AS A CONCLUSION THE COMPLICATIONS THAT OCCURRED WERE RELATED TO THROMBOSIS OF THE A NEURYSM, CAUSING FURTHER COMPRESSION TO THE BRAINSTEM AND THE PONTINE REGION, CAUSING FURTHER COMPRESSION TO THE BRAINSTEM AND THE PONTINE REGION, RESULTING IN INJURIES TO THAT AREA OF THE BRAIN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22338806&pc=OZD | 1220648-2025-29819 | 2024-08-23 04:00 | Death | ABIOMED, INC. | 2025-06-26 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Arrhythmia | P140003 |  | 1 | Event Description: THE COMPLAINTS TEAM RECEIVED A NOTIFICATION VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED PERICARDIAL EFFUSION WITH A "POSSIBLE" RELATIONSHIP TO THE IMPELLA CP DEVICE USED: ¿PERICARDIAL EFFUSION PRESENT ON ECHO DURING PERCUTANEOUS CORONARY INTERVENTION PROCEDURE AND IMPELLA CP IMPLANTATION. AFTER SECOND PEA ARREST, THE SUBCOSTAL AREA WAS RE-PREPPED AND PERICARDIOCENTESIS WAS PERFORMED WITH DRAINAGE OF AROUND 800 ML OF BLOOD.¿. THE PUMP WAS EXPLANTED AND LATER, THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22338821&pc=DSQ | 2916596-2025-03968 | 2025-06-10 04:00 | Death | THORATEC CORPORATION | 2025-06-26 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Adverse Event Without Identified Device or Use Problem | Hypoxia; Renal Failure; Respiratory Failure; Cardiovascular Insufficiency; Gastrointestinal Hemorrhage; Osteomyelitis | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO RESPIRATORY FAILURE. THE SITE WAS UNABLE TO FIND THE DEVICE TRACKING FORM. NO LOG FILES WERE SUBMITTED.Event Description: ADDITIONAL INFORMATION WAS PROVIDED THAT THE CAUSE OF THE RESPIRATORY FAILURE WAS UNKNOWN. THE PATIENT DEVELOPED HYPOXIA, FURTHER DECLINED AND TRANSITIONED TO INPATIENT HOSPICE, THEN PLACED ON COMFORT MEASURE AND STOPPED BREATHING. THE PATIENT HAD END-STAGE HEART FAILURE, WITH OSTEOMYELITIS, CHRONIC KIDNEY DISEASE, A HISTORY OF GASTROINTESTINAL BLEEDING (GIB), AND PERIPHERAL VASCULAR DISEASE (PVD). THE PATIENT HAD A LONG HISTORY OF MEDICATION AND DIETARY NONCOMPLIANCE, AS WELL AS PALLIATIVE CARE. THE DEVICE OPERATED AS EXPECTED AND WAS NOT CONSIDERED TO BE RELATED TO THE DEATH. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER REGARDING THE EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS, INCLUDING RESPIRATORY FAILURE, RENAL DYSFUNCTION, INFECTION (LOCAL, DRIVELINE, OR PUMP POCKET INFECTION), BLEEDING, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. SECTION 6 OF THE IFU, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES INFORMATION REGARDING THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) VALUES, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THERE IS A RISK OF BLEEDING. THIS SECTION ALSO LISTS INFECTION AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. CARE INSTRUCTIONS REGARDING INFECTION PREVENTION ARE PROVIDED IN VARIOUS SECTIONS OF THIS IFU AND PATIENT HANDBOOK. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A: PATIENT WEIGHT AND GENDER HAVE NOT YET BEEN PROVIDED BY THE SITE. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22338911&pc=MHX | 9710602-2025-02004 | 2025-04-13 04:00 | Death | GE MEDICAL SYSTEMS (CHINA) CO., LTD (WUXI) | 2025-06-26 04:00 | MHX | B125 V1.5 | Insufficient Information | Insufficient Information | K201941 |  | 1 | Event Description: IT WAS REPORTED THAT THE MONITOR STILL DISPLAYED A "NORMAL" ECG AND SPO2 WHEN THE PATIENT WAS IN CARDIAC ARREST. THE PATIENT WAS CODED BUT DID NOT SURVIVE. Manufacturer Narrative: LEGAL MANUFACTURER: HCS RESEARCH PARK - 9900 INNOVATION DRIVE USA WAUWATOSA, WI 53226. A1-6: NOT PROVIDED BY THE CUSTOMER. GE HEALTHCARE'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22338937&pc=DXJ | 3008729547-2025-00063 | 2025-03-31 04:00 | Death | CRITIKON DE MEXICO S. DE R.L. DE C.V. | 2025-06-26 04:00 | DXJ | CSCS V1 | Insufficient Information | Insufficient Information | K133882 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT DEATH OCCURRED, ALLEGEDLY WITHOUT APPROPRIATE ALARMS PROVIDED BY THE DEVICE. Manufacturer Narrative: LEGAL MANUFACTURER: GEHC TROUT - 3114 N GRANDVIEW BLVD. USA WAUKESHA, WI 53188. A1-6: NOT PROVIDED BY CUSTOMER. GE HEALTHCARE'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22339332&pc=QEW | 3005168196-2025-00219 | 2025-01-31 05:00 | Death | PENUMBRA, INC. | 2025-06-26 04:00 | QEW | INDIGO SYSTEM FLASH ASPIRATION CATHETER | Adverse Event Without Identified Device or Use Problem | Cardiac Arrest | K241399 |  | 1 | Event Description: THE PATIENT UNDERWENT A THROMBECTOMY PROCEDURE IN THE LEFT PULMONARY ARTERY USING AN INDIGO SYSTEM FLASH ASPIRATION CATHETER (FLASH CATHETER). IT WAS REPORTED THAT THERE WAS DIFFICULTY EXTRACTING CLOT IN THE LEFT PULMONARY ARTERY, WHICH WAS REMOVED VIA ASPIRATION. FOLLOWING THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO CARDIAC CARE. DURING THE RELEASE TO CARDIAC CARE, THE PATIENT EXPERIENCED A PERI-ARREST EVENT AND WAS NOTED TO BE COLD AND CLAMMY. THE PATIENT THEN SUFFERED A PULSELESS ELECTRICAL ACTIVITY (PEA) CARDIAC ARREST, PROMPTING THE EMERGENCY RESPONSE TEAM TO ADMINISTER ALTEPLASE AND ADRENALINE ALONG WITH CARDIOPULMONARY RESUSCITATION (CPR). DESPITE THE ATTEMPTS, THE PATIENT EXPIRED. THE INDEPENDENT MEDICAL REVIEWER (IMR) ADJUDICATED THE CARDIAC ARREST AS A SERIOUS ADVERSE EVENT WITH A POSSIBLE RELATIONSHIP TO THE INDIGO ASPIRATION SYSTEM, AND A POSSIBLE RELATIONSHIP TO THE INDEX PROCEDURE. Manufacturer Narrative: THE PRODUCT WAS NOT RETURNED FOR EVALUATION. THEREFORE, A PHYSICAL DEVICE INVESTIGATION WAS UNABLE TO BE PERFORMED. FROM THE INFORMATION PROVIDED, THERE IS NO INDICATION THAT THERE WAS ANY DEVICE MALFUNCTION, NONCONFORMANCE, OR MISUSE THAT CONTRIBUTED TO THE REPORTED EVENT. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22340158&pc=OZD | 1220648-2025-29824 | 2025-01-29 05:00 | Death | ABIOMED, INC. | 2025-06-26 04:00 | OZD | IMPELLA 5.5 | Malposition of Device | Hemorrhage/Bleeding | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT AN IMPELLA 5.5 WAS IMPLANTED INTO THE LEFT VENTRICLE AND TURNED TOWARD MITRAL. ATTEMPTS TO REPOSITION THE DEVICE WERE UNSUCCESSFUL. THE PUMP WAS REMOVED, REWIRED, AND SUCCESSFULLY REINSERTED. THE PATIENT EXHIBITED BLEEDING AT THE ANASTOMOSIS AND THE GRAFT SITE WAS SURGICALLY REPAIRED. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. THERE IS NOT ENOUGH EVIDENCE TO EXCLUDE IMPELLA AS AN ASSOCIATED FACTOR IN THE PATIENT'S OUTCOME. Manufacturer Narrative: A4 IS UNKNOWN. THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22340523&pc=OZD | 1220648-2025-29827 | 2025-06-05 04:00 | Death | ABIOMED, INC. | 2025-06-26 04:00 | OZD | IMPELLA CP | Adverse Event Without Identified Device or Use Problem | Tachycardia | P140003 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT WAS IMPLANTED WITH AN IMPELLA CP AND EXPERIENCED VENTRICULAR TACHYCARDIA THAT REQUIRED CARDIOVERSION AND PLACEMENT OF AN IMPLANTABLE CARDIOVERTER-DEFIBRILLATOR. AFTER FIVE HOURS, THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22340616&pc=DSQ | 2916596-2025-03997 | 2025-06-06 04:00 | Death | THORATEC CORPORATION | 2025-06-26 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS UNKNOWN AND THE LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS WORKING AS EXPECTED. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AN AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. DUE TO HOSPITAL PRIVACY POLICY, NO ADDITIONAL INFORMATION WOULD BE PROVIDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22340704&pc=NKE | 3004209178-2025-10983 | 2025-05-28 04:00 | Death | MEDTRONIC PUERTO RICO OPERATIONS CO. | 2025-06-26 04:00 | NKE | PERCEPTA QUAD CRT-P MRI SURESCAN | Appropriate Term/Code Not Available | Insufficient Information | P010015 |  | 1 | Event Description: IT WAS REPORTED BY A REPRESENTATIVE OF THE PATIENT THAT THE PATIENT PASSED AWAY BECAUSE OF THEIR CARDIAC RESYNCHRONIZATION THERAPY PACEMAKER (CRT-P). ADDITIONAL INFORMATION RELATED TO THE CAUSE OF DEATH WAS REQUESTED AND IS NOT AVAILABLE. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22340716&pc=DSQ | 2916596-2025-03903 | 2025-05-30 04:00 | Death | THORATEC CORPORATION | 2025-06-26 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Heart Failure/Congestive Heart Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY SECONDARY TO CONGESTIVE HEART FAILURE (CHF). NO ADDITIONAL INFORMATION WAS PROVIDED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6) AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION HAS BEEN PROVIDED BY THE ACCOUNT, AND THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING (DEATH), THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: SECTION A4: PATIENT WEIGHT WAS NOT PROVIDED. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22340845&pc=MIH | 2017233-2025-06358 | 2025-06-20 04:00 | Death | W. L. GORE & ASSOCIATES, INC. | 2025-06-26 04:00 | MIH | GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS | Adverse Event Without Identified Device or Use Problem | Ischemia; Embolism/Embolus | P210032 |  | 1 | Event Description: ON (B)(6) 2025, THE PATIENT UNDERWENT ENDOVASCULAR PROCEDURE TO TREAT A THORACIC ANEURYSM USING A GORE® TAG® THORACIC BRANCH ENDOPROSTHESIS DEVICE (TBE). IT WAS REPORTED THAT THE GORE TBE DEVICE WAS IMPLANTED SUCCESSFULLY. ON (B)(6) 2025, THE PATIENT EXPIRED. THE CAUSE OF DEATH WAS THROMBUS MIGRATION FROM THE DESCENDING AORTA INTO THE SUPERIOR MESENTERIC ARTERY CAUSING ISCHEMIA. IT WAS REPORTED TO GORE THAT THE PATIENT¿S DEATH WAS NOT DUE TO THE GORE DEVICE. REPORTEDLY, THIS INFORMATION WAS NOTIFIED BY ASSISTING SURGICAL CONSULTANT. Manufacturer Narrative: H6: CODE C19: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. THE DEVICE REMAINS IMPLANTED AND IS NOT AVAILABLE FOR ANALYSIS. CODE A24: WAS USED TO COVER THAT THE PATIENT¿S DEATH WAS NOT DUE TO THE GORE DEVICE. THE CAUSE OF DEATH WAS THROMBUS MIGRATION FROM THE DESCENDING AORTA INTO THE SUPERIOR MESENTERIC ARTERY CAUSING ISCHEMIA. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22340992&pc=DSQ | 2916596-2025-04127 | 2025-05-30 04:00 | Death | THORATEC CORPORATION | 2025-06-26 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2025 AT 18:15. THE PATIENT REMAINED UNCONSCIOUS POSTOPERATIVELY, AND A BRAIN COMPUTED TOPOGRAPHY (CT) SCAN CONFIRMED THE PRESENCE OF INTRACEREBRAL HEMORRHAGE (ICH). THE SITE NOTED THAT DUE TO THE PATIENT'S UNDERLYING COMORBIDITIES, THE SITE WAS UNABLE TO DETERMINE IF THE EVENT WAS ASSOCIATED WITH THE LEFT VENTRICULAR ASSIST DEVICE (LVAD). THE PATIENT'S FAMILY SIGNED A DO NOT RESUSCITATE ORDER AND DECIDED TO FORGO AGGRESSIVE TREATMENT. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22341211&pc=DSQ | 2916596-2025-04224 | 2025-05-17 04:00 | Death | THORATEC CORPORATION | 2025-06-26 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY AND THE CAUSE WAS UNKNOWN. DUE TO PATIENT PRIVACY LAWS THE MANAGING HOSPITAL WOULD NOT COMMUNICATE PATIENT OUTCOME INFORMATION. BASED ON A REVIEW OF THE AVAILABLE PATIENT¿S RECORD AND A REQUEST FOR PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. AUTOPSY WAS NOT PERFORMED, AND THE DEVICE WAS NOT EXPLANTED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22341221&pc=DSQ | 2916596-2025-04264 | 2025-02-27 05:00 | Death | THORATEC CORPORATION | 2025-06-26 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO MULTI-SYSTEM ORGAN FAILURE. THE PATIENT SUFFERED FROM A REFRACTORY MULTI ORGAN FAILURE WHICH WAS RELATED TO THE DEATH. NO DEVICE RELATED ISSUE WAS RELATED TO THE MULTI SYSTEM ORGAN FAILURE. THE DEVICE OPERATED AS EXPECTED. THE DEATH WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE PATIENT'S OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REVISION, REV. D IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22341541&pc=PNJ | 2017865-2025-90402 | 2025-06-12 04:00 | Death | ABBOTT | 2025-06-26 04:00 | PNJ | AVEIR¿ LEADLESS PACEMAKER, RIGHT VENTRICLE | Material Too Rigid or Stiff; Adverse Event Without Identified Device or Use Problem | Arrhythmia; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P150035 |  | 1 | Event Description: DURING AN IMPLANT PROCEDURE, CONTRAST WAS OBSERVED TO BE IN THE PERICARDIAL SPACE AND A CARDIAC PERFORATION OF THE ANTERIOR GROOVE RESULTING IN TAMPONADE, BILATERAL FIXATED MYDRIASIS, AND CARDIAC ARREST WERE OBSERVED DUE TO THE DELIVERY CATHETER WITH THE LEADLESS PACEMAKER (LP) INSIDE. CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED. FOLLOWING CPR, THE HEART WAS SUTURED AND THE CATHETER AND THE LP WERE REMOVED FROM THE PATIENT. THE EVENT OCCURRED PRIOR TO FIXATION AND THE CAUSE OF THE EVENT WAS ALLEGED TO BE DUE TO THE CATHETER BEING TOO STIFF. THE LEADLESS PACEMAKER WAS NOT IMPLANTED. THE PATIENT WAS HOSPITALIZED AND WILL CONTINUE TO BE MONITORED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT EXPERIENCED CEREBRAL HYPOXIA FOLLOWED BY PATHOLOGICAL AWAKENING UPON CESSATION OF SEDATION, CHARACTERIZED BY SEIZURES. PERSISTENT HYPERLACTATEMIA WAS OBSERVED DESPITE AMINERGIC SUPPORT. FURTHERMORE, SUPRATENTORIAL CEREBRAL EDEMA WAS OBSERVED ON A BRAIN SCAN. FOLLOWING THE PROCEDURE, TREATMENT WAS DISCONTINUED AND THE PATIENT PASSED AWAY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22341987&pc=NJK | 3007797756-2025-00026 | 2025-05-20 04:00 | Death | PULMONX CORPORATION | 2025-06-26 04:00 | NJK | ZEPHYR ENDOBRONCHIAL VALVE (EBV) | Adverse Event Without Identified Device or Use Problem | Pneumonia; Pneumothorax; Respiratory Failure | P180002 |  | 1 | Event Description: THE PATIENT RECEIVED 4 VALVES ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT WAS ADMITTED TO THE HOSPITAL WITH COPD EXACERBATIONS, PNEUMONIA AND RAPID DECLINE. THE PATIENT WAS RELEASED FROM THE HOSPITAL ON (B)(6) 2025. PATIENT EXPIRED (B)(6) 2025.Event Description: THREE VALVES WERE IMPLANTED IN THE RIGHT UPPER LOBE. PATIENT DID GET A PNEUMOTHORAX AFTER THE PROCEDURE. THE VALVES WERE ALL EVENTUALLY REMOVED AND THE PATIENT REFUSED INTUBATION. MD CLARIFIED THE VALVES THEMSELVES DID NOT CAUSE THE PATIENT DEATH. A NARRATIVE HAS BEEN REQUESTED FROM THE MD FOR ADDITIONAL INFORMATION. Manufacturer Narrative: CORRECTION OF REPORT AS THIS INITIALLY SUBMITTED REPORT WAS INCORRECT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22342535&pc=NIK | 2017865-2025-90416 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-26 04:00 | NIK | PROMOTE CRT-D | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22342536&pc=NVY | 2017865-2025-90419 |  | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-26 04:00 | NVY | RIATA ACTIVE FIXATION | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P950022 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT HAD DECEASED DUE TO AN UNKNOWN CAUSE. NO ADDITIONAL INFORMATION WAS REPORTED. Manufacturer Narrative: THE LEAD WAS RETURNED DUE TO PATIENT DEATH. AS RECEIVED, A PARTIAL LEAD (ONLY CONNECTOR PORTION) WAS RETURNED IN ONE PIECE. ADDITIONAL FINDINGS UNRELATED TO THE REPORTED EVENT. FOUND INSULATION DAMAGE NOTED AT 20.4 ¿ 20.6 CM FROM THE CONNECTOR PIN CONSISTENT WITH FRICTION TO DEVICE CAN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22342694&pc=NVZ | 2017865-2025-90421 | 2009-11-29 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-26 04:00 | NVZ | ACCENT DR | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P880086 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS CARDIAC INSUFFICIENCY.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22342696&pc=NVN | 2017865-2025-90422 | 2009-11-29 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-26 04:00 | NVN | TENDRIL ST | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS CARDIAC INSUFFICIENCY. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, A COMPLETE LEAD WAS RETURNED STRETCHED FOR ANALYSIS. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES ALTHOUGH INTERNAL SHORT WAS NOTED DUE TO PROCEDURAL DAMAGE. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE. Manufacturer Narrative: CORRECTION: H3: FIELD SHOULD BE MARKED "YES".
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22342701&pc=NVN | 2017865-2025-90424 | 2009-11-29 05:00 | Death | ST. JUDE MEDICAL, INC.(CRM-SYLMAR) | 2025-06-26 04:00 | NVN | TENDRIL ST | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P960013 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DECEASED. THERE IS NO ALLEGATION FROM A HEALTHCARE PROFESSIONAL THAT THE DEATH WAS DEVICE RELATED. THE CAUSE OF DEATH WAS CARDIAC INSUFFICIENCY. Manufacturer Narrative: THE REPORTED EVENT WAS PATIENT DEATH. AS RECEIVED, A COMPLETE LEAD WAS RETURNED FOR ANALYSIS. ELECTRICAL TESTING DID NOT FIND ANY INDICATION OF CONDUCTOR FRACTURES OR INTERNAL SHORTS. VISUAL INSPECTION OF THE LEAD DID NOT FIND ANY ANOMALIES WITH THE EXCEPTION OF PROCEDURAL DAMAGE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22342813&pc=QMG | 3015053858-2025-00080 | 2025-05-28 04:00 | Death | SHOCKWAVE MEDICAL, INC. | 2025-06-26 04:00 | QMG | SHOCKWAVE C2 CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | Material Split, Cut or Torn | Air Embolism; Cardiac Arrest | P200039 |  | 1 | Event Description: A SHOCKWAVE C2 CORONARY INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER WAS USED TO TREAT A HIGHLY ECCENTRIC AND CONCENTRIC CALCIFIED LESION IN THE CORONARY VASCULATURE. THE SHOCKWAVE CATHETER SUCCESSFULLY DELIVERED 9 CYCLES OF PULSES. THE PHYSICIAN THEN DECIDED TO PERFORM AN ANGIOPLASTY WITH A LUCERNE OPEN NON-COMPLIANT BALLOON WITHOUT REMOVING THE SHOCKWAVE CATHETER. FOLLOWING ANGIOPLASTY, THE PHYSICIAN ATTEMPTED TO DELIVER THE LAST CYCLE OF PULSES FROM THE SHOCKWAVE CATHETER. HOWEVER, THE BALLOON RUPTURED, RESULTING IN AN AIR EMBOLISM EXPERIENCED BY THE PATIENT AND SUBSEQUENT PULSELESS ELECTRICAL ACTIVITY (PEA) AND ASYSTOLE OF THE HEART REQUIRING IMMEDIATE MECHANICAL RESUSCITATION EFFORTS USING IMPELLA. THE PATIENT'S HEMODYNAMICS WERE SUCCESSFULLY RESTORED FOLLOWING ADVANCED LIFE SUPPORT, AND THE PHYSICIAN WAS ABLE TO PLACE A STENT ACHIEVING A GOOD OUTCOME IN THE VESSEL. HOWEVER, THE PATIENT SUFFERED THORACIC TRAUMA DUE TO THE RESUSCITATION, WHICH ULTIMATELY LED TO DEATH. THE PATIENT VERBALLY EXPRESSED A DESIRE TO DISCONTINUE LIFE-SUSTAINING EFFORTS AND STATED A PREFERENCE TO SUCCUMB TO FIGHTING THE ONGOING INFECTION; INFORMATION REGARDING THE INFECTION IS BEING REQUESTED. THE PROCEDURE WAS PERFORMED ON (B)(6) 2025 AND PATIENT DEATH OCCURRED ON (B)(6) 2025. Manufacturer Narrative: THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION CANNOT BE PERFORMED. MULTIPLE ATTEMPTS WERE MADE TO GATHER ADDITIONAL INFORMATION RELATED TO THE EVENT, IN RESPONSE, NO ADDITIONAL INFORMATION WILL BE MADE AVAILABLE. THE ROOT CAUSE OF THE BALLOON RUPTURE AND SUBSEQUENT DEATH COULD NOT BE DEFINITIVELY DETERMINED BASED ON THE INFORMATION AVAILABLE. BALLOON RUPTURES ARE A KNOWN FAILURE MODE WITH A LOW PROBABILITY OF OCCURRENCE. CONTRIBUTING FACTORS INCLUDE PATIENT CALCIUM, DAMAGE CAUSED WHEN THE BALLOON WALL COMES INTO CLOSE CONTACT WITH THE DEVICE EMITTERS, LIKELY CAUSES INCLUDE (A) AN IRREGULAR CALCIUM LESION CAPABLE OF COMPRESSING THE BALLOON WALL INTO CLOSE PROXIMITY WITH THE EMITTER(S) AND/OR (B) AN INCOMPLETELY INFLATED BALLOON. THE SHOCKWAVE INTRAVASCULAR LITHOTRIPSY (IVL) SYSTEM WITH THE SHOCKWAVE C2 CORONARY IVL CATHETER GENERATES ACOUSTIC PRESSURE PULSES WITHIN THE TARGET TREATMENT SITE, DISRUPTING CALCIUM WITHIN THE LESION ALLOWING SUBSEQUENT DILATATION OF A CORONARY ARTERY STENOSIS USING LOW BALLOON PRESSURE NOT TO EXCEED 4 ATM DURING IVL TREATMENT, AND 6 ATM FOR POST DILATATION. THE PRESSURE USED FOR THE IVL BALLOON CATHETER IS MUCH LOWER THAN OTHER BALLOON CATHETERS USED IN PERCUTANEOUS CORONARY INTERVENTION (PCI), THUS THE RISK ASSOCIATED WITH LOSS OF BALLOON PRESSURE DURING IVL CATHETER USE IS NEGLIGIBLE. THE ASSOCIATED PATIENT RISK FOR THIS FAILURE MODE IS LOW AND ADEQUATELY CAPTURED IN SWMI'S RISK DOCUMENTATION THE LOT NUMBER OF THE DEVICE WAS NOT PROVIDED. THEREFORE, REVIEW OF THE DEVICE MANUFACTURING AND TEST DOCUMENTATION COULD NOT BE COMPLETED. HOWEVER, SHOCKWAVE MEDICAL INC. HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO DISTRIBUTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22342909&pc=DSQ | 2916596-2025-04005 | 2025-06-09 04:00 | Death | THORATEC CORPORATION | 2025-06-26 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE ¿LOW DOSE APIXABAN IN HEARTMATE 3 LVAD PATIENTS, INTERIM ANALYSIS OF THE APIXIVAD TRIAL, A RANDOMIZED CONTROLLED STUDY¿ THAT HEARTMATE 3 (HM3) LEFT VENTRICULAR ASSIST DEVICE (LVAD) MAY BE ASSOCIATED WITH BLEEDING, THROMBOEMBOLIC EVENTS, ISCHEMIC AND HEMORRHAGIC STROKE, SEPTIC SHOCK, AND DEATH. THE APIXIVAD TRIAL WAS LAUNCHED TO TEST THE FEASIBILITY AND SAFETY OF LOW-DOSE APIXABAN IN STABLE HM3 PATIENTS, WITH A FOLLOW-UP OF UP TO 24 MONTHS OR UNTIL TRANSPLANTATION. THE APIXIVAD STUDY IS AN INVESTIGATOR-INITIATED, RANDOMIZED, CONTROLLED, OPEN-LABEL PILOT TRIAL CONDUCTED AT TWO INTERNATIONAL CENTERS. IT AIMS TO EVALUATE THE SAFETY OF LOW-DOSE APIXABAN (2.5 MG TWICE DAILY) COMPARED TO STANDARD VITAMIN K ANTAGONIST (VKA) THERAPY IN PATIENTS WITH HM3. A TOTAL OF 50 PATIENTS WERE TARGETED FOR ENROLLMENT OVER THREE YEARS. THE PRIMARY ENDPOINT WAS A COMPOSITE OF DEATH (1 PATIENT), PUMP THROMBOSIS/THROMBOEMBOLIC EVENTS, STROKES (1 PATIENT), AND MAJOR BLEEDING (3 PATIENTS), WITH TREATMENT CONTINUING UNTIL HEART TRANSPLANTATION OR A MAXIMUM OF 24 MONTHS. THE APIXIVAD INTERIM ANALYSIS EVALUATED THE SAFETY OF LOW-DOSE APIXABAN VERSUS STANDARD VITAMIN K ANTAGONIST (VKA) THERAPY IN PATIENTS WITH HM3. BETWEEN JANUARY 2022 AND MARCH 2025, 44 PATIENTS WERE ENROLLED¿21 RECEIVED APIXABAN AND 23 RECEIVED VKA. MOST PATIENTS (84%) WERE IMPLANTED AS A BRIDGE TO TRANSPLANT, AND THE MEDIAN TIME FROM LVAD IMPLANTATION TO RANDOMIZATION WAS SIX MONTHS. REGARDING HEMOCOMPATIBILITY-RELATED OUTCOMES, FIVE PATIENTS EXPERIENCED A PRIMARY EVENT: THREE HEMORRHAGIC EVENTS (TWO IN THE APIXABAN GROUP AND ONE IN THE VKA GROUP), ALL OF WHOM WERE ALSO ON ASPIRIN. ONE VKA PATIENT HAD A MINOR STROKE, AND ONE APIXABAN PATIENT DIED FROM PNEUMONIA AND SEPTIC SHOCK AFTER CHOOSING PALLIATIVE CARE. THE RATE OF HEMOCOMPATIBILITY-RELATED EVENTS WAS NEARLY IDENTICAL BETWEEN GROUPS¿16.9 PER 100 PATIENT-YEARS FOR APIXABAN AND 16.4 FOR VKA. THE ABSOLUTE DIFFERENCE IN EVENT RATES WAS 5.3%, WHICH WAS NOT STATISTICALLY SIGNIFICANT AND REMAINED BELOW THE PREDEFINED THRESHOLD FOR CONCERN. FOURTEEN PATIENTS REMAINED ON ACTIVE TREATMENT AT THE TIME OF THIS ANALYSIS. THE INTERIM RESULTS OF THE APIXIVAD STUDY SUGGEST THAT USING A REDUCED DOSE OF APIXABAN (2.5 MG TWICE DAILY) IN PATIENTS WITH HM3 IS PROMISING. THIS DOSAGE APPEARED TO PROVIDE ADEQUATE ANTICOAGULATION IN THE SHORT TO MEDIUM TERM AND DID NOT LEAD TO INCREASED BLEEDING DURING HEART TRANSPLANTATION WHEN MANAGED UNDER CURRENT CLINICAL PROTOCOLS. HOWEVER, DUE TO THE STUDY¿S SMALL SAMPLE SIZE, THESE FINDINGS ARE CONSIDERED PRELIMINARY AND SHOULD BE VIEWED AS HYPOTHESIS-GENERATING RATHER THAN DEFINITIVE. FURTHER RESEARCH WITH LARGER COHORTS IS NEEDED TO CONFIRM THESE OUTCOMES. Manufacturer Narrative: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DEVICE WAS IMPLANTED AT TIME OF EVENT. DATE OF EVENT IS APPROXIMATE AS THE DATA WERE COLLECTED BETWEEN JANUARY 2022 AND MARCH 2025. AUTHOR CITATION: SCHNEGG, B., DEVEZA, R., BARUA, S., SANJAY CHAVALI, LO, P., CAPEK, L., BRUNO, J., PAVLICEK-BAHLO, M., XOURGIA, E., NEAGOE, A., ZÜRCHER, K., KAVITHA MUTHIAH, LAM, K., REINEKE, D., MATTHIAS SIEPE, HUNZIKER, L., MARTINELLI, M., MACDONALD, P., & HAYWARD, C. (2025). LOW DOSE APIXABAN IN HEARTMATE 3 LVAD PATIENTS, INTERIM ANALYSIS OF THE APIXIVAD TRIAL, A RANDOMIZED CONTROLLED STUDY. THE JOURNAL OF HEART AND LUNG TRANSPLANTATION. HTTPS://DOI.ORG/10.1016/J.HEALUN.2025.04.012. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE DEVICE AND THE REPORTED OUTCOMES COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. C IS CURRENTLY AVAILABLE. SECTION 1, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING SEPSIS AND DEATH. ALTHOUGH ONLY SEPSIS WAS REPORTED BY THE ACCOUNT, THE IFU ALSO LISTS INFECTION AS AN ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, AS WELL AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. ADDITIONALLY, SEVERAL SECTIONS OF THE HM3 LVAS IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. THE SERIAL NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22349353&pc=QBJ | MW5172077 |  | Death | ABBOTT DIABETES CARE INC. | 2025-06-26 04:00 | QBJ | FREESTYLE LIBRE 2 SENSOR | Insufficient Information | Insufficient Information |  |  | 1 | Event Description: USER REACHED OUT ASKING FOR THE REQUEST TO BE DELETED BECAUSE THE PATIENT IS DECEASED HCP DID NOT CONSENT TO FOLLOW UP. (B)(4) NO FURTHER INFORMATION PROVIDED OR AVAILABLE REGARDING THIS REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22325460&pc=NAY | 2955842-2025-27068 | 2025-05-26 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-06-25 04:00 | NAY | DAVINCI X | Adverse Event Without Identified Device or Use Problem | Peritonitis; Bowel Perforation; Multiple Organ Failure | K171294 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED UMBILICAL HERNIA TAPP PROCEDURE, THE COLON WAS INADVERTENTLY INJURED, LEADING TO AN INFECTION AND A SUBSEQUENT PROCEDURE, AFTER WHICH THE PATIENT EXPIRED. THE INJURY OCCURRED DURING THE INITIAL PORT INSERTION USING THE BLUNT OBTURATOR AND CANNULA [TROCAR] ACCESSORY. THERE WAS NO BLEEDING, AND THE INJURY WAS IMMEDIATELY IDENTIFIED AND CLOSED WITH SUTURES. THE PROCEDURE WAS COMPLETED ROBOTICALLY WITHOUT FURTHER COMPLICATIONS AND THE SURGEON DOES NOT BELIEVE THAT A MALFUNCTION OF A DA VINCI SYSTEM, INSTRUMENT, OR ACCESSORY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE PATIENT WAS REPORTEDLY MORBIDLY OBESE AND THE HERNIA WAS LARGE AND OBSTRUCTED. SHORTLY AFTER THE PROCEDURE, THE PATIENT DEVELOPED ABDOMINAL PAIN, SWELLING, AND DIFFICULTY BREATHING, AND WAS ADMITTED TO THE ICU WITH SIGNS OF PERITONITIS. ON POSTOPERATIVE DAY THREE, AN EXPLORATORY LAPAROTOMY REVEALED AN INTESTINAL INFECTION CAUSED BY LEAKAGE AT THE SITE OF THE INITIAL INJURY. ACCORDING TO THE HOSPITAL¿S STATEMENT, A TEAM OF "SPECIALIST DOCTORS" PROVIDED ALL POSSIBLE MEDICAL CARE. HOWEVER, THE PATIENT¿S CONDITION DETERIORATED DUE TO INFECTION AND MULTI-ORGAN FAILURE, ULTIMATELY RESULTING IN DEATH. Manufacturer Narrative: A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, MEGA SUTURECUT NEEDLE DRIVER, FORCE BIPOLAR WITH DUAL GRIP. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT SUFFERED AN INJURY TO THE BOWEL UPON PLACING A BLUNT OBTURATOR AND PORT. HOW THE INJURY OCCURRED AND HOW THE SURGEON¿S TECHNIQUE MAY HAVE CONTRIBUTED TO THIS EVENT WERE NOT PROVIDED. THIS WAS RECOGNIZED AND REPAIRED AT THE TIME OF THE INITIAL INJURY. HOW THE REPAIR WAS DONE OR ANY SAFEGUARDS TAKEN TO CONFIRM THE REPAIR WAS ADEQUATE WERE NOT PROVIDED. THE PATIENT SUBSEQUENTLY DEVELOPED A LEAK AND REQUIRED RE-EXPLORATION. THEY ULTIMATELY DIED AFTER ADDITIONAL COMPLICATIONS. NO AUTOPSY WAS PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22325548&pc=NAY | 2955842-2025-26593 | 2025-05-21 04:00 | Death | INTUITIVE SURGICAL, INC | 2025-06-25 04:00 | NAY | DAVINCI XI | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K131861 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A DA VINCI-ASSISTED ROUX-EN-Y GASTRIC BYPASS PROCEDURE. THE PATIENT SUBSEQUENTLY EXPIRED ON POST-OPERATIVE DAY 6. NO INFORMATION REGARDING THE EVENTS LEADING UP TO THE DEATH OR THE CAUSE OF DEATH WERE PROVIDED. THERE WAS NO REPORT OF ANY INTRAOPERATIVE COMPLICATIONS. THE SITE VICE PRESIDENT OF BARIATRICS REPORTED THAT NO SPECIFIC INFORMATION REGARDING THE POST-OPERATIVE COURSE WOULD BE AVAILABLE; HOWEVER, THERE WAS A POTENTIAL THAT THE PATIENT'S CAUSE OF DEATH COULD BE ATTRIBUTED TO A POST-OPERATIVE OVERDOSE. NO ADDITIONAL INFORMATION WAS PROVIDED. THERE WAS NO ALLEGATION THAT A DA VINCI SYSTEM, INSTRUMENT, OR ACCESSORY MALFUNCTION OCCURRED DURING THE PROCEDURE. Manufacturer Narrative: THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, PERMANENT CAUTERY HOOK, TIP-UP FENESTRATED GRASPER, CADIERE FORCEPS, AND MEGA SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. DHR REVIEW FOR THE DEVICE(S) INVOLVED WITH THE REPORTED EVENT FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS COMPLAINT. REVIEW OF THE SUREFORM 60 STAPLER LOGS FOUND USED IN THIS EVENT FOUND THE INSTRUMENT WAS INSTALLED ON THE SYSTEM 8 TIMES AND FIRED 7 WHITE RELOADS. ON INSTALL 1, A WHITE RELOAD WAS LOADED, HOWEVER NO CLAMPING OR FIRING WAS ATTEMPTED. ON INSTALL 2, THE FIRST CLAMP ATTEMPT STOPPED AT APPROXIMATELY 94% COMPLETION. THE NEXT TWO CLAMPS WERE SUCCESSFUL, AND THE FIRING WAS COMPLETED WITH 2-3 PAUSES FOR COMPRESSION. ON INSTALLS 3-5, THE FIRST CLAMP WAS SUCCESSFUL, AND THE FIRINGS WERE COMPLETED WITH 1 PAUSE FOR COMPRESSION EACH. ON INSTALLS 6-8, THE FIRST CLAMP WAS SUCCESSFUL, AND THE FIRINGS WERE COMPLETED WITH NO PAUSES FOR COMPRESSION EACH. THERE WERE NO STAPLER RELATED ERRORS IN THE LOGS, AND THERE WAS NO REPORT OF ANY ISSUES FROM THE OR TEAM REGARDING THE STAPLER INSTRUMENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT PATIENT IN THIS REPORT DIED 6 DAYS FOLLOWING A REPORTEDLY UNEVENTFUL GASTRIC BYPASS PROCEDURE. HOW THEY DIED AND WHAT EVENTS LED TO THEIR DEATH ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22325626&pc=PPN | 3015053858-2025-00081 | 2025-05-19 04:00 | Death | SHOCKWAVE MEDICAL, INC. | 2025-06-25 04:00 | PPN | SHOCKWAVE M5+ PERIPHERAL INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K203365 |  | 1 | Event Description: A SHOCKWAVE M5+ PERIPHERAL INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER WAS USED WAS USED TO TREAT THE SUPERFICIAL FEMORAL ARTERY (SFA) AND POPLITEAL ARTERY. THE SHOCKWAVE CATHETER DELIVERED 300 PULSES AT A PULSE PRESSURE OF 4ATM SUCCESSFULLY. IT WAS REPORTED THAT THE PATIENT EXPERIENCED COMPLICATIONS FOLLOWING STENT DELIVERY. IT WAS REPORTED THAT THE SHOCKWAVE DEVICE HAD NO IMPACT ON THE PATIENT'S COMPLICATIONS. ACCORDING TO THE MANAGEMENT TEAM, THE PATIENT DEATH OCCURRED POST PROCEDURE, AFTER BEING TRANSFERRED TO THE HEART CENTER. THERE WAS NO ADDITIONAL INFORMATION KNOWN AND/OR PROVIDED REGARDING THE EVENT. BASED ON THE INFORMATION PROVIDED, THE PROCEDURE WAS PERFORMED ON (B)(6) 2025 AND PATIENT DEATH OCCURRED ON THE SAME DAY. Manufacturer Narrative: THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION CANNOT BE PERFORMED. BASED ON THE REPORTED INFORMATION, THERE WAS NO IVL MALFUNCTION AND NO INFORMATION SUGGESTING THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. MULTIPLE ATTEMPTS WERE MADE TO GATHER ADDITIONAL INFORMATION RELATED TO THE EVENT WITH LIMITED INFORMATION TO DATE. SHOCKWAVE MEDICAL HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO BEING DISTRIBUTED. A REVIEW OF THE MANUFACTURING AND TEST DOCUMENTATION FOR SUBJECT LOT DOES NOT REVEAL ANY ISSUES WITH THE MANUFACTURING OF THE DEVICE. THE DEVICE PASSED ALL OF SHOCKWAVE MEDICAL, INC. ACCEPTANCE CRITERIA PRIOR TO SHIPPING. COMPLAINT TRENDS WILL CONTINUE TO BE MONITORED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22326038&pc=BZD | 2518422-2025-047714 | 2024-07-16 04:00 | Death | RESPIRONICS, INC. | 2025-06-25 04:00 | BZD | DS2ADV AUTO CPAP | Degraded | No Clinical Signs, Symptoms or Conditions | K200480 |  | 1 | Event Description: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22326784&pc=OZD | 1220648-2025-29762 | 2025-06-03 04:00 | Death | ABIOMED, INC. | 2025-06-25 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Stroke/CVA; Hemolysis | P140003 |  | 1 | Event Description: THE USER FACILITY REPORTED A PATIENT WHO HAD ESCALATED CARE FROM AN INTRA-AORTIC BALLOON PUMP TO AN IMPELLA CP TO THIS 5.5 PUMP AND ECMO THE PUMP WAS SUPPORTING WHEN SIGNS OF HEMOLYSIS WERE PRESENT IN ELEVATED FREE HEMOGLOBIN AND THE PATIENT SUFFERED A BRAIN BLEED. FAMILY WITHDREW CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE HAS NOT BEEN RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿ SECTION: USE OF ECHOCARDIOGRAPHY FOR POSITIONING OF THE IMPELLA CATHETER ¿EVALUATE THE POSITION OF THE IMPELLA CATHETER IF THE AUTOMATED IMPELLA CONTROLLER DISPLAYS POSITION ALARMS OR IF YOU OBSERVE LOWER THAN EXPECTED FLOWS OR SIGNS OF HEMOLYSIS. IF THE CATHETER DOES NOT APPEAR TO BE CORRECTLY POSITIONED, INITIATE STEPS TO REPOSITION IT.¿ IMPELLA CATHETER TOO FAR INTO THE LEFT VENTRICLE: ¿OBSTRUCTION OF THE IMPELLA CATHETER INLET AREA CAN LEAD TO INCREASED MECHANICAL FORCES ON BLOOD CELL WALLS AND SUBSEQUENT HEMOLYSIS, WHICH OFTEN PRESENTS AS DARK OR BLOOD COLORED URINE.¿ SECTION: HEMOLYSIS: ¿HEMOLYSIS SHOULD BE MONITORED DURING SUPPORT. PATIENTS WHO DEVELOP HIGH LEVELS OF HEMOLYSIS MAY SHOW SIGNS OF DECREASED HEMOGLOBIN LEVELS, DARK OR BLOOD-COLORED URINE, AND IN SOME CASES, ACUTE RENAL FAILURE. PLASMA-FREE HEMOGLOBIN (PFHGB) IS THE BEST INDICATOR TO CONFIRM WHETHER A PATIENT IS EXPOSED TO AN UNACCEPTABLE LEVEL OF HEMOLYSIS¿ ¿PATIENT CONDITIONS¿ INCLUDING CATHETER POSITION, PRE-EXISTING MEDICAL CONDITIONS, AND SMALL LEFT VENTRICULAR VOLUMES¿ MAY ALSO PLAY A ROLE IN PATIENT SUSCEPTIBILITY TO HEMOLYSIS.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22327018&pc=DSQ | 2916596-2025-04000 | 2025-04-07 04:00 | Death | THORATEC CORPORATION | 2025-06-25 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO AN UNKNOWN ETIOLOGY. THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED. NO ADDITIONAL PATIENT OUTCOME INFORMATION WOULD BE PROVIDED. Manufacturer Narrative: SECTION E1: REPORTER PHONE NUMBER (B)(6). MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), REV C, IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22327040&pc=NPT | 2025587-2025-04490 | 2025-03-30 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-25 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Stroke/CVA | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT 5 DAYS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, A CEREBRAL INFARCTION OCCURRED. IT WAS FURTHER REPORTED THAT 9 DAYS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT DIED DUE TO A CEREBRAL INFARCTION LOCALIZED IN THE BRAIN STEM. Manufacturer Narrative: UPDATED DATA: H6. H11. THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: BRAND NAME: DELIV SYS D-EVOLUTFX-2329, PRODUCT ID: D-EVOLUTFX-2329, SERIAL/LOT: UNKNOWN, USE BY DATE: UNKNOWN, UDI: UNKNOWN. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: THIS IS A SYSTEM REPORT. THE SECTION D INFORMATION IS FOR THE PRIMARY DEVICE, WHICH WAS IN USE WITH THE FOLLOWING: PRODUCT ID: D-EVO LUTFX-2329, SERIAL/LOT: (B)(6), USE BY DATE: 2026-08-30, UDI: (B)(4). UPDATED H.6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22327211&pc=NPT | 2025587-2025-04493 | 2025-06-24 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-25 04:00 | NPT | EVOLUT FX PLUS VALVE | Device Dislodged or Dislocated; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Perforation of Vessels; Rupture | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF AN EVOLUT TRANSCATHETER AORTIC VALVE, AFTER THE VALVE WAS DEPLOYED, THE VALVE DISLODGED ABOVE THE ANNULAR PLANE. A SNARE WAS USED IN AN ATTEMPT TO WITHDRAW THE DISLODGED VALVE HIGHER TOWARDS THE ASCENDING AORTA. A SECOND VALVE WAS LOADED INTO A SECOND DELIVERY CATHETER SYSTEM AND INSERTED INTO THE PATIENT. ALL ATTEMPTS TO CROSS THE AORTIC ARCH AND THE FIRST DISLODGED VALVE FAILED, DESPITE USING A SECOND SNARE AND STIFFER GUIDEWIRES, AND THE DELIVERY OF THE SECOND VALVE WAS UNSUCCESSFUL. DURING THESE ATTEMPTS, THE PATIENT BECAME UNSTABLE AND EXPERIENCED A DROP IN BLOOD PRESSURE. CARDIOPULMONARY RESUSCITATION WAS PERFORMED, AND CARDIAC SURGEONS PERFORMED AN EMERGENCY SURGERY; HOWEVER, THE PATIENT DID NOT SURVIVE.Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF AN EVOLUT TRANSCATHETER AORTIC VALVE, AFTER THE VALVE WAS DEPLOYED, THE VALVE DISLODGED ABOVE THE ANNULAR PLANE. A SNARE WAS USED IN AN ATTEMPT TO WITHDRAW THE DISLODGED VALVE HIGHER TOWARDS THE ASCENDING AORTA. A SECOND VALVE WAS LOADED INTO A SECOND DELIVERY CATHETER SYSTEM (DCS) AND INSERTED INTO THE PATIENT. ALL ATTEMPTS TO CROSS THE AORTIC ARCH AND THE FIRST DISLODGED VALVE FAILED, DESPITE USING A SECOND SNARE AND STIFFER GUIDEWIRES, AND THE DELIVERY OF THE SECOND VALVE WAS UNSUCCESSFUL. DURING THESE ATTEMPTS, THE PATIENT BECAME UNSTABLE AND EXPERIENCED A DROP IN BLOOD PRESSURE. CARDIOPULMONARY RESUSCITATION WAS PERFORMED, AND CARDIAC SURGEONS PERFORMED AN EMERGENCY SURGERY; HOWEVER, THE PATIENT DID NOT SURVIVE. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT A PRE-IMPLANT DILATION WAS NOT PERFORMED AND NON-MEDTRONIC GUIDEWIRES WERE USED FOR THE CASE. THE STATING POINT OF DEPLOYMENT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER; THE VALVE WAS DEPLOYED AT AN IMPLANT DEPTH OF 5MM ON THE NON-CORONARY CUSP (NCC) AND 6MM ON THE LEFT CORONARY CUSP (LCC). FOLLOWING DISLODGEMENT, THE VALVE DEPTH WAS APPROXIMATELY -2 OR -3MM; HOWEVER, THE DEPTH WAS DIFFICULT TO ESTIMATE AS THE PHYSICIAN DID NOT INJECT CONTRAST AT THAT TIME. IT WAS NOTED THE PATIENT'S CALCIFIED AORTIC ANNULUS AND LEFT VENTRICULAR OUTFLOW TRACT WERE POSSIBLE CONTRIBUTING FACTORS TO THE DISLODGEMENT. THE CAUSE OF DEATH WAS RUPTURE IN THE AORTIC ARCH. PER THE PHYSICIAN; THE DEATH WAS PROCEDURE RELATED AND NOT RELATED TO THE DCS OR THE VALVE. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. H6. PATIENT CODES WERE ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4 H2 H3 H6 IMAGE REVIEW: INTRAPROCEDURAL FLUOROSCOPIC IMAGES WERE PROVIDED FOR REVIEW OF THE EVENT. PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. THE IMAGES PROVIDED DO NOT SHOW THE VALVE IMPLANTATION NOR THE DISLODGEMENT EVENT. THE IMAGES SHOWED THE VALVE ALREADY DISLODGED AND SNARED TO THE ASCENDING AORTA. SINCE THE DISLODGEMENT EVENT WAS NOT CAPTURED IT IS NOT POSSIBLE TO VALIDATE CAUSE OF THE DISLODGEMENT. IT WAS REPORTED THAT SEVERAL ATTEMPTS TO ADVANCE A SECOND DELIVERY SYSTEM FAILED, AND CARDIOPULMONARY RESUSCITATION EFFORTS WERE INITIATED. FINAL ANGIOGRAM REVEALED EVIDENCE OF A POSSIBLE DISSECTION IN THE ASCENDING AORTA. IT WAS REPORTED THAT EMERGENCY SURGERY WAS PERFORMED BUT THE PATIENT SUBSEQUENTLY DIED. AS STATED IN THE INSTRUCTIONS FOR USE (IFU), POTENTIAL RISKS ASSOCIATED WITH THE IMPLANTATION OF THE EVOLUT FX+ BIOPROSTHESIS MAY INCLUDE, BUT ARE NOT LIMITED TO, THE FOLLOWING: PROSTHETIC VALVE MIGRATION/EMBOLIZATION; EMERGENT SURGICAL OR TRANSCATHETER INTERVENTION (FOR EXAMPLE, CORONARY ARTERY BYPASS, HEART VALVE REPLACEMENT, VALVE EXPLANT, PERCUTANEOUS CORONARY INTERVENTION [PCI], BALLOON VALVULOPLASTY); DEATH; CARDIAC TAMPONADE, CARDIOVASCULAR INJURY (INCLUDING RUPTURE, PERFORATION, TISSUE EROSION, OR DISSECTION OF VESSELS, ASCENDING AORTA TRAUMA, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012291300); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID EVFXPLUS-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID D-EVOLUTFX-2329 (0012294205); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22327223&pc=NPT | 2025587-2025-04495 | 2025-06-24 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-25 04:00 | NPT | EVOLUT FX DCS | Failure to Advance; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; Perforation of Vessels; Rupture | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF AN EVOLUT TRANSCATHETER AORTIC VALVE, AFTER THE VALVE WAS DEPLOYED, THE VALVE DISLODGED ABOVE THE ANNULAR PLANE. A SNARE WAS USED IN AN ATTEMPT TO WITHDRAW THE DISLODGED VALVE HIGHER TOWARDS THE ASCENDING AORTA. A SECOND VALVE WAS LOADED INTO A SECOND DELIVERY CATHETER SYSTEM AND INSERTED INTO THE PATIENT. ALL ATTEMPTS TO CROSS THE AORTIC ARCH AND THE FIRST DISLODGED VALVE FAILED, DESPITE USING A SECOND SNARE AND STIFFER GUIDEWIRES, AND THE DELIVERY OF THE SECOND VALVE WAS UNSUCCESSFUL. DURING THESE ATTEMPTS, THE PATIENT BECAME UNSTABLE AND EXPERIENCED A DROP IN BLOOD PRESSURE. CARDIOPULMONARY RESUSCITATION WAS PERFORMED, AND CARDIAC SURGEONS PERFORMED AN EMERGENCY SURGERY; HOWEVER, THE PATIENT DID NOT SURVIVE.Event Description: IT WAS REPORTED THAT DURING THE PROCEDURE INVOLVING THE IMPLANT OF AN EVOLUT TRANSCATHETER AORTIC VALVE, AFTER THE VALVE WAS DEPLOYED, THE VALVE DISLODGED ABOVE THE ANNULAR PLANE. A SNARE WAS USED IN AN ATTEMPT TO WITHDRAW THE DISLODGED VALVE HIGHER TOWARDS THE ASCENDING AORTA. A SECOND VALVE WAS LOADED INTO A SECOND DELIVERY CATHETER SYSTEM (DCS) AND INSERTED INTO THE PATIENT. ALL ATTEMPTS TO CROSS THE AORTIC ARCH AND THE FIRST DISLODGED VALVE FAILED, DESPITE USING A SECOND SNARE AND STIFFER GUIDEWIRES, AND THE DELIVERY OF THE SECOND VALVE WAS UNSUCCESSFUL. DURING THESE ATTEMPTS, THE PATIENT BECAME UNSTABLE AND EXPERIENCED A DROP IN BLOOD PRESSURE. CARDIOPULMONARY RESUSCITATION WAS PERFORMED, AND CARDIAC SURGEONS PERFORMED AN EMERGENCY SURGERY; HOWEVER, THE PATIENT DID NOT SURVIVE. ADDITIONAL INFORMATION WAS RECEIVED TO REPORT A PRE-IMPLANT DILATION WAS NOT PERFORMED AND NON-MEDTRONIC GUIDEWIRES WERE USED FOR THE CASE. THE STATING POINT OF DEPLOYMENT WAS AT THE BOTTOM OF THE PIGTAIL CATHETER, THE VALVE WAS DEPLOYED AT AN IMPLANT DEPTH OF 5MM ON THE NON-CORONARY CUSP (NCC) AND 6MM ON THE LEFT CORONARY CUSP (LCC). FOLLOWING DISLODGEMENT, THE VALVE DEPTH WAS APPROXIMATELY -2 OR -3MM; HOWEVER, THE DEPTH WAS DIFFICULT TO ESTIMATE AS THE PHYSICIAN DID NOT INJECT CONTRAST AT THAT TIME. IT WAS NOTED THE PATIENT'S CALCIFIED AORTIC ANNULUS AND LEFT VENTRICULAR OUTFLOW TRACT WERE POSSIBLE CONTRIBUTING FACTORS TO THE DISLODGEMENT. THE CAUSE OF DEATH WAS RUPTURE IN THE AORTIC ARCH. PER THE PHYSICIAN, THE DEATH WAS PROCEDURE RELATED AND NOT RELATED TO THE DCS OR THE VALVE. Manufacturer Narrative: IMAGE REVIEW: INTRAPROCEDURAL FLUOROSCOPIC IMAGES WERE PROVIDED FOR REVIEW OF THE EVENT. PATIENT¿S EXECUTIVE SUMMARY WAS NOT PROVIDED FOR ANATOMICAL REVIEW. THE IMAGES PROVIDED DO NOT SHOW THE VALVE IMPLANTATION NOR THE DISLODGEMENT EVENT. THE IMAGES SHOWED THE VALVE ALREADY DISLODGED AND SNARED TO THE ASCENDING AORTA. SINCE THE DISLODGEMENT EVENT WAS NOT CAPTURED IT IS NOT POSSIBLE TO VALIDATE CAUSE OF THE DISLODGEMENT. IT WAS REPORTED THAT SEVERAL ATTEMPTS TO ADVANCE A SECOND DELIVERY SYSTEM FAILED, AND CARDIOPULMONARY RESUSCITATION EFFORTS WERE INITIATED. FINAL ANGIOGRAM REVEALED EVIDENCE OF A POSSIBLE DISSECTION IN THE ASCENDING AORTA. IT WAS REPORTED THAT EMERGENCY SURGERY WAS PERFORMED BUT THE PATIENT SUBSEQUENTLY DIED. AS STATED IN THE INSTRUCTIONS FOR USE (IFU), POTENTIAL RISKS ASSOCIATED WITH THE IMPLANTATION OF THE EVOLUT FX+ BIOPROSTHESIS MAY INCLUDE, BUT ARE NOT LIMITED TO, THE FOLLOWING: PROSTHETIC VALVE MIGRATION/EMBOLIZATION; EMERGENT SURGICAL OR TRANSCATHETER INTERVENTION (FOR EXAMPLE, CORONARY ARTERY BYPASS, HEART VALVE REPLACEMENT, VALVE EXPLANT, PERCUTANEOUS CORONARY INTERVENTION [PCI], BALLOON VALVULOPLASTY); DEATH; CARDIAC TAMPONADE, CARDIOVASCULAR INJURY (INCLUDING RUPTURE, PERFORATION, TISSUE EROSION, OR DISSECTION OF VESSELS, ASCENDING AORTA TRAUMA, VENTRICLE, MYOCARDIUM, OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED H.6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B5. A SECOND PARAGRAPH WAS ADDED. H6. PATIENT CODES WERE ADDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012291300); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: NON-IMPLANTABLE DEVICE PRODUCT ID EVFXPLUS-26 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVFXPLUS-26 (J388195); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22327344&pc=NPT | 2025587-2025-04498 | 2024-10-01 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-25 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Unspecified Infection; Sepsis | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 6 MONTHS FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT WAS HOSPITALIZED FOR THE INITIATION OF DIALYSIS. IT WAS FURTHER REPORTED THAT APPROXIMATELY 2 MONTHS AFTER THE HOSPITALIZATION, THE PATIENT DIED FROM SEPSIS DUE TO A CATHETER INFECTION. Manufacturer Narrative: SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4. H4. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22327766&pc=OCE | 3002836339-2025-00046 | 2025-06-03 04:00 | Death | VERICEL | 2025-06-25 04:00 | OCE | EPICEL | Leak/Splash | Cardiac Arrest; Septic Shock; No Clinical Signs, Symptoms or Conditions | 990200 |  | 1 | Event Description: MEDIA NOTED IN CELLOPHANE SELF-SEALING POUCH IN BOX #2 [PRODUCT LEAKAGE].Event Description: CARDIAC ARREST [CARDIAC ARREST]. PT CONDITION DETERIORATED [CONDITION AGGRAVATED]. MEDIA NOTED IN CELLOPHANE SELF SEALING POUCH IN BOX #2 [PRODUCT LEAKAGE].Event Description: CARDIAC ARREST [CARDIAC ARREST]. SEPTIC SHOCK [SEPTIC SHOCK]. PT CONDITION DETERIORATED [CONDITION AGGRAVATED]. MEDIA NOTED IN CELLOPHANE SELF SEALING POUCH IN BOX #2 [PRODUCT LEAKAGE]. Manufacturer Narrative: CASE REFERENCE NUMBER (B)(4) IS A SPONTANEOUS REPORT INITIALLY RECEIVED FROM A COMPANY EMPLOYEE ON 03-JUN-2025, CONCERNING AN 84-YEAR-OLD FEMALE PATIENT WHO RECEIVED EPICEL GRAFTS. HOWEVER, EPICEL GRAFTS WERE NOTED WITH LIQUID MEDIA IN CELLOPHANE SELF-SEALING POUCH IN SECOND LAYER (PT: PRODUCT LEAKAGE). THE PATIENT'S MEDICAL HISTORY AND CONCOMITANT MEDICATION DETAILS WERE NOT REPORTED. ON 03-JUN-2025, THE PATIENT RECEIVED EPICEL GRAFTS (CULTURED EPIDERMAL AUTOGRAFTS) (LOT NUMBER: EE03423-31, EXPIRATION DATE: 03-JUN-2025) FOR A MORE THAN 30% TOTAL BODY SURFACE AREA (TBSA). ON THE SAME DATE, A LIQUID MEDIA WAS NOTED IN CELLOPHANE SELF SEALING POUCH IN BOX#2 IN SECOND LAYER (FULL PACK OF 24 GRAFT DISHES) (PT: PRODUCT LEAKAGE). IT WAS REPORTED THAT THERE WERE NO OBVIOUS OPENED OR LEAKING CASSETTES. THE PRODUCT WAS NOT AFFECTED AS THERE WAS ONLY SOME SPILLAGE IN THE POUCH IN WHICH THE CASSETTES WERE PLACED. AT THE TIME OF THE REPORT, THE OUTCOME OF THE REPORTED EVENT WAS RESOLVED. NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED ON 04-JUN-2025 AND PROCESSED TOGETHER WITH THE INITIAL. NO FURTHER INFORMATION WAS PROVIDED. COMPANY COMMENT: VERICEL ASSESSED THE EVENT ON PRODUCT LEAKAGE AS NON-SERIOUS, POSSIBLY RELATED AND EXPECTED. THE CASE QUALIFIES AS A 30-DAY MDR REPORT. Manufacturer Narrative: CASE REFERENCE NUMBER (B)(4) (FOLLOW-UP) IS A SPONTANEOUS REPORT *INITIALLY* RECEIVED FROM A COMPANY EMPLOYEE ON 03-JUN-2025, CONCERNING AN 85-YEAR-OLD FEMALE PATIENT WHO RECEIVED EPICEL GRAFTS. HOWEVER, EPICEL GRAFTS WERE NOTED WITH LIQUID MEDIA IN CELLOPHANE SELF-SEALING POUCH IN SECOND LAYER (PT: PRODUCT LEAKAGE). ADDITIONALLY, THE PATIENT'S CONDITION DETERIORATED (PT: CONDITION AGGRAVATED), THE PATIENT EXPERIENCED *SEPTIC SHOCK (PT: SEPTIC SHOCK)*, CARDIAC ARREST (PT: CARDIAC ARREST) AND EXPIRED. ON 02-JUN-2025, QC LOT RELEASE ASSAYS (GRAFT INSPECTION QC2-027, DUAL STAIN QC2-094, ENDOTOXIN QC2-010, STERILITY PRE-RELEASE QC2-005, STERILITY FINAL PRODUCT QC2-012) AND ENVIRONMENTAL RESULTS (MANUFACTURING: GRADE A (FINGERTIPS - LEFT AND RIGHT, BSC HOOD SURFACE, SETTLING PLATE - LEFT AND RIGHT, SLEEVE - LEFT AND RIGHT) AND GRADE B (VIABLE AIR PARTICULATES AND NONVIABLE AIR PARTICULATES) AND QC STERILITY: GRADE A (FINGERTIPS - LEFT AND RIGHT, BSC HOOD SURFACE, SETTLING PLATE - LEFT AND RIGHT, SLEEVE - LEFT AND RIGHT)) WERE PERFORMED. ALL RESULTS WERE REPORTED AS ''PASS''. QC LOT RELEASE ASSAY (IMPLEMENTED (B)(6) 2019) 3T3 RESIDUAL ASSAY Q2C-136 WAS REPORTED AS 0.12 - PASS. THE PATIENT'S MEDICAL HISTORY INCLUDED CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD) AND HYPERTENSION (HTN). THE PATIENT'S CONCOMITANT MEDICATION DETAILS WERE NOT REPORTED. ON (B)(6) 2025, THE PATIENT RECEIVED EPICEL GRAFTS (CULTURED EPIDERMAL AUTOGRAFTS) (LOT NUMBER: EE03423, EXPIRATION DATE: 03-JUN-2025), UDI NUMBER (B)(4) FOR BURN GREATER THAN 30% TOTAL BODY SURFACE AREA (TBSA). ON THE SAME DATE, A LIQUID MEDIA WAS NOTED IN CELLOPHANE SELF SEALING POUCH IN BOX#2 IN SECOND LAYER (FULL PACK OF 24 GRAFT DISHES) (PT: PRODUCT LEAKAGE). IT WAS REPORTED THAT THERE WERE NO OBVIOUS OPENED OR LEAKING CASSETTES. THE PRODUCT WAS NOT AFFECTED AS THERE WAS ONLY SOME SPILLAGE IN THE POUCH IN WHICH THE CASSETTES WERE PLACED. ON 19-JUN-2025, DOCUMENTATION INDICATED THAT THE SHEET GRAFTS WERE INTACT WITH NO DRAINAGE OR SEROMAS OBSERVED. A 65% TAKE OF THE CULTURED EPITHELIAL AUTOGRAFTS (CEA) AND A 96% TAKE OF THE MEEK GRAFTS WERE NOTED, WITH DRYING ONGOING. ADDITIONAL CEAS WERE REQUESTED FOR BILATERAL EXTREMITIES, WITH AN OFFERING DATE RECORDED AS (B)(6) 2025. BASED ON THIS, IT WAS CONCLUDED THAT AN ACTIVE INFECTION OF THE EPICEL GRAFTS WAS UNLIKELY, AS NO GRAFT LOSS WAS NOTED. THE DEATH WAS CONSIDERED UNRELATED TO EPICEL USE AND MORE LIKELY ATTRIBUTED TO THE PATIENT'S COMORBIDITIES AND BURN SEQUELAE. FURTHERMORE, THERE WOULD BE NO REQUEST FOR ADDITIONAL GRAFTS IF THERE WAS SUSPECTED GRAFT SITE INFECTION.* ON AN UNKNOWN DATE, SHORTLY AFTER SURGERY, THE PATIENT DEVELOPED PROFOUND SEPTIC SHOCK (PT: SEPTIC SHOCK). THE PATIENT WAS ON PRESSORS, REQUIRING INTUBATION AND FULL HEMODYNAMIC SUPPORT DUE TO ORGANS FAILURE. ALTHOUGH THE PHYSICIAN COULD NOT CONFIRM THE STATUS OF THE GRAFTS, HE ATTRIBUTED THE CARDIAC ARREST TO TYPICAL BURN-RELATED COMPLICATIONS, NOTING THAT THE PATIENT'S ADVANCED AGE AND SEVERITY OF INJURIES CONTRIBUTED TO AN OVERALL POOR PROGNOSIS. GIVEN THE PATIENT'S SEPTIC CONDITION, INFECTION WAS PRESUMED; HOWEVER, THE STATUS OF THE GRAFTS REMAINED UNCLEAR, AND IT COULD NOT BE DETERMINED WHETHER THE GRAFT SITES WERE THE SOURCE. ON (B)(6) 2025, ONE DAY POST-SURGERY FOLLOWING A MEEK TECHNIQUE SKIN GRAFT PROCEDURE, THE PATIENT'S CONDITION DETERIORATED (PT: CONDITION AGGRAVATED). THE FAMILY PLACED THE PATIENT ON COMFORT MEASURES (HOSPICE). ON THE SAME DATE, THE PATIENT EXPERIENCED CARDIAC ARREST (PT: CARDIAC ARREST). ON THE SAME DATE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS CARDIAC ARREST. IT WAS UNKNOWN IF THE AUTOPSY WAS PERFORMED. AT THE TIME OF THE REPORT, THE OUTCOME OF THE EVENT LIQUID MEDIA WAS NOTED IN CELLOPHANE SELF SEALING POUCH IN BOX#2 IN SECOND LAYER (FULL PACK OF 24 GRAFT DISHES) (PT: PRODUCT LEAKAGE) WAS RESOLVED, THE OUTCOME OF THE EVENT CARDIAC ARREST (PT: CARDIAC ARREST) WAS FATAL AND THE OUTCOME OF THE EVENT CONDITION DETERIORATED (PT: CONDITION AGGRAVATED) AND *SEPTIC SHOCK (PT:SEPTIC SHOCK)* WERE UNKNOWN. *IT WAS CONFIRMED THAT DEATH WAS NOT RELATED TO THE USE OF EPICEL , IT WAS MORE LIKELY DUE TO THE PATIENT'S COMORBIDITIES AND BURN-RELATED COMPLICATIONS.* NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED ON 04-JUN-2025 AND PROCESSED TOGETHER WITH THE INITIAL. FOLLOW-UP INFORMATION WAS RECEIVED FROM THE REPORTER ON 20-JUN-2025 AND INCLUDED: CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD) AND HYPERTENSION (HTN) WERE ADDED AS MEDICAL HISTORY. INDICATION WAS UPDATED. TWO ADDITIONAL EVENTS OF CARDIAC ARREST AND CONDITION DETERIORATED WERE ADDED. DEATH WAS REPORTED. QC RESULTS WERE ADDED. LOT NUMBER WAS UPDATED FROM EE03423-31 TO EE03423.THE PATIENT'S INITIALS WERE UPDATED FROM PGG TO PG. ADDITIONAL INFORMATION WAS RECEIVED ON 23-JUN-2025, 25-JUN-2025,26-JUN-2025 AND PROCESSED TOGETHER WITH THE FOLLOW-UP. FOLLOW-UP INFORMATION WAS RECEIVED FROM THE REPORTER ON 02-JUL-2025 AND INCLUDED: NEW EVENT OF SEPTIC SHOCK WAS ADDED TO THE CASE. IT WAS CONFIRMED THAT CARDIAC ARREST WAS NOT RELATED TO THE USE OF EPICEL. IT WAS CONFIRMED THAT THERE WAS NO INFECTION AT THE EPICEL GRAFT SITE, AND NO GRAFT LOSS. NEW REPORTERS WERE ADDED. THE PATIENT'S AGE WAS UPDATED FROM 84 TO 85 YEARS. NARRATIVE UPDATED ACCORDINGLY. NO FURTHER INFORMATION WAS PROVIDED. COMPANY COMMENT: VERICEL ASSESSED THE EVENT OF CARDIAC ARREST AS SERIOUS (DUE TO FATAL OUTCOME), NOT RELATED AND UNEXPECTED. THE EVENT OF SEPTIC SHOCK IS ASSESSED AS SERIOUS (DUE TO MEDICAL SIGNIFICANCE), NOT RELATED AND UNEXPECTED. THE EVENT OF CONDITION AGGRAVATED IS ASSESSED AS NON-SERIOUS, NOT RELATED AND UNEXPECTED. THE EVENT OF PRODUCT LEAKAGE IS ASSESSED AS NON-SERIOUS, POSSIBLY RELATED AND EXPECTED. THE CASE QUALIFIES AS A 15-DAY-REPORT AND 30-DAY MDR REPORT. Manufacturer Narrative: CASE REFERENCE NUMBER (B)(4) (FOLLOW-UP) IS A SPONTANEOUS REPORT INITIALLY RECEIVED FROM A COMPANY EMPLOYEE ON 03-JUN-2025, CONCERNING AN 84-YEAR-OLD FEMALE PATIENT WHO RECEIVED EPICEL GRAFTS. HOWEVER, EPICEL GRAFTS WERE NOTED WITH LIQUID MEDIA IN CELLOPHANE SELF-SEALING POUCH IN SECOND LAYER (PT: PRODUCT LEAKAGE). ADDITIONALLY, THE PATIENT'S CONDITION DETERIORATED (PT: CONDITION AGGRAVATED), THE PATIENT EXPERIENCED CARDIAC ARREST (PT: CARDIAC ARREST) AND EXPIRED. ON 02-JUN-2025, QC LOT RELEASE ASSAYS (GRAFT INSPECTION QC2-027, DUAL STAIN QC2-094, ENDOTOXIN QC2-010, STERILITY PRE-RELEASE QC2-005, STERILITY FINAL PRODUCT QC2-012) AND ENVIRONMENTAL RESULTS (MANUFACTURING: GRADE A (FINGERTIPS - LEFT AND RIGHT, BSC HOOD SURFACE, SETTLING PLATE - LEFT AND RIGHT, SLEEVE - LEFT AND RIGHT) AND GRADE B (VIABLE AIR PARTICULATES AND NONVIABLE AIR PARTICULATES) AND QC STERILITY: GRADE A (FINGERTIPS - LEFT AND RIGHT, BSC HOOD SURFACE, SETTLING PLATE - LEFT AND RIGHT, SLEEVE - LEFT AND RIGHT) WERE PERFORMED. ALL RESULTS WERE REPORTED AS ''PASS''. QC LOT RELEASE ASSAY (IMPLEMENTED 29-SEP-2019) 3T3 RESIDUAL ASSAY Q2C-136 WAS REPORTED AS 0.12 - PASS. THE PATIENT'S MEDICAL HISTORY *INCLUDED CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD) AND HYPERTENSION (HTN). THE PATIENT'S CONCOMITANT MEDICATION DETAILS WERE NOT REPORTED. ON (B)(6) 2025, THE PATIENT RECEIVED EPICEL GRAFTS (CULTURED EPIDERMAL AUTOGRAFTS) (LOT NUMBER: EE03423, EXPIRATION DATE: 03-JUN-2025), UDI NUMBER: (B)(4) FOR BURN GREATER THAN 30% TOTAL BODY SURFACE AREA (TBSA). ON THE SAME DATE, A LIQUID MEDIA WAS NOTED IN CELLOPHANE SELF SEALING POUCH IN BOX#2 IN SECOND LAYER (FULL PACK OF 24 GRAFT DISHES) (PT: PRODUCT LEAKAGE). IT WAS REPORTED THAT THERE WERE NO OBVIOUS OPENED OR LEAKING CASSETTES. THE PRODUCT WAS NOT AFFECTED AS THERE WAS ONLY SOME SPILLAGE IN THE POUCH IN WHICH THE CASSETTES WERE PLACED. ON (B)(6) 2025, ONE DAY POST-SURGERY FOLLOWING A MEEK TECHNIQUE SKIN GRAFT PROCEDURE, THE PATIENT'S CONDITION DETERIORATED (PT: CONDITION AGGRAVATED). THE FAMILY PLACED THE PATIENT ON COMFORT MEASURES (HOSPICE). ON THE SAME DATE, THE PATIENT EXPERIENCED CARDIAC ARREST (PT: CARDIAC ARREST). ON THE SAME DATE, THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS CARDIAC ARREST. IT WAS UNKNOWN IF THE AUTOPSY WAS PERFORMED. AT THE TIME OF THE REPORT, THE OUTCOME OF THE EVENT LIQUID MEDIA WAS NOTED IN CELLOPHANE SELF SEALING POUCH IN BOX#2 IN SECOND LAYER (FULL PACK OF 24 GRAFT DISHES) (PT: PRODUCT LEAKAGE) WAS RESOLVED, THE OUTCOME OF THE EVENT CARDIAC ARREST (PT: CARDIAC ARREST) WAS FATAL AND THE OUTCOME OF THE EVENT CONDITION DETERIORATED (PT: CONDITION AGGRAVATED) WAS UNKNOWN. NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED ON 04-JUN-2025 AND PROCESSED TOGETHER WITH THE INITIAL. FOLLOW-UP INFORMATION WAS RECEIVED FROM THE REPORTER ON 20-JUN-2025 AND INCLUDED: CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD) AND HYPERTENSION (HTN) WERE ADDED AS MEDICAL HISTORY. INDICATION WAS UPDATED. TWO ADDITIONAL EVENTS OF CARDIAC ARREST AND CONDITION DETERIORATED WERE ADDED. DEATH WAS REPORTED. QC RESULTS WERE ADDED. LOT NUMBER WAS UPDATED FROM EE03423-31 TO EE03423. THE PATIENT'S INITIALS WERE UPDATED FROM PGG TO PG. ADDITIONAL INFORMATION WAS RECEIVED ON 23-JUN-2025, 25-JUN-2025,26-JUN-2025 AND PROCESSED TOGETHER WITH THE FOLLOW-UP. NO FURTHER INFORMATION WAS PROVIDED. COMPANY COMMENT: VERICEL ASSESSED THE EVENT OF CARDIAC ARREST AS SERIOUS (DUE TO FATAL OUTCOME), UNLIKELY RELATED AND UNEXPECTED. THE EVENT OF CONDITION AGGRAVATED IS ASSESSED AS NON-SERIOUS, UNLIKELY RELATED AND UNEXPECTED. THE EVENT OF PRODUCT LEAKAGE IS ASSESSED AS NON-SERIOUS, POSSIBLY RELATED AND EXPECTED. THE CASE QUALIFIES AS A 15-DAY-REPORT AND 30-DAY MDR REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22327979&pc=LPB | 2029046-2025-02071 | 2025-05-22 04:00 | Death | BIOSENSE WEBSTER INC | 2025-06-25 04:00 | LPB | THERMOCOOL SMARTTOUCH SF | Adverse Event Without Identified Device or Use Problem | Cardiac Perforation | P030031 |  | 1 | Event Description: IT WAS REPORTED THAT A PATIENT UNDERWENT A VENTRICULAR TACHYCARDIA (VT) PROCEDURE WITH A THERMOCOOL SMARTTOUCH SF AND THE PATIENT EXPERIENCED CARDIAC TAMPONADE TREATED WITH PERCUTANEOUS PERICARDIAL DRAINAGE, THEN SURGICAL DRAINAGE. SUBSEQUENTLY, THE PATIENT EXPIRED. ONE HOUR AFTER THE PROCEDURE, THE PATIENT HAD TAMPONADE, AND PERCUTANEOUS PERICARDIAL DRAINAGE WAS PERFORMED. SUBSEQUENTLY, A SURGICAL DRAINAGE WAS PERFORMED. THE PATIENT DIED DESPITE THE TWO DRAINAGES. THE DOCTORS OF THE DEPARTMENT DO NOT BLAME THE EQUIPMENT AND BELIEVE THAT THE COMPLICATION IS RELATED TO THE PATIENT'S CONDITION AND THE TYPE OF PROCEDURE. NO PRE-PROCEDURE EVENTS. ADDITIONAL INFORMATION WAS RECEIVED. CARDIAC PERFORATION WAS CONFIRMED IN THE POSTERIOR LEFT VENTRICLE. THE PHYSICIAN¿S OPINION ON THE CAUSE OF THIS ADVERSE EVENT WAS PROCEDURE AND PATIENT CONDITION, AND TAMPONADE AFTER PERFORATION. THE PROCEDURAL ACT WAS 300 SECONDS. PRIOR TO NOTING THE CT, ABLATION WAS PERFORMED DURING THE PROCEDURE. ADVERSE EVENT SYMPTOMS OCCURRED AFTER THE END OF PROCEDURE. ONE TRANSSEPTAL PUNCTURE SITE WAS PERFORMED DURING THE PROCEDURE. NUMBER OF PERFORMED RADIO FREQUENCY (RF) ABLATIONS WERE 29. NO ABLATIONS WERE PERFORMED WITHIN THE PULMONARY VEINS IT WAS A VT PROCEDURE. TRANSSEPTAL PUNCTURE WAS PERFORMED WITH ABBOTT BRK. NO EVIDENCE OF A STEAM POP. THE FLOW SETTING FOR IRRIGATION WAS 15ML/MIN. CORRECT CATHETER SETTINGS WERE SELECTED ON THE GENERATOR. NO ISSUES WITH FLOW RATE CHANGES AT THE START OF ABLATION. NO ERROR MESSAGES WERE OBSERVED ON BIOSENSE WEBSTER EQUIPMENT DURING THE PROCEDURE. FORCE VISUALIZATION FEATURES USED WERE DASHBOARD AND VECTOR. PARAMETERS FOR STABILITY USED WERE RANGE 3MM, TIME 3 SECONDS, 30% AND 3G SIZE: 3. ADDITIONAL FILTER USED WITH THE VISITAG WAS RESPIRATION. COLOR OPTION USED PROSPECTIVELY WAS ABLATION INDEX. Manufacturer Narrative: D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. E1. INITIAL REPORTER PHONE: (B)(6). AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22328045&pc=NPT | 2025587-2025-04500 | 2025-06-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-25 04:00 | NPT | EVOLUT FX PLUS VALVE | Difficult or Delayed Positioning; Material Invagination; Patient Device Interaction Problem | Low Blood Pressure/ Hypotension; No Clinical Signs, Symptoms or Conditions | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION PROCEDURE, A PATIENT WITH DIALYSIS DEPENDENCE, AN EJECTION FRACTION OF 30%, A HISTORY OF SINGLE BYPASS LEFT INTERNAL MAMMARY ARTERY (LIMA), AND PREVIOUS ASCENDING AORTA RECONSTRUCTION PRESENTED WITH A HEAVILY CALCIFIED AND DEGENERATED NON-MEDTRONIC BIOPROSTHETIC VALVE (EDWARDS SAPIEN S3), SEVERE AORTIC INSUFFICIENCY, AND MITRAL INSUFFICIENCY. THE INITIAL ATTEMPT INVOLVED PRE-DILATION WITH A 24 X 60 MILLIMETER (MM) NON-MEDTRONIC BALLOON (VACS III), DURING WHICH NO FULL EXPANSION WAS ACHIEVED. THE FIRST DEPLOYMENT OF THE EVOLUT FX+ 29 MM VALVE WAS TOO LOW, RESULTING IN MALPOSITION. DURING THE SECOND DEPLOYMENT ATTEMPT, SEVERE PRESSURE LOSS, HEMODYNAMIC INSTABILITY, AND LOSS OF BLOOD PRESSURE WAS OBSERVED. IT WAS RECOMMENDED TO RETRACT AND STABILIZE THE VALVE. THE PATIENT REQUIRED TRANSITION TO A HEART-LUNG MACHINE, RESUSCITATION, AND INTUBATION. OPEN SURGERY WAS NOT PURSUED DUE TO THE PATIENT¿S MEDICAL HISTORY. DURING THE SECOND ATTEMPT, A NEW EVOLUT FX 29 MILLIMETER PROSTHESIS WAS DEPLOYED VIA CONTRALATERAL ACCESS, BUT A SUDDEN PRESSURE DROP AND ABSENCE OF VENOUS RETURN OCCURRED DUE TO DISLOCATION OF THE ARTERIAL ACCESS FOR THE HEART-LUNG MACHINE, LEADING TO VESSEL RUPTURE AND RETROPERITONEAL BLEEDING. A VISCERAL SURGEON RECANNULATED THE VESSEL. THE PATIENT EXPERIENCED PERSISTENT VENTRICULAR FIBRILLATION, POOR LABORATORY RESULTS, AND REQUIRED SEVEN BLOOD TRANSFUSIONS FOR STABILIZATION. IT WAS DECIDED TO IMPLANT THE SECOND VALVE. CHALLENGES IN DEPTH ADJUSTMENTS WERE OBSERVED AS THE SECOND VALVE DISLODGED FURTHER DOWN, NECESSITATING FOUR RECAPTURE ATTEMPTS BEFORE FINAL DEPLOYMENT AT APPROXIMATELY 4 MM. TRANSESOPHAGEAL ECHOCARDIOGRAPHY REVEALED MODERATE PARAVALVULAR LEAK (PVL), AND POST-DILATION WITH A NEW 24 MM NON-MEDTRONIC BALLOON (VACS III) WAS PERFORMED. ANGIOGRAPHY AND ECHOCARDIOGRAPHY SHOWED ONLY PARTIAL IMPROVEMENT OF THE PVL.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION PROCEDURE, A PATIENT WITH DIALYSIS DEPENDENCE, AN EJECTION FRACTION OF 30%, A HISTORY OF SINGLE BYPASS LEFT INTERNAL MAMMARY ARTERY (LIMA), AND PREVIOUS ASCENDING AORTA RECONSTRUCTION PRESENTED WITH A HEAVILY CALCIFIED AND DEGENERATED NON-MEDTRONIC BIOPROSTHETIC VALVE (EDWARDS SAPIEN S3), SEVERE AORTIC INSUFFICIENCY, AND MITRAL INSUFFICIENCY. THE INITIAL ATTEMPT INVOLVED PRE-DILATION WITH A 24 X 60 MILLIMETER (MM) NON-MEDTRONIC BALLOON (VACS III), DURING WHICH NO FULL EXPANSION WAS ACHIEVED. THE FIRST DEPLOYMENT OF THE EVOLUT FX+ 29 MM VALVE WAS TOO LOW, RESULTING IN MALPOSITION. DURING THE SECOND DEPLOYMENT ATTEMPT, SEVERE PRESSURE LOSS, HEMODYNAMIC INSTABILITY, AND LOSS OF BLOOD PRESSURE WAS OBSERVED. IT WAS RECOMMENDED TO RETRACT AND STABILIZE THE VALVE. THE PATIENT REQUIRED TRANSITION TO A HEART-LUNG MACHINE (HLM), RESUSCITATION, AND INTUBATION. OPEN SURGERY WAS NOT PURSUED DUE TO THE PATIENT¿S MEDICAL HISTORY. DURING THE SECOND ATTEMPT, A NEW EVOLUT FX 29 MILLIMETER PROSTHESIS WAS DEPLOYED VIA CONTRALATERAL ACCESS, BUT A SUDDEN PRESSURE DROP AND ABSENCE OF VENOUS RETURN OCCURRED DUE TO DISLOCATION OF THE ARTERIAL ACCESS FOR THE HLM, LEADING TO VESSEL RUPTURE AND RETROPERITONEAL BLEEDING. A VISCERAL SURGEON RECANNULATED THE VESSEL. THE PATIENT EXPERIENCED PERSISTENT VENTRICULAR FIBRILLATION, POOR LABORATORY RESULTS, AND REQUIRED SEVEN BLOOD TRANSFUSIONS FOR STABILIZATION. IT WAS DECIDED TO IMPLANT THE SECOND VALVE. CHALLENGES IN DEPTH ADJUSTMENTS WERE OBSERVED AS THE SECOND VALVE DISLODGED FURTHER DOWN, NECESSITATING FOUR RECAPTURE ATTEMPTS BEFORE FINAL DEPLOYMENT AT APPROXIMATELY 4 MM. TRANSESOPHAGEAL ECHOCARDIOGRAPHY REVEALED MODERATE PARAVALVULAR LEAK (PVL), AND POST-DILATION WITH A NEW 24 MM NON-MEDTRONIC BALLOON (VACS III) WAS PERFORMED. ANGIOGRAPHY AND ECHOCARDIOGRAPHY SHOWED ONLY PARTIAL IMPROVEMENT OF THE PVL. ADDITIONAL INFORMATION WAS RECEIVED THAT A NON-MEDTRONIC (SAFARI) GUIDEWIRE WAS USED FOR THE PROCEDURE. THE FIRST VALVE WAS SUCCESSFULLY RECAPTURED AND RETRIEVED FROM THE PATIENT. A PIGTAIL CATHETER WAS NOT USED, THE PHYSICIAN OBTAINED ORIENTATION FOR DEPLOYMENT FROM THE BOTTOM STRUTS OF THE NON-MEDTRONIC VALVE. THE IMPLANT DEPTH PRIOR TO DISLODGEMENT WAS APPROXIMATELY 4 MM ON THE NON-CORONARY CUSP (NCC) AND 7 MM ON THE LEFT CORONARY CUSP (LCC). THE POOR LAB RESULTS OBSERVED PERTAINED TO BLOOD VALUES SUCH AS LACTATE, LOW PH, AND HEMOGLOBIN (HB). PER THE PHYSICIAN, THE VESSEL RUPTURE/RETROPERITONEAL BLEEDING WAS CAUSED BY THE CANNULATION OF THE HLM AND NOT THE FIRST DELIVERY CATHETER SYSTEM (DCS). THE PATIENT SUBSEQUENTLY DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CORRECTION: SECTION D10 INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012281972), USE BY DATE 2026-05-22, UDI (B)(4). UPDATED DATA: THE FOLLOWING FIELDS WERE UPDATED AS THEY PERTAIN TO THE SYSTEM REPORTED DEVICE D-EVOLUTFX-2329 (LOT: 0012281972) D9 H2 H3 H6 PRODUCT ANALYSIS: UPON RECEIPT OF THE SUSPECT COMPLAINT DEVICE AT MEDTRONIC¿S QUALITY LABORATORY, VISUAL ANALYSIS SHOWED THE DELIVERY CATHETER SYSTEM (DCS) WAS RECEIVED WITHOUT A VALVE LOADED WITHIN THE CAPSULE. THE DEVICE WAS RECEIVED WITH THE HANDLE FULLY CLOSED AND CAPSULE FULLY CLOSED. THE NOSECONE WAS OVERCAPTURED BY THE CAPSULE. THE HANDLE APPEARED TO RETRACT AND ADVANCE THE CAPSULE. THE TRIGGER MOVED TO FULLY ADVANCE AND RETRACT POSITIONS AND LOCKED IN PLACED WHEN RELEASED. DELAMINATION WAS EVIDENT TO THE PROXIMAL END OF THE CAPSULE. KINKS WERE EVIDENT ON THE INNER MEMBER. NO OTHER DEFORMATION EVIDENT TO THE REMAINDER OF THE DEVICE. THE REPORTED POSITIONING DIFFICULT COULD NOT BE CONFIRMED THROUGH ANALYSIS MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012281972); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVOLUTFX-29 (J306801); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025 PRODUCT ID D-EVOLUTFX-2329 (0012740591); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EDWARDS SAPIEN S3 (N/A); PRODUCT TYPE: HEART VALVES; PRODUCT ID 24 X 60 VACS III BALLOON (N/A); PRODUCT TYPE: DILATION BALLOON; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: B1 D4 H2 H3 H6 IMAGE REVIEW: 13 IMAGES OF THE EVENT WERE PROVIDED. THERE WAS NO COMPUTED TOMOGRAPHY (CT) OR EXECUTIVE SUMMARY PROVIDED FOR ANATOMICAL CONSIDERATIONS. FLUOROSCOPIC VALVE LOAD INSPECTION WAS PERFORMED AND A MISLOAD APPEARED TO BE PRESENT BY OVERLAP BEYOND NODE 4. PER MEDTRONIC BEST PRACTICES, OVERLAP PRIOR TO NODE 4 IS CONSIDERED A GOOD LOAD; OVERLAP AT NODE 4 AND BEYOND IS CONSIDERED A MISLOAD. AS STATED IN THE INSTRUCTIONS FOR USE (IFU), IF A MISLOAD IS DETECTED, DO NOT ATTEMPT TO RELOAD THE BIOPROSTHESIS. DISCARD THE ENTIRE SYSTEM. THE VALVE, CATHETER, LOADING SYSTEM, LOADING TRAY, AND SALINE MUST ALL BE REPLACED WITH NEW STERILE COMPONENTS. AFTER CLOSE INVESTIGATION OF IMAGE 4, THE VALVE APPEARS TO HAVE AN INFOLD. PER MEDTRONIC BEST PRACTICE, AN INFOLD APPEARS AS AN INWARD FOLD OR CREASE OF THE VALVE FRAME IDENTIFIED AS A DARK LINE UNDER INSPECTION, WHEN THIS HAPPENS PORTION IS NO LONGER IN CONTACT WITH THE ANATOMY AND CAN CREATE A DROP IN BLOOD PRESSURE. THIS CAN BE INDUCED UPON DEPLOYMENT OR RECAPTURED IN THE PRESENCE OF DENSE FOCAL CALCIFICATION. THE VALVE WAS RECAPTURED AND IT WAS REPORTED THAT DURING THE SECOND ATTEMPT, A NEW EVOLUT FX 29 MM PROSTHESIS WAS DEPLOYED VIA CONTRALATERAL ACCESS. FLUOROSCOPIC LOAD INSPECTION WAS PROVIDED AND VALVE APPEARS TO BE LOADED PROPERLY WITH NO MISLOAD. IT WAS REPORTED THAT IT WAS DECIDED TO IMPLANT THE SECOND VALVE. CHALLENGES IN DEPTH ADJUSTMENTS WERE OBSERVED AS THE SECOND VALVE DISLODGED FURTHER DOWN, NECESSITATING FOUR RECAPTURE ATTEMPTS BEFORE FINAL DEPLOYMENT AT APPROXIMATELY 4 MM. TRANSESOPHAGEAL ECHOCARDIOGRAPHY REVEALED MODERATE PARAVALVULAR LEAK (PVL), AND POST-DILATION WITH A NEW 24 MM NON-MEDTRONIC BALLOON (VACS III) WAS PERFORMED. ANGIOGRAPHY AND ECHOCARDIOGRAPHY SHOWED ONLY PARTIAL IMPROVEMENT OF THE PVL. ANGIOGRAPHIC EVIDENCE SHOWS PVL AFTER POST BALLOON DILATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22328249&pc=NPT | 2025587-2025-04504 | 2025-06-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-25 04:00 | NPT | EVOLUT FX VALVE | Difficult or Delayed Positioning; Perivalvular Leak; Patient Device Interaction Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Rupture; Valvular Insufficiency/ Regurgitation; No Clinical Signs, Symptoms or Conditions | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION PROCEDURE, A PATIENT WITH DIALYSIS DEPENDENCE, AN EJECTION FRACTION OF 30%, A HISTORY OF SINGLE BYPASS LEFT INTERNAL MAMMARY ARTERY (LIMA), AND PREVIOUS ASCENDING AORTA RECONSTRUCTION PRESENTED WITH A HEAVILY CALCIFIED AND DEGENERATED NON-MEDTRONIC BIOPROSTHETIC VALVE (EDWARDS SAPIEN S3), SEVERE AORTIC INSUFFICIENCY, AND MITRAL INSUFFICIENCY. THE INITIAL ATTEMPT INVOLVED PRE-DILATION WITH A 24 X 60 MILLIMETER (MM) NON-MEDTRONIC BALLOON (VACS III), DURING WHICH NO FULL EXPANSION WAS ACHIEVED. THE FIRST DEPLOYMENT OF THE EVOLUT FX+ 29 MM VALVE WAS TOO LOW, RESULTING IN MALPOSITION. DURING THE SECOND DEPLOYMENT ATTEMPT, SEVERE PRESSURE LOSS, HEMODYNAMIC INSTABILITY, AND LOSS OF BLOOD PRESSURE WAS OBSERVED. IT WAS RECOMMENDED TO RETRACT AND STABILIZE THE VALVE. THE PATIENT REQUIRED TRANSITION TO A HEART-LUNG MACHINE, RESUSCITATION, AND INTUBATION. OPEN SURGERY WAS NOT PURSUED DUE TO THE PATIENT¿S MEDICAL HISTORY. DURING THE SECOND ATTEMPT, A NEW EVOLUT FX 29 MILLIMETER PROSTHESIS WAS DEPLOYED VIA CONTRALATERAL ACCESS, BUT A SUDDEN PRESSURE DROP AND ABSENCE OF VENOUS RETURN OCCURRED DUE TO DISLOCATION OF THE ARTERIAL ACCESS FOR THE HEART-LUNG MACHINE, LEADING TO VESSEL RUPTURE AND RETROPERITONEAL BLEEDING. A VISCERAL SURGEON RECANNULATED THE VESSEL. THE PATIENT EXPERIENCED PERSISTENT VENTRICULAR FIBRILLATION, POOR LABORATORY RESULTS, AND REQUIRED SEVEN BLOOD TRA NSFUSIONS FOR STABILIZATION. IT WAS DECIDED TO IMPLANT THE SECOND VALVE. CHALLENGES IN DEPTH ADJUSTMENTS WERE OBSERVED AS THE SECOND VALVE DISLODGED FURTHER DOWN, NECESSITATING FOUR RECAPTURE ATTEMPTS BEFORE FINAL DEPLOYMENT AT APPROXIMATELY 4 MM. TRANSESOPHAGEAL ECHOCARDIOGRAPHY REVEALED MODERATE PARAVALVULAR LEAK (PVL), AND POST-DILATION WITH A NEW 24 MM NON-MEDTRONIC BALLOON (VACS III) WAS PERFORMED. ANGIOGRAPHY AND ECHOCARDIOGRAPHY SHOWED ONLY PARTIAL IMPROVEMENT OF THE PVL.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION PROCEDURE, A PATIENT WITH DIALYSIS DEPENDENCE, AN EJECTION FRACTION OF 30%, A HISTORY OF SINGLE BYPASS LEFT INTERNAL MAMMARY ARTERY (LIMA), AND PREVIOUS ASCENDING AORTA RECONSTRUCTION PRESENTED WITH A HEAVILY CALCIFIED AND DEGENERATED NON-MEDTRONIC BIOPROSTHETIC VALVE (EDWARDS SAPIEN S3), SEVERE AORTIC INSUFFICIENCY, AND MITRAL INSUFFICIENCY. THE INITIAL ATTEMPT INVOLVED PRE-DILATION WITH A 24 X 60 MILLIMETER (MM) NON-MEDTRONIC BALLOON (VACS III), DURING WHICH NO FULL EXPANSION WAS ACHIEVED. THE FIRST DEPLOYMENT OF THE EVOLUT FX+ 29 MM VALVE WAS TOO LOW, RESULTING IN MALPOSITION. DURING THE SECOND DEPLOYMENT ATTEMPT, SEVERE PRESSURE LOSS, HEMODYNAMIC INSTABILITY, AND LOSS OF BLOOD PRESSURE WAS OBSERVED. IT WAS RECOMMENDED TO RETRACT AND STABILIZE THE VALVE. THE PATIENT REQUIRED TRANSITION TO A HEART-LUNG MACHINE (HLM), RESUSCITATION, AND INTUBATION. OPEN SURGERY WAS NOT PURSUED DUE TO THE PATIENT¿S MEDICAL HISTORY. DURING THE SECOND ATTEMPT, A NEW EVOLUT FX 29 MILLIMETER PROSTHESIS WAS DEPLOYED VIA CONTRALATERAL ACCESS, BUT A SUDDEN PRESSURE DROP AND ABSENCE OF VENOUS RETURN OCCURRED DUE TO DISLOCATION OF THE ARTERIAL ACCESS FOR THE HLM, LEADING TO VESSEL RUPTURE AND RETROPERITONEAL BLEEDING. A VISCERAL SURGEON RECANNULATED THE VESSEL. THE PATIENT EXPERIENCED PERSISTENT VENTRICULAR FIBRILLATION, POOR LABORATORY RESULTS, AND REQUIRED SEVEN BLOOD TRANSFUSIONS FOR STABILIZATION. IT WAS DECIDED TO IMPLANT THE SECOND VALVE. CHALLENGES IN DEPTH ADJUSTMENTS WERE OBSERVED AS THE SECOND VALVE DISLODGED FURTHER DOWN, NECESSITATING FOUR RECAPTURE ATTEMPTS BEFORE FINAL DEPLOYMENT AT APPROXIMATELY 4 MM. TRANSESOPHAGEAL ECHOCARDIOGRAPHY REVEALED MODERATE PARAVALVULAR LEAK (PVL), AND POST-DILATION WITH A NEW 24 MM NON-MEDTRONIC BALLOON (VACS III) WAS PERFORMED. ANGIOGRAPHY AND ECHOCARDIOGRAPHY SHOWED ONLY PARTIAL IMPROVEMENT OF THE PVL. ADDITIONAL INFORMATION WAS RECEIVED THAT A NON-MEDTRONIC (SAFARI) GUIDEWIRE WAS USED FOR THE PROCEDURE. THE FIRST VALVE WAS SUCCESSFULLY RECAPTURED AND RETRIEVED FROM THE PATIENT. A PIGTAIL CATHETER WAS NOT USED, THE PHYSICIAN OBTAINED ORIENTATION FOR DEPLOYMENT FROM THE BOTTOM STRUTS OF THE NON-MEDTRONIC VALVE. THE IMPLANT DEPTH PRIOR TO DISLODGEMENT WAS APPROXIMATELY 4 MM ON THE NON-CORONARY CUSP (NCC) AND 7 MM ON THE LEFT CORONARY CUSP (LCC). THE POOR LAB RESULTS OBSERVED PERTAINED TO BLOOD VALUES SUCH AS LACTATE, LOW PH, AND HEMOGLOBIN (HB). PER THE PHYSICIAN, THE VESSEL RUPTURE/RETROPERITONEAL BLEEDING WAS CAUSED BY THE CANNULATION OF THE HLM AND NOT THE FIRST DELIVERY CATHETER SYSTEM (DCS). THE PATIENT SUBSEQUENTLY DIED. Manufacturer Narrative: UPDATED DATA: D4 H2 H3 H6. IMAGE REVIEW: 13 IMAGES OF THE EVENT WERE PROVIDED. THERE WAS NO COMPUTED TOMOGRAPHY (CT) OR EXECUTIVE SUMMARY PROVIDED FOR ANATOMICAL CONSIDERATIONS. FLUOROSCOPIC VALVE LOAD INSPECTION WAS PERFORMED AND A MISLOAD APPEARED TO BE PRESENT BY OVERLAP BEYOND NODE 4. PER MEDTRONIC BEST PRACTICES, OVERLAP PRIOR TO NODE 4 IS CONSIDERED A GOOD LOAD; OVERLAP AT NODE 4 AND BEYOND IS CONSIDERED A MISLOAD. AS STATED IN THE INSTRUCTIONS FOR USE (IFU), IF A MISLOAD IS DETECTED, DO NOT ATTEMPT TO RELOAD THE BIOPROSTHESIS. DISCARD THE ENTIRE SYSTEM. THE VALVE, CATHETER, LOADING SYSTEM, LOADING TRAY, AND SALINE MUST ALL BE REPLACED WITH NEW STERILE COMPONENTS. AFTER CLOSE INVESTIGATION OF IMAGE 4, THE VALVE APPEARS TO HAVE AN INFOLD. PER MEDTRONIC BEST PRACTICE, AN INFOLD APPEARS AS AN INWARD FOLD OR CREASE OF THE VALVE FRAME IDENTIFIED AS A DARK LINE UNDER INSPECTION, WHEN THIS HAPPENS PORTION IS NO LONGER IN CONTACT WITH THE ANATOMY AND CAN CREATE A DROP IN BLOOD PRESSURE. THIS CAN BE INDUCED UPON DEPLOYMENT OR RECAPTURED IN THE PRESENCE OF DENSE FOCAL CALCIFICATION. THE VALVE WAS RECAPTURED AND IT WAS REPORTED THAT DURING THE SECOND ATTEMPT, A NEW EVOLUT FX 29 MM PROSTHESIS WAS DEPLOYED VIA CONTRALATERAL ACCESS. FLUOROSCOPIC LOAD INSPECTION WAS PROVIDED AND VALVE APPEARS TO BE LOADED PROPERLY WITH NO MISLOAD. IT WAS REPORTED THAT IT WAS DECIDED TO IMPLANT THE SECOND VALVE. CHALLENGES IN DEPTH ADJUSTMENTS WERE OBSERVED AS THE SECOND VALVE DISLODGED FURTHER DOWN, NECESSITATING FOUR RECAPTURE ATTEMPTS BEFORE FINAL DEPLOYMENT AT APPROXIMATELY 4 MM. TRANSESOPHAGEAL ECHOCARDIOGRAPHY REVEALED MODERATE PARAVALVULAR LEAK (PVL), AND POST-DILATION WITH A NEW 24 MM NON-MEDTRONIC BALLOON (VACS III) WAS PERFORMED. ANGIOGRAPHY AND ECHOCARDIOGRAPHY SHOWED ONLY PARTIAL IMPROVEMENT OF THE PVL. ANGIOGRAPHIC EVIDENCE SHOWS PVL AFTER POST BALLOON DILATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED: B2, B5, H1, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012281972); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVFXPLUS-29 (K089923); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025 PRODUCT ID D-EVOLUTFX-2329 (0012740591); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EDWARDS SAPIEN S3 (N/A); PRODUCT TYPE: HEART VALVES; PRODUCT ID 24 X 60 VACS III BALLOON (N/A); PRODUCT TYPE: DILATION BALLOON; SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22328331&pc=NPT | 2025587-2025-04505 | 2025-06-18 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-25 04:00 | NPT | EVOLUT FX DCS | Difficult or Delayed Positioning; Device Dislodged or Dislocated; Patient Device Interaction Problem | Hemorrhage/Bleeding; Low Blood Pressure/ Hypotension; Rupture; No Clinical Signs, Symptoms or Conditions | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION PROCEDURE, A PATIENT WITH DIALYSIS DEPENDENCE, AN EJECTION FRACTION OF 30%, A HISTORY OF SINGLE BYPASS LEFT INTERNAL MAMMARY ARTERY (LIMA), AND PREVIOUS ASCENDING AORTA RECONSTRUCTION PRESENTED WITH A HEAVILY CALCIFIED AND DEGENERATED NON-MEDTRONIC BIOPROSTHETIC VALVE (EDWARDS SAPIEN S3), SEVERE AORTIC INSUFFICIENCY, AND MITRAL INSUFFICIENCY. THE INITIAL ATTEMPT INVOLVED PRE-DILATION WITH A 24 X 60 MILLIMETER (MM) NON-MEDTRONIC BALLOON (VACS III), DURING WHICH NO FULL EXPANSION WAS ACHIEVED. THE FIRST DEPLOYMENT OF THE EVOLUT FX+ 29 MM VALVE WAS TOO LOW, RESULTING IN MALPOSITION. DURING THE SECOND DEPLOYMENT ATTEMPT, SEVERE PRESSURE LOSS, HEMODYNAMIC INSTABILITY, AND LOSS OF BLOOD PRESSURE WAS OBSERVED. IT WAS RECOMMENDED TO RETRACT AND STABILIZE THE VALVE. THE PATIENT REQUIRED TRANSITION TO A HEART-LUNG MACHINE, RESUSCITATION, AND INTUBATION. OPEN SURGERY WAS NOT PURSUED DUE TO THE PATIENT¿S MEDICAL HISTORY. DURING THE SECOND ATTEMPT, A NEW EVOLUT FX 29 MILLIMETER PROSTHESIS WAS DEPLOYED VIA CONTRALATERAL ACCESS, BUT A SUDDEN PRESSURE DROP AND ABSENCE OF VENOUS RETURN OCCURRED DUE TO DISLOCATION OF THE ARTERIAL ACCESS FOR THE HEART-LUNG MACHINE, LEADING TO VESSEL RUPTURE AND RETROPERITONEAL BLEEDING. A VISCERAL SURGEON RECANNULATED THE VESSEL. THE PATIENT EXPERIENCED PERSISTENT VENTRICULAR FIBRILLATION, POOR LABORATORY RESULTS, AND REQUIRED SEVEN BLOOD TRA INFUSIONS FOR STABILIZATION. IT WAS DECIDED TO IMPLANT THE SECOND VALVE. CHALLENGES IN DEPTH ADJUSTMENTS WERE OBSERVED AS THE SECOND VALVE DISLODGED FURTHER DOWN, NECESSITATING FOUR RECAPTURE ATTEMPTS BEFORE FINAL DEPLOYMENT AT APPROXIMATELY 4 MM. TRANSESOPHAGEAL ECHOCARDIOGRAPHY REVEALED MODERATE PARAVALVULAR LEAK (PVL), AND POST-DILATION WITH A NEW 24 MM NON-MEDTRONIC BALLOON (VACS III) WAS PERFORMED. ANGIOGRAPHY AND ECHOCARDIOGRAPHY SHOWED ONLY PARTIAL IMPROVEMENT OF THE PVL.Event Description: IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION PROCEDURE, A PATIENT WITH DIALYSIS DEPENDENCE, AN EJECTION FRACTION OF 30%, A HISTORY OF SINGLE BYPASS LEFT INTERNAL MAMMARY ARTERY (LIMA), AND PREVIOUS ASCENDING AORTA RECONSTRUCTION PRESENTED WITH A HEAVILY CALCIFIED AND DEGENERATED NON-MEDTRONIC BIOPROSTHETIC VALVE (EDWARDS SAPIEN S3), SEVERE AORTIC INSUFFICIENCY, AND MITRAL INSUFFICIENCY. THE INITIAL ATTEMPT INVOLVED PRE-DILATION WITH A 24 X 60 MILLIMETER (MM) NON-MEDTRONIC BALLOON (VACS III), DURING WHICH NO FULL EXPANSION WAS ACHIEVED. THE FIRST DEPLOYMENT OF THE EVOLUT FX+ 29 MM VALVE WAS TOO LOW, RESULTING IN MALPOSITION. DURING THE SECOND DEPLOYMENT ATTEMPT, SEVERE PRESSURE LOSS, HEMODYNAMIC INSTABILITY, AND LOSS OF BLOOD PRESSURE WAS OBSERVED. IT WAS RECOMMENDED TO RETRACT AND STABILIZE THE VALVE. THE PATIENT REQUIRED TRANSITION TO A HEART-LUNG MACHINE (HLM), RESUSCITATION, AND INTUBATION. OPEN SURGERY WAS NOT PURSUED DUE TO THE PATIENT¿S MEDICAL HISTORY. DURING THE SECOND ATTEMPT, A NEW EVOLUT FX 29 MILLIMETER PROSTHESIS WAS DEPLOYED VIA CONTRALATERAL ACCESS, BUT A SUDDEN PRESSURE DROP AND ABSENCE OF VENOUS RETURN OCCURRED DUE TO DISLOCATION OF THE ARTERIAL ACCESS FOR THE HLM, LEADING TO VESSEL RUPTURE AND RETROPERITONEAL BLEEDING. A VISCERAL SURGEON RECANNULATED THE VESSEL. THE PATIENT EXPERIENCED PERSISTENT VENTRICULAR FIBRILLATION, POOR LABORATORY RESULTS, AND REQUIRED SEVEN BLOOD TRANSFUSIONS FOR STABILIZATION. IT WAS DECIDED TO IMPLANT THE SECOND VALVE. CHALLENGES IN DEPTH ADJUSTMENTS WERE OBSERVED AS THE SECOND VALVE DISLODGED FURTHER DOWN, NECESSITATING FOUR RECAPTURE ATTEMPTS BEFORE FINAL DEPLOYMENT AT APPROXIMATELY 4 MM. TRANSESOPHAGEAL ECHOCARDIOGRAPHY REVEALED MODERATE PARAVALVULAR LEAK (PVL), AND POST-DILATION WITH A NEW 24 MM NON-MEDTRONIC BALLOON (VACS III) WAS PERFORMED. ANGIOGRAPHY AND ECHOCARDIOGRAPHY SHOWED ONLY PARTIAL IMPROVEMENT OF THE PVL. ADDITIONAL INFORMATION WAS RECEIVED THAT A NON-MEDTRONIC (SAFARI) GUIDEWIRE WAS USED FOR THE PROCEDURE. THE FIRST VALVE WAS SUCCESSFULLY RECAPTURED AND RETRIEVED FROM THE PATIENT. A PIGTAIL CATHETER WAS NOT USED, THE PHYSICIAN OBTAINED ORIENTATION FOR DEPLOYMENT FROM THE BOTTOM STRUTS OF THE NON-MEDTRONIC VALVE. THE IMPLANT DEPTH PRIOR TO DISLODGEMENT WAS APPROXIMATELY 4 MM ON THE NON-CORONARY CUSP (NCC) AND 7 MM ON THE LEFT CORONARY CUSP (LCC). THE POOR LAB RESULTS OBSERVED PERTAINED TO BLOOD VALUES SUCH AS LACTATE, LOW PH, AND HEMOGLOBIN (HB). PER THE PHYSICIAN, THE VESSEL RUPTURE/RETROPERITONEAL BLEEDING WAS CAUSED BY THE CANNULATION OF THE HLM AND NOT THE FIRST DELIVERY CATHETER SYSTEM (DCS). THE PATIENT SUBSEQUENTLY DIED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: UPDATED DATA: D4 H2 H3 H6. IMAGE REVIEW: 13 IMAGES OF THE EVENT WERE PROVIDED. THERE WAS NO COMPUTED TOMOGRAPHY (CT) OR EXECUTIVE SUMMARY PROVIDED FOR ANATOMICAL CONSIDERATIONS. FLUOROSCOPIC VALVE LOAD INSPECTION WAS PERFORMED AND A MISLOAD APPEARED TO BE PRESENT BY OVERLAP BEYOND NODE 4. PER MEDTRONIC BEST PRACTICES, OVERLAP PRIOR TO NODE 4 IS CONSIDERED A GOOD LOAD; OVERLAP AT NODE 4 AND BEYOND IS CONSIDERED A MISLOAD. AS STATED IN THE INSTRUCTIONS FOR USE (IFU), IF A MISLOAD IS DETECTED, DO NOT ATTEMPT TO RELOAD THE BIOPROSTHESIS. DISCARD THE ENTIRE SYSTEM. THE VALVE, CATHETER, LOADING SYSTEM, LOADING TRAY, AND SALINE MUST ALL BE REPLACED WITH NEW STERILE COMPONENTS. AFTER CLOSE INVESTIGATION OF IMAGE 4, THE VALVE APPEARS TO HAVE AN INFOLD. PER MEDTRONIC BEST PRACTICE, AN INFOLD APPEARS AS AN INWARD FOLD OR CREASE OF THE VALVE FRAME IDENTIFIED AS A DARK LINE UNDER INSPECTION, WHEN THIS HAPPENS PORTION IS NO LONGER IN CONTACT WITH THE ANATOMY AND CAN CREATE A DROP IN BLOOD PRESSURE. THIS CAN BE INDUCED UPON DEPLOYMENT OR RECAPTURED IN THE PRESENCE OF DENSE FOCAL CALCIFICATION. THE VALVE WAS RECAPTURED AND IT WAS REPORTED THAT DURING THE SECOND ATTEMPT, A NEW EVOLUT FX 29 MM PROSTHESIS WAS DEPLOYED VIA CONTRALATERAL ACCESS. FLUOROSCOPIC LOAD INSPECTION WAS PROVIDED AND VALVE APPEARS TO BE LOADED PROPERLY WITH NO MISLOAD. IT WAS REPORTED THAT IT WAS DECIDED TO IMPLANT THE SECOND VALVE. CHALLENGES IN DEPTH ADJUSTMENTS WERE OBSERVED AS THE SECOND VALVE DISLODGED FURTHER DOWN, NECESSITATING FOUR RECAPTURE ATTEMPTS BEFORE FINAL DEPLOYMENT AT APPROXIMATELY 4 MM. TRANSESOPHAGEAL ECHOCARDIOGRAPHY REVEALED MODERATE PARAVALVULAR LEAK (PVL), AND POST-DILATION WITH A NEW 24 MM NON-MEDTRONIC BALLOON (VACS III) WAS PERFORMED. ANGIOGRAPHY AND ECHOCARDIOGRAPHY SHOWED ONLY PARTIAL IMPROVEMENT OF THE PVL. ANGIOGRAPHIC EVIDENCE SHOWS PVL AFTER POST BALLOON DILATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: CONTINUATION OF D10: PRODUCT ID: D-EVOLUTFX-2329 (LOT: 0012281972); PRODUCT TYPE: 0195-HEART VALVES; PRODUCT ID EVFXPLUS-29 (K089923); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025; PRODUCT ID EVOLUTFX-29 (J306801); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2025; PRODUCT ID EDWARDS SAPIEN S3 (N/A); PRODUCT TYPE: HEART VALVES; PRODUCT ID: 24 X 60 VACS III BALLOON (N/A); PRODUCT TYPE: DILATION BALLOON. SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22328467&pc=NKM | 2135147-2025-03413 | 2019-05-01 04:00 | Death | ABBOTT MEDICAL | 2025-06-25 04:00 | NKM | MITRACLIP® SYSTEM CLIP DELIVERY SYSTEM | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P100009 |  | 1 | Event Description: THE ARTICLE "CONCOMITANT TRANSCATHETER EDGE-TO-EDGE REPAIR AND LEFT ATRIAL APPENDAGE OCCLUSION" WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE SINGLE CENTER STUDY, TO EVALUATE THE SAFETY AND FEASIBILITY OF CONCOMITANT MITRAL TRANSCATHETER EDGE TO EDGE REPAIR (M-TEER) AND LEFT ATRIAL APPENDAGE OCCLUSION (LAAO). DEVICES MENTIONED INCLUDE MITRACLIP, AMPLATZER AMULET, AMPLATZER TORQVUE 45X45 DELIVERY SYSTEM. THE ARTICLE CONCLUDED THAT CONCOMITANT LAAO WITH M-TEER IS AN ATTRACTIVE OPTION FOR PATIENTS WITH INDICATIONS FOR BOTH PROCEDURES TO SIMPLIFY RECOVERY TIME, REDUCE COST AND MITIGATE THE CUMULATIVE RISKS ASSOCIATED WITH VASCULAR ACCESS AND TRANSSEPTAL PUNCTURE WHEN PERFORMED ON SEPARATE OCCASIONS. (B)(6). THE TIME FRAME OF THE STUDY WAS MAY 2019 AND SEPTEMBER 2024. A TOTAL OF 15 PATIENTS WERE INCLUDED IN THIS STUDY, OF WHICH ALL RECEIVED AN ABBOTT DEVICE. AVERAGE WAS 80 YEARS AND MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDE PRIOR MYOCARDIAL INFARCTION, DIABETES, HYPERTENSION, DYSLIPIDEMIA, SMOKING, PRIOR STROKE, PERIPHERAL ARTERIAL DISEASE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, DEGENERATIVE AND FUNCTIONAL MITRAL REGURGITATION, ATRIAL FIBRILLATION, HEART FAILURE. Manufacturer Narrative: SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF MITRACLIP PROCEDURES WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PRIOR MYOCARDIAL INFARCTION, DIABETES, HYPERTENSION, DYSLIPIDEMIA, SMOKING, PRIOR STROKE, PERIPHERAL ARTERIAL DISEASE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, DEGENERATIVE AND FUNCTIONAL MITRAL REGURGITATION, ATRIAL FIBRILLATION, HEART FAILURE. COMPLICATIONS REPORTED INCLUDED COMPLICATIONS INCLUDE DEATH, MAJOR BLEEDING, PERICARDIAL EFFUSION, PERICARDIOCENTESIS, HOSPITALIZATION, RECURRENT MITRAL REGURGITATION; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. B2: DEATH DATE IS ESTIMATED B3: EVENT DATE IS ESTIMATED D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. ATTACHMENT: CONCOMITANT TRANSCATHETER EDGE-TO-EDGE REPAIR AND LEFT ATRIAL APPENDAGE OCCLUSION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22328752&pc=DSQ | 2916596-2025-04074 | 2025-06-14 04:00 | Death | THORATEC CORPORATION | 2025-06-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY ONE DAY AFTER LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION DUE TO MULTI ORGAN FAILURE. THE DEVICE FUNCTIONED AS EXPECTED. THE OUTCOME WAS NOT CONSIDERED DEVICE OR THERAPY RELATED. AN AUTOPSY WAS NOT PERFORMED. Manufacturer Narrative: SECTION A: PATIENT INFORMATION CANNOT BE PROVIDED DUE TO PERSONAL DATA PRIVACY LEGISLATION/POLICY. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED MULTI ORGAN FAILURE AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING MULTIPLE TYPES OF ORGAN FAILURE AND DYSFUNCTION (RESPIRATORY FAILURE, RIGHT HEART FAILURE, RENAL FAILURE, HEPATIC DYSFUNCTION) AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22329001&pc=DSQ | 2916596-2025-04189 | 2025-06-18 04:00 | Death | THORATEC CORPORATION | 2025-06-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Hemoptysis; Pneumonia; Septic Shock; Respiratory Failure; Multiple Organ Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO SEPTIC SHOCK. THE SOURCE OF INFECTION WAS BACTEREMIA STATUS POST-ACUTE CHOLECYSTITIS. THE PATIENT HAD E. COLI AND CLOSTRIDIUM BACTEREMIA, A CHOLE DRAIN PLACED ON (B)(6) 2025 BY INTERVENTIONAL RADIOLOGY (IR), RESPIRATORY FAILURE WITH INTUBATION AND THEN TRACE. COLI PNEUMONIA, HEMOPTYSIS, HEPATIC ABSCESS, PORTAL VEIN THROMBOSIS, THROMBOCYTOPENIA, ACUTE KIDNEY INJURY (AKI), TRANSAMINITIS, HYPERBILIRUBINEMIA, AND COAGULOPATHY. THE PATIENT HAD MULTISYSTEM ORGAN FAILURE, ELECTED DNR AND WITHDRAWAL OF CARE. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: ADDITIONAL CODES FOR H6: 1903 - HYPERBILIRUBINEMIA, 4431 - THROMBOCYTOPENIA, 4440 - THROMBOSIS/THROMBUS. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS IFU AND THE HEARTMATE 3 PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING INFECTION (LOCAL, DRIVELINE, PUMP POCKET), SEPSIS, MULTI ORGAN FAILURE (INCLUDING RESPIRATORY FAILURE AND RENAL AND HEPATIC DYSFUNCTION), THROMBOEMBOLISM, BLEEDING, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. THIS SECTION ALSO LISTS INFECTION AND THROMBOEMBOLISM AS POTENTIAL LATE POSTIMPLANT COMPLICATIONS. SEVERAL SECTIONS OF THE IFU AND PATIENT HANDBOOK PROVIDE CARE INSTRUCTIONS REGARDING HOW TO PREVENT INFECTION AS WELL AS SUGGESTED RESPONSES IN THE EVENT OF INFECTION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22329032&pc=DSQ | 3007042319-2025-01323 | 2025-06-19 04:00 | Death | HEARTWARE, INC. | 2025-06-25 04:00 | DSQ | HEARTWARE VENTRICULAR ASSIST SYSTEM - PUMP | Adverse Event Without Identified Device or Use Problem | Fall; Intracranial Hemorrhage | P100047 |  | 1 | Event Description: IT WAS REPORTED THAT A VENTRICULAR ASSIST DEVICE (VAD) PATIENT ACCIDENTALLY FELL AND SUFFERED A BRAIN HEMORRHAGE. THE PATIENT SUBSEQUENTLY DIED. Manufacturer Narrative: INVESTIGATION OF THIS EVENT IS PENDING AND A SUPPLEMENTAL REPORT WILL BE SENT UPON ITS COMPLETION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: A SUPPLEMENTAL REPORT IS BEING SUBMITTED FOR INVESTIGATION COMPLETION. PRODUCT EVENT SUMMARY: THE VENTRICULAR ASSIST DEVICE (VAD) (B)(6) WAS NOT RETURNED FOR EVALUATION. THE COMPLAINT IS ASSOCIATED WITH A CLINICAL ADVERSE EVENT. A REVIEW OF THE MANUFACTURING DOCUMENTATION WAS NOT PERFORMED AS THE REPORTED EVENT AND ANALYSIS ARE NOT RELATED TO A MANUFACTURING OR SERVICING ISSUE. OF NOTE, INFORMATION PROVIDED BY THE SITE INDICATED THAT THE PATIENT DIED AFTER ACCIDENTALLY FALLING AND SUFFERED A BRAIN HEMORRHAGE. BASED ON THE LIMITED INFORMATION AVAILABLE, THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. PER THE INSTRUCTIONS FOR USE, BLEEDING AND DEATH ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE IMPLANTATION OF A VAD. POSSIBLE CLINICAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT INCLUDE THE PATIENT¿S PRE-EXISTING HISTORY AND RELATED COMORBIDITIES, THE PROGRESSION OF THEIR UNDERLYING DISEASE, ISSUES RELATED TO THE THERAPEUTIC USE OF ANTICOAGULANT AND ANTIPLATELET MEDICATIONS AND THE PATIENT'S COMPLEX POST-OPERATIVE COURSE. THERE ARE POSSIBLE PATIENT, PHARMACOLOGICAL AND PROCEDURAL FACTORS THAT MAY HAVE CONTRIBUTED TO THIS EVENT. INVESTIGATION OF THIS EVENT IS COMPLETED AND THE FILE WILL BE CLOSED. IF NEW INFORMATION IS RECEIVED, THE FILE WILL BE RE-OPENED AND A SUPPLEMENTAL WILL BE SUBMITTED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22329087&pc=GKT | 3004548776-2025-00291 | 2025-05-14 04:00 | Death | FENWAL INTERNATIONAL INC. | 2025-06-25 04:00 | GKT | PLASMACELL XI DISPOSABLE SET WITH SPIKESMART | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | BK160028 |  | 1 | Event Description: THE CUSTOMER REPORTED THAT ON (B)(6) 2025, BIOLIFE PLASMA SERVICES RECEIVED NOTIFICATION OF A DONOR'S DEATH. A BIOLIFE CENTER EMPLOYEE REPORTED VIEWING A SOCIAL MEDIA POST ABOUT THE DEATH, ALTHOUGH THE EMPLOYEE WAS NOT DIRECTLY AFFILIATED WITH THE DONOR. BIOLIFE MEDICAL AFFAIRS COORDINATED A RECORDS REVIEW WITH REGULATORY AFFAIRS AND QUALITY OPERATIONS WHICH CONFIRMED THAT THE DONOR HAD LAST DONATED ON (B)(6) 2025, THAT A PUBLICLY AVAILABLE OBITUARY STATED THE DEATH OCCURRED AT 15:55 ON (B)(6) 2025 AND THAT THE FATALITY HAD NOT BEEN PREVIOUSLY REPORTED TO BIOLIFE. THE BIOLIFE CENTER PHYSICIAN CONTACTED THE MEDICAL EXAMINER'S OFFICE ON (B)(6) 2025 AND THE CAUSE OF DEATH WAS NOT PROVIDED. THIS 31-YEAR-OLD MAN HAD A TOTAL OF 183 DONATIONS BEGINNING ON (B)(6) 2013 WITH 64 DONATIONS IN THE PAST 24 MONTHS. THE DONOR UNDERWENT THE STANDARD APPLICANT DONOR PROCESS: ELECTRONIC DONOR HISTORY QUESTIONNAIRE (EDQ), PHYSICAL EXAMINATION, EDUCATION AND CONSENT, DONOR SCREENING (VITAL SIGNS, HEMATOCRIT AND PROTEIN DETERMINATION), DONOR PHLEBOTOMY WITH COLLECTION OF ALL LABORATORY TESTING SAMPLES INCLUDING THOSE FOR STS AND SERUM PROTEIN ELECTROPHORESIS (SPE). THE DONOR RECORD INDICATED NO MEDICATION ENTERED OR DISCONTINUED. THE DONOR'S PHYSICAL EXAM WAS LAST PERFORMED ON (B)(6) 2024 AND WAS NOTABLE FOR TATTOOS ONLY. THE DONOR LAST HAD A SPE DRAWN ON (B)(6) 2025, CENTER PHYSICIAN REVIEWED SPE RESULTS ON (B)(6) 2025, WITH AN ACCEPTABLE TOTAL PROTEIN OF 6.3 G/DL AND WITH ALL PROTEIN FRACTIONS NUMERICALLY AND VISUALLY ACCEPTABLE. THE DONOR'S PRIOR SPE DRAWN ON (B)(6) 2025, CENTER PHYSICIAN REVIEWED SPE RESULTS ON (B)(6) 2025, WAS NOTABLE LOW TOTAL PROTEIN OF 5.8G/DL. THE DONOR DEFERRAL HISTORY IS NOTABLE FOR, RED CELL LOSS WITHOUT DONOR REACTION ON (B)(6) 2025 AND OUT OF RANGE INITIAL PULSE READINGS WITH SECOND READING IN RANGE AND DONOR ACCEPTABLE TO DONATE ON (B)(6) 2025 AND DEFERRED FOR OUT-OF-RANGE PULSE ON (B)(6) 2025. ON (B)(6) 2025, THE DONOR PRESENTED FOR PLASMAPHERESIS WITH APPROPRIATE DONOR EDUCATION AND CONSENT DOCUMENTED, DONOR EDQ WITH NO ABERRANT RESPONSES COMPLETED AND ALL DONOR SCREENING CRITERIA WITHIN ACCEPTABLE RANGES (BP 123/75 MMHG; PULSE 98 BPM; TEMPERATURE 98F; WEIGHT 284LBS; HCT 48%; TOTAL PROTEIN 7.2 G/DL). A TOTAL OF 980ML OF PLASMA WAS COLLECTED FROM THE DONOR'S RIGHT ARM WITH NO RED CELL LOSS AND WITH PROCEDURAL SALINE GIVEN AND WITH NO DONOR REACTIONS NOTED. THE BIOLIFE STAFF MEMBER WHO PERFORMED THE PLASMAPHERESIS PROCEDURE WAS QUESTIONED ON (B)(6) 2025 AND REPORTED NO NOTABLE OCCURRENCES RELATED TO THE COLLECTION. ON (B)(6) 2025, THE CENTER STAFF REMOVED THE AURORA XI INSTRUMENT FROM SERVICE THAT WAS USED IN THE DONOR'S LAST PLASMA COLLECTION ON (B)(6) 2025. THE PREVENTATIVE MAINTENANCE TPDS WAS PERFORMED ON (B)(6) 2025 AND FOUND THE INSTRUMENT WAS WORKING AS INTENDED. THE FOLLOWING INSTRUMENT MAINTENANCE DOCUMENTS WERE REVIEWED FROM (B)(6) 2023 TO (B)(6) 2025. LAST ANNUAL AND REQUIRED PMS, PLASMAPHERESIS INSTRUMENT TROUBLESHOOTING LOGS, THE PREVENTATIVE AND CORRECTIVE MAINTENANCE RECORDS AND THE INSTRUMENT WEIGHT SCALE VERIFICATIONS WERE REVIEWED. THE DEVICE DID NOT INDICATE ANY CONCERNS. THE DONOR DID NOT DISCLOSE ANY RELEVANT MEDICAL HISTORY, WAS NOT TAKING ANY MEDICATIONS AND MET ALL DONOR QUALIFICATIONS THROUGHOUT THE TIME OF THEIR PARTICIPATION IN THE PLASMAPHERESIS PROGRAM. WE ARE REPORTING THIS INCIDENT CONSERVATIVELY GIVEN THAT THE DONOR DIED THE DAY AFTER AN UNCOMPLICATED DONATION AND THE CAUSE OF DEATH IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22329436&pc=OZD | 1220648-2025-29783 | 2025-06-09 04:00 | Death | ABIOMED, INC. | 2025-06-25 04:00 | OZD | IMPELLA 5.5 | Adverse Event Without Identified Device or Use Problem | Hemorrhage/Bleeding; Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS ON IMPELLA 5.5 SUPPORT AND DURING SUPPORT, THE PATIENT HAD ACCESS SITE BLEEDING. TWO UNITS OF PACKED RED BLOOD CELLS WERE GIVEN TO THE PATIENT IN THE OPERATING ROOM. THE PLAN WAS TO MONITOR HEMOGLOBIN, HEMATOCRIT AND PLATELETS AND FOR SIGNS OF BLEEDING. HEMOGLOBLIN WAS TO BE REPEATED POST TRANSFUSION. ANGIOMAX WAS HELD UNTIL THE JP DRAINAGE SLOWS. SUPPORT WAS CONTINUED. ADDITIONALLY, AN AORTIC THROMBUS WAS NOTED AT THE LEVEL OF THE DIAPHRAGM. CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER. THEREFORE, THE INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿ SECTION: WARNINGS & CAUTIONS: WARNINGS: SECTION: PRE-SUPPORT EVALUATION SECTION: AXILLARY INSERTION OF THE IMPELLA 5.5 CATHETER SECTION: DIRECT AORTIC INSERTION SECTION: USE OF IMPELLA WITH TRANSCATHETER AORTIC VALVES "IN PATIENTS WITH TRANSCATHETER AORTIC VALVES POSITION THE IMPELLA SYSTEM CAREFULLY TO AVOID INTERACTION WITH THE TRANSCATHETER AORTIC VALVE PROSTHESIS. UNINTENTIONAL INTERACTION OF THE IMPELLA MOTOR HOUSING WITH THE TAVR DEVICE MAY RESULT IN DESTRUCTION OF THE IMPELLER BLADES. THIS CAN LEAD TO SYSTEMIC EMBOLIZATION, SERIOUS INJURY, OR DEATH.". IN THIS SITUATION, AVOID REPOSITIONING WHILE THE DEVICE IS RUNNING; TURN THE DEVICE TO P0 DURING REPOSITIONING OR ANY MOVEMENT THAT COULD BRING THE OUTLET WINDOWS INTO PROXIMITY TO THE VALVE STENT STRUCTURES. IF THERE IS LOW FLOW OBSERVED IN A PATIENT IMPLANTED WITH A TRANSCATHETER AORTIC VALVE PROSTHESIS, CONSIDER DAMAGE OF THE IMPELLER AND REPLACE THE IMPELLA AS SOON AS POSSIBLE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22329500&pc=KNO | 1525712-2025-00009 | 2023-04-16 04:00 | Death | UNKNOWN | 2025-06-25 04:00 | KNO | UNKNOWN | Appropriate Term/Code Not Available | Strangulation |  |  | 1 | Event Description: A LEGAL DOCUMENT WAS RECEIVED ALLEGING AN INCIDENT OCCURRED ON (B)(6) 2023. THE DOCUMENT STATES THE END-USER WAS SITTING IN HER WHEELCHAIR IN THE LIVING SPACE OF HER ROOM AT THE FACILITY. SHE WAS SLIGHTLY RECLINED IN THE WHEELCHAIR AND HAD A LAP BELT SECURED BY VELCRO POSITIONED ACROSS HER LAP. THE END-USER WAS LEFT ALONE IN HER ROOM. OVER THE NEXT TWO HOURS, THE END-USER BEGAN TO SLOWLY SLIDE FROM THE SEATED POSITION IN HER WHEELCHAIR TO THE GROUND. THE LAP BELT THAT WAS INITIALLY ATTACHED ACROSS HER HIPS, BECAME TIGHTER AROUND HER ABDOMEN, CHEST AND FINALLY HER NECK. SHE WAS STRANGLED AND DIED OF ASPHYXIATION. Manufacturer Narrative: THE LEGAL DOCUMENT RECEIVED BY INVACARE INDICATED THAT INVACARE IS THE MANUFACTURER OF THE LAP BELT. INVACARE WAS NOT LISTED AT THE MANUFACTURER OF THE WHEELCHAIR THE SUBJECT LAP BELT WAS INSTALLED ON. THE MODEL NUMBER OF THE LAP BELT WAS NOT PROVIDED. FOR FILING PURPOSES, TAYLOR STREET MANUFACTURING (ELYRIA, OHIO) CFN NUMBER IS BEING UTILIZED FOR THIS MEDWATCH. THE LEGAL DOCUMENT ADVISED A CAMERA WAS INSTALLED IN THE END-USER¿S ROOM WHICH PROVIDED A LIVE AUDIO/VIDEO FEED TO THE MONITORING STATION AT THE FACILITY THE END-USER LIVED AT. THERE HAD BEEN NO FACILITY EMPLOYEE AT THE MONITORING STATION DURING THE TIME OF THE INCIDENT. FURTHERMORE, IT WAS STATED THE FACILITY FAILED TO CHECK ON THE END-USER EVERY HOUR AS REQUIRED. THERE WAS NO INDICATION OF ANY MALFUNCTION OR DEFICIENCY OF THE LAP BELT. ATTEMPTS ARE ONGOING TO OBTAIN ADDITIONAL INFORMATION REGARDING THIS INCIDENT WHICH INCLUDES MODEL NUMBER OF THE LAP BELT AND MODEL OF THE WHEELCHAIR THE LAP BELT WAS INSTALLED ON. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, A FOLLOW-UP MEDWATCH WILL BE FILED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22329669&pc=LXH | 2027148-2025-00002 | 2025-06-03 04:00 | Death | KINAMED, INC. | 2025-06-25 04:00 | LXH | SUPERCABLE ISO-ELASTIC CERCLAGE SYSTEM | Deformation Due to Compressive Stress | Cardiac Arrest | 510(K)EXEMPT |  | 1 | Event Description: REPORTER'S OWN WORDS: STRAIGHT TENSIONER [35-800-2020] WAS BROKEN BEFORE THE START OF THE CASE. THE DEPLOYMENT POST WAS BENT SO UNUSABLE. WE THEN HAD TO USE THE ANGLED TENSIONER WHICH IS DIFFICULT TO USE AND NOT IDEAL. IN THE MIDDLE OF THE CASE IT [ASSUMED TO REFER TO 35-800-7000] FELL APART. WE THEN HAD TO PUT IT BACK TOGETHER ENOUGH SO WE COULD USE IT TO COMPLETE THE CASE. IN TOTAL I [REPORTER] FEEL IT ADDED AN HOUR AT LEAST TO A VERY DIFFICULT PROCEDURE. PATIENT WAS ELDERLY AND UNWELL SO TIME IS IMPORTANT TO GET HIM OUT OF THEATRE AS QUICKLY AS POSSIBLE. PATIENT HAS A PERI PROSTHETIC FRACTURE THAT NEEDED TO BE REPAIRED. AS THE SURGEON WAS APPLYING STAPLES TO CLOSE THE WOUND AT THE END OF THE CASE, THE PATIENT'S CARDIAC STATUS BECAME UNSTABLE. THE TOTAL OPERATION WAS LONG AND COMPLEX, TOOK ABOUT 5 HOURS AND INVOLVED REMOVAL OF A CEMENTED STEM AND IMPLANTATION OF A REVISION STEM. SUBSEQUENT REPORT FROM DISTRIBUTOR: ELDERLY PATIENT UNDERWENT FEMORAL PERI-PROSTHETIC FRACTURE REPAIR PROCEDURE USING CABLE FIXATION, PLATE, AND HIP STEM. DURING THE PROCEDURE THERE WAS A DELAY OF APPROXIMATELY 60 MINUTES DUE TO INSTRUMENT MALFUNCTION. THE REPAIR WAS COMPLETED AND DURING CLOSURE, THE PATIENT WENT INTO CARDIAC ARREST AND PASSED AWAY SOME TIME AFTER THE PROCEDURE. TO DATE, OVER 240,000 SUPERCABLE CERCLAGE IMPLANTS HAVE BEEN USED IN APPROXIMATELY 48,000 SURGERIES SINCE INTRODUCTION IN 2004. THIS IS THE FIRST REPORT OF THIS TYPE RECEIVED FROM ALL POST-MARKET SURVEILLANCE ACTIVITIES INVOLVING THE SUPERCABLE SYSTEM. THE INCIDENCE RATE FOR THIS TYPE OF EVENT IS 0.002% (1 OUT OF 48,000). A LITERATURE SEARCH (CITATIONS REFERENCED BELOW) WAS CONDUCTED TO DETERMINE THE MORTALITY RATE DURING AND FOLLOWING REVISION TOTAL HIP ARTHROPLASTY (THA) PERFORMED FOR PERIPROSTHETIC FRACTURE. THE MORTALITY RATE REPORTED IN THE LITERATURE IS SUBSTANTIALLY HIGHER THAN THE RATE ASSOCIATED WITH THE SUPERCABLE SYSTEM. BERSTOCK ET AL [1] REVIEWED 32 STUDIES INVOLVING OVER 1 MILLION PATIENTS AND REPORTED A MORTALITY RATE OF 0.30% DURING THE FIRST 30 DAYS FOLLOWING HIP REPLACEMENT. THEY FOUND THAT THE MOST COMMON RISK FACTORS FOR EARLY MORTALITY WERE INCREASING AGE, MALE GENDER, AND CO-MORBID CONDITIONS SUCH AS CARDIOVASCULAR DISEASE. THEY REPORTED THAT THE 30-DAY MORTALITY FOR THE > 80 AGE GROUP WAS 1.43%. HAUGHOM ET AL [2] REVIEWED 523 PERIPROSTHETIC HIP FRACTURE CASES AND FOUND THAT 14.5% OF THE PATIENTS EXPERIENCED A SEVERE ADVERSE EVENT, AND THAT THE 30-DAY MORTALITY RATE WAS 2.7%. THEY REPORTED THAT AGE GREATER THAN 85 WAS AN INDEPENDENT RISK FACTOR FOR MORTALITY FOLLOWING PERIPROSTHETIC FRACTURE. KARLIDAG ET AL [3] REVIEWED 13,203 REVISION TOTAL HIP ARTHROPLASTY CASES AND REPORTED A 0.53% MORTALITY RATE (70 PATIENTS) DURING HOSPITALIZATION. KHAN ET AL [4] REVIEWED 6,131 REVISIONS PERFORMED FOR PERIPROSTHETIC FRACTURE AND REPORTED THAT THE PROBABILITY OF DEATH IN MALES OVER 75 YEARS WITH ASA CLASS >= 3 (ASA CLASS IS A MEASURE OF CO-MORBIDITY) WAS APPROXIMATELY 9% AT 90 DAYS. LAMB ET AL [5] REPORTED AN INPATIENT MORTALITY RATE OF 2.4% BASED ON DATA FROM 10 STUDIES THAT INCLUDED 1293 PATIENTS WITH MEAN AGE OF 76.3 YEARS. LAUGHLIN ET AL [6] REVIEWED 596 REVISION THA CASES AND REPORTED A 30-DAY MORTALITY RATE OF 0.34%. THEY FOUND THAT PATIENTS WITH A PERIPROSTHETIC FRACTURE HAD A HIGHER RISK OF DEATH COMPARED TO ALL OTHER INDICATIONS FOR REVISION THA, AND THAT INCREASED AGE AT THE TIME OF REVISION THA AND MALE GENDER WERE ASSOCIATED WITH A HIGHER RISK OF DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22329672&pc=LXH | 2027148-2025-00001 | 2025-06-03 04:00 | Death | KINAMED, INC. | 2025-06-25 04:00 | LXH | SUPERCABLE ISO-ELASTIC CERCLAGE SYSTEM | Mechanical Problem | Cardiac Arrest | 510(K)EXEMPT |  | 1 | Event Description: REPORTERS OWN WORDS: STRAIGHT TENSIONER [35-800-2020] WAS BROKEN BEFORE THE START OF THE CASE. THE DEPLOYMENT POST WAS BENT SO UNUSABLE. WE THEN HAD TO USE THE ANGLED TENSIONER WHICH IS DIFFICULT TO USE AND NOT IDEAL. IN THE MIDDLE OF THE CASE IT [ASSUMED TO REFER TO 35-800-7000] FELL APART. WE THEN HAD TO PUT IT BACK TOGETHER ENOUGH SO WE COULD USE IT TO COMPLETE THE CASE. IN TOTAL I [REPORTER] FEEL IT ADDED AN HOUR AT LEAST TO A VERY DIFFICULT PROCEDURE. PATIENT WAS ELDERLY AND UNWELL SO TIME IS IMPORTANT TO GET HIM OUT OF THEATRE AS QUICKLY AS POSSIBLE. PATIENT HAS A PERI PROSTHETIC FRACTURE THAT NEEDED TO BE REPAIRED. AS THE SURGEON WAS APPLYING STAPLES TO CLOSE THE WOUND AT THE END OF THE CASE, THE PATIENT'S CARDIAC STATUS BECAME UNSTABLE. THE TOTAL OPERATION WAS LONG AND COMPLEX, TOOK ABOUT 5 HOURS AND INVOLVED REMOVAL OF A CEMENTED STEM AND IMPLANTATION OF A REVISION STEM. SUBSEQUENT REPORT FROM DISTRIBUTOR: ELDERLY PATIENT UNDERWENT FEMORAL PERI-PROSTHETIC FRACTURE REPAIR PROCEDURE USING CABLE FIXATION, PLATE, AND HIP STEM. DURING THE PROCEDURE THERE WAS A DELAY OF APPROXIMATELY 60 MINUTES DUE TO INSTRUMENT MALFUNCTION. THE REPAIR WAS COMPLETED AND DURING CLOSURE, THE PATIENT WENT INTO CARDIAC ARREST AND PASSED AWAY SOME TIME AFTER THE PROCEDURE. TO DATE, OVER 240,000 SUPERCABLE CERCLAGE IMPLANTS HAVE BEEN USED IN APPROXIMATELY 48,000 SURGERIES SINCE INTRODUCTION IN 2004. THIS IS THE FIRST REPORT OF THIS TYPE RECEIVED FROM ALL POST-MARKET SURVEILLANCE ACTIVITIES INVOLVING THE SUPERCABLE SYSTEM. THE INCIDENCE RATE FOR THIS TYPE OF EVENT IS 0.002% (1 OUT OF 48,000). A LITERATURE SEARCH (CITATIONS REFERENCED BELOW) WAS CONDUCTED TO DETERMINE THE MORTALITY RATE DURING AND FOLLOWING REVISION TOTAL HIP ARTHROPLASTY (THA) PERFORMED FOR PERIPROSTHETIC FRACTURE. THE MORTALITY RATE REPORTED IN THE LITERATURE IS SUBSTANTIALLY HIGHER THAN THE RATE ASSOCIATED WITH THE SUPERCABLE SYSTEM. BERSTOCK ET AL [1] REVIEWED 32 STUDIES INVOLVING OVER 1 MILLION PATIENTS AND REPORTED A MORTALITY RATE OF 0.30% DURING THE FIRST 30 DAYS FOLLOWING HIP REPLACEMENT. THEY FOUND THAT THE MOST COMMON RISK FACTORS FOR EARLY MORTALITY WERE INCREASING AGE, MALE GENDER, AND CO-MORBID CONDITIONS SUCH AS CARDIOVASCULAR DISEASE. THEY REPORTED THAT THE 30-DAY MORTALITY FOR THE > 80 AGE GROUP WAS 1.43%. HAUGHOM ET AL [2] REVIEWED 523 PERIPROSTHETIC HIP FRACTURE CASES AND FOUND THAT 14.5% OF THE PATIENTS EXPERIENCED A SEVERE ADVERSE EVENT, AND THAT THE 30-DAY MORTALITY RATE WAS 2.7%. THEY REPORTED THAT AGE GREATER THAN 85 WAS AN INDEPENDENT RISK FACTOR FOR MORTALITY FOLLOWING PERIPROSTHETIC FRACTURE. KARLIDAG ET AL [3] REVIEWED 13,203 REVISION TOTAL HIP ARTHROPLASTY CASES AND REPORTED A 0.53% MORTALITY RATE (70 PATIENTS) DURING HOSPITALIZATION. KHAN ET AL [4] REVIEWED 6,131 REVISIONS PERFORMED FOR PERIPROSTHETIC FRACTURE AND REPORTED THAT THE PROBABILITY OF DEATH IN MALES OVER 75 YEARS WITH ASA CLASS >= 3 (ASA CLASS IS A MEASURE OF CO-MORBIDITY) WAS APPROXIMATELY 9% AT 90 DAYS. LAMB ET AL [5] REPORTED AN INPATIENT MORTALITY RATE OF 2.4% BASED ON DATA FROM 10 STUDIES THAT INCLUDED 1293 PATIENTS WITH MEAN AGE OF 76.3 YEARS. LAUGHLIN ET AL [6] REVIEWED 596 REVISION THA CASES AND REPORTED A 30-DAY MORTALITY RATE OF 0.34%. THEY FOUND THAT PATIENTS WITH A PERIPROSTHETIC FRACTURE HAD A HIGHER RISK OF DEATH COMPARED TO ALL OTHER INDICATIONS FOR REVISION THA, AND THAT INCREASED AGE AT THE TIME OF REVISION THA AND MALE GENDER WERE ASSOCIATED WITH A HIGHER RISK OF DEATH.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22329688&pc=DSQ | 2916596-2025-03914 | 2025-06-04 04:00 | Death | THORATEC CORPORATION | 2025-06-25 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM, SERIAL NUMBER (B)(6) AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WILL NOT BE EXPLANTED AND WILL NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. THE CURRENT REVISION OF THE IFU CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, "INTRODUCTION," LISTS POTENTIAL ADVERSE EVENTS WHICH MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LVAS, INCLUDING DEATH. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22329704&pc=DSQ | 2916596-2025-03871 | 2025-04-22 04:00 | Death | THORATEC CORPORATION | 2025-06-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Abdominal Pain; Anemia; Dyspnea; High Blood Pressure/ Hypertension; Muscle Weakness; Renal Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT ON THE PATIENT'S FOLLOW UP APPOINTMENT ON AROUND (B)(6) 2025, THEY HAD COMPLAINTS OF ABDOMINAL PAIN AND DISTENSION. AN ABDOMEN/PELVIS COMPUTED TOMOGRAPHY (CT) WAS DONE AS THE PATIENT HAD NOT HAD ONE IN A WHILE. IT WAS NEGATIVE FOR ANY ACUTE ISSUES, BUT DID SHOW MODERATE FECES IN COLON, ALTHOUGH THE PATIENT HAD DENIED ANY CONSTIPATION AND HAD REGULAR BOWEL MOVEMENTS (BMS). THE PATIENT CALLED AND SCHEDULED A CLINIC APPOINTMENT ABOUT 5-6 WEEKS LATER BECAUSE THEY HAD BEEN FEELING WORSE AND SAID THEIR RIGHT UPPER QUADRANT (RUQ) OF THE ABDOMEN WAS IN A LOT MORE PAIN. THE PATIENT SAID THEY HAVE BEEN EATING ALL RIGHT. ALTHOUGH THE PATIENT HAD SOME NAUSEA AT TIMES, THEY WERE NOT VOMITING. THE PATIENT HAD THE FEELING OF FOOD GETTING CAUGHT AS THEY DID IN (B)(6), AND REGLAN WAS TRIED, BUT THEY REPORTED IT DID NOT HELP. THE PATIENT HAD BEEN FEELING WEAKER AND HAD TROUBLE WALKING INTO THE CLINIC AND ALSO HAD SHORTNESS OF BREATH (SOB). THE PATIENT SAID THEY WERE MAKING URINE AND HAD REGULAR BMS. THEIR WEIGHT WAS ABOUT THE SAME FROM THEIR LAST CLINIC VISIT. THE PATIENT WAS ONLY ON ANTI-HYPERTENSIVE MEDS AND THEIR BLOOD PRESSURE (BP) IN THE CLINIC WAS 120/0 AND TENDED TO RUN HIGH. THE PATIENT HAD MANY TIMES IN THE PAST BEEN VERY HYPERTENSIVE NEAR 150S WHEN THEY HAD BEEN SEEN. THEY DENIED HAVING INCREASED DIZZINESS. THE PATIENT HAD NOT BEEN ON DIURETICS IN A WHILE DUE TO RENAL DYSFUNCTION AND HAD BEEN MOSTLY EUVOLEMIC IN THE PAST. THE PATIENT UNDERWENT EXTENSIVE TESTING TO DETERMINE THE CAUSE OF THEIR CHRONIC DISCOMFORT. THE PATIENT WAS FOUND TO BE ANEMIC WITH HEMOGLOBIN LEVEL OF 7.2 ON ADMISSION AND WAS GIVEN A UNIT OF BLOOD AND STARTED ON RETACRIT 3 TIMES WEEKLY AND THE HEMOGLOBIN AND HEMATOCRIT STABILIZED. THE PATIENT HAD CHRONIC KIDNEY DISEASE. THEIR BLOOD UREA NITROGEN (BUN) AND CREATININE HAD WORSENED PRIOR TO THIS ADMISSION AND CONTINUED TO WORSEN DURING THEIR ADMISSION. NEPHROLOGY WAS CONSULTED AND THE POSSIBILITY OF NEEDING DIALYSIS WAS DISCUSSED. THE PATIENT DID NOT WANT TO HAVE DIALYSIS, EVEN IF IT WAS DEEMED NECESSARY. THE PATIENT WAS FOUND TO HAVE NON-OBSTRUCTING RENAL STONES, AND UROLOGY WAS CONSULTED AND RECOMMENDED OUTPATIENT URETEROSCOPY (URS)/LASER/STENT. THE PATIENT WAS ALSO TREATED FOR KLEBSIELLA PNEUMONIAE CYSTITIS WITH ROCEPHIN. THEY UNDERWENT A RIGHT HEART CATHETERIZATION (RHC) WITH SPEED OPTIMIZATION AS WELL. DESPITE ALL OF THESE FINDINGS BEING ADDRESSED THE PATIENT CONTINUED TO COMPLAIN THAT THEY DID NOT FEEL WELL. THE PATIENT HAS HAD VARIOUS VAGUE COMPLAINTS SINCE THEIR LEFT VENTRICULAR ASSIST DEVICE (LVAD) IMPLANTATION. THEY HAVE STATED ON A REGULAR BASIS THAT THEY HAD NOT FELT RIGHT SINCE IMPLANTATION AND THAT THEY REGRETTED GETTING THE LVAD. FROM A MEDICAL OPINION, THE PATIENT HAD BEEN SEVERELY DEPRESSED FOR A LONG TIME, AND THIS WAS DISCUSSED WITH THE PATIENT MANY TIMES AND ANTIDEPRESSANTS WERE RECOMMENDED WHICH THEY HAD REFUSED. DURING THE PATIENT'S HOSPITALIZATION THEY WERE SEEN BY PSYCHIATRY AND STARTED ON ZOLOFT, BUT THEY TOOK IT ONLY FOR A SHORT TIME AND THEN BEGAN REFUSING IT. THE PATIENT STATED MULTIPLE TIMES DURING THIS HOSPITALIZATION THAT THEY WANTED TO BE ABLE TO "GO HOME AND DIE" SINCE THEY THOUGHT THEY WOULD NEVER START TO FEEL BETTER. PALLIATIVE CARE WAS CONSULTED, AND ARRANGEMENTS WERE MADE FOR THE PATIENT TO BE DISCHARGED HOME WITH HOME HOSPICE. THE PATIENT ELECTED TO HAVE THEIR PUMP TURNED OFF ON (B)(6) 2025 AND PASSED AWAY PEACEFULLY SHORTLY AFTER DUE TO SUPPORT WITHDRAWAL. THE DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED AND THE DEVICE OPERATED AS EXPECTED. NO DEVICE WOULD BE RETURNED. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. HEARTMATE 3 LVAS, SERIAL NUMBER (B)(6), WAS NOT RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS, INCLUDING DEATH, RENAL DYSFUNCTION, AND HYPERTENSION. SECTION 6 OF THE IFU (UNDER ¿ANTICOAGULATION¿) PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22329996&pc=DSQ | 2916596-2025-03799 | 2025-06-02 04:00 | Death | THORATEC CORPORATION | 2025-06-25 04:00 | DSQ | THORATEC® HEARTMATE®, LVAS IMPLANT KIT (WITH SEALED GRAFTS) | Infusion or Flow Problem; Adverse Event Without Identified Device or Use Problem | Anemia; Muscle Weakness | P060040 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WAS ADMITTED ON (B)(6) 2025 WITH ANEMIA, WEAKNESS AND LOW FLOW ALARMS. THE PATIENT WAS GIVEN 2 UNITS OF PACKED RED BLOOD CELLS (PRBC) AT OUTSIDE HOSPITAL FOR HEMOGLOBIN (HGB) OF 7.8 AND HGB 11 ON (B)(6) 2025. HOLDING COUMADIN, INTERNATIONAL NORMALIZED RATIO (INR) WAS 2.7. THE PATIENT'S INR GOAL WAS 1.8-2.2 DUE TO PREVIOUS GASTROINTESTINAL (GI) BLEEDING EVENTS AND OFF ASPIRIN PREVIOUSLY. THERE WAS NO GI CONSULT DUE TO PATIENT NOT WANTING ANY SCOPES DONE. PALLIATIVE CARE CONSULT WAS PERFORMED. THE LOG FILE FROM (B)(6) 2025 UNTIL (B)(6) 2025 CAPTURED SOME LOW FLOW ALARMS. THE CALCULATED FLOW APPEARED TO HAVE FLUCTUATED BELOW THE ALARM THRESHOLD OF 2.5 LPM DURING THE EVENTS. THE LOG FILE ALSO CONTAINED LOW FLOW HAZARD EVENTS WHEN THE PULSATILITY INDEX (PI) WAS DYNAMIC AND ELEVATED. THE LOW FLOW READINGS APPEARED TO BE A PATIENT RELATED ISSUE. THE EQUIPMENT WAS OPERATING AS INTENDED. THE PATIENT DECLINED GI EVALUATION. THE SOURCE OF BLEEDING WAS UNKNOWN. THE BLEEDING DID NOT RESOLVE. BLOOD PRESSURE WAS OPTIMIZED, THE PATIENT WAS HYDRATED, HEMOGLOBIN STABILIZED, AND THE PATIENT CONTINUED TO HAVE LOW FLOW ALARMS WITH STANDING AND ACTIVITY. THE PATIENT WAS DISCHARGED HOME WITH HOSPICE CARE.Event Description: THE PATIENT PASSED AWAY UNDER HOSPICE CARE ON (B)(6) 2025. THE PUMP WOULD NOT BE RETRIEVED. Manufacturer Narrative: PREVIOUS GI BLEEDING CAPTURED UNDER MFR. REPORT # 2916596-2025-03951. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE II LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER (B)(6) AND THE REPORTED BLEEDING AND SUBSEQUENT PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. REVIEW OF THE SUBMITTED LOG FILE CONFIRMED THE REPORTED LOW FLOW ALARMS; HOWEVER, A SPECIFIC CAUSE FOR THE ALARMS COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE LOG FILE CONTAINED EVENTS DATA FROM (B)(6) 2025 THROUGH (B)(6) 2025. INTERMITTENT LOW FLOW HAZARD ALARMS WERE CAPTURED THROUGHOUT THE LOG FILE WHEN THE ESTIMATED FLOW DECREASED BELOW THE LOW FLOW ALARM THRESHOLD. NO OTHER NOTABLE EVENTS OR ALARMS WERE CAPTURED, AND THE SYSTEM APPEARED TO FUNCTION AS INTENDED. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE II LVAS IFU AND THE HEARTMATE II PATIENT HANDBOOK ARE CURRENTLY AVAILABLE. SECTION 1, ¿INTRODUCTION¿, LISTS BLEEDING AND DEATH AS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE II LVAS. SECTION 1 OF THE IFU ALSO ADDRESSES ALL PUMP PARAMETERS, INCLUDING PUMP FLOW. SECTION 4 OF THE IFU, ¿HEARTMATE TOUCH COMMUNICATION SYSTEM¿, DESCRIBES THE PUMP FLOW DISPLAY AND THE HAZARD ALARMS. IT ALSO EXPLAINS THAT CHANGES IN PATIENT CONDITION CAN RESULT IN LOW FLOW. PER DESIGN, WHEN THE ESTIMATED FLOW VALUE IS CALCULATED AT LESS THAN 2.5 LITERS PER MINUTE (LPM), A LOW FLOW STATUS IS POSTED TO THE LOG FILE. IF THE FLOW REMAINS BELOW 2.5 LPM FOR 10 SECONDS, A LOW FLOW HAZARD ALARM IS TRIGGERED. SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿ (UNDER ¿ANTICOAGULATION¿), OUTLINES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO RANGE, AS WELL AS THE SUGGESTED ANTICOAGULATION MODIFICATIONS IN THE EVENT THAT THERE IS A RISK OF BLEEDING. SECTION 7 OF THE IFU, ¿ALARMS AND TROUBLESHOOTING¿, AND SECTION 5 OF THE PATIENT HANDBOOK, ¿ALARMS AND TROUBLESHOOTING¿, OUTLINES ALL SYSTEM ALARM CONDITIONS AS WELL AS HOW TO RESPOND TO EACH ALARM CONDITION. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22330007&pc=DSQ | 2916596-2025-03982 | 2025-06-03 04:00 | Death | THORATEC CORPORATION | 2025-06-25 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Insufficient Information | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE CAUSE WAS NOT PROVIDED. IT WAS REPORTED NO, THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED; THE DEVICE FUNCTIONED AS EXPECTED. THE DEVICE WAS NOT EXPLANTED AND WOULD NOT BE RETURNED FOR EVALUATION. THE ACCOUNT HAD DECLINED TO PROVIDE PATIENT OUTCOME INFORMATION FOR THIS EVENT, DUE TO PATIENT PRIVACY LAWS. BASED ON A REVIEW OF AVAILABLE PATIENT¿S RECORD AND THE REPORTED PATIENT OUTCOME, THERE WERE NO DEVICE ISSUES AT THE TIME OF THE PATIENT OUTCOME. THE DEVICE OPERATED AS EXPECTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) SERIAL NUMBER: (B)(6), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY DETERMINED THROUGH THIS EVALUATION. THE DEVICE WAS NOT EXPLANTED FOR EVALUATION. DUE TO HOSPITAL PRIVACY POLICY, THE ACCOUNT WILL NOT PROVIDE ADDITIONAL INFORMATION REGARDING THE EVENT. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LVAS INSTRUCTIONS FOR USE (IFU) IS ARE CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION¿, LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22330149&pc=DSQ | 2916596-2025-03852 | 2025-06-04 04:00 | Death | THORATEC CORPORATION | 2025-06-25 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Abscess; Unspecified Infection; Hematuria; Hemorrhagic Stroke; Thrombosis/Thrombus; Skin Infection; Drug Resistant Bacterial Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THE RESEARCH ARTICLE TITLED, ¿HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE SYSTEM IN PATIENTS WITH GLUCOSE-6-PHOSPHATE DEHYDROGENASE DEFICIENCY¿, GLUCOSE-6-PHOSPHATE DEHYDROGENASE (G6PD) DEFICIENCY IS A GENETIC ENZYMATIC DISORDER THAT AFFECTS MILLIONS OF PEOPLE WORLDWIDE AND CHARACTERIZED BY HEMOLYSIS UNDER OXIDATIVE STRESS. LEFT VENTRICULAR ASSIST DEVICES (LVADS) HAVE SUBSTANTIALLY ENHANCED SURVIVAL AND QUALITY OF LIFE FOR INDIVIDUALS WITH ADVANCED HEART FAILURE. HOWEVER, THEIR USE IS ASSOCIATED WITH THE RISK OF HEMOLYSIS, THROMBOSIS, AND EMBOLIC EVENTS. THESE RISKS MAY BE HEIGHTENED IN PATIENTS WITH GLUCOSE-6 PHOSPHATE DEHYDROGENASE (G6PD) DEFICIENCY. PATIENT 2: THERE WAS A NOTICEABLE HEMATURIA IN THE ICU. THEN, THE PATIENT DEVELOPED DRIVELINE INFECTION, AND HIS CULTURE REVEALED METHICILLIN RESISTANT STAPHYLOCOCCAL (MRSA) INFECTION. SUBSEQUENTLY, HIS CONDITION WAS COMPLICATED WITH ABDOMINAL ABSCESS AND THEN CEREBRAL HEMORRHAGE WHICH WAS THE DIRECT CAUSE OF DEATH (DIED AFTER 32 MONTHS OF FOLLOW-UP). IT IS ESTIMATED THAT ABOUT 400 MILLION PEOPLE ARE AFFECTED BY G6PD DEFICIENCY WORLDWIDE.14 THIS GENETIC CONDITION IS INHERITED AS AN X-LINKED RECESSIVE STATE AND PRESENTS ONE OF THE MOST FREQUENTLY ENCOUNTERED RED-CELL ENZYMOPATHIES. THE MAJOR INTEREST IN THE G6PD-DEFICIENT STATE RESULTS FROM THE ASSOCIATED HEMOLYTIC ANEMIA RESULTING FROM OXIDATIVE STRESS.15,16 IT IS WORTH MENTIONING THAT HEMOLYSIS-INDUCED THROMBOTIC EVENTS IN G6PD PATIENTS OCCUR WHEN HEMOLYSIS RELEASES FREE HEMOGLOBIN INTO THE BLOODSTREAM. THIS FREE HEMOGLOBIN BINDS TO HAPTOGLOBIN, FORMING METHEMOGLOBIN. WHEN HAPTO GLOBIN IS SATURATED, FREE METHEMOGLOBIN CAN PRECIPITATE IN THE MICROCIRCULATION. THESE PRECIPITATES CAN ACTIVATE PLATELETS AND CLOTTING FACTORS, LEADING TO THROMBUS FORMATION. AS MENTIONED, MECHANICAL CIRCULATORY DEVICES HAVE BEEN SHOWN TO INDUCE, ON SOME OCCASION, MARKED SHEAR STRESS AND, SUBSEQUENTLY, BLOOD ELEMENTS DAMAGE. A RETROSPECTIVE, OBSERVATIONAL, SINGLE-CENTER STUDY WAS CONDUCTED ON ADULT (>18 YEARS OF AGE) PATIENTS WITH G6PD DEFICIENCY, WHO UNDERWENT LVAD IMPLANTATION USING THE HM3 LVAD BETWEEN 2017 AND 2022. HEMOLYSIS-RELATED INVESTIGATION AND FINDINGS AS WELL AS IN-HOSPITAL OUTCOME WERE ASSESSED. WHILE THE STUDY WAS LIMITED IN SIZE, LVADS SEEM SAFE FOR G6PD-DEFICIENT PATIENTS AND OFFER SIGNIFICANT CLINICAL BENEFITS. LARGER STUDIES ARE NEEDED TO CONFIRM THIS AND ASSESS LONG-TERM INTERACTIONS. RESULTS SUGGEST THAT G6PD DEFICIENCY DOES NOT NEGATIVELY IMPACT POST-IMPLANT OUTCOMES OR INCREASE THE RISK OF HEMOLYSIS-RELATED COMPLICATIONS IN LVAD PATIENTS. NO SIGNIFICANT INCREASE IN HEMOLYSIS BIOMARKERS WAS OBSERVED DURING IN-HOSPITAL FOLLOW-UP OR POST-DISCHARGE, AND NO MAJOR HEMOLYSIS-RELATED COMPLICATIONS OCCURRED IN THESE PATIENTS. Manufacturer Narrative: SECTIONS A AND D: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. GRASSO, E., LORUSSO, R., IBRAHIM, A., ABDELHAMED, M.I., ABDALLAH, H., AND SAYIN, O.A. PERFUSION 2025, VOL. 0(0) 1¿8 © THE AUTHOR(S) 2025. DOI: 10.1177/02676591251334897 DEPARTMENT OF ADULT CARDIAC SURGERY, PRINCE SULTAN CARDIAC CENTER- AL HASSA, SAUDI ARABIA CARDIO-THORACIC SURGERY DEPARTMENT, HEART & VASCULAR CENTRE, MAASTRICHT UNIVERSITY MEDICAL CENTRE (MUMC+), MAASTRICHT, THE NETHERLANDS. CARDIOVASCULAR RESEARCH INSTITUTE MAASTRICHT (CARIM), MAASTRICHT, THE NETHERLANDS DEPARTMENT RESEARCH& BIOSTATISTICS, PRINCE SULTAN CARDIAC CENTER- AL HASSA, SAUDI ARABIA 5DEPARTMENT OF VASCULAR SURGERY, KANTONSSPITAL, AARAU, SWITZERLAND. NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED. Manufacturer Narrative: SECTION D4: MODEL AND CATALOG NUMBERS CORRECTED. MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. THE REPORTED INFORMATION WAS OBTAINED THROUGH A LITERATURE REVIEW. THE HEARTMATE 3 DEVICE SERIAL NUMBERS AND OTHER SPECIFIC CASE/PATIENT INFORMATION ARE NOT AVAILABLE. THE CURRENT REVISION OF THE INSTRUCTIONS FOR USE (IFU) CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LEFT VENTRICLE ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU) IS CURRENTLY AVAILABLE. SECTION 1 OF THE IFU, ¿INTRODUCTION,¿ LISTS POTENTIAL ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22330227&pc=DSQ | 2916596-2025-03988 | 2025-06-09 04:00 | Death | THORATEC CORPORATION | 2025-06-25 04:00 | DSQ | HEARTMATE 3 LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Ischemia Stroke | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT PATIENT PASSED AWAY DUE TO ISCHEMIC STROKE WITH INFAUST PROGNOSIS. THE OUTCOME WAS NOT DEVICE OR THERAPY RELATED; THE PATIENT WAS STRAIGHT FORWARD IMPLANTED ON (B)(6) 2025. POSTOPERATIVE PHASE WAS HEMODYNAMICALLY NO PROBLEM, APART FROM AWAKENING. CARDIAC COMPUTED TOMOGRAPHY (CCT) SCAN SHOWED THEN TWO SEVERE STROKE AREAS WITH AN INFAUST PROGNOSIS. THE DEVICE OPERATED AS EXPECTED. THERE WERE NO CHANGES TO PATIENT CONDITION OR ANTICOAGULATION STATUS THAT MAY HAVE CONTRIBUTED. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: THE PUMP REMAINS IN USE SUPPORTING THE PATIENT. A CORRELATION BETWEEN THE DEVICE AND THE REPORT OF STROKE COULD NOT BE CONCLUSIVELY DETERMINED. STROKE IS LISTED IN THE INSTRUCTIONS FOR USE AS A POTENTIAL ADVERSE EVENT THAT MAY BE ASSOCIATED WITH THE USE OF HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. STROKE HAS BEEN PREVIOUSLY INVESTIGATED AND WILL CONTINUE TO BE MONITORED THROUGH QUALITY DATA REVIEWS, WHICH ARE CONDUCTED ON PRODUCTION AND POST PRODUCT SIGNALS TO EVALUATE IF PRODUCTS ARE CONFORMING TO PRODUCT REQUIREMENTS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22330479&pc=PYX | 1220648-2025-29787 | 2025-03-22 04:00 | Death | ABIOMED, INC. | 2025-06-25 04:00 | PYX | IMPELLA RP FLEX | Adverse Event Without Identified Device or Use Problem | Ischemia | P170011 |  | 1 | Event Description: NOTIFICATION WAS RECEIVED VIA ABIOMED¿S LONG-TERM OUTCOME AND QUALITY INDICATOR (LOQI) IMPELLA REGISTRY REGARDING A PATIENT THAT ENDURED RIGHT SAPHENOFEMORAL ARTERY OCCLUSION WITH A " PROBABLE" RELATIONSHIP TO THE IMPELLA DEVICE USED: IT WAS REPORTED AS ¿DECREASED TEMPERATURE, INCREASED TENSION, AND DECREASED PULSES WERE APPRECIATED IN THE LOWER RIGHT EXTREMITY. THE PATIENT UNDERWENT ULTRASOUND AND COMPUTED TOMOGRAPHY (CT) IMAGING OF THE RIGHT LEG. ULTRASOUND REVEALED ARTERIAL OCCLUSION OF THE SUPERFICIAL FEMORAL AND POPLITEAL ARTERIES. CT REVEALED THAT THE LUMEN OF THE RIGHT EXTERNAL ILIAC AND PROXIMAL FEMORAL ARTERY WAS COMPLETELY FILLED BY THE IMPELLA CP CATHETER (PLACED AT OUTSIDE FACILITY). VASCULAR SURGERY ASSESSED THE LEG AND DETERMINED THAT THE LIMB WAS UNSALVAGEABLE AND WOULD OTHERWISE NEED TO BE AMPUTATED, HOWEVER, GIVEN THE PATIENT'S STATE AND CONCERN FOR DEVELOPING WORSENING METABOLIC PROBLEMS, THE PATIENT WAS TOO UNSTABLE TO UNDERGO SURGERY. VASCULAR SURGERY ALSO EXPLAINED THAT THE PATIENT LIKELY HAD CHRONIC RIGHT SUPERIOR FEMORAL ARTERY (SFA) OCCLUSION WHICH WAS EXACERBATED BY OCCLUSION OF THE EXTERNAL ILIAC FROM THE IMPELLA CP PLACEMENT.¿ ULTIMATELY, CARE WAS WITHDRAWN AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS DISCARDED BY THE CUSTOMER AND THEREFORE, AN EVALUATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MANUFACTURER DEVICE REPORT WILL BE FILED. AS NOTED IN THE IMPELLA INSTRUCTIONS FOR USE: IMPELLA RP FLEX WITH SMART ASSIST SYSTEM WITH AUTOMATED IMPELLA CONTROLLER SECTION: CONTRAINDICATIONS (UNITED STATES) ¿THE IMPELLA RP FLEX WITH SMARTASSIST SYSTEM IS CONTRAINDICATED FOR USE WITH PATIENTS EXPERIENCING ANY OF THE FOLLOWING CONDITIONS: DISORDERS OF THE PULMONARY ARTERY WALL THAT WOULD PRECLUDE PLACEMENT OR CORRECT POSITIONING OF THE IMPELLA RP FLEX WITH SMARTASSIST DEVICE; MECHANICAL VALVES, SEVERE VALVULAR STENOSIS OR VALVULAR REGURGITATION OF THE TRICUSPID OR PULMONARY VALVE; MURAL THROMBUS OF THE RIGHT ATRIUM OR VENA CAVA; ANATOMIC CONDITIONS PRECLUDING INSERTION OF THE PUMP; PRESENCE OF A VENA CAVA FILTER OR CAVAL INTERRUPTION DEVICE, UNLESS THERE IS CLEAR ACCESS FROM THE INTERNAL JUGULAR VEIN TO THE RIGHT ATRIUM THAT IS LARGE ENOUGH TO ACCOMMODATE A 22 FR CATHETER.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22330562&pc=DSQ | 2916596-2025-03895 | 2025-01-01 05:00 | Death | THORATEC CORPORATION | 2025-06-25 04:00 | DSQ | HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM | Adverse Event Without Identified Device or Use Problem | Intracranial Hemorrhage; Thrombosis/Thrombus; Liver Failure | P160054 |  | 1 | Event Description: IT WAS REPORTED THROUGH THE RESEARCH ARTICLE TITLED, ¿HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE SYSTEM IN PATIENTS WITH GLUCOSE-6-PHOSPHATE DEHYDROGENASE DEFICIENCY¿, THAT HEARTMATE 3 MAY BE ASSOCIATED WITH HEMODYNAMIC INSTABILITY, HEPATIC FAILURE, STROKE/HEMORRHAGE, THROMBOEMBOLIC EVENTS AND DEATH.GLUCOSE-6-PHOSPHATE DEHYDROGENASE (G6PD) DEFICIENCY IS A GENETIC ENZYMATIC DISORDER THAT AFFECTS MILLIONS OF PEOPLE WORLDWIDE AND CHARACTERIZED BY HEMOLYSIS UNDER OXIDATIVE STRESS. LEFT VENTRICULAR ASSIST DEVICES (LVADS) HAVE SUBSTANTIALLY ENHANCED SURVIVAL AND QUALITY OF LIFE FOR INDIVIDUALS WITH ADVANCED HEART FAILURE. HOWEVER, THEIR USE WAS ASSOCIATED WITH THE RISK OF HEMOLYSIS, THROMBOSIS, AND EMBOLIC EVENTS. THESE RISKS MAY BE HEIGHTENED IN PATIENTS WITH GLUCOSE-6 PHOSPHATE DEHYDROGENASE (G6PD) DEFICIENCY. THE PATIENT (PATIENT 4) SUFFERED RIGHT SIDED HEART FAILURE AND ADVANCED TO HEPATIC FAILURE AND CEREBRAL HEMORRHAGE WHICH CAUSED HIS DEATH AFTER 16 MONTHS OF SURGERY. IT WAS ESTIMATED THAT ABOUT 400 MILLION PEOPLE ARE AFFECTED BY G6PD DEFICIENCY WORLDWIDE.14 THIS GENETIC CONDITION IS INHERITED AS AN X-LINKED RECESSIVE STATE AND PRESENTS ONE OF THE MOST FREQUENTLY ENCOUNTERED RED-CELL ENZYMOPATHIES. THE MAJOR INTEREST IN THE G6PD-DEFICIENT STATE RESULTS FROM THE ASSOCIATED HEMOLYTIC ANEMIA RESULTING FROM OXIDATIVE STRESS.15,16 IT IS WORTH MENTIONING THAT HEMOLYSIS-INDUCED THROMBOTIC EVENTS IN G6PD PATIENTS OCCUR WHEN HEMOLYSIS RELEASES FREE HEMOGLOBIN INTO THE BLOODSTREAM. THIS FREE HEMOGLOBIN BINDS TO HAPTOGLOBIN, FORMING METHEMOGLOBIN. WHEN HAPTOGLOBIN IS SATURATED, FREE METHEMOGLOBIN CAN PRECIPITATE IN THE MICROCIRCULATION. THESE PRECIPITATES CAN ACTIVATE PLATELETS AND CLOTTING FACTORS, LEADING TO THROMBUS FORMATION. AS MENTIONED, MECHANICAL CIRCULATORY DEVICES HAVE BEEN SHOWN TO INDUCE, IN SOME OCCASION, MARKED SHEAR STRESS AND, SUBSEQUENTLY, BLOOD ELEMENTS DAMAGE. A RETROSPECTIVE, OBSERVATIONAL, SINGLE-CENTER STUDY WAS CONDUCTED ON ADULT (>18 YEARS OF AGE) PATIENTS WITH G6PD DEFICIENCY, WHO UNDERWENT LVAD IMPLANTATION USING THE HM3 LVAD BETWEEN 2017 AND 2022. HEMOLYSIS-RELATED INVESTIGATION AND FINDINGS AS WELL AS IN-HOSPITAL OUTCOME WERE ASSESSED. HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICES (LVADS) WERE SUCCESSFULLY IMPLANTED IN FIVE ADULT PATIENTS WITH G6PD DEFICIENCY, INCLUDING ONE INDIVIDUAL WITH ASSOCIATED SICKLE CELL TRAIT (SCT). THERE WERE NO MAJOR COMPLICATIONS OR FATALITIES DURING THE HOSPITALIZATION PERIOD. THE AVERAGE FOLLOW-UP DURATION WAS 30 MONTHS (MEAN 30.4 ± 13). DURING THE FOLLOW-UP PERIOD, TWO PATIENTS DIED, TWO LVADS WERE EXPLANTED, AND ONE PATIENT RECEIVED A HEART TRANSPLANT. NO INSTANCES OF MACROSCOPIC HEMOLYSIS WERE OBSERVED THROUGHOUT THE FOLLOW-UP PERIOD. HEMOLYSIS HAS BECOME AN INCREASINGLY RECOGNIZED COMPLICATION OF LVADS SUPPORT. THE INCIDENCE OF HEMOLYSIS VARIES (5%-18%) DEPENDING ON THE DEFINITION AND AMONG DIFFERENT GENERATIONS OF LVADS, BEING SLIGHTLY HIGHER IN CONTINUOUS-FLOW THAN TO PULSATILE DEVICES. CAREFUL CLINICAL AND LABORATORY MONITORING REMAIN THE CORNERSTONE OF PATIENT MANAGEMENT IN THIS RESPECT. G6PDDEFICIENCYIS ONE OF THE COMMON HUMAN ENZYME DEFICIENCY DISORDERS WITH VARIABLE PREVALENCE IN THE MIDDLE EAST AND MEDITERRANEAN COUNTRIES. THIS CONDITION MIGHT REPRESENT A RISK FACTOR OF EXACERBATION OF HEMOLYSIS IN LVAD RECIPIENTS. THE STUDY DEFINED HEMOLYSIS ACCORDING TO THE INSTITUTIONAL PROTOCOL, WHICH INCLUDES PATIENTS WITH THE FOLLOWING 3 PARAMETERS: UNEXPLAINED ANEMIA (HGB250U/L) IN THE ABSENCE OF OBVIOUS BLEEDING, TOTAL BILIRUBIN(=1MG/DL) AND HIGH LDH (>250U/L). PLASMA FREE HGB DATA WAS NOT AVAILABLE FOR MOST PATIENTS, AS THEY WERE REFERRED TO A LABORATORY AND WERE NOT ROUTINELY MONITORED DURING THE STUDY PERIOD. THE ONSET OF MACROSCOPIC HEMATURIA WAS ADDITIONALLY EVALUATED. PERIOPERATIVELY BLOOD GAS CHECKS WERE REPETITIVELY EXECUTED TO DETECT ACIDOSIS AND HYPERGLYCEMIA, WHICH ARE POTENTIAL PRECIPITATING FACTORS FOR HEMOLYSIS. ALL PATIENTS UNDERWENT CARDIAC SURGERY THROUGH FULL MEDIAN STERNOTOMY AND RECEIVED THE VAD SYSTEM HM3. IMMEDIATELY AFTER COMING OFF CPB, THE PATIENT URINE WAS CHECKED REGARDING ANY SIGNS OF HEMATURIA. POSTOPERATIVELY, ANTICOAGULATION WAS STARTED WITH LOW MOLECULAR WEIGHT HEPARIN ON THE ICU AFTER 12¿24 HOURS, DEPENDING ON BLEEDING AMOUNT. ANTIPLATELET DRUGS AND ORAL ANTICOAGULANTS WERE STARTED ON THE SECOND POSTOPERATIVE DAY, UNLESS BLEEDING COMPLICATIONS OCCUR. THERE WAS NEITHER MORTALITY NOR COMPLICATION WITHIN 30 DAYS, 3 MONTHS AND 6 MONTHS POSTOPERATIVELY. THERE WERE NO IN-HOSPITAL DEATHS. IN CONCLUSION, WHILE THE STUDY WAS LIMITED IN SIZE, LVADS SEEM SAFE FOR G6PD-DEFICIENT PATIENTS AND OFFER SIGNIFICANT CLINICAL BENEFITS. LARGER STUDIES ARE NEEDED TO CONFIRM THIS AND ASSESS LONG-TERM INTERACTIONS. RESULTS SUGGEST THAT G6PD DEFICIENCY DOES NOT NEGATIVELY IMPACT POST-IMPLANT OUTCOMES OR INCREASE THE RISK OF HEMOLYSIS-RELATED COMPLICATIONS IN LVAD PATIENTS. NO SIGNIFICANT INCREASE IN HEMOLYSIS BIOMARKERS WAS OBSERVED DURING IN-HOSPITAL FOLLOW-UP OR POST-DISCHARGE, AND NO MAJOR HEMOLYSIS-RELATED COMPLICATIONS OCCURRED IN THESE PATIENTS. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) AND THE REPORTED EVENTS COULD NOT CONCLUSIVELY BE ESTABLISHED THROUGH THIS EVALUATION. THE SERIAL NUMBERS OF THE HEARTMATE 3 LEFT VENTRICULAR SYSTEMS IN USE WERE NOT PROVIDED. THE CURRENT REVISIONS OF THE IFU AND PATIENT HANDBOOK CAN BE FOUND ON THE EIFU PAGE OF THE ABBOTT WEBSITE. THE HEARTMATE 3 LVAS IFU, REV. D AND THE HEARTMATE 3 LVAS PATIENT HANDBOOK, REV. D ARE CURRENTLY AVAILABLE. SECTION 1 OF THIS DOCUMENT LISTS ADVERSE EVENTS, INCLUDING RIGHT HEART FAILURE, HEPATIC FAILURE, THROMBOEMBOLISM, HEMOLYSIS, STROKE, BLEEDING, AND DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM. SECTION 6 "PATIENT CARE AND MANAGEMENT" (UNDER "RIGHT HEART FAILURE") DISCUSSES THE POTENTIAL DEVELOPMENT OF RIGHT HEART FAILURE FOLLOWING IMPLANT AND OUTLINES THE ASSOCIATED TREATMENT OPTIONS, INCLUDING INOTROPES AND RIGHT VENTRICULAR ASSIST DEVICE (RVAD) PLACEMENT. SECTION 6 (UNDER CAUTION!) STATES: "RIGHT HEART FAILURE CAN OCCUR FOLLOWING IMPLANTATION OF THE PUMP. RIGHT VENTRICULAR DYSFUNCTION, ESPECIALLY WHEN COMBINED WITH ELEVATED PULMONARY VASCULAR RESISTANCE, MAY LIMIT THE EFFECTIVENESS OF THE LVAS DUE TO REDUCED FILLING OF THE PUMP." SECTION 6, ¿PATIENT CARE AND MANAGEMENT¿, ALSO LISTS THROMBOEMBOLISM AS A POTENTIAL LATE POSTIMPLANT COMPLICATION. ADDITIONALLY, THIS SECTION, UNDER ¿ANTICOAGULATION¿, PROVIDES THE RECOMMENDED ANTICOAGULATION REGIMEN, INCLUDING INTERNATIONAL NORMALIZED RATIO (INR) RANGE, AS WELL AS SUGGESTED ANTICOAGULATION MODIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: B5: SPECIFIC PATIENT INFORMATION AND DEVICE SERIAL NUMBER ARE DOCUMENTED AS UNKNOWN. DETAILS ARE LISTED IN THE ATTACHED ARTICLE. DEVICE WAS IMPLANTED AT TIME OF EVENT. AUTHOR CITATION: GRASSO E, LORUSSO R, IBRAHIM A, ABDELHAMED MI, ABDALLAH H, SAYIN OA. HEARTMATE 3 LEFT VENTRICULAR ASSIST DEVICE SYSTEM IN PATIENTS WITH GLUCOSE-6-PHOSPHATE DEHYDROGENASE DEFICIENCY. PERFUSION. 2025;0(0). DOI:10.1177/02676591251334897. DEPARTMENT OF ADULT CARDIAC SURGERY, PRINCE SULTAN CARDIAC CENTER- AL HASSA, SAUDI ARABIA CARDIO-THORACIC SURGERY DEPARTMENT, HEART & VASCULAR CENTRE, MAASTRICHT UNIVERSITY MEDICAL CENTRE (MUMC+), MAASTRICHT, THE NETHERLANDS CARDIOVASCULAR RESEARCH INSTITUTE MAASTRICHT (CARIM), MAASTRICHT, THE NETHERLANDS DEPARTMENT RESEARCH& BIOSTATISTICS, PRINCE SULTAN CARDIAC CENTER- AL HASSA, SAUDI ARABIA DEPARTMENT OF VASCULAR SURGERY, KANTONSSPITAL, AARAU, SWITZERLAND. NO FURTHER INFORMATION PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22330630&pc=LWQ | 2025587-2025-04529 | 2025-06-10 04:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-25 04:00 | LWQ | HALL EASY FIT | Insufficient Information | Insufficient Information | P790018 |  | 1 | Event Description: IT WAS REPORTED THAT THE HALL MECHANICAL HEART VALVE WAS FAILING, WITH ALLEGATIONS THAT THE VALVE LEAFLET WAS MALFUNCTIONING AND REQUIRED REPLACEMENT. THE PATIENT REPRESENTATIVE REPORTED THAT THE PHYSICIAN INDICATED THE HEART VALVE, WHICH WAS EXPECTED TO LAST 40 YEARS, ONLY LASTED 12 TO 13 YEARS. AN AUDIBLE VALVE CLICKING WAS NOTED. THE PATIENT WAS HOSPITALIZED DUE TO THE VALVE FAILURE, AND SUBSEQUENTLY, THE PATIENT DIED AFTER THE REPORTED VALVE FAILURE. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE VALVE COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22330696&pc=LWS | 2649622-2025-17215 | 2025-05-26 04:00 | Death | MPRI | 2025-06-25 04:00 | LWS | SPRINT FIDELIS | Signal Artifact/Noise; Fracture; Over-Sensing; Impedance Problem | Cardiac Arrest; Hypoxia; Distress | P920015 |  | 1 | Event Description: IT WAS REPORTED THAT THE LEAD INTEGRITY ALERT (LIA) WAS TRIGGERED FOR THE RIGHT VENTRICULAR (RV) LEAD DUE TO HIGH-RATE NON-SUSTAINED EPISODES AND HIGH SENSING INTEGRITY COUNTER (SIC). IN ADDITION, VARIABLE PACING IMPEDANCE AND NOISE WERE OBSERVED. AN RV TIP FRACTURE WAS SUSPECTED. THE PATIENT WAS BROUGHT TO THE HOSPITAL TO REPLACE THE RIGHT VENTRICULAR (RV) LEAD. AFTER START OF THE PROCEDURE AND AFTER LEFT SUBCLAVIAN VENOUS ACCESS, THE PATIENT BEGAN TO SHOW SIGNS OF DISTRESS AND HYPOXIA, WITH SUBSEQUENT RESPIRATORY AND CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION WAS PERFORMED SEVERAL TIMES IN THE PROCEDURAL ROOM, HOWEVER THE PATIENT EXPIRED. THE RV LEAD REMAINS IN-SITU. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22331031&pc=FSA | 9681684-2025-00052 | 2025-05-28 04:00 | Death | ARJOHUNTLEIGH MAGOG INC. | 2025-06-25 04:00 | FSA | SARA FLEX | Use of Device Problem | Bone Fracture(s) |  |  | 1 | Event Description: THE RESIDENT FELL FROM THE SLING USED WITH SARA FLEX. THE INFORMATION RECEIVED SUGGESTS THAT THE SLING MIGHT HAVE BEEN LARGE. THE RESIDENT STATED THAT SHE FELT THAT THE SLING WAS TOO LOOSE. THE RESIDENT THEN FELL, BROKE HER BACK, AND SUBSEQUENTLY DIED. THE TIME BETWEEN AN INCIDENT AND PATIENT DEATH IS UNKNOWN. Manufacturer Narrative: THE INVESTIGATION PROCESS IS STILL ONGOING. ADDITIONAL INFORMATION WILL BE PROVIDED IN NEXT REPORT. Manufacturer Narrative: PROCESS OF GATHERING AND ANALYZING INFORMATION IS ONGOING. ADDITIONAL INFORMATION WILL BE PROVIDED UPON CONCLUSION OF THE INVESTIGATION. Manufacturer Narrative: THE RESIDENT FELL FROM THE SLING USED WITH SARA FLEX. THE INFORMATION RECEIVED SUGGESTS THAT THE SLING MIGHT HAVE BEEN LARGE BECAUSE THE RESIDENT STATED THAT SHE FELT THAT THE SLING WAS TOO LOOSE. THE RESIDENT THEN FELL, BROKE HER BACK, AND SUBSEQUENTLY DIED. THE TIME BETWEEN AN INCIDENT AND PATIENT DEATH IS UNKNOWN. IN THE COMPLAINT THE REPORTER DID NOT DIRECTLY LINK THE PATIENT'S DEATH TO THE FALL OR BROKEN BACK. NO MALFUNCTION WAS FOUND IN THE LIFT OR SLING THAT MAY HAVE CONTRIBUTED TO THE CLAIMED EVENT. THE REPORTED INDICATED THAT THE SLING SIZE WAS SELECTED ACCORDING TO THE RESIDENT¿S WEIGHT BASED ON THE NURSING ASSESSMENT AND ARJO SIZING GUIDE. THE RESIDENT EXPRESSED CONCERN THAT THE SLING FELT TOO LOOSE, WHICH MAY INDICATE THAT WEIGHT-BASED SIZING ALONE WAS INSUFFICIENT TO ENSURE PROPER FIT AND SUPPORT FOR THIS INDIVIDUAL. SLING INSTRUCTIONS FOR USE (04.SF.00-INT1A_8), GUIDES THROUGH THE SIZE SELECTION (THERE IS A SEPARATE CHAPTER), TO ENSURE THE SELECTION OF THE SLING IS BASED ON THE MEASURING OF THE PATIENT'S WAIST. THERE IS ALSO INFORMATION THAT THE SAFE WORKING LOAD FOR ALL SLINGS (S, M, L, XL) IS 200 KG. RELYING ONLY ON BODY MASS MAY RESULT IN INADEQUATE SUPPORT OR A LOOSE FIT, ESPECIALLY FOR INDIVIDUALS WITH ATYPICAL BODY PROPORTIONS. APPROPRIATE SIZING IS ESSENTIAL TO ENSURE COMFORT, SAFETY, AND EFFECTIVE FUNCTIONALITY DURING TRANSFERS. ACCORDING TO THE SARA FLEX INSTRUCTIONS FOR USE (IFU, 04.KL.00.EN_8), ¿BEFORE USE THE CAREGIVER SHOULD ALWAYS CONSIDER THE PATIENT¿S/RESIDENT¿S MEDICAL CONDITION, PHYSICAL AND MENTAL CAPABILITIES. IN ADDITION THE PATIENT/RESIDENT MUST BE ABLE TO BEAR WEIGHT ON AT LEAST ONE LEG AND HAVE SOME TRUNK STABILITY, BE ABLE TO SIT ON THE EDGE OF THE BED, WEIGH IN THE RANGE OF 45 TO 200 KG (99 LB ¿ 440 LB), AND HAVE A HEIGHT WITHIN THE RANGE OF 145 TO 195 CM (4¿8¿ ¿ 6¿4¿). IF THE PATIENT DOES NOT MEET THESE CRITERIA AN ALTERNATIVE EQUIPMENT/SYSTEM SHALL BE USED.¿ ARJO RECEIVED INFORMATION THAT THE RESIDENT¿S ARMS WERE POSITIONED INSIDE THE SLING, RATHER THAN ON THE DESIGNATED GRIP HANDLES DURING THE EVENT. IT MAY HAVE REDUCED UPPER-BODY STABILITY DURING TRANSFER AS THE GRIP HANDLES ARE DESIGNED SPECIFICALLY FOR THE PATIENT TO HOLD ONTO DURING THE SIT-TO-STAND MOVEMENT. ACCORDING TO IFU, THE CAREGIVER SHOULD ¿ASK OR ASSIST PATIENT TO PLACE HIS/HER HANDS ON THE PATIENT/RESIDENT HANDLES.¿ SUM UP, THE INCIDENT APPEARS TO HAVE RESULTED FROM A POTENTIAL UNINTENTIONAL USE ERROR (SLING MAY NOT HAVE BEEN A PROPER SIZE BASED ON THE MEASURING OF THE PATIENT'S WAIST). NO DEFICIENCIES WERE IDENTIFIED IN THE LIFT OR SLING; BOTH MET THEIR SPECIFICATIONS. THEY WERE USED AS A SYSTEM FOR PATIENT TRANSFER AND THEREFORE PLAYED A ROLE IN THE INCIDENT. THIS COMPLAINT DECIDED TO BE REPORTABLE DUE TO THE PATIENT¿S FALL.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22331057&pc=OZD | 1220648-2025-29795 | 2022-03-03 05:00 | Death | ABIOMED, INC. | 2025-06-25 04:00 | OZD | IMPELLA CP | Positioning Failure | Arrhythmia | P140003 |  | 1 | Event Description: LEGAL FORWARDED DOCUMENTS TO THE COMPLAINT TEAM ON THAT SPEAKS TO THE IMPELLA CP SUPPORT IN WHICH THERE IS AN ALLEGATION/LITIGATION. THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA CP DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, AFTER ARRIVING IN THE ICU UNIT, SUCTION ALARMS WERE PRESENT, AND BEDSIDE ECHO WAS ORDERED TO VERIFY POSITIONING. IMPELLA WAS PULLED BACK AND PROPER POSITIONING WAS VERIFIED. IMPELLA POSITION CONTINUED TO NEED ADJUSTMENT DURING SUPPORT. THE PATIENT¿S ARTERIAL BLOOD PRESSURE STARTED TRENDING DOWN TO MAPS OF 50 AND THEN 30. AS EXPECTED, THE PATIENT DID NOT HAVE A PULSE, HOWEVER, THE IMPELLA SUGGESTED DECREASED FLOW WITH LOW PRESSURE READINGS. PATIENT'S NOREPINEPHRINE WAS INCREASED TO 30 WITHOUT EFFECT. GIVEN THAT THE PATIENT HAD DROPPING MAPS TO 30S, A CODE WAS CALLED. EPINEPHRINE 1MG IV WAS ADMINISTERED, THE IMPELLA PUMP WAS REDUCED TO P2, AND CHEST COMPRESSIONS WERE PERFORMED; PATIENT HAD VT ON THE MONITOR AND 1 SHOCK WAS DELIVERED FOLLOWING WHICH PATIENT HAD RECOVERY OF THE MAPS TO >70S. AN EPINEPHRINE DRIP WAS STARTED IN ADDITION TO THE NOREPINEPHRINE GTT. PATIENT WAS ALREADY ON AMIODARONE GTT AND A LIDOCAINE GTT WAS ALSO INITIATED. PATIENT'S IMPELLA WAS REPOSITIONED. A BEDSIDE ECHO WAS PERFORMED BY THE CCU FELLOW WHICH SHOWED POOR VENTRICULAR CONTRACTILITY WITH AN EF OF 5-10%. PATIENT CONTINUED TO HAVE DROP IN HIS MAPS (20S) WHILE REMAINING PULSELESS AND REQUIRING ANOTHER 1.5 EPINEPHRINE BOLUSES AS WELL AS SHOCK X 2. ON EXAM, THE PATIENT HAD LEFT MID-DILATED PUPIL. THE FAMILY DECIDED FOR COMFORT CARE AND THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. IMPELLA CP WITH SMARTASSIST FOR USE DURING CARDIOGENIC SHOCK AND HIGH-RISK PCI INSTRUCTIONS FOR USE FOR THE RELATED EVENT ARE AS FOLLOWS: SECTION: WARNINGS, CAUTIONS AND PRECAUTIONS: WARNINGS. ¿TO REDUCE THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING PERFORATION) WHEN MANIPULATING THE HEART DURING CARDIAC SURGERY, EVALUATE THE POSITION OF THE PUMP USING IMAGING GUIDANCE PRIOR TO MANIPULATING THE HEART, AND MONITOR POSITION.¿ ¿PHYSICIANS SHOULD EXERCISE SPECIAL CARE WHEN INSERTING THE IMPELLA CATHETER DURING ACTIVE CARDIOPULMONARY RESUSCITATION (CPR). IN ADDITION, ACTIVE CPR MANEUVERS MAY CHANGE THE POSITION OF THE IMPELLA DEVICE, INTRODUCING THE RISK OF CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION). CHECK THAT THE PUMP IS POSITIONED CORRECTLY AFTER CPR WITH CHEST X-RAY GUIDANCE.¿ UNDERSTANDING AND MANAGING IMPELLA CATHETER POSITION ALARMS ¿IF THE IMPELLA CATHETER IS COMPLETELY OUT OF THE VENTRICLE, DO NOT ATTEMPT TO REPOSITION THE CATHETER ACROSS THE VALVE WITHOUT A GUIDEWIRE.¿ CORRECT IMPELLA CATHETER POSITION: ¿FOR OPTIMAL POSITIONING OF THE IMPELLA CATHETER, THE INLET AREA OF THE CATHETER SHOULD BE 3.5 CM BELOW THE AORTIC VALVE ANNULUS AND WELL AWAY FROM PAPILLARY MUSCLE AND SUBANNULAR STRUCTURES. THE OUTLET AREA SHOULD BE WELL ABOVE THE AORTIC VALVE.¿
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22331319&pc=OZD | 1220648-2025-29798 | 2025-06-11 04:00 | Death | ABIOMED, INC. | 2025-06-25 04:00 | OZD | IMPELLA 5.5 | Positioning Failure; Restricted Flow rate | Thrombosis/Thrombus | P140003 |  | 1 | Event Description: THE COMPLAINANT REPORTED A PATIENT WAS IMPLANTED WITH AN IMPELLA 5.5 DEVICE FOR MECHANICAL CIRCULATORY SUPPORT. WHILE ON SUPPORT, THERE WAS A RED ALARM IN IC "IMPELLA POSITION IN AORTA" ISSUE THAT RESOLVED INDEPENDENTLY. ADDITIONALLY, THE PHYSICIAN REPORTED A SUCTION ALARM WITH A SPIKE THAT WAS VISIBLE IN THE MC, AND SINCE THEN THE MC HAD BEEN FLAT, THE IMPELLA WAS REPOSITIONED SEVERAL TIMES WITHOUT SUCCESS. HEPARIN WAS PAUSED FOR SEVERAL HOURS DUE TO THE SPIKE. IT WAS NOTED OF A CLOT BEING SUSPECTED. HOWEVER, AFTER SEVERAL DAYS OF SUPPORT, THE PATIENT EXPIRED. Manufacturer Narrative: THE IMPELLA DEVICE WAS NOT RECEIVED FROM THE CUSTOMER, THEREFORE, INVESTIGATION OF THE DEVICE WAS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES) ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION)¿.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22331323&pc=FSA | 1419652-2025-00027 | 2025-05-28 04:00 | Death | ARJOHUNTLEIGH MAGOG INC. | 2025-06-25 04:00 | FSA | SARA FLEX | Use of Device Problem | Bone Fracture(s) |  |  | 1 | Event Description: THE RESIDENT FELL FROM THE SLING USED WITH SARA FLEX. THE INFORMATION RECEIVED SUGGESTS THAT THE SLING MIGHT HAVE BEEN LARGE. THE RESIDENT STATED THAT SHE FELT THAT THE SLING WAS TOO LOOSE. THE RESIDENT THEN FELL, BROKE HER BACK, AND SUBSEQUENTLY DIED. THE TIME BETWEEN AN INCIDENT AND PATIENT DEATH IS UNKNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22331458&pc=DSQ | 2916596-2025-03910 | 2025-04-11 04:00 | Death | THORATEC CORPORATION | 2025-06-25 04:00 | DSQ | THORATEC® HEARTMATE 3® LVAS IMPLANT KIT | Adverse Event Without Identified Device or Use Problem | Bacterial Infection; Localized Skin Lesion; Skin Infection | P160054 |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT PASSED AWAY. THE PATIENT COMMUNICATED THAT THEY WERE NOT EXPERIENCING ACCEPTABLE QUALITY OF LIFE WITH CONTINUED LEFT VENTRICULAR ASSIST DEVICE (LVAD) THERAPY GIVEN INABILITY TO MAINTAIN FULL FUNCTIONAL INDEPENDENCE. IT WAS ADDITIONALLY REPORTED THAT THE PATIENT WAS ADMITTED AFTER WORK-UP FOR BACK PAIN AND PROGRESSING DEBILITY REVEALED SPINAL LESIONS, FELT TO BE INFECTION, WITH BIOPSY PRELIMINARY RESULTS FROM (B)(6) 2025 REVEALING ENTOEROBACTER. ON (B)(6) 2025 THE PATIENT VALIDATED THAT THEY WERE NO LONGER EXPERIENCING ACCEPTABLE QUALITY OF LIFE WITH THE LVAD GIVEN THE PATIENT'S FUNCTIONAL DECLINE. THE PATIENT REQUESTED THE LVAD BE DECOMMISSIONED, AND THE PATIENT WAS MOVED TO HOSPICE. THE PATIENT PASSED AWAY ON (B)(6) 2025. Manufacturer Narrative: MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS), SERIAL NUMBER: (B)(6), AND THE REPORTED EVENTS COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. ADDITIONALLY, A SPECIFIC CAUSE FOR THE PATIENT¿S INFECTION COULD NOT BE CONCLUSIVELY DETERMINED. IT WAS REPORTED THAT THE PATIENT¿S DEATH WAS NOT CONSIDERED TO BE DEVICE RELATED, AND IT WAS REPORTED THAT THE LVAD OPERATED AS EXPECTED. THE DEVICE WILL REPORTEDLY NOT BE RETURNED FOR EVALUATION. THE RELEVANT SECTIONS OF THE DEVICE HISTORY RECORDS FOR (B)(6) WERE REVIEWED AND SHOWED NO DEVIATIONS FROM MANUFACTURING OR QUALITY ASSURANCE SPECIFICATIONS. THE HEARTMATE 3 LEFT VENTRICULAR ASSIST SYSTEM (LVAS) INSTRUCTIONS FOR USE (IFU), IS CURRENTLY AVAILABLE. SECTION 1 ¿INTRODUCTION¿ OF THIS DOCUMENT LISTS LOCALIZED INFECTION AND DEATH AS ADVERSE EVENTS THAT MAY BE ASSOCIATED WITH THE USE OF THE HEARTMATE 3 LVAS. ADDITIONALLY, SECTION 6 ¿PATIENT CARE AND MANAGEMENT¿ INCLUDES INFORMATION REGARDING HOW TO PREVENT AND CONTROL INFECTION. THE HEARTMATE 3 LVAS PATIENT HANDBOOK, CONTAINS SEVERAL SECTIONS WITH INFORMATION ABOUT HOW TO PREVENT AND CONTROL INFECTION. THE CURRENT REVISIONS OF THE INSTRUCTIONS FOR USE (IFU) AND PATIENT HANDBOOK CAN ALSO BE FOUND ON THE ELECTRONIC IFU (EIFU) PAGE OF THE ABBOTT WEBSITE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT. Manufacturer Narrative: NO FURTHER INFORMATION WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22331650&pc=GAM | 2210968-2025-07449 | 2025-01-06 05:00 | Death | ETHICON INC. | 2025-06-25 04:00 | GAM | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | Adverse Event Without Identified Device or Use Problem | Sepsis | K022269 |  | 1 | Event Description: TITLE: INSTITUTIONAL EXPERIENCE ON COMPARISON OF ALLOGRAFTS WITH MULTIPLE-RENAL ARTERY VS. SINGLE RENAL ARTERY IN LIVING AND DECEASED DONOR TRANSPLANTATION THE AIM OF THIS STUDY IS TO COMPARE UROLOGICAL (URINE LEAK AND STENOSIS) AND VASCULAR COMPLICATIONS (RENAL ARTERY AND VEIN THROMBOSIS) IN RENAL ALLOGRAFTS WITH A SINGLE ARTERY VS. MULTIPLE ARTERIES. ADDITIONALLY, THE SECONDARY OBJECTIVES INCLUDED ASSESSING DELAYED GRAFT FUNCTION AND GRAFT SURVIVAL AT 3- AND 6-MONTHS POST-TRANSPLANTATION. BETWEEN APRIL 2022 TO APRIL 2024, A TOTAL OF E 95 PATIENTS, 67.4% (64) WERE MALE, AND 32.6% (31) WERE FEMALE. THE MEAN AGE OF THE PARTICIPANTS WAS 34.65 ± 10.41 YEARS. FOR MRA GRAFTS A POLYPROPYLENE SUTURES (PROLENE 6-0; ETHICON) WAS USED FOR THE VASCULAR CONNECTIONS. AT THE SAME TIME IN ALL CASE, MODIFIED LICH-GREGOIR URETERONEOCYSTOSTOMY WAS PERFORMED USING 5-0 POLYGLACTIN SUTURES (VICRYL; ETHICON). REPORTED COMPLICATIONS ARE : POLYPROPYLENE SUTURES (PROLENE 6-0; ETHICON). -ACUTE TUBULAR NECROSIS (N=10). TREATMENT: NOT PROVIDED. -VASCULAR COMPLICATIONS (N=3). TREATMENT: NOT PROVIDED. -UROLOGICAL COMPLICATION (N=5). TREATMENT: NOT PROVIDED. 5-0 POLYGLACTIN SUTURES (VICRYL; ETHICON). -ACUTE TUBULAR NECROSIS (N=10). TREATMENT: NOT PROVIDED. -VASCULAR COMPLICATIONS (N=3). TREATMENT: NOT PROVIDED. -EXPERIENCING GRAFT FAILURE (N=1). TREATMENT: NOT PROVIDED. -UROLOGICAL COMPLICATION (N=5). TREATMENT: NOT PROVIDED. POLYPROPYLENE SUTURES (PROLENE 6-0; ETHICON) -MORTALITY WAS OBSERVED, ALL OF WHICH WERE ATTRIBUTED TO SEPSIS (N=3). TREATMENT: NOT PROVIDED. 5-0 POLYGLACTIN SUTURES (VICRYL; ETHICON). -MORTALITY WAS OBSERVED, ALL OF WHICH WERE ATTRIBUTED TO SEPSIS (N=3). TREATMENT: NOT PROVIDED. IN CONCLUSION, THIS STUDY DEMONSTRATES THAT RENAL TRANSPLANTATION OUTCOMES ARE SUPERIOR TO THOSE OF SINGLE RENAL ARTERY ALLOGRAFTS COMPARED TO THOSE OF MULTIPLE RENAL ARTERIES. IMMEDIATE GRAFT FUNCTION PLAYS A CRITICAL ROLE IN ENSURING GRAFT SURVIVAL AND MULTIPLE RENAL ARTERIES ARE ASSOCIATED WITH HIGHER VASCULAR COMPLICATIONS AND MORTALITY. THESE FINDINGS HIGHLIGHT THE IMPORTANCE OF SURGICAL EXPERTISE, METICULOUS PERIOPERATIVE CARE, AND TARGETED POSTOPERATIVE MANAGEMENT OF PATIENTS WITH ANATOMICALLY COMPLEX ALLOGRAFTS. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DOES THE SURGEON BELIEVE THAT ANY OF THE ETHICON PRODUCTS INVOLVED CAUSED AND/OR CONTRIBUTED TO THE POST-OPERATIVE COMPLICATIONS/DEATH DESCRIBED IN THE ARTICLE? WHICH SPECIFIC ETHICON PRODUCTS HAVE BEEN USED DURING THE PROCEDURES (PRODUCT CODE, LOT NUMBER)? DOES THE SURGEON BELIEVE THERE WAS ANY DEFICIENCY WITH ANY OF THE ETHICON PRODUCTS USED IN THIS PROCEDURE? IF SO, PLEASE PROVIDE DETAILS. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE A COMPLAINT REFERENCE NUMBER. PATIENT DEMOGRAPHICS? HAS AN AUTOPSY BEEN PERFORMED FOR DEATH CASE? IF SO, CAN THE RESULTS BE SHARED? WOULD THE JOURNAL AUTHOR / PHYSICIAN LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. H6 COMPONENT CODE: G07002 ¿ DEVICE NOT RETURNED. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. CITATION: INT J ACAD MED PHARM 2025; 7 (1); 6-10. DOI: 10.47009/JAMP.2025.7.1.2.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22331652&pc=GAW | 2210968-2025-07448 | 2025-01-06 05:00 | Death | ETHICON INC. | 2025-06-25 04:00 | GAW | PROLENE POLYPROPYLENE SUTURE UNKNOWN PRODUCT | Adverse Event Without Identified Device or Use Problem | Sepsis | K133356 |  | 1 | Event Description: TITLE: INSTITUTIONAL EXPERIENCE ON COMPARISON OF ALLOGRAFTS WITH MULTIPLE-RENAL ARTERY VS. SINGLE RENAL ARTERY IN LIVING AND DECEASED DONOR TRANSPLANTATION THE AIM OF THIS STUDY IS TO COMPARE UROLOGICAL (URINE LEAK AND STENOSIS) AND VASCULAR COMPLICATIONS (RENAL ARTERY AND VEIN THROMBOSIS) IN RENAL ALLOGRAFTS WITH A SINGLE ARTERY VS. MULTIPLE ARTERIES. ADDITIONALLY, THE SECONDARY OBJECTIVES INCLUDED ASSESSING DELAYED GRAFT FUNCTION AND GRAFT SURVIVAL AT 3- AND 6-MONTHS POST-TRANSPLANTATION. BETWEEN APRIL 2022 TO APRIL 2024, A TOTAL OF E 95 PATIENTS, 67.4% (64) WERE MALE, AND 32.6% (31) WERE FEMALE. THE MEAN AGE OF THE PARTICIPANTS WAS 34.65 ± 10.41 YEARS. FOR MRA GRAFTS A POLYPROPYLENE SUTURES (PROLENE 6-0; ETHICON) WAS USED FOR THE VASCULAR CONNECTIONS. AT THE SAME TIME IN ALL CASE, MODIFIED LICH-GREGOIR URETERONEOCYSTOSTOMY WAS PERFORMED USING 5-0 POLYGLACTIN SUTURES (VICRYL; ETHICON). REPORTED COMPLICATIONS ARE: POLYPROPYLENE SUTURES (PROLENE 6-0; ETHICON) ACUTE TUBULAR NECROSIS (N=10). TREATMENT: NOT PROVIDED VASCULAR COMPLICATIONS (N=3). TREATMENT: NOT PROVIDED UROLOGICAL COMPLICATION (N=5). TREATMENT: NOT PROVIDED 5-0 POLYGLACTIN SUTURES (VICRYL; ETHICON) . -ACUTE TUBULAR NECROSIS (N=10). TREATMENT : NOT PROVIDED -VASCULAR COMPLICATIONS (N=3). TREATMENT : NOT PROVIDED -EXPERIENCING GRAFT FAILURE (N=1). TREATMENT : NOT PROVIDED -UROLOGICAL COMPLICATION (N=5). TREATMENT : NOT PROVIDED POLYPROPYLENE SUTURES (PROLENE 6-0; ETHICON) -MORTALITY WAS OBSERVED, ALL OF WHICH WERE ATTRIBUTED TO SEPSIS (N=3) TREATMENT : NOT PROVIDED. 5-0 POLYGLACTIN SUTURES (VICRYL; ETHICON) -MORTALITY WAS OBSERVED, ALL OF WHICH WERE ATTRIBUTED TO SEPSIS (N=3) TREATMENT : NOT PROVIDED . IN CONCLUSION, THIS STUDY DEMONSTRATES THAT RENAL TRANSPLANTATION OUTCOMES ARE SUPERIOR TO THOSE OF SINGLE RENAL ARTERY ALLOGRAFTS COMPARED TO THOSE OF MULTIPLE RENAL ARTERIES. IMMEDIATE GRAFT FUNCTION PLAYS A CRITICAL ROLE IN ENSURING GRAFT SURVIVAL AND MULTIPLE RENAL ARTERIES ARE ASSOCIATED WITH HIGHER VASCULAR COMPLICATIONS AND MORTALITY. THESE FINDINGS HIGHLIGHT THE IMPORTANCE OF SURGICAL EXPERTISE, METICULOUS PERIOPERATIVE CARE, AND TARGETED POSTOPERATIVE MANAGEMENT OF PATIENTS WITH ANATOMICALLY COMPLEX ALLOGRAFTS. Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DOES THE SURGEON BELIEVE THAT ANY OF THE ETHICON PRODUCTS INVOLVED CAUSED AND/OR CONTRIBUTED TO THE POST-OPERATIVE COMPLICATIONS/DEATH DESCRIBED IN THE ARTICLE? WHICH SPECIFIC ETHICON PRODUCTS HAVE BEEN USED DURING THE PROCEDURES (PRODUCT CODE, LOT NUMBER)? DOES THE SURGEON BELIEVE THERE WAS ANY DEFICIENCY WITH ANY OF THE ETHICON PRODUCTS USED IN THIS PROCEDURE? IF SO, PLEASE PROVIDE DETAILS. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE A COMPLAINT REFERENCE NUMBER. PATIENT DEMOGRAPHICS? HAS AN AUTOPSY BEEN PERFORMED FOR DEATH CASE? IF SO, CAN THE RESULTS BE SHARED? WOULD THE JOURNAL AUTHOR / PHYSICIAN LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. H6 COMPONENT CODE: G07002 ¿ DEVICE NOT RETURNED. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. CITATION: INT J ACAD MED PHARM 2025; 7 (1); 6-10 DOI: 10.47009/JAMP.2025.7.1.2 . Manufacturer Narrative: PRODUCT COMPLAINT # (B)(4). CORRECTED INFORMATION: B1, B2, H1 - ADDITIONAL INFORMATION WAS RECEIVED AND IT WAS REPORTED THAT THIS EVENT NO LONGER MEETS REPORTABLE CRITERIA. THEREFORE, THIS MEDWATCH REPORT IS BEING VOIDED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22331675&pc=FSA | 9681684-2025-00053 | 2025-05-28 04:00 | Death | ARJOHUNTLEIGH MAGOG INC. | 2025-06-25 04:00 | FSA | MAXI MOVE | Use of Device Problem | Head Injury |  |  | 1 | Event Description: DURING A TRANSFER FROM BED TO CHAIR USING A MAXI MOVE FLOOR LIFT THE RESIDENT SLIPPED OUT OF THE SLING, HIT THEIR HEAD AND AS A RESULT, DIED AT THE SCENE. Manufacturer Narrative: PROCESS OF GATHERING AND ANALYZING INFORMATION IS ONGOING. ADDITIONAL INFORMATION WILL BE PROVIDED UPON CONCLUSION OF THE INVESTIGATION. Manufacturer Narrative: THE INVOLVED DEVICE WAS NOT AVAILABLE FOR AN EVALUATION, BECAUSE THE POLICE HAVE CONFISCATED IT AS A STANDARD PART OF THE INVESTIGATION. BASED ON THE COLLECTED INFORMATION, THE CUSTOMER REPRESENTATIVE BELIEVES THAT THE EVENT WAS NOT CAUSED BY ANY MALFUNCTION OR DEFECT WITH THE LIFT OR THE SLING. ALSO, NO MALFUNCTION WAS CLAIMED. THE CLIENT SPECULATED THAT THE CAREGIVERS COULD HAVE TILTED THE FRONT OF THE LIFT, WHICH MAY HAVE CONTRIBUTED TO THE EVENT. GIVEN THE LACK OF DIRECT OBSERVATION, THIS REMAINS SPECULATIVE AND CANNOT BE CONFIRMED. THE MAXI MOVE INSTRUCTIONS FOR USE (001.25060.EN) CONTAINS STEP-BY-STEP INSTRUCTIONS FOR PERFORMING THE TRANSFER. IFU CONTAINS ALSO WARNING: "DO NOT ATTEMPT TO MOVE THE LIFT BY PULLING OR PUSHING ON THE MAST, THE JIB, THE SPREADER BAR OR THE PATIENT AS THIS CAN CAUSE THE LIFT TO TOPPLE OVER, AND CAUSE INJURIES. ALWAYS USE THE MANEUVERING HANDLE WHEN DISPLACING THE LIFT." BASED ON VERY LIMITED INFORMATION AVAILABLE, WE WERE UNABLE TO DETERMINE THE ROOT CAUSE OF THE REPORTED PROBLEM. NO DEFICIENCIES WERE IDENTIFIED IN THE LIFT OR SLING BY THE CUSTOMER. THEY WERE USED AS A SYSTEM FOR PATIENT TRANSFER AND, THEREFORE, PLAYED A ROLE IN THE INCIDENT. THIS COMPLAINT WAS DECIDED TO BE REPORTABLE DUE TO THE PATIENT¿S FALL. Manufacturer Narrative: NO UDI INFORMATION IS AVAILABLE, THE DEVICE WAS MANUFACTURED BEFORE UDI IMPLEMENTATION. THE INVESTIGATION IS STILL ONGOING. RESULTS WILL BE UPDATED IN THE NEXT REPORT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22331688&pc=FSA | 1419652-2025-00028 | 2025-05-28 04:00 | Death | ARJOHUNTLEIGH MAGOG INC. | 2025-06-25 04:00 | FSA | MAXI MOVE | Use of Device Problem | Head Injury |  |  | 1 | Event Description: DURING A TRANSFER FROM BED TO CHAIR USING A MAXI MOVE FLOOR LIFT THE RESIDENT SLIPPED OUT OF THE SLING, HIT THEIR HEAD AND AS A RESULT, DIED AT THE SCENE.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22331925&pc=NPT | 2025587-2025-04545 | 2025-02-27 05:00 | Death | MEDTRONIC HEART VALVES DIVISION | 2025-06-25 04:00 | NPT | EVOLUT FX VALVE | Patient Device Interaction Problem | Paresis; Renal Failure | P130021 |  | 1 | Event Description: IT WAS REPORTED THAT APPROXIMATELY 1 MONTH FOLLOWING THE IMPLANT OF A TRANSCATHETER VALVE, THE PATIENT HAD DIFFICULTY MOVING, WAS HOSPITALIZED DUE TO CHRONIC RENAL FAILURE, AND WAS ADMINISTERED AN UNKNOWN TREATMENT. APPROXIMATELY 1 MONTH AFTER THE HOSPITALIZATION, THE PATIENT WAS TRANSFERRED TO A DIFFERENT FACILITY; HOWEVER, 3 MONTHS AND 12 DAYS FOLLOWING THE IMPLANT OF THE TRANSCATHETER VALVE, THE PATIENT DIED DUE TO RENAL FAILURE. Manufacturer Narrative: UPDATED DATA: D4. H4. H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN. Manufacturer Narrative: MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22332880&pc=MVK | 3008642652-2025-06172 | 2025-05-11 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-25 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED TWO INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 10:59:31 ON (B)(6) 2025. AT 22:31:48, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 22:33:11, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 30 BPM DEGRADING TO ASYSTOLE. AT 22:36:56, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 22:38:01, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. POST SHOCK RHYTHM WAS IDIOVENTRICULAR RHYTHM @ 10 BPM DEGRADING TO ASYSTOLE. THE DEVICE WAS SHUT DOWN AT 23:05:53 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR AND ELECTRODE BELT HAS BEEN COMPLETED. THE EVALUATION INCLUDED REVIEW OF DOWNLOADED SOFTWARE FLAG FILES ON THE DAY OF THE EVENT AND INCOMING FUNCTIONAL TESTING. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22332884&pc=MVK | 3008642652-2025-06174 | 2025-06-17 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-25 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Over-Sensing | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED THREE INAPPROPRIATE TREATMENTS. THE DEVICE WAS STARTED UP AT 23:21:02 ON (B)(6) 2025. AT 00:26:57 ON (B)(6) 2025, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS ASYSTOLE. AT 00:27:31, THE PATIENT RECEIVED THE FIRST INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. AT 00:27:57, THE PATIENT RECEIVED THE SECOND INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. AT 00:28:24, THE PATIENT RECEIVED THE THIRD INAPPROPRIATE TREATMENT. OVERSENSING OF LOW AMPLITUDE CARDIAC SIGNAL CONTRIBUTED TO THE FALSE DETECTION. THE PATIENT WAS IN A NON-LIFE-SUSTAINING RHYTHM PRIOR TO THE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS ASYSTOLE. POST SHOCK RHYTHM CONTINUED TO BE IN ASYSTOLE, WHICH IS A NON-LIFE-SUSTAINING RHYTHM. THE DEVICE WAS SHUT DOWN AT 03:41:03 ON (B)(6) 2025. Manufacturer Narrative: THE ELECTRODE BELT AND MONITOR HAVE BEEN RETURNED AND THE EVALUATION IS UNDERWAY. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22332885&pc=MVK | 3008642652-2025-06173 | 2025-05-04 04:00 | Death | ZOLL MANUFACTURING CORPORATION | 2025-06-25 04:00 | MVK | LIFEVEST WCD 4000 SYSTEM | Failure to Convert Rhythm | Arrhythmia | P010030 |  | 1 | Event Description: A US DISTRIBUTOR CONTACTED ZOLL TO REPORT THAT A PATIENT PASSED AWAY ON (B)(6) 2025 WHILE REPORTEDLY WEARING THE LIFEVEST. THE PATIENT RECEIVED AN APPROPRIATE TREATMENT. THE DEVICE WAS STARTED UP AT 02:27:35 ON (B)(6) 2025. AT 18:29:11, AN ARRHYTHMIA WAS DETECTED. ECG SHOWS SINUS BRADYCARDIA @ 50 BPM WITH BIGEMINY DEGRADING TO VT @ 130 BPM WITH PACEMAKER SPIKES. AT 18:30:22, THE PATIENT RECEIVED THE APPROPRIATE TREATMENT. THE RHYTHM AT THE TIME OF TREATMENT WAS VT @ 130 BPM. POST SHOCK RHYTHM WAS VF WITH PACEMAKER SPIKES, MOTION ARTIFACT, AND ELECTRODE LEAD FALL OFF. THE DEVICE WAS SHUT DOWN AT 19:15:21 ON (B)(6) 2025. Manufacturer Narrative: DEVICE EVALUATION OF THE MONITOR HAS BEEN COMPLETED. DURING THE INCOMING FUNCTIONAL TESTING, A 1HZ SIMULATED NORMAL SINUS RHYTHM SIGNAL WAS APPLIED TO THE ECG ELECTRODES, FOLLOWED BY A 5HZ SIMULATED TREATABLE ARRHYTHMIA SIGNAL WHICH VERIFIED PROPER PERFORMANCE OF THE DETECTION ALGORITHM. DURING THE TRANSITION TO THE 5HZ SIGNAL, THE DEVICE WAS CONFIRMED TO PROPERLY ENTER INTO A TREATMENT SEQUENCE WHICH INCLUDES A VERIFICATION OF THE TACTILE VIBRATION ALARM, AUDIO MESSAGING, AND SIREN ALARMS, AS WELL AS A TEST OF THE PULSE DELIVERY CIRCUITRY. THE PULSE DELIVERY CIRCUITRY TEST VERIFIED PROPER CHARGING OF THE HIGH VOLTAGE CAPACITORS AND PROPER DELIVERY OF FIVE FULL ENERGY 150J BIPHASIC PULSES. THE FUNCTIONAL TESTING CONFIRMED PROPER RESPONSE BUTTON FUNCTIONALITY, ECG ACQUISITION, DETECTION ALGORITHM PERFORMANCE, AND PULSE DELIVERY FUNCTIONALITY. THERE IS NO INDICATION OF A PRODUCT MALFUNCTION. THE ELECTRODE BELT HAS NOT BEEN RECOVERED. THE DEVICE FLAG DATA FROM THE LAST DOWNLOAD DOES NOT INDICATE ANY DEVICE MALFUNCTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22333037&pc=DWM | 1625685-2025-00060 | 2025-03-06 05:00 | Death | CAREFUSION, INC | 2025-06-25 04:00 | DWM | UNKNOWN PLEURX | Adverse Event Without Identified Device or Use Problem | Insufficient Information | NA |  | 1 | Event Description: IT WAS REPORTED THAT DURING THE PLEURX CATHETER DRAINAGE PROCESS, THE INSTRUCTIONS FOR USE WERE NOT FOLLOWED AS THE INCORRECT DRAINING METHOD AND NON-MEDICAL (BOILING KETTLE WATER MIXED WITH STERILE WATER PACKS, MEASURING JUG FROM THE KITCHEN) ACCESSORY PRODUCTS WERE BEING USED. IT WAS FURTHER REPORTED THAT THE PATIENT EXPERIENCED A LEVEL OF PAIN AND DISCOMFORT IMMEDIATELY AFTER THE PLACEMENT PROCEDURE. ON (B)(6) 2025 THE PATIENT HAD A PRE-PLANNED DRAINING PROCEDURE TO DRAIN THE FLUID, AND THE FAMILY OBSERVED THE TUBING WAS LEFT OPEN AND DRAGGING ALONG THE FLOOR IN THE WARD AND THE TOILET. IT WAS FURTHER REPORTED THAT THERE WAS A NOTICEABLE DETERIORATION IN THE PATIENT AFTER EACH HOSPITAL STAY, BECOMING WEAKER AND LESS MOBILE. AT WHICH TIME THE FAMILY WAS INFORMED THAT THE PATIENT WAS NOW CRITICALLY ILL. PRIOR TO THIS SEVERAL ATTEMPTS WERE MADE TO REMOVE THE TUBING, THIS APPEARED TO BE DONE BY PULLING IT OUT, WHICH PLACED THE PATIENT IN PAIN AND DISCOMFORT. ANOTHER UNSUCCESSFUL REMOVAL PROCEDURE WAS THEN ATTEMPTED AGAIN AND, THE FAMILY STATED THAT THE TUBING HAD INTERNAL STITCHING AND COULD NOT BE REMOVED BY SIMPLY PULLING OUT. ON (B)(6) 2025, MEDICAL STAFF DISCUSSED FITTING OF FEEDING TUBE; HOWEVER, THE PATIENT DECLINED CONSENT. THE DOCTOR INFORMED THE FAMILY THAT ALL INTERVENTIONS WOULD BE WITHDRAWN. THE PATIENT REPORTEDLY EXPIRED ON (B)(6) 2025 AND APPROXIMATELY 36 HOURS AFTER THE PATIENT EXPIRED THE DRAIN WAS SURGICALLY REMOVED FROM THE DECEASED. THE RELATIONSHIP BETWEEN THE PATIENT'S DEATH, THE DEVICE USE, AND THE PROCEDURE WAS NOT PROVIDED. Manufacturer Narrative: H10: MANUFACTURING REVIEW: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. H10: NO PHOTOS OR SAMPLES WERE RECEIVED BY OUR QUALITY TEAM FOR EVALUATION THEREFORE THE FAILURE MODE COULD NOT BE VERIFIED, AND THE ROOT CAUSE COULD NOT BE DETERMINED. WE ARE UNABLE TO CONFIRM ANY FAILURES THAT WOULD HAVE CONTRIBUTED TO PATIENT DEATH. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22333099&pc=MNS | 2518422-2025-106857 | 2025-06-24 04:00 | Death | RESPIRONICS, INC. | 2025-06-25 04:00 | MNS | DREAMSTATION ST30 | Adverse Event Without Identified Device or Use Problem | Insufficient Information | K102465 |  | 1 | Event Description: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. Manufacturer Narrative: THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE PREVIOUS REPORT HAD THE INCORRECT HEALTH IMPACT CODE. THIS HAS BEEN UPDATED TO DEATH UNDER THE HEALTH IMPACT GRID.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22333138&pc=NPT | 2015691-2025-05176 | 2025-05-20 04:00 | Death | EDWARDS LIFESCIENCES | 2025-06-25 04:00 | NPT | EDWARDS SAPIEN 3 TRANSCATHETER HEART VALVE | Central Regurgitation | Insufficient Information | P140031 |  | 1 | Event Description: AS REPORTED THROUGH EDWARDS LIFESCIENCES IMPLANT PATIENT REGISTRY (IPR), APPROXIMATELY 5 YEARS AND 8 MONTHS POST TRANSFEMORAL TAVR (TRANSCATHETER AORTIC VALVE REPLACEMENT) PROCEDURE WITH A 23MM SAPIEN 3 VALVE IN THE NATIVE AORTIC POSITION, THE PATIENT PASSED AWAY ON AN UNKNOWN DATE DUE TO AN UNKNOWN CAUSE. ACCORDING TO THE PROVIDED MEDICAL RECORDS, THE TAVR BIOPROSTHETIC VALVE WAS NOTED TO HAVE SEVERE CENTRAL/AORTIC REGURGITATION (AI), WHICH MAY HAVE CONTRIBUTED TO THIS PATIENT'S DEATH. THE PROVIDED 3-YEAR POST TAVR ECHOCARDIOGRAM ALSO INDICATED THAT THE PATIENT'S LEFT VENTRICULAR EJECTION FRACTION (LVEF) WAS LESS THAN 20% WITH SEVERELY REDUCED FUNCTION AND GLOBAL HYPOKINESIS. THE PEAK AND MEAN GRADIENTS OF THE TAVR WERE 11MMHG AND 8MMHG, RESPECTIVELY, WITH MODERATE-TO-SEVERE AI AND AN AORTIC VALVE AREA (AVA) OF 1.8 CM^2. IN ADDITION, THERE WAS AN ELEVATED RIGHT VENTRICULAR (RV) SYSTOLIC PRESSURE, LEFT ATRIAL (LA) DILATION (MODERATE-TO-SEVERE), A POSTERIOR MITRAL REGURGITANT JET, SEVERE MITRAL REGURGITATION, AND MODERATE-TO-SEVERE TRICUSPID REGURGITATION. Manufacturer Narrative: A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO ENGINEERING EVALUATION FINDINGS. SECTIONS B4, G3, G6, H2, H6: TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS HAVE BEEN UPDATED. ADDITION TO SECTION H6: TYPE OF INVESTIGATION. CORRECTIONS TO SECTIONS H6: INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS. THE EVENT(S) REPORTED IS/ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. ADDITIONAL ASSESSMENT OF THE FAILURE MODE(S) IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. IN THIS CASE, THE COMPLAINT OF CENTRAL REGURGITATION WAS CONFIRMED BASED ON PROVIDED MEDICAL REPORT. PER MEDICAL RECORD, PATIENT HAD VAST COMORBIDITIES (E.G. HYPERLIPIDEMIA, HYPERTENSION, CORONARY ARTERY DISEASE, AND ETC.) WHICH ARE KNOWN FACTORS THAT MAY INCREASE THE RISK OF VALVE DEGENERATION, LEADING TO CENTRAL REGURGITATION. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (PRE-EXISTING COMORBIDITIES) MAY HAVE CONTRIBUTED TO THE COMPLAINT EVENT; HOWEVER, A DEFINITIVE ROOT CAUSE IS UNABLE TO BE DETERMINED. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED. Manufacturer Narrative: THE INVESTIGATION IS ONGOING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22336298&pc=OYC | MW5171912 | 2021-06-11 04:00 | Death | MEDTRONIC MINIMED | 2025-06-25 04:00 | OYC | MINIMED 600 SERIES | Defective Alarm; Excess Flow or Over-Infusion; Nonstandard Device; Battery Problem; Improper Flow or Infusion | Diabetic Ketoacidosis; Confusion/ Disorientation |  |  | 1 | Event Description: PATIENT'S SISTER CALLED IN TO REPORT THE MALFUNCTION OF THE DEVICE WHICH SHE BELIEVES IS THE CAUSE OF DEATH AS THE DEVICE HAS MULTIPLE ISSUES WITH THE BATTERY, OVER INJECTING OR UNDER INJECTING INSULIN, NO ALARMS WHEN THE READINGS WERE NOT WITHIN THE NORMAL RANGE. SHE REPORTED THAT THE DEVICE WAS RECALLED TWICE BY FDA IN 2018 AND 2024. THE PATIENT'S FAMILY HAD FILED A LAWSUIT AGAINST THE DEVICE AND MANUFACTURE. SHE STATES THAT THE PATIENT HAD JUST CHANGED THE PUMP ON THE DAY HER BROTHER DIED, THE DEVICE OVER INJECTED HIM WITH INSULIN AND THE CAUSE OF HIS DEATH IS DKA. HE WAS DISORIENTED AND WAS NOT FEELING WELL SINCE 12 HOURS BEFORE HIS DEATH. THE REPORTER ALSO STATES THAT THE PERSON WHO DID THE AUTOPSY HAD AGREED THAT THE DEATH WAS DUE TO THE DEVICE MALFUNCTION. THE MANUFACTURE HAD INTERVIEWED THE FAMILY AND THEY WERE HOSTILE WITH THEM WHILE INTERVIEWING THEM. THE REPORTED DOESN'T WANT ANYONE ELSE TO GO THROUGH THIS AND WANTS THE PRODUCT TO BE RECALLED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22337635&pc=NIK | MW5171939 | 2025-06-07 04:00 | Death | MEDTRONIC, INC. | 2025-06-25 04:00 | NIK | COBALT HF CRT-D | Failure to Deliver Shock/Stimulation | Tachycardia; Loss of consciousness |  |  | 1 | Event Description: MY HUSBAND HAD THE MEDTRONIC COBALT HF CRT-D DEVICE TO TREAT TACHYCARDIA. THE DEVICE WAS SUPPOSED TO PACE OUT EPISODES AND IF NECESSARY, DELIVER SHOCK THERAPY WHEN NECESSARY. ON THE DAY OF HIS DEATH, HE COLLAPSED AND THE DEVICE FAILED TO DELIVER THERAPY. I REQUESTED A DEVICE CHECK THRU HIS CARDIOLOGIST'S OFFICE THAT PLACED THE DEVICE AND MONITORS SAME. SINCE REQUEST I HAVE BEEN LOCKED OUT THE MYCHART APP THAT HOUSES HIS DEVICE AND MEDICAL RECORDS. THE DOCTORS OFFICE CALLED ME AND SAID ON THE DAY OF HIS DEATH (B)(6) 2025 HE HAD ONE UNSUSTAINED EPISODE OF TACHYCARDIA AT 3:11PM AND NO THERAPIES WERE DELIVERED. IN (B)(6) 2025 HE HAD ONE UNTREATED EPISODE OF TACHYCARDIA. I ASKED HOW THAT WAS POSSIBLE AS HE DIED ON (B)(6) 2025. NO REASONABLE EXPLANATION WAS PROVIDED. I REQUESTED ACCESS TO THE REPORT. THEY REFUSED TO EMAIL THE REPORT BUT SAID THEY WOULD PRINT IT OUT AND MAIL IT TO ME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22337637&pc=MOZ | MW5171940 | 2025-06-22 04:00 | Death | GENZYME CORPORATION | 2025-06-25 04:00 | MOZ | SYNVISC ONE PFS 8MG/ML | Adverse Event Without Identified Device or Use Problem | Insufficient Information |  |  | 1 | Event Description: PT DECEASED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=22338159&pc=OHT | MW5171942 | 2025-05-30 04:00 | Death | SHENZHEN IONKA MEDICAL TECHNOLOGY CO., LTD. | 2025-06-25 04:00 | OHT | HOME LASER HAIR REMOVAL DEVICE | Nonstandard Device; Off-Label Use; Device Markings/Labelling Problem | No Clinical Signs, Symptoms or Conditions |  |  | 1 | Event Description: FORMAL COMPLAINT AGAINST "UBROO" HOME LASER HAIR REMOVAL DEVICE: I AM A DEDICATED SKINCARE ENTHUSIAST WITH A FOCUS ON CONSUMER RIGHTS, WRITING TO FORMALLY REPORT SERIOUS REGULATORY VIOLATIONS BY UBROO REGARDING ITS "UBROO HOME HAIR REMOVER MACHINE /HAND-HLED HAIR REMOVAL DEVICE", SOLD ON AMAZON PLATFORM (HTTPS://WWW.AMAZON.COM/DP/B0D172GRNQ/REF=TWISTER_B0DRJJTQDC?_ENCODING=UTF8&TH=1) AND ITS DTC WEBSITE (HTTPS://UBROOSKIN.COM/PRODUCTS/LASER-HAIR-REMOVAL-WITH-5-ICE-COOLING-PAINLESS-IPL-LASER-HAIR-REMOVAL-DEVICE-3-IN-1-UBROO-AT-HOME-HAIR-REMOVER-MACHINE-FOR-WOMEN-MEN-ELECTROLYSIS-PERMANENT-EPILATOR-ON-FACE-BODY-DEPILATION-GOLD). MY COMPLAINTS INCLUDE "UNLAWFUL MARKETING OF AN UNREGISTERED CLASS II MEDICAL DEVICE, FALSE ADVERTISING, AND DANGEROUS MISBRANDING", DETAILED BELOW WITH LEGAL CITATIONS AND EVIDENCE. UNREGISTERED CLASS II MEDICAL DEVICE (FDA 510(K) VIOLATION) THE "UBROO" DEVICE IS MARKETED AS A "IPL HAIR REMOVAL DEVICE", WHICH IS CLASSIFIED AS A CLASS II MEDICAL DEVICE BY FDA . DESPITE THIS, NO 510(K) PREMARKET NOTIFICATION IS LISTED IN FDA'S DATABASE FOR THIS PRODUCT AND NO FDA ESTABLISHMENT REGISTRATION OF MANUFACTURER. MISBRANDING & INADEQUATE LABELING THE DEVICE'S LABELING VIOLATES FDA'S REQUIREMENTS FOR "TRANSPARENCY AND SAFETY": MISSING MANUFACTURER INFORMATION: THE DEVICE, PACKAGING, AND USER MANUAL LACK THE "NAME/ADDRESS OF THE MANUFACTURER/DISTRIBUTOR", VIOLATING 21 CFR 801.1(A). - INCOMPLETE INSTRUCTIONS FOR USE: - NO CONTRAINDICATIONS (E.G., FOR DARKER SKIN TONES OR EPILEPSY) ARE LISTED, VIOLATING 21 CFR 801.109(C). - NO EMC (ELECTROMAGNETIC COMPATIBILITY) COMPLIANCE INFORMATION IS PROVIDED, A CRITICAL SAFETY STANDARD FOR MEDICAL DEVICES (*IEC 60601-1-2*, ADOPTED BY FDA). CONTRADICTORY TECHNOLOGY CLAIMS & RADIATION SAFETY VIOLATIONS THE PRODUCT PACKAGING AND MANUAL CONTAIN CONFLICTING TECHNOLOGY DESCRIPTIONS, CREATING MATERIAL MISREPRESENTATION: IPL VS. LASER MISBRANDING: THE PACKAGING STATES "IPL TECHNOLOGY," BUT THE MANUAL CLAIMS "LASER HAIR REMOVAL." THESE ARE DISTINCT TECHNOLOGIES: - IPL (INTENSE PULSED LIGHT): A NON-LASER DEVICE, TYPICALLY REGULATED AS A CLASS II DEVICE. - LASER: A CLASS II MEDICAL DEVICE REQUIRING 510(K) CLEARANCE AND RADIATION SAFETY COMPLIANCE (21 CFR 1040.10). - IF THE DEVICE IS IPL BUT MARKETED AS "LASER," THIS VIOLATES FTC ACT § 5 (DECEPTIVE ADVERTISING). IF IT IS TRULY A LASER, THE LACK OF 510(K) REGISTRATION VIOLATES FD&C ACT § 510(K). - UNSUBSTANTIATED LASER CLAIMS: - THE MANUAL'S "LASER" LABELING LACKS FDA-MANDATED RADIATION SAFETY DISCLOSURES, INCLUDING: - LASER CLASS DESIGNATION (E.G., CLASS IIIB) PER 21 CFR 1040.10(G). - WAVELENGTH, POWER OUTPUT, AND WARNING STATEMENTS (E.G., "DANGER: EYE HAZARD") PER 21 CFR 1040.10(E). - MISSING RADIATION SAFETY REPORT: UNLAWFUL MEDICAL DEVICE MISBRANDING & UNAPPROVED THERAPEUTIC CLAIMS UBROO'S AMAZON AND DTC LISTING AND PRODUCT MATERIALS CONTAIN UNSUBSTANTIATED MEDICAL CLAIMS THAT VIOLATE FDA'S PREMARKET APPROVAL REQUIREMENTS AND LABELING REGULATIONS. - UNAUTHORIZED MEDICAL DEVICE MARKETING: THE DEVICE IS MARKETED WITH THERAPEUTIC CLAIMS EXCEEDING THE SCOPE OF A GENERAL CONSUMER PRODUCT: - PAINLESS TREATMENT: THE CLAIM "PAINLESS" CONSTITUTES A MEDICAL EFFICACY STATEMENT REQUIRING FDA CLASSIFICATION AS A CLASS II DEVICE. - TEMPERATURE CONTROL: THE "KEEPS SKIN BELOW 5°C" CLAIM RELATES TO THERMAL SAFETY, A CRITICAL PARAMETER FOR MEDICAL DEVICES. - SMOOTH SKIN CLAIMS: STATEMENTS LIKE "75% HAIR REDUCTION AFTER 4 WEEKS", "SMOOTH SKIN AFTER 8 WEEKS", AND "92.4% OF WOMEN REPORTED 96% HAIR REDUCTION AFTER 2 MONTHS" CONSTITUTE MEDICAL DEVICE CLAIMS THAT MAY ONLY BE MADE AFTER OBTAINING 510(K) CLEARANCE AS A CLASS II DEVICE UNDER 21 CFR 807.20(A) AND 21 CFR 801.1. 5. REQUEST FOR ACTION - ENFORCE 510(K) REGISTRATION: REMOVE "UBROO" FROM THE MARKET UNTIL 510(K) CLEARANCE AND ESTABLISHMENT REGISTRATION ARE CONFIRMED. INFORMATION ABOUT MANUFACTURERS AND SELLER SINCE THE MANUFACTURER INFORMATION FOR THIS PRODUCT IS COMPLETELY UNTRACEABLE, THE ONLY CONTACT DETAILS THAT CAN BE FOUND ARE IN THE INSTRUCTION MANUAL: PHONE: (B)(6). - PRODUCTS S/N: (B)(6). - THE SELLER MARKED AS 'HOUNYTRY' ON THE AMAZON WEBSITE, WHICH IS SHENZHEN ZEHONGYUAN TECHNOLOGY CO., LTD. REGISTERED ADDRESS IS: BUILDING 69, AOTOU NEW VILLAGE, SHANGFEN COMMUNITY, MINZHI STREET, LONGHUA DISTRICT, SHENZHEN, CHINA.